## Slide 1
Principles of Quantitation:
Chromatography
Lane C. Sander, Ph.D.
Chemical Sciences Division
Material Measurement Laboratory
National Institute of Standards and Technology1
1Contribution of the National Institute of Standards and Technology. Not subject to copyright. Certain commercial equipment, instruments, or materials are identified to specify adequately the experimental procedure. Such identification does not imply recommendation or endorsement by the National Institute of Standards and Technology, nor does it imply that the materials or equipment identified are the best available for the purpose.
Notes: The accurate characterization of the chemical composition of a sample depends on a variety of factors that must be considered by the analyst. Knowledge of how different approaches to quantitation affect the measurement accuracy is requisite to the development of methods of analysis. This presentation will describe some of the principles of quantitative analysis, particularly as applicable to chromatographic methods.

## Slide 2: Accuracy
Accuracy has been defined as:
closeness of agreement between a measured quantity value and a true quantity value of a measurand
Accuracy can be thought of as the combination of “trueness” and “precision”
Good trueness
Good precision
Good trueness
Poor precision
Poor trueness
Good precision
Poor trueness
Poor precision
Notes: Accuracy has been defined by the international community as the closeness of agreement between a measured quantity value and a true quantity value of a measurand.
Accuracy can be thought of as a combination of two components of trueness and precision. The difference between trueness and precision can be illustrated by target plots.
So, in the first case, this is an illustration of good trueness and good precision;
at the right the example shows poor trueness and good precision;
good trueness and poor precision (bottom left);
and poor trueness and poor precision, bottom right.

## Slide 3: Approaches to Quantitation
External Standard Calibration Approach
Internal Standard Calibration Approach
Standard Additions Approach
Notes: Three approaches are commonly used in quantitative analysis: External Standard Calibration Approach, Internal Standard Calibration Approach, and the Standard Additions. For these methods, instrumental response is related to the amount of substance through a calibration procedure using reference standards, for each analyte to be determined.

## Slide 4: Approaches to Quantitation
External Standard Calibration Approach
  Calibration is based on direct comparison with instrumental response (via calibration curves or response factors)
  Advantage: simple
  Disadvantages:
    Skill required to minimize biases from sample processing;
    Not well suited to complex matrix samples
    Need to know exact extract volumes (or masses)
    Quantitative transfers are required for all processing steps
Notes: The External Standard Calibration approach is well suited to the analysis of solutions or samples that do not require extensive processing steps. It is less well suited to samples that need to be extracted, since all steps must recover 100% of the analyte.
The main advantage to this approach is that it is simple, and there are fewer things to go wrong.
The external standard approach has several disadvantages:
skill is required to minimize biases from sample processing
The approach is not well suited to complex matrix samples
Knowledge of the volumes or masses associated with all transfers or processing steps is required
And, quantitative transfers are required for all processing steps.

## Slide 5: Approaches to Quantitation
External Standard Calibration Approach
Calibrants
Unknown
Calibration Curve
Notes: The basis for external standard calibration is conceptually very simple. Instrumental response can be related to the quantity of analyte present. Usually this response is linearly related to the mass of the analyte. If one or more samples are prepared with known composition, the response of the instrument can be calibrated. The figure on the left shows chromatograms for four calibrants with different concentrations. Chromatographic response is characterized by the area of the peak corresponding to the analyte. This area can be plotted as a function of the amount of substance present in the calibrant, usually in terms of concentration or mass fraction units, as is shown in the figure at the right. The area response for the analyte in an unknown sample can then be related to concentration through this calibration curve.

## Slide 6: Approaches to Quantitation
Internal Standard Calibration Approach
  One or more compounds are added to both calibrants and unknowns as the internal standard(s)
  Calibration is based on the ratio of responses for analytes and internal standards
  Advantage: losses from transfers, dilutions, etc. are compensated; may compensate changes in instrumental response; less skill is required
  Disadvantage: calculations more complex; internal standards must be identified and used
  Internal standards are added at the earliest opportunity
  Knowledge of volumes is not required
  Quantitative transfers are not required
Notes: 1 The internal standard calibration approach is similar to the external standard approach, except that analyte response is determined relative to the response of a reference compound that is added to the calibrants and samples in known quantities.
2 With this approach, calibration is based on the ratio of responses for analytes relative to the internal standards, rather than the absolute response of the analytes.
3 The primary advantage in this approach is that losses from transfers, dilutions, and other manipulations are to some degree compensated. It may help to compensate for changes in instrumental response, and less skill is required by the analyst.
4 Disadvantages include: it is a more complex approach in terms of calculations; and internal standards must be identified and used – the internal standards must be added to both the calibrants and the unknown samples
5 Internal standards are added at the earliest opportunity in the sample workup to provide the best overall compensation for any of the manipulations
6 And knowledge of the volumes associated with calibrants and samples is not required;
7 Quantitative transfers are not required.

## Slide 7: Approaches to Quantitation
Internal Standard Calibration Approach
Calibrants
Unknown
Calibration Curve
Notes: One way of thinking of the internal standard is as a volume corrector. For example, with use of an internal standard, the response ratio remains constant even if some of the sample is not recovered from a processing step.
Typically, a single internal standard solution is prepared for addition to calibrants and unknown samples. It is also common to add the same quantity of the internal standard to the calibrants and unknowns, to achieve a constant response for this component. The calibration curve is prepared by plotting the ratio of the response for the analyte and associated internal standard as a function of the ratio of the concentrations of the analyte and internal standard. If a single internal standard solution is used, the actual concentration of the internal standard does not need to be known since the values cancel mathematically in calculation of the unknown concentration.

## Slide 8: Approaches to Quantitation
Standard Additions Approach
  A known quantity of the analyte is added to the unknown sample; the response is compared with the unspiked sample
  Advantages: quantitation is based on the sample matrix; similar advantages to using calibrants prepared in a “blank” sample matrix
  Disadvantages: calculations are complex, sample processing is labor intensive; separate calibrations needed for each sample; “N” for calibration is small compared with other approaches; linear response is a requirement.
Notes: For the standard additions approach: a known quantity of the analyte is added to the unknown to provide at least two levels. The responses are plotted against the quantities added, and an extrapolation is made to zero response. This negative quantity can be thought of as the amount of analyte that would have to be removed from the sample to achieve zero response, and it represents the amount of analyte present in the unknown. A linear instrumental response is requisite for this approach to work.
Advantages: include that quantitation is based on the sample matrix; Similar advantages exist to using calibrants prepared in a blank sample matrix
Disadvantages: calculations are complex, sample processing is labor intensive; separate calibrations needed for each sample; and because of this, “N” for calibration is small compared with other approaches; and linear response is a requirement for this method.

## Slide 9: Approaches to Quantitation
Standard Additions Approach
Samples
Notes: This figure illustrates the standard additions approach to quantitation. The standard additions method combines calibration and sample analysis into a single process. At least two sample analyses are required for quantitation. Typically, the sample is analyzed without addition of the analyte, and one or more additional samples are analyzed which have been spiked with known quantities of the analyte. The resulting calibration curve can be extrapolated to intersect the x-axis at a point that corresponds to the amount of analyte that would need to be removed to achieve a zero response. This quantity represents the level of the analyte in the sample.

## Slide 10: Internal Standards: Choice & Number
Isotopically labeled compounds are preferable
  Perdeuterated compounds: resolved in RPLC separations (e.g., PAHs)
  Typically for MS, m/z difference should be at least 3
  Separation of analyte and IS is not desirable for MS
  2H vs 13C labels:
    Deuterium / hydrogen exchange may occur with 2H labeled reagents
    13C retention is similar to unlabeled reagents; separation of perdeuterated and unlabeled reagents may be possible
Notes: A few recommendations can be made concerning the development of analytical methods that utilize internal standards.
For mass spectrometry, by far the best approach is to use an isotopically labeled version of the analyte. Carbon 13 and deuterium labels are most commonly used for this purpose, but other isotopes may also be suitable if they are available. It should be noted that isotopically labeled compounds often have small differences in retention compared with the corresponding unlabeled compounds. Perdeuterated compounds can be fully resolved from the unlabeled compounds in both liquid chromatography and gas chromatography, and this is useful for methods that use nonselective detection, such as absorbance detection in LC, or flame ionization detection in GC.
for mass spectrometry, the rule of thumb is that the mass to charge difference should be at least 3 units,
…but not so great that chromatographic separation of the labeled and unlabeled compounds occurs. Short term fluctuations in source conditions may alter ionization efficiency, and the internal standard should co-elute with the analyte to achieve the best compensation for these fluctuations.
When considering whether to use 13C or deuterium labeled compounds, several factors should be considered.
The stability of deuterium labels should be evaluated since deuterium/hydrogen exchange can occur under certain conditions. 13C labels are not subject to exchange with unlabeled carbon.
It may be possible to separate deuterated (and especially, perdeuterated) compounds from the corresponding unlabeled compounds by reversed phase liquid chromatography and gas chromatography. This is usually not desirable with detection with mass spectrometry, but it may prove useful for absorbance detection and flame ionization detection, where separation of the internal standard is essential. However, for these applications, there are few advantages for selecting isotopically labeled compounds over other unlabeled compounds, such as a methyl-substituted isomers.

## Slide 11: Internal Standards: Choice & Number
If labeled reagent is not available, try to select a compound as similar as possible to the analyte (i.e., methyl analog or isomer)
Internal standards should be resolved from analytes and matrix components, either by chromatography and/or by selective detection.
For determination of multiple analytes, use several internal standards to cover ranges in volatility and/or isomer sets
Notes: If labeled reagent is not available, try to select a compound as similar as possible to the analyte (such as a methyl analog or an isomer) that is not present in the sample.
Internal standards need to be resolved from analytes and matrix components, either chromatographically or through selective detection, as is the case with mass spectrometry.
For determination of multiple analytes, it is good practice to use several internal standards to cover ranges in volatility and/or different isomer sets, to achieve the best results. This is commonly done with complex environmental samples, which span a range of properties for the analytes of interest.

## Slide 12: Internal Standards
For the internal standard approach, both the ratios and the absolute levels of analytes and internal standards should be similar for calibrants and unknowns
Internal standard level should be greater than contributions from interferences (noise, matrix constituents)
High levels of internal standard may be tolerable... (?)
Internal standards have different consequences for short term fluctuations (“noise”) and longer term fluctuations (“drift”)
Isotopically labeled internal standards that coelute with the analyte improve MS precision by compensating for short term noise associated with ionization (this is the basis of IDMS)
Internal standards that are separated from the analyte do not compensate as well for short term fluctuations, but may compensate for drift
Notes: For the internal standard approach to quantitation, a few recommendations can be made that can help to increase overall method accuracy.
Both the ratios and the absolute levels of analytes and internal standards should be similar for calibrants and unknowns
The Internal standard level should be greater than contributions from any interferences (such as noise, or matrix constituents)
And for this reason, it may be better to use high levels of an internal standard than low levels that are similar to these interferences.
Internal standards have different consequences for short term fluctuations (“noise”), and longer term fluctuations (such as “drift”)
For mass spectrometry, isotopically labeled internal standards coelute with the analyte; this provides the best compensation for noise and drift, and it is the basis for isotope dilution mass spectrometry
For liquid chromatography with detection that is not as selective as mass spectrometry, internal standards must be separated from the analyte, and they do not compensate as well for short term fluctuations, but they may still compensate for drift

## Slide 13: Internal Standards
Unknown
Analyte
IS
Calibrant
“case 1”
Analyte
IS
Calibrant
“case 2”
Analyte
IS
Calibrant
“case 3”
Analyte
IS
Good: level of internal standard matches the level of the analyte in the unknown
Good: level of internal standard matches the level of the analyte, and both levels are comparable between calibrant and unknown
Less desirable: levels are not comparable between calibrant and unknown even though the ratios of analyte/IS are similar
Less desirable: levels of the IS are not comparable between calibrant and unknown
Notes: The following figures show some of the desirable, and less than desirable, examples for use of an internal standard. In principle, levels of the IS added should not affect the precision and accuracy of the IS approach.
In practice, better precision and accuracy often results when the levels of the IS are added are similar to levels of the analyte in the unknown. In the first example which is desirable, The internal standard matches the level of the analyte in the unknown. Contributions from noise and matrix interferences should be small compared with the response for the internal standard. The need for linear instrumental response is reduced when the levels of internal standards and analytes are matched between calibrants and unknowns.
For the calibrant, it is desirable for the ratio of the internal standard and the analyte to be similar to this ratio in unknown samples, and it is also desirable for the internal standard and analytes to have similar absolute levels.
The next example is less desirable. Here, the levels of the internal standard and analyte are smaller than levels in the unknown. It can be noted that the ratios of the internal standard and analyte are similar between the two cases.
The last case, which is also less than desirable shows a calibrant where the internal standard level is low compared with the internal standard in the unknown; however, the absolute response of the analyte is similar to that in the unknown.

## Slide 14: Calibration
cal
cal
cal
cal
unknown
calibration
clustered near
unknown
cal
cal
cal
cal
unknown
Case 1
(unknown levels)
Case 2
(bracketed)
Case 3
(extrapolation)
Case 1 is the best approach if the target level of the unknown can be only broadly estimated
Case 2 is the best approach if the approximate level of the unknown is known
Case 3 is less desirable; valid results are obtained by extrapolation if the instrumental response is linear
response
amount
response
response
amount
amount
Notes: In planning the preparation of calibration solutions, the concentration range to be targeted should be guided by knowledge or assumptions about the concentration of the unknown samples.
Ideally the levels in the calibrants should be similar to the levels in the unknowns. This minimizes any issues of linearity of instrumental response.
When the target level of the analyte can be estimated to fall within a reasonable concentration range, calibrants can be created to bracket this range.
If more specific information is known about the target level of the unknown, then the calibrants can be clustered to more closely bracket this level.
Finally, if the level of the unknown cannot be estimated, it may happened that the calibrants do not bracket the unknown. Valid results can still be obtained by extrapolation if the instrument response is linear.

## Slide 15: Calibration Models
averaged
response factors
zero intercept
linear regression
calculated intercept
linear regression
y
mx
=
slope of regression line
- unweighted: dilute calibrants
are less important than
concentrated calibrants
- linear response assumed
y
mx
b
=
+
slope of regression line
“y” intercept
- appropriate if a bias is
suspected, e.g., calibrants
prepared by spiking a blank
matrix (such as blank urine)
which might exhibit a nonzero
blank response
Ave RF = (RF1 + RF2 + RF3 + RF4)/4
- all calibrants are
weighted equally
- linear response assumed
- zero intercept assumed
Cal 1
Cal 2
Cal
3
Cal
4
RF2
RF3
RF4
RF1
- zero intercept is forced
response
amount
response
response
amount
amount
Notes: Several mathematical models are commonly used in calibration. The selection of an appropriate model should be based on knowledge of the measurement system – such as the anticipated analyte levels, matrix interferences, noise, and drift.
One approach is to use averaged response factors. In this figure, 4 calibration solutions are shown. A response factor can be calculated for each calibrant by dividing the analyte response by the level of that analyte. These response factors can be averaged to give an averaged response factor. This approach weights all calibrants equally. It assumes a linear instrumental response, and a zero intercept.
Another common approach is to use a linear regression model with a zero intercept; this model is described by the function y=mx. For the four calibrants levels shown in the second example, a line can be calculated that passes through zero. In the unweighted version of this model, dilute calibrants are less important than concentrated calibrants, so if one of the calibrants is very concentrated compared with the other calibrants, the concentrated calibrant will dominate the calculation of the linear regression more than the dilute calibrants. As with averaged response factors, linear response is assumed, and a zero intercept is assumed.
In the third example, a linear regression model with a calculated intercept, y = mx + b, is fit to the data. This approach is appropriate if a bias is expected for the calibrants. A good example might be represented by calibrants that are prepared in a complex matrix material that exhibits a nonzero blank response.

## Slide 16: Outliers
High level calibrants are weighted more than lower level calibrants for linear regression
Large effect for extrapolation to higher levels
For the same data [expanded scale], a low level unknown is only slightly affected
Zero intercept strongly influences low level analytes.
response
amount
response
amount
Notes: Outliers are data observations that are distant from other observations or that appear to differ from a modeled relationship. A detailed discussion of the identification and treatment of outliers is beyond the scope of the current presentation, but in certain instances, the exclusion of such data may be warranted. The effect that inclusion or exclusion of a potential outlier has on calibration depends in part on the type of calibration model used.
In the first example, a potential outlier is shown for data modeled by unweighted linear regression model of the type y=mx.
*If the outlier is excluded, the calibration line is modified in a way that significant effects extrapolated data at high levels.
In the second example, the same data and model is used to determine a low level analyte.
*here, if the outlier is excluded, the modified regression line only slightly changes the extrapolated data at low levels.

## Slide 17: Miscellaneous Considerations
Sample handling
Add the internal standard at the earliest opportunity, i.e., before extraction
Consider employing mass fraction based quantitation (use fluid masses rather than volumes)
Devise weighing schemes so “weight by difference” does not involve the difference of two large numbers
  20 mg – 19 mg is better than 100.020 g – 100.019 g
Notes: A few other recommendations can be made to help achieve accurate results.
Add the internal standard at the earliest opportunity in the measurement scheme. For example, if there is an extraction step, it is best to add the internal standard before the extraction. This helps to provide the best compensation for any manipulations in processing the sample.
Consider employing mass fraction based quantitation as apposed to volumetric based quantitation. It is much easier to accurately determine masses using modern electronic balances, than it is to use volumetric glassware. Typically, measurements with volumetric glassware are accurate to only about 3 significant figures, whereas masses can be routinely determined to 5 or 6 significant figures.
Devise weighing schemes so that “weight by difference” does not involve the difference of two large numbers. For example, for the measurement of a milligram quantity, it is better determine the mass as the difference of 20 mg minus 19 mg, rather than 100.020 g minus 100.019 g.

## Slide 18: Miscellaneous Considerations
Calibration solutions
Use at least two independently weighed calibrants
Avoid the use of a single “stock solution” – what if an error is made in the original solution?
Dilutions are OK if analyte masses are too small to be accurately weighted
Best accuracy results if calibrants closely match unknowns
Notes: Use at least two independently weighed calibrants
Avoid the use of a single “stock solution”. The question can be asked “what if an error is made in the original solution?’ Dilutions made from this stock solution might exhibit good linearity, but a bias would exist if the original stock solution were biased.
Dilutions are OK if analyte masses are too small to be accurately weighted in a single step. Unnecessary steps should be avoided; however if a fractional milligram quantity is required to directly prepare a dilute solution, it is probably better to prepare the solution by diluting a more concentrated solution, using gravimetric based measurements.
Best accuracy typically results if calibrants closely match the levels of the unknowns

## Slide 19: Miscellaneous Considerations
Sample Introduction (Liquid Chromatography)
Injection is a volumetric process; quantitation with mass fraction units assumes that the density of calibrants and unknowns are the same (external standard)
If densities differ, a bias will be introduced through the injection of different masses
Injection volumes should be as small as practical
  10 μL for 4.6 mm i.d. columns
  3 μL for 2 mm i.d., columns
If possible, the sample and calibrant solvent should match the mobile phase composition
Notes: A few special considerations apply to liquid chromatographic methods.
Injection is a volumetric process; and quantitation with mass fraction units assumes that the density of calibrants and unknowns are the same. This is true for the external approach to quantitation. This is not a problem with the internal standard approach to quantitation.
However, for external standards, if densities do differ, a bias will be introduced through the injection of different masses of the samples and unknowns.
In general, injection volumes should be kept as small as practical.
  for standard 4.6 mm i.d. columns, a good rule of thumb is to use a 10 μL injection volume
  for smaller inside diameter columns, for example 2 mm i.d., 3 μL is suitable.
If possible, the sample and calibrant solvent should match the mobile phase composition. If this is the case, or if the injection solvent is even weaker than the mobile phase, then the peak shape for the separation will be maintained.

## Slide 20: Sample Queue Design
When practical, the number of vials with calibrants measured should be comparable to the number of vials measured with unknown samples
Intersperse calibrants and unknowns in the sample queue
Order samples and calibrants using a random selection scheme
To permit assessment of within sample and between sample effects, process subsamples from a single bottle for comparison with samples from multiple bottles
Plot measurements
  Levels vs run order
  Levels vs sample processing order (or bottle fill order)
Notes: In planning the certification of a reference material, consideration should be given to the order of analysis of calibrants and samples. For example, it is easy to imagine how measurement bias might be introduced through the experimental design in the presence of system drift. The following guidelines can help to reduce biases and enhance the assessment of the sources of measurement uncertainty.
When practical, the number of vials with calibrants should be comparable to the number of vials with unknown samples. Obviously, this is only a guideline, but it is a good starting point.
Intersperse calibrants and unknowns throughout the sample queue. This is a better approach than to group the calibrants together and the samples together, particularly in the presence of continuing drift in instrumental response.
Order samples and calibrants using a random selection scheme
To permit assessment of within sample and between sample effects, process subsamples from a single bottle for comparison with samples from multiple bottles
It is often quite useful to plot measurements as a diagnostic tool. Trends in data can become apparent by plotting levels as a function of run order, or as a function of processing or bottle fill order.

## Slide 21: Drift and Inhomogeneity
Randomize samples and calibrants in processing stream
Plot data as a function of run order and processing order
7.0
8.0
9.0
10.0
11.0
12.0
0
200
400
600
800
1000
7.0
8.0
9.0
10.0
11.0
12.0
0
5
10
15
20
25
No Instrumental Drift
No Bottle to Bottle Inhomogeneity
ave = 10
s = 0.6
Run Order
Fill Order /
Processing Order
Run number
Fill or Processing Order
7.0
8.0
9.0
10.0
11.0
12.0
0
200
400
600
800
1000
7.0
8.0
9.0
10.0
11.0
12.0
0
5
10
15
20
25
Instrumental Drift
No Bottle to Bottle Inhomogeneity
ave = 10
s = 0.6
Run number
Fill or Processing Order
7.0
8.0
9.0
10.0
11.0
12.0
0
5
10
15
20
25
7.0
8.0
9.0
10.0
11.0
12.0
0
200
400
600
800
1000
No Instrumental Drift
ave = 10
s = 0.6
Bottle to Bottle Inhomogeneity or
Sample processing drift
Run number
Fill or Processing Order
Notes: The following figures illustrate how the measurements of a reference material might change as a function of instrumental drift, sample inhomogeneity, or sample processing. The modeled data were generated to represent samples and calibrants that were ordered randomly in the processing stream, and data have been plotted as a function of run order and processing order.
The first figure illustrates a case in which there is no experimental drift and no bottle to bottle inhomogeneity. No trends are apparent in either of the two plots.
The second figure illustrates a trend in instrumental drift as a function of run order, but when this data is plotted as a function of fill or processing order, the drift is not as apparent.
In the third figure, a case is illustrated with data acquired in the absence of instrumental drift. A trend is shown that results during production (or filling) of the bottles. A similar trend could also result from changes that occurred during processing of the samples prior to analysis.

## Slide 22: Integration
Chromatographic resolution and/or selective detection is always the best solution!
Always inspect how baselines are set by the data system
If appropriate, adjust integration settings and/or manually reintegrate to achieve best estimate of peak area
Use peak areas rather than peak heights: peak area does not change with changes in retention
Notes: A few words can be said about the integration of chromatograms.
Although the accuracy of measurements can be enhanced by appropriate integration of chromatographic peaks, the resolution of the analytes either by chromatography or by selective detection is always better than any post run data manipulation.
Always inspect how baselines are set by the data system. The analyst can often recognize problems with integration that data system algorithms do not handle properly. Common issues include noise spikes, negative peaks, poorly resolved interferences, and instrumental drift.
If appropriate, adjust integration settings or manually reintegrate to achieve best estimate of peak areas. It should be noted that restrictions on manipulation of integration may be imposed by quality system provisions or regulatory requirements.
Use peak areas rather than peak heights: peak area does not change with changes in retention

## Slide 23: Integration and Baselines
How should these peaks be integrated?
area ratio
4:1
Notes: Unfortunately, it is not always possible separate the constituents in mixtures with full baseline resolution. This figure illustrates a synthetically generated chromatogram with components of known response ratios. The highlighted peaks were generated with an area ratio of 4 to 1, with a resolution of 1.0. The question can be asked: What is the best approach for integrating these peaks to provide the most accurate results?

## Slide 24
Notes: Different approaches to integration can be demonstrated using a chromatographic data system. When reintegrating data, it is useful to zoom in on peaks to be considered for more accurate results.
The first approach illustrated is valley baseline, where the integration lines are drawn from the minimum in the valley between adjacent peaks, to the baseline of the chromatogram. Once the integration lines are drawn, adjustments can be made to the start and stop points for each peak.
Another possibility is to consider the smaller peak as a rider on the first peak. To do this, you need to set the integration to encompass both peaks, and then split the selection into two peaks. The second peak can then be identified as a rider. When integrated in this way, the area from rider peak is subtracted from the area for both peaks to give the area of the first peak.
The integration for the rider can be converted to an exponentially skimmed rider. After doing so, the integration line well approximates the peak shape of the major component, and perhaps better describes the area of the rider as well.
Another possibility is to drop a vertical from the valley point to the baseline

## Slide 25: Integration and Baselines
-13.5%
-49.0%
Valley - Baseline
0.9 %
-3.7 %
Vertical - Baseline
11.8%
-47.3%
Skimmed Rider
2.4%
-9.7%
Exponentially Skimmed
Rider
Peak Deconvolution
-0.6 %
-2.1 %
Known area ratio: 4:1
Resolution: 1.0
Notes: The following results were obtained for different integration methods for the mathematically generated chromatogram. The biases for these examples are dependent on both the resolution of the peaks, here set at 1.0, and the relative levels of the hypothetical components, which is set at 4 to 1 for each example.
In the first case, “valley – baseline”, shown by the red integration lines, there is an under representation of the areas of both of the components. The larger component is under represented by about 13%, and the smaller component is under represented by nearly 50%.
It is also possible to integrate the two peaks by considering the smaller peak as a “rider” on the tail portion of the first peak. With this approach, the area of the large peak is determined as the difference of the combined areas of the peaks minus the skimmed area of the smaller peak. In the example, the area of the larger peak is over estimated by about 12%, and the area of the skimmed rider is underestimated by about 47%.
The next example illustrates an exponentially skimmed rider. A Gaussian model is used to estimate the peak shape of the first component, and the area of the second component is determined as the difference of the combined areas for the two peaks minus the area of the larger peak. This approach provides a better estimate of the larger component based on the assumed Gaussian peak shape, and improved estimate of the area of the smaller component, with a negative bias of about 10%.
Peak deconvolution is another approach to estimating the areas of unresolved peaks. A mathematical model of the assumed peak shape is fit to the peaks, to provide an estimate of each peak in the absence of the other. Most commonly, a Gaussian model is used to describe the chromatographic peak shape, however, in principle, this approach is not limited to any specific model. The accuracy of peak deconvolution depends on the assumptions of the model, and how well the peak shape is described by the model. In the current example, the areas of both peaks are only slightly underestimated.
The last example illustrated on this slide is based on area estimates defined by a vertical line drawn from the valley between the two peaks. This simple approach often produces relatively unbiased results, particularly if the peak responses are comparable. The reason that this works as well as it does is that the area of the first peak lost by truncating the tail portion of the peak is to some degree compensated by the area under the tail of the second peak, which is added to the first peak. A similar rational can be made for the second peak. In the current example, the area of the major component is only slightly biased positively, and the area of the minor component is only slightly under represented at about 4%.

## Slide 26: Peak Resolution and Bias
Take home message:
If possible, fully separate the constituents in either the chromatographic or detection domains
Resolution:
1
2
% error peak 1
% error peak 2
Notes: Another way of examining biases for unresolved components is to evaluate the magnitude of the biases as a function of resolution.
The series of separations shown in this example are again based on synthetically generated chromatograms for two components with the same area ratios (that is, 1 to 1), shown for different resolutions ranging from 0.6 to 2.0.
The resolution equation is shown at the right. There are slightly different forms of this equation that depend on how retention is defined, but each provide similar results. It is important to realize that a resolution of about 1.5 is usually considered representative of full, or baseline resolution, as can be observed in the examples at the left.
This figure shows the biases in the areas determined, plotted as a function of peak resolution. There is a negative bias in the area determined for peak two, and a positive bias in the area determined for peak 1. This has to do with the nature of the chromatographic elution process. As components elute, peak band broadening occurs from diffusion and other processes. More strongly retained components are broader than components that are less retained. Even though the two peaks have the same area, the second, more strongly retained component is broader than the first component. When these components are not fully resolved, the vertical baseline approach to integration does not equally estimate areas due to differences in the peak widths, and biases result.
The take home message is that if possible, fully separate the constituents in either the chromatographic or detection domains. This of course is always the approach that is desired.

## Slide 27: Sources of Bias
Identification of the sources of bias is requisite to unbiased metrology
Quantitation can be biased by problems originating with calibration or sample analysis
Controls provide indication of the “direction” of the bias (i.e., positive vs negative biases)
Clues to biases
  Consistent bias across measurement
  Consistent bias among sample matrix type
Notes: Even when components are fully resolved, measurement biases are still possible.
An understanding of the causal relationships in quantitation is requisite to identifying the sources of measurement bias.
Biases can result from problems that originate with calibration, or with the sample analysis, or both. Interferences in these two aspects of an analysis can bias results in different ways, and it is useful to understand these differences to pinpoint the sources of measurement error.
Control samples provide an indication of the presence of measurement bias. The determination of the “direction” of the bias (that is, whether the bias is positive or negative), is an indicator of potential bias origins.
Other aspects of measurement bias can provide clues to the source of the bias.
For example, if measurements appear to be consistently biased either low or high, this may indicate a problem with calibration.
If the biases appear to fluctuate among different types of samples, then coelution or matrix effects may be at issue.
In the next few slides, the effect of different types of response incongruities are evaluated for calibration and sample analysis components, using an external standard approach to quantitation.

## Slide 28: Sources of Bias: Calibration
Case 1: Unbiased
Assume 1 mass unit analyte
gives a response of 1 area unit
A response of 1 area unit
corresponds to 1 mass unit
Calibrant
Sample
Notes: In the following slides, “truth” denotes theoretical knowledge of the response or amount of substance that is present, in the absence of any measurement bias. In this simple example, 1 mass unit of the analyte gives a response of 1 area unit.
Thus, if a response of 1 area unit is observed, in the absence of any biases, an amount of substance of 1 mass unit is known to be present.

## Slide 29: Sources of Bias: Calibration
Case 2: Positive Bias
Calibrant
Sample
response < truth for a mass of 1 unit
loss of analyte in calibrant due to liquid transfers
analyte volatilizes in calibrant
true reference standard purity < estimate
injection volume < specified volume
answer greater than truth
Notes: In the second case, a situation is represented in which an unknown factor causes the response of the analyte to be less than truth for the calibrant, but the response of the analyte is correct for the sample. The question can be asked: “will the measurement of the analyte in the sample be greater or less than truth”?
In this example, the response factor from calibration is less than truth. A response of 1 area unit from the analysis of the sample corresponds to a larger mass quantity than truth, and the measurement bias is positive.
Several factors could potentially result in the reduced response for the calibrants. During manipulation of the calibrants, liquid transfers might result in loss of the analyte. Likewise, analyte volatility could reduce the quantity of analyte in calibrants. If the purity of the reference standard is less than stated, then the quantity of analyte in the calibrants is reduced. Finally, injection volume can directly affect response, when differences between the actual and assumed injection volumes differ. Other factors may influence response. It should be noted that other modes of quantitation, such as the use of internal standards, may compensate for some of these factors.

## Slide 30: Sources of Bias: Calibration
Case 3: Negative Bias
Calibrant
Sample
response > truth for a mass of 1 unit
concentration of analyte in calibrant
evaporation of calibrant solvent
true reference standard purity > estimate
coeluting interferences in calibrant
injection volume > specified volume
answer less than truth
Notes: In the third case, the opposite situation is represented in which the response of the analyte in the calibrant is greater than truth. Logically, it can be anticipated that the measurement of the analyte in the sample will be biased in the opposite direction compared with the previous case.
In the current example, a response of 1 area unit from the analysis of the sample corresponds to a smaller mass quantity than truth, and the measurement bias is negative.
Factors that could potentially result in increased response for the calibrants include steps that concentrate the analyte in the calibrant, such as solvent evaporation. If the true reference standard purity is greater than the stated or measured purity, this would also result in a calibrant that contains more mass of the analyte than assumed. The response for the analyte would be greater than truth if an interference is present that coelutes with the analyte. If a larger quantity of the calibrant is injected than is assumed, this too would result in an increased analyte response.

## Slide 31: Sources of Bias: Sample
Case 1: Unbiased
Assume 1 mass unit analyte
gives a response of 1 area unit
A response of 1 area unit
corresponds to 1 mass unit
Calibrant
Sample
Notes: In the next series of slides, the effects of biases that originate in the sample are considered. For this discussion, it is assumed that the response factor determined from the calibrants represents truth.
Thus, as before, a response of 1 area unit corresponds to 1 mass unit.

## Slide 32: Sources of Bias: Sample
Case 2: Positive Bias
response of sample > truth
coeluting interferences in the sample
evaporation of sample solvent
injection volume > specified volume
Calibrant
Sample
answer is greater than truth
Notes: Biases that result from issues with the sample are intuitive.
Perhaps most commonly, interferences that are present in the sample and that coelute with the analyte result in analyte levels that are greater than truth.
Likewise, evaporation of the sample, or injection of a larger sample volume than assumed will result in positive measurement biases.

## Slide 33: Sources of Bias: Sample
Case 3: Negative bias
response of sample < truth
volatile analyte in sample
injection volume < specified volume
Calibrant
Sample
answer is less than truth
Notes: When the response of the analyte is somehow reduced for samples, the measurement is negatively biased.
Negative biases can result from analyte volatility, or from the injection of a volume less than the specified quantity.

## Slide 34: Use of Controls
Ideally, a measurement control should offer the same analytical challenges as the sample
  Matched matrix
  Matched analyte levels
  Interferences
  Bulk properties
  Commutability
Analyte levels determined should overlap the certified or reference levels (within measurement uncertainty)
If suitable SRMs are not available, use other commercial or in-house controls (e.g., previously characterized samples, or spiked blanks)
Notes: The use of control samples is an essential component of chemical metrology. Controls offer evidence for the validity of measurements, and provide tools for developing and troubleshooting analytical methods.
Ideally, a measurement control should offer the same analytical challenges as the sample. For example, the control should have a matrix that is matched to the sample, analyte levels should be similar to the sample, the control should have similar interferences, similar bulk properties, and similar commutability.
After analysis of the control, the analyte levels determined should overlap the certified or reference levels (within measurement uncertainty). At the 95% level of confidence, some of the measurements may fall outside this overlapping range and still be within control.
If suitable standard reference materials are not available, the use other commercial or in-house controls is appropriate (e.g., previously characterized samples, or spiked blanks)

## Slide 35: Traceability
Traceability has been defined as:
a “property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty”
Measurements of unknown “C” are traceable to the SI through the CRM
Measurement uncertainty increases with each comparison
Traceability depends on an “unbroken chain of calibrations”
(suppose B is “damaged”)
Notes: The term Traceability is defined differently for different situations. In chemical metrology, the term has been defined by the international community as … a “property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty”
This can be visualized by an illustration showing steps in the determination of mass. In the first step, a certified reference material is used to assign a value to a reference standard denoted as “A”. This reference standard can then be used to assign values a working standard or calibrant “B”. Intermediate comparisons can continue until ultimately, the calibrant is used to assign a value to the unknown sample. Because there is an unbroken chain of comparisons, the measurement of “C” is said to be traceable to the Certified Reference Material. In most cases, the CRM is itself traceable to units of the SI, the highest metrological reference. In this way, measurement traceability can be claimed all of the way to units of the SI. It should be emphasized that each of the comparisons involve measurement uncertainty, and these uncertainty contributions are combined to provide an assessment of uncertainty for the level assigned to the unknown C.
Measurement traceability depends on an “unbroken chain of comparisons” and the next illustration shows what can happen if one of the calibrations is not valid. For example, if the calibrant B were in some way damaged, then the chain of comparisons from the CRM to the unknown sample is interrupted, and the validity of the measurement is compromised. Traceability to the SI is no longer valid.

## Slide 36: Traceability
In principle, measurements that are traceable are also comparable
To assess potential breaks in the traceability chain, ask the question:
“What could go wrong with this measurement?”
  Method suitability?
    Specificity, sensitivity, repeatability, dynamic range, sources of bias
  Instrumentation?
  CRM commutability?
  Extraction efficiency
  Matrix interferences
  Analyte stability
  Calibration
Notes: In principle, measurements that are traceable are also comparable, across time and across different laboratories.
To assess potential breaks in the traceability chain, ask the question: “What could go wrong with this measurement?”
Any number of issues can be listed as possible problems in an analytical method. One can ask the question: Is the method suitable to characterize the properties in question? Does the method provide sufficient Specificity, sensitivity, repeatability, and dynamic range for the samples to be determined? What are the sources of bias for the method?
Is the instrument functioning correctly?
Is the reference material that is being used for claims of traceability, commutable? In this context, the word commutability implies that the reference material is suitable for the application. Materials based on dissimilar sample types may not provide the same analytical challenges, and may not be commutable for the application
Extraction efficiency. There are cases for which less than quantitative recoveries are achieved during sample extraction, and depending on the method, this may result in measurement bias.
Matrix interferences can cause problems for sample analysis
Analyte instability may provide another source of measurement bias
And finally, problems with calibration can affect the accuracy and ultimately the traceability of a measurement.

## Slide 37: Concluding Thoughts…
The practice of higher order chemical metrology requires a significant commitment by the analyst:
  Knowledge of measurement principles
  Laboratory skills
  Implementation of technology
  Attention to detail
Notes: The practice of higher order chemical metrology requires a significant commitment by the analyst that include:
Knowledge of measurement principles
Laboratory skills
Implementation of technology
Attention to detail
Analytical expertise is developed through practice and experience and is a measure of this commitment.

## Slide 38: www.nist.gov
Notes: To learn more about the National Institute of Standards and Technology, visit www.nist.gov