## Software Technical Assessment
Software Technical Assessment
Software Description | WAMRF-014 (Rev. 29), SOP 21 Workbook | Approval Date | 2014-10-22
Assessed By | Dan Wright | Approved By | Dan Wright
Codes | Assessment | Pass/Fail | Evidence
A. Software Inspection | Spreadsheet is clear and makes sense | Pass | All worksheet are appropriately named and easy to understand
| There are instructions for use | Pass | Instructions worksheet
| Instructions and data input appear on the visible portion of the first worksheet | Pass | The workbook opens to the instructions worksheet
| Data-entry fields are labeled and color coded (it is recommended to avoid red and green) | Pass | Light yellow for blank, light blue for non-blank cells
| The Standard Operating Procedure (SOP) used is clearly specified | Pass | Specified on Uncertainty worksheet
| The number of digits to be rounded to is specified | Pass | Water Density is rounded to 9 places and the volume correction and uncertainty are rounded according to NISTIR 6969, GLP 9 rounding rules
| The user is warned/notified whenever data-entry fields are left blank | Pass | In data validation cells only
| Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations | Pass | Data entry fields on the workbook template are blank. Instructions tell user to break links before saving after a calibration is performed
| The software opens at the right location within the file | Pass | Opens to the Instruction worksheet
| Unused fields/cells are locked | Pass | Unused cells are locked and worksheets are protected
| Rows/columns that the operator need not see are hidden | N/A | No hidden cells or worksheets
| Unused sheets are removed | Pass | No unused worksheets
| Worksheets are named appropriately | Pass | Worksheets are appropriately named
B. Mathematical Specification | The correct SOP is used | Pass | SOP 21 in embedded
| The formulae and methods chosen from that SOP are specified | Pass | Volume Transfer method
| Sources and references for formulae are specified | Pass | As specified in SOP
| The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty | Pass | For Volume II calibrations
C. Code review | The formulae in the fields exactly match the SOP | Pass | Formulae on calculation worksheet match SOP
| Repeated calculations appropriately reference the correct cells | Pass | Calculation worksheet
| Calculations, when tested using standard data or reference test data, show appropriate accuracy | Pass | Tested using SOP example
| Rounding is done at the appropriate locations in the file | Pass | Water Density and end volume results rounded
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences | Pass | No issues encountered
| Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa | Pass | Formatted correctly throughout workbook
| “Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used | Pass | Where appropriate
E. Component Testing | Each macro used is functional | Pass | FORMAT_CX and ROUND_UNC
| Each command/button is functional | N/A | No command buttons
| Combinations of interdependent macros are functional | Pass | No issues encountered
| Plotted graphs are accurate | Pass | Graphs on uncertainty worksheet display properly
| Worksheets/reports print properly, if needed to | Pass | Worksheets formatted for printing as necessary
| Conditional (color and non-color) formatting is functional | Pass | On data entry and calculation worksheets
F. Numerical Reference | Look-up tables and lists match the latest calibration report. | Pass | Linked tables are validated and up to date
| Uncertainties match the latest Scope | Pass | Test on uncertainty worksheet will not allow an uncertainty that is better than the latest published scope
| Values that reference another workbook or spreadsheet are dated | Pass | Linked references have been validated and are up to date
| When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user | N/A
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct | Pass | All embedded data has been validated
| The evaluation of the embedded data is dated and documented | Pass | Records keep in Software Validation folder on computer hard drive
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented | Pass | Evaluated and filed
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty | N/A | No hand calculations were done
J. Security | Equation and calculation cells are protected against inadvertent editing | Pass | Cells are locked and worksheets are password protected
| Cells are locked in place; they cannot be moved/dragged | Pass | Locked and password protected
| Confidentiality of passwords is appropriate | Pass | State Metrologist and Program Manager only have access
| Files are backed up automatically | Pass | Backed up to agency servers weekly
| Additional back-up is available at alternate facilities | Pass | Agency IT backs up servers weekly
| Files on network drives cannot be accidentally deleted | Pass | Agency IT keeps an additional back up at a remote location

## Documentation
Documentation
Date | Description | Initials
2004-01-30 | This workbook was developed by Dan Wright, Washington State Department of Agriculture Metrology Laboratory. It follows the procedure documented in NISTIR 7383, SOP 21. NIST/WMD has evaluated this spreadsheet and provided useful feedback in its development. | DAW
2004-03-03 | Updated water density formula to the Patterson/Morris equation. | DAW
2004-03-05 | Updated water density formula to the Patterson/Morris equation on the 'Uncertainty Analysis' worksheet. | DAW
2004-03-11 | Made suggested changed from NIST, WMD. Revised pressure measurements. | DAW
2004-07-06 | Updated uncertainty analysis to new format. | DAW
2004-09-10 | Updated uncertainty analysis. | DAW
2005-03-22 | Corrected Z60 formula on 'Uncertainty Analysis' worksheet. | DAW
2005-10-16 | Removed unnecessary information on 'NSCV Data Entry' worksheet. | DAW
2006-09-01 | Complete revision. | DAW
2006-11-14 | Changed reference to ASTM 674 to ASTM 694 on 'Uncertainty Analysis' worksheet. | DAW
2006-11-14 | Deleted uncertainty for viscosity and added statement to report of calibration. | DAW
2006-11-14 | Deleted note about standard reference temperature. | DAW
2007-03-16 | Major revision. Updated to Kragten uncertainty spreadsheet, added tolerances, added metric vessel capability, etc. | DAW
2007-03-30 | Reformatted standards tables, corrected errors in 'LPG Data Entry' worksheet. | DAW
2007-04-05 | Fixed conditional formatting in uncertainty test on the 'Calculations' worksheet. | DAW
2007-09-24 | Rounded water density formulas to 8 places, revised tolerance evaluation, and revised 'NSCV Data Entry' worksheet. | DAW
2008-01-10 | Revised 'NSCV Data Entry' worksheet. | DAW
2008-02-07 | Removed 'NSCV Data Entry' worksheet. | DAW
2008-06-10 | Added 'Control Data' worksheet, removed 'Standards' worksheet, reformatted vlookup formulas to global standards workbooks, and revised 'Documentation' worksheet. | DAW
2009-01-21 | Color coded data entry cells. Edited all comments. | DAW
2009-01-21 | Added Standards Used table on Data Entry worksheet. | DAW
2009-10-30 | Added Information worksheet. | DAW
2009-10-30 | Validated ranges. | DAW
2009-10-30 | Validated workbook, file name "WA2009-10-30 WAMRF-014 (Rev. 20), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2010-03-26 | Updated criteria for tolerance and uncertainty evaluations on the "Calculations' sheet to NIST HB 105-4(2010) specifications. | DAW
2010-03-26 | Removed NSCV value. The neck scale either meets the new specification tolerances or it is rejected. | DAW
2011-03-26 | Removed '/NSCV' from cell H82 formula on the Data Entry worksheet. | DAW
2011-03-26 | Added drain time and drain method data to 'Report Pages 1-3', 'Data Entry', and 'Tables and Lists' worksheets. | DAW
2010-03-26 | Added Software Technical Assessment and SOP 21 worksheets, renamed Information worksheet to Instructions. | DAW
2011-03-26 | Validated ranges. | DAW
2010-03-26 | Validated workbook, file name "WA2010-03-26 WAMRF-014 (Rev. 21), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2012-02-13 | Added FORMAT_CX and ROUND_UNC visual basic modules for rounding results and uncertainties IAW NISTIR 6969, GLP 9 using the FmtCX() and Rnd2() functions. | DAW
2012-02-13 | Added SI statement and conversion to the Traceability paragraph on the Report Pages 1-3 spreadsheet.
2012-02-13 | Validated ranges. | DAW
2012-02-13 | Validated workbook, file name "WA2012-02-13 WAMRF-014 (Rev. 22), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2012-04-06 | Changed "Certification Statement" to "Accreditation Statement" on Report Pages 1-3 worksheet to be in compliance with NVLAP requirement Annex A, l). | DAW
2012-05-28 | Embedded 2012 SOP | DAW
2012-06-03 | Updated water density formula on Calculations worksheet | DAW
2012-06-03 | Updated reference to the new water density formula on the Uncertainty Analysis worksheet | DAW
2012-06-03 | Added degrees of freedom and best uncertainty k factor to the Data Entry worksheet | DAW
2012-06-03 | Updated k factor based on degrees of freedom IAW SOP 29 (2012 edition) | DAW
2012-06-03 | Added k factor column to report page 4 and updated uncertainty statement on report page 1-3 | DAW
2012-06-03 | Validated ranges. | DAW
2012-06-03 | Validated workbook, file name "WA2012-06-03 WAMRF-013 (Rev. 23), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2012-08-13 | Inserted new WSDA Logo on "Report Pages 1-3" worksheet. | DAW
2013-02-18 | Removed GLP Rounding macros | DAW
2013-02-18 | Entered new rounding technique and corrected formulas | DAW
2013-02-18 | Validated range names. | DAW
2013-02-18 | Validated workbook, file name "WA2013-02-18 WAMRF-013 (Rev. 24), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2013-04-02 | Updated the tolerance test formulas and conditional formatting to include the uncertainty on Report Pages 1-3, row 82, Report Page 4-5, cells B5 & C5, Calculations, row 42, cells H45 & H44. | DAW
2013-04-02 | Validated range names. | DAW
2013-04-02 | Validated workbook, file name "WA2013-04-02 WAMRF-014 (Rev. 25), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2013-06-26 | Corrected gallon to cubic meter conversion on line 93 of the Report Pages 1-3 spreadsheet. | DAW
2013-09-01 | Changed all ISBLANK() statements. | DAW
2013-09-01 | Redefined uncertainty components, added the Welch-Satterthwaite formula for effective degrees of freedom. | DAW
2013-09-01 | Revised the Control Data worksheet | DAW
2013-09-01 | Validated links and range names and spell checked. | DAW
2013-09-01 | Validated workbook, file name "WA2013-09-01 WAMRF-014 (Rev. 26), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-01-10 | Linked LPG Data Entry cells F18, H18 & J18 to the YYYY_BEST_UNC workbook. | DAW
2014-06-30 | Updated data input and uncertainties in accordance with NISTIR 7383, SOP 21 (2013 edition). | DAW
2014-06-30 | Validated range names and spell checked. | DAW
2014-06-30 | Validated workbook, file name "WA2014-06-30 WAMRF-014 (Rev. 27), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-08-25 | Changed the reference in cell B31 on the Pressure Corrections worksheet from J96 to E102 of the Prover Data Entry worksheet. | DAW
2014-08-25 | Validated range names and spell checked. | DAW
2014-08-25 | Validated workbook, file name "WA2014-08-25 WAMRF-014 (Rev. 28), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-10-22 | Deleted viscosity uncertainty, not enough information to analyze this uncertainty. | DAW
2014-10-22 | Changed statistical df from 100 to 10000. | DAW
2014-10-22 | Validated range names and spell checked. | DAW
2014-10-22 | Validated workbook, file name "WA2014-10-22 WAMRF-014 (Rev. 29), SOP 21 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW

## SOP 21
Note on A1: Unprotect the worksheet, then double click the document to open in Adobe Acrobat
Unprotect the worksheet, then double click the document to open in Adobe Acrobat

## Instructions
Instructions
No. | Description
1 | This workbook must be validated after any changes or modifications.
2 | The worksheets are password protected. Password is password. You should protect each sheet with your own unique password. Place a password hint here so you don't forget it. Password hint is: "?"
3 | Validate all links and ranges after any changes and during annual reviews.
4 | Modify the report pages with your laboratory and accreditation authority information.
5 | Modify the footer information on the report sheets and the header information on all other sheets if you want to use your own form number for the template.
Range Names
Date | Description
_unc1 | ='Prover Data Entry'!$H$37
_unc2 | ='Prover Data Entry'!$H$38
_unc3 | ='Prover Data Entry'!$H$39
_unc4 | ='Prover Data Entry'!$H$40
a_1 | ='Prover Data Entry'!$F$43
a_2 | ='Prover Data Entry'!$F$44
a_3 | ='Prover Data Entry'!$F$45
a_4 | ='Prover Data Entry'!$F$46
Address_1 | ='Prover Data Entry'!$C$5
Adjusted | ='Prover Data Entry'!$J$97
AirTemp | ='Prover Data Entry'!$D$54
AirTemp_F | ='Prover Data Entry'!$D$108
answer_list | ='Tables & Lists'!$D$14:$D$15
AsFound | =Calculations!$E$63
AsLeft | =Calculations!$E$64
B | ='Prover Data Entry'!$I$14
BestU | ='Prover Data Entry'!$F$20
BestUdf | ='Prover Data Entry'!$K$20
BestUk | ='Prover Data Entry'!$I$20
Bottom_Neck_id | ='Prover Data Entry'!$G$15
BottomZero | ='Prover Data Entry'!$J$85
Cal_Date | ='Prover Data Entry'!$C$22
CCE_Table | ='Tables & Lists'!$A$5:$B$7
Condition | ='Prover Data Entry'!$C$16
CustomerID | ='Prover Data Entry'!$K$4
Description | ='Prover Data Entry'!$C$10
df | ='Prover Data Entry'!$K$19
drops1 | ='Prover Data Entry'!$E$37
drops2 | ='Prover Data Entry'!$E$38
drops3 | ='Prover Data Entry'!$E$39
drops4 | ='Prover Data Entry'!$E$40
Eff_df | =Calculations!$E$65
Final_k | ='Uncertainty Analysis'!$K$106
Final_U | ='Uncertainty Analysis'!$K$107
Final_Veff | ='Uncertainty Analysis'!$K$105
GradNeckVol | ='Prover Data Entry'!$J$12
Humidity | ='Prover Data Entry'!$G$54
Humidity_F | ='Prover Data Entry'!$G$108
increment | ='Prover Data Entry'!$J$13
interval | ='Prover Data Entry'!$K$21
intervalQ | ='Prover Data Entry'!$F$21
JobNo | ='Prover Data Entry'!$J$16
k | ='Uncertainty Analysis'!$K$100
k_1 | ='Prover Data Entry'!$J$37
k_2 | ='Prover Data Entry'!$J$38
k_3 | ='Prover Data Entry'!$J$39
k_4 | ='Prover Data Entry'!$J$40
kfactor | =Calculations!$E$66
Line_w | ='Prover Data Entry'!$K$15
Material | ='Prover Data Entry'!$C$14
Material_list | ='Tables & Lists'!$A$5:$A$7
MeasuredU | ='Uncertainty Analysis'!$K$101
MFG | ='Prover Data Entry'!$C$11
mmHg | ='Prover Data Entry'!$J$18
Nom_Val | ='Prover Data Entry'!$C$12
NomVal_list | ='Tables & Lists'!$D$17:$D$18
NomValUnit | ='Prover Data Entry'!$A$12
PO_No | ='Prover Data Entry'!$I$5
POC_Name | ='Prover Data Entry'!$I$6
POC_Phone | ='Prover Data Entry'!$I$7
PooledSD | ='Prover Data Entry'!$F$19
Print_Titles | =Instructions!$5:$6
ProverTempCorr | ='Prover Data Entry'!$F$51
ProverTempk | ='Prover Data Entry'!$J$51
ProverTempU | ='Prover Data Entry'!$H$51
RangeMean_list | ='Tables & Lists'!$D$9:$D$12
RefT | ='Prover Data Entry'!$J$17
RefTemp_list | ='Tables & Lists'!$D$20:$D$21
RefTempUnit | ='Prover Data Entry'!$A$17
Rmean | ='Prover Data Entry'!$F$19
RptNo | ='Prover Data Entry'!$I$1
Scale_list | ='Tables & Lists'!$D$4:$D$7
Scale_Unit | ='Prover Data Entry'!$E$12
SealNoBottom | ='Prover Data Entry'!$I$105
SealNoTop | ='Prover Data Entry'!$D$105
SN | ='Prover Data Entry'!$J$11
Standards_Table_A | ='Prover Data Entry'!$A$36:$K$40
Standards_Table_B | ='Prover Data Entry'!$A$42:$K$46
StdTempCorr | ='Prover Data Entry'!$F$50
StdTempk | ='Prover Data Entry'!$J$50
StdTempU | ='Prover Data Entry'!$H$50
t_1 | ='Prover Data Entry'!$E$81
t_2 | ='Prover Data Entry'!$J$81
Tech | ='Prover Data Entry'!$F$22
Tolerance | =Calculations!$E$62
Top_Neck_id | ='Prover Data Entry'!$C$15
U60_1 | =Calculations!$E$21
U60_2 | =Calculations!$J$21
Unc | =Calculations!$E$67
V_1 | ='Prover Data Entry'!$F$37
V_2 | ='Prover Data Entry'!$F$38
V_3 | ='Prover Data Entry'!$F$39
V_4 | ='Prover Data Entry'!$F$40
Veff | ='Uncertainty Analysis'!$K$99
WODate | ='Prover Data Entry'!$J$10
Z60_1 | =Calculations!$E$27
Z60_2 | =Calculations!$E$28

## Report Pages 1-3
REPORT OF CALIBRATION
Issued To:
Point of Contact:
Purchase Order Number:
Report Number:
Calibration Date
Calibration Due Date
This is to certify that the information contained in this report is true and correct as of the date of calibration.
Dan Wright, Program Specialist 3 |  |  |  |  | Date of Issue
Accreditation LOGO |  | Accredited by the Accreditation Authority for the specific scope of accreditation under lab code XXXXXX-X. This report may not be used to claim product endorsement by Accreditation Authority or any other government agency, and may not be reproduced, except in full, without written approval from the laboratory.
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date
Artifact(s) Description
Test Item: | L LPG Prover |  |  | Date Received:
Serial Number: |  |  |  | Manufacture:
Material: |  |  |  | Material CCE:
Condition: |  |  |  | Specification: | NIST HB 105-4
Calibration Information
Job Order #: |  |  |  | Temperature:
Metrologist: |  |  |  | Humidity:
Procedure: | NISTIR 7383, SOP 21 |  |  | Water Temperature:
Laboratory Reference Standards Used
Description |  | Serial Number | Cert. Number |  | Cal Date | Cal Due
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
Traceability Statement
The artifact(s) described in this report have been compared to the Standards of the State of Washington. The Standards of the State of Washington are traceable to the National Institute of Standards and Technology (NIST) and are part of a comprehensive measurement assurance program for ensuring continued accuracy and measurement traceability within the level of uncertainty reported by this laboratory. The International System of Units (SI) for volume is the cubic meter (m³) (see Conversion Factors on page 3). The report number for this report is the only unique report number to be used in referencing measurement traceability for the artifact(s) described in this report.
Uncertainty Statement
The combined standard uncertainty includes uncertainties for the standard(s), for the measurement process, for viscosity, for the material cubical coefficient of expansion, for reading meniscus, for the pressure gauge, for graduated neck errors and for the thermometer(s) used for measuring the water temperature. The combined standard uncertainty is multiplied by a coverage factor, k, to give the expanded uncertainty, which defines an interval with a 95.45 % level of confidence. The expanded uncertainty presented in this report is consistent with NIST Technical Note 1297. A component for the effects of viscosity was not included in the uncertainty budget.
Accreditation Statement
Accredited by the Accreditation Authority for the specified scope of accreditation under lab code XXXXXX-X. This laboratory meets the requirements of ISO/IEC 17025 and ANSI/NCSL Z540-1.
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date
Pertinent Information
● In-accordance-with ISO/IEC FDIS 17025, General Requirements for the Competence of Testing and Calibration Laboratories, paragraph 5.10.4.4 ‘A calibration certificate (or calibration label) shall not contain any recommendation on the calibration interval except where this has been agreed with the client. This requirement may be superseded by legal regulations.’
● In-accordance-with Washington Administrative Code (WAC) Chapter 16-663, Service Agents -- Reporting, Test Procedures, Standards And Calibration Of Weighing And Measuring Devices, Section 16-663-130, Adequacy of standards and submission of standards for certification, paragraph 2, ‘… All standards used for servicing, repairing and/or calibrating commercial weighing and measuring devices must be submitted at least every two years for examination and certification…’
● The artifact is considered in-tolerance when the error is equal to or less than the specified tolerance minus the measurement uncertainty. RED print indicates an out-of-tolerance reading.
● The LPG Prover 'As Left' volume is 'In-Tolerance', the prover may be used in meter testing without a correction.
● Enter the Pressure Correction from Table 1 that corresponds with the pressure being tested on your LPG Meter Test form.
● The calibration item was calibrated in a 'wet down' condition using water. The calibration data above applies when the prover is drained for a 30 (± 5) second period after cessation of the main flow.
● The drain time (using ) was approximately minute(s) seconds.
● The Top Securty Seal Number is and the Bottom Security Seal Number is .
● The results listed in this report relate only to the artifacts described and extent of calibrations performed.
Conversion Factors
From NIST Special Publication 811, Guide for the Use of the International System of Units (SI)
Factors in boldface are exact
To convert from |  |  | to |  | multiply by
gallon (U.S.) (gal) |  |  | to cubic meter (m³) |  | 3.875 412 E-03
cubic inch (in³) |  |  | to cubic meter (m³) |  | 1.638 706 4 E-05
liter (L) |  |  | to cubic meter (m³) |  | 1.0 E-03
Note on A1: Break links to lock in data.
Break links to lock in data.

## Report Pages 4-5
WSDA Weights and Measures Metrology Laboratory
Report Number: |  |  |  | Calibration Date
Calibration Results
Nominal Volume (at zero mark on gauge) | Prover Volume As Found FALSE | Prover Volume As Left FALSE | NIST HB 105-4 Specification Tolerance ± (gal) | Uncertainty (k=) (df=) ± (gal)
Table 1 - LPG Prover Corrections @
psig | Prover Scale Reading (in³) | Pressure Correction (Pcorr) (in³) | Prover Error (Not Adjusted) (gal) | Prover Volume (Not Adjusted) (gal)
Chart 1 - Pressure Correction (Pcorr) (in³) |  |  |  | Correction (in³)

## Report Page 6
WSDA Weights and Measures Metrology Laboratory
Report Number: |  |  |  | Calibration Date
Table 2 - Prover Temperature Correction Chart
Coefficient of Cubical Expansion = / ºF |  |  | A negative error means the prover is small A prover is smaller at temperatures lower than the reference temperature
Temperature ºF | Correction (in³) | Correction (gal)
-10
-5
0
5
10
15
20
25
30
35
40
45
50
55
60
65
70
75
80
85
90
95
100
105
110
115
120
125
130

## Prover Data Entry
Prover Data Entry | [6 empty] | Report Number:
Customer Information
Customer ID
PO#
POC
Phone
Artifact and Measurement Information
Description |  | LPG Prover |  |  |  | Date Received
Manufacture | [5 empty] | Serial Number
Graduated Neck Volume (L)
Did the Neck Scale Plate Verification pass?
Prover Material | [5 empty] | CCE (β) (/ ºF) =
Top Neck Inside Dia. (mm) |  |  | Bottom Neck Inside Dia. (mm) |  |  |  | Graduation Line Width (mm)
Artifact Condition | [6 empty] | Job Order Number
Normal Barometric Pressure for Location of Calibration (mm Hg)
Process Standard Deviation of Trials from Control Chart (L) | [6 empty] | s(p) Degrees of Freedom
Accredited Best Uncertainty for this Calibration (L) | [6 empty] | k factor |  | df
Was a Calibration Interval Requested or is one Required? | [5 empty] | What is the Calibration Interval (months)?
Date of Calibration |  |  |  | Technician |  |  |  | Technician Initials
Standards Used
Standard Item ID | Description |  |  |  | Serial Number |  | Report Number |  | Calibration Date | Calibration Due
Volume Standard(s) Data
ID Code | Description |  |  | Drops | Volume (Vs @ trefs) (gal) |  | Unc (gal) |  | k-factor | df
[5 empty] | 0 |  | 0 |  | 1 | 1
[5 empty] | 0 |  | 0 |  | 1 | 1
[5 empty] | 0 |  | 0 |  | 1 | 1
[5 empty] | 0 |  | 0 |  | 1 | 1
ID Code | Description |  |  | Std. Ref. Temp. (trefS ºF) | CCE (α) (/ ºF) |  | Graduated Neck Volume (gal) | Inside Neck Diameter (mm) | Sight Glass Inside Dia. (mm) | Graduation Line Width (mm)
|  |  |  | 0 | 0 |  | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 |  | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 |  | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 |  | 0 | 0 | 0 | 0
Temperature Standard(s) Data
Description |  |  |  | ID Code | Correction (ºC) |  | Uncertainty (ºC) |  | k factor
Temp. Standard Used For Volume Standard(s):
Temperature Standard Used For Unit-Under-Test:
Laboratory Test Conditions
Start | Temperature (ºC) |  |  | Humidity (% RH)
Standard(s) Observations
Trial 1 |  |  |  |  | Trial 2
Drop Number | ID Code | Volume (Vs @ tref) (gal) | Water Temp. (ºC) | Delta (Δ) (in³) | Drop Number | ID Code | Volume (Vs @ tref) (gal) | Water Temp. (ºC) | Delta (Δ) (in³)
1 |  |  |  |  | 1
2 |  |  |  |  | 2
3 |  |  |  |  | 3
4 |  |  |  |  | 4
5 |  |  |  |  | 5
6 |  |  |  |  | 6
7 |  |  |  |  | 7
8 |  |  |  |  | 8
9 |  |  |  |  | 9
10 |  |  |  |  | 10
11 |  |  |  |  | 11
12 |  |  |  |  | 12
13 |  |  |  |  | 13
14 |  |  |  |  | 14
15 |  |  |  |  | 15
Artifact Observations
Trial 1 |  |  |  |  | Trial 2
Neck Scale Reading @ 0 psig (in³) |  |  |  |  | Neck Scale Reading @ 0 psig (in³)
Water Temperature, Bottom (ºC) |  |  |  |  | Water Temperature, Bottom (ºC)
Water Temperature, Middle (ºC) |  |  |  |  | Water Temperature, Middle (ºC)
Water Temperature, Top (ºC) |  |  |  |  | Water Temperature, Top (ºC)
Water Temperature, Average (ºC) |  |  |  |  | Water Temperature, Average (ºC)
Pressure Measurements
Adjust prover for Nominal Volume at ºF and 100 psig
Does the prover have a bottom scale plate?:
If bottom gauge scale plate is adjustable, enter inside diameter of the lower neck (in)
Neck Scale Reading As Found @ 100 psig (in³)
Desired Neck Scale Reading at ºF and 100 psig is | [5 empty] | TBD
If the upper and/or lower gauge scale plate is adjustable perform only ONE of the following adjustments:
A. | If the top scale is adjustable, then adjust the upper scale to read the Desired Gauge Reading at ºF and 100 psig.
B. | If the bottom scale is adjustable, adjust the bottom scale: | [6 empty] | N/A
Note: If lower scale was adjusted, print this form, including calculations sheet, drain, rezero, and retest before the Increasing / Decreasing Pressure Test.
i | It is recommended that LPG provers be adjusted, as close as possible, to the nominal volume at 100 psig. If neither scale is adjustable, a prover volume at 100 psig must be calculated and reported.
i | It is best to report the actual volume at 100 psig to ensure that the owner/user applies the correction properly. (A high neck scale reading at 100 psig means the prover is small and the volume reported at ºF and 100 psig should be less than the nominal value of the prover.) The prover error is added to the nominal volume to determine the prover volume (see the Pressure Corrections worksheet).
i | Include a Pressure Correction Table 1 and Chart with your Report of Calibration for the owner/user to use during their testing of LPG Meters (see the Report spreadsheet).
Neck Scale Reading As Left @ 100 psig (in³)
Was upper or lower scale adjusted for nominal volume @ 100 psig during this calibration?
Increasing / Decreasing Pressure Test.
Increasing Pressure |  |  |  |  | Decreasing Pressure
Neck Scale Reading @ 0 psig (in³) |  |  |  |  | Neck Scale Reading @ 150 psig (in³)
Neck Scale Reading @ 50 psig (in³) |  |  |  |  | Neck Scale Reading @ 100 psig (in³)
Neck Scale Reading @ 100 psig (in³) |  |  |  |  | Neck Scale Reading @ 50 psig (in³)
Neck Scale Reading @ 150 psig (in³) |  |  |  |  | Neck Scale Reading @ 0 psig (in³)
Neck Scale Reading @ 200 psig (in³)
Top Scale Security Seal # |  |  |  |  | Bottom Scale Security Seal #
Laboratory Test Conditions
Finish | Temperature (ºC) |  |  | Humidity (% RH)
Prover Drain Time/Method
The prover drain time was: |  |  |  |  | minute(s) |  | seconds(s)
The prover was drained using:

## Control Data
Control Chart Information
Note: Copy and Paste Appropriate Row(s) to Control Chart
R-Chart Information
Tech Initials | Date | Test Number | Volume (Trial 1) FALSE | Volume (Trial 2) FALSE | Start Humidity (% RH) | Finish Humidity (% RH) | Start Air Temp. (ºC) | Finish Air Temp. (ºC) | Water Temp. (Trial 1) (ºC) | Water Temp. (Trial 2) (ºC)

## Calculations
Calculations
Measurement Trial Calculations
Volume Standard(s) Trial 1 |  |  |  |  | Volume Standard(s) Trial 2
Drop # | Std. Ref. Temp. (trefS ºF) | Water Temp. (ºF) | Water Density (g/cm³) | Delivered Volume | Drop # | Std. Ref. Temp. (trefS ºF) | Water Temp. (ºF) | Water Density (g/cm³) | Delivered Volume
1 |  |  |  |  | 1
2 |  |  |  |  | 2
3 |  |  |  |  | 3
4 |  |  |  |  | 4
5 |  |  |  |  | 5
6 |  |  |  |  | 6
7 |  |  |  |  | 7
8 |  |  |  |  | 8
9 |  |  |  |  | 9
10 |  |  |  |  | 10
11 |  |  |  |  | 11
12 |  |  |  |  | 12
13 |  |  |  |  | 13
14 |  |  |  |  | 14
15 |  |  |  |  | 15
Trial 1 Delivered Volume (gal) = |  |  |  |  | Trial 2 Delivered Volume (gal) =
LPG Prover Trial 1 |  |  |  |  | LPG Prover Trial 2
Water Density (g/cm³) = |  |  |  |  | Water Density (g/cm³) =
Water Temperature (ºF) = |  |  |  |  | Water Temperature (ºF) =
LPG Prover Volume Determinations At The Designated Reference Temperature
Trial 1 VXtrefX =
Trial 2 VXtrefX =
Trial 1 LPG Prover Volume =
Trial 2 LPG Prover Volume =
LPG Volume Determinations At The Designated Reference Temperature And 100 psig
Prover Volume As Found =
Prover Volume As Left =
Uncertainty Evaluation
NIST HB 105-4 Criteria: The expanded uncertainty for calibration must be less than 0.2 % of the measured volume.
0.2 % of Measured Volume =
Expanded Uncertainty =
Tolerance Evaluation
NIST HB 105-4 Criteria: The maximum permissible errors shall be ± 0.2 % of the nominal capacity minus the uncertainty.
Tolerance =
Volume Error As Found =
Volume Error As Left =
Measurement Evaluation
NIST HB 105-4 Criteria: Must be capable of repeating within 0.02 % of the test volume.
0.02 % of measure Volume =
Measurement Range =
Measurement Control Calculation
Plot on Control Chart
Std. Dev. of the Trials =
Increasing / Decreasing Pressure Test
Criteria: Must agree within 0.02 % of prover nominal volume.
Results Formatted For Calibration Report
LPG Prover Nominal Value =
Tolerance =
LPG Prover Volume As Found =
LPG Prover Volume As Left =
Effective Degrees of Freedom =
k factor =
Uncertainty (uc*k) =
Note on A57: Metric Prover Pressure Test Formula
Note on J57: Customary Prover Pressure Test Formula
Metric Prover Pressure Test Formula
Customary Prover Pressure Test Formula

## Pressure Corrections
Pressure Corrections
Pressure Correction Formulas
Prover Error at Reference Temperature and Selected Pressure
where
| VXtrefX = | Prover volume at reference temperature and 0 psig
| VNominal = | Nominal volume of prover
| Wcf = | Water compressibility correction factor (See SOP 21, paragraph 3.3)
| Sr = | Scale reading
| Npsig = | Selected applied pressure
Pressure Correction
where
| Pcorr @ Npsig= | Prover correction at selected psig
| Wcv = | Water compressibility value gal/gal at 100 psig
Pressure Correction Calculations
psig | Prover Scale Reading (in³) | Pressure Correction (Pcorr) (in³) | Prover Error (Adjusted) (gal) | Prover Error (Not Adjusted) (gal) | Prover Volume (Not Adjusted) (gal)
0
10
20
30
40
50
60
70
80
90
100
110
120
130
140
150
160
170
180
190
200

## Water Compressability
Compressibility of Water
Compressibility of Water Equations and Calculations
The Thermal Compressibility of Water (kT) equation is from the article by Frank E. Jones and Georgia L. Harris, "ITS-90 Density of Water Formulation for Volumetric Standards Calibration", as published in the Journal of Research of the National Institution of Standards and Technology, Vol. 97, No. 3, 1992.
kT (kPa)-1 = | 50.831 01x10-8 - 3.682 93x10-9t + 7.263 725x10-11t2 - 6.597 702x10-13t3 + 2.877 67x10-15t4
kT (kPa)-1 =
where
t = temperature of water
Thermal Compressibility of Water Based on Calibration Location Atmospheric Pressure
kT (atm) =
Note: This worksheet takes into account the typical atmospheric pressure at the location of the calibration.
Water Compressibility Value
Wcv = ( Vnom ) ( kT ) ( P )
Water Compressibility Factor
Wcf = 1 + ( kT ) ( P )
where
Vnom = | Nominal volume of the prover in gallons or liters
kT = | Thermal compressibility of water in atmospheres (atm)
P = | Pressure in atm
Conversion Factors (from NIST SP 811)
To Convert From |  | To | Multiply By
atmosphere (standard) |  | Pascal (Pa) (exactly) | 101325
pounds per square inch (psi) |  | Pascal Pa | 6894.757
millimeter of mercury (0 ºC) (mm Hg) |  | Pascal Pa | 133.322
Calibration Data
=
Designated Reference Temperature (ºC) =
Typical Barometric Pressure at Your Location (mm Hg) =
or Pascal (Pa) =
Compressibility of Water Versus Pressure Calculations
Compressability Value per gallon at mm Hg =
Pressure (psig) | Pressure (atm) | Wcf | Wcv (in³) | Wcv (gal)
0 | 0
10 | 0.68046
20 | 1.3609
30 | 2.0414
40 | 2.7218
50 | 3.4023
60 | 4.0828
70 | 4.7632
80 | 5.4437
90 | 6.1241
100 | 6.8046
110 | 7.4851
120 | 8.1655
130 | 8.846
140 | 9.5264
150 | 10.2069
160 | 10.8874
170 | 11.5678
180 | 12.2483
190 | 12.9287
200 | 13.6092

## Uncertainty Analysis
Uncertainty Analysis
Specify Procedure
NISTIR 7383, SOP 21, Calibration of LPG Provers
Specify Equation(s)
Volume at the Reference Temperature
Volume = VXtrefX - gauge reading
where
| VXtrefX = | Volume of the unknown vessel, VX at its designated reference temperature, trefX
| VStrefS = | Volume of the standard vessel, VS at its designated reference temperature, trefS
| ρ1, ρ2, …, ρN = | Density of the water in the standard where ρ1 is the density of the water for the first delivery, ρ2 is the density of the water for the second delivery, and so on until all N deliveries are completed using the M. Tanaka, G. Girard, R. Davis, A. Peuto, and N. Bignell equation, Metrologia, 38, 301-309 (2001) as referenced in NISTIR 7383, GLP 10.
| Δ1, Δ2,…,ΔN = | Volume differences between the water level and the reference mark on the standard where the subscripts 1, 2,…,N, represent each delivery as above. If the water level is below the reference line, Δ is negative. If the water level is above the reference line, Δ is positive. If the water level is at the reference line, Δ is zero NOTE: units must match volume units for the standard
| t1, t2, …, tN = | Temperature of water for each delivery with the subscripts as above
| α = | Coefficient of cubical expansion for the standard in its designated units
| β = | Coefficient of cubical expansion for the unknown vessel in its designated units
| tx = | Temperature of water in the filled unknown vessel in designated units
| ρx = | Density of water in the prover in g/cm3
Identify Components
Component | Description |  |  |  |  | Reference
s(p) | Uncertainty associated with the measurement process from the average range of several sets of measurements, normal distribution, 1s |  |  |  |  | NISTIR 6969, SOP 29, Para. 3.2.1.2. Standard deviation from a set of replicate measurements and NISTIR 6969, Statistical Techniques, paragraphs 8.3 & 8.4, Pooled standard deviation from standard deviation for each set of replicates from a control chart, Type A evaluation
u(s) | Uncertainty associated with the weights used, normal distribution, k =1 |  |  |  |  | Calibration Report divided by the k-factor, Type B Evaluation
u(m) | Uncertainty associated with reading the meniscus of the standard and unknown vessel, triangular distribution |  |  |  |  | NISTIR 7383, GMP 3, page 4, equation for volume of a cylinder, Type B evaluation
u(n) | Uncertainty associated with the graduated neck calibration of the standard and unknown vessel, triangular distribution |  |  |  |  | NISTIR 7383, SOP 19, paragraph 2.4.2.7 and SOP 21, paragraph 2.4.3.7, 0.5 % of graduated neck volume (if passes neck scale verification), Type B evaluation
u(CCE) | Uncertainty of the Cubic Coefficient of Expansion for the vessel material of the standard and the unknown, rectangular distribution |  |  |  |  | EURAMET cg-19, paragraph 5.3.6 5 % to 10 % of CCE
u(tw) | Uncertainty of the thermometer used for measuring water temperature, rectangular distribution |  |  |  |  | NIST HB 143, Table 11, accurate to ± 0.5 ºC, Type B evaluation
u(psig) | Uncertainty associated with the pressure gauge, rectangular distribution |  |  |  |  | NIST HB 105-4, paragraph 3.12.2, 5 psig or smaller graduations (used 2.5 psig for calculation), Type B evaluation
Uncertainty Intermediate Calculations
Standard Meniscus |  |  | Neck (mL) |  | u(mS1) = | 0 |  | Sight Glass (mL) |  | u(mS1) = | 0
[5 empty] | u(mS2) = | 0 |  |  |  | u(mS2) = | 0
[5 empty] | u(mS3) = | 0 |  |  |  | u(mS3) = | 0
[5 empty] | u(mS4) = | 0 |  |  |  | u(mS4) = | 0
Unknown Meniscus |  |  | Top Neck (mL) |  | u(mU1) = | 0 |  | Bottom Neck (mL) |  | u(mU1) = | 0
Standard Graduated Neck |  |  | u(nS1) = | 0
|  |  | u(nS2) = | 0
|  |  | u(nS3) = | 0
|  |  | u(nS4) = | 0
Standard Coefficient of Expansion |  |  | Drop # 1 |  | Volume [α + u(CCEα)] = | 0 |  |  | Delivered Volume [base value] =
|  |  | Drop # 2 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 3 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 4 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 5 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 6 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 7 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 8 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 9 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 10 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 11 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 12 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 13 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 14 |  | Volume [α + u(CCEα)] = | 0
|  |  | Drop # 15 |  | Volume [α + u(CCEα)] = | 0
Unknown Coefficient of Expansion |  |  | Volume Trial 1 [β + u(CCEβ)] = |  |  | 0 |  |  | Volume Trial 1 [base value] =
Standard Water Temperature |  |  | Drop # 1 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 2 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 3 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 4 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 5 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 6 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 7 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 8 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 9 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 10 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 11 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 12 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 13 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 14 |  | Volume [t + u(txS)] = | 0
|  |  | Drop # 15 |  | Volume [t + u(txS)] = | 0
Unknown Water Temperature |  |  |  |  | Volume Trial 1[t + u(txX)] = | 0
Pressure Gauge |  |  |  |  | Pcorr @ (100psig + 2.5psig) = | 0 |  |  | Base Value Pcorr @ 100psig =
Quantity, Convert, Combined, Expand & Report
Component Description |  | Component Symbol | Estimated Uncertainty | Units | Reported Unc (L) | Type | Distribution | Factor to Normalize | Standard Unc (L) | % Contribution | Degrees of Freedom
Process Standard Deviation |  | s(p) | 0 | L | 0 | A | Normal | 1 | 0 |  | 1
Volume Standards |  | u(s) | 0 | L | 0 | B | Normal | 1 | 0 |  | 1
Standard Meniscus |  | u(mS) | 0 | L | 0 | B | Triangular | 2.4495 | 0 |  | 1
Unknown Meniscus |  | u(mU) | 0 | L | 0 | B | Triangular | 2.4495 | 0 |  | 1
Std Graduated Neck |  | u(nS) | 0 | L | 0 | B | Triangular | 2.4495 | 0 |  | 1
Ukn Graduated Neck |  | u(nU) | 0 | L | 0 | B | Triangular | 2.4495 | 0 |  | 1
Std Coefficient of Expansion |  | u(CCEα) | 0 | L | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Ukn Coefficient of Expansion |  | u(CCEβ) | 0 | / ºF | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Std Water Temperature |  | u(twS) | 0 | ºC | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Ukn Water Temperature |  | u(twU) | 0 | ºC | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Pressure Gauge |  | u(psig) | 0 | L | 0 | B | Rectangular | 1.7321 | 0 |  | 1
[9 empty] | Combined U (uc) = |  | L
Effective Degrees of Freedom, Welch-Satterthwaite formula, NIST Technical Note 1297, (B-1) (veff) =
k factor =
[9 empty] | Expanded U (uc*k) = |  | L
[9 empty] | NVLAP Effective Degrees of Freedom (veff) = | 0
[9 empty] | NVLAP k factor = | 0
[9 empty] | NVLAP U = | 0 | L
Final Effective Degrees of Freedom (veff) =
Final k factor =
[9 empty] | Final U (uc*k) = |  | L
Values in L

## Tables & Lists
Tables & Lists
Cubical Coefficient of Expansion Table (CCE_Table) & Material List (Material.List) |  | Pick Lists for Data Validation Cells
Material | cm³/ºF | Scale Graduations (in³)
Steel, Pressure Vessel, Low Carbon | 1.6e-05 | Scale Graduations (gal)
Steel, Prover, Low Carbon | 1.86e-05 | Scale Graduations (L)
Steel, Stainless | 2.65e-05 | Scale Graduations (mL)
Range Mean from CC (in³)
Range Mean from CC (gal)
Range Mean from CC (L)
Range Mean from CC (mL)
Yes
No
Nominal Volume (gal)
Nominal Volume (L)
Designated Reference Temperature For This Calibration (ºF)
Designated Reference Temperature For This Calibration (ºC)