Evaluating the use of UVC light devices in a clinical setting to reduce pathogens on computer workstations

Steve Reinecke, MT(CLS), CPHIMS, VP of Business Development, Proximity, Tomball, Texas, USA

Hospitals are challenged with trying to enforce infection control policies.  Traditional chemical disinfection practices have been used in hospitals for decades. Typical chemical disinfection products rely on a surface to be completely covered, and to remain wet for upwards of 5 minutes, to effectively kill viruses and microorganisms.

UVC is an extremely effective form of disinfection, killing a wide array of harmful organisms including bacteria, viruses, moulds, and spores.  This method was studied as an addition to traditional surface disinfection practices.

Small UVC devices are starting to be made available to disinfect high touch surfaces like touchscreens, kiosks, keyboards, and mice. These devices mount on a screen or over a keyboard and emit a low dose UVC light. These devices run routinely (5 minutes / hour) and have a motion sensor that will turn off the light when a user approaches, and will turn on again when there is no activity.  Testing in laboratory settings have seen a 99.996% reduction in C. diff after 30 minutes of UVC exposure and a 99.998% reduction in MRSA after 15 minutes of UVC exposure.

In March of 2019, HCA Houston South East, in Houston, Texas, installed 53 UVC devices on computer workstations throughout their MICU.  These devices were installed on 16 mobile carts and 37 in room computer workstations.   Prior to installation of the UVC devices, 53 keyboards were swabbed and cultured revealing 71% of the keyboards having bacteria (Staphylococcus, Bacillus, Diphtheroid and Enterococcus).  One keyboard contained a Vancomycin-resistant Enterococci also referred to as a “super-bug”.  The UVC devices ran for 2 weeks, and all subsequent swabbing turned up negative for any organisms.

Continuing advancements in automation in all industries have led to increased productivity, efficiency, reliability, and confidence in effectiveness.  Automation in infection control, specifically in using UVC automated disinfection devices, is an enhancement to traditional infection control practices that require human intervention, accuracy, and reliability.  Ultimately this automation has the ability to significantly reduce HAIs.