## Cone Summary
Date: | 2019-03-29
Time: | 13:39
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B3S-Test3
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0407 |  | Start Test | 0 | 2019-03-29 13:53
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 5 | 2019-03-29 13:53
Sample Area | 1 | m^2 | Wetting | 264 | 2019-03-29 13:57
Cone Heat Flux | 75 | kW/m^2 | Dripping | 321 | 2019-03-29 13:58
Cone Orientation |  |  | Bottom Ignition | 612 | 2019-03-29 14:03
Sample Grid |  |  | Flame Out | 903 | 2019-03-29 14:08
Sample Thickness | 106 | mm
Sample Initial Mass | 51.3 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 23.7596 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 4.5956 | kW/m2 | Average Heat Release Rate | 0.98379 | kW/m2
Total Heat Release | 0.98629 | MJ/m2 | Average Mass Loss Rate | 0.045459 | g/m2s
Peak Mass Loss Rate | 0.19734 | g/m2s | Effective Heat of Combustion | 21.76 | MJ/kg
Total Mass Lost | 43.459 | g | Average Specific Extinction Area | 80.8418 | m2/kg
Time to Ignition | 5 | s | Peak CO Yield | 0.66192 | kg/kg
Time to Peak Heat Release | 645 | s | Peak CO2 Yield | 25.7061 | kg/kg
Time to Flame Out | 903 | s | Average CO Yield | 0.040272 | kg/kg
Average Heat Release Rate | 1.0971 | kW/m2 | Average CO2 Yield | 2.3458 | kg/kg
Average Mass Loss Rate | 0.048039 | g/m2s
Effective Heat of Combustion | 22.6947 | MJ/kg
Dripping | 321 | s | Time to dripping
HRRD | 0.43329 | kW | HRR at dripping time
THRD | 0.34559 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.0936 | kW | AHRR from Ignition to Dripping
Wetting | 264 | s | Time to wetting
HRRW | 0.48533 | kW | HRR at wetting time
THRW | 0.31935 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.2147 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-29 13:39 | 1553881160.5237 | -827 | 0.2085 | 0.00047 | 2e-06 | 101.0257 | 100.9666 | 125.9084 | 21.0164 | 0.001903 | 0.001758 | -151.87 | 0.23302 | 8.4027 | 0.002746 | 0.004994 | 0.004826 | 5.7e-05 | 0.001104 | 74.8637 | 1.3415 | 2.5471 | 100.22 | 20.73 | 37.395 | 67.3676 | 87.6478 | 63.9015 | 0.009137 | 1.0475 | 0.024876 | 23.7487 | 0.005232 | 0.13862 | -1.9e-05 | -0.001503 | -0.001503 | -3.7528
2019-03-29 13:39 | 1553881161.5239 | -826 | 0.2085 | 0.00047 | 2e-06 | 101.0257 | 100.9666 | 125.9084 | 21.0164 | 0.001903 | 0.001758 | -151.87 | 0.23302 | 8.4027 | 0.002746 | 0.004994 | 0.004826 | 5.7e-05 | 0.001104 | 74.8637 | 1.3415 | 2.5471 | 100.22 | 20.73 | 37.394 | 67.3676 | 87.6478 | 63.9015 | 0.009137 | 1.0475 | 0.024876 | 23.7487 | 0.005232 | 0.13862 | -1.9e-05 | -0.001504 | -0.001504 | -3.752
2019-03-29 13:39 | 1553881162.5236 | -825 | 0.20853 | 0.000474 | 2e-06 | 101.0455 | 101.0199 | 125.9084 | 21.0172 | 0.001931 | 0.001791 | -151.87 | 0.23181 | 8.4039 | 0.002771 | 0.004615 | 0.004497 | 5.9e-05 | 0.001108 | 74.8692 | 1.3878 | 2.5472 | 100.22 | 20.73 | 37.393 | 68.0446 | 88.2349 | 64.6877 | 0.009136 | 1.045 | 0.024847 | 23.7764 | 0.002266 | 0.13791 | -2.3e-05 | -0.001875 | -0.001875 | -2.994
2019-03-29 13:39 | 1553881163.5228 | -824 | 0.2085 | 0.000472 | 2e-06 | 101.1221 | 101.0397 | 125.9084 | 21.0174 | 0.001876 | 0.00175 | -151.87 | 0.23305 | 8.4023 | 0.002786 | 0.003998 | 0.003815 | 5.7e-05 | 0.001112 | 74.898 | 1.3794 | 2.5473 | 100.22 | 20.74 | 37.391 | 66.732 | 88.6068 | 63.6877 | 0.009142 | 1.0481 | 0.024883 | 23.7411 | 0.007294 | 0.13864 | -2.2e-05 | -0.001746 | -0.001746 | -3.2316
2019-03-29 13:39 | 1553881164.5226 | -823 | 0.2085 | 0.000473 | 2e-06 | 101.1377 | 101.0144 | 125.9084 | 21.0182 | 0.0019 | 0.001772 | -151.87 | 0.23219 | 8.4038 | 0.00274 | 0.004577 | 0.00437 | 5.7e-05 | 0.001111 | 74.8552 | 1.3419 | 2.5471 | 100.22 | 20.74 | 37.388 | 67.2983 | 87.5058 | 64.2312 | 0.009141 | 1.0464 | 0.024863 | 23.7603 | 0.010907 | 0.13813 | -2.3e-05 | -0.001819 | -0.001819 | -3.0903
2019-03-29 13:39 | 1553881165.5237 | -822 | 0.20849 | 0.00047 | 1e-06 | 101.1625 | 101.0567 | 125.9084 | 21.0184 | 0.001864 | 0.001797 | -151.87 | 0.23219 | 8.4023 | 0.002738 | 0.003906 | 0.003771 | 5.9e-05 | 0.001103 | 74.8675 | 1.3502 | 2.5471 | 100.22 | 20.73 | 37.38 | 66.4458 | 87.4504 | 64.8293 | 0.009133 | 1.0446 | 0.024841 | 23.7813 | 0.009357 | 0.13813 | -1.8e-05 | -0.001426 | -0.001426 | -3.9437
2019-03-29 13:39 | 1553881166.5237 | -821 | 0.20852 | 0.000471 | 2e-06 | 101.229 | 101.168 | 125.9084 | 21.0184 | 0.001886 | 0.001752 | -151.87 | 0.23314 | 8.4015 | 0.00278 | 0.003972 | 0.003711 | 5.9e-05 | 0.001099 | 74.8589 | 1.3657 | 2.5472 | 100.22 | 20.73 | 37.38 | 66.9715 | 88.4713 | 63.7552 | 0.009133 | 1.0479 | 0.024881 | 23.7435 | 0.005389 | 0.13869 | -2e-05 | -0.001622 | -0.001622 | -3.4792
2019-03-29 13:39 | 1553881167.5227 | -820 | 0.20856 | 0.000468 | 2e-06 | 101.2775 | 101.1912 | 125.9084 | 21.0185 | 0.001909 | 0.001744 | -151.87 | 0.23267 | 8.4014 | 0.002825 | 0.003776 | 0.003647 | 5.7e-05 | 0.001095 | 74.9142 | 1.3266 | 2.5471 | 100.23 | 20.73 | 37.369 | 67.5326 | 89.5329 | 63.5655 | 0.00913 | 1.0485 | 0.024888 | 23.7368 | 0.007621 | 0.13842 | -1.6e-05 | -0.001317 | -0.001317 | -4.2782
2019-03-29 13:39 | 1553881168.5228 | -819 | 0.20853 | 0.000475 | 2e-06 | 101.2855 | 101.1584 | 125.9084 | 21.0185 | 0.001906 | 0.001746 | -151.87 | 0.23131 | 8.403 | 0.002818 | 0.005134 | 0.004974 | 6.1e-05 | 0.001095 | 74.8847 | 1.3141 | 2.5472 | 100.23 | 20.73 | 37.369 | 67.4446 | 89.3861 | 63.596 | 0.00913 | 1.0484 | 0.024887 | 23.7379 | 0.011231 | 0.13761 | -2.4e-05 | -0.001909 | -0.001909 | -2.934
2019-03-29 13:39 | 1553881169.524 | -818 | 0.20852 | 0.000471 | 2e-06 | 101.3843 | 101.4291 | 125.9084 | 21.0197 | 0.001925 | 0.001778 | -151.87 | 0.23258 | 8.4021 | 0.00276 | 0.003947 | 0.003765 | 6e-05 | 0.001086 | 74.8688 | 1.3399 | 2.5472 | 100.23 | 20.73 | 37.364 | 67.9035 | 87.9843 | 64.3857 | 0.009128 | 1.0459 | 0.024858 | 23.7657 | -0.003948 | 0.13836 | -2e-05 | -0.001611 | -0.001611 | -3.4963
2019-03-29 13:39 | 1553881170.5238 | -817 | 0.20855 | 0.000471 | 2e-06 | 101.3595 | 101.3523 | 125.9084 | 21.0195 | 0.001887 | 0.001745 | -151.87 | 0.23306 | 8.4009 | 0.00278 | 0.004473 | 0.004269 | 6.3e-05 | 0.001088 | 74.849 | 1.3397 | 2.5473 | 100.23 | 20.73 | 37.361 | 67.0024 | 88.4518 | 63.5899 | 0.009128 | 1.0484 | 0.024887 | 23.7377 | 0.000634 | 0.13865 | -2e-05 | -0.001647 | -0.001647 | -3.4267
2019-03-29 13:40 | 1553881171.5229 | -816 | 0.20855 | 0.000472 | 2e-06 | 101.3732 | 101.4086 | 126.2065 | 21.0194 | 0.001917 | 0.001791 | -151.8737 | 0.23286 | 8.4023 | 0.002823 | 0.003915 | 0.003776 | 6e-05 | 0.001085 | 74.8796 | 1.3826 | 2.5473 | 100.23 | 20.74 | 37.353 | 67.7122 | 89.5033 | 64.6899 | 0.009131 | 1.045 | 0.024876 | 23.8046 | -0.00312 | 0.13853 | -2.2e-05 | -0.001755 | -0.001755 | -3.2133
2019-03-29 13:40 | 1553881172.5229 | -815 | 0.20854 | 0.000472 | 2e-06 | 101.4939 | 101.4359 | 126.0356 | 21.0193 | 0.001904 | 0.001788 | -151.8732 | 0.23394 | 8.4028 | 0.002826 | 0.004311 | 0.004103 | 5.9e-05 | 0.001087 | 74.7518 | 1.3661 | 2.5473 | 100.23 | 20.74 | 37.353 | 67.3977 | 89.5713 | 64.6141 | 0.009131 | 1.0452 | 0.024862 | 23.7858 | 0.00511 | 0.13917 | -2.1e-05 | -0.001659 | -0.001659 | -3.4142
2019-03-29 13:40 | 1553881173.522 | -814 | 0.20856 | 0.000466 | 2e-06 | 101.4355 | 101.4885 | 126.0908 | 21.0198 | 0.001905 | 0.001762 | -151.8657 | 0.23347 | 8.4013 | 0.002745 | 0.003548 | 0.003367 | 6.1e-05 | 0.001088 | 74.8966 | 1.3472 | 2.5471 | 100.23 | 20.74 | 37.35 | 67.4396 | 87.6149 | 63.9976 | 0.009131 | 1.0472 | 0.02489 | 23.7693 | -0.004669 | 0.13889 | -1.3e-05 | -0.001087 | -0.001087 | -5.1996
2019-03-29 13:40 | 1553881174.5231 | -813 | 0.20855 | 0.00047 | 2e-06 | 101.4681 | 101.4206 | 126.452 | 21.0199 | 0.001911 | 0.001808 | -151.871 | 0.23328 | 8.4015 | 0.002789 | 0.004574 | 0.004379 | 6.3e-05 | 0.001092 | 74.9084 | 1.3632 | 2.5472 | 100.23 | 20.74 | 37.347 | 67.5816 | 88.6764 | 65.09 | 0.00913 | 1.0438 | 0.024885 | 23.8418 | 0.004195 | 0.13878 | -1.9e-05 | -0.00152 | -0.00152 | -3.7154
2019-03-29 13:40 | 1553881175.522 | -812 | 0.20854 | 0.000473 | 2e-06 | 101.5364 | 101.5294 | 126.3356 | 21.0198 | 0.001898 | 0.001774 | -151.8564 | 0.23252 | 8.4021 | 0.002712 | 0.00458 | 0.004378 | 6.1e-05 | 0.001094 | 74.9689 | 1.3403 | 2.5473 | 100.23 | 20.74 | 37.339 | 67.26 | 86.8148 | 64.2765 | 0.009128 | 1.0463 | 0.024904 | 23.8022 | 0.000613 | 0.13833 | -2.3e-05 | -0.001796 | -0.001796 | -3.1354
2019-03-29 13:40 | 1553881176.5231 | -811 | 0.20853 | 0.000467 | 2e-06 | 101.5555 | 101.5192 | 126.0483 | 21.0203 | 0.001852 | 0.001751 | -151.8707 | 0.23286 | 8.4002 | 0.002647 | 0.003975 | 0.003786 | 6e-05 | 0.001093 | 74.8973 | 1.3632 | 2.5473 | 100.23 | 20.74 | 37.339 | 66.1574 | 85.2549 | 63.7243 | 0.009128 | 1.048 | 0.024896 | 23.7556 | 0.003202 | 0.13853 | -1.5e-05 | -0.001211 | -0.001211 | -4.6575
2019-03-29 13:40 | 1553881177.5221 | -810 | 0.20853 | 0.000473 | 2e-06 | 101.5882 | 101.5035 | 126.019 | 21.0206 | 0.001914 | 0.001807 | -151.8421 | 0.23295 | 8.4022 | 0.002693 | 0.004115 | 0.003944 | 6.2e-05 | 0.001102 | 74.8973 | 1.3796 | 2.5472 | 100.23 | 20.74 | 37.333 | 67.6393 | 86.3596 | 65.0837 | 0.009127 | 1.0438 | 0.024843 | 23.8007 | 0.007459 | 0.13858 | -2.2e-05 | -0.001752 | -0.001752 | -3.2198
2019-03-29 13:40 | 1553881178.5221 | -809 | 0.20853 | 0.000467 | 1e-06 | 101.6373 | 101.4888 | 126.0251 | 21.0202 | 0.001907 | 0.001735 | -151.8338 | 0.23332 | 8.4023 | 0.002621 | 0.004393 | 0.004277 | 5.8e-05 | 0.001103 | 74.8591 | 1.3141 | 2.5473 | 100.23 | 20.74 | 37.331 | 67.4835 | 84.6265 | 63.332 | 0.009126 | 1.0492 | 0.024908 | 23.7396 | 0.013073 | 0.13881 | -1.5e-05 | -0.001182 | -0.001182 | -4.7812
2019-03-29 13:40 | 1553881179.5231 | -808 | 0.20854 | 0.000478 | 2e-06 | 101.5693 | 101.4758 | 126.3441 | 21.0219 | 0.001863 | 0.001778 | -151.8358 | 0.23349 | 8.4014 | 0.002677 | 0.004529 | 0.004309 | 6.1e-05 | 0.00111 | 74.8426 | 1.3168 | 2.5472 | 100.23 | 20.74 | 37.328 | 66.4141 | 85.9916 | 64.3792 | 0.009125 | 1.046 | 0.024901 | 23.8066 | 0.008235 | 0.1389 | -2.9e-05 | -0.002281 | -0.002281 | -2.4782
2019-03-29 13:40 | 1553881180.5221 | -807 | 0.20849 | 0.000476 | 2e-06 | 101.6642 | 101.6124 | 126.5534 | 21.0211 | 0.001856 | 0.001758 | -151.8346 | 0.23321 | 8.4017 | 0.002634 | 0.003617 | 0.003547 | 6e-05 | 0.001111 | 74.9335 | 1.3135 | 2.5473 | 100.23 | 20.74 | 37.324 | 66.2462 | 84.9427 | 63.9038 | 0.009124 | 1.0474 | 0.024939 | 23.8095 | 0.004562 | 0.13873 | -2.6e-05 | -0.002058 | -0.002058 | -2.7434
2019-03-29 13:40 | 1553881181.5222 | -806 | 0.20854 | 0.000472 | 2e-06 | 101.6094 | 101.6083 | 126.5628 | 21.0212 | 0.001892 | 0.001697 | -151.8267 | 0.2337 | 8.4028 | 0.002651 | 0.003926 | 0.003686 | 5.9e-05 | 0.001112 | 74.9515 | 1.3147 | 2.5472 | 100.22 | 20.74 | 37.323 | 67.1272 | 85.3556 | 62.4206 | 0.009125 | 1.0521 | 0.024995 | 23.7579 | 0.000102 | 0.13903 | -2.1e-05 | -0.001707 | -0.001707 | -3.3149
2019-03-29 13:40 | 1553881182.5226 | -805 | 0.20852 | 0.000474 | 2e-06 | 101.675 | 101.6035 | 126.4948 | 21.0221 | 0.001868 | 0.001731 | -151.8264 | 0.23288 | 8.4 | 0.002611 | 0.003749 | 0.003611 | 5.7e-05 | 0.001113 | 74.9537 | 1.2827 | 2.5472 | 100.22 | 20.74 | 37.314 | 66.5391 | 84.3967 | 63.2566 | 0.009123 | 1.0495 | 0.024957 | 23.7811 | 0.006291 | 0.13854 | -2.3e-05 | -0.001872 | -0.001872 | -3.0124
2019-03-29 13:40 | 1553881183.5233 | -804 | 0.20855 | 0.000478 | 2e-06 | 101.6561 | 101.7332 | 126.3194 | 21.0218 | 0.001909 | 0.001784 | -151.8311 | 0.23365 | 8.4019 | 0.002594 | 0.003969 | 0.0038 | 6e-05 | 0.001112 | 74.9156 | 1.3067 | 2.5471 | 100.22 | 20.74 | 37.314 | 67.5326 | 83.9999 | 64.5314 | 0.009123 | 1.0455 | 0.024893 | 23.8096 | -0.006778 | 0.139 | -2.9e-05 | -0.002278 | -0.002278 | -2.4831
2019-03-29 13:40 | 1553881184.5223 | -803 | 0.20855 | 0.000475 | 2e-06 | 101.6935 | 101.6018 | 126.2833 | 21.023 | 0.001923 | 0.00177 | -151.8114 | 0.23273 | 8.4026 | 0.002661 | 0.003685 | 0.003535 | 6.2e-05 | 0.001107 | 75.0003 | 1.3441 | 2.5471 | 100.22 | 20.74 | 37.307 | 67.8685 | 85.6118 | 64.1813 | 0.009121 | 1.0466 | 0.024902 | 23.7939 | 0.008074 | 0.13845 | -2.5e-05 | -0.001971 | -0.001971 | -2.859
2019-03-29 13:40 | 1553881185.5224 | -802 | 0.20857 | 0.000476 | 2e-06 | 101.7178 | 101.6298 | 126.6736 | 21.0219 | 0.001875 | 0.001785 | -151.8099 | 0.23298 | 8.4019 | 0.002655 | 0.003783 | 0.003595 | 6.2e-05 | 0.001107 | 74.9115 | 1.3841 | 2.5474 | 100.22 | 20.74 | 37.307 | 66.7191 | 85.4609 | 64.5496 | 0.009121 | 1.0454 | 0.024927 | 23.8436 | 0.007743 | 0.1386 | -2.5e-05 | -0.002007 | -0.002007 | -2.8101
2019-03-29 13:40 | 1553881186.5224 | -801 | 0.20855 | 0.000476 | 2e-06 | 101.6729 | 101.7127 | 127.2332 | 21.0221 | 0.001902 | 0.001785 | -151.7927 | 0.23284 | 8.4036 | 0.002659 | 0.004204 | 0.004073 | 6.2e-05 | 0.0011 | 74.8142 | 1.3807 | 2.5473 | 100.22 | 20.74 | 37.302 | 67.3498 | 85.5516 | 64.5568 | 0.00912 | 1.0454 | 0.024982 | 23.8965 | -0.003503 | 0.13851 | -2.6e-05 | -0.002015 | -0.002015 | -2.7978
2019-03-29 13:40 | 1553881187.5224 | -800 | 0.20855 | 0.000477 | 2e-06 | 101.7124 | 101.582 | 127.5977 | 21.0231 | 0.001885 | 0.00177 | -151.8221 | 0.23294 | 8.4042 | 0.002661 | 0.003949 | 0.003817 | 5.8e-05 | 0.001107 | 74.8577 | 1.3717 | 2.5473 | 100.22 | 20.74 | 37.302 | 66.9419 | 85.5924 | 64.1962 | 0.00912 | 1.0465 | 0.025031 | 23.9179 | 0.011475 | 0.13858 | -2.8e-05 | -0.002251 | -0.002251 | -2.506
2019-03-29 13:40 | 1553881188.5225 | -799 | 0.20855 | 0.000476 | 2e-06 | 101.7268 | 101.7294 | 128.0648 | 21.0229 | 0.001869 | 0.00174 | -151.822 | 0.23256 | 8.4051 | 0.002696 | 0.003496 | 0.003226 | 6e-05 | 0.001117 | 74.9117 | 1.3649 | 2.5473 | 100.22 | 20.74 | 37.296 | 66.5663 | 86.4554 | 63.4581 | 0.009118 | 1.0488 | 0.025104 | 23.9354 | -0.000234 | 0.13835 | -2.6e-05 | -0.002084 | -0.002084 | -2.7026
2019-03-29 13:40 | 1553881189.5215 | -798 | 0.20855 | 0.000479 | 2e-06 | 101.7082 | 101.7533 | 127.9313 | 21.0238 | 0.001878 | 0.001749 | -151.8289 | 0.23276 | 8.399 | 0.002579 | 0.004243 | 0.004016 | 5.8e-05 | 0.001114 | 74.8174 | 1.3821 | 2.5473 | 100.22 | 20.74 | 37.296 | 66.795 | 83.6424 | 63.6863 | 0.009118 | 1.0481 | 0.025083 | 23.931 | -0.003967 | 0.13847 | -3e-05 | -0.002337 | -0.002337 | -2.411
2019-03-29 13:40 | 1553881190.5216 | -797 | 0.20856 | 0.000476 | 2e-06 | 101.6148 | 101.6963 | 127.8174 | 21.023 | 0.001874 | 0.001758 | -151.8335 | 0.23338 | 8.402 | 0.002678 | 0.004444 | 0.004291 | 5.9e-05 | 0.001113 | 74.9034 | 1.4079 | 2.5475 | 100.22 | 20.75 | 37.292 | 66.6944 | 86.0005 | 63.9064 | 0.009123 | 1.0474 | 0.025063 | 23.9282 | -0.007172 | 0.13884 | -2.6e-05 | -0.002087 | -0.002087 | -2.7081
2019-03-29 13:40 | 1553881191.5226 | -796 | 0.20854 | 0.000478 | 2e-06 | 101.5924 | 101.4762 | 127.2853 | 21.0235 | 0.001929 | 0.001786 | -151.8434 | 0.2323 | 8.403 | 0.002696 | 0.004112 | 0.003928 | 6e-05 | 0.001105 | 74.9361 | 1.3375 | 2.5473 | 100.22 | 20.75 | 37.292 | 68.0195 | 86.456 | 64.5682 | 0.009123 | 1.0454 | 0.024987 | 23.9018 | 0.010234 | 0.1382 | -2.8e-05 | -0.002202 | -0.002202 | -2.5544
2019-03-29 13:40 | 1553881192.5225 | -795 | 0.20856 | 0.000472 | 2e-06 | 101.616 | 101.8076 | 126.9813 | 21.0231 | 0.00192 | 0.00173 | -151.8325 | 0.23321 | 8.4017 | 0.002618 | 0.003608 | 0.00343 | 6e-05 | 0.0011 | 74.9048 | 1.3105 | 2.5474 | 100.22 | 20.75 | 37.289 | 67.7821 | 84.5623 | 63.2124 | 0.009122 | 1.0496 | 0.025007 | 23.8253 | -0.016847 | 0.13874 | -2.1e-05 | -0.001732 | -0.001732 | -3.2593
2019-03-29 13:40 | 1553881193.5215 | -794 | 0.20852 | 0.000475 | 2e-06 | 101.5654 | 101.6014 | 126.672 | 21.0239 | 0.001879 | 0.001761 | -151.8407 | 0.23319 | 8.4013 | 0.002613 | 0.003868 | 0.003745 | 5.9e-05 | 0.0011 | 74.8512 | 1.3348 | 2.5473 | 100.22 | 20.75 | 37.287 | 66.81 | 84.45 | 63.9675 | 0.009122 | 1.0472 | 0.024948 | 23.8229 | -0.003172 | 0.13873 | -2.5e-05 | -0.002046 | -0.002046 | -2.7596
2019-03-29 13:40 | 1553881194.5218 | -793 | 0.20854 | 0.000473 | 2e-06 | 101.5648 | 101.6356 | 126.3903 | 21.0244 | 0.001927 | 0.001764 | -151.8461 | 0.23396 | 8.4032 | 0.002556 | 0.003403 | 0.003199 | 6.1e-05 | 0.001098 | 74.9183 | 1.3203 | 2.5473 | 100.22 | 20.75 | 37.286 | 67.9706 | 83.072 | 64.0515 | 0.009122 | 1.047 | 0.024918 | 23.7994 | -0.006228 | 0.13918 | -2.3e-05 | -0.001845 | -0.001845 | -3.0708
2019-03-29 13:40 | 1553881195.5217 | -792 | 0.20854 | 0.000476 | 2e-06 | 101.5894 | 101.4724 | 125.7706 | 21.0236 | 0.001865 | 0.001701 | -151.8481 | 0.23331 | 8.4015 | 0.002603 | 0.003935 | 0.00379 | 5.9e-05 | 0.001102 | 74.8955 | 1.3787 | 2.5474 | 100.22 | 20.75 | 37.282 | 66.4742 | 84.2008 | 62.5221 | 0.009121 | 1.0518 | 0.024913 | 23.6871 | 0.010301 | 0.1388 | -2.6e-05 | -0.002028 | -0.002028 | -2.785
2019-03-29 13:40 | 1553881196.5226 | -791 | 0.20853 | 0.000473 | 1e-06 | 101.5091 | 101.5786 | 125.3951 | 21.0228 | 0.001877 | 0.00173 | -151.8497 | 0.23321 | 8.4024 | 0.002662 | 0.004589 | 0.004309 | 6.2e-05 | 0.001096 | 74.9054 | 1.3521 | 2.5474 | 100.22 | 20.75 | 37.281 | 66.7543 | 85.6328 | 63.2111 | 0.00912 | 1.0496 | 0.02485 | 23.6759 | -0.006116 | 0.13873 | -2.2e-05 | -0.001696 | -0.001696 | -3.3293
2019-03-29 13:40 | 1553881197.5229 | -790 | 0.20854 | 0.000476 | 2e-06 | 101.4511 | 101.4417 | 125.2275 | 21.024 | 0.00193 | 0.001756 | -151.8355 | 0.23332 | 8.4014 | 0.002635 | 0.003485 | 0.003342 | 6.3e-05 | 0.001101 | 74.9633 | 1.3513 | 2.5474 | 100.22 | 20.75 | 37.279 | 68.0248 | 84.9852 | 63.8457 | 0.00912 | 1.0476 | 0.02481 | 23.6824 | 0.000825 | 0.1388 | -2.6e-05 | -0.002056 | -0.002056 | -2.7473
2019-03-29 13:40 | 1553881198.5227 | -789 | 0.20858 | 0.000477 | 2e-06 | 101.4876 | 101.384 | 125.1453 | 21.0243 | 0.001891 | 0.001736 | -151.8383 | 0.2325 | 8.4022 | 0.002734 | 0.003891 | 0.003673 | 5.9e-05 | 0.001097 | 74.9075 | 1.3737 | 2.5473 | 100.22 | 20.75 | 37.276 | 67.1076 | 87.3608 | 63.3705 | 0.009119 | 1.0491 | 0.02482 | 23.6579 | 0.009131 | 0.13832 | -2.7e-05 | -0.002085 | -0.002085 | -2.7001
2019-03-29 13:40 | 1553881199.5217 | -788 | 0.20856 | 0.000474 | 2e-06 | 101.3903 | 101.3199 | 125.2134 | 21.025 | 0.001956 | 0.001818 | -151.8261 | 0.23258 | 8.4031 | 0.002791 | 0.004343 | 0.004102 | 6e-05 | 0.001087 | 74.9124 | 1.3778 | 2.5472 | 100.22 | 20.76 | 37.275 | 68.6676 | 88.7268 | 65.3486 | 0.009125 | 1.043 | 0.024754 | 23.7338 | 0.006216 | 0.13836 | -2.4e-05 | -0.001952 | -0.001952 | -2.8848
2019-03-29 13:40 | 1553881200.5219 | -787 | 0.20854 | 0.000472 | 2e-06 | 101.3176 | 101.1878 | 125.2768 | 21.0246 | 0.001893 | 0.001791 | -151.8311 | 0.23394 | 8.4012 | 0.00274 | 0.004051 | 0.003936 | 6e-05 | 0.001083 | 74.9271 | 1.3959 | 2.5474 | 100.22 | 20.76 | 37.272 | 67.1517 | 87.5111 | 64.6817 | 0.009124 | 1.045 | 0.024784 | 23.7164 | 0.011464 | 0.13917 | -2.2e-05 | -0.001758 | -0.001758 | -3.221
2019-03-29 13:40 | 1553881201.5219 | -786 | 0.20853 | 0.000477 | 2e-06 | 101.2933 | 101.1451 | 125.5404 | 21.025 | 0.001867 | 0.001767 | -151.8364 | 0.2328 | 8.402 | 0.00271 | 0.004048 | 0.003801 | 6.3e-05 | 0.001083 | 74.9068 | 1.381 | 2.5474 | 100.22 | 20.76 | 37.268 | 66.5263 | 86.7753 | 64.1081 | 0.009123 | 1.0468 | 0.024832 | 23.7212 | 0.013095 | 0.13849 | -2.9e-05 | -0.002269 | -0.002269 | -2.484
2019-03-29 13:40 | 1553881202.5219 | -785 | 0.20855 | 0.000475 | 2e-06 | 101.2245 | 101.2178 | 125.7206 | 21.025 | 0.00196 | 0.001768 | -151.8417 | 0.23258 | 8.402 | 0.002724 | 0.004538 | 0.004415 | 6.1e-05 | 0.001081 | 74.9102 | 1.334 | 2.5473 | 100.22 | 20.76 | 37.268 | 68.758 | 87.1141 | 64.1317 | 0.009123 | 1.0467 | 0.024849 | 23.7391 | 0.000588 | 0.13836 | -2.5e-05 | -0.001958 | -0.001958 | -2.8758
2019-03-29 13:40 | 1553881203.5229 | -784 | 0.20858 | 0.000473 | 2e-06 | 101.1089 | 101.1182 | 125.8319 | 21.0258 | 0.001916 | 0.001789 | -151.8437 | 0.23278 | 8.4015 | 0.002657 | 0.004107 | 0.003903 | 6.6e-05 | 0.001088 | 74.9504 | 1.3272 | 2.5473 | 100.22 | 20.76 | 37.26 | 67.6888 | 85.5138 | 64.6383 | 0.009121 | 1.0452 | 0.024841 | 23.7674 | -0.000817 | 0.13848 | -2.3e-05 | -0.001818 | -0.001818 | -3.1009
2019-03-29 13:40 | 1553881204.5209 | -783 | 0.20852 | 0.000476 | 2e-06 | 101.0694 | 100.839 | 125.6608 | 21.0263 | 0.001952 | 0.001802 | -151.848 | 0.2328 | 8.4044 | 0.002621 | 0.003667 | 0.003542 | 6.2e-05 | 0.0011 | 74.8903 | 1.3396 | 2.5473 | 100.22 | 20.76 | 37.26 | 68.5745 | 84.6481 | 64.9699 | 0.009121 | 1.0441 | 0.024812 | 23.7629 | 0.020416 | 0.1385 | -2.7e-05 | -0.002139 | -0.002139 | -2.6352
2019-03-29 13:40 | 1553881205.5221 | -782 | 0.20855 | 0.000475 | 2e-06 | 100.9617 | 100.7455 | 125.7594 | 21.0256 | 0.001925 | 0.001773 | -151.865 | 0.23271 | 8.4021 | 0.002665 | 0.003107 | 0.002821 | 6.2e-05 | 0.001111 | 74.8594 | 1.3298 | 2.5473 | 100.22 | 20.76 | 37.251 | 67.9124 | 85.7017 | 64.2673 | 0.009119 | 1.0463 | 0.024847 | 23.7475 | 0.019173 | 0.13844 | -2.5e-05 | -0.00195 | -0.00195 | -2.889
2019-03-29 13:40 | 1553881206.5221 | -781 | 0.20855 | 0.000472 | 2e-06 | 100.8746 | 100.7851 | 125.7392 | 21.0266 | 0.001977 | 0.0018 | -151.8579 | 0.23295 | 8.402 | 0.002719 | 0.003555 | 0.003517 | 5.9e-05 | 0.001115 | 74.945 | 1.3558 | 2.5473 | 100.22 | 20.76 | 37.251 | 69.1598 | 87.004 | 64.9218 | 0.009119 | 1.0443 | 0.024821 | 23.7686 | 0.007938 | 0.13858 | -2.2e-05 | -0.001764 | -0.001764 | -3.1978
2019-03-29 13:40 | 1553881207.521 | -780 | 0.20853 | 0.000477 | 2e-06 | 100.8087 | 100.7608 | 125.6843 | 21.0263 | 0.001861 | 0.001815 | -151.8537 | 0.23397 | 8.4041 | 0.002609 | 0.004208 | 0.003908 | 6.1e-05 | 0.001118 | 75.0344 | 1.3696 | 2.5474 | 100.22 | 20.76 | 37.246 | 66.3729 | 84.3658 | 65.279 | 0.009117 | 1.0432 | 0.024803 | 23.776 | 0.004248 | 0.13919 | -2.7e-05 | -0.002153 | -0.002153 | -2.6316
2019-03-29 13:40 | 1553881208.5211 | -779 | 0.20854 | 0.000475 | 2e-06 | 100.8072 | 100.8451 | 125.7171 | 21.0271 | 0.001863 | 0.001765 | -151.8313 | 0.23355 | 8.4039 | 0.002672 | 0.003835 | 0.003638 | 6.1e-05 | 0.001118 | 74.9398 | 1.3684 | 2.5475 | 100.22 | 20.76 | 37.237 | 66.4292 | 85.8811 | 64.0572 | 0.009115 | 1.047 | 0.024851 | 23.7361 | -0.003365 | 0.13894 | -2.6e-05 | -0.002021 | -0.002021 | -2.7975
2019-03-29 13:40 | 1553881209.5221 | -778 | 0.20853 | 0.000478 | 2e-06 | 100.736 | 100.5738 | 125.6587 | 21.027 | 0.001905 | 0.001779 | -151.8354 | 0.23328 | 8.4007 | 0.002678 | 0.004241 | 0.004032 | 6.4e-05 | 0.001122 | 74.9423 | 1.3356 | 2.5475 | 100.22 | 20.76 | 37.23 | 67.4372 | 86.0044 | 64.4084 | 0.009114 | 1.0459 | 0.024832 | 23.7429 | 0.014409 | 0.13878 | -2.9e-05 | -0.002282 | -0.002282 | -2.4754
2019-03-29 13:40 | 1553881210.5212 | -777 | 0.20854 | 0.000477 | 2e-06 | 100.6671 | 100.3943 | 125.3066 | 21.0266 | 0.001892 | 0.001716 | -151.8422 | 0.2336 | 8.4 | 0.002671 | 0.004393 | 0.004231 | 6.2e-05 | 0.001123 | 74.9205 | 1.3237 | 2.5472 | 100.22 | 20.76 | 37.23 | 67.1185 | 85.8556 | 62.8877 | 0.009114 | 1.0506 | 0.024853 | 23.6562 | 0.024274 | 0.13897 | -2.7e-05 | -0.002084 | -0.002084 | -2.7137
2019-03-29 13:40 | 1553881211.5211 | -776 | 0.20856 | 0.00048 | 3e-06 | 100.507 | 100.4209 | 124.9378 | 21.028 | 0.001894 | 0.001791 | -151.8297 | 0.23239 | 8.4039 | 0.002634 | 0.003535 | 0.00334 | 5.7e-05 | 0.001125 | 74.9247 | 1.3898 | 2.5474 | 100.22 | 20.76 | 37.22 | 67.1667 | 84.9612 | 64.6873 | 0.009111 | 1.045 | 0.024751 | 23.6845 | 0.007662 | 0.13825 | -3.2e-05 | -0.002504 | -0.002504 | -2.2473
2019-03-29 13:40 | 1553881212.5226 | -775 | 0.20854 | 0.000476 | 2e-06 | 100.4609 | 100.1117 | 124.8454 | 21.0285 | 0.001905 | 0.001769 | -151.838 | 0.23278 | 8.4005 | 0.002624 | 0.003562 | 0.003313 | 6.1e-05 | 0.001117 | 74.896 | 1.378 | 2.5474 | 100.22 | 20.76 | 37.22 | 67.437 | 84.7173 | 64.156 | 0.009111 | 1.0467 | 0.024761 | 23.6571 | 0.031142 | 0.13848 | -2.7e-05 | -0.002123 | -0.002123 | -2.6544
2019-03-29 13:40 | 1553881213.5212 | -774 | 0.20852 | 0.000476 | 2e-06 | 100.366 | 100.2049 | 125.1599 | 21.0284 | 0.00188 | 0.001773 | -151.8235 | 0.23364 | 8.4021 | 0.002634 | 0.004487 | 0.004308 | 6.1e-05 | 0.001115 | 74.9922 | 1.3533 | 2.5473 | 100.22 | 20.76 | 37.213 | 66.8488 | 84.9513 | 64.2717 | 0.009109 | 1.0463 | 0.024788 | 23.691 | 0.01437 | 0.13899 | -2.6e-05 | -0.002056 | -0.002056 | -2.751
2019-03-29 13:40 | 1553881214.5222 | -773 | 0.20852 | 0.000475 | 2e-06 | 100.2178 | 100.0493 | 125.4669 | 21.0286 | 0.001931 | 0.001778 | -151.8202 | 0.2333 | 8.4016 | 0.002635 | 0.004089 | 0.004024 | 6.3e-05 | 0.001112 | 74.9212 | 1.3388 | 2.5473 | 100.22 | 20.76 | 37.213 | 68.0574 | 84.9781 | 64.3905 | 0.009109 | 1.0459 | 0.024814 | 23.7242 | 0.015058 | 0.13879 | -2.4e-05 | -0.00194 | -0.00194 | -2.912
2019-03-29 13:40 | 1553881215.5224 | -772 | 0.20853 | 0.000477 | 1e-06 | 100.2603 | 100.039 | 125.5507 | 21.0285 | 0.001823 | 0.001767 | -151.807 | 0.23368 | 8.4009 | 0.002623 | 0.003651 | 0.003435 | 6.3e-05 | 0.001118 | 74.9373 | 1.3407 | 2.5474 | 100.22 | 20.76 | 37.212 | 65.4652 | 84.6878 | 64.1139 | 0.009109 | 1.0468 | 0.024832 | 23.7224 | 0.019766 | 0.13902 | -2.7e-05 | -0.002083 | -0.002083 | -2.7163
2019-03-29 13:40 | 1553881216.5213 | -771 | 0.20852 | 0.000476 | 2e-06 | 100.1433 | 99.9807 | 125.2878 | 21.0278 | 0.001891 | 0.001769 | -151.8174 | 0.23241 | 8.4031 | 0.002593 | 0.003829 | 0.003643 | 6.1e-05 | 0.001117 | 74.9219 | 1.3075 | 2.5472 | 100.22 | 20.76 | 37.207 | 67.1071 | 83.981 | 64.172 | 0.009108 | 1.0466 | 0.024804 | 23.6996 | 0.014538 | 0.13826 | -2.6e-05 | -0.002055 | -0.002055 | -2.738
2019-03-29 13:40 | 1553881217.5213 | -770 | 0.2085 | 0.000471 | 2e-06 | 100.087 | 99.9131 | 125.2465 | 21.0288 | 0.001884 | 0.0018 | -151.8078 | 0.23275 | 8.4024 | 0.002623 | 0.00387 | 0.003639 | 6.1e-05 | 0.001116 | 74.8833 | 1.2822 | 2.5471 | 100.22 | 20.76 | 37.206 | 66.9279 | 84.6832 | 64.9239 | 0.009108 | 1.0443 | 0.024773 | 23.7221 | 0.015558 | 0.13846 | -2e-05 | -0.001581 | -0.001581 | -3.5652
2019-03-29 13:40 | 1553881218.5214 | -769 | 0.20854 | 0.000475 | 2e-06 | 99.9472 | 99.753 | 124.9416 | 21.0285 | 0.001902 | 0.001792 | -151.8167 | 0.23281 | 8.4023 | 0.002579 | 0.00407 | 0.003992 | 6.1e-05 | 0.001111 | 74.8891 | 1.2859 | 2.5472 | 100.22 | 20.76 | 37.205 | 67.3574 | 83.6306 | 64.7254 | 0.009107 | 1.0449 | 0.02475 | 23.6862 | 0.017397 | 0.1385 | -2.5e-05 | -0.002011 | -0.002011 | -2.8025
2019-03-29 13:40 | 1553881219.5204 | -768 | 0.20855 | 0.000474 | 2e-06 | 99.8694 | 99.6721 | 124.8397 | 21.0285 | 0.001851 | 0.001774 | -151.8106 | 0.23312 | 8.4003 | 0.002602 | 0.003841 | 0.003674 | 6.3e-05 | 0.00111 | 74.8973 | 1.2941 | 2.5473 | 100.22 | 20.76 | 37.205 | 66.1302 | 84.193 | 64.2763 | 0.009107 | 1.0463 | 0.024756 | 23.6608 | 0.017685 | 0.13868 | -2.4e-05 | -0.001893 | -0.001893 | -2.9823
2019-03-29 13:40 | 1553881220.5214 | -767 | 0.20853 | 0.000476 | 2e-06 | 99.7964 | 99.4261 | 125.2697 | 21.0287 | 0.001859 | 0.00177 | -151.801 | 0.23298 | 8.4024 | 0.002627 | 0.003568 | 0.003449 | 6e-05 | 0.001104 | 74.9247 | 1.3467 | 2.5474 | 100.22 | 20.76 | 37.207 | 66.3216 | 84.794 | 64.1906 | 0.009108 | 1.0466 | 0.024802 | 23.6985 | 0.033251 | 0.1386 | -2.7e-05 | -0.002111 | -0.002111 | -2.6722
2019-03-29 13:40 | 1553881221.5206 | -766 | 0.20856 | 0.000478 | 2e-06 | 99.6964 | 99.4431 | 125.5439 | 21.0288 | 0.001867 | 0.001776 | -151.8047 | 0.23276 | 8.403 | 0.002684 | 0.004002 | 0.003768 | 6.3e-05 | 0.0011 | 74.9152 | 1.3359 | 2.5474 | 100.22 | 20.76 | 37.208 | 66.5199 | 86.1498 | 64.3274 | 0.009108 | 1.0461 | 0.024824 | 23.7293 | 0.022751 | 0.13847 | -3e-05 | -0.002342 | -0.002342 | -2.4068
2019-03-29 13:40 | 1553881222.5216 | -765 | 0.20855 | 0.000473 | 2e-06 | 99.615 | 99.4093 | 125.8117 | 21.0303 | 0.001853 | 0.001722 | -151.7877 | 0.23336 | 8.4033 | 0.002581 | 0.004363 | 0.004078 | 6.6e-05 | 0.001096 | 74.9359 | 1.3233 | 2.5473 | 100.22 | 20.76 | 37.208 | 66.1966 | 83.6891 | 63.0475 | 0.009108 | 1.0501 | 0.024898 | 23.7095 | 0.018483 | 0.13882 | -2.2e-05 | -0.001799 | -0.001799 | -3.1399
2019-03-29 13:40 | 1553881223.5217 | -764 | 0.20854 | 0.000475 | 2e-06 | 99.5129 | 99.1752 | 125.6995 | 21.0299 | 0.001911 | 0.001779 | -151.7964 | 0.23301 | 8.4042 | 0.00263 | 0.004372 | 0.004313 | 6.6e-05 | 0.00109 | 74.9066 | 1.3661 | 2.5473 | 100.22 | 20.76 | 37.207 | 67.5819 | 84.8717 | 64.4019 | 0.009108 | 1.0459 | 0.024837 | 23.7466 | 0.030404 | 0.13861 | -2.5e-05 | -0.001987 | -0.001987 | -2.8394
2019-03-29 13:40 | 1553881224.5215 | -763 | 0.20854 | 0.000479 | 2e-06 | 99.4099 | 99.2568 | 125.9833 | 21.0294 | 0.001902 | 0.001766 | -151.8007 | 0.23327 | 8.4009 | 0.002624 | 0.004189 | 0.004017 | 6.7e-05 | 0.001089 | 74.9115 | 1.3813 | 2.5474 | 100.22 | 20.76 | 37.207 | 67.3739 | 84.7199 | 64.089 | 0.009108 | 1.0469 | 0.024876 | 23.7623 | 0.01379 | 0.13877 | -3e-05 | -0.002331 | -0.002331 | -2.4233
2019-03-29 13:40 | 1553881225.5206 | -762 | 0.20855 | 0.000478 | 2e-06 | 99.3118 | 98.9568 | 126.3732 | 21.0299 | 0.00189 | 0.001745 | -151.7896 | 0.23325 | 8.4048 | 0.002618 | 0.004298 | 0.004089 | 6.5e-05 | 0.001094 | 74.935 | 1.3221 | 2.5473 | 100.22 | 20.76 | 37.208 | 67.0671 | 84.5814 | 63.5889 | 0.009108 | 1.0484 | 0.024933 | 23.7814 | 0.032026 | 0.13876 | -2.9e-05 | -0.002252 | -0.002252 | -2.5076
2019-03-29 13:40 | 1553881226.5216 | -761 | 0.20854 | 0.000476 | 2e-06 | 99.2058 | 98.9797 | 126.8567 | 21.0309 | 0.001835 | 0.001801 | -151.786 | 0.23343 | 8.4016 | 0.00253 | 0.004139 | 0.004042 | 6.5e-05 | 0.001098 | 74.9186 | 1.3141 | 2.5473 | 100.22 | 20.76 | 37.208 | 65.7453 | 82.4573 | 64.9329 | 0.009108 | 1.0443 | 0.024931 | 23.8744 | 0.020411 | 0.13887 | -2.6e-05 | -0.002041 | -0.002041 | -2.7691
2019-03-29 13:40 | 1553881227.5208 | -760 | 0.20856 | 0.000472 | 2e-06 | 99.0801 | 98.7213 | 126.548 | 21.0305 | 0.001827 | 0.001768 | -151.7988 | 0.23292 | 8.4025 | 0.002619 | 0.003639 | 0.003437 | 6.6e-05 | 0.001095 | 74.9553 | 1.3057 | 2.5473 | 100.22 | 20.77 | 37.21 | 65.554 | 84.6032 | 64.1433 | 0.009114 | 1.0467 | 0.02493 | 23.8175 | 0.032446 | 0.13856 | -2.2e-05 | -0.001717 | -0.001717 | -3.2835
2019-03-29 13:40 | 1553881228.5218 | -759 | 0.20855 | 0.000473 | 2e-06 | 98.9103 | 98.6329 | 126.5781 | 21.0303 | 0.00191 | 0.001774 | -151.8063 | 0.23349 | 8.4041 | 0.00261 | 0.003775 | 0.003546 | 6.4e-05 | 0.001094 | 74.9262 | 1.3575 | 2.5474 | 100.22 | 20.77 | 37.21 | 67.5628 | 84.3742 | 64.2951 | 0.009114 | 1.0462 | 0.024927 | 23.8256 | 0.025124 | 0.13891 | -2.3e-05 | -0.001849 | -0.001849 | -3.0575
2019-03-29 13:40 | 1553881229.5207 | -758 | 0.20854 | 0.000477 | 2e-06 | 98.7254 | 98.25 | 126.6408 | 21.0305 | 0.001925 | 0.001774 | -151.8142 | 0.23228 | 8.4009 | 0.002602 | 0.004432 | 0.004256 | 6.5e-05 | 0.001108 | 74.9818 | 1.3217 | 2.5473 | 100.22 | 20.77 | 37.213 | 67.9311 | 84.1989 | 64.2986 | 0.009115 | 1.0462 | 0.024933 | 23.8317 | 0.043173 | 0.13818 | -2.7e-05 | -0.002104 | -0.002104 | -2.6729
2019-03-29 13:40 | 1553881230.5207 | -757 | 0.20853 | 0.000476 | 2e-06 | 98.6131 | 98.1981 | 126.2328 | 21.0322 | 0.001861 | 0.001758 | -151.8183 | 0.23221 | 8.4046 | 0.002687 | 0.004382 | 0.004167 | 6.5e-05 | 0.00111 | 74.929 | 1.3349 | 2.5473 | 100.22 | 20.77 | 37.214 | 66.3905 | 86.2266 | 63.9123 | 0.009115 | 1.0474 | 0.024907 | 23.7796 | 0.037721 | 0.13814 | -2.6e-05 | -0.002026 | -0.002026 | -2.7749
2019-03-29 13:40 | 1553881231.522 | -756 | 0.20857 | 0.000474 | 2e-06 | 98.552 | 98.2217 | 126.1443 | 21.0319 | 0.001908 | 0.001748 | -151.8143 | 0.23269 | 8.4022 | 0.002571 | 0.003484 | 0.00339 | 6.3e-05 | 0.001113 | 74.9054 | 1.3229 | 2.5474 | 100.22 | 20.77 | 37.214 | 67.523 | 83.4467 | 63.6737 | 0.009115 | 1.0482 | 0.024907 | 23.7629 | 0.030036 | 0.13843 | -2.4e-05 | -0.00191 | -0.00191 | -2.9494
2019-03-29 13:41 | 1553881232.521 | -755 | 0.20854 | 0.000473 | 2e-06 | 98.4443 | 97.7893 | 125.93 | 21.0312 | 0.001855 | 0.001802 | -151.8299 | 0.2328 | 8.3999 | 0.002636 | 0.003135 | 0.002887 | 6.4e-05 | 0.001113 | 74.9669 | 1.3019 | 2.5473 | 100.22 | 20.77 | 37.214 | 66.242 | 85.0102 | 64.9628 | 0.009115 | 1.0442 | 0.024839 | 23.7881 | 0.059715 | 0.1385 | -2.2e-05 | -0.00174 | -0.00174 | -3.2395
2019-03-29 13:41 | 1553881233.5209 | -754 | 0.20856 | 0.000473 | 2e-06 | 98.3033 | 97.9691 | 126.1126 | 21.0312 | 0.001901 | 0.001767 | -151.8145 | 0.23356 | 8.4011 | 0.002634 | 0.004005 | 0.003836 | 6.5e-05 | 0.001118 | 75.0175 | 1.36 | 2.5474 | 100.22 | 20.77 | 37.215 | 67.3499 | 84.954 | 64.1298 | 0.009115 | 1.0467 | 0.024887 | 23.776 | 0.03046 | 0.13894 | -2.2e-05 | -0.001737 | -0.001737 | -3.2553
2019-03-29 13:41 | 1553881234.522 | -753 | 0.20853 | 0.000471 | 2e-06 | 98.2674 | 97.7551 | 125.8746 | 21.0322 | 0.001867 | 0.001781 | -151.804 | 0.23297 | 8.4033 | 0.002544 | 0.004169 | 0.003907 | 6.3e-05 | 0.001121 | 75.0135 | 1.3678 | 2.5472 | 100.22 | 20.77 | 37.215 | 66.5239 | 82.7901 | 64.4532 | 0.009115 | 1.0457 | 0.024852 | 23.7649 | 0.046749 | 0.13859 | -2e-05 | -0.001614 | -0.001614 | -3.4949
2019-03-29 13:41 | 1553881235.5201 | -752 | 0.20853 | 0.000477 | 2e-06 | 98.0671 | 97.6521 | 125.5793 | 21.0314 | 0.001925 | 0.001806 | -151.8088 | 0.23306 | 8.4035 | 0.002622 | 0.004399 | 0.004152 | 6.7e-05 | 0.001117 | 74.9845 | 1.3657 | 2.5473 | 100.22 | 20.77 | 37.216 | 67.9212 | 84.6795 | 65.0715 | 0.009116 | 1.0438 | 0.0248 | 23.7588 | 0.037939 | 0.13865 | -2.7e-05 | -0.002134 | -0.002134 | -2.6444
2019-03-29 13:41 | 1553881236.5209 | -751 | 0.20852 | 0.000479 | 2e-06 | 97.9387 | 97.3155 | 125.7333 | 21.0317 | 0.001934 | 0.00179 | -151.8212 | 0.23278 | 8.404 | 0.002593 | 0.004275 | 0.004129 | 6.5e-05 | 0.001113 | 74.9707 | 1.3563 | 2.5473 | 100.22 | 20.76 | 37.215 | 68.1366 | 83.9824 | 64.6715 | 0.00911 | 1.0451 | 0.02483 | 23.7592 | 0.057095 | 0.13848 | -3e-05 | -0.002287 | -0.002287 | -2.4639
2019-03-29 13:41 | 1553881237.5199 | -750 | 0.20856 | 0.000476 | 2e-06 | 97.8381 | 97.4299 | 125.9819 | 21.0329 | 0.001899 | 0.001835 | -151.8222 | 0.23384 | 8.4007 | 0.002576 | 0.00384 | 0.003742 | 6.4e-05 | 0.001106 | 75.0048 | 1.3788 | 2.5474 | 100.22 | 20.76 | 37.215 | 67.3025 | 83.5575 | 65.7629 | 0.00911 | 1.0417 | 0.024814 | 23.8211 | 0.037403 | 0.13911 | -2.6e-05 | -0.002059 | -0.002059 | -2.7503
2019-03-29 13:41 | 1553881238.521 | -749 | 0.20854 | 0.00048 | 2e-06 | 97.7199 | 97.2473 | 126.1915 | 21.033 | 0.001875 | 0.00175 | -151.8226 | 0.23327 | 8.4042 | 0.002597 | 0.00434 | 0.004143 | 6.3e-05 | 0.001096 | 75.0119 | 1.3853 | 2.5472 | 100.22 | 20.77 | 37.213 | 66.7221 | 84.0678 | 63.7077 | 0.009115 | 1.0481 | 0.024911 | 23.7685 | 0.043365 | 0.13877 | -3.1e-05 | -0.002378 | -0.002378 | -2.3755
2019-03-29 13:41 | 1553881239.5213 | -748 | 0.20856 | 0.000474 | 1e-06 | 97.6053 | 97.0586 | 126.1011 | 21.0328 | 0.001869 | 0.001754 | -151.8223 | 0.23271 | 8.4015 | 0.0026 | 0.004244 | 0.00404 | 6.5e-05 | 0.001096 | 74.9877 | 1.3558 | 2.5473 | 100.22 | 20.77 | 37.213 | 66.5784 | 84.1448 | 63.8089 | 0.009115 | 1.0477 | 0.024898 | 23.7636 | 0.050241 | 0.13844 | -2.3e-05 | -0.001801 | -0.001801 | -3.128
2019-03-29 13:41 | 1553881240.5202 | -747 | 0.20856 | 0.000473 | 2e-06 | 97.4574 | 96.9886 | 126.0572 | 21.0328 | 0.001861 | 0.001785 | -151.8073 | 0.23331 | 8.4033 | 0.002635 | 0.004184 | 0.00404 | 6.5e-05 | 0.00109 | 74.9974 | 1.3289 | 2.5474 | 100.22 | 20.77 | 37.209 | 66.3771 | 84.9906 | 64.5508 | 0.009114 | 1.0454 | 0.024866 | 23.7856 | 0.04313 | 0.13879 | -2.2e-05 | -0.001755 | -0.001755 | -3.219
2019-03-29 13:41 | 1553881241.5201 | -746 | 0.20853 | 0.000475 | 1e-06 | 97.2021 | 96.6504 | 126.0802 | 21.0337 | 0.001897 | 0.001755 | -151.7974 | 0.23304 | 8.4039 | 0.002601 | 0.004138 | 0.003983 | 6.5e-05 | 0.001093 | 75.0299 | 1.3339 | 2.5473 | 100.22 | 20.77 | 37.209 | 67.2463 | 84.18 | 63.8239 | 0.009114 | 1.0477 | 0.024895 | 23.7621 | 0.050911 | 0.13863 | -2.4e-05 | -0.001872 | -0.001872 | -3.0149
2019-03-29 13:41 | 1553881242.5202 | -745 | 0.2085 | 0.00048 | 2e-06 | 97.1563 | 96.5579 | 126.1392 | 21.0336 | 0.001837 | 0.001758 | -151.7946 | 0.23317 | 8.4015 | 0.002574 | 0.003584 | 0.003375 | 6.7e-05 | 0.001091 | 74.9895 | 1.323 | 2.5473 | 100.22 | 20.77 | 37.204 | 65.7973 | 83.5098 | 63.8972 | 0.009113 | 1.0475 | 0.024899 | 23.7703 | 0.05526 | 0.13872 | -3.1e-05 | -0.002386 | -0.002386 | -2.3659
2019-03-29 13:41 | 1553881243.5212 | -744 | 0.20847 | 0.000474 | 2e-06 | 96.9862 | 96.1473 | 126.3518 | 21.0344 | 0.00188 | 0.00178 | -151.8093 | 0.23351 | 8.4008 | 0.002547 | 0.003459 | 0.00334 | 6.5e-05 | 0.001098 | 75.0278 | 1.3568 | 2.5473 | 100.22 | 20.77 | 37.2 | 66.8413 | 82.8773 | 64.4305 | 0.009112 | 1.0458 | 0.0249 | 23.8091 | 0.077707 | 0.13892 | -2.3e-05 | -0.001829 | -0.001829 | -3.0912
2019-03-29 13:41 | 1553881244.5202 | -743 | 0.20849 | 0.000476 | 2e-06 | 96.7914 | 96.1145 | 126.5208 | 21.0339 | 0.0019 | 0.001764 | -151.8209 | 0.23276 | 8.4026 | 0.002608 | 0.004585 | 0.004367 | 6.5e-05 | 0.001102 | 74.9325 | 1.3699 | 2.5474 | 100.22 | 20.77 | 37.191 | 67.323 | 84.3457 | 64.0412 | 0.00911 | 1.047 | 0.024931 | 23.8113 | 0.062766 | 0.13847 | -2.6e-05 | -0.002023 | -0.002023 | -2.7859
2019-03-29 13:41 | 1553881245.5213 | -742 | 0.20848 | 0.000475 | 1e-06 | 96.5947 | 95.8606 | 126.3751 | 21.035 | 0.001892 | 0.00171 | -151.8263 | 0.2329 | 8.4011 | 0.002599 | 0.003708 | 0.003552 | 6.6e-05 | 0.001103 | 74.8914 | 1.3356 | 2.5472 | 100.22 | 20.77 | 37.186 | 67.1211 | 84.1128 | 62.7497 | 0.009108 | 1.051 | 0.024964 | 23.752 | 0.068235 | 0.13855 | -2.5e-05 | -0.001957 | -0.001957 | -2.882
2019-03-29 13:41 | 1553881246.5209 | -741 | 0.20848 | 0.000477 | 2e-06 | 96.3517 | 95.5466 | 125.9192 | 21.0341 | 0.001899 | 0.001777 | -151.8288 | 0.23289 | 8.4039 | 0.002716 | 0.004196 | 0.004012 | 6.4e-05 | 0.001106 | 74.9881 | 1.3672 | 2.5473 | 100.22 | 20.77 | 37.18 | 67.2902 | 86.9456 | 64.3669 | 0.009107 | 1.046 | 0.02486 | 23.7661 | 0.075046 | 0.13855 | -2.8e-05 | -0.002176 | -0.002176 | -2.591
2019-03-29 13:41 | 1553881247.5207 | -740 | 0.20849 | 0.000476 | 2e-06 | 96.1475 | 95.4391 | 125.8814 | 21.0354 | 0.001918 | 0.001757 | -151.8174 | 0.23228 | 8.401 | 0.00272 | 0.003869 | 0.003769 | 6.5e-05 | 0.001104 | 74.9508 | 1.363 | 2.5473 | 100.22 | 20.77 | 37.16 | 67.762 | 87.0297 | 63.891 | 0.009102 | 1.0475 | 0.024873 | 23.7458 | 0.066144 | 0.13818 | -2.6e-05 | -0.002027 | -0.002027 | -2.7748
2019-03-29 13:41 | 1553881248.5204 | -739 | 0.20854 | 0.000473 | 1e-06 | 95.9036 | 95.0685 | 125.6162 | 21.0344 | 0.001871 | 0.00178 | -151.8092 | 0.23273 | 8.4031 | 0.002666 | 0.003985 | 0.003782 | 6.4e-05 | 0.001108 | 74.9861 | 1.3267 | 2.5473 | 100.22 | 20.77 | 37.16 | 66.6196 | 85.7289 | 64.4476 | 0.009102 | 1.0458 | 0.024827 | 23.7403 | 0.078224 | 0.13845 | -2.2e-05 | -0.001741 | -0.001741 | -3.2365
2019-03-29 13:41 | 1553881249.5214 | -738 | 0.2085 | 0.000473 | 1e-06 | 95.7294 | 94.9825 | 125.5207 | 21.0343 | 0.001932 | 0.001726 | -151.8016 | 0.23271 | 8.4023 | 0.002663 | 0.003024 | 0.002865 | 6.7e-05 | 0.001111 | 74.9664 | 1.3506 | 2.5473 | 100.22 | 20.77 | 37.139 | 68.0908 | 85.6502 | 63.1286 | 0.009097 | 1.0499 | 0.024866 | 23.6849 | 0.070057 | 0.13844 | -2.2e-05 | -0.001672 | -0.001672 | -3.3706
2019-03-29 13:41 | 1553881250.5204 | -737 | 0.20852 | 0.000475 | 1e-06 | 95.5435 | 94.5695 | 125.8019 | 21.0352 | 0.001889 | 0.001741 | -151.8156 | 0.23326 | 8.4042 | 0.002711 | 0.003508 | 0.003285 | 6.4e-05 | 0.001112 | 74.9339 | 1.34 | 2.5473 | 100.22 | 20.77 | 37.139 | 67.0503 | 86.8093 | 63.4879 | 0.009097 | 1.0487 | 0.02488 | 23.7241 | 0.091645 | 0.13877 | -2.5e-05 | -0.001967 | -0.001967 | -2.8717
2019-03-29 13:41 | 1553881251.5194 | -736 | 0.20852 | 0.000475 | 2e-06 | 95.3067 | 94.7289 | 126.1143 | 21.035 | 0.001948 | 0.001839 | -151.8303 | 0.23328 | 8.4029 | 0.002706 | 0.003804 | 0.003673 | 6.5e-05 | 0.001112 | 74.923 | 1.3698 | 2.5475 | 100.22 | 20.77 | 37.121 | 68.4855 | 86.7078 | 65.8539 | 0.009092 | 1.0414 | 0.024824 | 23.8368 | 0.05439 | 0.13878 | -2.5e-05 | -0.001919 | -0.001919 | -2.9432
2019-03-29 13:41 | 1553881252.5197 | -735 | 0.20856 | 0.000476 | 1e-06 | 95.0903 | 94.1405 | 126.318 | 21.0347 | 0.001923 | 0.00179 | -151.8326 | 0.23198 | 8.4031 | 0.002695 | 0.003983 | 0.003784 | 6.1e-05 | 0.001115 | 74.9459 | 1.4059 | 2.5475 | 100.22 | 20.77 | 37.121 | 67.876 | 86.4404 | 64.6768 | 0.009092 | 1.045 | 0.024887 | 23.8146 | 0.089786 | 0.138 | -2.6e-05 | -0.002009 | -0.002009 | -2.7962
2019-03-29 13:41 | 1553881253.5205 | -734 | 0.2085 | 0.000475 | 2e-06 | 95.0023 | 94.0597 | 126.0014 | 21.0356 | 0.001893 | 0.001779 | -151.8374 | 0.23308 | 8.4031 | 0.002665 | 0.004042 | 0.003813 | 6.2e-05 | 0.001119 | 74.9883 | 1.3877 | 2.5473 | 100.22 | 20.77 | 37.113 | 67.1498 | 85.7177 | 64.4036 | 0.00909 | 1.0459 | 0.024866 | 23.7751 | 0.089192 | 0.13866 | -2.5e-05 | -0.001975 | -0.001975 | -2.8581
2019-03-29 13:41 | 1553881254.5206 | -733 | 0.20852 | 0.000475 | 1e-06 | 94.6628 | 93.83 | 125.9026 | 21.0348 | 0.001917 | 0.001744 | -151.8378 | 0.23265 | 8.4034 | 0.002648 | 0.004478 | 0.004309 | 6.5e-05 | 0.001124 | 74.9429 | 1.3815 | 2.5474 | 100.22 | 20.77 | 37.113 | 67.7277 | 85.3096 | 63.5679 | 0.00909 | 1.0485 | 0.024887 | 23.7364 | 0.079045 | 0.13841 | -2.4e-05 | -0.001859 | -0.001859 | -3.0309
2019-03-29 13:41 | 1553881255.5207 | -732 | 0.20853 | 0.000477 | 1e-06 | 94.459 | 93.6235 | 125.8936 | 21.0359 | 0.001866 | 0.001726 | -151.8465 | 0.23329 | 8.4018 | 0.002714 | 0.003845 | 0.003677 | 6.3e-05 | 0.001126 | 74.9224 | 1.3527 | 2.5474 | 100.22 | 20.77 | 37.107 | 66.5157 | 86.8951 | 63.1302 | 0.009089 | 1.0499 | 0.024903 | 23.7201 | 0.079468 | 0.13878 | -2.7e-05 | -0.002104 | -0.002104 | -2.6846
2019-03-29 13:41 | 1553881256.5196 | -731 | 0.20851 | 0.000473 | 1e-06 | 94.2824 | 93.1546 | 126.0511 | 21.036 | 0.001883 | 0.001734 | -151.845 | 0.23314 | 8.4016 | 0.002655 | 0.004115 | 0.003933 | 6.3e-05 | 0.001131 | 74.9151 | 1.335 | 2.5472 | 100.22 | 20.77 | 37.104 | 66.9083 | 85.4645 | 63.3263 | 0.009088 | 1.0492 | 0.024911 | 23.7418 | 0.10764 | 0.1387 | -2.2e-05 | -0.001734 | -0.001734 | -3.2547
2019-03-29 13:41 | 1553881257.5197 | -730 | 0.20854 | 0.000476 | 1e-06 | 94.0528 | 93.1071 | 126.2016 | 21.0361 | 0.00191 | 0.001829 | -151.864 | 0.23384 | 8.402 | 0.002612 | 0.00454 | 0.00431 | 6.2e-05 | 0.001133 | 74.9829 | 1.3427 | 2.5472 | 100.22 | 20.77 | 37.102 | 67.5677 | 84.4351 | 65.6272 | 0.009088 | 1.0421 | 0.024841 | 23.8371 | 0.090391 | 0.13911 | -2.6e-05 | -0.002028 | -0.002028 | -2.7917
2019-03-29 13:41 | 1553881258.5211 | -729 | 0.20851 | 0.000474 | 1e-06 | 93.8118 | 92.8583 | 126.2204 | 21.036 | 0.001838 | 0.001777 | -151.8712 | 0.2328 | 8.4014 | 0.002639 | 0.003609 | 0.003532 | 6.3e-05 | 0.00113 | 75.0003 | 1.3437 | 2.5473 | 100.22 | 20.77 | 37.103 | 65.8392 | 85.0805 | 64.3719 | 0.009088 | 1.046 | 0.024889 | 23.7947 | 0.091379 | 0.13849 | -2.3e-05 | -0.001785 | -0.001785 | -3.1574
2019-03-29 13:41 | 1553881259.5197 | -728 | 0.2085 | 0.000476 | 2e-06 | 93.5679 | 92.4 | 126.3652 | 21.0353 | 0.001893 | 0.001798 | -151.8611 | 0.23236 | 8.402 | 0.002675 | 0.003851 | 0.003656 | 6e-05 | 0.001134 | 74.9154 | 1.3389 | 2.5475 | 100.22 | 20.77 | 37.103 | 67.1596 | 85.9406 | 64.8631 | 0.009088 | 1.0445 | 0.024885 | 23.8256 | 0.11235 | 0.13823 | -2.6e-05 | -0.002007 | -0.002007 | -2.8029
2019-03-29 13:41 | 1553881260.5198 | -727 | 0.20854 | 0.000475 | 1e-06 | 93.356 | 92.1925 | 126.3053 | 21.0367 | 0.001902 | 0.001747 | -151.8425 | 0.23385 | 8.4013 | 0.00274 | 0.003378 | 0.003277 | 6.5e-05 | 0.001137 | 74.9312 | 1.3049 | 2.5473 | 100.22 | 20.77 | 37.105 | 67.374 | 87.5183 | 63.6355 | 0.009089 | 1.0483 | 0.024925 | 23.7767 | 0.11218 | 0.13912 | -2.5e-05 | -0.001939 | -0.001939 | -2.9202
2019-03-29 13:41 | 1553881261.5212 | -726 | 0.20852 | 0.000475 | 1e-06 | 93.1799 | 92.1215 | 125.9871 | 21.0359 | 0.00189 | 0.001775 | -151.8315 | 0.23297 | 8.4022 | 0.002641 | 0.00356 | 0.003395 | 6.3e-05 | 0.001139 | 74.8842 | 1.3205 | 2.5472 | 100.22 | 20.77 | 37.105 | 67.0862 | 85.1241 | 64.3117 | 0.009089 | 1.0462 | 0.024868 | 23.7705 | 0.10219 | 0.1386 | -2.5e-05 | -0.001939 | -0.001939 | -2.9085
2019-03-29 13:41 | 1553881262.5209 | -725 | 0.20852 | 0.000475 | 2e-06 | 92.9192 | 91.7829 | 125.5775 | 21.0374 | 0.00192 | 0.001775 | -151.8356 | 0.23269 | 8.4027 | 0.002714 | 0.003952 | 0.003797 | 6e-05 | 0.001143 | 74.9253 | 1.3941 | 2.5473 | 100.22 | 20.77 | 37.108 | 67.7991 | 86.8967 | 64.3139 | 0.009089 | 1.0462 | 0.024828 | 23.7319 | 0.11006 | 0.13843 | -2.5e-05 | -0.001954 | -0.001954 | -2.8827
2019-03-29 13:41 | 1553881263.5199 | -724 | 0.2085 | 0.000475 | 1e-06 | 92.7785 | 91.6792 | 125.6094 | 21.0371 | 0.001902 | 0.001757 | -151.8301 | 0.23272 | 8.4031 | 0.002606 | 0.004114 | 0.003941 | 6.1e-05 | 0.001145 | 74.9198 | 1.3675 | 2.5474 | 100.22 | 20.77 | 37.108 | 67.3814 | 84.3005 | 63.8811 | 0.009089 | 1.0475 | 0.024847 | 23.7197 | 0.10662 | 0.13845 | -2.5e-05 | -0.001902 | -0.001902 | -2.9621
2019-03-29 13:41 | 1553881264.5209 | -723 | 0.2085 | 0.000474 | 1e-06 | 92.5867 | 91.5713 | 125.3343 | 21.0375 | 0.001872 | 0.001773 | -151.8086 | 0.23278 | 8.4053 | 0.00258 | 0.003962 | 0.003849 | 5.9e-05 | 0.001144 | 74.945 | 1.3807 | 2.5475 | 100.22 | 20.77 | 37.108 | 66.65 | 83.6596 | 64.2827 | 0.009089 | 1.0463 | 0.024805 | 23.7079 | 0.098632 | 0.13848 | -2.3e-05 | -0.001782 | -0.001782 | -3.1633
2019-03-29 13:41 | 1553881265.521 | -722 | 0.20853 | 0.00047 | 1e-06 | 92.3837 | 90.9898 | 125.3477 | 21.0369 | 0.001869 | 0.001812 | -151.7941 | 0.23292 | 8.401 | 0.002662 | 0.004389 | 0.004128 | 6.2e-05 | 0.001146 | 74.9963 | 1.3571 | 2.5475 | 100.22 | 20.77 | 37.109 | 66.5768 | 85.6299 | 65.2142 | 0.009089 | 1.0434 | 0.024772 | 23.7418 | 0.13599 | 0.13856 | -1.8e-05 | -0.001411 | -0.001411 | -3.9963
2019-03-29 13:41 | 1553881266.52 | -721 | 0.20857 | 0.000475 | 1e-06 | 92.1925 | 90.8505 | 125.059 | 21.0378 | 0.001911 | 0.001793 | -151.8005 | 0.23219 | 8.4006 | 0.002627 | 0.004823 | 0.004523 | 6.1e-05 | 0.00115 | 75.0302 | 1.3438 | 2.5473 | 100.22 | 20.77 | 37.108 | 67.5841 | 84.7897 | 64.7414 | 0.009089 | 1.0448 | 0.024761 | 23.6979 | 0.13115 | 0.13813 | -2.5e-05 | -0.001912 | -0.001912 | -2.9395
2019-03-29 13:41 | 1553881267.5201 | -720 | 0.20854 | 0.000471 | 1e-06 | 91.9461 | 90.6744 | 125.0719 | 21.0377 | 0.001858 | 0.001741 | -151.7995 | 0.23297 | 8.4008 | 0.002561 | 0.003614 | 0.003556 | 6.1e-05 | 0.001154 | 74.9381 | 1.3581 | 2.5475 | 100.22 | 20.77 | 37.108 | 66.3116 | 83.2159 | 63.4952 | 0.009089 | 1.0487 | 0.024808 | 23.6554 | 0.12457 | 0.1386 | -2e-05 | -0.001573 | -0.001573 | -3.5853
2019-03-29 13:41 | 1553881268.52 | -719 | 0.20855 | 0.000476 | 1e-06 | 91.6722 | 90.3065 | 125.251 | 21.0387 | 0.001846 | 0.001811 | -151.8045 | 0.23213 | 8.4035 | 0.002584 | 0.003916 | 0.003694 | 5.7e-05 | 0.001156 | 74.9689 | 1.3311 | 2.5474 | 100.22 | 20.77 | 37.108 | 66.023 | 83.7581 | 65.1785 | 0.009089 | 1.0435 | 0.024764 | 23.7314 | 0.13426 | 0.13809 | -2.6e-05 | -0.001991 | -0.001991 | -2.8226
2019-03-29 13:41 | 1553881269.5201 | -718 | 0.20856 | 0.000476 | 2e-06 | 91.5118 | 90.0232 | 125.4505 | 21.0386 | 0.001859 | 0.001758 | -151.8127 | 0.23313 | 8.4036 | 0.002598 | 0.004146 | 0.003988 | 6.2e-05 | 0.001159 | 74.9305 | 1.3282 | 2.5474 | 100.22 | 20.77 | 37.106 | 66.3483 | 84.0913 | 63.9161 | 0.009089 | 1.0474 | 0.02483 | 23.706 | 0.14669 | 0.13869 | -2.6e-05 | -0.00202 | -0.00202 | -2.7948
2019-03-29 13:41 | 1553881270.5191 | -717 | 0.20856 | 0.000473 | 1e-06 | 91.2115 | 89.8505 | 125.4594 | 21.0398 | 0.001872 | 0.001728 | -151.8345 | 0.23312 | 8.4033 | 0.002651 | 0.003916 | 0.003752 | 6.1e-05 | 0.001161 | 74.9208 | 1.3128 | 2.5474 | 100.22 | 20.77 | 37.104 | 66.6568 | 85.3837 | 63.1997 | 0.009088 | 1.0496 | 0.024857 | 23.6816 | 0.13447 | 0.13869 | -2.2e-05 | -0.001691 | -0.001691 | -3.3377
2019-03-29 13:41 | 1553881271.5192 | -716 | 0.20856 | 0.000474 | 2e-06 | 90.9975 | 89.6454 | 125.6873 | 21.0386 | 0.001894 | 0.001774 | -151.8365 | 0.23372 | 8.4017 | 0.002636 | 0.003835 | 0.003641 | 5.8e-05 | 0.001159 | 75.0231 | 1.3155 | 2.5473 | 100.22 | 20.77 | 37.097 | 67.1893 | 85.0198 | 64.3056 | 0.009087 | 1.0462 | 0.024839 | 23.742 | 0.1339 | 0.13904 | -2.4e-05 | -0.001884 | -0.001884 | -3.0038
2019-03-29 13:41 | 1553881272.5201 | -715 | 0.20855 | 0.000473 | 1e-06 | 90.8179 | 89.4069 | 125.9896 | 21.0381 | 0.001928 | 0.001798 | -151.8148 | 0.23349 | 8.4039 | 0.002711 | 0.004111 | 0.003874 | 6.1e-05 | 0.001159 | 75.0171 | 1.3025 | 2.5474 | 100.22 | 20.77 | 37.097 | 67.994 | 86.8278 | 64.8784 | 0.009087 | 1.0444 | 0.024848 | 23.7907 | 0.14006 | 0.13891 | -2.2e-05 | -0.001679 | -0.001679 | -3.3665
2019-03-29 13:41 | 1553881273.5205 | -714 | 0.20856 | 0.000473 | 1e-06 | 90.7057 | 89.3495 | 126.2589 | 21.04 | 0.001837 | 0.001773 | -151.8014 | 0.2334 | 8.4022 | 0.00269 | 0.004517 | 0.004388 | 5.9e-05 | 0.001159 | 74.9714 | 1.3147 | 2.5473 | 100.22 | 20.77 | 37.094 | 65.8066 | 86.3169 | 64.2704 | 0.009086 | 1.0463 | 0.024897 | 23.7947 | 0.13474 | 0.13885 | -2.3e-05 | -0.001778 | -0.001778 | -3.1782
2019-03-29 13:41 | 1553881274.5192 | -713 | 0.20854 | 0.000477 | 1e-06 | 90.4716 | 89.0983 | 126.6647 | 21.0395 | 0.001884 | 0.001791 | -151.8138 | 0.23397 | 8.4016 | 0.002646 | 0.004487 | 0.004265 | 6.1e-05 | 0.001154 | 74.9456 | 1.311 | 2.5474 | 100.22 | 20.77 | 37.09 | 66.9523 | 85.2478 | 64.6992 | 0.009085 | 1.045 | 0.024921 | 23.8481 | 0.13681 | 0.13919 | -2.7e-05 | -0.002096 | -0.002096 | -2.7022
2019-03-29 13:41 | 1553881275.5196 | -712 | 0.20855 | 0.000476 | 1e-06 | 90.321 | 88.8506 | 126.6626 | 21.0395 | 0.001968 | 0.001763 | -151.8287 | 0.23313 | 8.4038 | 0.002733 | 0.004314 | 0.004127 | 6e-05 | 0.001155 | 74.986 | 1.3708 | 2.5474 | 100.22 | 20.77 | 37.088 | 68.9524 | 87.3516 | 64.0279 | 0.009084 | 1.0471 | 0.024945 | 23.8242 | 0.14682 | 0.13869 | -2.6e-05 | -0.002028 | -0.002028 | -2.7834
2019-03-29 13:41 | 1553881276.5196 | -711 | 0.20854 | 0.000473 | 1e-06 | 90.1884 | 88.7321 | 126.7506 | 21.0402 | 0.001876 | 0.001754 | -151.8187 | 0.23312 | 8.4014 | 0.002644 | 0.003851 | 0.003608 | 5.9e-05 | 0.001152 | 74.9554 | 1.3351 | 2.5473 | 100.22 | 20.77 | 37.084 | 66.7635 | 85.2117 | 63.817 | 0.009083 | 1.0477 | 0.024962 | 23.825 | 0.14561 | 0.13868 | -2.3e-05 | -0.001789 | -0.001789 | -3.1556
2019-03-29 13:41 | 1553881277.5202 | -710 | 0.20857 | 0.000473 | 1e-06 | 90.2663 | 88.6328 | 126.7051 | 21.0405 | 0.001847 | 0.001809 | -151.8143 | 0.23236 | 8.4006 | 0.00258 | 0.003715 | 0.003572 | 5.7e-05 | 0.001151 | 74.984 | 1.3378 | 2.5473 | 100.22 | 20.77 | 37.084 | 66.0606 | 83.6663 | 65.1359 | 0.009083 | 1.0436 | 0.024909 | 23.8673 | 0.16334 | 0.13823 | -2.2e-05 | -0.00172 | -0.00172 | -3.2702
2019-03-29 13:41 | 1553881278.5194 | -709 | 0.20856 | 0.000473 | 1e-06 | 90.2181 | 88.5854 | 126.5776 | 21.0411 | 0.001921 | 0.001783 | -151.8107 | 0.23289 | 8.402 | 0.00267 | 0.003904 | 0.003778 | 5.8e-05 | 0.001153 | 74.9988 | 1.3933 | 2.5474 | 100.22 | 20.77 | 37.083 | 67.8315 | 85.8356 | 64.5049 | 0.009083 | 1.0456 | 0.024919 | 23.833 | 0.16335 | 0.13855 | -2.3e-05 | -0.001761 | -0.001761 | -3.2029
2019-03-29 13:41 | 1553881279.5193 | -708 | 0.20856 | 0.000476 | 1e-06 | 90.1962 | 88.7649 | 126.5722 | 21.0415 | 0.001893 | 0.001719 | -151.8362 | 0.23364 | 8.4015 | 0.002654 | 0.003642 | 0.003534 | 6.1e-05 | 0.00115 | 74.9334 | 1.3584 | 2.5474 | 100.22 | 20.77 | 37.079 | 67.1681 | 85.4472 | 62.9795 | 0.009082 | 1.0503 | 0.024975 | 23.7786 | 0.14308 | 0.13899 | -2.5e-05 | -0.001965 | -0.001965 | -2.8787
2019-03-29 13:41 | 1553881280.5204 | -707 | 0.20851 | 0.000473 | 1e-06 | 90.2689 | 88.67 | 126.7241 | 21.0411 | 0.001835 | 0.001761 | -151.8352 | 0.23359 | 8.4031 | 0.002726 | 0.004225 | 0.00404 | 5.8e-05 | 0.001149 | 74.9309 | 1.3407 | 2.5474 | 100.22 | 20.77 | 37.078 | 65.7615 | 87.1844 | 63.9842 | 0.009082 | 1.0472 | 0.024953 | 23.8284 | 0.15985 | 0.13896 | -2.3e-05 | -0.001774 | -0.001774 | -3.1885
2019-03-29 13:41 | 1553881281.5199 | -706 | 0.20855 | 0.000476 | 1e-06 | 90.4587 | 88.9287 | 126.4236 | 21.0405 | 0.001923 | 0.00175 | -151.8435 | 0.232 | 8.4058 | 0.002732 | 0.003724 | 0.003532 | 6.1e-05 | 0.001147 | 74.923 | 1.3705 | 2.5474 | 100.22 | 20.77 | 37.078 | 67.8817 | 87.3384 | 63.7126 | 0.009082 | 1.048 | 0.024933 | 23.7905 | 0.15258 | 0.13801 | -2.5e-05 | -0.001978 | -0.001978 | -2.8397
2019-03-29 13:41 | 1553881282.519 | -705 | 0.20845 | 0.000473 | 1e-06 | 90.6386 | 89.1424 | 126.1183 | 21.0412 | 0.001902 | 0.00179 | -151.8432 | 0.23275 | 8.4032 | 0.002675 | 0.003715 | 0.003572 | 6.3e-05 | 0.001149 | 74.9592 | 1.3887 | 2.5474 | 100.22 | 20.77 | 37.078 | 67.3846 | 85.965 | 64.6915 | 0.009082 | 1.045 | 0.024867 | 23.7963 | 0.14889 | 0.13846 | -2.2e-05 | -0.001692 | -0.001692 | -3.33
2019-03-29 13:41 | 1553881283.5195 | -704 | 0.20852 | 0.000474 | 2e-06 | 90.9903 | 89.4567 | 126.0131 | 21.0414 | 0.001912 | 0.001743 | -151.8477 | 0.23254 | 8.402 | 0.002615 | 0.004277 | 0.004011 | 6.3e-05 | 0.00115 | 74.9323 | 1.3559 | 2.5474 | 100.22 | 20.77 | 37.078 | 67.6243 | 84.5074 | 63.5518 | 0.009082 | 1.0485 | 0.024899 | 23.7462 | 0.15205 | 0.13834 | -2.4e-05 | -0.001866 | -0.001866 | -3.0169
2019-03-29 13:41 | 1553881284.5191 | -703 | 0.20851 | 0.000475 | 1e-06 | 91.2151 | 89.8949 | 125.8962 | 21.0424 | 0.001878 | 0.001787 | -151.8449 | 0.23324 | 8.4011 | 0.002652 | 0.003532 | 0.003391 | 6.2e-05 | 0.001152 | 74.9743 | 1.3678 | 2.5474 | 100.22 | 20.77 | 37.078 | 66.7929 | 85.4019 | 64.6186 | 0.009082 | 1.0452 | 0.024848 | 23.7728 | 0.13041 | 0.13875 | -2.5e-05 | -0.001929 | -0.001929 | -2.9282
2019-03-29 13:41 | 1553881285.52 | -702 | 0.2085 | 0.000474 | 1e-06 | 91.4771 | 90.0669 | 126.1368 | 21.0422 | 0.001896 | 0.001774 | -151.8432 | 0.23221 | 8.4011 | 0.002663 | 0.003678 | 0.003467 | 6.1e-05 | 0.001153 | 74.9151 | 1.4108 | 2.5475 | 100.22 | 20.77 | 37.078 | 67.2366 | 85.6601 | 64.3029 | 0.009082 | 1.0462 | 0.024883 | 23.7844 | 0.13895 | 0.13814 | -2.4e-05 | -0.001886 | -0.001886 | -2.9808
2019-03-29 13:41 | 1553881286.5196 | -701 | 0.20852 | 0.000475 | 1e-06 | 91.6662 | 90.384 | 126.2187 | 21.0416 | 0.001901 | 0.001732 | -151.8455 | 0.23208 | 8.4006 | 0.002691 | 0.003772 | 0.003581 | 6.2e-05 | 0.001156 | 74.8896 | 1.3151 | 2.5474 | 100.22 | 20.77 | 37.077 | 67.3444 | 86.3369 | 63.2965 | 0.009082 | 1.0493 | 0.024929 | 23.7566 | 0.126 | 0.13806 | -2.5e-05 | -0.001897 | -0.001897 | -2.9623
2019-03-29 13:41 | 1553881287.5191 | -700 | 0.20852 | 0.000474 | 1e-06 | 91.8596 | 90.4775 | 126.2845 | 21.0424 | 0.001887 | 0.001764 | -151.8426 | 0.23235 | 8.4035 | 0.002573 | 0.003864 | 0.003741 | 5.9e-05 | 0.001162 | 74.942 | 1.3149 | 2.5472 | 100.22 | 20.77 | 37.077 | 67.0199 | 83.5069 | 64.062 | 0.009082 | 1.047 | 0.024907 | 23.7898 | 0.1356 | 0.13822 | -2.3e-05 | -0.001833 | -0.001833 | -3.0692
2019-03-29 13:41 | 1553881288.5197 | -699 | 0.20851 | 0.000471 | 2e-06 | 91.9449 | 90.8058 | 126.2934 | 21.0428 | 0.001922 | 0.001775 | -151.8284 | 0.23225 | 8.4041 | 0.002694 | 0.004519 | 0.004252 | 6.1e-05 | 0.001168 | 74.9633 | 1.3276 | 2.5474 | 100.22 | 20.77 | 37.076 | 67.8519 | 86.4132 | 64.3136 | 0.009081 | 1.0462 | 0.024898 | 23.7995 | 0.11151 | 0.13817 | -2e-05 | -0.001617 | -0.001617 | -3.4783
2019-03-29 13:41 | 1553881289.5198 | -698 | 0.20854 | 0.000479 | 1e-06 | 92.1323 | 90.9048 | 126.4962 | 21.0432 | 0.001847 | 0.00181 | -151.8031 | 0.23308 | 8.402 | 0.002667 | 0.003981 | 0.003837 | 6.1e-05 | 0.001163 | 74.9991 | 1.3582 | 2.5474 | 100.22 | 20.77 | 37.074 | 66.0663 | 85.7688 | 65.1643 | 0.009081 | 1.0435 | 0.024887 | 23.8486 | 0.11997 | 0.13866 | -3e-05 | -0.002281 | -0.002281 | -2.4741
2019-03-29 13:41 | 1553881290.5188 | -697 | 0.20853 | 0.000474 | 2e-06 | 92.3547 | 91.0826 | 126.5584 | 21.0429 | 0.001919 | 0.001814 | -151.8106 | 0.23323 | 8.4018 | 0.002634 | 0.003983 | 0.003841 | 6.1e-05 | 0.001158 | 74.912 | 1.3424 | 2.5475 | 100.22 | 20.77 | 37.075 | 67.7856 | 84.9653 | 65.2683 | 0.009081 | 1.0432 | 0.024889 | 23.8581 | 0.12406 | 0.13875 | -2.4e-05 | -0.001893 | -0.001893 | -2.9825
2019-03-29 13:41 | 1553881291.5188 | -696 | 0.20853 | 0.000472 | 2e-06 | 92.4633 | 91.2136 | 126.7896 | 21.0431 | 0.00191 | 0.001792 | -151.804 | 0.23269 | 8.4028 | 0.002671 | 0.003608 | 0.003473 | 6.1e-05 | 0.00116 | 74.9495 | 1.3953 | 2.5474 | 100.22 | 20.77 | 37.074 | 67.5635 | 85.8592 | 64.7239 | 0.009081 | 1.0449 | 0.024932 | 23.8607 | 0.12172 | 0.13843 | -2.1e-05 | -0.001653 | -0.001653 | -3.4079
2019-03-29 13:42 | 1553881292.5199 | -695 | 0.20853 | 0.000473 | 1e-06 | 92.651 | 91.5529 | 127.0553 | 21.0428 | 0.00187 | 0.001774 | -151.8173 | 0.23265 | 8.4045 | 0.002644 | 0.003927 | 0.003733 | 6e-05 | 0.001164 | 74.9983 | 1.3622 | 2.5473 | 100.22 | 20.77 | 37.074 | 66.6215 | 85.2173 | 64.3105 | 0.009081 | 1.0462 | 0.024973 | 23.8711 | 0.10665 | 0.13841 | -2.2e-05 | -0.001728 | -0.001728 | -3.2608
2019-03-29 13:42 | 1553881293.5198 | -694 | 0.20852 | 0.000479 | 2e-06 | 92.8728 | 91.685 | 126.9673 | 21.0439 | 0.001886 | 0.001779 | -151.8212 | 0.23252 | 8.4024 | 0.002594 | 0.003365 | 0.003176 | 6.1e-05 | 0.001166 | 74.9488 | 1.3571 | 2.5474 | 100.22 | 20.77 | 37.073 | 66.9955 | 83.9997 | 64.4169 | 0.009081 | 1.0459 | 0.024961 | 23.8666 | 0.11514 | 0.13833 | -3e-05 | -0.002346 | -0.002346 | -2.3996
2019-03-29 13:42 | 1553881294.519 | -693 | 0.20855 | 0.000476 | 2e-06 | 93.0411 | 91.8566 | 126.9501 | 21.0436 | 0.00188 | 0.001764 | -151.8413 | 0.23227 | 8.403 | 0.002675 | 0.004358 | 0.004088 | 6e-05 | 0.001161 | 74.9441 | 1.3496 | 2.5474 | 100.22 | 20.77 | 37.073 | 66.844 | 85.9659 | 64.0552 | 0.009081 | 1.047 | 0.024973 | 23.8522 | 0.1146 | 0.13818 | -2.6e-05 | -0.002033 | -0.002033 | -2.7662
2019-03-29 13:42 | 1553881295.5193 | -692 | 0.20853 | 0.000473 | 2e-06 | 93.253 | 92.0525 | 126.8039 | 21.0439 | 0.001948 | 0.001816 | -151.8233 | 0.23301 | 8.4017 | 0.002669 | 0.00377 | 0.003631 | 6.2e-05 | 0.001156 | 75.0247 | 1.3342 | 2.5474 | 100.22 | 20.78 | 37.073 | 68.4824 | 85.8133 | 65.3014 | 0.009086 | 1.0431 | 0.024912 | 23.8824 | 0.11589 | 0.13861 | -2.2e-05 | -0.00177 | -0.00177 | -3.1877
2019-03-29 13:42 | 1553881296.519 | -691 | 0.20853 | 0.000475 | 1e-06 | 93.4167 | 92.2897 | 126.5384 | 21.0439 | 0.001848 | 0.001761 | -151.8113 | 0.23256 | 8.4014 | 0.002651 | 0.003592 | 0.003466 | 6.1e-05 | 0.001153 | 75.0699 | 1.3604 | 2.5474 | 100.22 | 20.77 | 37.073 | 66.0709 | 85.3762 | 63.9799 | 0.009081 | 1.0472 | 0.024935 | 23.8108 | 0.10856 | 0.13835 | -2.5e-05 | -0.001904 | -0.001904 | -2.9568
2019-03-29 13:42 | 1553881297.519 | -690 | 0.20852 | 0.000476 | 1e-06 | 93.5433 | 92.5467 | 126.5234 | 21.0454 | 0.001867 | 0.001777 | -151.8039 | 0.23351 | 8.4012 | 0.002555 | 0.003857 | 0.003594 | 6e-05 | 0.001157 | 74.9956 | 1.3336 | 2.5474 | 100.22 | 20.78 | 37.072 | 66.5469 | 83.0844 | 64.3638 | 0.009086 | 1.046 | 0.024919 | 23.8229 | 0.095807 | 0.13891 | -2.6e-05 | -0.002006 | -0.002006 | -2.8189
2019-03-29 13:42 | 1553881298.5193 | -689 | 0.20853 | 0.000472 | 2e-06 | 93.819 | 92.8795 | 126.3983 | 21.0449 | 0.001884 | 0.001759 | -151.8336 | 0.23283 | 8.4017 | 0.002563 | 0.003566 | 0.003302 | 5.9e-05 | 0.001157 | 74.991 | 1.3271 | 2.5474 | 100.22 | 20.77 | 37.072 | 66.9506 | 83.2758 | 63.9496 | 0.00908 | 1.0473 | 0.024922 | 23.7965 | 0.090026 | 0.13851 | -2.1e-05 | -0.001687 | -0.001687 | -3.3413
2019-03-29 13:42 | 1553881299.518 | -688 | 0.20855 | 0.000477 | 2e-06 | 93.9423 | 92.8924 | 126.256 | 21.0459 | 0.001845 | 0.001816 | -151.8508 | 0.23291 | 8.4018 | 0.002573 | 0.003934 | 0.003854 | 6e-05 | 0.001154 | 74.9825 | 1.315 | 2.5475 | 100.22 | 20.77 | 37.072 | 66.0119 | 83.5109 | 65.3221 | 0.00908 | 1.0431 | 0.024858 | 23.8315 | 0.10053 | 0.13856 | -2.8e-05 | -0.00217 | -0.00217 | -2.5986
2019-03-29 13:42 | 1553881300.519 | -687 | 0.20852 | 0.000471 | 2e-06 | 94.1465 | 93.1088 | 126.3269 | 21.0454 | 0.001867 | 0.001779 | -151.833 | 0.2329 | 8.4008 | 0.00256 | 0.003715 | 0.003528 | 6.1e-05 | 0.001154 | 74.9624 | 1.32 | 2.5473 | 100.22 | 20.77 | 37.072 | 66.5493 | 83.1967 | 64.4232 | 0.00908 | 1.0458 | 0.024898 | 23.8065 | 0.099132 | 0.13855 | -2e-05 | -0.001584 | -0.001584 | -3.5589
2019-03-29 13:42 | 1553881301.5191 | -686 | 0.20849 | 0.000472 | 2e-06 | 94.3099 | 93.3017 | 126.4962 | 21.0451 | 0.001898 | 0.001751 | -151.8296 | 0.23304 | 8.403 | 0.002543 | 0.00407 | 0.003864 | 6.1e-05 | 0.001156 | 74.9533 | 1.3222 | 2.5474 | 100.22 | 20.77 | 37.072 | 67.2964 | 82.7918 | 63.7592 | 0.00908 | 1.0479 | 0.024939 | 23.799 | 0.096141 | 0.13863 | -2.1e-05 | -0.001671 | -0.001671 | -3.3776
2019-03-29 13:42 | 1553881302.5191 | -685 | 0.20855 | 0.000475 | 1e-06 | 94.5174 | 93.5119 | 126.3225 | 21.0448 | 0.001929 | 0.001787 | -151.844 | 0.23286 | 8.405 | 0.002673 | 0.004305 | 0.004106 | 5.9e-05 | 0.001149 | 74.9969 | 1.3489 | 2.5475 | 100.22 | 20.77 | 37.071 | 68.0409 | 85.9123 | 64.6272 | 0.00908 | 1.0452 | 0.02489 | 23.8133 | 0.095662 | 0.13853 | -2.4e-05 | -0.00189 | -0.00189 | -2.983
2019-03-29 13:42 | 1553881303.5181 | -684 | 0.20853 | 0.000475 | 1e-06 | 94.5641 | 93.5668 | 126.1443 | 21.0462 | 0.001844 | 0.001819 | -151.8391 | 0.23327 | 8.4027 | 0.002653 | 0.00322 | 0.003167 | 5.9e-05 | 0.001136 | 74.9457 | 1.3407 | 2.5474 | 100.22 | 20.77 | 37.07 | 65.9803 | 85.4432 | 65.3943 | 0.00908 | 1.0428 | 0.024844 | 23.8235 | 0.094825 | 0.13877 | -2.5e-05 | -0.001904 | -0.001904 | -2.9662
2019-03-29 13:42 | 1553881304.5192 | -683 | 0.20932 | 0.000478 | 2e-06 | 94.7071 | 93.8709 | 125.8702 | 21.046 | 0.001857 | 0.001806 | -151.8294 | 0.23305 | 8.4029 | 0.002704 | 0.003462 | 0.003184 | 5.9e-05 | 0.001125 | 74.9022 | 1.3351 | 2.5475 | 100.22 | 20.77 | 37.069 | 66.2908 | 86.6645 | 65.0818 | 0.00908 | 1.0438 | 0.024828 | 23.7866 | 0.079327 | 0.13864 | -3e-05 | -0.002349 | -0.002349 | -2.4025
2019-03-29 13:42 | 1553881305.5186 | -682 | 0.21107 | 0.000479 | 2e-06 | 94.8475 | 93.888 | 125.7577 | 21.0468 | 0.001867 | 0.001794 | -151.842 | 0.23303 | 8.402 | 0.002711 | 0.004463 | 0.004242 | 6.1e-05 | 0.001124 | 74.9398 | 1.3666 | 2.5475 | 100.22 | 20.77 | 37.066 | 66.5436 | 86.8285 | 64.7828 | 0.009079 | 1.0447 | 0.024828 | 23.7655 | 0.09095 | 0.13863 | -3e-05 | -0.002403 | -0.002403 | -2.3479
2019-03-29 13:42 | 1553881306.5196 | -681 | 0.21137 | 0.000479 | 2e-06 | 95.0415 | 94.03 | 125.9751 | 21.0459 | 0.0019 | 0.001778 | -151.857 | 0.23254 | 8.4036 | 0.002686 | 0.004509 | 0.004301 | 5.9e-05 | 0.001126 | 74.9081 | 1.3651 | 2.5474 | 100.22 | 20.77 | 37.064 | 67.3259 | 86.2367 | 64.3931 | 0.009078 | 1.0459 | 0.024864 | 23.7723 | 0.095707 | 0.13834 | -3e-05 | -0.002378 | -0.002378 | -2.3676
2019-03-29 13:42 | 1553881307.5183 | -680 | 0.21135 | 0.00048 | 1e-06 | 95.1918 | 94.2532 | 125.8648 | 21.0462 | 0.001929 | 0.001696 | -151.8546 | 0.23326 | 8.401 | 0.002705 | 0.00398 | 0.003778 | 6.1e-05 | 0.001127 | 74.9394 | 1.3457 | 2.5473 | 100.22 | 20.77 | 37.06 | 68.0322 | 86.6857 | 62.438 | 0.009077 | 1.052 | 0.024925 | 23.6929 | 0.088631 | 0.13877 | -3.1e-05 | -0.002456 | -0.002456 | -2.3
2019-03-29 13:42 | 1553881308.5194 | -679 | 0.21147 | 0.000478 | 0 | 95.3094 | 94.2487 | 125.9678 | 21.0471 | 0.001893 | 0.00173 | -151.8397 | 0.23347 | 8.4048 | 0.002668 | 0.004131 | 0.003986 | 5.9e-05 | 0.001127 | 74.9506 | 1.3535 | 2.5473 | 100.22 | 20.77 | 37.059 | 67.179 | 85.7834 | 63.2552 | 0.009077 | 1.0495 | 0.024905 | 23.7315 | 0.1001 | 0.13889 | -2.8e-05 | -0.002155 | -0.002155 | -2.6232
2019-03-29 13:42 | 1553881309.5188 | -678 | 0.2116 | 0.000473 | 1e-06 | 95.4902 | 94.7064 | 125.6615 | 21.0461 | 0.001922 | 0.001831 | -151.8241 | 0.23362 | 8.4015 | 0.002727 | 0.003873 | 0.003673 | 5.9e-05 | 0.001134 | 74.9468 | 1.3583 | 2.5474 | 100.22 | 20.77 | 37.056 | 67.8645 | 87.2136 | 65.6766 | 0.009077 | 1.042 | 0.024786 | 23.7878 | 0.073721 | 0.13898 | -2.1e-05 | -0.001653 | -0.001653 | -3.4216
2019-03-29 13:42 | 1553881310.5194 | -677 | 0.2117 | 0.000466 | 1e-06 | 95.4797 | 94.8395 | 125.3061 | 21.0465 | 0.00194 | 0.001769 | -151.8385 | 0.23377 | 8.4026 | 0.002716 | 0.004152 | 0.003856 | 5.9e-05 | 0.001133 | 74.9169 | 1.3728 | 2.5472 | 100.22 | 20.77 | 37.052 | 68.3048 | 86.9468 | 64.1932 | 0.009076 | 1.0465 | 0.024805 | 23.7021 | 0.060177 | 0.13907 | -1.3e-05 | -0.001019 | -0.001019 | -5.557
2019-03-29 13:42 | 1553881311.5188 | -676 | 0.21181 | 0.000456 | 1e-06 | 95.7946 | 94.9238 | 125.0557 | 21.0461 | 0.001894 | 0.001764 | -151.8397 | 0.23306 | 8.4017 | 0.002686 | 0.004304 | 0.004222 | 5.9e-05 | 0.001135 | 74.9375 | 1.3895 | 2.5475 | 100.22 | 20.77 | 37.05 | 67.1944 | 86.215 | 64.0639 | 0.009075 | 1.0469 | 0.024785 | 23.6738 | 0.081682 | 0.13865 | 0 | -4.5e-05 | -4.5e-05 | -126.4884
2019-03-29 13:42 | 1553881312.5184 | -675 | 0.2105 | 0.000433 | 1e-06 | 95.9221 | 95.0941 | 125.0813 | 21.047 | 0.00187 | 0.001753 | -151.8315 | 0.23356 | 8.4035 | 0.002705 | 0.003615 | 0.003405 | 5.9e-05 | 0.001135 | 74.94 | 1.3271 | 2.5474 | 100.22 | 20.77 | 37.046 | 66.6242 | 86.6819 | 63.8056 | 0.009074 | 1.0478 | 0.024797 | 23.6672 | 0.077547 | 0.13895 | 2.8e-05 | 0.002056 | 0.002056 | 2.7503
2019-03-29 13:42 | 1553881313.5184 | -674 | 0.20918 | 0.000423 | 1e-06 | 96.0401 | 95.2586 | 125.3341 | 21.0466 | 0.00191 | 0.001776 | -151.8411 | 0.23304 | 8.4009 | 0.002702 | 0.004268 | 0.00406 | 6e-05 | 0.001138 | 74.8413 | 1.4071 | 2.5474 | 100.22 | 20.77 | 37.043 | 67.581 | 86.6197 | 64.3548 | 0.009073 | 1.046 | 0.024802 | 23.7104 | 0.073078 | 0.13863 | 4.2e-05 | 0.003065 | 0.003065 | 1.8412
2019-03-29 13:42 | 1553881314.5177 | -673 | 0.20921 | 0.000406 | 1e-06 | 96.1837 | 95.3067 | 125.6361 | 21.047 | 0.001893 | 0.001813 | -151.8419 | 0.23276 | 8.4027 | 0.002677 | 0.003763 | 0.003556 | 6e-05 | 0.001138 | 74.9278 | 1.3867 | 2.5474 | 100.22 | 20.77 | 37.041 | 67.1733 | 86.0152 | 65.2396 | 0.009073 | 1.0433 | 0.0248 | 23.77 | 0.081932 | 0.13847 | 6.4e-05 | 0.00471 | 0.00471 | 1.1965
2019-03-29 13:42 | 1553881315.5185 | -672 | 0.20922 | 0.000393 | 0 | 96.2119 | 95.5651 | 125.6897 | 21.0478 | 0.001934 | 0.001799 | -151.8205 | 0.23344 | 8.4015 | 0.002718 | 0.004362 | 0.00423 | 6e-05 | 0.001143 | 74.9712 | 1.3476 | 2.5474 | 100.22 | 20.78 | 37.035 | 68.1648 | 86.9996 | 64.9094 | 0.009077 | 1.0443 | 0.024817 | 23.7635 | 0.060334 | 0.13887 | 8e-05 | 0.00596 | 0.00596 | 0.94835
2019-03-29 13:42 | 1553881316.5176 | -671 | 0.20927 | 0.000389 | 1e-06 | 96.3172 | 95.6719 | 125.846 | 21.048 | 0.001854 | 0.001805 | -151.8149 | 0.23228 | 8.4021 | 0.002779 | 0.003933 | 0.003747 | 6e-05 | 0.001141 | 74.935 | 1.3782 | 2.5475 | 100.22 | 20.78 | 37.035 | 66.2357 | 88.4624 | 65.044 | 0.009077 | 1.0439 | 0.024827 | 23.783 | 0.060132 | 0.13819 | 8.5e-05 | 0.006272 | 0.006272 | 0.89671
2019-03-29 13:42 | 1553881317.5176 | -670 | 0.20933 | 0.000382 | 1e-06 | 96.4178 | 95.6511 | 126.0138 | 21.0472 | 0.001851 | 0.001764 | -151.8164 | 0.23345 | 8.4037 | 0.002776 | 0.003727 | 0.003603 | 6.2e-05 | 0.001146 | 74.8724 | 1.3719 | 2.5474 | 100.22 | 20.78 | 37.034 | 66.1498 | 88.3976 | 64.0533 | 0.009077 | 1.047 | 0.02488 | 23.764 | 0.071413 | 0.13888 | 9.4e-05 | 0.006976 | 0.006976 | 0.81021
2019-03-29 13:42 | 1553881318.5176 | -669 | 0.20932 | 0.000383 | 0 | 96.578 | 95.9019 | 126.2394 | 21.048 | 0.001922 | 0.001751 | -151.7994 | 0.23292 | 8.4011 | 0.002697 | 0.003589 | 0.003409 | 6.1e-05 | 0.001145 | 74.923 | 1.3227 | 2.5472 | 100.22 | 20.78 | 37.03 | 67.8615 | 86.501 | 63.7644 | 0.009076 | 1.0479 | 0.024913 | 23.775 | 0.062831 | 0.13856 | 9.3e-05 | 0.006956 | 0.006956 | 0.81068
2019-03-29 13:42 | 1553881319.5186 | -668 | 0.20935 | 0.000381 | 0 | 96.6226 | 95.9446 | 126.5058 | 21.0483 | 0.001896 | 0.001742 | -151.7973 | 0.23381 | 8.402 | 0.00266 | 0.004476 | 0.004226 | 6e-05 | 0.001138 | 74.9398 | 1.3378 | 2.5473 | 100.22 | 20.78 | 37.029 | 67.2525 | 85.5918 | 63.5358 | 0.009076 | 1.0486 | 0.024948 | 23.792 | 0.062982 | 0.13909 | 9.5e-05 | 0.007134 | 0.007134 | 0.79351
2019-03-29 13:42 | 1553881320.5183 | -667 | 0.20934 | 0.00038 | 0 | 96.6995 | 96.1023 | 126.6293 | 21.048 | 0.00185 | 0.001774 | -151.7867 | 0.23255 | 8.4033 | 0.002711 | 0.003129 | 0.002993 | 5.7e-05 | 0.001139 | 74.949 | 1.3234 | 2.5475 | 100.22 | 20.78 | 37.024 | 66.1342 | 86.8241 | 64.3032 | 0.009074 | 1.0462 | 0.024932 | 23.8307 | 0.055413 | 0.13834 | 9.7e-05 | 0.007245 | 0.007245 | 0.77721
2019-03-29 13:42 | 1553881321.5176 | -666 | 0.20935 | 0.000384 | 0 | 96.812 | 96.0514 | 126.7018 | 21.0487 | 0.001867 | 0.001768 | -151.7971 | 0.23308 | 8.4025 | 0.002746 | 0.003635 | 0.003522 | 5.8e-05 | 0.001141 | 74.9815 | 1.3459 | 2.5475 | 100.22 | 20.78 | 37.022 | 66.5368 | 87.6642 | 64.1634 | 0.009074 | 1.0466 | 0.024944 | 23.8326 | 0.070552 | 0.13866 | 9.3e-05 | 0.006961 | 0.006961 | 0.81079
2019-03-29 13:42 | 1553881322.5177 | -665 | 0.20937 | 0.000389 | 0 | 96.918 | 96.1487 | 126.9903 | 21.0487 | 0.001915 | 0.001774 | -151.7846 | 0.23379 | 8.401 | 0.002626 | 0.00377 | 0.003586 | 5.7e-05 | 0.001143 | 74.9122 | 1.3788 | 2.5473 | 100.22 | 20.78 | 37.015 | 67.6985 | 84.7908 | 64.3155 | 0.009072 | 1.0462 | 0.024967 | 23.8651 | 0.071284 | 0.13908 | 8.5e-05 | 0.006371 | 0.006371 | 0.88853
2019-03-29 13:42 | 1553881323.519 | -664 | 0.20936 | 0.000395 | 0 | 96.9578 | 96.2651 | 126.844 | 21.0489 | 0.001872 | 0.001792 | -151.7736 | 0.2335 | 8.4018 | 0.00265 | 0.003912 | 0.003771 | 5.9e-05 | 0.001149 | 74.8975 | 1.3645 | 2.5473 | 100.22 | 20.78 | 37.015 | 66.6635 | 85.3724 | 64.742 | 0.009072 | 1.0448 | 0.024937 | 23.8665 | 0.064137 | 0.13891 | 7.8e-05 | 0.00586 | 0.00586 | 0.96484
2019-03-29 13:42 | 1553881324.5188 | -663 | 0.20938 | 0.000392 | 0 | 97.059 | 96.4514 | 126.4797 | 21.0486 | 0.001883 | 0.001769 | -151.7568 | 0.23259 | 8.4017 | 0.002608 | 0.004004 | 0.003748 | 6.1e-05 | 0.001153 | 74.9091 | 1.3204 | 2.5473 | 100.22 | 20.78 | 37.011 | 66.9191 | 84.3443 | 64.1921 | 0.009071 | 1.0466 | 0.024921 | 23.8128 | 0.056168 | 0.13837 | 8.1e-05 | 0.006061 | 0.006061 | 0.92917
2019-03-29 13:42 | 1553881325.5187 | -662 | 0.2094 | 0.0004 | 0 | 97.1649 | 96.662 | 126.7163 | 21.051 | 0.001901 | 0.001787 | -151.772 | 0.23402 | 8.4026 | 0.002672 | 0.003476 | 0.003375 | 5.9e-05 | 0.001149 | 74.972 | 1.3473 | 2.5474 | 100.22 | 20.78 | 37.011 | 67.373 | 85.9051 | 64.6262 | 0.009071 | 1.0452 | 0.024929 | 23.8503 | 0.046422 | 0.13922 | 7.2e-05 | 0.005359 | 0.005359 | 1.0573
2019-03-29 13:42 | 1553881326.5178 | -661 | 0.20935 | 0.000402 | 0 | 97.2293 | 96.5685 | 126.5398 | 21.0503 | 0.001867 | 0.001762 | -151.7729 | 0.23262 | 8.4025 | 0.002651 | 0.003879 | 0.003609 | 5.9e-05 | 0.001144 | 74.966 | 1.3148 | 2.5473 | 100.22 | 20.78 | 37.009 | 66.5366 | 85.3908 | 64.0265 | 0.009071 | 1.0471 | 0.024933 | 23.8126 | 0.061001 | 0.13839 | 6.9e-05 | 0.005186 | 0.005186 | 1.0861
2019-03-29 13:42 | 1553881327.5189 | -660 | 0.20947 | 0.000403 | 0 | 97.3987 | 96.9111 | 126.3035 | 21.0501 | 0.001914 | 0.001762 | -151.7766 | 0.23295 | 8.402 | 0.00262 | 0.00378 | 0.003573 | 6e-05 | 0.001139 | 74.9583 | 1.299 | 2.5474 | 100.22 | 20.78 | 37.008 | 67.6866 | 84.6479 | 64.01 | 0.00907 | 1.0471 | 0.024911 | 23.7897 | 0.044892 | 0.13858 | 6.8e-05 | 0.005058 | 0.005058 | 1.1152
2019-03-29 13:42 | 1553881328.5173 | -659 | 0.20938 | 0.000411 | 0 | 97.4804 | 97.0715 | 125.8502 | 21.0506 | 0.001973 | 0.001763 | -151.7989 | 0.23436 | 8.4042 | 0.002607 | 0.003054 | 0.00288 | 6e-05 | 0.001134 | 74.9361 | 1.3175 | 2.5474 | 100.22 | 20.78 | 37.004 | 69.0965 | 84.3314 | 64.0494 | 0.009069 | 1.047 | 0.024864 | 23.7484 | 0.0376 | 0.13942 | 5.8e-05 | 0.004375 | 0.004375 | 1.2969
2019-03-29 13:42 | 1553881329.518 | -658 | 0.2094 | 0.000409 | 0 | 97.561 | 97.1367 | 125.4007 | 21.05 | 0.001869 | 0.001747 | -151.7899 | 0.23377 | 8.4025 | 0.002652 | 0.004154 | 0.003913 | 6.1e-05 | 0.00113 | 74.9423 | 1.3525 | 2.5473 | 100.22 | 20.78 | 37.003 | 66.6035 | 85.404 | 63.6561 | 0.009069 | 1.0482 | 0.024834 | 23.6921 | 0.038981 | 0.13907 | 6.1e-05 | 0.004538 | 0.004538 | 1.2474
2019-03-29 13:42 | 1553881330.517 | -657 | 0.20941 | 0.00041 | 0 | 97.6579 | 97.1203 | 125.3385 | 21.0501 | 0.001885 | 0.001815 | -151.7981 | 0.23289 | 8.4034 | 0.002598 | 0.003478 | 0.003343 | 5.9e-05 | 0.001131 | 74.9764 | 1.3434 | 2.5474 | 100.22 | 20.78 | 36.999 | 66.9672 | 84.115 | 65.299 | 0.009068 | 1.0431 | 0.024768 | 23.7439 | 0.049377 | 0.13855 | 5.9e-05 | 0.004356 | 0.004356 | 1.2946
2019-03-29 13:42 | 1553881331.517 | -656 | 0.20939 | 0.000417 | 0 | 97.725 | 97.2446 | 125.1747 | 21.0515 | 0.001884 | 0.001794 | -151.7933 | 0.2339 | 8.4016 | 0.002607 | 0.003555 | 0.003352 | 6.2e-05 | 0.001133 | 74.9775 | 1.3598 | 2.5474 | 100.22 | 20.78 | 36.998 | 66.9632 | 84.3207 | 64.7922 | 0.009068 | 1.0447 | 0.02477 | 23.7106 | 0.044082 | 0.13915 | 5e-05 | 0.003713 | 0.003713 | 1.5252
2019-03-29 13:42 | 1553881332.5182 | -655 | 0.20943 | 0.000423 | 0 | 97.8378 | 97.4302 | 124.9679 | 21.0502 | 0.001914 | 0.001763 | -151.7934 | 0.23322 | 8.4039 | 0.002613 | 0.004055 | 0.003825 | 6.1e-05 | 0.001127 | 74.9689 | 1.3416 | 2.5474 | 100.22 | 20.78 | 36.998 | 67.6663 | 84.4763 | 64.0383 | 0.009068 | 1.047 | 0.024777 | 23.6646 | 0.037344 | 0.13874 | 4.3e-05 | 0.003217 | 0.003217 | 1.7554
2019-03-29 13:42 | 1553881333.5172 | -654 | 0.20943 | 0.000417 | 0 | 97.9947 | 97.4903 | 124.744 | 21.0502 | 0.001877 | 0.001782 | -151.8128 | 0.23406 | 8.4021 | 0.002546 | 0.00431 | 0.004117 | 6.3e-05 | 0.001129 | 74.8406 | 1.3229 | 2.5473 | 100.22 | 20.78 | 36.995 | 66.7816 | 82.8614 | 64.4974 | 0.009067 | 1.0456 | 0.024738 | 23.6595 | 0.04616 | 0.13924 | 5.1e-05 | 0.003777 | 0.003777 | 1.5006
2019-03-29 13:42 | 1553881334.5171 | -653 | 0.20941 | 0.000425 | 0 | 98.1348 | 97.7367 | 124.8482 | 21.0509 | 0.001911 | 0.001793 | -151.8159 | 0.23389 | 8.4029 | 0.002638 | 0.003466 | 0.003308 | 6.1e-05 | 0.001126 | 74.9628 | 1.3653 | 2.5474 | 100.22 | 20.78 | 36.995 | 67.5953 | 85.0856 | 64.775 | 0.009067 | 1.0447 | 0.024739 | 23.6791 | 0.036358 | 0.13914 | 4e-05 | 0.002987 | 0.002987 | 1.8957
2019-03-29 13:42 | 1553881335.5184 | -652 | 0.2094 | 0.000426 | 0 | 98.2443 | 97.8753 | 124.3543 | 21.0515 | 0.001844 | 0.001772 | -151.8113 | 0.23316 | 8.403 | 0.002597 | 0.003783 | 0.003592 | 6.2e-05 | 0.001129 | 74.9567 | 1.3496 | 2.5473 | 100.22 | 20.78 | 36.988 | 66.0003 | 84.1 | 64.2597 | 0.009065 | 1.0463 | 0.024708 | 23.6142 | 0.033666 | 0.13871 | 3.8e-05 | 0.002797 | 0.002797 | 2.0184
2019-03-29 13:42 | 1553881336.5186 | -651 | 0.20937 | 0.000425 | 0 | 98.3365 | 97.9493 | 124.3771 | 21.0522 | 0.001867 | 0.00178 | -151.8022 | 0.23274 | 8.4024 | 0.002664 | 0.003688 | 0.003594 | 5.9e-05 | 0.001137 | 74.876 | 1.3351 | 2.5474 | 100.22 | 20.78 | 36.988 | 66.5565 | 85.7017 | 64.4507 | 0.009065 | 1.0457 | 0.024704 | 23.623 | 0.03529 | 0.13846 | 4.1e-05 | 0.00306 | 0.00306 | 1.8415
2019-03-29 13:42 | 1553881337.5173 | -650 | 0.20942 | 0.000427 | 0 | 98.5332 | 98.1473 | 124.4991 | 21.0516 | 0.001908 | 0.00176 | -151.8083 | 0.23351 | 8.4039 | 0.00263 | 0.004185 | 0.003904 | 5.8e-05 | 0.001138 | 74.9378 | 1.3059 | 2.5473 | 100.22 | 20.78 | 36.981 | 67.5241 | 84.8788 | 63.9677 | 0.009064 | 1.0472 | 0.024734 | 23.6177 | 0.035102 | 0.13892 | 3.8e-05 | 0.002794 | 0.002794 | 2.0236
2019-03-29 13:42 | 1553881338.5173 | -649 | 0.2094 | 0.000428 | 0 | 98.7239 | 98.2684 | 124.943 | 21.0523 | 0.001896 | 0.001808 | -151.8161 | 0.23387 | 8.4007 | 0.002628 | 0.004159 | 0.003966 | 5.8e-05 | 0.00114 | 74.8261 | 1.3251 | 2.5473 | 100.22 | 20.78 | 36.981 | 67.2511 | 84.8325 | 65.1339 | 0.009064 | 1.0436 | 0.024735 | 23.7007 | 0.04136 | 0.13913 | 3.6e-05 | 0.002643 | 0.002643 | 2.1422
2019-03-29 13:42 | 1553881339.5186 | -648 | 0.2094 | 0.00043 | 0 | 98.7523 | 98.424 | 125.1169 | 21.0521 | 0.001824 | 0.001847 | -151.8176 | 0.23267 | 8.4012 | 0.002555 | 0.00412 | 0.003997 | 6e-05 | 0.00114 | 74.881 | 1.3385 | 2.5472 | 100.22 | 20.78 | 36.979 | 65.5092 | 83.0815 | 66.0697 | 0.009063 | 1.0408 | 0.024718 | 23.7499 | 0.029782 | 0.13841 | 3.3e-05 | 0.002413 | 0.002413 | 2.3348
2019-03-29 13:42 | 1553881340.5183 | -647 | 0.20941 | 0.000435 | 0 | 98.9184 | 98.607 | 124.9414 | 21.0525 | 0.001905 | 0.001799 | -151.8155 | 0.23341 | 8.4015 | 0.002642 | 0.004488 | 0.004287 | 5.8e-05 | 0.001141 | 74.9649 | 1.3212 | 2.5473 | 100.22 | 20.78 | 36.976 | 67.4673 | 85.1621 | 64.9118 | 0.009063 | 1.0443 | 0.024743 | 23.6927 | 0.028207 | 0.13886 | 2.8e-05 | 0.002077 | 0.002077 | 2.7211
2019-03-29 13:42 | 1553881341.5187 | -646 | 0.2094 | 0.000434 | 0 | 99.1247 | 98.6643 | 124.9341 | 21.0531 | 0.001872 | 0.00176 | -151.815 | 0.23296 | 8.4025 | 0.002597 | 0.004238 | 0.004005 | 5.7e-05 | 0.001134 | 74.9381 | 1.3658 | 2.5474 | 100.22 | 20.78 | 36.972 | 66.6759 | 84.0974 | 63.9676 | 0.009062 | 1.0472 | 0.024777 | 23.6589 | 0.041637 | 0.13859 | 2.9e-05 | 0.002124 | 0.002124 | 2.6551
2019-03-29 13:42 | 1553881342.5174 | -645 | 0.20941 | 0.000432 | 0 | 99.2405 | 98.9268 | 124.7937 | 21.0531 | 0.001895 | 0.001724 | -151.8195 | 0.23336 | 8.4026 | 0.002602 | 0.003895 | 0.003789 | 6e-05 | 0.001135 | 74.9068 | 1.3807 | 2.5474 | 100.22 | 20.78 | 36.97 | 67.2183 | 84.2069 | 63.108 | 0.009061 | 1.0499 | 0.024794 | 23.6155 | 0.028323 | 0.13882 | 3.1e-05 | 0.002334 | 0.002334 | 2.4203
2019-03-29 13:42 | 1553881343.5175 | -644 | 0.20942 | 0.000431 | 0 | 99.3091 | 99.1957 | 124.7226 | 21.0532 | 0.001885 | 0.001753 | -151.8147 | 0.23271 | 8.4037 | 0.002603 | 0.003573 | 0.003449 | 5.8e-05 | 0.001132 | 74.9178 | 1.3684 | 2.5474 | 100.22 | 20.78 | 36.967 | 66.9868 | 84.231 | 63.7945 | 0.00906 | 1.0478 | 0.024762 | 23.6328 | 0.010216 | 0.13844 | 3.2e-05 | 0.002339 | 0.002339 | 2.4094
2019-03-29 13:42 | 1553881344.5178 | -643 | 0.20942 | 0.000436 | 0 | 99.4812 | 99.1278 | 124.8918 | 21.0534 | 0.001951 | 0.001777 | -151.8124 | 0.23293 | 8.402 | 0.002571 | 0.004468 | 0.004225 | 6.1e-05 | 0.001128 | 74.9457 | 1.3506 | 2.5475 | 100.22 | 20.78 | 36.966 | 68.5611 | 83.4612 | 64.3816 | 0.00906 | 1.046 | 0.024757 | 23.6694 | 0.031831 | 0.13857 | 2.6e-05 | 0.001937 | 0.001937 | 2.9113
2019-03-29 13:42 | 1553881345.5165 | -642 | 0.20941 | 0.000434 | 0 | 99.5922 | 99.3148 | 125.2618 | 21.0524 | 0.00193 | 0.00177 | -151.8182 | 0.23362 | 8.4032 | 0.002594 | 0.003684 | 0.003477 | 6.1e-05 | 0.001127 | 74.9266 | 1.3505 | 2.5474 | 100.22 | 20.78 | 36.964 | 68.0671 | 84.0251 | 64.2185 | 0.00906 | 1.0465 | 0.0248 | 23.6987 | 0.02495 | 0.13898 | 2.9e-05 | 0.002135 | 0.002135 | 2.6497
2019-03-29 13:42 | 1553881346.5175 | -641 | 0.20943 | 0.000441 | 0 | 99.6778 | 99.4629 | 125.2045 | 21.0538 | 0.001957 | 0.001714 | -151.8296 | 0.23346 | 8.401 | 0.00265 | 0.00368 | 0.003541 | 6e-05 | 0.001125 | 74.9906 | 1.3193 | 2.5474 | 100.22 | 20.78 | 36.963 | 68.7227 | 85.3764 | 62.8609 | 0.009059 | 1.0507 | 0.024844 | 23.6456 | 0.019305 | 0.13889 | 2e-05 | 0.001474 | 0.001474 | 3.8338
2019-03-29 13:42 | 1553881347.5165 | -640 | 0.20941 | 0.000441 | 0 | 99.7628 | 99.6595 | 125.2251 | 21.0536 | 0.001896 | 0.001821 | -151.8192 | 0.23288 | 8.3998 | 0.002635 | 0.004233 | 0.004005 | 5.6e-05 | 0.001131 | 74.9238 | 1.3686 | 2.5474 | 100.22 | 20.78 | 36.964 | 67.2514 | 85.0156 | 65.4397 | 0.00906 | 1.0427 | 0.024752 | 23.7381 | 0.00927 | 0.13854 | 2e-05 | 0.00152 | 0.00152 | 3.7106
2019-03-29 13:42 | 1553881348.5178 | -639 | 0.20941 | 0.000439 | 0 | 99.8907 | 99.5756 | 125.3101 | 21.0543 | 0.001908 | 0.001798 | -151.8054 | 0.23356 | 8.4022 | 0.002621 | 0.003666 | 0.003604 | 5.9e-05 | 0.001133 | 74.9425 | 1.3624 | 2.5474 | 100.22 | 20.78 | 36.964 | 67.5354 | 84.6784 | 64.8837 | 0.00906 | 1.0444 | 0.02478 | 23.7267 | 0.02826 | 0.13895 | 2.2e-05 | 0.001638 | 0.001638 | 3.4533
2019-03-29 13:42 | 1553881349.517 | -638 | 0.20942 | 0.000443 | 0 | 99.9631 | 99.8854 | 125.2057 | 21.0542 | 0.001857 | 0.001699 | -151.8084 | 0.2331 | 8.4025 | 0.002752 | 0.003826 | 0.003578 | 6e-05 | 0.001135 | 74.8379 | 1.3093 | 2.5473 | 100.22 | 20.78 | 36.963 | 66.3165 | 87.8226 | 62.5013 | 0.009059 | 1.0518 | 0.024858 | 23.6331 | 0.006951 | 0.13867 | 1.7e-05 | 0.001305 | 0.001305 | 4.3263
2019-03-29 13:42 | 1553881350.5166 | -637 | 0.20942 | 0.000443 | 0 | 99.9329 | 99.7189 | 125.112 | 21.0542 | 0.001871 | 0.001753 | -151.8032 | 0.23336 | 8.4033 | 0.002679 | 0.003463 | 0.003318 | 5.7e-05 | 0.001138 | 74.7907 | 1.3619 | 2.5473 | 100.22 | 20.78 | 36.963 | 66.6475 | 86.073 | 63.8055 | 0.009059 | 1.0478 | 0.0248 | 23.6701 | 0.019172 | 0.13882 | 1.8e-05 | 0.001317 | 0.001317 | 4.2912
2019-03-29 13:42 | 1553881351.518 | -636 | 0.20941 | 0.000442 | 0 | 100.1232 | 99.8988 | 125.5223 | 21.0544 | 0.001908 | 0.001779 | -151.8036 | 0.23332 | 8.4016 | 0.002662 | 0.004238 | 0.003975 | 5.6e-05 | 0.001143 | 74.8891 | 1.3341 | 2.5474 | 100.22 | 20.78 | 36.964 | 67.5362 | 85.6666 | 64.4319 | 0.00906 | 1.0458 | 0.024818 | 23.7309 | 0.020077 | 0.13881 | 1.9e-05 | 0.001401 | 0.001401 | 4.0313
2019-03-29 13:43 | 1553881352.5169 | -635 | 0.20941 | 0.000441 | 0 | 100.1511 | 99.7861 | 125.6394 | 21.0552 | 0.001872 | 0.001823 | -151.807 | 0.23324 | 8.3991 | 0.00261 | 0.003994 | 0.003787 | 6.1e-05 | 0.00115 | 74.8998 | 1.3393 | 2.5475 | 100.22 | 20.78 | 36.965 | 66.6617 | 84.4032 | 65.4873 | 0.00906 | 1.0425 | 0.024791 | 23.779 | 0.032653 | 0.13875 | 1.9e-05 | 0.001385 | 0.001385 | 4.0785
2019-03-29 13:43 | 1553881353.5168 | -634 | 0.20943 | 0.000446 | 0 | 100.2379 | 100.0708 | 125.9831 | 21.054 | 0.001932 | 0.001816 | -151.7905 | 0.23295 | 8.4022 | 0.002686 | 0.003656 | 0.003454 | 5.9e-05 | 0.001149 | 74.9389 | 1.3157 | 2.5475 | 100.22 | 20.78 | 36.966 | 68.11 | 86.2335 | 65.3278 | 0.00906 | 1.043 | 0.02483 | 23.8059 | 0.014925 | 0.13858 | 1.3e-05 | 0.000933 | 0.000933 | 6.0463
2019-03-29 13:43 | 1553881354.5178 | -633 | 0.2094 | 0.000445 | 0 | 100.2846 | 99.9885 | 126.0089 | 21.0545 | 0.001854 | 0.001788 | -151.7834 | 0.23374 | 8.403 | 0.002568 | 0.003074 | 0.002884 | 5.6e-05 | 0.001144 | 74.9712 | 1.2999 | 2.5472 | 100.22 | 20.78 | 36.966 | 66.2351 | 83.3851 | 64.6498 | 0.00906 | 1.0451 | 0.024858 | 23.7845 | 0.026446 | 0.13905 | 1.5e-05 | 0.001145 | 0.001145 | 4.9413
2019-03-29 13:43 | 1553881355.5168 | -632 | 0.20942 | 0.000443 | 0 | 100.3771 | 100.1861 | 125.8742 | 21.0556 | 0.001894 | 0.00178 | -151.7646 | 0.23375 | 8.4027 | 0.002565 | 0.00355 | 0.003421 | 5.9e-05 | 0.001135 | 74.9637 | 1.3416 | 2.5472 | 100.22 | 20.78 | 36.966 | 67.1935 | 83.3174 | 64.4533 | 0.00906 | 1.0457 | 0.024852 | 23.7649 | 0.017032 | 0.13906 | 1.7e-05 | 0.001238 | 0.001238 | 4.5715
2019-03-29 13:43 | 1553881356.5178 | -631 | 0.20942 | 0.000444 | 0 | 100.4968 | 100.2875 | 125.7843 | 21.0559 | 0.001882 | 0.001815 | -151.7678 | 0.233 | 8.4029 | 0.00265 | 0.004381 | 0.004151 | 5.7e-05 | 0.00113 | 74.9219 | 1.3258 | 2.5475 | 100.22 | 20.78 | 36.966 | 66.9105 | 85.3753 | 65.3077 | 0.00906 | 1.0431 | 0.024812 | 23.7864 | 0.01865 | 0.13861 | 1.7e-05 | 0.001243 | 0.001243 | 4.537
2019-03-29 13:43 | 1553881357.5169 | -630 | 0.20942 | 0.000447 | 0 | 100.489 | 100.3738 | 125.8063 | 21.0553 | 0.001848 | 0.001737 | -151.7632 | 0.23302 | 8.4029 | 0.002635 | 0.003154 | 0.00304 | 5.9e-05 | 0.001124 | 74.9676 | 1.3357 | 2.5474 | 100.22 | 20.78 | 36.966 | 66.0954 | 85.0103 | 63.427 | 0.00906 | 1.0489 | 0.024883 | 23.7223 | 0.01026 | 0.13862 | 1.2e-05 | 0.000884 | 0.000884 | 6.3797
2019-03-29 13:43 | 1553881358.58 | -629 | 0.20942 | 0.000443 | 0 | 100.5103 | 100.368 | 125.826 | 21.0559 | 0.001882 | 0.001812 | -151.76 | 0.2326 | 8.4035 | 0.002669 | 0.003987 | 0.003791 | 6e-05 | 0.001126 | 74.9773 | 1.3491 | 2.5473 | 100.22 | 20.78 | 36.966 | 66.9214 | 85.8256 | 65.2264 | 0.00906 | 1.0434 | 0.024819 | 23.7875 | 0.012668 | 0.13837 | 1.7e-05 | 0.00124 | 0.00124 | 4.5425
2019-03-29 13:43 | 1553881359.5297 | -628 | 0.2094 | 0.000447 | 0 | 100.5947 | 100.4305 | 126.1241 | 21.0564 | 0.001945 | 0.001789 | -151.7816 | 0.23254 | 8.4014 | 0.002635 | 0.003588 | 0.003404 | 5.7e-05 | 0.001129 | 75.0127 | 1.3255 | 2.5473 | 100.22 | 20.78 | 36.964 | 68.4219 | 85.0066 | 64.676 | 0.00906 | 1.0451 | 0.024868 | 23.7963 | 0.014613 | 0.13834 | 1.3e-05 | 0.000969 | 0.000969 | 5.8114
2019-03-29 13:43 | 1553881360.5295 | -627 | 0.20966 | 0.000449 | 0 | 100.6192 | 100.483 | 126.3058 | 21.0566 | 0.001887 | 0.001816 | -151.7633 | 0.23329 | 8.4033 | 0.002625 | 0.004134 | 0.003924 | 5.8e-05 | 0.001133 | 74.9933 | 1.3218 | 2.5475 | 100.22 | 20.78 | 36.963 | 67.0266 | 84.7783 | 65.3278 | 0.009059 | 1.043 | 0.024862 | 23.8364 | 0.012115 | 0.13878 | 1e-05 | 0.000697 | 0.000697 | 8.1032
2019-03-29 13:43 | 1553881361.5292 | -626 | 0.20939 | 0.000447 | 0 | 100.6638 | 100.5431 | 126.2312 | 21.0565 | 0.001922 | 0.001797 | -151.7506 | 0.23253 | 8.4011 | 0.002617 | 0.004104 | 0.00388 | 5.8e-05 | 0.001128 | 75.0008 | 1.3364 | 2.5473 | 100.22 | 20.78 | 36.961 | 67.8676 | 84.5721 | 64.875 | 0.009059 | 1.0444 | 0.024872 | 23.8134 | 0.010734 | 0.13833 | 1.2e-05 | 0.000953 | 0.000953 | 5.9098
2019-03-29 13:43 | 1553881362.5292 | -625 | 0.20949 | 0.000445 | 0 | 100.7695 | 100.5138 | 126.3976 | 21.0571 | 0.001837 | 0.001765 | -151.7534 | 0.23271 | 8.402 | 0.002664 | 0.003875 | 0.003649 | 5.7e-05 | 0.001126 | 75.0073 | 1.3153 | 2.5475 | 100.22 | 20.79 | 36.953 | 65.8239 | 85.7097 | 64.102 | 0.009062 | 1.0468 | 0.024916 | 23.8018 | 0.022729 | 0.13844 | 1.5e-05 | 0.001126 | 0.001126 | 5.0026
2019-03-29 13:43 | 1553881363.5307 | -624 | 0.2094 | 0.000447 | 0 | 100.7521 | 100.4564 | 126.5239 | 21.0564 | 0.001887 | 0.001797 | -151.7525 | 0.23345 | 8.4018 | 0.002614 | 0.003829 | 0.003637 | 5.9e-05 | 0.001124 | 74.9743 | 1.343 | 2.5471 | 100.22 | 20.79 | 36.953 | 67.0296 | 84.4994 | 64.8711 | 0.009062 | 1.0444 | 0.024901 | 23.8409 | 0.026292 | 0.13888 | 1.2e-05 | 0.000913 | 0.000913 | 6.1886
2019-03-29 13:43 | 1553881364.5293 | -623 | 0.20946 | 0.000446 | 0 | 100.7494 | 100.8745 | 126.3859 | 21.0575 | 0.001939 | 0.001745 | -151.7557 | 0.23347 | 8.4025 | 0.002637 | 0.003738 | 0.003604 | 5.8e-05 | 0.001126 | 74.9553 | 1.3537 | 2.5474 | 100.22 | 20.79 | 36.951 | 68.2818 | 85.0591 | 63.6251 | 0.009062 | 1.0483 | 0.024933 | 23.7839 | -0.011095 | 0.13889 | 1.4e-05 | 0.001082 | 0.001082 | 5.2262
2019-03-29 13:43 | 1553881365.5303 | -622 | 0.2095 | 0.000442 | 0 | 100.8653 | 100.5015 | 126.4708 | 21.0573 | 0.001899 | 0.001743 | -151.7723 | 0.23394 | 8.4013 | 0.002678 | 0.004349 | 0.004117 | 5.6e-05 | 0.001131 | 74.9041 | 1.3378 | 2.5472 | 100.22 | 20.79 | 36.948 | 67.3127 | 86.0502 | 63.5751 | 0.009061 | 1.0485 | 0.024943 | 23.7901 | 0.032321 | 0.13917 | 1.9e-05 | 0.00136 | 0.00136 | 4.1655
2019-03-29 13:43 | 1553881366.5294 | -621 | 0.20947 | 0.000444 | 0 | 100.8554 | 100.666 | 126.7006 | 21.058 | 0.001895 | 0.00179 | -151.7711 | 0.23363 | 8.4014 | 0.00265 | 0.004313 | 0.004167 | 5.9e-05 | 0.001132 | 74.8454 | 1.3506 | 2.5473 | 100.22 | 20.79 | 36.943 | 67.2372 | 85.3766 | 64.6938 | 0.00906 | 1.045 | 0.024924 | 23.8513 | 0.016816 | 0.13899 | 1.6e-05 | 0.00119 | 0.00119 | 4.7552
2019-03-29 13:43 | 1553881367.5304 | -620 | 0.20943 | 0.000453 | 0 | 100.932 | 100.8281 | 126.5718 | 21.0578 | 0.001929 | 0.001733 | -151.7725 | 0.23345 | 8.4023 | 0.002565 | 0.004014 | 0.003833 | 5.8e-05 | 0.001131 | 74.8776 | 1.3653 | 2.5475 | 100.22 | 20.79 | 36.94 | 68.0381 | 83.3368 | 63.3303 | 0.009059 | 1.0492 | 0.024962 | 23.791 | 0.009217 | 0.13888 | 5e-06 | 0.00039 | 0.00039 | 14.4791
2019-03-29 13:43 | 1553881368.5293 | -619 | 0.20945 | 0.000445 | 0 | 100.8967 | 100.7448 | 126.1953 | 21.0579 | 0.001897 | 0.001799 | -151.7718 | 0.23368 | 8.4013 | 0.002584 | 0.00357 | 0.003413 | 5.7e-05 | 0.001131 | 74.8487 | 1.3316 | 2.5473 | 100.22 | 20.79 | 36.936 | 67.2799 | 83.7808 | 64.9184 | 0.009058 | 1.0443 | 0.024866 | 23.8116 | 0.013477 | 0.13901 | 1.5e-05 | 0.001127 | 0.001127 | 5.0222
2019-03-29 13:43 | 1553881369.5295 | -618 | 0.20949 | 0.000449 | 0 | 100.9961 | 100.7165 | 126.2157 | 21.0581 | 0.001942 | 0.001755 | -151.7495 | 0.23353 | 8.4011 | 0.002608 | 0.003674 | 0.00343 | 5.8e-05 | 0.001135 | 74.8582 | 1.3316 | 2.5473 | 100.22 | 20.79 | 36.934 | 68.3526 | 84.3685 | 63.8623 | 0.009058 | 1.0476 | 0.024907 | 23.7763 | 0.024798 | 0.13892 | 1e-05 | 0.000764 | 0.000764 | 7.4055
2019-03-29 13:43 | 1553881370.5305 | -617 | 0.20945 | 0.000449 | 0 | 100.9925 | 100.8216 | 126.3016 | 21.0585 | 0.00183 | 0.001785 | -151.7589 | 0.23248 | 8.4027 | 0.002597 | 0.004387 | 0.004245 | 5.7e-05 | 0.00114 | 74.9032 | 1.3676 | 2.5473 | 100.22 | 20.79 | 36.931 | 65.6677 | 84.0974 | 64.5789 | 0.009057 | 1.0454 | 0.024889 | 23.8096 | 0.015149 | 0.1383 | 9e-06 | 0.000697 | 0.000697 | 8.0762
2019-03-29 13:43 | 1553881371.5306 | -616 | 0.20946 | 0.00045 | 0 | 101.0368 | 101.0666 | 126.1873 | 21.0583 | 0.001959 | 0.001821 | -151.762 | 0.23276 | 8.4029 | 0.002647 | 0.004166 | 0.004019 | 6e-05 | 0.001141 | 74.8856 | 1.39 | 2.5474 | 100.22 | 20.79 | 36.929 | 68.774 | 85.2942 | 65.4568 | 0.009057 | 1.0426 | 0.024846 | 23.8297 | -0.002641 | 0.13847 | 9e-06 | 0.000723 | 0.000723 | 7.7949
2019-03-29 13:43 | 1553881372.5295 | -615 | 0.20944 | 0.000452 | 0 | 101.1371 | 101.0144 | 126.0016 | 21.0592 | 0.001882 | 0.001768 | -151.7587 | 0.23377 | 8.4004 | 0.002675 | 0.004331 | 0.004166 | 6e-05 | 0.001143 | 74.9178 | 1.3952 | 2.5474 | 100.22 | 20.79 | 36.928 | 66.9246 | 85.9677 | 64.1837 | 0.009056 | 1.0466 | 0.024874 | 23.7674 | 0.010854 | 0.13907 | 6e-06 | 0.000464 | 0.000464 | 12.2023
2019-03-29 13:43 | 1553881373.5299 | -614 | 0.20946 | 0.000449 | 0 | 101.1434 | 100.9571 | 125.7167 | 21.0588 | 0.00188 | 0.001795 | -151.7493 | 0.23377 | 8.4027 | 0.002659 | 0.004481 | 0.004299 | 5.8e-05 | 0.001135 | 74.9436 | 1.3646 | 2.5473 | 100.22 | 20.79 | 36.928 | 66.8766 | 85.585 | 64.8209 | 0.009056 | 1.0446 | 0.024823 | 23.7629 | 0.016488 | 0.13907 | 9e-06 | 0.000666 | 0.000666 | 8.4963
2019-03-29 13:43 | 1553881374.529 | -613 | 0.20947 | 0.000452 | 0 | 101.1212 | 100.9462 | 125.9239 | 21.0589 | 0.001836 | 0.001784 | -151.7544 | 0.23315 | 8.4027 | 0.00266 | 0.00376 | 0.003494 | 6.1e-05 | 0.00113 | 74.9558 | 1.353 | 2.5473 | 100.22 | 20.79 | 36.928 | 65.8139 | 85.6154 | 64.5636 | 0.009056 | 1.0454 | 0.024853 | 23.7735 | 0.015497 | 0.1387 | 5e-06 | 0.0004 | 0.0004 | 14.104
2019-03-29 13:43 | 1553881375.5286 | -612 | 0.20946 | 0.000448 | -1e-06 | 101.1787 | 101.3028 | 125.9356 | 21.0597 | 0.001876 | 0.001745 | -151.7579 | 0.23301 | 8.4019 | 0.002659 | 0.003573 | 0.003473 | 5.8e-05 | 0.001133 | 74.9753 | 1.39 | 2.5474 | 100.22 | 20.78 | 36.929 | 66.7731 | 85.585 | 63.6102 | 0.009051 | 1.0484 | 0.024889 | 23.741 | -0.010967 | 0.13861 | 1.1e-05 | 0.000887 | 0.000887 | 6.3569
2019-03-29 13:43 | 1553881376.5289 | -611 | 0.20946 | 0.000453 | 1e-06 | 101.2158 | 101.0151 | 125.8568 | 21.0597 | 0.001898 | 0.00177 | -151.752 | 0.23317 | 8.401 | 0.002652 | 0.004044 | 0.003861 | 5.8e-05 | 0.00113 | 75.0123 | 1.3871 | 2.5475 | 100.22 | 20.78 | 36.929 | 67.3059 | 85.4172 | 64.213 | 0.009051 | 1.0465 | 0.024859 | 23.7548 | 0.017753 | 0.13871 | 4e-06 | 0.000257 | 0.000257 | 21.9915
2019-03-29 13:43 | 1553881377.5297 | -610 | 0.20948 | 0.000454 | 0 | 101.2143 | 101.1154 | 125.9389 | 21.0587 | 0.001935 | 0.00182 | -151.7558 | 0.23343 | 8.3992 | 0.002609 | 0.003595 | 0.003456 | 5.8e-05 | 0.001132 | 75.0396 | 1.3684 | 2.5473 | 100.22 | 20.79 | 36.93 | 68.194 | 84.3784 | 65.4329 | 0.009057 | 1.0427 | 0.024822 | 23.8054 | 0.008741 | 0.13887 | 3e-06 | 0.000216 | 0.000216 | 26.1579
2019-03-29 13:43 | 1553881378.5291 | -609 | 0.20946 | 0.000451 | 0 | 101.2248 | 101.2185 | 126.0547 | 21.0596 | 0.001858 | 0.001807 | -151.7442 | 0.23306 | 8.4008 | 0.002646 | 0.004389 | 0.004111 | 6.2e-05 | 0.001134 | 74.9665 | 1.3789 | 2.5474 | 100.22 | 20.79 | 36.93 | 66.3359 | 85.2735 | 65.1044 | 0.009057 | 1.0437 | 0.024846 | 23.8048 | 0.000554 | 0.13865 | 7e-06 | 0.000477 | 0.000477 | 11.8241
2019-03-29 13:43 | 1553881379.5297 | -608 | 0.20946 | 0.000451 | 0 | 101.255 | 101.2387 | 125.4937 | 21.0597 | 0.001861 | 0.001757 | -151.7553 | 0.23213 | 8.4026 | 0.002743 | 0.004229 | 0.004091 | 6e-05 | 0.00114 | 74.9632 | 1.3698 | 2.5474 | 100.22 | 20.79 | 36.933 | 66.4046 | 87.6172 | 63.9178 | 0.009058 | 1.0474 | 0.024834 | 23.7101 | 0.001446 | 0.13809 | 8e-06 | 0.000624 | 0.000624 | 9.0133
2019-03-29 13:43 | 1553881380.5288 | -607 | 0.20947 | 0.000449 | -1e-06 | 101.2008 | 101.1844 | 125.3465 | 21.0601 | 0.0019 | 0.001735 | -151.7649 | 0.23277 | 8.4047 | 0.002698 | 0.003492 | 0.00339 | 5.9e-05 | 0.001143 | 75.024 | 1.3293 | 2.5474 | 100.22 | 20.79 | 36.933 | 67.3537 | 86.5392 | 63.3911 | 0.009058 | 1.049 | 0.024839 | 23.6777 | 0.001454 | 0.13848 | 1e-05 | 0.000759 | 0.000759 | 7.424
2019-03-29 13:43 | 1553881381.5288 | -606 | 0.20944 | 0.000452 | 0 | 101.2568 | 101.3349 | 125.2312 | 21.0606 | 0.001895 | 0.001753 | -151.7821 | 0.23384 | 8.4031 | 0.00276 | 0.00355 | 0.003345 | 5.9e-05 | 0.001133 | 75.0028 | 1.3391 | 2.5474 | 100.22 | 20.79 | 36.936 | 67.2342 | 88.0327 | 63.8031 | 0.009058 | 1.0478 | 0.024812 | 23.6813 | -0.006895 | 0.13911 | 6e-06 | 0.000483 | 0.000483 | 11.7326
2019-03-29 13:43 | 1553881382.5289 | -605 | 0.20946 | 0.00045 | 0 | 101.3206 | 101.0393 | 125.1871 | 21.0599 | 0.001921 | 0.001772 | -151.7852 | 0.2327 | 8.4023 | 0.002683 | 0.00413 | 0.003952 | 6e-05 | 0.001134 | 74.9753 | 1.395 | 2.5475 | 100.22 | 20.79 | 36.938 | 67.8606 | 86.1675 | 64.2724 | 0.009059 | 1.0463 | 0.024791 | 23.6936 | 0.024861 | 0.13843 | 9e-06 | 0.000696 | 0.000696 | 8.1009
2019-03-29 13:43 | 1553881383.5299 | -604 | 0.20943 | 0.000451 | -1e-06 | 101.2748 | 101.2574 | 125.1083 | 21.0606 | 0.001888 | 0.001801 | -151.807 | 0.23277 | 8.4045 | 0.002722 | 0.003623 | 0.003352 | 6.2e-05 | 0.001131 | 74.9863 | 1.3622 | 2.5474 | 100.22 | 20.79 | 36.938 | 67.0517 | 87.1003 | 64.9663 | 0.009059 | 1.0442 | 0.024757 | 23.7105 | 0.001532 | 0.13848 | 8e-06 | 0.000628 | 0.000628 | 8.973
2019-03-29 13:43 | 1553881384.529 | -603 | 0.20947 | 0.000456 | 0 | 101.2817 | 101.1291 | 124.847 | 21.0605 | 0.001864 | 0.001803 | -151.8128 | 0.23341 | 8.4022 | 0.002653 | 0.004247 | 0.004087 | 6e-05 | 0.00113 | 74.9969 | 1.3371 | 2.5472 | 100.22 | 20.79 | 36.941 | 66.4843 | 85.4382 | 65.0163 | 0.00906 | 1.044 | 0.02473 | 23.6874 | 0.013483 | 0.13886 | 1e-06 | 2.2e-05 | 2.2e-05 | 256.1466
2019-03-29 13:43 | 1553881385.5294 | -602 | 0.20944 | 0.000455 | 0 | 101.2406 | 101.227 | 124.5139 | 21.0615 | 0.0019 | 0.001788 | -151.8069 | 0.23388 | 8.4029 | 0.002663 | 0.003665 | 0.003586 | 5.8e-05 | 0.001125 | 74.9669 | 1.3551 | 2.5474 | 100.21 | 20.79 | 36.943 | 67.3387 | 85.6976 | 64.6677 | 0.009061 | 1.0451 | 0.024709 | 23.6436 | 0.001201 | 0.13914 | 3e-06 | 0.000194 | 0.000194 | 29.1913
2019-03-29 13:43 | 1553881386.529 | -601 | 0.20943 | 0.000456 | 0 | 101.2793 | 101.2755 | 124.5102 | 21.0616 | 0.001915 | 0.001731 | -151.8089 | 0.23405 | 8.4003 | 0.002623 | 0.00415 | 0.003839 | 5.9e-05 | 0.001127 | 74.9581 | 1.354 | 2.5474 | 100.21 | 20.79 | 36.947 | 67.7103 | 84.7199 | 63.2871 | 0.009062 | 1.0494 | 0.02476 | 23.5949 | 0.000331 | 0.13923 | 1e-06 | 7.5e-05 | 7.5e-05 | 75.1786
2019-03-29 13:43 | 1553881387.529 | -600 | 0.20945 | 0.000455 | 0 | 101.2829 | 101.0881 | 124.2705 | 21.0609 | 0.001925 | 0.001766 | -151.8153 | 0.23222 | 8.4014 | 0.002692 | 0.004205 | 0.004088 | 6.2e-05 | 0.001121 | 74.9644 | 1.4053 | 2.5474 | 100.21 | 20.79 | 36.949 | 67.9428 | 86.3964 | 64.1172 | 0.009062 | 1.0468 | 0.024705 | 23.6013 | 0.017212 | 0.13815 | 3e-06 | 0.000217 | 0.000217 | 25.9498
2019-03-29 13:43 | 1553881388.5291 | -599 | 0.2094 | 0.000454 | 0 | 101.1661 | 101.1711 | 124.2395 | 21.0607 | 0.001932 | 0.001779 | -151.8197 | 0.23426 | 8.4024 | 0.002676 | 0.003962 | 0.003893 | 6e-05 | 0.001125 | 74.9262 | 1.3758 | 2.5474 | 100.21 | 20.79 | 36.95 | 68.1112 | 86.0035 | 64.435 | 0.009063 | 1.0458 | 0.024691 | 23.6094 | -0.000439 | 0.13936 | 3e-06 | 0.00023 | 0.00023 | 24.6486
2019-03-29 13:43 | 1553881389.5291 | -598 | 0.20944 | 0.000455 | 0 | 101.2113 | 101.1591 | 124.5808 | 21.0613 | 0.001902 | 0.001835 | -151.8199 | 0.23411 | 8.4016 | 0.002762 | 0.004005 | 0.00374 | 6.5e-05 | 0.001122 | 74.9459 | 1.3695 | 2.5474 | 100.21 | 20.79 | 36.953 | 67.3893 | 88.078 | 65.7953 | 0.009063 | 1.0416 | 0.024675 | 23.6894 | 0.004613 | 0.13927 | 2e-06 | 0.000179 | 0.000179 | 31.5876
2019-03-29 13:43 | 1553881390.5291 | -597 | 0.20943 | 0.000454 | 0 | 101.1661 | 101.067 | 124.6731 | 21.0623 | 0.001856 | 0.001772 | -151.8145 | 0.23324 | 8.4029 | 0.002705 | 0.004268 | 0.004076 | 5.9e-05 | 0.001131 | 74.954 | 1.349 | 2.5474 | 100.22 | 20.79 | 36.955 | 66.3018 | 86.6881 | 64.2741 | 0.009063 | 1.0463 | 0.02474 | 23.6449 | 0.008769 | 0.13875 | 4e-06 | 0.000304 | 0.000304 | 18.601
2019-03-29 13:43 | 1553881391.5285 | -596 | 0.20943 | 0.000452 | 0 | 101.1751 | 101.0912 | 124.831 | 21.0616 | 0.001884 | 0.00177 | -151.8023 | 0.23354 | 8.401 | 0.002667 | 0.004286 | 0.004085 | 5.9e-05 | 0.001129 | 74.9447 | 1.3315 | 2.5474 | 100.22 | 20.79 | 36.958 | 66.963 | 85.7891 | 64.2243 | 0.009064 | 1.0465 | 0.024757 | 23.6582 | 0.007418 | 0.13893 | 6e-06 | 0.000447 | 0.000447 | 12.6593
2019-03-29 13:43 | 1553881392.5292 | -595 | 0.20943 | 0.000458 | 0 | 101.199 | 101.0585 | 124.8768 | 21.0622 | 0.001896 | 0.001804 | -151.8056 | 0.23294 | 8.4028 | 0.002808 | 0.003545 | 0.003403 | 6.3e-05 | 0.001129 | 74.8973 | 1.3313 | 2.5473 | 100.22 | 20.79 | 36.959 | 67.2615 | 89.1879 | 65.0468 | 0.009064 | 1.0439 | 0.024732 | 23.6913 | 0.012435 | 0.13858 | -2e-06 | -0.000105 | -0.000105 | -53.9184
2019-03-29 13:43 | 1553881393.5282 | -594 | 0.20943 | 0.000452 | 0 | 101.1658 | 101.1352 | 125.0402 | 21.0626 | 0.001935 | 0.001756 | -151.8003 | 0.23415 | 8.4037 | 0.002749 | 0.004122 | 0.004 | 6.1e-05 | 0.00113 | 74.8995 | 1.3148 | 2.5472 | 100.22 | 20.79 | 36.96 | 68.1854 | 87.7699 | 63.8934 | 0.009064 | 1.0475 | 0.02479 | 23.6664 | 0.002703 | 0.1393 | 6e-06 | 0.000466 | 0.000466 | 12.1609
2019-03-29 13:43 | 1553881394.5287 | -593 | 0.20943 | 0.000453 | 0 | 101.2134 | 101.0567 | 125.1451 | 21.0622 | 0.00188 | 0.001771 | -151.8072 | 0.2337 | 8.4034 | 0.002777 | 0.003969 | 0.003765 | 6.1e-05 | 0.001127 | 74.9616 | 1.2814 | 2.5473 | 100.22 | 20.8 | 36.961 | 66.8784 | 88.4271 | 64.2364 | 0.00907 | 1.0464 | 0.024788 | 23.6883 | 0.013855 | 0.13903 | 5e-06 | 0.000378 | 0.000378 | 14.9508
2019-03-29 13:43 | 1553881395.5292 | -592 | 0.20944 | 0.000458 | 0 | 101.1239 | 100.9789 | 125.266 | 21.0628 | 0.001926 | 0.00179 | -151.7995 | 0.23335 | 8.4029 | 0.00272 | 0.004387 | 0.004189 | 6e-05 | 0.001131 | 74.9377 | 1.3303 | 2.5475 | 100.22 | 20.8 | 36.961 | 67.9805 | 87.0531 | 64.7079 | 0.00907 | 1.045 | 0.024782 | 23.7163 | 0.012832 | 0.13882 | -2e-06 | -0.000114 | -0.000114 | -49.744
2019-03-29 13:43 | 1553881396.5293 | -591 | 0.20946 | 0.00046 | 0 | 101.0407 | 101.0096 | 125.5526 | 21.0633 | 0.001971 | 0.001817 | -151.787 | 0.23336 | 8.4016 | 0.002654 | 0.004062 | 0.003934 | 6e-05 | 0.001136 | 74.9664 | 1.3288 | 2.5473 | 100.22 | 20.8 | 36.961 | 69.0572 | 85.4768 | 65.3459 | 0.00907 | 1.043 | 0.024787 | 23.7659 | 0.002749 | 0.13883 | -5e-06 | -0.00036 | -0.00036 | -15.6887
2019-03-29 13:43 | 1553881397.5292 | -590 | 0.20948 | 0.000463 | 0 | 101.0422 | 101.0707 | 125.8286 | 21.0637 | 0.001863 | 0.00177 | -151.7877 | 0.23338 | 8.4033 | 0.002741 | 0.004359 | 0.004101 | 6.1e-05 | 0.001133 | 74.9273 | 1.3394 | 2.5473 | 100.22 | 20.8 | 36.961 | 66.4624 | 87.5671 | 64.2263 | 0.00907 | 1.0464 | 0.024856 | 23.7526 | -0.002527 | 0.13884 | -7e-06 | -0.000505 | -0.000505 | -11.1886
2019-03-29 13:43 | 1553881398.5284 | -589 | 0.20947 | 0.000458 | -1e-06 | 101.0416 | 101.0168 | 125.9206 | 21.0634 | 0.001857 | 0.001793 | -151.798 | 0.23416 | 8.4001 | 0.002842 | 0.004161 | 0.004043 | 6.1e-05 | 0.001141 | 74.9081 | 1.3258 | 2.5474 | 100.22 | 20.8 | 36.961 | 66.3171 | 90.0031 | 64.7819 | 0.00907 | 1.0447 | 0.024844 | 23.7808 | 0.002193 | 0.1393 | -1e-06 | -6.1e-05 | -6.1e-05 | -92.4369
2019-03-29 13:43 | 1553881399.5293 | -588 | 0.20944 | 0.000465 | 0 | 100.9736 | 100.9789 | 125.8033 | 21.0632 | 0.001899 | 0.001809 | -151.7944 | 0.23391 | 8.3997 | 0.002703 | 0.003851 | 0.003716 | 6.3e-05 | 0.001154 | 74.8982 | 1.3214 | 2.5475 | 100.22 | 20.8 | 36.961 | 67.3257 | 86.6476 | 65.1688 | 0.00907 | 1.0435 | 0.024819 | 23.7833 | -0.000468 | 0.13915 | -1.1e-05 | -0.000791 | -0.000791 | -7.159
2019-03-29 13:43 | 1553881400.5294 | -587 | 0.20947 | 0.000456 | 0 | 100.9455 | 100.9868 | 125.8413 | 21.0649 | 0.001937 | 0.001745 | -151.796 | 0.23319 | 8.4006 | 0.002722 | 0.004539 | 0.00432 | 6e-05 | 0.001154 | 74.8702 | 1.4242 | 2.5474 | 100.22 | 20.8 | 36.96 | 68.2515 | 87.1161 | 63.6177 | 0.00907 | 1.0483 | 0.024879 | 23.7323 | -0.003657 | 0.13872 | 1e-06 | 7.5e-05 | 7.5e-05 | 75.3423
2019-03-29 13:43 | 1553881401.5295 | -586 | 0.20949 | 0.000456 | -1e-06 | 100.9629 | 100.8189 | 125.6406 | 21.064 | 0.001867 | 0.001775 | -151.7989 | 0.23315 | 8.4009 | 0.002722 | 0.004108 | 0.003965 | 6e-05 | 0.001153 | 74.9138 | 1.3659 | 2.5474 | 100.22 | 20.8 | 36.96 | 66.5478 | 87.1044 | 64.3539 | 0.00907 | 1.046 | 0.024832 | 23.7393 | 0.012766 | 0.1387 | 2e-06 | 0.000195 | 0.000195 | 28.9603
2019-03-29 13:43 | 1553881402.5299 | -585 | 0.20945 | 0.000462 | -1e-06 | 100.9084 | 100.8438 | 125.3871 | 21.0646 | 0.001973 | 0.001815 | -151.7874 | 0.23362 | 8.4049 | 0.002753 | 0.003517 | 0.00336 | 5.9e-05 | 0.001151 | 74.921 | 1.3439 | 2.5474 | 100.22 | 20.8 | 36.959 | 69.1109 | 87.8609 | 65.3152 | 0.00907 | 1.0431 | 0.024772 | 23.7491 | 0.005728 | 0.13898 | -7e-06 | -0.000469 | -0.000469 | -12.0686
2019-03-29 13:43 | 1553881403.5295 | -584 | 0.20949 | 0.000459 | 0 | 100.9338 | 100.8704 | 125.3683 | 21.064 | 0.00192 | 0.001772 | -151.8008 | 0.23334 | 8.4017 | 0.002737 | 0.003842 | 0.003695 | 6e-05 | 0.001139 | 74.8672 | 1.3076 | 2.5473 | 100.22 | 20.8 | 36.958 | 67.8311 | 87.4689 | 64.2733 | 0.009069 | 1.0463 | 0.024809 | 23.7108 | 0.005623 | 0.13882 | -3e-06 | -0.000171 | -0.000171 | -33.0299
2019-03-29 13:43 | 1553881404.5296 | -583 | 0.20947 | 0.000461 | 0 | 100.8356 | 100.8806 | 125.4786 | 21.0646 | 0.001919 | 0.001835 | -151.7929 | 0.23248 | 8.4036 | 0.002742 | 0.004424 | 0.004224 | 6e-05 | 0.001136 | 74.9283 | 1.3462 | 2.5475 | 100.22 | 20.8 | 36.957 | 67.8098 | 87.592 | 65.8016 | 0.009069 | 1.0416 | 0.024763 | 23.7748 | -0.00399 | 0.13831 | -6e-06 | -0.000406 | -0.000406 | -13.8625
2019-03-29 13:43 | 1553881405.5287 | -582 | 0.2095 | 0.000459 | -1e-06 | 100.8841 | 100.6827 | 125.5329 | 21.0653 | 0.001934 | 0.001755 | -151.7994 | 0.23292 | 8.4016 | 0.002745 | 0.003722 | 0.003527 | 5.9e-05 | 0.001131 | 74.9325 | 1.3322 | 2.5474 | 100.22 | 20.8 | 36.957 | 68.1661 | 87.6748 | 63.8602 | 0.009069 | 1.0476 | 0.02484 | 23.7118 | 0.017879 | 0.13856 | -2e-06 | -0.000125 | -0.000125 | -45.1286
2019-03-29 13:43 | 1553881406.5285 | -581 | 0.20946 | 0.000458 | 0 | 100.8255 | 100.752 | 125.6333 | 21.0651 | 0.00192 | 0.001786 | -151.8015 | 0.23374 | 8.4008 | 0.002663 | 0.004236 | 0.004048 | 6.1e-05 | 0.001129 | 74.9831 | 1.316 | 2.5474 | 100.22 | 20.8 | 36.956 | 67.8423 | 85.6854 | 64.6225 | 0.009069 | 1.0452 | 0.024822 | 23.7481 | 0.006523 | 0.13905 | -2e-06 | -0.000173 | -0.000173 | -32.7791
2019-03-29 13:43 | 1553881407.5277 | -580 | 0.20947 | 0.000461 | 0 | 100.7114 | 100.6172 | 125.7124 | 21.0653 | 0.001897 | 0.001762 | -151.7876 | 0.23377 | 8.4014 | 0.002657 | 0.003539 | 0.003324 | 6.1e-05 | 0.001126 | 74.982 | 1.3383 | 2.5473 | 100.21 | 20.8 | 36.957 | 67.2722 | 85.5583 | 64.0272 | 0.00907 | 1.0471 | 0.024852 | 23.7346 | 0.008376 | 0.13907 | -6e-06 | -0.000412 | -0.000412 | -13.7288
2019-03-29 13:43 | 1553881408.5277 | -579 | 0.20947 | 0.000458 | -1e-06 | 100.7201 | 100.6008 | 125.7718 | 21.0653 | 0.001923 | 0.00182 | -151.7952 | 0.23338 | 8.403 | 0.002713 | 0.004411 | 0.004212 | 5.9e-05 | 0.001124 | 74.9597 | 1.3088 | 2.5473 | 100.21 | 20.8 | 36.957 | 67.8956 | 86.9035 | 65.4392 | 0.00907 | 1.0427 | 0.024806 | 23.7899 | 0.010603 | 0.13884 | -1e-06 | 7e-06 | 7e-06 | 762.7069
2019-03-29 13:43 | 1553881409.5279 | -578 | 0.20944 | 0.000461 | -1e-06 | 100.6905 | 100.5308 | 125.8702 | 21.0659 | 0.001887 | 0.001816 | -151.7981 | 0.23408 | 8.401 | 0.002763 | 0.003779 | 0.003577 | 6e-05 | 0.001124 | 74.9696 | 1.33 | 2.5473 | 100.21 | 20.8 | 36.959 | 67.0333 | 88.0933 | 65.3345 | 0.00907 | 1.043 | 0.024819 | 23.7955 | 0.014194 | 0.13925 | -5e-06 | -0.000354 | -0.000354 | -15.9945
2019-03-29 13:43 | 1553881410.5278 | -577 | 0.20944 | 0.000457 | 0 | 100.6713 | 100.6346 | 125.7314 | 21.0654 | 0.001893 | 0.001774 | -151.7888 | 0.23276 | 8.399 | 0.00275 | 0.004168 | 0.003972 | 6.2e-05 | 0.001121 | 74.9486 | 1.3564 | 2.5472 | 100.21 | 20.8 | 36.959 | 67.1943 | 87.7858 | 64.3215 | 0.00907 | 1.0461 | 0.024843 | 23.7468 | 0.003265 | 0.13847 | 0 | 1.1e-05 | 1.1e-05 | 535.2343
2019-03-29 13:44 | 1553881411.5287 | -576 | 0.20945 | 0.000458 | 0 | 100.645 | 100.5267 | 125.6392 | 21.0654 | 0.001817 | 0.001781 | -151.7957 | 0.23253 | 8.4034 | 0.002781 | 0.004288 | 0.004098 | 6.1e-05 | 0.001126 | 74.9669 | 1.3379 | 2.5474 | 100.21 | 20.8 | 36.959 | 65.3667 | 88.5309 | 64.5028 | 0.00907 | 1.0456 | 0.024827 | 23.7444 | 0.010515 | 0.13833 | -2e-06 | -0.000109 | -0.000109 | -51.4666
2019-03-29 13:44 | 1553881412.5282 | -575 | 0.20946 | 0.000453 | 0 | 100.6028 | 100.4243 | 125.8561 | 21.0664 | 0.001888 | 0.001799 | -151.7937 | 0.23309 | 8.4041 | 0.002765 | 0.004449 | 0.004272 | 6e-05 | 0.001131 | 74.9942 | 1.3169 | 2.5472 | 100.21 | 20.8 | 36.959 | 67.072 | 88.1577 | 64.9274 | 0.00907 | 1.0443 | 0.024833 | 23.7798 | 0.015878 | 0.13867 | 4e-06 | 0.000339 | 0.000339 | 16.638
2019-03-29 13:44 | 1553881413.5292 | -574 | 0.20947 | 0.000459 | 0 | 100.5755 | 100.4144 | 125.6162 | 21.0664 | 0.001937 | 0.001754 | -151.7825 | 0.23331 | 8.4014 | 0.00277 | 0.003384 | 0.003256 | 6.5e-05 | 0.001137 | 74.9624 | 1.3317 | 2.5474 | 100.21 | 20.8 | 36.959 | 68.2523 | 88.2601 | 63.8502 | 0.00907 | 1.0476 | 0.024849 | 23.7193 | 0.014338 | 0.1388 | -3e-06 | -0.000193 | -0.000193 | -29.2449
2019-03-29 13:44 | 1553881414.5279 | -573 | 0.20945 | 0.00046 | 0 | 100.5034 | 100.465 | 125.6392 | 21.0661 | 0.001904 | 0.001841 | -151.7749 | 0.23314 | 8.4017 | 0.002731 | 0.003823 | 0.003563 | 6.2e-05 | 0.00114 | 74.9413 | 1.3238 | 2.5475 | 100.21 | 20.8 | 36.96 | 67.4494 | 87.3367 | 65.9435 | 0.009071 | 1.0411 | 0.024774 | 23.795 | 0.003422 | 0.1387 | -4e-06 | -0.000294 | -0.000294 | -19.2046
2019-03-29 13:44 | 1553881415.5279 | -572 | 0.20942 | 0.000456 | -1e-06 | 100.5142 | 100.2731 | 125.83 | 21.0663 | 0.001913 | 0.001805 | -151.7758 | 0.23347 | 8.4041 | 0.002701 | 0.005012 | 0.004793 | 6.1e-05 | 0.001142 | 74.9635 | 1.336 | 2.5473 | 100.21 | 20.8 | 36.961 | 67.6766 | 86.6078 | 65.0594 | 0.009071 | 1.0439 | 0.024825 | 23.782 | 0.021474 | 0.13889 | 1e-06 | 0.000157 | 0.000157 | 36.0287
2019-03-29 13:44 | 1553881416.5288 | -571 | 0.20945 | 0.000459 | 0 | 100.5312 | 100.4042 | 125.7333 | 21.0671 | 0.001888 | 0.001772 | -151.7709 | 0.23351 | 8.4006 | 0.002655 | 0.004384 | 0.004261 | 6e-05 | 0.001144 | 74.8654 | 1.3526 | 2.5473 | 100.21 | 20.8 | 36.962 | 67.0758 | 85.5002 | 64.2818 | 0.009071 | 1.0463 | 0.024844 | 23.7455 | 0.011309 | 0.13891 | -2e-06 | -0.000126 | -0.000126 | -44.9401
2019-03-29 13:44 | 1553881417.5294 | -570 | 0.20943 | 0.000455 | 0 | 100.5001 | 100.1755 | 125.4653 | 21.0675 | 0.001919 | 0.001801 | -151.7962 | 0.23379 | 8.4013 | 0.002756 | 0.00376 | 0.003592 | 6.1e-05 | 0.001144 | 74.8856 | 1.3224 | 2.5472 | 100.21 | 20.8 | 36.965 | 67.8096 | 87.934 | 64.9692 | 0.009072 | 1.0441 | 0.024793 | 23.7444 | 0.028933 | 0.13908 | 2e-06 | 0.000166 | 0.000166 | 34.1206
2019-03-29 13:44 | 1553881418.529 | -569 | 0.20942 | 0.000456 | 0 | 100.3977 | 100.3223 | 125.3329 | 21.0671 | 0.001868 | 0.001832 | -151.7889 | 0.23306 | 8.4039 | 0.002721 | 0.003994 | 0.003743 | 6.2e-05 | 0.00115 | 74.82 | 1.3365 | 2.5475 | 100.21 | 20.8 | 36.966 | 66.5898 | 87.1004 | 65.729 | 0.009072 | 1.0418 | 0.024752 | 23.7585 | 0.006723 | 0.13865 | 0 | 3.2e-05 | 3.2e-05 | 174.0434
2019-03-29 13:44 | 1553881419.529 | -568 | 0.20946 | 0.00046 | 0 | 100.3477 | 100.2011 | 125.5002 | 21.0675 | 0.00193 | 0.001801 | -151.79 | 0.23321 | 8.4036 | 0.002788 | 0.003655 | 0.003521 | 5.9e-05 | 0.001156 | 74.8383 | 1.3175 | 2.5474 | 100.21 | 20.8 | 36.966 | 68.0747 | 88.6969 | 64.9692 | 0.009072 | 1.0441 | 0.024796 | 23.7477 | 0.013077 | 0.13874 | -4e-06 | -0.000275 | -0.000275 | -20.5647
2019-03-29 13:44 | 1553881420.5283 | -567 | 0.20945 | 0.000452 | 0 | 100.3938 | 100.1322 | 125.3758 | 21.069 | 0.001867 | 0.001817 | -151.8093 | 0.23285 | 8.4013 | 0.002685 | 0.003994 | 0.003717 | 6e-05 | 0.001153 | 74.8826 | 1.3324 | 2.5473 | 100.21 | 20.8 | 36.968 | 66.5596 | 86.2223 | 65.3657 | 0.009073 | 1.0429 | 0.024769 | 23.7498 | 0.023341 | 0.13852 | 6e-06 | 0.000483 | 0.000483 | 11.6832
2019-03-29 13:44 | 1553881421.5271 | -566 | 0.20949 | 0.000457 | 0 | 100.375 | 100.0755 | 125.2815 | 21.0679 | 0.001926 | 0.00183 | -151.7979 | 0.23319 | 8.4022 | 0.002695 | 0.004061 | 0.003944 | 6.3e-05 | 0.001156 | 74.8993 | 1.3562 | 2.5474 | 100.21 | 20.8 | 36.968 | 67.9785 | 86.4648 | 65.6682 | 0.009073 | 1.042 | 0.024749 | 23.7515 | 0.026723 | 0.13872 | 0 | 9e-06 | 9e-06 | 624.6982
2019-03-29 13:44 | 1553881422.5286 | -565 | 0.20945 | 0.000459 | 0 | 100.2888 | 100.2407 | 125.4498 | 21.0681 | 0.001876 | 0.001733 | -151.8005 | 0.23419 | 8.402 | 0.002659 | 0.004477 | 0.004237 | 5.9e-05 | 0.001158 | 74.9233 | 1.3212 | 2.5473 | 100.21 | 20.8 | 36.971 | 66.774 | 85.6024 | 63.3295 | 0.009073 | 1.0492 | 0.024851 | 23.6853 | 0.004286 | 0.13932 | -3e-06 | -0.000206 | -0.000206 | -27.524
2019-03-29 13:44 | 1553881423.5282 | -564 | 0.20947 | 0.000456 | 0 | 100.2277 | 100.2247 | 125.7821 | 21.0685 | 0.001875 | 0.001812 | -151.804 | 0.2329 | 8.4009 | 0.002671 | 0.004364 | 0.004141 | 6.4e-05 | 0.00116 | 74.9068 | 1.2855 | 2.5472 | 100.21 | 20.8 | 36.971 | 66.7566 | 85.8799 | 65.2379 | 0.009073 | 1.0433 | 0.024814 | 23.7838 | 0.00027 | 0.13855 | 1e-06 | 6e-05 | 6e-05 | 94.2607
2019-03-29 13:44 | 1553881424.5282 | -563 | 0.20948 | 0.000456 | 0 | 100.2657 | 100.0745 | 125.8563 | 21.0692 | 0.001886 | 0.001757 | -151.8205 | 0.23264 | 8.4018 | 0.002734 | 0.004158 | 0.004012 | 6.1e-05 | 0.001157 | 74.8657 | 1.3772 | 2.5474 | 100.21 | 20.8 | 36.978 | 67.0225 | 87.403 | 63.9217 | 0.009075 | 1.0474 | 0.02487 | 23.7445 | 0.017073 | 0.13839 | 1e-06 | 0.000107 | 0.000107 | 52.5303
2019-03-29 13:44 | 1553881425.5275 | -562 | 0.20949 | 0.000457 | -1e-06 | 100.2487 | 100.153 | 125.7732 | 21.0685 | 0.001875 | 0.001756 | -151.8193 | 0.23361 | 8.4022 | 0.002742 | 0.003977 | 0.003791 | 6.1e-05 | 0.001151 | 74.8891 | 1.3219 | 2.5474 | 100.21 | 20.8 | 36.978 | 66.7612 | 87.5944 | 63.8998 | 0.009075 | 1.0475 | 0.024862 | 23.7359 | 0.008538 | 0.13897 | 0 | 3.3e-05 | 3.3e-05 | 172.2267
2019-03-29 13:44 | 1553881426.5273 | -561 | 0.20946 | 0.000462 | -1e-06 | 100.2822 | 100.1127 | 125.8129 | 21.0692 | 0.001936 | 0.001771 | -151.8376 | 0.23268 | 8.4011 | 0.002705 | 0.004171 | 0.004004 | 5.9e-05 | 0.001149 | 74.8716 | 1.2944 | 2.5473 | 100.21 | 20.8 | 36.982 | 68.2247 | 86.7018 | 64.2612 | 0.009076 | 1.0463 | 0.024853 | 23.7523 | 0.015126 | 0.13842 | -6e-06 | -0.00037 | -0.00037 | -15.2162
2019-03-29 13:44 | 1553881427.5273 | -560 | 0.20941 | 0.000456 | -1e-06 | 100.2588 | 100.0506 | 125.8439 | 21.0682 | 0.00196 | 0.001825 | -151.8175 | 0.23334 | 8.4023 | 0.002736 | 0.003427 | 0.003247 | 6.1e-05 | 0.001155 | 74.8731 | 1.3378 | 2.5473 | 100.21 | 20.8 | 36.982 | 68.7923 | 87.4505 | 65.5468 | 0.009076 | 1.0424 | 0.024809 | 23.8005 | 0.018595 | 0.13882 | 1e-06 | 8.2e-05 | 8.2e-05 | 68.6628
2019-03-29 13:44 | 1553881428.5278 | -559 | 0.20944 | 0.000457 | 0 | 100.225 | 100.2598 | 125.8664 | 21.0692 | 0.001898 | 0.001801 | -151.8128 | 0.23327 | 8.4024 | 0.002843 | 0.004201 | 0.003963 | 5.9e-05 | 0.001156 | 74.8399 | 1.3222 | 2.5473 | 100.21 | 20.8 | 36.988 | 67.3197 | 90.0198 | 64.9882 | 0.009078 | 1.0441 | 0.024831 | 23.783 | -0.003107 | 0.13877 | 0 | 2e-06 | 2e-06 | 3702.435
2019-03-29 13:44 | 1553881429.5274 | -558 | 0.20946 | 0.000459 | -1e-06 | 100.1945 | 100.0994 | 125.5319 | 21.0692 | 0.00186 | 0.001773 | -151.8015 | 0.23373 | 8.4051 | 0.002718 | 0.004337 | 0.004093 | 6.2e-05 | 0.001154 | 74.8525 | 1.3348 | 2.5474 | 100.21 | 20.8 | 36.988 | 66.4006 | 87.011 | 64.3072 | 0.009078 | 1.0462 | 0.024823 | 23.7274 | 0.008489 | 0.13905 | -3e-06 | -0.000156 | -0.000156 | -36.3271
2019-03-29 13:44 | 1553881430.5275 | -557 | 0.20944 | 0.000462 | -2e-06 | 100.1951 | 100.1213 | 125.6441 | 21.069 | 0.001939 | 0.001833 | -151.7783 | 0.2333 | 8.4011 | 0.002706 | 0.002968 | 0.002884 | 6.2e-05 | 0.001156 | 74.8682 | 1.3743 | 2.5475 | 100.21 | 20.8 | 36.99 | 68.3024 | 86.7204 | 65.7472 | 0.009078 | 1.0417 | 0.024782 | 23.7886 | 0.006591 | 0.13879 | -5e-06 | -0.000257 | -0.000257 | -21.939
2019-03-29 13:44 | 1553881431.5274 | -556 | 0.20944 | 0.000459 | -1e-06 | 100.2355 | 99.8223 | 125.7462 | 21.0692 | 0.00195 | 0.001794 | -151.7816 | 0.23297 | 8.402 | 0.002728 | 0.004016 | 0.003829 | 6.4e-05 | 0.001152 | 74.8424 | 1.3732 | 2.5474 | 100.21 | 20.8 | 36.991 | 68.5499 | 87.2687 | 64.8178 | 0.009078 | 1.0446 | 0.024826 | 23.7656 | 0.036947 | 0.13859 | -2e-06 | -7.1e-05 | -7.1e-05 | -79.5759
2019-03-29 13:44 | 1553881432.5288 | -555 | 0.20945 | 0.00046 | -1e-06 | 100.1714 | 99.9796 | 125.6451 | 21.069 | 0.001854 | 0.001814 | -151.7859 | 0.23275 | 8.4026 | 0.002687 | 0.004632 | 0.004399 | 5.8e-05 | 0.001151 | 74.8266 | 1.3458 | 2.5473 | 100.21 | 20.8 | 36.992 | 66.2568 | 86.2769 | 65.2812 | 0.009079 | 1.0432 | 0.024799 | 23.7723 | 0.017142 | 0.13846 | -3e-06 | -0.000196 | -0.000196 | -28.7871
2019-03-29 13:44 | 1553881433.5279 | -554 | 0.20945 | 0.00046 | -1e-06 | 100.1409 | 99.9718 | 125.5378 | 21.07 | 0.001915 | 0.001733 | -151.7759 | 0.23299 | 8.4036 | 0.002668 | 0.004158 | 0.004036 | 6.2e-05 | 0.00115 | 74.872 | 1.3349 | 2.5474 | 100.21 | 20.8 | 36.991 | 67.7167 | 85.826 | 63.3328 | 0.009078 | 1.0492 | 0.02486 | 23.6937 | 0.015118 | 0.1386 | -4e-06 | -0.000282 | -0.000282 | -20.0243
2019-03-29 13:44 | 1553881434.5281 | -553 | 0.20942 | 0.000457 | -1e-06 | 100.1858 | 99.893 | 125.3113 | 21.0693 | 0.001883 | 0.001748 | -151.762 | 0.23319 | 8.4024 | 0.002628 | 0.004277 | 0.004099 | 5.9e-05 | 0.001153 | 74.9002 | 1.337 | 2.5474 | 100.21 | 20.8 | 36.991 | 66.9524 | 84.864 | 63.6999 | 0.009078 | 1.0481 | 0.024824 | 23.6852 | 0.026184 | 0.13872 | 0 | 5.8e-05 | 5.8e-05 | 97.4106
2019-03-29 13:44 | 1553881435.5286 | -552 | 0.20946 | 0.000456 | -1e-06 | 100.1532 | 99.9984 | 125.2017 | 21.0698 | 0.001987 | 0.001774 | -151.7549 | 0.23325 | 8.4031 | 0.002768 | 0.003519 | 0.00336 | 6.2e-05 | 0.001152 | 74.8704 | 1.3833 | 2.5475 | 100.21 | 20.8 | 36.991 | 69.438 | 88.2119 | 64.3283 | 0.009078 | 1.0461 | 0.02479 | 23.6969 | 0.013833 | 0.13876 | 2e-06 | 0.000175 | 0.000175 | 32.319
2019-03-29 13:44 | 1553881436.5286 | -551 | 0.20946 | 0.000458 | 0 | 100.1208 | 99.981 | 124.7616 | 21.0713 | 0.001974 | 0.0018 | -151.7425 | 0.23365 | 8.4012 | 0.002801 | 0.00348 | 0.003236 | 6.1e-05 | 0.001151 | 74.8314 | 1.3566 | 2.5475 | 100.21 | 20.8 | 36.988 | 69.1383 | 89.0092 | 64.9495 | 0.009078 | 1.0442 | 0.024724 | 23.677 | 0.012502 | 0.139 | -1e-06 | -6.3e-05 | -6.3e-05 | -90.2297
2019-03-29 13:44 | 1553881437.5276 | -550 | 0.20946 | 0.000458 | -1e-06 | 100.1472 | 99.9523 | 124.9179 | 21.0691 | 0.001918 | 0.001773 | -151.7359 | 0.23314 | 8.4038 | 0.002791 | 0.004018 | 0.003806 | 6.1e-05 | 0.001153 | 74.8702 | 1.3836 | 2.5474 | 100.21 | 20.8 | 36.986 | 67.7972 | 88.7635 | 64.2987 | 0.009077 | 1.0462 | 0.024763 | 23.669 | 0.01742 | 0.13869 | -1e-06 | -2.9e-05 | -2.9e-05 | -197.8491
2019-03-29 13:44 | 1553881438.5271 | -549 | 0.20948 | 0.000458 | -1e-06 | 100.1586 | 100.0714 | 125.0306 | 21.0713 | 0.001911 | 0.001764 | -151.7384 | 0.23332 | 8.4017 | 0.002763 | 0.003329 | 0.003176 | 6.3e-05 | 0.001155 | 74.891 | 1.3899 | 2.5474 | 100.21 | 20.8 | 36.983 | 67.6175 | 88.1062 | 64.079 | 0.009076 | 1.0469 | 0.024782 | 23.672 | 0.007783 | 0.13881 | -1e-06 | -5.8e-05 | -5.8e-05 | -97.216
2019-03-29 13:44 | 1553881439.5277 | -548 | 0.20947 | 0.000456 | -1e-06 | 100.0667 | 99.7977 | 125.1169 | 21.0711 | 0.001931 | 0.001786 | -151.7514 | 0.23311 | 8.4015 | 0.002803 | 0.003642 | 0.00349 | 5.9e-05 | 0.001155 | 74.8521 | 1.3562 | 2.5474 | 100.21 | 20.8 | 36.98 | 68.0955 | 89.0574 | 64.6262 | 0.009076 | 1.0452 | 0.024771 | 23.6994 | 0.024071 | 0.13868 | 1e-06 | 0.000133 | 0.000133 | 42.3177
2019-03-29 13:44 | 1553881440.5276 | -547 | 0.20949 | 0.000459 | -1e-06 | 99.993 | 99.9445 | 125.45 | 21.0715 | 0.001965 | 0.00178 | -151.7706 | 0.23353 | 8.4034 | 0.002823 | 0.0041 | 0.003865 | 5.9e-05 | 0.001155 | 74.8428 | 1.3184 | 2.5474 | 100.21 | 20.8 | 36.977 | 68.9163 | 89.557 | 64.4788 | 0.009075 | 1.0457 | 0.024809 | 23.7257 | 0.004344 | 0.13892 | -3e-06 | -0.000207 | -0.000207 | -27.3593
2019-03-29 13:44 | 1553881441.5277 | -546 | 0.20946 | 0.000457 | 0 | 100.0493 | 100.0363 | 125.4603 | 21.0708 | 0.001948 | 0.001818 | -151.762 | 0.23311 | 8.4028 | 0.002815 | 0.003945 | 0.00373 | 6e-05 | 0.001166 | 74.8254 | 1.3537 | 2.5474 | 100.21 | 20.8 | 36.974 | 68.5061 | 89.3589 | 65.3946 | 0.009074 | 1.0428 | 0.024776 | 23.7588 | 0.001163 | 0.13868 | 0 | 9e-06 | 9e-06 | 626.037
2019-03-29 13:44 | 1553881442.5266 | -545 | 0.20943 | 0.000461 | 0 | 100.1382 | 99.9127 | 125.1207 | 21.0713 | 0.001921 | 0.001772 | -151.7589 | 0.23257 | 8.4039 | 0.00282 | 0.004468 | 0.004275 | 6.3e-05 | 0.001165 | 74.8566 | 1.3437 | 2.5474 | 100.21 | 20.8 | 36.969 | 67.8726 | 89.4733 | 64.272 | 0.009073 | 1.0463 | 0.024784 | 23.6873 | 0.020161 | 0.13836 | -5e-06 | -0.000377 | -0.000377 | -14.9459
2019-03-29 13:44 | 1553881443.5278 | -544 | 0.20955 | 0.000456 | 0 | 100.0715 | 99.795 | 124.8132 | 21.0711 | 0.001964 | 0.001778 | -151.7847 | 0.23325 | 8.4048 | 0.002754 | 0.00373 | 0.003504 | 5.9e-05 | 0.001169 | 74.8405 | 1.3162 | 2.5474 | 100.21 | 20.8 | 36.968 | 68.8941 | 87.8788 | 64.431 | 0.009073 | 1.0458 | 0.024748 | 23.6637 | 0.024744 | 0.13876 | 1e-06 | 9.6e-05 | 9.6e-05 | 58.6821
2019-03-29 13:44 | 1553881444.5278 | -543 | 0.20947 | 0.000453 | 0 | 99.9634 | 99.7899 | 124.7137 | 21.0704 | 0.001916 | 0.001771 | -151.7752 | 0.23343 | 8.4014 | 0.002819 | 0.004237 | 0.004115 | 6.3e-05 | 0.001169 | 74.8372 | 1.3738 | 2.5474 | 100.21 | 20.8 | 36.967 | 67.7429 | 89.4623 | 64.26 | 0.009072 | 1.0463 | 0.024744 | 23.6483 | 0.015539 | 0.13887 | 5e-06 | 0.00041 | 0.00041 | 13.7854
2019-03-29 13:44 | 1553881445.5269 | -542 | 0.20943 | 0.00046 | -1e-06 | 100.0173 | 99.7742 | 124.7066 | 21.0715 | 0.001959 | 0.001811 | -151.7809 | 0.23342 | 8.4043 | 0.002807 | 0.004786 | 0.004655 | 6.4e-05 | 0.001165 | 74.8673 | 1.3552 | 2.5473 | 100.21 | 20.8 | 36.966 | 68.7861 | 89.1525 | 65.2207 | 0.009072 | 1.0434 | 0.024708 | 23.6813 | 0.021766 | 0.13886 | -4e-06 | -0.000246 | -0.000246 | -22.9828
2019-03-29 13:44 | 1553881446.5269 | -541 | 0.20947 | 0.000455 | -1e-06 | 100.0125 | 99.8401 | 124.6886 | 21.0716 | 0.001895 | 0.00181 | -151.7877 | 0.23403 | 8.3992 | 0.002703 | 0.003332 | 0.00304 | 6.2e-05 | 0.001165 | 74.8987 | 1.3351 | 2.5474 | 100.21 | 20.8 | 36.966 | 67.2432 | 86.6551 | 65.1847 | 0.009072 | 1.0435 | 0.024708 | 23.6783 | 0.015435 | 0.13922 | 2e-06 | 0.000214 | 0.000214 | 26.446
2019-03-29 13:44 | 1553881447.527 | -540 | 0.20943 | 0.000458 | 0 | 100.0697 | 99.8889 | 124.5944 | 21.0702 | 0.001898 | 0.001728 | -151.7935 | 0.23337 | 8.4032 | 0.002715 | 0.004374 | 0.004221 | 6e-05 | 0.001166 | 74.8525 | 1.3639 | 2.5473 | 100.21 | 20.8 | 36.966 | 67.3176 | 86.9518 | 63.2095 | 0.009072 | 1.0496 | 0.024771 | 23.6002 | 0.016176 | 0.13883 | -1e-06 | -6.7e-05 | -6.7e-05 | -84.1881
2019-03-29 13:44 | 1553881448.5269 | -539 | 0.20949 | 0.000458 | 0 | 99.9245 | 99.7728 | 124.4144 | 21.0702 | 0.001902 | 0.001834 | -151.7908 | 0.23327 | 8.4 | 0.002777 | 0.003965 | 0.003835 | 5.9e-05 | 0.00116 | 74.8243 | 1.3346 | 2.5473 | 100.21 | 20.8 | 36.966 | 67.401 | 88.4419 | 65.775 | 0.009072 | 1.0417 | 0.024659 | 23.6729 | 0.013586 | 0.13877 | -2e-06 | -0.000106 | -0.000106 | -53.0527
2019-03-29 13:44 | 1553881449.527 | -538 | 0.20945 | 0.000456 | 0 | 99.9496 | 99.6851 | 124.6111 | 21.0717 | 0.001926 | 0.0018 | -151.7871 | 0.23292 | 8.4024 | 0.002805 | 0.003586 | 0.00346 | 6e-05 | 0.001163 | 74.8788 | 1.3824 | 2.5475 | 100.21 | 20.8 | 36.966 | 67.9846 | 89.1057 | 64.95 | 0.009072 | 1.0442 | 0.024709 | 23.6627 | 0.023703 | 0.13856 | 1e-06 | 9.1e-05 | 9.1e-05 | 62.2126
2019-03-29 13:44 | 1553881450.527 | -537 | 0.20947 | 0.000459 | 0 | 99.9511 | 99.6203 | 124.2417 | 21.072 | 0.001883 | 0.001783 | -151.784 | 0.23342 | 8.4029 | 0.002744 | 0.004887 | 0.004666 | 6.5e-05 | 0.001157 | 74.865 | 1.3642 | 2.5474 | 100.21 | 20.8 | 36.968 | 66.9394 | 87.6494 | 64.5475 | 0.009073 | 1.0454 | 0.024687 | 23.6136 | 0.029658 | 0.13886 | -3e-06 | -0.000197 | -0.000197 | -28.6504
2019-03-29 13:44 | 1553881451.5274 | -536 | 0.20947 | 0.000458 | 0 | 99.9418 | 99.7247 | 124.2381 | 21.0716 | 0.001937 | 0.001768 | -151.7938 | 0.23285 | 8.4027 | 0.002763 | 0.004276 | 0.00411 | 6.2e-05 | 0.001162 | 74.8752 | 1.3641 | 2.5473 | 100.21 | 20.8 | 36.968 | 68.2549 | 88.1128 | 64.1801 | 0.009073 | 1.0466 | 0.0247 | 23.6004 | 0.019456 | 0.13852 | -1e-06 | -5.6e-05 | -5.6e-05 | -100.2064
2019-03-29 13:44 | 1553881452.5271 | -535 | 0.20942 | 0.000456 | -1e-06 | 99.9368 | 99.6022 | 124.3766 | 21.0725 | 0.001885 | 0.001764 | -151.8002 | 0.23359 | 8.4027 | 0.002713 | 0.004846 | 0.004707 | 6e-05 | 0.001159 | 74.8677 | 1.3207 | 2.5474 | 100.21 | 20.8 | 36.971 | 66.9969 | 86.9026 | 64.0927 | 0.009073 | 1.0469 | 0.024717 | 23.6105 | 0.029996 | 0.13896 | 1e-06 | 0.000135 | 0.000135 | 41.7991
2019-03-29 13:44 | 1553881453.5261 | -534 | 0.20949 | 0.000458 | -1e-06 | 99.8529 | 99.7469 | 124.8066 | 21.0726 | 0.001922 | 0.001813 | -151.7728 | 0.23336 | 8.4009 | 0.002733 | 0.003707 | 0.00358 | 6.4e-05 | 0.001163 | 74.9045 | 1.3606 | 2.5474 | 100.21 | 20.8 | 36.972 | 67.8833 | 87.3712 | 65.2763 | 0.009074 | 1.0432 | 0.024716 | 23.6927 | 0.009506 | 0.13883 | -1e-06 | -7.1e-05 | -7.1e-05 | -79.3729
2019-03-29 13:44 | 1553881454.5276 | -533 | 0.20946 | 0.000458 | 0 | 99.9508 | 99.7612 | 124.9428 | 21.0737 | 0.001939 | 0.001807 | -151.7736 | 0.23345 | 8.4024 | 0.002742 | 0.003968 | 0.003725 | 6e-05 | 0.001161 | 74.8927 | 1.3223 | 2.5475 | 100.21 | 20.8 | 36.971 | 68.3 | 87.5971 | 65.1143 | 0.009073 | 1.0437 | 0.024736 | 23.7 | 0.016984 | 0.13888 | -2e-06 | -0.000149 | -0.000149 | -38.0604
2019-03-29 13:44 | 1553881455.5272 | -532 | 0.20949 | 0.000457 | -1e-06 | 99.8913 | 99.6015 | 124.9132 | 21.0726 | 0.001913 | 0.001826 | -151.7686 | 0.23292 | 8.4019 | 0.002724 | 0.004008 | 0.003909 | 6.1e-05 | 0.001158 | 74.9054 | 1.3108 | 2.5473 | 100.21 | 20.8 | 36.971 | 67.6796 | 87.1712 | 65.5869 | 0.009073 | 1.0422 | 0.024715 | 23.7137 | 0.025984 | 0.13856 | 0 | 3.6e-05 | 3.6e-05 | 156.3836
2019-03-29 13:44 | 1553881456.5272 | -531 | 0.20944 | 0.00046 | 0 | 99.8685 | 99.7308 | 125.0125 | 21.0727 | 0.001858 | 0.001785 | -151.766 | 0.23297 | 8.4019 | 0.002789 | 0.003686 | 0.003504 | 5.9e-05 | 0.001157 | 74.8709 | 1.3268 | 2.5475 | 100.21 | 20.8 | 36.97 | 66.3627 | 88.7419 | 64.6011 | 0.009073 | 1.0453 | 0.024761 | 23.6886 | 0.012339 | 0.1386 | -4e-06 | -0.000313 | -0.000313 | -17.9986
2019-03-29 13:44 | 1553881457.5272 | -530 | 0.20947 | 0.000458 | 0 | 99.8892 | 99.5861 | 124.9442 | 21.0718 | 0.001884 | 0.001781 | -151.766 | 0.23362 | 8.4006 | 0.002742 | 0.003635 | 0.003496 | 6e-05 | 0.001155 | 74.864 | 1.2998 | 2.5473 | 100.21 | 20.8 | 36.97 | 66.982 | 87.6046 | 64.4948 | 0.009073 | 1.0456 | 0.024758 | 23.6784 | 0.027176 | 0.13898 | -2e-06 | -9.1e-05 | -9.1e-05 | -61.8357
2019-03-29 13:44 | 1553881458.5269 | -529 | 0.20949 | 0.000459 | -1e-06 | 99.8757 | 99.6305 | 124.8428 | 21.0731 | 0.00191 | 0.001795 | -151.7722 | 0.23345 | 8.4024 | 0.002725 | 0.003945 | 0.003765 | 6.1e-05 | 0.00116 | 74.8776 | 1.3539 | 2.5473 | 100.21 | 20.8 | 36.968 | 67.595 | 87.189 | 64.8286 | 0.009073 | 1.0446 | 0.024736 | 23.6805 | 0.021986 | 0.13888 | -3e-06 | -0.00018 | -0.00018 | -31.4825
2019-03-29 13:44 | 1553881459.5273 | -528 | 0.20952 | 0.000457 | 0 | 99.9008 | 99.7738 | 124.7562 | 21.0737 | 0.001903 | 0.001802 | -151.7658 | 0.23355 | 8.4017 | 0.002661 | 0.003671 | 0.003469 | 6e-05 | 0.001156 | 74.8614 | 1.3639 | 2.5475 | 100.21 | 20.8 | 36.968 | 67.4434 | 85.6587 | 65.0166 | 0.009073 | 1.044 | 0.024721 | 23.6788 | 0.011378 | 0.13894 | -1e-06 | -2.6e-05 | -2.6e-05 | -218.3956
2019-03-29 13:44 | 1553881460.5263 | -527 | 0.20948 | 0.000456 | 0 | 99.8248 | 99.6831 | 124.8984 | 21.0734 | 0.001885 | 0.001859 | -151.761 | 0.23294 | 8.4028 | 0.002792 | 0.004006 | 0.003792 | 6.2e-05 | 0.001153 | 74.933 | 1.3662 | 2.5473 | 100.21 | 20.8 | 36.968 | 66.9984 | 88.8121 | 66.3744 | 0.009073 | 1.0398 | 0.024685 | 23.7398 | 0.01271 | 0.13858 | 0 | 2.4e-05 | 2.4e-05 | 232.3072
2019-03-29 13:44 | 1553881461.5273 | -526 | 0.20946 | 0.000458 | 0 | 99.8428 | 99.5855 | 125.077 | 21.0731 | 0.001927 | 0.001734 | -151.7589 | 0.23395 | 8.4005 | 0.002724 | 0.004353 | 0.004184 | 5.8e-05 | 0.001149 | 74.8632 | 1.3399 | 2.5474 | 100.21 | 20.8 | 36.968 | 68.0076 | 87.1633 | 63.3623 | 0.009073 | 1.0491 | 0.024813 | 23.6512 | 0.023082 | 0.13918 | -1e-06 | -7.2e-05 | -7.2e-05 | -79.1136
2019-03-29 13:44 | 1553881462.5264 | -525 | 0.20947 | 0.000457 | -1e-06 | 99.8062 | 99.5988 | 125.2826 | 21.0732 | 0.001898 | 0.001796 | -151.7652 | 0.23241 | 8.4026 | 0.002823 | 0.004155 | 0.003975 | 6e-05 | 0.001151 | 74.8828 | 1.3704 | 2.5474 | 100.21 | 20.8 | 36.97 | 67.3073 | 89.5461 | 64.8607 | 0.009073 | 1.0445 | 0.024778 | 23.7233 | 0.018614 | 0.13826 | -1e-06 | 1e-05 | 1e-05 | 547.1538
2019-03-29 13:44 | 1553881463.5273 | -524 | 0.20953 | 0.000459 | 0 | 99.8254 | 99.6868 | 125.0888 | 21.074 | 0.001858 | 0.001748 | -151.7727 | 0.23353 | 8.4019 | 0.002773 | 0.004152 | 0.003984 | 6.2e-05 | 0.001154 | 74.8767 | 1.3603 | 2.5473 | 100.21 | 20.79 | 36.972 | 66.3452 | 88.3572 | 63.7083 | 0.009068 | 1.0481 | 0.024802 | 23.6645 | 0.012427 | 0.13892 | -4e-06 | -0.000328 | -0.000328 | -17.2418
2019-03-29 13:44 | 1553881464.5264 | -523 | 0.20963 | 0.000459 | -1e-06 | 99.8113 | 99.6087 | 125.0162 | 21.0735 | 0.001867 | 0.001773 | -151.7922 | 0.23347 | 8.4014 | 0.002762 | 0.003206 | 0.002997 | 5.9e-05 | 0.00115 | 74.8774 | 1.3254 | 2.5474 | 100.21 | 20.79 | 36.973 | 66.5592 | 88.071 | 64.302 | 0.009068 | 1.0462 | 0.024773 | 23.6784 | 0.018181 | 0.13889 | -2e-06 | -8.7e-05 | -8.7e-05 | -65.0636
2019-03-29 13:44 | 1553881465.5267 | -522 | 0.20937 | 0.000457 | -1e-06 | 99.8239 | 99.5936 | 124.9439 | 21.0739 | 0.001873 | 0.001758 | -151.8121 | 0.23321 | 8.4011 | 0.002768 | 0.004463 | 0.004238 | 6.1e-05 | 0.001152 | 74.9115 | 1.3294 | 2.5474 | 100.21 | 20.79 | 36.978 | 66.7243 | 88.2339 | 63.9496 | 0.00907 | 1.0473 | 0.024778 | 23.6592 | 0.020657 | 0.13873 | 0 | 3.6e-05 | 3.6e-05 | 158.0326
2019-03-29 13:44 | 1553881466.5269 | -521 | 0.20924 | 0.000455 | -1e-06 | 99.847 | 99.5162 | 125.0242 | 21.0732 | 0.001897 | 0.001848 | -151.8137 | 0.23366 | 8.4019 | 0.002763 | 0.003806 | 0.003715 | 6e-05 | 0.001154 | 74.872 | 1.3235 | 2.5475 | 100.21 | 20.79 | 36.981 | 67.2993 | 88.1152 | 66.1087 | 0.00907 | 1.0406 | 0.024707 | 23.7425 | 0.029682 | 0.13901 | 2e-06 | 0.000205 | 0.000205 | 27.6394
2019-03-29 13:44 | 1553881467.5266 | -520 | 0.20958 | 0.00046 | -1e-06 | 99.7691 | 99.6721 | 125.1144 | 21.074 | 0.001992 | 0.001763 | -151.8212 | 0.23287 | 8.4032 | 0.002804 | 0.004568 | 0.004397 | 6.3e-05 | 0.001156 | 74.8478 | 1.3484 | 2.5474 | 100.21 | 20.79 | 36.989 | 69.5786 | 89.1009 | 64.0715 | 0.009072 | 1.0469 | 0.024791 | 23.6796 | 0.0087 | 0.13853 | -4e-06 | -0.000286 | -0.000286 | -19.7265
2019-03-29 13:44 | 1553881468.5265 | -519 | 0.20946 | 0.000456 | -1e-06 | 99.779 | 99.608 | 125.0639 | 21.0745 | 0.001882 | 0.001805 | -151.816 | 0.23331 | 8.4029 | 0.002799 | 0.004475 | 0.004285 | 6e-05 | 0.001161 | 74.8128 | 1.3293 | 2.5474 | 100.21 | 20.79 | 36.989 | 66.9247 | 88.9685 | 65.0768 | 0.009072 | 1.0438 | 0.024749 | 23.7101 | 0.015345 | 0.1388 | 1e-06 | 0.000132 | 0.000132 | 42.7844
2019-03-29 13:44 | 1553881469.5257 | -518 | 0.2095 | 0.000457 | -1e-06 | 99.7643 | 99.5947 | 124.9069 | 21.0743 | 0.001883 | 0.001789 | -151.8141 | 0.23305 | 8.404 | 0.002792 | 0.00394 | 0.003777 | 6.2e-05 | 0.001161 | 74.9025 | 1.3265 | 2.5473 | 100.21 | 20.79 | 37 | 66.9597 | 88.7967 | 64.6924 | 0.009075 | 1.045 | 0.024747 | 23.6818 | 0.015226 | 0.13864 | 0 | 2.9e-05 | 2.9e-05 | 195.2357
2019-03-29 13:44 | 1553881470.5267 | -517 | 0.20956 | 0.000459 | -1e-06 | 99.7916 | 99.68 | 125.212 | 21.0753 | 0.001915 | 0.001839 | -151.814 | 0.23305 | 8.4019 | 0.002721 | 0.004004 | 0.003859 | 6e-05 | 0.001154 | 74.8751 | 1.3353 | 2.5474 | 100.21 | 20.79 | 37 | 67.7351 | 87.0927 | 65.9079 | 0.009075 | 1.0413 | 0.024733 | 23.7533 | 0.010008 | 0.13864 | -2e-06 | -0.000104 | -0.000104 | -54.1081
2019-03-29 13:44 | 1553881471.5261 | -516 | 0.20943 | 0.000456 | -1e-06 | 99.7754 | 99.4974 | 125.0188 | 21.0745 | 0.001928 | 0.001787 | -151.8179 | 0.234 | 8.4028 | 0.002778 | 0.003534 | 0.003381 | 6.3e-05 | 0.001156 | 74.8647 | 1.3281 | 2.5475 | 100.21 | 20.79 | 37.006 | 68.0339 | 88.4719 | 64.6442 | 0.009076 | 1.0451 | 0.02476 | 23.6907 | 0.024958 | 0.13921 | 1e-06 | 8.8e-05 | 8.8e-05 | 64.5108
2019-03-29 13:45 | 1553881472.5268 | -515 | 0.20953 | 0.00046 | -1e-06 | 99.773 | 99.4974 | 125.1688 | 21.0754 | 0.001895 | 0.001714 | -151.8157 | 0.23303 | 8.3995 | 0.00269 | 0.004303 | 0.004052 | 6.1e-05 | 0.001156 | 74.893 | 1.3728 | 2.5474 | 100.21 | 20.79 | 37.006 | 67.2479 | 86.3419 | 62.8968 | 0.009076 | 1.0506 | 0.024839 | 23.6435 | 0.024744 | 0.13863 | -3e-06 | -0.00017 | -0.00017 | -33.1444
2019-03-29 13:45 | 1553881473.5267 | -514 | 0.20945 | 0.000456 | -1e-06 | 99.7446 | 99.3865 | 125.3498 | 21.0752 | 0.001833 | 0.001768 | -151.8146 | 0.23343 | 8.4022 | 0.002747 | 0.00446 | 0.004391 | 6.1e-05 | 0.001158 | 74.8718 | 1.335 | 2.5474 | 100.21 | 20.79 | 37.014 | 65.7417 | 87.7126 | 64.1836 | 0.009078 | 1.0466 | 0.02481 | 23.7058 | 0.032171 | 0.13887 | 2e-06 | 0.00015 | 0.00015 | 37.7515
2019-03-29 13:45 | 1553881474.5271 | -513 | 0.20951 | 0.00046 | -1e-06 | 99.7772 | 99.609 | 125.527 | 21.0758 | 0.001971 | 0.00185 | -151.8112 | 0.23289 | 8.4051 | 0.002697 | 0.0037 | 0.00359 | 6.1e-05 | 0.001154 | 74.9005 | 1.3252 | 2.5473 | 100.21 | 20.79 | 37.014 | 69.0593 | 86.5148 | 66.1613 | 0.009078 | 1.0405 | 0.024755 | 23.792 | 0.015092 | 0.13855 | -3e-06 | -0.000202 | -0.000202 | -27.8956
2019-03-29 13:45 | 1553881475.5271 | -512 | 0.20952 | 0.000455 | -1e-06 | 99.8092 | 99.6752 | 125.8514 | 21.0755 | 0.001891 | 0.001801 | -151.797 | 0.23251 | 8.4007 | 0.002675 | 0.003946 | 0.003724 | 6.5e-05 | 0.001142 | 74.911 | 1.3559 | 2.5474 | 100.21 | 20.79 | 37.018 | 67.1395 | 85.982 | 64.9894 | 0.009079 | 1.0441 | 0.02483 | 23.7816 | 0.01202 | 0.13832 | 3e-06 | 0.000257 | 0.000257 | 21.8996
2019-03-29 13:45 | 1553881476.5269 | -511 | 0.20949 | 0.000462 | -1e-06 | 99.7802 | 99.5957 | 126.0577 | 21.0751 | 0.001916 | 0.001777 | -151.8074 | 0.23431 | 8.4007 | 0.00271 | 0.00435 | 0.004209 | 6.4e-05 | 0.001138 | 74.923 | 1.3799 | 2.5474 | 100.21 | 20.79 | 37.018 | 67.7491 | 86.8303 | 64.395 | 0.009079 | 1.0459 | 0.024872 | 23.7801 | 0.016556 | 0.13939 | -6e-06 | -0.000397 | -0.000397 | -14.2768
2019-03-29 13:45 | 1553881477.5269 | -510 | 0.2095 | 0.000465 | -1e-06 | 99.7449 | 99.5005 | 126.2241 | 21.0761 | 0.001915 | 0.001776 | -151.8119 | 0.23328 | 8.4017 | 0.002725 | 0.003822 | 0.003632 | 6.7e-05 | 0.001144 | 74.8797 | 1.4014 | 2.5475 | 100.21 | 20.79 | 37.021 | 67.7251 | 87.2014 | 64.3797 | 0.00908 | 1.046 | 0.024889 | 23.7953 | 0.021945 | 0.13878 | -9e-06 | -0.000605 | -0.000605 | -9.3354
2019-03-29 13:45 | 1553881478.5263 | -509 | 0.20946 | 0.000462 | 0 | 99.7979 | 99.566 | 126.5267 | 21.0745 | 0.001895 | 0.001834 | -151.8325 | 0.23301 | 8.4033 | 0.002625 | 0.003797 | 0.003678 | 6.1e-05 | 0.001146 | 74.9066 | 1.3652 | 2.5475 | 100.21 | 20.79 | 37.021 | 67.2351 | 84.7908 | 65.7768 | 0.00908 | 1.0417 | 0.024868 | 23.873 | 0.020806 | 0.13861 | -7e-06 | -0.000479 | -0.000479 | -11.7671
2019-03-29 13:45 | 1553881479.526 | -508 | 0.20953 | 0.000463 | -1e-06 | 99.744 | 99.5431 | 126.7964 | 21.0758 | 0.001937 | 0.00181 | -151.824 | 0.23368 | 8.4003 | 0.002566 | 0.003983 | 0.003763 | 6.5e-05 | 0.001143 | 74.8609 | 1.3764 | 2.5473 | 100.21 | 20.79 | 37.022 | 68.2646 | 83.3575 | 65.2099 | 0.00908 | 1.0434 | 0.024915 | 23.8785 | 0.01803 | 0.13901 | -8e-06 | -0.000523 | -0.000523 | -10.8114
2019-03-29 13:45 | 1553881480.527 | -507 | 0.2095 | 0.000464 | -1e-06 | 99.791 | 99.5285 | 126.5189 | 21.0765 | 0.001903 | 0.001794 | -151.8263 | 0.23332 | 8.404 | 0.002773 | 0.004019 | 0.003834 | 6.5e-05 | 0.00114 | 74.9235 | 1.3524 | 2.5475 | 100.21 | 20.79 | 37.024 | 67.436 | 88.3565 | 64.8061 | 0.009081 | 1.0446 | 0.024902 | 23.8381 | 0.023562 | 0.1388 | -9e-06 | -0.000648 | -0.000648 | -8.7164
2019-03-29 13:45 | 1553881481.5273 | -506 | 0.20949 | 0.000459 | -1e-06 | 99.8113 | 99.6435 | 126.7485 | 21.0767 | 0.001905 | 0.001782 | -151.8299 | 0.23306 | 8.4011 | 0.00278 | 0.004465 | 0.004291 | 6.1e-05 | 0.001136 | 74.8826 | 1.3469 | 2.5474 | 100.21 | 20.79 | 37.026 | 67.4752 | 88.5092 | 64.5215 | 0.009081 | 1.0455 | 0.024936 | 23.8497 | 0.015056 | 0.13865 | -3e-06 | -0.000204 | -0.000204 | -27.6244
2019-03-29 13:45 | 1553881482.526 | -505 | 0.20948 | 0.00046 | -1e-06 | 99.7093 | 99.6421 | 126.5232 | 21.077 | 0.001915 | 0.001758 | -151.813 | 0.23343 | 8.4016 | 0.002697 | 0.004031 | 0.003775 | 6.6e-05 | 0.001133 | 74.8715 | 1.3026 | 2.5473 | 100.21 | 20.79 | 37.028 | 67.726 | 86.5275 | 63.9377 | 0.009082 | 1.0473 | 0.024935 | 23.8079 | 0.006027 | 0.13887 | -4e-06 | -0.000281 | -0.000281 | -20.1114
2019-03-29 13:45 | 1553881483.5251 | -504 | 0.20951 | 0.000467 | 0 | 99.7994 | 99.5302 | 126.5793 | 21.0769 | 0.001913 | 0.00179 | -151.8209 | 0.23301 | 8.4041 | 0.002765 | 0.004528 | 0.004433 | 6.3e-05 | 0.001132 | 74.9174 | 1.2899 | 2.5472 | 100.21 | 20.79 | 37.034 | 67.6674 | 88.1499 | 64.7199 | 0.009083 | 1.0449 | 0.024912 | 23.8407 | 0.02416 | 0.13862 | -1.3e-05 | -0.000918 | -0.000918 | -6.147
2019-03-29 13:45 | 1553881484.5264 | -503 | 0.20947 | 0.00046 | 0 | 99.7724 | 99.5848 | 126.5933 | 21.0769 | 0.001923 | 0.001816 | -151.8284 | 0.23171 | 8.4029 | 0.002705 | 0.004257 | 0.004047 | 6e-05 | 0.001136 | 74.9266 | 1.3397 | 2.5474 | 100.21 | 20.79 | 37.037 | 67.9271 | 86.7124 | 65.3507 | 0.009084 | 1.043 | 0.02489 | 23.8643 | 0.016839 | 0.13784 | -5e-06 | -0.000317 | -0.000317 | -17.6907
2019-03-29 13:45 | 1553881485.5252 | -502 | 0.2095 | 0.000465 | -1e-06 | 99.7772 | 99.5288 | 126.679 | 21.0776 | 0.001894 | 0.00183 | -151.8394 | 0.23258 | 8.4022 | 0.00275 | 0.00363 | 0.003493 | 6.3e-05 | 0.001141 | 74.9224 | 1.3365 | 2.5473 | 100.21 | 20.79 | 37.044 | 67.2201 | 87.7852 | 65.6939 | 0.009086 | 1.0419 | 0.024886 | 23.8845 | 0.022297 | 0.13836 | -1e-05 | -0.000711 | -0.000711 | -7.9154
2019-03-29 13:45 | 1553881486.5262 | -501 | 0.20952 | 0.000464 | -1e-06 | 99.7251 | 99.5704 | 126.5478 | 21.0774 | 0.001855 | 0.001817 | -151.829 | 0.23312 | 8.3992 | 0.002676 | 0.00386 | 0.003736 | 6.3e-05 | 0.00114 | 74.8979 | 1.3323 | 2.5474 | 100.21 | 20.79 | 37.049 | 66.2884 | 86.0058 | 65.3669 | 0.009087 | 1.0429 | 0.024885 | 23.8606 | 0.013886 | 0.13868 | -9e-06 | -0.000639 | -0.000639 | -8.8378
2019-03-29 13:45 | 1553881487.5251 | -500 | 0.20951 | 0.000458 | -1e-06 | 99.7371 | 99.425 | 126.3898 | 21.0772 | 0.001913 | 0.001769 | -151.8097 | 0.23265 | 8.4027 | 0.002766 | 0.003994 | 0.003828 | 6.3e-05 | 0.001137 | 74.8876 | 1.3306 | 2.5473 | 100.21 | 20.79 | 37.053 | 67.6661 | 88.1791 | 64.2137 | 0.009088 | 1.0465 | 0.024912 | 23.8051 | 0.028029 | 0.13841 | -1e-06 | -3.8e-05 | -3.8e-05 | -147.8837
2019-03-29 13:45 | 1553881488.5258 | -499 | 0.20949 | 0.00046 | -1e-06 | 99.7829 | 99.6141 | 126.2762 | 21.0769 | 0.001898 | 0.00179 | -151.7894 | 0.23252 | 8.4033 | 0.002697 | 0.00403 | 0.003867 | 6.6e-05 | 0.00114 | 74.9226 | 1.3064 | 2.5473 | 100.21 | 20.79 | 37.062 | 67.3078 | 86.5096 | 64.7099 | 0.00909 | 1.0449 | 0.024882 | 23.8118 | 0.015142 | 0.13833 | -3e-06 | -0.000206 | -0.000206 | -27.3785
2019-03-29 13:45 | 1553881489.5253 | -498 | 0.20949 | 0.00046 | -1e-06 | 99.7584 | 99.5227 | 126.1359 | 21.0772 | 0.001926 | 0.001755 | -151.794 | 0.23252 | 8.4017 | 0.002774 | 0.003768 | 0.003509 | 6.4e-05 | 0.00114 | 74.8804 | 1.3316 | 2.5475 | 100.21 | 20.79 | 37.067 | 67.9938 | 88.3667 | 63.8795 | 0.009091 | 1.0475 | 0.024899 | 23.7693 | 0.021158 | 0.13833 | -3e-06 | -0.000185 | -0.000185 | -30.4853
2019-03-29 13:45 | 1553881490.5263 | -497 | 0.20951 | 0.000463 | -1e-06 | 99.77 | 99.4926 | 126.1539 | 21.0781 | 0.001881 | 0.001692 | -151.7945 | 0.2327 | 8.4057 | 0.002795 | 0.004688 | 0.004508 | 6.3e-05 | 0.001141 | 74.8657 | 1.3411 | 2.5473 | 100.21 | 20.79 | 37.072 | 66.9016 | 88.8728 | 62.3551 | 0.009093 | 1.0523 | 0.024957 | 23.7172 | 0.024905 | 0.13843 | -8e-06 | -0.000592 | -0.000592 | -9.515
2019-03-29 13:45 | 1553881491.5254 | -496 | 0.20949 | 0.000458 | -1e-06 | 99.767 | 99.7172 | 126.0077 | 21.0787 | 0.001868 | 0.001772 | -151.792 | 0.23269 | 8.4008 | 0.002816 | 0.004205 | 0.003973 | 6.4e-05 | 0.001149 | 74.8533 | 1.3304 | 2.5474 | 100.21 | 20.79 | 37.079 | 66.591 | 89.3986 | 64.2876 | 0.009094 | 1.0463 | 0.024871 | 23.7716 | 0.00447 | 0.13843 | -1e-06 | -3.9e-05 | -3.9e-05 | -145.3501
2019-03-29 13:45 | 1553881492.5264 | -495 | 0.20949 | 0.00046 | -1e-06 | 99.7239 | 99.5459 | 126.2152 | 21.0783 | 0.001986 | 0.001793 | -151.7989 | 0.23232 | 8.4027 | 0.002772 | 0.004244 | 0.004024 | 6.5e-05 | 0.001152 | 74.8636 | 1.3209 | 2.5475 | 100.21 | 20.79 | 37.089 | 69.4237 | 88.3217 | 64.7923 | 0.009097 | 1.0447 | 0.024873 | 23.809 | 0.015986 | 0.1382 | -4e-06 | -0.000261 | -0.000261 | -21.5164
2019-03-29 13:45 | 1553881493.5267 | -494 | 0.20947 | 0.00046 | -1e-06 | 99.7129 | 99.5865 | 126.0138 | 21.078 | 0.001908 | 0.001777 | -151.804 | 0.23276 | 8.3993 | 0.002682 | 0.004669 | 0.004447 | 6.6e-05 | 0.001153 | 74.9179 | 1.3631 | 2.5474 | 100.21 | 20.79 | 37.089 | 67.5538 | 86.1678 | 64.3954 | 0.009097 | 1.0459 | 0.024868 | 23.776 | 0.011344 | 0.13847 | -4e-06 | -0.00028 | -0.00028 | -20.1008
2019-03-29 13:45 | 1553881494.5255 | -493 | 0.2095 | 0.000459 | -1e-06 | 99.7476 | 99.5994 | 126.1488 | 21.0783 | 0.001918 | 0.001759 | -151.7818 | 0.23278 | 8.4038 | 0.002722 | 0.004568 | 0.004396 | 6.5e-05 | 0.001155 | 74.928 | 1.3385 | 2.5473 | 100.21 | 20.79 | 37.095 | 67.7873 | 87.1134 | 63.9632 | 0.009098 | 1.0473 | 0.024897 | 23.7735 | 0.013294 | 0.13848 | -3e-06 | -0.000151 | -0.000151 | -37.4116
2019-03-29 13:45 | 1553881495.5265 | -492 | 0.20947 | 0.000458 | -1e-06 | 99.808 | 99.408 | 126.2539 | 21.0796 | 0.001911 | 0.001785 | -151.7805 | 0.23217 | 8.4025 | 0.002665 | 0.003965 | 0.003799 | 6.3e-05 | 0.001156 | 74.8544 | 1.3369 | 2.5474 | 100.21 | 20.79 | 37.095 | 67.62 | 85.7541 | 64.6032 | 0.009098 | 1.0453 | 0.024884 | 23.806 | 0.035925 | 0.13812 | -2e-06 | -8.7e-05 | -8.7e-05 | -64.6088
2019-03-29 13:45 | 1553881496.5255 | -491 | 0.20952 | 0.000462 | 0 | 99.7739 | 99.4175 | 126.253 | 21.0792 | 0.001964 | 0.001802 | -151.7868 | 0.23198 | 8.4011 | 0.002745 | 0.004039 | 0.003904 | 6.4e-05 | 0.00116 | 74.9185 | 1.3376 | 2.5475 | 100.21 | 20.78 | 37.1 | 68.9058 | 87.6775 | 65.0017 | 0.009094 | 1.044 | 0.024869 | 23.8199 | 0.032006 | 0.138 | -7e-06 | -0.000499 | -0.000499 | -11.2578
2019-03-29 13:45 | 1553881497.5266 | -490 | 0.20949 | 0.000458 | 0 | 99.7359 | 99.6953 | 126.1453 | 21.0785 | 0.001903 | 0.001795 | -151.7897 | 0.23295 | 8.4022 | 0.002721 | 0.004067 | 0.003856 | 6.5e-05 | 0.001158 | 74.8803 | 1.3652 | 2.5475 | 100.21 | 20.78 | 37.1 | 67.4317 | 87.0927 | 64.8411 | 0.009094 | 1.0445 | 0.024864 | 23.8041 | 0.003638 | 0.13858 | -2e-06 | -0.000152 | -0.000152 | -37.1457
2019-03-29 13:45 | 1553881498.5267 | -489 | 0.20949 | 0.000461 | -1e-06 | 99.7697 | 99.4701 | 125.9542 | 21.079 | 0.001868 | 0.001775 | -151.7923 | 0.23258 | 8.4029 | 0.002809 | 0.004214 | 0.004055 | 6.2e-05 | 0.00116 | 74.8919 | 1.3649 | 2.5476 | 100.21 | 20.78 | 37.104 | 66.5982 | 89.2074 | 64.3612 | 0.009095 | 1.046 | 0.024863 | 23.7692 | 0.026903 | 0.13836 | -5e-06 | -0.000332 | -0.000332 | -16.9848
2019-03-29 13:45 | 1553881499.5249 | -488 | 0.2095 | 0.000458 | -1e-06 | 99.7332 | 99.4523 | 126.2394 | 21.0787 | 0.001892 | 0.001827 | -151.8093 | 0.23316 | 8.4021 | 0.002788 | 0.004226 | 0.004051 | 6.8e-05 | 0.001154 | 74.8943 | 1.3112 | 2.5475 | 100.21 | 20.78 | 37.106 | 67.1785 | 88.7244 | 65.6077 | 0.009095 | 1.0422 | 0.024845 | 23.84 | 0.025225 | 0.1387 | -1e-06 | -1.9e-05 | -1.9e-05 | -296.8456
2019-03-29 13:45 | 1553881500.5252 | -487 | 0.20946 | 0.000457 | 0 | 99.7811 | 99.5899 | 126.4396 | 21.0792 | 0.00191 | 0.001833 | -151.8021 | 0.23223 | 8.4004 | 0.002766 | 0.004016 | 0.003809 | 6.8e-05 | 0.00115 | 74.9158 | 1.322 | 2.5472 | 100.21 | 20.78 | 37.108 | 67.6095 | 88.1883 | 65.7633 | 0.009096 | 1.0417 | 0.024859 | 23.8644 | 0.017157 | 0.13816 | 0 | 3e-06 | 3e-06 | 1827.9065
2019-03-29 13:45 | 1553881501.5246 | -486 | 0.2095 | 0.000457 | -1e-06 | 99.7625 | 99.6803 | 126.3917 | 21.0791 | 0.001892 | 0.001806 | -151.7981 | 0.23273 | 8.403 | 0.002795 | 0.004244 | 0.004098 | 6.1e-05 | 0.001155 | 74.8849 | 1.3634 | 2.5474 | 100.21 | 20.78 | 37.111 | 67.1843 | 88.8824 | 65.1101 | 0.009097 | 1.0437 | 0.024879 | 23.8368 | 0.007374 | 0.13845 | -1e-06 | -2.9e-05 | -2.9e-05 | -194.387
2019-03-29 13:45 | 1553881502.525 | -485 | 0.20951 | 0.000457 | -1e-06 | 99.7751 | 99.7186 | 126.0462 | 21.0802 | 0.001876 | 0.001777 | -151.7964 | 0.23246 | 8.4015 | 0.002834 | 0.004289 | 0.00413 | 6.1e-05 | 0.00116 | 74.8553 | 1.3724 | 2.5474 | 100.21 | 20.78 | 37.114 | 66.7992 | 89.8116 | 64.4062 | 0.009097 | 1.0459 | 0.024871 | 23.7795 | 0.005072 | 0.13829 | 0 | 9.6e-05 | 9.6e-05 | 58.7556
2019-03-29 13:45 | 1553881503.5257 | -484 | 0.20947 | 0.000455 | -1e-06 | 99.8033 | 99.4902 | 125.8885 | 21.0796 | 0.001887 | 0.001803 | -151.7704 | 0.23205 | 8.4023 | 0.00287 | 0.004227 | 0.004085 | 6.3e-05 | 0.00117 | 74.9003 | 1.3936 | 2.5473 | 100.21 | 20.78 | 37.118 | 67.0653 | 90.6863 | 65.0435 | 0.009098 | 1.0439 | 0.024832 | 23.787 | 0.028098 | 0.13805 | 2e-06 | 0.000227 | 0.000227 | 24.7689
2019-03-29 13:45 | 1553881504.5258 | -483 | 0.20949 | 0.000463 | -1e-06 | 99.741 | 99.6581 | 125.7627 | 21.0797 | 0.001919 | 0.001806 | -151.7819 | 0.2329 | 8.4005 | 0.002803 | 0.003608 | 0.003438 | 6.4e-05 | 0.00117 | 74.8882 | 1.3428 | 2.5473 | 100.21 | 20.78 | 37.126 | 67.8183 | 89.0697 | 65.1069 | 0.0091 | 1.0437 | 0.024817 | 23.7773 | 0.007433 | 0.13855 | -8e-06 | -0.00053 | -0.00053 | -10.6333
2019-03-29 13:45 | 1553881505.5258 | -482 | 0.20949 | 0.000457 | -1e-06 | 99.7931 | 99.5083 | 125.3754 | 21.0796 | 0.001896 | 0.001789 | -151.7642 | 0.232 | 8.403 | 0.002884 | 0.004104 | 0.003953 | 6.3e-05 | 0.001164 | 74.9027 | 1.3205 | 2.5474 | 100.21 | 20.78 | 37.13 | 67.2791 | 91.0224 | 64.6905 | 0.009101 | 1.045 | 0.024794 | 23.7261 | 0.02556 | 0.13802 | 0 | 6.4e-05 | 6.4e-05 | 87.8812
2019-03-29 13:45 | 1553881506.5249 | -481 | 0.20945 | 0.00046 | -1e-06 | 99.7302 | 99.408 | 125.2437 | 21.0796 | 0.001881 | 0.001827 | -151.738 | 0.23277 | 8.4017 | 0.002869 | 0.003694 | 0.003421 | 6.3e-05 | 0.001166 | 74.9267 | 1.3146 | 2.5474 | 100.21 | 20.78 | 37.14 | 66.9116 | 90.6567 | 65.6218 | 0.009104 | 1.0421 | 0.024747 | 23.7463 | 0.028947 | 0.13848 | -4e-06 | -0.000241 | -0.000241 | -23.3387
2019-03-29 13:45 | 1553881507.5253 | -480 | 0.20948 | 0.00046 | -1e-06 | 99.7673 | 99.5237 | 125.2496 | 21.0796 | 0.001898 | 0.001797 | -151.752 | 0.2328 | 8.4023 | 0.002847 | 0.004483 | 0.004377 | 6.1e-05 | 0.001166 | 74.8727 | 1.3503 | 2.5474 | 100.21 | 20.79 | 37.145 | 67.3306 | 90.1253 | 64.8842 | 0.00911 | 1.0444 | 0.024774 | 23.721 | 0.021872 | 0.13849 | -3e-06 | -0.00021 | -0.00021 | -26.8533
2019-03-29 13:45 | 1553881508.5248 | -479 | 0.20947 | 0.000457 | -1e-06 | 99.8176 | 99.4308 | 125.5491 | 21.0798 | 0.001904 | 0.001737 | -151.7502 | 0.23269 | 8.4025 | 0.002821 | 0.003868 | 0.00365 | 6.1e-05 | 0.001164 | 74.8806 | 1.3694 | 2.5476 | 100.21 | 20.79 | 37.15 | 67.4548 | 89.4988 | 63.4523 | 0.009112 | 1.0489 | 0.024857 | 23.699 | 0.034726 | 0.13843 | 0 | 4.1e-05 | 4.1e-05 | 139.0449
2019-03-29 13:45 | 1553881509.5253 | -478 | 0.20947 | 0.00046 | 0 | 99.7997 | 99.5411 | 125.3082 | 21.0802 | 0.001987 | 0.001758 | -151.7403 | 0.23193 | 8.4034 | 0.002826 | 0.003719 | 0.003682 | 6.4e-05 | 0.001168 | 74.8959 | 1.3452 | 2.5473 | 100.21 | 20.78 | 37.168 | 69.4668 | 89.6406 | 63.9596 | 0.00911 | 1.0473 | 0.024814 | 23.694 | 0.023205 | 0.13797 | -4e-06 | -0.000279 | -0.000279 | -20.0948
2019-03-29 13:45 | 1553881510.5263 | -477 | 0.20948 | 0.00046 | -1e-06 | 99.8128 | 99.6151 | 124.8444 | 21.0804 | 0.001939 | 0.001802 | -151.7382 | 0.23295 | 8.4018 | 0.002723 | 0.004151 | 0.003928 | 6.3e-05 | 0.001172 | 74.8688 | 1.321 | 2.5472 | 100.21 | 20.78 | 37.168 | 68.3072 | 87.1506 | 65.0036 | 0.00911 | 1.044 | 0.02473 | 23.6867 | 0.017733 | 0.13858 | -4e-06 | -0.000282 | -0.000282 | -20.0347
2019-03-29 13:45 | 1553881511.526 | -476 | 0.20951 | 0.000456 | -1e-06 | 99.8122 | 99.4916 | 124.9165 | 21.0803 | 0.00192 | 0.001834 | -151.723 | 0.23308 | 8.4023 | 0.002803 | 0.004188 | 0.004003 | 6.1e-05 | 0.00117 | 74.9212 | 1.3248 | 2.5475 | 100.21 | 20.78 | 37.18 | 67.847 | 89.064 | 65.7902 | 0.009113 | 1.0416 | 0.024708 | 23.7211 | 0.02878 | 0.13866 | 1e-06 | 0.000107 | 0.000107 | 52.5155
2019-03-29 13:45 | 1553881512.525 | -475 | 0.2095 | 0.000459 | -1e-06 | 99.7539 | 99.4093 | 125.1054 | 21.0806 | 0.001894 | 0.001821 | -151.7151 | 0.23326 | 8.4002 | 0.002811 | 0.004559 | 0.004423 | 5.8e-05 | 0.001174 | 74.8386 | 1.3368 | 2.5474 | 100.21 | 20.78 | 37.186 | 67.2146 | 89.2682 | 65.4729 | 0.009115 | 1.0426 | 0.024739 | 23.7279 | 0.030945 | 0.13877 | -3e-06 | -0.000181 | -0.000181 | -31.1403
2019-03-29 13:45 | 1553881513.5263 | -474 | 0.20946 | 0.000458 | -1e-06 | 99.791 | 99.5766 | 125.2648 | 21.0806 | 0.001892 | 0.001795 | -151.7107 | 0.2333 | 8.4048 | 0.002737 | 0.004032 | 0.003869 | 6.1e-05 | 0.001173 | 74.8266 | 1.3516 | 2.5475 | 100.21 | 20.78 | 37.203 | 67.1692 | 87.4953 | 64.8508 | 0.009119 | 1.0445 | 0.024777 | 23.7212 | 0.019238 | 0.13879 | -2e-06 | -8.2e-05 | -8.2e-05 | -68.4831
2019-03-29 13:45 | 1553881514.525 | -473 | 0.2095 | 0.000456 | 0 | 99.8491 | 99.5991 | 124.8998 | 21.0803 | 0.001908 | 0.001756 | -151.7098 | 0.23269 | 8.4032 | 0.002835 | 0.003809 | 0.003555 | 5.5e-05 | 0.001171 | 74.8654 | 1.3634 | 2.5473 | 100.21 | 20.78 | 37.203 | 67.5606 | 89.836 | 63.9012 | 0.009119 | 1.0475 | 0.024776 | 23.6534 | 0.022419 | 0.13843 | 0 | 2e-05 | 2e-05 | 283.3879
2019-03-29 13:45 | 1553881515.5254 | -472 | 0.20951 | 0.000463 | 0 | 99.8179 | 99.6895 | 124.8477 | 21.0812 | 0.00189 | 0.001767 | -151.7158 | 0.23319 | 8.404 | 0.002819 | 0.003978 | 0.003821 | 5.7e-05 | 0.001169 | 74.8609 | 1.3382 | 2.5473 | 100.21 | 20.78 | 37.208 | 67.1176 | 89.4665 | 64.1651 | 0.00912 | 1.0466 | 0.024761 | 23.6577 | 0.011511 | 0.13873 | -8e-06 | -0.000605 | -0.000605 | -9.332
2019-03-29 13:45 | 1553881516.5252 | -471 | 0.20951 | 0.000459 | 0 | 99.8006 | 99.5544 | 124.8306 | 21.0811 | 0.001958 | 0.001753 | -151.7382 | 0.23295 | 8.4032 | 0.002772 | 0.003984 | 0.00387 | 6e-05 | 0.001176 | 74.8675 | 1.3378 | 2.5474 | 100.21 | 20.78 | 37.211 | 68.7572 | 88.3283 | 63.8403 | 0.009121 | 1.0476 | 0.024771 | 23.6447 | 0.02209 | 0.13858 | -3e-06 | -0.000206 | -0.000206 | -27.3819
2019-03-29 13:45 | 1553881517.5252 | -470 | 0.20952 | 0.000459 | 0 | 99.7997 | 99.6964 | 125.2244 | 21.0808 | 0.001892 | 0.001817 | -151.7445 | 0.233 | 8.4021 | 0.002734 | 0.004189 | 0.003963 | 5.7e-05 | 0.001171 | 74.9198 | 1.364 | 2.5475 | 100.21 | 20.78 | 37.22 | 67.1835 | 87.4143 | 65.3691 | 0.009123 | 1.0429 | 0.024754 | 23.7356 | 0.009262 | 0.13861 | -4e-06 | -0.00026 | -0.00026 | -21.6944
2019-03-29 13:45 | 1553881518.5242 | -469 | 0.20952 | 0.000459 | -1e-06 | 99.8634 | 99.65 | 125.1381 | 21.0814 | 0.001946 | 0.001825 | -151.7431 | 0.23338 | 8.4023 | 0.002685 | 0.003464 | 0.003262 | 5.6e-05 | 0.001166 | 74.8957 | 1.3935 | 2.5474 | 100.21 | 20.78 | 37.224 | 68.4791 | 86.2298 | 65.5588 | 0.009124 | 1.0423 | 0.024739 | 23.734 | 0.01914 | 0.13884 | -3e-06 | -0.000149 | -0.000149 | -37.8782
2019-03-29 13:45 | 1553881519.5247 | -468 | 0.20949 | 0.000454 | -1e-06 | 99.844 | 99.7435 | 124.9557 | 21.0809 | 0.001915 | 0.001792 | -151.7408 | 0.23369 | 8.4037 | 0.002769 | 0.004323 | 0.004197 | 5.6e-05 | 0.001165 | 74.935 | 1.3798 | 2.5475 | 100.21 | 20.78 | 37.233 | 67.7196 | 88.2484 | 64.7825 | 0.009126 | 1.0447 | 0.024749 | 23.6896 | 0.009008 | 0.13902 | 3e-06 | 0.000263 | 0.000263 | 21.5148
2019-03-29 13:45 | 1553881520.5243 | -467 | 0.20948 | 0.000458 | 0 | 99.8488 | 99.611 | 125.0209 | 21.0813 | 0.001926 | 0.001821 | -151.7531 | 0.23146 | 8.4032 | 0.002754 | 0.003468 | 0.003224 | 5.4e-05 | 0.001169 | 74.9106 | 1.3398 | 2.5473 | 100.21 | 20.78 | 37.237 | 68.0061 | 87.8872 | 65.4791 | 0.009127 | 1.0426 | 0.02473 | 23.7201 | 0.021319 | 0.13769 | -2e-06 | -0.000134 | -0.000134 | -41.8622
2019-03-29 13:45 | 1553881521.5243 | -466 | 0.20948 | 0.000457 | -1e-06 | 99.8227 | 99.6012 | 124.9254 | 21.0815 | 0.00196 | 0.001806 | -151.7633 | 0.23243 | 8.402 | 0.002751 | 0.003834 | 0.003649 | 5.5e-05 | 0.001172 | 74.8763 | 1.3318 | 2.5474 | 100.21 | 20.78 | 37.243 | 68.816 | 87.8297 | 65.1156 | 0.009129 | 1.0437 | 0.024734 | 23.6984 | 0.019875 | 0.13827 | 0 | 2e-06 | 2e-06 | 3303.0309
2019-03-29 13:45 | 1553881522.5244 | -465 | 0.20949 | 0.000456 | -1e-06 | 99.861 | 99.6049 | 125.0205 | 21.082 | 0.001976 | 0.001782 | -151.7701 | 0.2331 | 8.403 | 0.00271 | 0.004403 | 0.004147 | 5.5e-05 | 0.001171 | 74.8297 | 1.3531 | 2.5474 | 100.21 | 20.78 | 37.247 | 69.1941 | 86.8432 | 64.5259 | 0.00913 | 1.0455 | 0.024765 | 23.6867 | 0.02297 | 0.13867 | 1e-06 | 0.000122 | 0.000122 | 46.4371
2019-03-29 13:45 | 1553881523.5245 | -464 | 0.20949 | 0.000453 | -1e-06 | 99.8272 | 99.5882 | 125.1251 | 21.0818 | 0.001948 | 0.001811 | -151.7839 | 0.23278 | 8.4036 | 0.002708 | 0.003587 | 0.00347 | 6e-05 | 0.001172 | 74.8577 | 1.327 | 2.5475 | 100.21 | 20.78 | 37.251 | 68.5247 | 86.793 | 65.2196 | 0.009131 | 1.0434 | 0.02475 | 23.7209 | 0.021442 | 0.13848 | 5e-06 | 0.000419 | 0.000419 | 13.4572
2019-03-29 13:45 | 1553881524.5255 | -463 | 0.20949 | 0.000453 | -1e-06 | 99.8954 | 99.5175 | 125.3632 | 21.083 | 0.001952 | 0.001841 | -151.7856 | 0.23271 | 8.4034 | 0.002787 | 0.005166 | 0.004935 | 5.7e-05 | 0.001177 | 74.8264 | 1.3372 | 2.5473 | 100.21 | 20.77 | 37.263 | 68.6256 | 88.6956 | 65.9412 | 0.009128 | 1.0412 | 0.024747 | 23.7688 | 0.033902 | 0.13844 | 4e-06 | 0.000368 | 0.000368 | 15.3264
2019-03-29 13:45 | 1553881525.5257 | -462 | 0.20948 | 0.000455 | 0 | 99.8077 | 99.6025 | 125.3115 | 21.0823 | 0.001892 | 0.001819 | -151.7996 | 0.23289 | 8.4032 | 0.002711 | 0.004747 | 0.004612 | 5.6e-05 | 0.001172 | 74.769 | 1.3161 | 2.5472 | 100.21 | 20.77 | 37.263 | 67.1819 | 86.8543 | 65.4168 | 0.009128 | 1.0428 | 0.024761 | 23.7455 | 0.018411 | 0.13855 | 1e-06 | 0.000129 | 0.000129 | 43.7577
2019-03-29 13:45 | 1553881526.5257 | -461 | 0.20945 | 0.000451 | -1e-06 | 99.8314 | 99.711 | 125.2763 | 21.0823 | 0.001949 | 0.001835 | -151.7954 | 0.23212 | 8.4012 | 0.00277 | 0.004479 | 0.004306 | 5.7e-05 | 0.001172 | 74.856 | 1.3293 | 2.5474 | 100.21 | 20.77 | 37.272 | 68.5516 | 88.2873 | 65.8149 | 0.00913 | 1.0415 | 0.024743 | 23.7561 | 0.010789 | 0.13809 | 8e-06 | 0.000649 | 0.000649 | 8.6536
2019-03-29 13:45 | 1553881527.5259 | -460 | 0.20946 | 0.000456 | -1e-06 | 99.852 | 99.7387 | 125.3357 | 21.0832 | 0.001934 | 0.001774 | -151.7682 | 0.23258 | 8.4042 | 0.002761 | 0.004269 | 0.003987 | 5.5e-05 | 0.001173 | 74.8506 | 1.35 | 2.5474 | 100.21 | 20.77 | 37.272 | 68.1976 | 88.0749 | 64.3451 | 0.00913 | 1.0461 | 0.024803 | 23.7102 | 0.01016 | 0.13836 | 1e-06 | 0.000139 | 0.000139 | 40.3743
2019-03-29 13:45 | 1553881528.5247 | -459 | 0.20946 | 0.000455 | -1e-06 | 99.8805 | 99.7158 | 125.4308 | 21.0828 | 0.00192 | 0.001794 | -151.7651 | 0.23236 | 8.4015 | 0.002746 | 0.004817 | 0.004635 | 5.5e-05 | 0.001181 | 74.8855 | 1.3259 | 2.5473 | 100.21 | 20.77 | 37.274 | 67.8439 | 87.7113 | 64.8178 | 0.009131 | 1.0446 | 0.024795 | 23.7358 | 0.014758 | 0.13823 | 2e-06 | 0.000189 | 0.000189 | 29.8079
2019-03-29 13:45 | 1553881529.5246 | -458 | 0.20949 | 0.000453 | -1e-06 | 99.8775 | 99.7138 | 125.8371 | 21.0828 | 0.00195 | 0.001758 | -151.7667 | 0.23321 | 8.4025 | 0.002668 | 0.004307 | 0.004123 | 5.8e-05 | 0.001183 | 74.9018 | 1.3373 | 2.5473 | 100.21 | 20.77 | 37.277 | 68.584 | 85.818 | 63.9606 | 0.009132 | 1.0473 | 0.024866 | 23.744 | 0.014673 | 0.13873 | 5e-06 | 0.000394 | 0.000394 | 14.3185
2019-03-29 13:45 | 1553881530.5239 | -457 | 0.20949 | 0.000456 | 0 | 99.9194 | 99.6302 | 125.8965 | 21.0827 | 0.001914 | 0.001812 | -151.7589 | 0.2331 | 8.4016 | 0.002699 | 0.003847 | 0.003709 | 5.9e-05 | 0.001185 | 74.8636 | 1.3144 | 2.5474 | 100.21 | 20.77 | 37.278 | 67.712 | 86.5862 | 65.258 | 0.009132 | 1.0433 | 0.024825 | 23.7953 | 0.025929 | 0.13867 | 1e-06 | 7.6e-05 | 7.6e-05 | 73.9645
2019-03-29 13:46 | 1553881531.5236 | -456 | 0.20949 | 0.000453 | -1e-06 | 99.9619 | 99.6735 | 126.1016 | 21.0827 | 0.001914 | 0.001828 | -151.7709 | 0.2334 | 8.4017 | 0.002683 | 0.004885 | 0.004686 | 5.7e-05 | 0.001186 | 74.8046 | 1.3217 | 2.5473 | 100.21 | 20.77 | 37.278 | 67.7002 | 86.199 | 65.6498 | 0.009132 | 1.042 | 0.02483 | 23.8284 | 0.025843 | 0.13885 | 5e-06 | 0.000403 | 0.000403 | 14.0092
2019-03-29 13:46 | 1553881532.5237 | -455 | 0.2095 | 0.000457 | 0 | 99.858 | 99.6322 | 126.1469 | 21.0834 | 0.001884 | 0.001759 | -151.7869 | 0.23342 | 8.402 | 0.002695 | 0.004262 | 0.004098 | 6e-05 | 0.001187 | 74.8358 | 1.3508 | 2.5474 | 100.21 | 20.77 | 37.278 | 66.981 | 86.4878 | 63.9902 | 0.009132 | 1.0472 | 0.024896 | 23.7743 | 0.020251 | 0.13886 | -1e-06 | -3.2e-05 | -3.2e-05 | -175.0958
2019-03-29 13:46 | 1553881533.5237 | -454 | 0.20953 | 0.000456 | 0 | 99.8829 | 99.8148 | 126.0054 | 21.083 | 0.001959 | 0.001754 | -151.7848 | 0.23289 | 8.4013 | 0.002687 | 0.003753 | 0.003586 | 6.1e-05 | 0.001193 | 74.8715 | 1.3509 | 2.5473 | 100.21 | 20.77 | 37.279 | 68.7792 | 86.2883 | 63.8523 | 0.009132 | 1.0476 | 0.024887 | 23.7561 | 0.006099 | 0.13855 | 1e-06 | 7.4e-05 | 7.4e-05 | 76.579
2019-03-29 13:46 | 1553881534.5252 | -453 | 0.20951 | 0.000455 | 0 | 99.9038 | 99.7588 | 125.726 | 21.0833 | 0.001921 | 0.001754 | -151.777 | 0.23304 | 8.403 | 0.002692 | 0.00425 | 0.004114 | 5.9e-05 | 0.001194 | 74.9023 | 1.3524 | 2.5473 | 100.21 | 20.77 | 37.277 | 67.8763 | 86.4034 | 63.8651 | 0.009132 | 1.0476 | 0.024859 | 23.7302 | 0.012992 | 0.13863 | 1e-06 | 0.000135 | 0.000135 | 41.6926
2019-03-29 13:46 | 1553881535.5248 | -452 | 0.20947 | 0.000454 | -1e-06 | 99.9508 | 99.6267 | 125.9518 | 21.0831 | 0.001928 | 0.001826 | -151.7727 | 0.23217 | 8.4008 | 0.002684 | 0.003762 | 0.003554 | 5.9e-05 | 0.001185 | 74.8487 | 1.3813 | 2.5474 | 100.21 | 20.77 | 37.276 | 68.0493 | 86.211 | 65.601 | 0.009131 | 1.0422 | 0.024817 | 23.8126 | 0.029049 | 0.13812 | 4e-06 | 0.000311 | 0.000311 | 18.0467
2019-03-29 13:46 | 1553881536.5238 | -451 | 0.2095 | 0.000452 | 0 | 99.9266 | 99.8025 | 126.2532 | 21.0833 | 0.001974 | 0.001756 | -151.774 | 0.2328 | 8.4036 | 0.002715 | 0.003911 | 0.003817 | 5.8e-05 | 0.001187 | 74.8151 | 1.3719 | 2.5474 | 100.21 | 20.77 | 37.272 | 69.1603 | 86.9506 | 63.9017 | 0.00913 | 1.0475 | 0.02491 | 23.7812 | 0.011113 | 0.1385 | 5e-06 | 0.000394 | 0.000394 | 14.3028
2019-03-29 13:46 | 1553881537.5242 | -450 | 0.2095 | 0.000454 | -1e-06 | 99.9814 | 99.7912 | 126.176 | 21.0835 | 0.001932 | 0.001774 | -151.7759 | 0.23318 | 8.4028 | 0.002737 | 0.004006 | 0.003824 | 5.6e-05 | 0.001185 | 74.7742 | 1.3676 | 2.5474 | 100.21 | 20.77 | 37.271 | 68.1449 | 87.4949 | 64.3322 | 0.00913 | 1.0461 | 0.024886 | 23.7891 | 0.017024 | 0.13872 | 3e-06 | 0.000282 | 0.000282 | 20.0237
2019-03-29 13:46 | 1553881538.5248 | -449 | 0.20949 | 0.000452 | -1e-06 | 99.9412 | 99.6807 | 126.2133 | 21.083 | 0.001958 | 0.001807 | -151.7644 | 0.23227 | 8.4007 | 0.002735 | 0.004431 | 0.004272 | 5.9e-05 | 0.001177 | 74.8119 | 1.3446 | 2.5474 | 100.21 | 20.77 | 37.27 | 68.7535 | 87.4328 | 65.131 | 0.00913 | 1.0436 | 0.02486 | 23.8207 | 0.023352 | 0.13818 | 6e-06 | 0.000454 | 0.000454 | 12.3754
2019-03-29 13:46 | 1553881539.5249 | -448 | 0.20954 | 0.000456 | -1e-06 | 99.9969 | 99.6452 | 126.041 | 21.0826 | 0.001897 | 0.001784 | -151.7601 | 0.23315 | 8.402 | 0.002769 | 0.00455 | 0.004392 | 5.8e-05 | 0.001177 | 74.8708 | 1.3428 | 2.5474 | 100.21 | 20.76 | 37.269 | 67.2913 | 88.2524 | 64.581 | 0.009124 | 1.0453 | 0.024864 | 23.7851 | 0.031518 | 0.1387 | 0 | 7.8e-05 | 7.8e-05 | 72.0891
2019-03-29 13:46 | 1553881540.5255 | -447 | 0.20952 | 0.000458 | 0 | 99.9784 | 99.8401 | 126.137 | 21.0833 | 0.001878 | 0.001843 | -151.7538 | 0.23241 | 8.4007 | 0.002701 | 0.004293 | 0.004091 | 5.9e-05 | 0.001168 | 74.825 | 1.3058 | 2.5473 | 100.21 | 20.76 | 37.269 | 66.8364 | 86.6289 | 66.0047 | 0.009124 | 1.041 | 0.024821 | 23.8443 | 0.012383 | 0.13826 | -1e-06 | -8.6e-05 | -8.6e-05 | -65.6496
2019-03-29 13:46 | 1553881541.5249 | -446 | 0.20952 | 0.000457 | 0 | 100.0269 | 99.8588 | 125.9044 | 21.0836 | 0.001964 | 0.001806 | -151.7438 | 0.23327 | 8.4001 | 0.002667 | 0.003973 | 0.00378 | 5.7e-05 | 0.00117 | 74.8607 | 1.3216 | 2.5472 | 100.21 | 20.76 | 37.269 | 68.9219 | 85.8025 | 65.1167 | 0.009124 | 1.0437 | 0.02483 | 23.7911 | 0.015038 | 0.13877 | 0 | -6e-06 | -6e-06 | -992.3322
2019-03-29 13:46 | 1553881542.5255 | -445 | 0.2095 | 0.000458 | -1e-06 | 100.0101 | 99.7796 | 125.5002 | 21.084 | 0.00187 | 0.001763 | -151.7302 | 0.2327 | 8.404 | 0.002702 | 0.003888 | 0.003704 | 5.4e-05 | 0.001169 | 74.8474 | 1.3329 | 2.5474 | 100.21 | 20.76 | 37.272 | 66.6499 | 86.6445 | 64.0874 | 0.009125 | 1.0469 | 0.024828 | 23.7167 | 0.020634 | 0.13843 | -2e-06 | -6.4e-05 | -6.4e-05 | -88.2992
2019-03-29 13:46 | 1553881543.525 | -444 | 0.20952 | 0.000459 | -1e-06 | 100.0083 | 99.5977 | 125.4977 | 21.0841 | 0.001896 | 0.001802 | -151.735 | 0.23225 | 8.4026 | 0.002764 | 0.004748 | 0.004579 | 5.7e-05 | 0.001164 | 74.8546 | 1.3608 | 2.5474 | 100.21 | 20.76 | 37.272 | 67.2758 | 88.1516 | 65.0058 | 0.009125 | 1.044 | 0.024794 | 23.7487 | 0.036795 | 0.13816 | -2e-06 | -0.000101 | -0.000101 | -55.7012
2019-03-29 13:46 | 1553881544.524 | -443 | 0.20947 | 0.000457 | -1e-06 | 99.9778 | 99.6114 | 125.8244 | 21.0835 | 0.001925 | 0.001836 | -151.7395 | 0.2325 | 8.4027 | 0.002694 | 0.004378 | 0.004215 | 5.7e-05 | 0.001166 | 74.8609 | 1.3401 | 2.5474 | 100.21 | 20.76 | 37.277 | 67.9703 | 86.4652 | 65.8245 | 0.009126 | 1.0415 | 0.024797 | 23.8084 | 0.032838 | 0.13832 | 0 | 1.7e-05 | 1.7e-05 | 339.1058
2019-03-29 13:46 | 1553881545.5245 | -442 | 0.20951 | 0.000457 | -1e-06 | 100.0047 | 99.708 | 125.6451 | 21.0843 | 0.001952 | 0.001828 | -151.7444 | 0.23151 | 8.4026 | 0.002703 | 0.004244 | 0.004063 | 5.7e-05 | 0.001165 | 74.8758 | 1.3775 | 2.5475 | 100.21 | 20.76 | 37.28 | 68.6182 | 86.6837 | 65.6326 | 0.009127 | 1.0421 | 0.024786 | 23.7847 | 0.026579 | 0.13772 | 0 | 6e-05 | 6e-05 | 93.2761
2019-03-29 13:46 | 1553881546.5231 | -441 | 0.2095 | 0.000456 | -1e-06 | 100.0595 | 99.8192 | 125.7751 | 21.0857 | 0.00188 | 0.001788 | -151.7386 | 0.23293 | 8.4007 | 0.002699 | 0.004642 | 0.004501 | 5.7e-05 | 0.001169 | 74.8908 | 1.3805 | 2.5475 | 100.21 | 20.76 | 37.287 | 66.888 | 86.5829 | 64.6739 | 0.009128 | 1.0451 | 0.024834 | 23.7633 | 0.021502 | 0.13857 | 1e-06 | 0.000108 | 0.000108 | 52.3837
2019-03-29 13:46 | 1553881547.5241 | -440 | 0.20948 | 0.000457 | -1e-06 | 100.0245 | 99.9554 | 125.2974 | 21.0849 | 0.001954 | 0.001792 | -151.749 | 0.2333 | 8.4023 | 0.002725 | 0.004393 | 0.004183 | 5.9e-05 | 0.001172 | 74.8286 | 1.3846 | 2.5474 | 100.21 | 20.76 | 37.29 | 68.6805 | 87.1911 | 64.7753 | 0.009129 | 1.0447 | 0.024783 | 23.7217 | 0.006177 | 0.13879 | -1e-06 | -3.2e-05 | -3.2e-05 | -178.8619
2019-03-29 13:46 | 1553881548.5241 | -439 | 0.20948 | 0.000455 | -1e-06 | 100.0798 | 99.9244 | 124.8118 | 21.0849 | 0.001913 | 0.001788 | -151.7547 | 0.23231 | 8.4034 | 0.002721 | 0.003833 | 0.003697 | 5.7e-05 | 0.001174 | 74.844 | 1.3502 | 2.5476 | 100.21 | 20.76 | 37.293 | 67.6898 | 87.1146 | 64.6747 | 0.00913 | 1.0451 | 0.024739 | 23.6721 | 0.013907 | 0.1382 | 2e-06 | 0.000223 | 0.000223 | 25.1945
2019-03-29 13:46 | 1553881549.5232 | -438 | 0.20954 | 0.000454 | -1e-06 | 100.0182 | 99.9189 | 124.7778 | 21.0848 | 0.001873 | 0.0018 | -151.7695 | 0.23345 | 8.4032 | 0.002727 | 0.003687 | 0.003546 | 5.7e-05 | 0.001169 | 74.8047 | 1.3183 | 2.5473 | 100.21 | 20.75 | 37.3 | 66.7279 | 87.2526 | 64.9588 | 0.009126 | 1.0442 | 0.024725 | 23.6789 | 0.008883 | 0.13888 | 4e-06 | 0.000338 | 0.000338 | 16.719
2019-03-29 13:46 | 1553881550.5233 | -437 | 0.20948 | 0.000452 | -1e-06 | 100.0358 | 99.8127 | 124.6818 | 21.0854 | 0.001907 | 0.001768 | -151.7716 | 0.2336 | 8.4023 | 0.002684 | 0.003885 | 0.003707 | 5.8e-05 | 0.001164 | 74.797 | 1.3094 | 2.5473 | 100.21 | 20.75 | 37.303 | 67.5398 | 86.218 | 64.1966 | 0.009127 | 1.0465 | 0.024743 | 23.643 | 0.019969 | 0.13897 | 6e-06 | 0.000462 | 0.000462 | 12.2466
2019-03-29 13:46 | 1553881551.5232 | -436 | 0.20951 | 0.000453 | -1e-06 | 100.1205 | 99.6902 | 124.6376 | 21.0859 | 0.001896 | 0.001795 | -151.7716 | 0.23293 | 8.4022 | 0.002714 | 0.004071 | 0.003887 | 5.2e-05 | 0.001157 | 74.773 | 1.3399 | 2.5473 | 100.21 | 20.75 | 37.316 | 67.2828 | 86.9312 | 64.8403 | 0.00913 | 1.0445 | 0.024715 | 23.6614 | 0.038527 | 0.13857 | 4e-06 | 0.000376 | 0.000376 | 14.9825
2019-03-29 13:46 | 1553881552.5233 | -435 | 0.20945 | 0.000456 | -1e-06 | 100.0475 | 99.8223 | 124.882 | 21.0857 | 0.001943 | 0.001788 | -151.777 | 0.23283 | 8.4023 | 0.002715 | 0.003788 | 0.003568 | 6e-05 | 0.001163 | 74.8281 | 1.2904 | 2.5473 | 100.21 | 20.75 | 37.316 | 68.4138 | 86.9544 | 64.6693 | 0.00913 | 1.0451 | 0.024746 | 23.6786 | 0.020157 | 0.13851 | 1e-06 | 0.000125 | 0.000125 | 45.0886
2019-03-29 13:46 | 1553881553.5242 | -434 | 0.20943 | 0.000455 | 0 | 100.1056 | 100.0261 | 125.0097 | 21.0861 | 0.001855 | 0.001803 | -151.7868 | 0.23222 | 8.4033 | 0.002696 | 0.003822 | 0.003676 | 5.7e-05 | 0.001165 | 74.8076 | 1.2997 | 2.5474 | 100.21 | 20.75 | 37.325 | 66.3016 | 86.4935 | 65.0373 | 0.009132 | 1.0439 | 0.024745 | 23.7036 | 0.007109 | 0.13815 | 1e-06 | 0.000141 | 0.000141 | 39.9831
2019-03-29 13:46 | 1553881554.5243 | -433 | 0.20948 | 0.000452 | -1e-06 | 100.0936 | 99.8032 | 125.1075 | 21.0862 | 0.001972 | 0.001773 | -151.797 | 0.23274 | 8.4038 | 0.002674 | 0.004524 | 0.004404 | 5.5e-05 | 0.001168 | 74.8453 | 1.3164 | 2.5473 | 100.21 | 20.75 | 37.325 | 69.1044 | 85.9775 | 64.33 | 0.009132 | 1.0461 | 0.024781 | 23.6881 | 0.02599 | 0.13846 | 6e-06 | 0.000498 | 0.000498 | 11.3263
2019-03-29 13:46 | 1553881555.5238 | -432 | 0.20949 | 0.000452 | -1e-06 | 100.0131 | 99.8496 | 125.1418 | 21.0849 | 0.001951 | 0.001798 | -151.7791 | 0.23353 | 8.4019 | 0.00261 | 0.003798 | 0.003606 | 5.5e-05 | 0.001175 | 74.8752 | 1.3425 | 2.5473 | 100.21 | 20.75 | 37.327 | 68.5885 | 84.4272 | 64.9159 | 0.009133 | 1.0443 | 0.024763 | 23.7119 | 0.014632 | 0.13893 | 7e-06 | 0.000565 | 0.000565 | 10.0122
2019-03-29 13:46 | 1553881556.5236 | -431 | 0.20952 | 0.000454 | -1e-06 | 100.0768 | 100.0568 | 124.8137 | 21.086 | 0.001906 | 0.001821 | -151.7737 | 0.23296 | 8.4024 | 0.00264 | 0.004196 | 0.004021 | 5.3e-05 | 0.001176 | 74.8122 | 1.3299 | 2.5473 | 100.21 | 20.75 | 37.33 | 67.514 | 85.1681 | 65.4743 | 0.009133 | 1.0426 | 0.02471 | 23.7003 | 0.001794 | 0.13859 | 4e-06 | 0.000358 | 0.000358 | 15.7444
2019-03-29 13:46 | 1553881557.5244 | -430 | 0.20945 | 0.000457 | -1e-06 | 100.1379 | 99.8441 | 124.7595 | 21.0864 | 0.001922 | 0.001765 | -151.7646 | 0.23239 | 8.4026 | 0.002645 | 0.004042 | 0.003867 | 5.7e-05 | 0.001182 | 74.872 | 1.342 | 2.5473 | 100.21 | 20.75 | 37.337 | 67.9144 | 85.2763 | 64.1336 | 0.009135 | 1.0467 | 0.024753 | 23.6482 | 0.026278 | 0.13825 | 0 | 5.4e-05 | 5.4e-05 | 103.6795
2019-03-29 13:46 | 1553881558.5245 | -429 | 0.20953 | 0.000453 | -1e-06 | 100.0556 | 99.8745 | 125.085 | 21.0866 | 0.001942 | 0.001805 | -151.7771 | 0.23377 | 8.4032 | 0.002634 | 0.004096 | 0.00392 | 5.3e-05 | 0.001178 | 74.8668 | 1.392 | 2.5474 | 100.21 | 20.75 | 37.34 | 68.3796 | 85.0157 | 65.0831 | 0.009136 | 1.0438 | 0.024751 | 23.7123 | 0.016202 | 0.13907 | 5e-06 | 0.000389 | 0.000389 | 14.5554
2019-03-29 13:46 | 1553881559.5245 | -428 | 0.20954 | 0.000457 | -1e-06 | 100.105 | 99.7834 | 125.0749 | 21.0863 | 0.001891 | 0.001793 | -151.7636 | 0.23273 | 8.401 | 0.002586 | 0.003676 | 0.003469 | 5.6e-05 | 0.001183 | 74.7974 | 1.3685 | 2.5474 | 100.21 | 20.75 | 37.346 | 67.1507 | 83.8654 | 64.8017 | 0.009137 | 1.0447 | 0.02476 | 23.7015 | 0.02878 | 0.13845 | -1e-06 | 0 | 0 | 94122.7495
2019-03-29 13:46 | 1553881560.5245 | -427 | 0.20951 | 0.000458 | -1e-06 | 100.0637 | 100.0677 | 125.1902 | 21.0853 | 0.001982 | 0.001808 | -151.762 | 0.23231 | 8.4023 | 0.002635 | 0.004355 | 0.004155 | 5.6e-05 | 0.001184 | 74.8099 | 1.3689 | 2.5474 | 100.21 | 20.75 | 37.35 | 69.3398 | 85.0409 | 65.1534 | 0.009138 | 1.0436 | 0.024759 | 23.7248 | -0.000359 | 0.1382 | -2e-06 | -0.000122 | -0.000122 | -46.2751
2019-03-29 13:46 | 1553881561.5246 | -426 | 0.20952 | 0.00045 | -1e-06 | 100.1286 | 99.8882 | 125.5603 | 21.087 | 0.001911 | 0.001807 | -151.7779 | 0.23258 | 8.4017 | 0.002641 | 0.00409 | 0.003938 | 5.2e-05 | 0.001182 | 74.7322 | 1.3665 | 2.5471 | 100.21 | 20.75 | 37.353 | 67.6491 | 85.1729 | 65.1387 | 0.009139 | 1.0436 | 0.024796 | 23.7593 | 0.021506 | 0.13836 | 8e-06 | 0.000646 | 0.000646 | 8.7205
2019-03-29 13:46 | 1553881562.5226 | -425 | 0.20948 | 0.000451 | -1e-06 | 100.096 | 99.9974 | 125.7174 | 21.087 | 0.001898 | 0.001787 | -151.7694 | 0.2331 | 8.4035 | 0.002647 | 0.004197 | 0.004015 | 5.3e-05 | 0.001183 | 74.7902 | 1.386 | 2.5475 | 100.21 | 20.75 | 37.36 | 67.3214 | 85.3368 | 64.6682 | 0.009141 | 1.0451 | 0.024829 | 23.7576 | 0.008817 | 0.13867 | 8e-06 | 0.000624 | 0.000624 | 9.0513
2019-03-29 13:46 | 1553881563.5235 | -424 | 0.20949 | 0.000453 | -1e-06 | 100.0969 | 99.9046 | 125.7087 | 21.0879 | 0.001898 | 0.001745 | -151.7572 | 0.23368 | 8.4012 | 0.002647 | 0.00343 | 0.003307 | 5.1e-05 | 0.001186 | 74.8883 | 1.3427 | 2.5475 | 100.21 | 20.75 | 37.363 | 67.3198 | 85.3268 | 63.6453 | 0.009141 | 1.0482 | 0.024865 | 23.7208 | 0.017204 | 0.13901 | 5e-06 | 0.000453 | 0.000453 | 12.499
2019-03-29 13:46 | 1553881564.5231 | -423 | 0.2095 | 0.000454 | -1e-06 | 100.1002 | 99.936 | 125.6211 | 21.0867 | 0.001924 | 0.001732 | -151.7581 | 0.23364 | 8.4018 | 0.002691 | 0.003993 | 0.003737 | 5.1e-05 | 0.001185 | 74.8628 | 1.3133 | 2.5472 | 100.21 | 20.75 | 37.37 | 67.9474 | 86.3832 | 63.3257 | 0.009143 | 1.0492 | 0.024868 | 23.7013 | 0.014688 | 0.13899 | 4e-06 | 0.000405 | 0.000405 | 13.9772
2019-03-29 13:46 | 1553881565.5237 | -422 | 0.20949 | 0.000454 | -1e-06 | 100.0774 | 99.8865 | 126.0507 | 21.0873 | 0.001982 | 0.001838 | -151.7778 | 0.23272 | 8.403 | 0.002668 | 0.003901 | 0.003829 | 5.4e-05 | 0.001188 | 74.8711 | 1.3408 | 2.5474 | 100.21 | 20.75 | 37.374 | 69.3379 | 85.8335 | 65.8922 | 0.009144 | 1.0413 | 0.024816 | 23.8322 | 0.017085 | 0.13845 | 4e-06 | 0.000365 | 0.000365 | 15.4569
2019-03-29 13:46 | 1553881566.5241 | -421 | 0.20946 | 0.000451 | -1e-06 | 100.0583 | 99.8571 | 126.2166 | 21.087 | 0.001971 | 0.001844 | -151.7829 | 0.23314 | 8.4034 | 0.002609 | 0.004285 | 0.004136 | 5.2e-05 | 0.001187 | 74.8383 | 1.3746 | 2.5474 | 100.21 | 20.75 | 37.379 | 69.0927 | 84.4019 | 66.0286 | 0.009145 | 1.0409 | 0.024828 | 23.8526 | 0.018001 | 0.1387 | 8e-06 | 0.00068 | 0.00068 | 8.3021
2019-03-29 13:46 | 1553881567.5228 | -420 | 0.20949 | 0.000448 | -1e-06 | 100.1089 | 99.8684 | 126.5943 | 21.0871 | 0.001893 | 0.001766 | -151.795 | 0.23403 | 8.4032 | 0.002633 | 0.004412 | 0.004202 | 5.3e-05 | 0.001186 | 74.8542 | 1.3383 | 2.5473 | 100.21 | 20.74 | 37.387 | 67.1953 | 84.9835 | 64.1468 | 0.009142 | 1.0467 | 0.024934 | 23.8219 | 0.021514 | 0.13922 | 1.2e-05 | 0.000974 | 0.000974 | 5.816
2019-03-29 13:46 | 1553881568.5228 | -419 | 0.20948 | 0.000446 | -1e-06 | 100.0484 | 99.7906 | 126.8675 | 21.0882 | 0.001882 | 0.001763 | -151.7804 | 0.23405 | 8.403 | 0.002688 | 0.004078 | 0.003887 | 5.4e-05 | 0.001186 | 74.8875 | 1.3381 | 2.5474 | 100.21 | 20.75 | 37.39 | 66.9513 | 86.3088 | 64.0853 | 0.009148 | 1.0469 | 0.024963 | 23.8455 | 0.023081 | 0.13923 | 1.3e-05 | 0.00105 | 0.00105 | 5.3949
2019-03-29 13:46 | 1553881569.5228 | -418 | 0.20953 | 0.000445 | -1e-06 | 100.0709 | 99.9479 | 126.9107 | 21.0881 | 0.001923 | 0.001818 | -151.7851 | 0.23314 | 8.4032 | 0.002774 | 0.004285 | 0.004049 | 5.5e-05 | 0.001185 | 74.8381 | 1.3834 | 2.5475 | 100.21 | 20.74 | 37.393 | 67.9199 | 88.382 | 65.3998 | 0.009143 | 1.0428 | 0.024919 | 23.8959 | 0.010997 | 0.13869 | 1.5e-05 | 0.00122 | 0.00122 | 4.6274
2019-03-29 13:46 | 1553881570.5238 | -417 | 0.20952 | 0.000448 | -1e-06 | 100.1678 | 99.979 | 126.8511 | 21.0879 | 0.001887 | 0.001777 | -151.7958 | 0.23289 | 8.4007 | 0.002723 | 0.003332 | 0.003201 | 5.2e-05 | 0.001176 | 74.8114 | 1.3163 | 2.5474 | 100.21 | 20.74 | 37.406 | 67.0546 | 87.1509 | 64.406 | 0.009146 | 1.0459 | 0.02495 | 23.8552 | 0.016882 | 0.13855 | 1.1e-05 | 0.000882 | 0.000882 | 6.3952
2019-03-29 13:46 | 1553881571.5229 | -416 | 0.2095 | 0.00045 | -1e-06 | 100.1693 | 99.9837 | 126.4215 | 21.0886 | 0.001964 | 0.001756 | -151.7819 | 0.23316 | 8.402 | 0.002685 | 0.003867 | 0.003705 | 5.4e-05 | 0.001177 | 74.8117 | 1.3696 | 2.5471 | 100.21 | 20.74 | 37.406 | 68.9197 | 86.2421 | 63.9093 | 0.009146 | 1.0474 | 0.024926 | 23.7973 | 0.016589 | 0.13871 | 9e-06 | 0.000765 | 0.000765 | 7.3755
2019-03-29 13:46 | 1553881572.5229 | -415 | 0.20951 | 0.000451 | -1e-06 | 100.1457 | 99.9551 | 126.2314 | 21.0882 | 0.001953 | 0.001826 | -151.8025 | 0.23403 | 8.4022 | 0.002648 | 0.004154 | 0.00399 | 5.3e-05 | 0.001184 | 74.8013 | 1.3684 | 2.5474 | 100.21 | 20.74 | 37.414 | 68.647 | 85.3636 | 65.6076 | 0.009148 | 1.0422 | 0.024845 | 23.8392 | 0.017042 | 0.13923 | 8e-06 | 0.000643 | 0.000643 | 8.8134
2019-03-29 13:46 | 1553881573.5229 | -414 | 0.20951 | 0.000449 | -1e-06 | 100.1684 | 99.9203 | 126.2366 | 21.0891 | 0.001964 | 0.001793 | -151.8263 | 0.2331 | 8.4019 | 0.002689 | 0.003679 | 0.003562 | 5.4e-05 | 0.00119 | 74.8336 | 1.4012 | 2.5475 | 100.21 | 20.74 | 37.418 | 68.9068 | 86.3379 | 64.7982 | 0.009149 | 1.0447 | 0.024875 | 23.8112 | 0.022189 | 0.13867 | 1e-05 | 0.000763 | 0.000763 | 7.3959
2019-03-29 13:46 | 1553881574.5239 | -413 | 0.20946 | 0.000449 | -1e-06 | 100.1702 | 99.9619 | 126.0997 | 21.0888 | 0.001859 | 0.0018 | -151.8216 | 0.23379 | 8.4028 | 0.002698 | 0.004716 | 0.00433 | 5.5e-05 | 0.001199 | 74.8095 | 1.3654 | 2.5474 | 100.21 | 20.74 | 37.422 | 66.4001 | 86.5625 | 64.9696 | 0.00915 | 1.0441 | 0.024855 | 23.8043 | 0.018623 | 0.13908 | 9e-06 | 0.000744 | 0.000744 | 7.6132
2019-03-29 13:46 | 1553881575.5233 | -412 | 0.20954 | 0.000448 | -1e-06 | 100.1559 | 100.0107 | 125.7512 | 21.0885 | 0.001923 | 0.001753 | -151.8159 | 0.23328 | 8.4013 | 0.002709 | 0.004094 | 0.003973 | 5.6e-05 | 0.001196 | 74.8207 | 1.378 | 2.5474 | 100.21 | 20.74 | 37.432 | 67.939 | 86.8127 | 63.8389 | 0.009153 | 1.0476 | 0.024862 | 23.7316 | 0.012974 | 0.13878 | 1.1e-05 | 0.000873 | 0.000873 | 6.4732
2019-03-29 13:46 | 1553881576.5231 | -411 | 0.20949 | 0.000451 | -1e-06 | 100.166 | 100.0575 | 125.5444 | 21.0896 | 0.001925 | 0.00181 | -151.8277 | 0.2335 | 8.4024 | 0.002679 | 0.00343 | 0.003285 | 5.2e-05 | 0.001185 | 74.834 | 1.3321 | 2.5474 | 100.21 | 20.74 | 37.432 | 67.9822 | 86.0893 | 65.2077 | 0.009153 | 1.0434 | 0.024792 | 23.7602 | 0.009702 | 0.13891 | 7e-06 | 0.000562 | 0.000562 | 10.0635
2019-03-29 13:46 | 1553881577.5222 | -410 | 0.2095 | 0.000449 | -1e-06 | 100.1804 | 99.9124 | 125.5364 | 21.0889 | 0.001912 | 0.001834 | -151.8298 | 0.23332 | 8.401 | 0.002666 | 0.004291 | 0.00414 | 5.6e-05 | 0.001177 | 74.8264 | 1.3177 | 2.5474 | 100.21 | 20.74 | 37.436 | 67.6666 | 85.785 | 65.7894 | 0.009154 | 1.0416 | 0.02477 | 23.7799 | 0.023961 | 0.1388 | 9e-06 | 0.000714 | 0.000714 | 7.9091
2019-03-29 13:46 | 1553881578.5221 | -409 | 0.2095 | 0.000451 | -1e-06 | 100.1529 | 99.9711 | 125.4655 | 21.0891 | 0.001896 | 0.001756 | -151.8404 | 0.23367 | 8.4041 | 0.002739 | 0.004402 | 0.004193 | 5.2e-05 | 0.001181 | 74.7896 | 1.3374 | 2.5475 | 100.21 | 20.74 | 37.438 | 67.272 | 87.5473 | 63.9112 | 0.009154 | 1.0474 | 0.024831 | 23.7072 | 0.016248 | 0.13901 | 8e-06 | 0.000648 | 0.000648 | 8.7297
2019-03-29 13:46 | 1553881579.5231 | -408 | 0.2095 | 0.000444 | -1e-06 | 100.1613 | 99.8803 | 125.3592 | 21.0888 | 0.001877 | 0.001808 | -151.8447 | 0.23359 | 8.3999 | 0.002692 | 0.003666 | 0.003537 | 5.3e-05 | 0.001182 | 74.755 | 1.3295 | 2.5474 | 100.21 | 20.74 | 37.444 | 66.8246 | 86.4172 | 65.1686 | 0.009155 | 1.0435 | 0.024775 | 23.7413 | 0.025123 | 0.13896 | 1.7e-05 | 0.00128 | 0.00128 | 4.4185
2019-03-29 13:46 | 1553881580.5231 | -407 | 0.2095 | 0.000448 | -1e-06 | 100.1933 | 99.8612 | 125.465 | 21.0892 | 0.001896 | 0.001769 | -151.8498 | 0.23402 | 8.4031 | 0.002673 | 0.004345 | 0.004175 | 5.3e-05 | 0.001185 | 74.8489 | 1.3301 | 2.5475 | 100.21 | 20.74 | 37.444 | 67.2876 | 85.9657 | 64.2236 | 0.009155 | 1.0465 | 0.02482 | 23.7181 | 0.029694 | 0.13922 | 1.1e-05 | 0.000882 | 0.000882 | 6.4215
2019-03-29 13:46 | 1553881581.5222 | -406 | 0.20948 | 0.000445 | -1e-06 | 100.1247 | 100.0953 | 125.8136 | 21.0906 | 0.001941 | 0.001821 | -151.8444 | 0.23372 | 8.4013 | 0.002597 | 0.003436 | 0.00329 | 5.4e-05 | 0.001187 | 74.7719 | 1.3454 | 2.5474 | 100.21 | 20.74 | 37.447 | 68.3748 | 84.1206 | 65.4725 | 0.009156 | 1.0426 | 0.024808 | 23.795 | 0.002627 | 0.13904 | 1.5e-05 | 0.001136 | 0.001136 | 4.9827
2019-03-29 13:46 | 1553881582.5223 | -405 | 0.20951 | 0.000448 | -1e-06 | 100.1637 | 99.9452 | 125.8568 | 21.0899 | 0.001902 | 0.00183 | -151.8279 | 0.23315 | 8.4046 | 0.002662 | 0.003734 | 0.003547 | 5.2e-05 | 0.001189 | 74.7823 | 1.341 | 2.5474 | 100.21 | 20.74 | 37.447 | 67.4351 | 85.6807 | 65.6889 | 0.009156 | 1.0419 | 0.024805 | 23.8067 | 0.019531 | 0.1387 | 1.1e-05 | 0.000879 | 0.000879 | 6.4187
2019-03-29 13:46 | 1553881583.5235 | -404 | 0.20952 | 0.000449 | -1e-06 | 100.2466 | 99.9295 | 125.9009 | 21.0901 | 0.00196 | 0.001802 | -151.8062 | 0.23383 | 8.4015 | 0.002712 | 0.003907 | 0.003714 | 5.6e-05 | 0.001185 | 74.8295 | 1.3431 | 2.5475 | 100.21 | 20.74 | 37.45 | 68.8198 | 86.8831 | 65.028 | 0.009157 | 1.044 | 0.024833 | 23.7876 | 0.028338 | 0.13911 | 1.1e-05 | 0.00085 | 0.00085 | 6.6626
2019-03-29 13:46 | 1553881584.5238 | -403 | 0.20951 | 0.000446 | -1e-06 | 100.2044 | 100.1544 | 126.0502 | 21.0905 | 0.001909 | 0.001784 | -151.7852 | 0.23347 | 8.4028 | 0.002713 | 0.003635 | 0.003432 | 5.6e-05 | 0.001182 | 74.864 | 1.3451 | 2.5474 | 100.21 | 20.74 | 37.452 | 67.6056 | 86.9294 | 64.5863 | 0.009157 | 1.0453 | 0.024864 | 23.7862 | 0.004462 | 0.13889 | 1.4e-05 | 0.001101 | 0.001101 | 5.1322
2019-03-29 13:46 | 1553881585.5224 | -402 | 0.20946 | 0.000454 | -1e-06 | 100.2035 | 100.0199 | 126.2098 | 21.0902 | 0.001898 | 0.001781 | -151.7857 | 0.23383 | 8.4012 | 0.002715 | 0.003942 | 0.003788 | 5.4e-05 | 0.001166 | 74.8169 | 1.3979 | 2.5474 | 100.21 | 20.74 | 37.454 | 67.3213 | 86.9619 | 64.5095 | 0.009158 | 1.0456 | 0.024883 | 23.7985 | 0.016399 | 0.1391 | 4e-06 | 0.000312 | 0.000312 | 18.1381
2019-03-29 13:46 | 1553881586.5223 | -401 | 0.2095 | 0.000449 | -1e-06 | 100.1427 | 100.0216 | 126.4598 | 21.0908 | 0.001913 | 0.00184 | -151.7757 | 0.23351 | 8.4018 | 0.002735 | 0.003782 | 0.003621 | 5.6e-05 | 0.001168 | 74.7945 | 1.3555 | 2.5474 | 100.21 | 20.74 | 37.457 | 67.6965 | 87.4365 | 65.9463 | 0.009159 | 1.0411 | 0.024855 | 23.8727 | 0.010821 | 0.13892 | 1e-05 | 0.000809 | 0.000809 | 6.9912
2019-03-29 13:46 | 1553881587.5238 | -400 | 0.20949 | 0.00045 | -1e-06 | 100.1948 | 100.0697 | 126.5687 | 21.0914 | 0.001941 | 0.00182 | -151.7678 | 0.23289 | 8.4029 | 0.002761 | 0.00431 | 0.004128 | 5.8e-05 | 0.00117 | 74.7988 | 1.323 | 2.5473 | 100.21 | 20.74 | 37.463 | 68.3568 | 88.0672 | 65.4444 | 0.00916 | 1.0427 | 0.024884 | 23.8653 | 0.01117 | 0.13855 | 8e-06 | 0.000682 | 0.000682 | 8.2687
2019-03-29 13:46 | 1553881588.5225 | -399 | 0.2095 | 0.000448 | -1e-06 | 100.1975 | 100.0196 | 126.5154 | 21.092 | 0.001911 | 0.001791 | -151.7533 | 0.23409 | 8.4016 | 0.00276 | 0.004624 | 0.004458 | 5.6e-05 | 0.001176 | 74.8438 | 1.3153 | 2.5473 | 100.21 | 20.74 | 37.466 | 67.6407 | 88.0615 | 64.7634 | 0.009161 | 1.0448 | 0.024904 | 23.8363 | 0.015896 | 0.13926 | 1.1e-05 | 0.000892 | 0.000892 | 6.3532
2019-03-29 13:46 | 1553881589.5225 | -398 | 0.20949 | 0.000453 | -1e-06 | 100.2196 | 99.8459 | 126.6448 | 21.0908 | 0.001924 | 0.001801 | -151.7407 | 0.23416 | 8.4015 | 0.002753 | 0.004207 | 0.004 | 5.6e-05 | 0.001175 | 74.7972 | 1.3471 | 2.5475 | 100.21 | 20.74 | 37.47 | 67.9524 | 87.8927 | 65.0005 | 0.009162 | 1.044 | 0.024908 | 23.8568 | 0.033421 | 0.1393 | 6e-06 | 0.00051 | 0.00051 | 11.1198
2019-03-29 13:46 | 1553881590.5235 | -397 | 0.20949 | 0.000451 | -1e-06 | 100.2499 | 100.0435 | 126.6241 | 21.0914 | 0.001983 | 0.001833 | -151.7378 | 0.23371 | 8.402 | 0.002647 | 0.00431 | 0.004155 | 5.5e-05 | 0.001181 | 74.8245 | 1.313 | 2.5473 | 100.21 | 20.74 | 37.487 | 69.3631 | 85.3333 | 65.7567 | 0.009166 | 1.0417 | 0.024878 | 23.8815 | 0.018434 | 0.13903 | 8e-06 | 0.000679 | 0.000679 | 8.3316
2019-03-29 13:46 | 1553881591.5225 | -396 | 0.20951 | 0.000446 | -1e-06 | 100.2002 | 100.0066 | 126.595 | 21.0922 | 0.001955 | 0.001772 | -151.7482 | 0.23309 | 8.4025 | 0.002683 | 0.004328 | 0.004132 | 5.6e-05 | 0.001187 | 74.8001 | 1.326 | 2.5473 | 100.21 | 20.74 | 37.487 | 68.6953 | 86.2034 | 64.3006 | 0.009166 | 1.0462 | 0.024929 | 23.8274 | 0.017296 | 0.13867 | 1.4e-05 | 0.001123 | 0.001123 | 5.0263
2019-03-29 13:47 | 1553881592.5226 | -395 | 0.2095 | 0.00045 | -1e-06 | 100.2771 | 99.9742 | 126.8466 | 21.092 | 0.001894 | 0.001822 | -151.7465 | 0.23421 | 8.4029 | 0.00273 | 0.003928 | 0.003856 | 5.2e-05 | 0.00119 | 74.8079 | 1.3325 | 2.5475 | 100.21 | 20.74 | 37.504 | 67.2404 | 87.3261 | 65.4965 | 0.00917 | 1.0425 | 0.024909 | 23.8933 | 0.027061 | 0.13933 | 9e-06 | 0.000725 | 0.000725 | 7.8221
2019-03-29 13:47 | 1553881593.5216 | -394 | 0.2095 | 0.000454 | -1e-06 | 100.2549 | 100.0011 | 126.8591 | 21.0921 | 0.001936 | 0.001727 | -151.7546 | 0.23382 | 8.4026 | 0.002642 | 0.004089 | 0.003846 | 5.5e-05 | 0.001191 | 74.8467 | 1.3496 | 2.5474 | 100.21 | 20.74 | 37.504 | 68.2397 | 85.2137 | 63.2262 | 0.00917 | 1.0496 | 0.024994 | 23.8143 | 0.022673 | 0.1391 | 4e-06 | 0.000332 | 0.000332 | 17.0416
2019-03-29 13:47 | 1553881594.5226 | -393 | 0.20949 | 0.000454 | -1e-06 | 100.237 | 99.9629 | 126.9809 | 21.0925 | 0.001886 | 0.001855 | -151.7754 | 0.23378 | 8.3994 | 0.002756 | 0.004065 | 0.003909 | 5.8e-05 | 0.001192 | 74.874 | 1.3708 | 2.5473 | 100.21 | 20.74 | 37.511 | 67.0389 | 87.948 | 66.2884 | 0.009172 | 1.0401 | 0.024893 | 23.9339 | 0.02449 | 0.13908 | 3e-06 | 0.00028 | 0.00028 | 20.1912
2019-03-29 13:47 | 1553881595.5226 | -392 | 0.2095 | 0.000445 | -1e-06 | 100.2166 | 100.0066 | 127.1053 | 21.0922 | 0.001916 | 0.001803 | -151.761 | 0.23346 | 8.4014 | 0.002637 | 0.003664 | 0.003535 | 5.8e-05 | 0.001186 | 74.9174 | 1.3468 | 2.5473 | 100.21 | 20.74 | 37.521 | 67.7719 | 85.0969 | 65.0456 | 0.009174 | 1.0439 | 0.024951 | 23.9017 | 0.018765 | 0.13889 | 1.5e-05 | 0.001209 | 0.001209 | 4.676
2019-03-29 13:47 | 1553881596.5227 | -391 | 0.20949 | 0.000448 | -1e-06 | 100.1837 | 100.0018 | 127.0173 | 21.0923 | 0.0019 | 0.001802 | -151.7579 | 0.23301 | 8.4025 | 0.002637 | 0.004352 | 0.004122 | 5.6e-05 | 0.001186 | 74.8891 | 1.3886 | 2.5474 | 100.21 | 20.74 | 37.526 | 67.3734 | 85.0883 | 65.0284 | 0.009176 | 1.044 | 0.024943 | 23.8929 | 0.016253 | 0.13861 | 1.1e-05 | 0.00087 | 0.00087 | 6.4836
2019-03-29 13:47 | 1553881597.5231 | -390 | 0.2095 | 0.000445 | -1e-06 | 100.1786 | 99.8633 | 126.8093 | 21.0933 | 0.001891 | 0.001775 | -151.7654 | 0.23412 | 8.3995 | 0.002687 | 0.004129 | 0.003943 | 5.6e-05 | 0.001182 | 74.8366 | 1.3101 | 2.5474 | 100.21 | 20.74 | 37.532 | 67.1721 | 86.2928 | 64.365 | 0.009177 | 1.046 | 0.024947 | 23.8499 | 0.028201 | 0.13928 | 1.5e-05 | 0.001229 | 0.001229 | 4.6141
2019-03-29 13:47 | 1553881598.5218 | -389 | 0.20952 | 0.000452 | -1e-06 | 100.2044 | 100.0977 | 126.7398 | 21.0931 | 0.001944 | 0.001784 | -151.786 | 0.23406 | 8.4015 | 0.00271 | 0.004227 | 0.004038 | 5.5e-05 | 0.001182 | 74.8836 | 1.3161 | 2.5472 | 100.21 | 20.74 | 37.541 | 68.4397 | 86.8499 | 64.5982 | 0.009179 | 1.0453 | 0.024932 | 23.8516 | 0.009524 | 0.13924 | 7e-06 | 0.000557 | 0.000557 | 10.1691
2019-03-29 13:47 | 1553881599.5219 | -388 | 0.2095 | 0.000448 | -1e-06 | 100.2002 | 99.9694 | 126.5213 | 21.0937 | 0.001963 | 0.001784 | -151.8061 | 0.2339 | 8.4012 | 0.002711 | 0.003875 | 0.003685 | 5.6e-05 | 0.001178 | 74.886 | 1.3457 | 2.5474 | 100.21 | 20.74 | 37.546 | 68.8942 | 86.8661 | 64.595 | 0.00918 | 1.0453 | 0.02491 | 23.8309 | 0.020624 | 0.13915 | 1.1e-05 | 0.000909 | 0.000909 | 6.2318
2019-03-29 13:47 | 1553881600.5218 | -387 | 0.20949 | 0.000449 | 0 | 100.2217 | 99.9005 | 126.3072 | 21.0931 | 0.001943 | 0.001794 | -151.7995 | 0.23377 | 8.4033 | 0.002629 | 0.004221 | 0.004097 | 5.6e-05 | 0.001179 | 74.9461 | 1.3441 | 2.5473 | 100.21 | 20.73 | 37.556 | 68.414 | 84.8972 | 64.8411 | 0.009177 | 1.0445 | 0.02488 | 23.8194 | 0.028718 | 0.13907 | 9e-06 | 0.000694 | 0.000694 | 8.1535
2019-03-29 13:47 | 1553881601.5232 | -386 | 0.20948 | 0.000449 | -1e-06 | 100.2451 | 99.8994 | 126.0427 | 21.0938 | 0.001886 | 0.00177 | -151.7909 | 0.23374 | 8.4016 | 0.002667 | 0.00389 | 0.003716 | 5.5e-05 | 0.001183 | 74.8688 | 1.3681 | 2.5475 | 100.21 | 20.73 | 37.56 | 67.0394 | 85.804 | 64.2601 | 0.009178 | 1.0463 | 0.024876 | 23.774 | 0.030893 | 0.13905 | 1e-05 | 0.000789 | 0.000789 | 7.1711
2019-03-29 13:47 | 1553881602.522 | -385 | 0.20955 | 0.000448 | -1e-06 | 100.2681 | 99.9868 | 125.6545 | 21.0934 | 0.001891 | 0.001803 | -151.7898 | 0.23348 | 8.402 | 0.002658 | 0.003787 | 0.003639 | 5.6e-05 | 0.001186 | 74.8051 | 1.362 | 2.5475 | 100.21 | 20.73 | 37.563 | 67.1576 | 85.5929 | 65.049 | 0.009179 | 1.0439 | 0.024808 | 23.7651 | 0.02513 | 0.1389 | 1.1e-05 | 0.00088 | 0.00088 | 6.4243
2019-03-29 13:47 | 1553881603.5229 | -384 | 0.20948 | 0.000452 | -1e-06 | 100.254 | 100.1039 | 125.3373 | 21.0946 | 0.001961 | 0.001861 | -151.7731 | 0.23544 | 8.4014 | 0.00261 | 0.003842 | 0.003685 | 5.5e-05 | 0.001192 | 74.8042 | 1.3389 | 2.5473 | 100.21 | 20.73 | 37.569 | 68.845 | 84.4484 | 66.4448 | 0.00918 | 1.0396 | 0.024726 | 23.784 | 0.013409 | 0.14007 | 6e-06 | 0.000475 | 0.000475 | 11.9917
2019-03-29 13:47 | 1553881604.5232 | -383 | 0.20954 | 0.00045 | -1e-06 | 100.2265 | 100.1575 | 125.1587 | 21.0948 | 0.001959 | 0.00181 | -151.7711 | 0.23461 | 8.403 | 0.002627 | 0.003666 | 0.003465 | 5.3e-05 | 0.001194 | 74.8248 | 1.3325 | 2.5474 | 100.21 | 20.73 | 37.574 | 68.7955 | 84.8474 | 65.2065 | 0.009182 | 1.0434 | 0.024754 | 23.7237 | 0.006165 | 0.13957 | 8e-06 | 0.00065 | 0.00065 | 8.7381
2019-03-29 13:47 | 1553881605.522 | -382 | 0.20955 | 0.000448 | 0 | 100.2035 | 100.153 | 124.8294 | 21.0951 | 0.00194 | 0.001794 | -151.787 | 0.23341 | 8.4037 | 0.002622 | 0.003815 | 0.003621 | 5.3e-05 | 0.001184 | 74.801 | 1.3293 | 2.5473 | 100.21 | 20.73 | 37.579 | 68.3333 | 84.7392 | 64.8329 | 0.009183 | 1.0446 | 0.024735 | 23.6793 | 0.004504 | 0.13886 | 1e-05 | 0.000807 | 0.000807 | 7
2019-03-29 13:47 | 1553881606.523 | -381 | 0.20949 | 0.000448 | -1e-06 | 100.2915 | 100.1646 | 124.7472 | 21.0951 | 0.001923 | 0.001763 | -151.7875 | 0.23344 | 8.4019 | 0.002638 | 0.003971 | 0.003845 | 5.7e-05 | 0.001178 | 74.8254 | 1.3122 | 2.5473 | 100.21 | 20.73 | 37.586 | 67.9354 | 85.12 | 64.0795 | 0.009185 | 1.0469 | 0.024754 | 23.6452 | 0.01132 | 0.13887 | 1.1e-05 | 0.000847 | 0.000847 | 6.676
2019-03-29 13:47 | 1553881607.5221 | -380 | 0.2095 | 0.000454 | -1e-06 | 100.3223 | 100.2086 | 124.9141 | 21.0957 | 0.001944 | 0.001816 | -151.7968 | 0.23433 | 8.403 | 0.002528 | 0.00418 | 0.004 | 5.5e-05 | 0.001174 | 74.7866 | 1.3193 | 2.5474 | 100.21 | 20.73 | 37.59 | 68.4384 | 82.468 | 65.3565 | 0.009186 | 1.043 | 0.024724 | 23.7057 | 0.010139 | 0.1394 | 4e-06 | 0.000323 | 0.000323 | 17.5574
2019-03-29 13:47 | 1553881608.5211 | -379 | 0.20949 | 0.000449 | -1e-06 | 100.3058 | 100.1339 | 124.8644 | 21.0951 | 0.001912 | 0.001766 | -151.7998 | 0.23385 | 8.4038 | 0.002608 | 0.00455 | 0.004326 | 5.2e-05 | 0.001173 | 74.8453 | 1.3608 | 2.5475 | 100.21 | 20.73 | 37.597 | 67.6757 | 84.3926 | 64.1528 | 0.009187 | 1.0467 | 0.024763 | 23.6588 | 0.015345 | 0.13912 | 9e-06 | 0.000732 | 0.000732 | 7.7403
2019-03-29 13:47 | 1553881609.522 | -378 | 0.20951 | 0.000451 | -1e-06 | 100.3816 | 100.3544 | 124.9942 | 21.0951 | 0.001959 | 0.001804 | -151.787 | 0.23384 | 8.4033 | 0.002573 | 0.003657 | 0.003445 | 5.2e-05 | 0.00117 | 74.8627 | 1.3588 | 2.5475 | 100.21 | 20.73 | 37.602 | 68.8114 | 83.5438 | 65.0826 | 0.009188 | 1.0438 | 0.024742 | 23.7037 | 0.002423 | 0.13911 | 7e-06 | 0.000601 | 0.000601 | 9.4284
2019-03-29 13:47 | 1553881610.5211 | -377 | 0.20949 | 0.000448 | -1e-06 | 100.3627 | 100.0786 | 125.1425 | 21.0958 | 0.001979 | 0.001765 | -151.7968 | 0.23447 | 8.3995 | 0.002555 | 0.003931 | 0.00385 | 5.1e-05 | 0.001175 | 74.8121 | 1.3711 | 2.5474 | 100.21 | 20.73 | 37.606 | 69.2808 | 83.1131 | 64.1365 | 0.009189 | 1.0467 | 0.024791 | 23.6846 | 0.025355 | 0.13949 | 1.1e-05 | 0.000881 | 0.000881 | 6.4408
2019-03-29 13:47 | 1553881611.5211 | -376 | 0.20953 | 0.000449 | -1e-06 | 100.3223 | 100.1551 | 125.3289 | 21.095 | 0.001936 | 0.001767 | -151.7995 | 0.23398 | 8.402 | 0.002642 | 0.004028 | 0.003844 | 5.4e-05 | 0.001172 | 74.8085 | 1.3818 | 2.5475 | 100.21 | 20.73 | 37.614 | 68.2592 | 85.2151 | 64.1848 | 0.009191 | 1.0466 | 0.024808 | 23.7039 | 0.014923 | 0.1392 | 1e-05 | 0.000823 | 0.000823 | 6.8821
2019-03-29 13:47 | 1553881612.5214 | -375 | 0.2095 | 0.00045 | -1e-06 | 100.3406 | 100.2127 | 125.1885 | 21.0959 | 0.001888 | 0.001801 | -151.7943 | 0.23382 | 8.4023 | 0.002537 | 0.00351 | 0.00337 | 5e-05 | 0.00117 | 74.8074 | 1.3882 | 2.5474 | 100.21 | 20.73 | 37.619 | 67.0828 | 82.6865 | 64.9999 | 0.009193 | 1.044 | 0.024764 | 23.7192 | 0.011401 | 0.1391 | 8e-06 | 0.000597 | 0.000597 | 9.4878
2019-03-29 13:47 | 1553881613.5212 | -374 | 0.20954 | 0.000446 | -1e-06 | 100.4067 | 100.1516 | 125.1831 | 21.0966 | 0.001966 | 0.001838 | -151.8173 | 0.23418 | 8.4 | 0.002576 | 0.004114 | 0.00395 | 5.1e-05 | 0.001178 | 74.8155 | 1.358 | 2.5474 | 100.21 | 20.73 | 37.627 | 68.9616 | 83.6241 | 65.9036 | 0.009195 | 1.0413 | 0.024731 | 23.7504 | 0.02275 | 0.13931 | 1.3e-05 | 0.001041 | 0.001041 | 5.4487
2019-03-29 13:47 | 1553881614.5223 | -373 | 0.20951 | 0.000449 | -1e-06 | 100.4483 | 100.1288 | 125.0061 | 21.096 | 0.001907 | 0.001792 | -151.8097 | 0.23393 | 8.4052 | 0.002617 | 0.004049 | 0.003859 | 5.4e-05 | 0.001167 | 74.7782 | 1.3193 | 2.5473 | 100.21 | 20.73 | 37.632 | 67.5446 | 84.6206 | 64.7799 | 0.009196 | 1.0447 | 0.024754 | 23.6942 | 0.028498 | 0.13917 | 9e-06 | 0.000742 | 0.000742 | 7.6319
2019-03-29 13:47 | 1553881615.5223 | -372 | 0.20952 | 0.000449 | -1e-06 | 100.4103 | 100.2199 | 125.2057 | 21.0968 | 0.001903 | 0.001836 | -151.7993 | 0.23445 | 8.3993 | 0.002613 | 0.003364 | 0.00318 | 5.4e-05 | 0.001151 | 74.7536 | 1.3676 | 2.5475 | 100.21 | 20.73 | 37.637 | 67.4556 | 84.5114 | 65.8523 | 0.009197 | 1.0414 | 0.024735 | 23.7507 | 0.016976 | 0.13947 | 1e-05 | 0.000787 | 0.000787 | 7.2107
2019-03-29 13:47 | 1553881616.5227 | -371 | 0.20954 | 0.000452 | -1e-06 | 100.4636 | 100.1622 | 125.5418 | 21.0973 | 0.001924 | 0.00179 | -151.8191 | 0.23396 | 8.4027 | 0.002594 | 0.003405 | 0.003209 | 5.2e-05 | 0.001149 | 74.787 | 1.3497 | 2.5474 | 100.21 | 20.73 | 37.646 | 67.968 | 84.053 | 64.7303 | 0.009199 | 1.0449 | 0.024809 | 23.7432 | 0.02687 | 0.13918 | 6e-06 | 0.00047 | 0.00047 | 12.0643
2019-03-29 13:47 | 1553881617.5224 | -370 | 0.20954 | 0.000449 | -1e-06 | 100.4079 | 100.2305 | 125.4211 | 21.0972 | 0.001838 | 0.001809 | -151.8043 | 0.23342 | 8.4029 | 0.0026 | 0.003322 | 0.003075 | 5.4e-05 | 0.00115 | 74.82 | 1.3223 | 2.5473 | 100.21 | 20.73 | 37.652 | 65.8987 | 84.2073 | 65.1869 | 0.009201 | 1.0435 | 0.02478 | 23.7478 | 0.015819 | 0.13886 | 1e-05 | 0.000775 | 0.000775 | 7.2962
2019-03-29 13:47 | 1553881618.5224 | -369 | 0.20954 | 0.000452 | -1e-06 | 100.4555 | 100.2697 | 125.3589 | 21.0971 | 0.001962 | 0.001729 | -151.7985 | 0.23365 | 8.405 | 0.002653 | 0.00352 | 0.003332 | 5.2e-05 | 0.00115 | 74.8017 | 1.3389 | 2.5474 | 100.21 | 20.73 | 37.662 | 68.8826 | 85.4825 | 63.2756 | 0.009203 | 1.0494 | 0.024844 | 23.6748 | 0.016556 | 0.139 | 6e-06 | 0.000517 | 0.000517 | 10.951
2019-03-29 13:47 | 1553881619.5217 | -368 | 0.20952 | 0.000447 | -1e-06 | 100.4875 | 100.2339 | 125.5258 | 21.0957 | 0.001913 | 0.001759 | -151.8149 | 0.23429 | 8.4031 | 0.002606 | 0.004284 | 0.004098 | 5.2e-05 | 0.001146 | 74.8001 | 1.3402 | 2.5473 | 100.21 | 20.73 | 37.667 | 67.684 | 84.3379 | 63.9927 | 0.009204 | 1.0472 | 0.024834 | 23.7158 | 0.022604 | 0.13938 | 1.2e-05 | 0.000951 | 0.000951 | 5.9668
2019-03-29 13:47 | 1553881620.5215 | -367 | 0.20949 | 0.000449 | -1e-06 | 100.4797 | 100.4547 | 125.861 | 21.0979 | 0.001952 | 0.001803 | -151.8011 | 0.23395 | 8.4038 | 0.002676 | 0.004025 | 0.003878 | 5e-05 | 0.001149 | 74.7687 | 1.3308 | 2.5474 | 100.21 | 20.73 | 37.673 | 68.6331 | 86.0271 | 65.0627 | 0.009206 | 1.0439 | 0.024828 | 23.7851 | 0.002229 | 0.13918 | 1e-05 | 0.000784 | 0.000784 | 7.2248
2019-03-29 13:47 | 1553881621.5225 | -366 | 0.20948 | 0.00045 | -1e-06 | 100.5234 | 100.467 | 126.0047 | 21.0976 | 0.001933 | 0.001782 | -151.7866 | 0.23308 | 8.4026 | 0.002629 | 0.003414 | 0.00321 | 5e-05 | 0.001152 | 74.8126 | 1.3403 | 2.5474 | 100.21 | 20.73 | 37.683 | 68.1727 | 84.911 | 64.5456 | 0.009208 | 1.0455 | 0.024861 | 23.7804 | 0.005024 | 0.13866 | 8e-06 | 0.000685 | 0.000685 | 8.2349
2019-03-29 13:47 | 1553881622.5225 | -365 | 0.20952 | 0.00045 | -1e-06 | 100.4519 | 100.3158 | 126.1985 | 21.0975 | 0.001917 | 0.00177 | -151.8025 | 0.23313 | 8.4016 | 0.00256 | 0.003284 | 0.003123 | 5.1e-05 | 0.001155 | 74.8083 | 1.3383 | 2.5474 | 100.21 | 20.73 | 37.683 | 67.7848 | 83.2466 | 64.2474 | 0.009208 | 1.0464 | 0.024891 | 23.7882 | 0.012127 | 0.13869 | 9e-06 | 0.000685 | 0.000685 | 8.2375
2019-03-29 13:47 | 1553881623.521 | -364 | 0.2095 | 0.000447 | -1e-06 | 100.4702 | 100.2626 | 126.2025 | 21.0972 | 0.001892 | 0.001757 | -151.7878 | 0.23405 | 8.4002 | 0.002669 | 0.003647 | 0.003483 | 5.5e-05 | 0.001155 | 74.7446 | 1.3546 | 2.5474 | 100.21 | 20.73 | 37.694 | 67.202 | 85.8752 | 63.9379 | 0.009211 | 1.0473 | 0.024903 | 23.7777 | 0.018501 | 0.13924 | 1.2e-05 | 0.00092 | 0.00092 | 6.1592
2019-03-29 13:47 | 1553881624.5206 | -363 | 0.20953 | 0.000454 | -1e-06 | 100.5594 | 100.4356 | 126.2037 | 21.0974 | 0.001918 | 0.001795 | -151.795 | 0.23386 | 8.4023 | 0.00258 | 0.00459 | 0.004362 | 5e-05 | 0.001156 | 74.7713 | 1.3618 | 2.5475 | 100.21 | 20.73 | 37.7 | 67.8128 | 83.7316 | 64.8615 | 0.009212 | 1.0445 | 0.024869 | 23.8103 | 0.011016 | 0.13913 | 3e-06 | 0.00027 | 0.00027 | 20.9364
2019-03-29 13:47 | 1553881625.5216 | -362 | 0.20953 | 0.000449 | -1e-06 | 100.5456 | 100.2581 | 126.0856 | 21.0983 | 0.001918 | 0.001743 | -151.8066 | 0.23433 | 8.4014 | 0.002608 | 0.003769 | 0.003603 | 5.3e-05 | 0.001155 | 74.83 | 1.3208 | 2.5474 | 100.21 | 20.73 | 37.709 | 67.8168 | 84.4099 | 63.6 | 0.009215 | 1.0484 | 0.024904 | 23.7548 | 0.02561 | 0.1394 | 1e-05 | 0.000759 | 0.000759 | 7.4765
2019-03-29 13:47 | 1553881626.5206 | -361 | 0.20953 | 0.00045 | -1e-06 | 100.4776 | 100.3097 | 125.6293 | 21.098 | 0.001903 | 0.001765 | -151.796 | 0.23385 | 8.4033 | 0.002678 | 0.003731 | 0.003507 | 5.1e-05 | 0.001154 | 74.8422 | 1.3697 | 2.5475 | 100.21 | 20.73 | 37.713 | 67.4624 | 86.0867 | 64.1456 | 0.009216 | 1.0467 | 0.024839 | 23.7309 | 0.014966 | 0.13912 | 8e-06 | 0.000618 | 0.000618 | 9.1634
2019-03-29 13:47 | 1553881627.5217 | -360 | 0.20952 | 0.000451 | -1e-06 | 100.5588 | 100.4752 | 125.5962 | 21.0992 | 0.001955 | 0.001816 | -151.8099 | 0.23281 | 8.4018 | 0.002581 | 0.00406 | 0.00388 | 5.3e-05 | 0.001153 | 74.821 | 1.3518 | 2.5472 | 100.21 | 20.73 | 37.728 | 68.7132 | 83.7608 | 65.3776 | 0.009219 | 1.0429 | 0.024791 | 23.7711 | 0.007437 | 0.1385 | 8e-06 | 0.000622 | 0.000622 | 9.0606
2019-03-29 13:47 | 1553881628.5221 | -359 | 0.20953 | 0.000449 | -1e-06 | 100.5288 | 100.2411 | 125.8242 | 21.0982 | 0.001922 | 0.001791 | -151.8198 | 0.23348 | 8.4013 | 0.002578 | 0.003883 | 0.003731 | 5.3e-05 | 0.001161 | 74.8919 | 1.3814 | 2.5475 | 100.21 | 20.73 | 37.728 | 67.9235 | 83.6733 | 64.7649 | 0.009219 | 1.0448 | 0.024835 | 23.7711 | 0.025641 | 0.1389 | 1e-05 | 0.000765 | 0.000765 | 7.3924
2019-03-29 13:47 | 1553881629.5218 | -358 | 0.20955 | 0.000451 | -1e-06 | 100.5549 | 100.4748 | 125.8861 | 21.0989 | 0.001933 | 0.001776 | -151.7967 | 0.2342 | 8.402 | 0.002655 | 0.00483 | 0.004635 | 5.3e-05 | 0.001168 | 74.8584 | 1.3626 | 2.5474 | 100.21 | 20.73 | 37.731 | 68.1754 | 85.5395 | 64.3917 | 0.00922 | 1.0459 | 0.024855 | 23.7638 | 0.007122 | 0.13933 | 7e-06 | 0.000571 | 0.000571 | 9.9242
2019-03-29 13:47 | 1553881630.5218 | -357 | 0.20955 | 0.000452 | 0 | 100.5677 | 100.2936 | 125.4735 | 21.0981 | 0.001971 | 0.001813 | -151.816 | 0.23414 | 8.4015 | 0.002691 | 0.003534 | 0.003372 | 5.4e-05 | 0.001172 | 74.8738 | 1.3406 | 2.5474 | 100.21 | 20.73 | 37.746 | 69.0879 | 86.3904 | 65.2878 | 0.009224 | 1.0432 | 0.024782 | 23.7563 | 0.024414 | 0.13929 | 5e-06 | 0.000408 | 0.000408 | 13.8877
2019-03-29 13:47 | 1553881631.5207 | -356 | 0.20953 | 0.000454 | 0 | 100.5695 | 100.4243 | 125.0489 | 21.1001 | 0.001969 | 0.001835 | -151.8275 | 0.23358 | 8.4011 | 0.002652 | 0.003753 | 0.003541 | 5.4e-05 | 0.001175 | 74.8453 | 1.355 | 2.5475 | 100.21 | 20.73 | 37.746 | 69.0369 | 85.469 | 65.8197 | 0.009224 | 1.0415 | 0.02472 | 23.7347 | 0.012924 | 0.13896 | 2e-06 | 0.000153 | 0.000153 | 36.8566
2019-03-29 13:47 | 1553881632.5218 | -355 | 0.20953 | 0.000447 | -1e-06 | 100.6058 | 100.2158 | 124.9723 | 21.0997 | 0.001894 | 0.001794 | -151.8268 | 0.23382 | 8.4009 | 0.002661 | 0.003764 | 0.003587 | 5.6e-05 | 0.001177 | 74.8221 | 1.3383 | 2.5476 | 100.21 | 20.73 | 37.75 | 67.2339 | 85.6652 | 64.8491 | 0.009225 | 1.0445 | 0.024748 | 23.6935 | 0.034737 | 0.1391 | 1.1e-05 | 0.000877 | 0.000877 | 6.458
2019-03-29 13:47 | 1553881633.5218 | -354 | 0.20954 | 0.000449 | -1e-06 | 100.5588 | 100.3663 | 125.0099 | 21.0996 | 0.001869 | 0.001798 | -151.834 | 0.23425 | 8.4005 | 0.00266 | 0.003704 | 0.003556 | 5.6e-05 | 0.001178 | 74.853 | 1.3007 | 2.5472 | 100.21 | 20.72 | 37.761 | 66.6464 | 85.6567 | 64.9388 | 0.009222 | 1.0442 | 0.024749 | 23.7002 | 0.017135 | 0.13936 | 9e-06 | 0.000741 | 0.000741 | 7.6587
2019-03-29 13:47 | 1553881634.5218 | -353 | 0.20954 | 0.000449 | 0 | 100.5396 | 100.4639 | 125.0827 | 21.0991 | 0.001931 | 0.001793 | -151.8437 | 0.2334 | 8.4019 | 0.002712 | 0.003944 | 0.003668 | 5.4e-05 | 0.001174 | 74.8343 | 1.342 | 2.5473 | 100.21 | 20.72 | 37.77 | 68.1217 | 86.9065 | 64.8211 | 0.009224 | 1.0446 | 0.02476 | 23.7029 | 0.006736 | 0.13885 | 9e-06 | 0.000676 | 0.000676 | 8.3643
2019-03-29 13:47 | 1553881635.5223 | -352 | 0.20955 | 0.000451 | -1e-06 | 100.6285 | 100.4817 | 125.3359 | 21.1002 | 0.001926 | 0.001799 | -151.8242 | 0.23363 | 8.4033 | 0.002631 | 0.004656 | 0.004532 | 5.2e-05 | 0.001169 | 74.8637 | 1.3177 | 2.5474 | 100.21 | 20.72 | 37.77 | 68.0043 | 84.9633 | 64.948 | 0.009224 | 1.0442 | 0.024781 | 23.7314 | 0.013061 | 0.13899 | 7e-06 | 0.000547 | 0.000547 | 10.3401
2019-03-29 13:47 | 1553881636.522 | -351 | 0.20948 | 0.00045 | -1e-06 | 100.648 | 100.4803 | 125.4688 | 21.1 | 0.001905 | 0.001788 | -151.8026 | 0.23378 | 8.401 | 0.002633 | 0.003696 | 0.003477 | 5.5e-05 | 0.001159 | 74.8157 | 1.3166 | 2.5474 | 100.21 | 20.72 | 37.781 | 67.4977 | 84.9981 | 64.6978 | 0.009226 | 1.045 | 0.024803 | 23.7352 | 0.014912 | 0.13908 | 7e-06 | 0.000585 | 0.000585 | 9.6805
2019-03-29 13:47 | 1553881637.521 | -350 | 0.2095 | 0.000449 | -1e-06 | 100.6447 | 100.5547 | 125.6812 | 21.1001 | 0.001894 | 0.001818 | -151.7904 | 0.23303 | 8.4024 | 0.002608 | 0.004275 | 0.00418 | 5.4e-05 | 0.001148 | 74.7473 | 1.324 | 2.5472 | 100.21 | 20.72 | 37.781 | 67.2327 | 84.4132 | 65.4231 | 0.009226 | 1.0427 | 0.024797 | 23.7807 | 0.007999 | 0.13863 | 9e-06 | 0.000705 | 0.000705 | 8.008
2019-03-29 13:47 | 1553881638.522 | -349 | 0.2095 | 0.00045 | -1e-06 | 100.7069 | 100.6049 | 125.7488 | 21.1006 | 0.001881 | 0.001758 | -151.7744 | 0.23438 | 8.4028 | 0.00266 | 0.003543 | 0.003404 | 5.4e-05 | 0.001146 | 74.7136 | 1.3542 | 2.5475 | 100.21 | 20.72 | 37.79 | 66.9359 | 85.6629 | 63.963 | 0.009229 | 1.0473 | 0.024858 | 23.7358 | 0.009069 | 0.13943 | 8e-06 | 0.000611 | 0.000611 | 9.2949
2019-03-29 13:47 | 1553881639.5201 | -348 | 0.20961 | 0.00045 | -1e-06 | 100.7315 | 100.6199 | 125.5286 | 21.1009 | 0.001868 | 0.001779 | -151.7831 | 0.23429 | 8.4024 | 0.002587 | 0.004079 | 0.003899 | 5.6e-05 | 0.001148 | 74.7265 | 1.335 | 2.5474 | 100.21 | 20.72 | 37.793 | 66.6303 | 83.9099 | 64.4661 | 0.009229 | 1.0457 | 0.024817 | 23.7327 | 0.009914 | 0.13938 | 8e-06 | 0.00062 | 0.00062 | 9.1434
2019-03-29 13:47 | 1553881640.5202 | -347 | 0.20957 | 0.000453 | -1e-06 | 100.7069 | 100.5114 | 125.4223 | 21.1013 | 0.001921 | 0.001783 | -151.7639 | 0.23422 | 8.4011 | 0.002618 | 0.003447 | 0.003244 | 5.3e-05 | 0.001158 | 74.7837 | 1.3276 | 2.5475 | 100.21 | 20.72 | 37.797 | 67.8861 | 84.6336 | 64.5836 | 0.00923 | 1.0453 | 0.024802 | 23.7268 | 0.017387 | 0.13933 | 4e-06 | 0.000333 | 0.000333 | 17.0461
2019-03-29 13:47 | 1553881641.5211 | -346 | 0.20947 | 0.000454 | -1e-06 | 100.6782 | 100.4667 | 125.7735 | 21.1015 | 0.001872 | 0.001831 | -151.7646 | 0.23312 | 8.4018 | 0.00261 | 0.004427 | 0.004251 | 5.4e-05 | 0.001162 | 74.7994 | 1.343 | 2.5474 | 100.21 | 20.72 | 37.811 | 66.7247 | 84.459 | 65.7203 | 0.009234 | 1.0418 | 0.024795 | 23.7999 | 0.018814 | 0.13869 | 3e-06 | 0.000289 | 0.000289 | 19.5237
2019-03-29 13:47 | 1553881642.5211 | -345 | 0.20951 | 0.000454 | -1e-06 | 100.7557 | 100.4527 | 126.1176 | 21.1013 | 0.001942 | 0.001786 | -151.7608 | 0.23276 | 8.4034 | 0.002572 | 0.003457 | 0.003304 | 5.1e-05 | 0.001166 | 74.7751 | 1.3392 | 2.5474 | 100.21 | 20.72 | 37.811 | 68.389 | 83.5442 | 64.6461 | 0.009234 | 1.0451 | 0.024869 | 23.7946 | 0.026945 | 0.13847 | 3e-06 | 0.000262 | 0.000262 | 21.5037
2019-03-29 13:47 | 1553881643.5213 | -344 | 0.20957 | 0.000457 | -1e-06 | 100.7171 | 100.6127 | 126.4727 | 21.102 | 0.001883 | 0.001723 | -151.7532 | 0.23338 | 8.4029 | 0.002585 | 0.003462 | 0.003287 | 5.1e-05 | 0.001167 | 74.8166 | 1.3098 | 2.5474 | 100.21 | 20.72 | 37.822 | 66.9755 | 83.8626 | 63.1391 | 0.009237 | 1.0498 | 0.02496 | 23.7749 | 0.009275 | 0.13884 | -1e-06 | -2e-06 | -2e-06 | -2645.3381
2019-03-29 13:47 | 1553881644.5202 | -343 | 0.20946 | 0.000458 | -1e-06 | 100.7147 | 100.4588 | 126.6403 | 21.102 | 0.001898 | 0.001849 | -151.7361 | 0.2337 | 8.4016 | 0.002609 | 0.003897 | 0.003748 | 5.7e-05 | 0.001162 | 74.846 | 1.3001 | 2.5474 | 100.21 | 20.72 | 37.822 | 67.3328 | 84.4196 | 66.1523 | 0.009237 | 1.0405 | 0.024865 | 23.897 | 0.022757 | 0.13903 | -2e-06 | -0.000103 | -0.000103 | -55.1146
2019-03-29 13:47 | 1553881645.5208 | -342 | 0.20951 | 0.000459 | -1e-06 | 100.6842 | 100.4609 | 126.5279 | 21.1015 | 0.001946 | 0.001797 | -151.7534 | 0.23333 | 8.401 | 0.00255 | 0.003784 | 0.003583 | 5.3e-05 | 0.001161 | 74.8098 | 1.3672 | 2.5475 | 100.21 | 20.72 | 37.834 | 68.5039 | 83.0157 | 64.9225 | 0.009239 | 1.0443 | 0.024899 | 23.8431 | 0.019862 | 0.13881 | -3e-06 | -0.000148 | -0.000148 | -38.1397
2019-03-29 13:47 | 1553881646.5203 | -341 | 0.20949 | 0.00046 | -1e-06 | 100.7189 | 100.6581 | 126.4769 | 21.1018 | 0.001864 | 0.001817 | -151.7767 | 0.23367 | 8.4028 | 0.002541 | 0.004429 | 0.004263 | 5.4e-05 | 0.001159 | 74.8609 | 1.3096 | 2.5475 | 100.21 | 20.72 | 37.834 | 66.5182 | 82.7992 | 65.3976 | 0.009239 | 1.0428 | 0.024877 | 23.855 | 0.0054 | 0.13901 | -4e-06 | -0.00027 | -0.00027 | -20.9589
2019-03-29 13:47 | 1553881647.5203 | -340 | 0.2095 | 0.000457 | -1e-06 | 100.6701 | 100.4697 | 126.0737 | 21.1026 | 0.001896 | 0.001737 | -151.7664 | 0.23424 | 8.4016 | 0.00259 | 0.003609 | 0.003427 | 5.6e-05 | 0.001159 | 74.8478 | 1.3513 | 2.5475 | 100.21 | 20.72 | 37.838 | 67.2865 | 83.9692 | 63.466 | 0.00924 | 1.0488 | 0.024908 | 23.7489 | 0.017823 | 0.13935 | -1e-06 | -3e-06 | -3e-06 | -1620.7026
2019-03-29 13:47 | 1553881648.5204 | -339 | 0.20951 | 0.000457 | -1e-06 | 100.6444 | 100.5039 | 125.7223 | 21.1025 | 0.001934 | 0.001729 | -151.776 | 0.23373 | 8.4022 | 0.002633 | 0.004589 | 0.004428 | 5.5e-05 | 0.001159 | 74.8386 | 1.335 | 2.5475 | 100.21 | 20.72 | 37.842 | 68.2046 | 85.0015 | 63.2871 | 0.009241 | 1.0494 | 0.02488 | 23.7095 | 0.012497 | 0.13905 | -1e-06 | 1e-06 | 1e-06 | 5387.7819
2019-03-29 13:47 | 1553881649.5214 | -338 | 0.2095 | 0.000458 | 0 | 100.7494 | 100.596 | 125.7462 | 21.1027 | 0.001882 | 0.00177 | -151.7843 | 0.23386 | 8.4008 | 0.002673 | 0.003427 | 0.003301 | 5.8e-05 | 0.001161 | 74.8473 | 1.3338 | 2.5472 | 100.21 | 20.72 | 37.846 | 66.9677 | 85.9678 | 64.251 | 0.009242 | 1.0464 | 0.024847 | 23.7457 | 0.013633 | 0.13913 | -3e-06 | -0.000167 | -0.000167 | -33.8462
2019-03-29 13:47 | 1553881650.5213 | -337 | 0.20951 | 0.000455 | -1e-06 | 100.7045 | 100.4561 | 125.7859 | 21.1032 | 0.001933 | 0.001729 | -151.7915 | 0.23377 | 8.4005 | 0.002609 | 0.004176 | 0.003934 | 5.5e-05 | 0.00116 | 74.813 | 1.3333 | 2.5473 | 100.21 | 20.72 | 37.849 | 68.1718 | 84.4293 | 63.2767 | 0.009243 | 1.0494 | 0.024887 | 23.7151 | 0.022096 | 0.13907 | 2e-06 | 0.000208 | 0.000208 | 27.1514
2019-03-29 13:48 | 1553881651.5214 | -336 | 0.20952 | 0.000456 | 0 | 100.7716 | 100.3438 | 125.6127 | 21.1029 | 0.001919 | 0.001777 | -151.7918 | 0.23437 | 8.402 | 0.002625 | 0.003865 | 0.00361 | 5.6e-05 | 0.00116 | 74.8144 | 1.3562 | 2.5474 | 100.21 | 20.72 | 37.854 | 67.8477 | 84.8045 | 64.4328 | 0.009244 | 1.0458 | 0.024827 | 23.7395 | 0.038053 | 0.13943 | 0 | 5e-06 | 5e-06 | 1054.1862
2019-03-29 13:48 | 1553881652.5215 | -335 | 0.20951 | 0.000456 | -1e-06 | 100.7318 | 100.8254 | 125.4357 | 21.104 | 0.001947 | 0.001808 | -151.7999 | 0.23379 | 8.4025 | 0.002568 | 0.004193 | 0.004031 | 5.6e-05 | 0.001157 | 74.7684 | 1.3345 | 2.5472 | 100.21 | 20.72 | 37.856 | 68.5125 | 83.4557 | 65.185 | 0.009245 | 1.0435 | 0.024782 | 23.7491 | -0.008305 | 0.13908 | 1e-06 | 9.1e-05 | 9.1e-05 | 62.0024
2019-03-29 13:48 | 1553881653.5219 | -334 | 0.2095 | 0.000458 | -1e-06 | 100.7402 | 100.6957 | 125.0967 | 21.1043 | 0.001884 | 0.001766 | -151.7977 | 0.23314 | 8.4026 | 0.002593 | 0.004651 | 0.004508 | 5.5e-05 | 0.00115 | 74.8171 | 1.3599 | 2.5473 | 100.21 | 20.71 | 37.857 | 67.0152 | 84.0504 | 64.1767 | 0.009239 | 1.0466 | 0.024785 | 23.6817 | 0.003952 | 0.1387 | -3e-06 | -0.000155 | -0.000155 | -36.3738
2019-03-29 13:48 | 1553881654.5216 | -333 | 0.20953 | 0.000456 | -1e-06 | 100.7973 | 100.596 | 124.8775 | 21.1036 | 0.00191 | 0.001823 | -151.8007 | 0.2345 | 8.4019 | 0.002573 | 0.003642 | 0.003484 | 5.7e-05 | 0.001138 | 74.8013 | 1.3445 | 2.5473 | 100.21 | 20.71 | 37.859 | 67.62 | 83.5559 | 65.5412 | 0.00924 | 1.0424 | 0.024714 | 23.7087 | 0.017884 | 0.1395 | 1e-06 | 8.3e-05 | 8.3e-05 | 68.7077
2019-03-29 13:48 | 1553881655.52 | -332 | 0.20952 | 0.000461 | -1e-06 | 100.7432 | 100.5677 | 124.8662 | 21.1041 | 0.001943 | 0.001723 | -151.7878 | 0.23401 | 8.4028 | 0.002622 | 0.003962 | 0.003848 | 5.6e-05 | 0.001134 | 74.8415 | 1.378 | 2.5473 | 100.21 | 20.71 | 37.86 | 68.4324 | 84.747 | 63.1433 | 0.00924 | 1.0498 | 0.0248 | 23.6236 | 0.015593 | 0.13921 | -7e-06 | -0.000461 | -0.000461 | -12.2837
2019-03-29 13:48 | 1553881656.5197 | -331 | 0.20954 | 0.000458 | -1e-06 | 100.7701 | 100.5472 | 124.7825 | 21.1043 | 0.00191 | 0.001825 | -151.7749 | 0.23413 | 8.4003 | 0.00261 | 0.003732 | 0.003547 | 5.7e-05 | 0.001131 | 74.7554 | 1.3774 | 2.5472 | 100.21 | 20.71 | 37.859 | 67.6432 | 84.4655 | 65.5981 | 0.00924 | 1.0422 | 0.024702 | 23.7017 | 0.019806 | 0.13928 | -2e-06 | -0.000102 | -0.000102 | -55.501
2019-03-29 13:48 | 1553881657.5197 | -330 | 0.20955 | 0.000459 | -1e-06 | 100.7414 | 100.667 | 124.9289 | 21.105 | 0.001941 | 0.001813 | -151.7821 | 0.23401 | 8.4022 | 0.002699 | 0.003887 | 0.003704 | 5.4e-05 | 0.001136 | 74.8162 | 1.3321 | 2.5473 | 100.21 | 20.71 | 37.859 | 68.3662 | 86.5898 | 65.3078 | 0.00924 | 1.0431 | 0.024727 | 23.7054 | 0.006605 | 0.13921 | -3e-06 | -0.000163 | -0.000163 | -34.7041
2019-03-29 13:48 | 1553881658.5206 | -329 | 0.20954 | 0.000459 | -1e-06 | 100.7602 | 100.6861 | 125.2336 | 21.1052 | 0.001904 | 0.001806 | -151.7727 | 0.23442 | 8.4041 | 0.002733 | 0.003929 | 0.003709 | 5.7e-05 | 0.00114 | 74.8166 | 1.3219 | 2.5475 | 100.21 | 20.71 | 37.858 | 67.478 | 87.4116 | 65.1244 | 0.00924 | 1.0437 | 0.024764 | 23.7279 | 0.006581 | 0.13946 | -3e-06 | -0.000183 | -0.000183 | -31.024
2019-03-29 13:48 | 1553881659.5889 | -328 | 0.20953 | 0.000454 | -1e-06 | 100.7979 | 100.6175 | 125.3047 | 21.1049 | 0.001972 | 0.001761 | -151.754 | 0.23447 | 8.4045 | 0.002701 | 0.003677 | 0.003564 | 5.6e-05 | 0.001142 | 74.8327 | 1.3668 | 2.5474 | 100.21 | 20.71 | 37.857 | 69.1273 | 86.6522 | 64.0577 | 0.009239 | 1.047 | 0.02481 | 23.6972 | 0.016025 | 0.13949 | 3e-06 | 0.000243 | 0.000243 | 23.3779
2019-03-29 13:48 | 1553881660.536 | -327 | 0.20952 | 0.000455 | 0 | 100.7952 | 100.5482 | 125.5127 | 21.1051 | 0.001929 | 0.001814 | -151.7451 | 0.23427 | 8.403 | 0.002677 | 0.004447 | 0.004175 | 5.4e-05 | 0.001143 | 74.8083 | 1.3021 | 2.5473 | 100.21 | 20.71 | 37.857 | 68.0799 | 86.0577 | 65.335 | 0.009239 | 1.043 | 0.024784 | 23.7617 | 0.021947 | 0.13937 | 2e-06 | 0.000168 | 0.000168 | 33.7118
2019-03-29 13:48 | 1553881661.535 | -326 | 0.20954 | 0.000458 | -1e-06 | 100.7563 | 100.5953 | 125.4575 | 21.1051 | 0.00193 | 0.001828 | -151.744 | 0.23392 | 8.4029 | 0.00272 | 0.004274 | 0.004134 | 5.8e-05 | 0.001139 | 74.8092 | 1.3178 | 2.5473 | 100.21 | 20.71 | 37.855 | 68.1213 | 87.1053 | 65.659 | 0.009239 | 1.042 | 0.024767 | 23.7678 | 0.014304 | 0.13916 | -2e-06 | -0.000112 | -0.000112 | -50.6588
2019-03-29 13:48 | 1553881662.5354 | -325 | 0.20952 | 0.00046 | -1e-06 | 100.7572 | 100.5329 | 125.2388 | 21.1052 | 0.001871 | 0.001784 | -151.7308 | 0.23461 | 8.4035 | 0.002678 | 0.004525 | 0.004315 | 5.7e-05 | 0.001133 | 74.8212 | 1.3174 | 2.5474 | 100.21 | 20.71 | 37.852 | 66.6939 | 86.1013 | 64.6038 | 0.009238 | 1.0453 | 0.024784 | 23.7101 | 0.019939 | 0.13957 | -4e-06 | -0.000255 | -0.000255 | -22.3052
2019-03-29 13:48 | 1553881663.536 | -324 | 0.20955 | 0.000456 | 0 | 100.7842 | 100.6841 | 125.3071 | 21.1056 | 0.001911 | 0.001827 | -151.7307 | 0.23431 | 8.4017 | 0.00275 | 0.003514 | 0.003367 | 5.5e-05 | 0.001134 | 74.8007 | 1.3429 | 2.5474 | 100.21 | 20.71 | 37.852 | 67.6484 | 87.8281 | 65.6478 | 0.009238 | 1.0421 | 0.024752 | 23.7532 | 0.008889 | 0.13939 | 0 | 4.5e-05 | 4.5e-05 | 125.5859
2019-03-29 13:48 | 1553881664.5361 | -323 | 0.20953 | 0.000451 | 0 | 100.7755 | 100.7383 | 125.2728 | 21.1046 | 0.001944 | 0.001801 | -151.7322 | 0.23398 | 8.4007 | 0.002761 | 0.004317 | 0.004127 | 5.4e-05 | 0.001141 | 74.8173 | 1.3349 | 2.5473 | 100.21 | 20.71 | 37.849 | 68.4541 | 88.0926 | 65.0226 | 0.009237 | 1.044 | 0.024772 | 23.728 | 0.003299 | 0.1392 | 6e-06 | 0.000482 | 0.000482 | 11.7631
2019-03-29 13:48 | 1553881665.5357 | -322 | 0.20955 | 0.000452 | -1e-06 | 100.8105 | 100.6172 | 125.3315 | 21.1057 | 0.001947 | 0.00182 | -151.7433 | 0.2336 | 8.4017 | 0.002727 | 0.004741 | 0.004534 | 5.7e-05 | 0.001139 | 74.7749 | 1.3555 | 2.5474 | 100.21 | 20.71 | 37.846 | 68.5173 | 87.2715 | 65.4597 | 0.009237 | 1.0426 | 0.024761 | 23.7489 | 0.017171 | 0.13897 | 6e-06 | 0.000495 | 0.000495 | 11.4245
2019-03-29 13:48 | 1553881666.5362 | -321 | 0.20954 | 0.000456 | 0 | 100.7952 | 100.8537 | 125.4887 | 21.1065 | 0.001931 | 0.001779 | -151.7548 | 0.23377 | 8.4043 | 0.002738 | 0.004153 | 0.003993 | 5.4e-05 | 0.001133 | 74.7484 | 1.3359 | 2.5473 | 100.21 | 20.71 | 37.845 | 68.1337 | 87.5336 | 64.4777 | 0.009236 | 1.0457 | 0.024813 | 23.7293 | -0.005184 | 0.13907 | 0 | 6e-05 | 6e-05 | 94.9795
2019-03-29 13:48 | 1553881667.5362 | -320 | 0.20951 | 0.000456 | -1e-06 | 100.7072 | 100.625 | 125.4944 | 21.1062 | 0.001919 | 0.001743 | -151.7502 | 0.23442 | 8.4023 | 0.002707 | 0.004133 | 0.003951 | 5.8e-05 | 0.001134 | 74.8047 | 1.3477 | 2.5475 | 100.21 | 20.71 | 37.845 | 67.8471 | 86.7893 | 63.6265 | 0.009236 | 1.0483 | 0.024845 | 23.6999 | 0.007306 | 0.13946 | 1e-06 | 0.000116 | 0.000116 | 49.136
2019-03-29 13:48 | 1553881668.5366 | -319 | 0.20957 | 0.000459 | -1e-06 | 100.7979 | 100.6858 | 125.3582 | 21.1053 | 0.001925 | 0.001818 | -151.7734 | 0.23385 | 8.4021 | 0.002734 | 0.00412 | 0.003937 | 5.4e-05 | 0.001132 | 74.8539 | 1.353 | 2.5475 | 100.21 | 20.71 | 37.844 | 68.0054 | 87.427 | 65.4312 | 0.009236 | 1.0427 | 0.024765 | 23.7504 | 0.009959 | 0.13912 | -3e-06 | -0.000175 | -0.000175 | -32.2875
2019-03-29 13:48 | 1553881669.5353 | -318 | 0.20951 | 0.000455 | -1e-06 | 100.8096 | 100.6158 | 125.4157 | 21.1065 | 0.001862 | 0.001784 | -151.7714 | 0.23431 | 8.3989 | 0.002671 | 0.003189 | 0.002995 | 5.6e-05 | 0.001132 | 74.8281 | 1.3387 | 2.5473 | 100.21 | 20.71 | 37.842 | 66.4844 | 85.9301 | 64.6028 | 0.009236 | 1.0453 | 0.024801 | 23.7268 | 0.017213 | 0.13939 | 2e-06 | 0.000191 | 0.000191 | 29.7322
2019-03-29 13:48 | 1553881670.5356 | -317 | 0.20953 | 0.000455 | -1e-06 | 100.8267 | 100.682 | 125.2315 | 21.1058 | 0.001888 | 0.001778 | -151.7723 | 0.23414 | 8.4016 | 0.002662 | 0.004557 | 0.004349 | 5.2e-05 | 0.001141 | 74.8352 | 1.3315 | 2.5474 | 100.21 | 20.71 | 37.842 | 67.0963 | 85.7117 | 64.4648 | 0.009236 | 1.0457 | 0.024788 | 23.7045 | 0.012842 | 0.13929 | 2e-06 | 0.0002 | 0.0002 | 28.4004
2019-03-29 13:48 | 1553881671.5343 | -316 | 0.20945 | 0.000455 | -1e-06 | 100.8081 | 100.666 | 125.3704 | 21.1063 | 0.001937 | 0.001803 | -151.773 | 0.23474 | 8.4002 | 0.002729 | 0.004081 | 0.003928 | 5.3e-05 | 0.001145 | 74.8677 | 1.313 | 2.5473 | 100.21 | 20.71 | 37.841 | 68.2864 | 87.3235 | 65.0593 | 0.009235 | 1.0439 | 0.02478 | 23.7386 | 0.012621 | 0.13965 | 2e-06 | 0.000221 | 0.000221 | 25.6776
2019-03-29 13:48 | 1553881672.5353 | -315 | 0.20949 | 0.000454 | -1e-06 | 100.8105 | 100.6236 | 125.4054 | 21.1069 | 0.001953 | 0.001787 | -151.7816 | 0.23358 | 8.4003 | 0.002742 | 0.003519 | 0.003318 | 5.3e-05 | 0.001141 | 74.8643 | 1.3225 | 2.5474 | 100.21 | 20.71 | 37.841 | 68.6731 | 87.6394 | 64.6826 | 0.009235 | 1.045 | 0.024797 | 23.7286 | 0.016594 | 0.13896 | 3e-06 | 0.000262 | 0.000262 | 21.6103
2019-03-29 13:48 | 1553881673.5357 | -314 | 0.20949 | 0.000454 | -1e-06 | 100.73 | 100.7021 | 125.2951 | 21.1068 | 0.00188 | 0.00181 | -151.7902 | 0.23448 | 8.4019 | 0.002734 | 0.004156 | 0.003954 | 5.6e-05 | 0.001135 | 74.8856 | 1.29 | 2.5474 | 100.21 | 20.71 | 37.84 | 66.9169 | 87.4316 | 65.2368 | 0.009235 | 1.0433 | 0.024766 | 23.7376 | 0.002472 | 0.13949 | 4e-06 | 0.000321 | 0.000321 | 17.7111
2019-03-29 13:48 | 1553881674.5355 | -313 | 0.20948 | 0.000448 | -1e-06 | 100.7677 | 100.6236 | 125.2885 | 21.1064 | 0.001932 | 0.001771 | -151.79 | 0.23397 | 8.4034 | 0.00283 | 0.004248 | 0.004093 | 5.2e-05 | 0.001131 | 74.8591 | 1.308 | 2.5473 | 100.21 | 20.71 | 37.84 | 68.1709 | 89.7395 | 64.2981 | 0.009235 | 1.0462 | 0.0248 | 23.7041 | 0.012796 | 0.13919 | 1.1e-05 | 0.000887 | 0.000887 | 6.3847
2019-03-29 13:48 | 1553881675.5347 | -312 | 0.20949 | 0.000448 | 0 | 100.8195 | 100.6407 | 125.0367 | 21.1064 | 0.001951 | 0.001757 | -151.7801 | 0.23447 | 8.401 | 0.002793 | 0.004316 | 0.004128 | 5.4e-05 | 0.00112 | 74.8548 | 1.3676 | 2.5474 | 100.21 | 20.71 | 37.84 | 68.6077 | 88.8683 | 63.9444 | 0.009235 | 1.0473 | 0.024788 | 23.6678 | 0.015875 | 0.13949 | 1e-05 | 0.000808 | 0.000808 | 7.0227
2019-03-29 13:48 | 1553881676.5361 | -311 | 0.2095 | 0.000452 | -1e-06 | 100.7806 | 100.6455 | 125.1017 | 21.1058 | 0.001896 | 0.001805 | -151.7866 | 0.23332 | 8.4012 | 0.002843 | 0.00409 | 0.003919 | 5.7e-05 | 0.001119 | 74.853 | 1.3407 | 2.5474 | 100.21 | 20.71 | 37.84 | 67.3084 | 90.068 | 65.101 | 0.009235 | 1.0437 | 0.024752 | 23.7146 | 0.011997 | 0.13881 | 5e-06 | 0.000412 | 0.000412 | 13.713
2019-03-29 13:48 | 1553881677.5345 | -310 | 0.20951 | 0.000451 | -1e-06 | 100.7973 | 100.6114 | 125.1758 | 21.1074 | 0.001929 | 0.00176 | -151.7839 | 0.23361 | 8.4031 | 0.002786 | 0.003907 | 0.00369 | 5.3e-05 | 0.00112 | 74.8946 | 1.36 | 2.5475 | 100.21 | 20.71 | 37.839 | 68.0884 | 88.6935 | 64.0228 | 0.009235 | 1.0471 | 0.024799 | 23.6837 | 0.016518 | 0.13897 | 7e-06 | 0.000625 | 0.000625 | 9.0432
2019-03-29 13:48 | 1553881678.5345 | -309 | 0.20948 | 0.000452 | 0 | 100.7285 | 100.6015 | 125.196 | 21.1071 | 0.00198 | 0.001794 | -151.7655 | 0.23425 | 8.4042 | 0.002798 | 0.004138 | 0.003976 | 5.6e-05 | 0.001125 | 74.8465 | 1.3672 | 2.5474 | 100.21 | 20.71 | 37.837 | 69.3064 | 88.9807 | 64.8493 | 0.009234 | 1.0445 | 0.02477 | 23.7147 | 0.011287 | 0.13935 | 6e-06 | 0.000446 | 0.000446 | 12.7265
2019-03-29 13:48 | 1553881679.5346 | -308 | 0.20953 | 0.000451 | 0 | 100.7518 | 100.7421 | 125.1221 | 21.1067 | 0.001973 | 0.001777 | -151.7566 | 0.2339 | 8.4005 | 0.002756 | 0.004198 | 0.004003 | 5.6e-05 | 0.001132 | 74.8411 | 1.3367 | 2.5474 | 100.21 | 20.71 | 37.835 | 69.1541 | 87.9603 | 64.4413 | 0.009234 | 1.0458 | 0.024778 | 23.6934 | 0.000866 | 0.13915 | 6e-06 | 0.000461 | 0.000461 | 12.2924
2019-03-29 13:48 | 1553881680.5346 | -307 | 0.2095 | 0.000448 | -1e-06 | 100.6659 | 100.5298 | 125.0761 | 21.1071 | 0.001947 | 0.001754 | -151.7646 | 0.23382 | 8.4017 | 0.002697 | 0.004099 | 0.0039 | 5.6e-05 | 0.001138 | 74.8372 | 1.3325 | 2.5474 | 100.21 | 20.71 | 37.835 | 68.5141 | 86.5591 | 63.8717 | 0.009234 | 1.0475 | 0.024794 | 23.669 | 0.012104 | 0.1391 | 1e-05 | 0.000839 | 0.000839 | 6.744
2019-03-29 13:48 | 1553881681.535 | -306 | 0.20951 | 0.000449 | 0 | 100.6632 | 100.553 | 124.812 | 21.1076 | 0.001954 | 0.00179 | -151.7642 | 0.23395 | 8.4036 | 0.002761 | 0.003663 | 0.00349 | 5.4e-05 | 0.001136 | 74.8372 | 1.4205 | 2.5474 | 100.21 | 20.71 | 37.833 | 68.7048 | 88.0885 | 64.7413 | 0.009233 | 1.0448 | 0.024736 | 23.6745 | 0.0098 | 0.13918 | 9e-06 | 0.000715 | 0.000715 | 7.9177
2019-03-29 13:48 | 1553881682.5349 | -305 | 0.2095 | 0.000457 | -1e-06 | 100.6764 | 100.3834 | 125.2444 | 21.1077 | 0.00191 | 0.001804 | -151.7608 | 0.23351 | 8.4011 | 0.002753 | 0.003311 | 0.003133 | 5.6e-05 | 0.001141 | 74.8186 | 1.3943 | 2.5475 | 100.21 | 20.72 | 37.831 | 67.6434 | 87.9003 | 65.0792 | 0.009239 | 1.0438 | 0.024767 | 23.7273 | 0.026071 | 0.13891 | -1e-06 | -3.7e-05 | -3.7e-05 | -153.2367
2019-03-29 13:48 | 1553881683.5348 | -304 | 0.2095 | 0.000452 | 0 | 100.6408 | 100.6766 | 125.5592 | 21.1069 | 0.001975 | 0.001875 | -151.7444 | 0.23358 | 8.4014 | 0.002816 | 0.004512 | 0.004228 | 5e-05 | 0.001144 | 74.8804 | 1.395 | 2.5475 | 100.21 | 20.71 | 37.83 | 69.1877 | 89.4212 | 66.7915 | 0.009233 | 1.0385 | 0.024735 | 23.8172 | -0.003178 | 0.13896 | 5e-06 | 0.000353 | 0.000353 | 16.0258
2019-03-29 13:48 | 1553881684.5348 | -303 | 0.20952 | 0.000452 | -1e-06 | 100.5698 | 100.5905 | 125.5657 | 21.1073 | 0.001933 | 0.001772 | -151.7397 | 0.23374 | 8.4039 | 0.002735 | 0.004074 | 0.003958 | 5.7e-05 | 0.001148 | 74.7808 | 1.3541 | 2.5474 | 100.21 | 20.71 | 37.829 | 68.1992 | 87.4713 | 64.317 | 0.009233 | 1.0462 | 0.024826 | 23.7309 | -0.001841 | 0.13905 | 5e-06 | 0.000437 | 0.000437 | 12.9478
2019-03-29 13:48 | 1553881685.5343 | -302 | 0.20953 | 0.000447 | -1e-06 | 100.6273 | 100.6499 | 125.2845 | 21.1083 | 0.001856 | 0.001802 | -151.732 | 0.23412 | 8.4008 | 0.002757 | 0.004111 | 0.003865 | 5.7e-05 | 0.001152 | 74.6816 | 1.3425 | 2.5474 | 100.21 | 20.71 | 37.826 | 66.3403 | 87.9932 | 65.0315 | 0.009232 | 1.044 | 0.024772 | 23.7294 | -0.00201 | 0.13928 | 1.1e-05 | 0.000857 | 0.000857 | 6.6129
2019-03-29 13:48 | 1553881686.5338 | -301 | 0.20947 | 0.000451 | -1e-06 | 100.5758 | 100.4254 | 125.0897 | 21.1075 | 0.001895 | 0.001795 | -151.7157 | 0.23366 | 8.4024 | 0.002835 | 0.004056 | 0.003916 | 5.3e-05 | 0.001145 | 74.8203 | 1.383 | 2.5474 | 100.21 | 20.71 | 37.826 | 67.2759 | 89.8695 | 64.8599 | 0.009232 | 1.0445 | 0.024759 | 23.705 | 0.013392 | 0.139 | 7e-06 | 0.000534 | 0.000534 | 10.5915
2019-03-29 13:48 | 1553881687.5338 | -300 | 0.20951 | 0.000452 | -1e-06 | 100.5743 | 100.4779 | 124.9946 | 21.1079 | 0.001926 | 0.001789 | -151.7362 | 0.23418 | 8.4007 | 0.002816 | 0.003961 | 0.003786 | 5.6e-05 | 0.001142 | 74.848 | 1.3279 | 2.5473 | 100.21 | 20.71 | 37.825 | 68.029 | 89.4033 | 64.718 | 0.009232 | 1.0449 | 0.024755 | 23.691 | 0.008579 | 0.13932 | 5e-06 | 0.000433 | 0.000433 | 13.0973
2019-03-29 13:48 | 1553881688.534 | -299 | 0.20947 | 0.00045 | -1e-06 | 100.5743 | 100.394 | 124.9193 | 21.108 | 0.001892 | 0.00179 | -151.7272 | 0.23379 | 8.4016 | 0.002752 | 0.004471 | 0.004304 | 5.8e-05 | 0.001142 | 74.8882 | 1.3224 | 2.5473 | 100.21 | 20.71 | 37.825 | 67.2037 | 87.8794 | 64.7448 | 0.009232 | 1.0448 | 0.024747 | 23.6848 | 0.016055 | 0.13908 | 8e-06 | 0.000649 | 0.000649 | 8.7202
2019-03-29 13:48 | 1553881689.5339 | -298 | 0.2095 | 0.000451 | -1e-06 | 100.5564 | 100.4441 | 124.8001 | 21.1082 | 0.001901 | 0.001828 | -151.742 | 0.23366 | 8.4015 | 0.002742 | 0.004361 | 0.0042 | 5.6e-05 | 0.001146 | 74.8751 | 1.3249 | 2.5474 | 100.21 | 20.71 | 37.824 | 67.4271 | 87.6354 | 65.6737 | 0.009231 | 1.042 | 0.024701 | 23.706 | 0.00999 | 0.13901 | 7e-06 | 0.000575 | 0.000575 | 9.8466
2019-03-29 13:48 | 1553881690.5343 | -297 | 0.2095 | 0.000453 | -1e-06 | 100.5809 | 100.3503 | 124.9524 | 21.1086 | 0.001882 | 0.001751 | -151.7326 | 0.23384 | 8.4018 | 0.002807 | 0.004338 | 0.004156 | 5.6e-05 | 0.00115 | 74.8248 | 1.2958 | 2.5473 | 100.21 | 20.71 | 37.823 | 66.9562 | 89.2042 | 63.8233 | 0.009231 | 1.0477 | 0.024784 | 23.6556 | 0.020534 | 0.13911 | 4e-06 | 0.000353 | 0.000353 | 16.0266
2019-03-29 13:48 | 1553881691.535 | -296 | 0.20948 | 0.00045 | -1e-06 | 100.5734 | 100.3786 | 124.8986 | 21.1092 | 0.001883 | 0.001785 | -151.7392 | 0.2336 | 8.4016 | 0.002741 | 0.003493 | 0.003271 | 5.7e-05 | 0.001156 | 74.8218 | 1.3004 | 2.5474 | 100.21 | 20.71 | 37.823 | 66.9911 | 87.6074 | 64.6264 | 0.009231 | 1.0452 | 0.024749 | 23.6787 | 0.017344 | 0.13897 | 8e-06 | 0.000637 | 0.000637 | 8.8762
2019-03-29 13:48 | 1553881692.5339 | -295 | 0.20949 | 0.000451 | -1e-06 | 100.5453 | 100.2909 | 124.4503 | 21.1086 | 0.001944 | 0.001778 | -151.7324 | 0.2345 | 8.403 | 0.002739 | 0.003735 | 0.003604 | 5.4e-05 | 0.001161 | 74.7906 | 1.3298 | 2.5474 | 100.21 | 20.71 | 37.824 | 68.46 | 87.5505 | 64.4571 | 0.009231 | 1.0457 | 0.024711 | 23.6302 | 0.022661 | 0.1395 | 6e-06 | 0.000512 | 0.000512 | 11.0802
2019-03-29 13:48 | 1553881693.5344 | -294 | 0.20949 | 0.000447 | -1e-06 | 100.5243 | 100.3353 | 124.2135 | 21.1103 | 0.001866 | 0.001761 | -151.7411 | 0.23332 | 8.4029 | 0.002769 | 0.004187 | 0.003986 | 5.6e-05 | 0.001163 | 74.8406 | 1.3672 | 2.5473 | 100.21 | 20.71 | 37.823 | 66.5786 | 88.2958 | 64.0536 | 0.009231 | 1.047 | 0.024702 | 23.5936 | 0.01684 | 0.13881 | 1.2e-05 | 0.000965 | 0.000965 | 5.8522
2019-03-29 13:48 | 1553881694.5351 | -293 | 0.20948 | 0.000445 | -1e-06 | 100.4726 | 100.3561 | 124.3301 | 21.1097 | 0.001868 | 0.001805 | -151.7612 | 0.2337 | 8.4 | 0.002762 | 0.003673 | 0.003494 | 5.9e-05 | 0.001164 | 74.7713 | 1.3352 | 2.5476 | 100.21 | 20.71 | 37.822 | 66.6203 | 88.1225 | 65.1234 | 0.009231 | 1.0437 | 0.024675 | 23.6421 | 0.010376 | 0.13903 | 1.3e-05 | 0.001035 | 0.001035 | 5.4696
2019-03-29 13:48 | 1553881695.5341 | -292 | 0.2095 | 0.000445 | -1e-06 | 100.5498 | 100.2854 | 124.6494 | 21.1091 | 0.001928 | 0.001773 | -151.7626 | 0.23351 | 8.4018 | 0.002693 | 0.004541 | 0.004356 | 5.7e-05 | 0.001168 | 74.7696 | 1.3611 | 2.5475 | 100.21 | 20.71 | 37.822 | 68.08 | 86.4595 | 64.339 | 0.009231 | 1.0461 | 0.024735 | 23.645 | 0.023547 | 0.13891 | 1.4e-05 | 0.001125 | 0.001125 | 5.0258
2019-03-29 13:48 | 1553881696.5341 | -291 | 0.20952 | 0.000448 | -1e-06 | 100.5309 | 100.3349 | 124.9961 | 21.1107 | 0.001911 | 0.001743 | -151.7594 | 0.23395 | 8.4023 | 0.002698 | 0.004718 | 0.004523 | 5.7e-05 | 0.001166 | 74.8566 | 1.3371 | 2.5473 | 100.21 | 20.71 | 37.822 | 67.6558 | 86.586 | 63.62 | 0.009231 | 1.0483 | 0.024796 | 23.6526 | 0.017457 | 0.13918 | 1e-05 | 0.000828 | 0.000828 | 6.8428
2019-03-29 13:48 | 1553881697.5345 | -290 | 0.20951 | 0.000445 | 0 | 100.4588 | 100.2762 | 124.8822 | 21.1097 | 0.001913 | 0.00179 | -151.7498 | 0.2342 | 8.4013 | 0.002626 | 0.003831 | 0.00367 | 5.7e-05 | 0.001166 | 74.9081 | 1.3443 | 2.5472 | 100.21 | 20.71 | 37.819 | 67.7024 | 84.8532 | 64.741 | 0.00923 | 1.0449 | 0.024743 | 23.6811 | 0.016271 | 0.13933 | 1.4e-05 | 0.001096 | 0.001096 | 5.1734
2019-03-29 13:48 | 1553881698.5352 | -289 | 0.20953 | 0.000447 | -1e-06 | 100.5357 | 100.2663 | 124.9763 | 21.1099 | 0.001974 | 0.00174 | -151.7513 | 0.23434 | 8.4007 | 0.002744 | 0.003445 | 0.00323 | 5.6e-05 | 0.001161 | 74.8474 | 1.379 | 2.5475 | 100.21 | 20.71 | 37.817 | 69.1673 | 87.6767 | 63.5403 | 0.00923 | 1.0486 | 0.024797 | 23.6479 | 0.024 | 0.13941 | 1.3e-05 | 0.001025 | 0.001025 | 5.5349
2019-03-29 13:48 | 1553881699.5342 | -288 | 0.20953 | 0.000447 | -1e-06 | 100.4654 | 100.5417 | 125.2202 | 21.1102 | 0.00191 | 0.001788 | -151.7678 | 0.23377 | 8.4024 | 0.002706 | 0.003987 | 0.003802 | 5.5e-05 | 0.001161 | 74.917 | 1.4004 | 2.5474 | 100.21 | 20.71 | 37.817 | 67.6358 | 86.7683 | 64.7046 | 0.00923 | 1.045 | 0.024778 | 23.7119 | -0.006796 | 0.13907 | 1.3e-05 | 0.001035 | 0.001035 | 5.4669
2019-03-29 13:48 | 1553881700.5343 | -287 | 0.20952 | 0.000445 | -1e-06 | 100.4022 | 100.1854 | 125.3089 | 21.1113 | 0.001932 | 0.001792 | -151.775 | 0.23321 | 8.4013 | 0.002678 | 0.004129 | 0.003892 | 5.6e-05 | 0.00116 | 75.0005 | 1.3537 | 2.5474 | 100.21 | 20.72 | 37.814 | 68.1766 | 86.0965 | 64.8121 | 0.009235 | 1.0446 | 0.024783 | 23.7241 | 0.019333 | 0.13874 | 1.5e-05 | 0.001163 | 0.001163 | 4.8561
2019-03-29 13:48 | 1553881701.5336 | -286 | 0.20954 | 0.000446 | -1e-06 | 100.4501 | 100.3121 | 125.6108 | 21.1104 | 0.001911 | 0.001826 | -151.7749 | 0.23422 | 8.401 | 0.002642 | 0.003424 | 0.003298 | 5.8e-05 | 0.001156 | 75.0076 | 1.3821 | 2.5474 | 100.21 | 20.72 | 37.814 | 67.6545 | 85.2416 | 65.6135 | 0.009235 | 1.0422 | 0.024783 | 23.7808 | 0.012302 | 0.13933 | 1.3e-05 | 0.00102 | 0.00102 | 5.5581
2019-03-29 13:48 | 1553881702.5347 | -285 | 0.20952 | 0.000446 | -1e-06 | 100.4627 | 100.1646 | 125.9507 | 21.1118 | 0.001946 | 0.001816 | -151.7675 | 0.2339 | 8.403 | 0.002681 | 0.003707 | 0.003515 | 5.5e-05 | 0.001146 | 75.0353 | 1.3384 | 2.5473 | 100.21 | 20.72 | 37.812 | 68.5124 | 86.1758 | 65.3822 | 0.009234 | 1.0429 | 0.024825 | 23.8048 | 0.026577 | 0.13915 | 1.3e-05 | 0.001002 | 0.001002 | 5.6495
2019-03-29 13:48 | 1553881703.5333 | -284 | 0.20952 | 0.000444 | -1e-06 | 100.3911 | 100.1957 | 126.0389 | 21.1115 | 0.001922 | 0.001786 | -151.7769 | 0.23383 | 8.4013 | 0.002684 | 0.004136 | 0.003954 | 5.6e-05 | 0.001138 | 75.0827 | 1.338 | 2.5474 | 100.21 | 20.72 | 37.811 | 67.9294 | 86.2395 | 64.6669 | 0.009234 | 1.0451 | 0.02486 | 23.7879 | 0.017433 | 0.13911 | 1.6e-05 | 0.001259 | 0.001259 | 4.4957
2019-03-29 13:48 | 1553881704.5333 | -283 | 0.20948 | 0.000448 | -1e-06 | 100.4058 | 100.1926 | 126.1244 | 21.112 | 0.001931 | 0.001783 | -151.7548 | 0.23401 | 8.405 | 0.002699 | 0.002884 | 0.002716 | 5.6e-05 | 0.001138 | 75.0787 | 1.3193 | 2.5473 | 100.21 | 20.72 | 37.81 | 68.1329 | 86.606 | 64.5909 | 0.009234 | 1.0453 | 0.024872 | 23.7933 | 0.019013 | 0.13922 | 1e-05 | 0.000786 | 0.000786 | 7.2077
2019-03-29 13:48 | 1553881705.5349 | -282 | 0.20951 | 0.000449 | -1e-06 | 100.354 | 100.1929 | 126.3129 | 21.112 | 0.001924 | 0.001791 | -151.7704 | 0.23297 | 8.4022 | 0.002602 | 0.003149 | 0.002907 | 5.8e-05 | 0.001143 | 75.1444 | 1.3791 | 2.5475 | 100.21 | 20.72 | 37.809 | 67.9768 | 84.2578 | 64.7713 | 0.009233 | 1.0448 | 0.024883 | 23.8175 | 0.014369 | 0.13859 | 1e-05 | 0.000763 | 0.000763 | 7.3882
2019-03-29 13:48 | 1553881706.5335 | -281 | 0.2095 | 0.000446 | 0 | 100.3618 | 100.2752 | 125.9518 | 21.1112 | 0.001891 | 0.001744 | -151.7807 | 0.23246 | 8.4003 | 0.002716 | 0.003888 | 0.003759 | 5.8e-05 | 0.00115 | 75.0414 | 1.3417 | 2.5475 | 100.21 | 20.72 | 37.808 | 67.1818 | 87.0042 | 63.6516 | 0.009233 | 1.0482 | 0.024889 | 23.744 | 0.007723 | 0.13829 | 1.2e-05 | 0.00094 | 0.00094 | 5.9847
2019-03-29 13:48 | 1553881707.5334 | -280 | 0.20953 | 0.000442 | -1e-06 | 100.3846 | 100.1888 | 126.3661 | 21.1114 | 0.001916 | 0.001848 | -151.7924 | 0.23348 | 8.4012 | 0.002677 | 0.004173 | 0.003951 | 5.9e-05 | 0.001149 | 75.1 | 1.3499 | 2.5473 | 100.21 | 20.72 | 37.806 | 67.785 | 86.0777 | 66.1568 | 0.009233 | 1.0405 | 0.024838 | 23.8713 | 0.017455 | 0.1389 | 1.8e-05 | 0.001404 | 0.001404 | 4.025
2019-03-29 13:48 | 1553881708.5348 | -279 | 0.20952 | 0.000447 | -1e-06 | 100.36 | 100.2131 | 126.3314 | 21.1118 | 0.00198 | 0.001799 | -151.7883 | 0.23384 | 8.4031 | 0.002694 | 0.004331 | 0.004177 | 5.6e-05 | 0.00115 | 75.1238 | 1.3631 | 2.5474 | 100.21 | 20.72 | 37.805 | 69.329 | 86.491 | 64.9764 | 0.009232 | 1.0441 | 0.024878 | 23.8264 | 0.013105 | 0.13911 | 1.1e-05 | 0.000897 | 0.000897 | 6.31
2019-03-29 13:48 | 1553881709.5345 | -278 | 0.20952 | 0.000448 | -1e-06 | 100.296 | 100.082 | 126.1408 | 21.1126 | 0.001896 | 0.001803 | -151.7759 | 0.23358 | 8.4032 | 0.002629 | 0.003941 | 0.003731 | 5.7e-05 | 0.001141 | 75.1625 | 1.3458 | 2.5473 | 100.21 | 20.72 | 37.805 | 67.3049 | 84.9145 | 65.0614 | 0.009232 | 1.0439 | 0.024856 | 23.8114 | 0.019099 | 0.13896 | 1e-05 | 0.000819 | 0.000819 | 6.9042
2019-03-29 13:48 | 1553881710.5335 | -277 | 0.20954 | 0.000443 | -1e-06 | 100.2633 | 100.0216 | 126.1086 | 21.1128 | 0.001965 | 0.001788 | -151.7677 | 0.2341 | 8.4028 | 0.002679 | 0.004383 | 0.004242 | 5.8e-05 | 0.001148 | 75.1494 | 1.2984 | 2.5474 | 100.21 | 20.72 | 37.804 | 68.954 | 86.1167 | 64.7175 | 0.009232 | 1.0449 | 0.024865 | 23.7963 | 0.021589 | 0.13926 | 1.7e-05 | 0.00138 | 0.00138 | 4.108
2019-03-29 13:48 | 1553881711.5339 | -276 | 0.20952 | 0.000442 | -1e-06 | 100.3022 | 100.1288 | 126.4344 | 21.113 | 0.001927 | 0.001792 | -151.7731 | 0.23319 | 8.4041 | 0.002708 | 0.004605 | 0.004354 | 6.2e-05 | 0.001155 | 75.1345 | 1.3658 | 2.5474 | 100.21 | 20.72 | 37.803 | 68.0441 | 86.8194 | 64.7939 | 0.009232 | 1.0447 | 0.024895 | 23.8297 | 0.015482 | 0.13873 | 1.8e-05 | 0.001395 | 0.001395 | 4.0478
2019-03-29 13:49 | 1553881712.5336 | -275 | 0.20954 | 0.000449 | 0 | 100.313 | 100.1865 | 126.6056 | 21.113 | 0.001952 | 0.001791 | -151.7926 | 0.23456 | 8.4007 | 0.002622 | 0.004105 | 0.003947 | 6.1e-05 | 0.001155 | 75.1869 | 1.3464 | 2.5473 | 100.21 | 20.72 | 37.803 | 68.6391 | 84.7587 | 64.7699 | 0.009232 | 1.0448 | 0.024912 | 23.845 | 0.011292 | 0.13954 | 8e-06 | 0.000661 | 0.000661 | 8.5869
2019-03-29 13:49 | 1553881713.535 | -274 | 0.20953 | 0.000446 | -1e-06 | 100.2421 | 100.0104 | 126.7253 | 21.1134 | 0.001944 | 0.001797 | -151.7913 | 0.23403 | 8.4031 | 0.002717 | 0.004741 | 0.004581 | 5.6e-05 | 0.001153 | 75.0781 | 1.3616 | 2.5473 | 100.21 | 20.73 | 37.801 | 68.4594 | 87.0446 | 64.9226 | 0.009237 | 1.0443 | 0.024918 | 23.8616 | 0.0207 | 0.13923 | 1.3e-05 | 0.000993 | 0.000993 | 5.7064
2019-03-29 13:49 | 1553881714.5342 | -273 | 0.20952 | 0.000445 | -1e-06 | 100.2077 | 100.0435 | 126.5405 | 21.113 | 0.001885 | 0.00182 | -151.7998 | 0.23395 | 8.4003 | 0.002702 | 0.004204 | 0.004042 | 6e-05 | 0.001149 | 75.1309 | 1.3514 | 2.5474 | 100.21 | 20.73 | 37.801 | 67.0282 | 86.6757 | 65.4863 | 0.009237 | 1.0426 | 0.02488 | 23.8641 | 0.014668 | 0.13918 | 1.4e-05 | 0.001049 | 0.001049 | 5.4004
2019-03-29 13:49 | 1553881715.5347 | -272 | 0.20953 | 0.000447 | -1e-06 | 100.2041 | 99.9418 | 126.307 | 21.1134 | 0.001858 | 0.001799 | -151.7796 | 0.23388 | 8.4048 | 0.002654 | 0.004587 | 0.004401 | 5.9e-05 | 0.001152 | 75.0672 | 1.3869 | 2.5473 | 100.21 | 20.73 | 37.799 | 66.3998 | 85.5163 | 64.9748 | 0.009237 | 1.0441 | 0.024875 | 23.8241 | 0.023445 | 0.13914 | 1.2e-05 | 0.000938 | 0.000938 | 6.0391
2019-03-29 13:49 | 1553881716.5347 | -271 | 0.20951 | 0.000445 | -1e-06 | 100.2211 | 99.8694 | 126.2044 | 21.1139 | 0.001913 | 0.001781 | -151.7802 | 0.23339 | 8.4028 | 0.002614 | 0.004155 | 0.004 | 5.8e-05 | 0.001151 | 75.0993 | 1.3494 | 2.5475 | 100.21 | 20.73 | 37.799 | 67.7086 | 84.5624 | 64.5366 | 0.009237 | 1.0455 | 0.024881 | 23.799 | 0.031445 | 0.13884 | 1.4e-05 | 0.001102 | 0.001102 | 5.1297
2019-03-29 13:49 | 1553881717.5328 | -270 | 0.20955 | 0.000445 | -1e-06 | 100.172 | 100.0346 | 125.6479 | 21.1141 | 0.001891 | 0.001803 | -151.7712 | 0.23369 | 8.4007 | 0.002684 | 0.003624 | 0.003429 | 5.7e-05 | 0.001143 | 75.0957 | 1.3856 | 2.5474 | 100.21 | 20.73 | 37.796 | 67.1815 | 86.2411 | 65.0596 | 0.009236 | 1.0439 | 0.024807 | 23.7648 | 0.012281 | 0.13902 | 1.4e-05 | 0.001111 | 0.001111 | 5.093
2019-03-29 13:49 | 1553881718.5327 | -269 | 0.20951 | 0.000442 | -1e-06 | 100.1427 | 99.9339 | 125.335 | 21.1143 | 0.001861 | 0.001817 | -151.7906 | 0.23401 | 8.4019 | 0.002652 | 0.004557 | 0.004402 | 5.9e-05 | 0.001143 | 75.1331 | 1.3536 | 2.5475 | 100.21 | 20.73 | 37.795 | 66.4695 | 85.4811 | 65.4167 | 0.009236 | 1.0428 | 0.024763 | 23.7477 | 0.018668 | 0.13921 | 1.8e-05 | 0.001354 | 0.001354 | 4.1852
2019-03-29 13:49 | 1553881719.5328 | -268 | 0.20951 | 0.000454 | -1e-06 | 100.1134 | 99.9674 | 125.1503 | 21.1146 | 0.001978 | 0.001764 | -151.7824 | 0.23418 | 8.4032 | 0.002594 | 0.004696 | 0.004464 | 5.9e-05 | 0.001144 | 75.0451 | 1.3341 | 2.5473 | 100.21 | 20.73 | 37.792 | 69.2817 | 84.0698 | 64.1297 | 0.009235 | 1.0467 | 0.024792 | 23.6851 | 0.013055 | 0.13932 | 2e-06 | 0.000214 | 0.000214 | 26.5277
2019-03-29 13:49 | 1553881720.5341 | -267 | 0.20948 | 0.000449 | -1e-06 | 100.0933 | 99.9295 | 125.0705 | 21.114 | 0.001885 | 0.001779 | -151.7842 | 0.23303 | 8.4046 | 0.00275 | 0.003676 | 0.003535 | 5.8e-05 | 0.001143 | 75.0873 | 1.3023 | 2.5475 | 100.21 | 20.73 | 37.79 | 67.041 | 87.8261 | 64.4867 | 0.009234 | 1.0456 | 0.024771 | 23.6901 | 0.014653 | 0.13863 | 9e-06 | 0.000752 | 0.000752 | 7.5013
2019-03-29 13:49 | 1553881721.533 | -266 | 0.20953 | 0.000452 | -1e-06 | 100.0296 | 99.7827 | 124.732 | 21.1132 | 0.001894 | 0.001797 | -151.7757 | 0.23338 | 8.4022 | 0.002704 | 0.003479 | 0.003393 | 6.1e-05 | 0.001139 | 75.0665 | 1.3022 | 2.5475 | 100.21 | 20.73 | 37.788 | 67.2446 | 86.7181 | 64.9294 | 0.009234 | 1.0443 | 0.024721 | 23.6735 | 0.022099 | 0.13884 | 6e-06 | 0.000546 | 0.000546 | 10.3556
2019-03-29 13:49 | 1553881722.533 | -265 | 0.20947 | 0.000451 | -1e-06 | 100.0598 | 99.8554 | 124.5674 | 21.1145 | 0.001979 | 0.001777 | -151.7736 | 0.23281 | 8.4022 | 0.002665 | 0.004253 | 0.004012 | 6.1e-05 | 0.001134 | 75.0817 | 1.3738 | 2.5474 | 100.21 | 20.73 | 37.784 | 69.3028 | 85.7827 | 64.4482 | 0.009233 | 1.0458 | 0.024723 | 23.641 | 0.018288 | 0.1385 | 7e-06 | 0.000537 | 0.000537 | 10.4879
2019-03-29 13:49 | 1553881723.5341 | -264 | 0.20948 | 0.000449 | -1e-06 | 100.0322 | 99.8216 | 124.6224 | 21.1142 | 0.001983 | 0.001841 | -151.7789 | 0.23237 | 8.4009 | 0.002746 | 0.004161 | 0.003953 | 6.1e-05 | 0.00113 | 75.0432 | 1.3253 | 2.5474 | 100.21 | 20.73 | 37.784 | 69.3984 | 87.7473 | 65.9916 | 0.009233 | 1.041 | 0.024672 | 23.7002 | 0.018853 | 0.13824 | 9e-06 | 0.000722 | 0.000722 | 7.7949
2019-03-29 13:49 | 1553881724.533 | -263 | 0.20955 | 0.000451 | -1e-06 | 100.0637 | 99.7145 | 124.9599 | 21.1149 | 0.001965 | 0.001744 | -151.7748 | 0.23292 | 8.4012 | 0.002766 | 0.004409 | 0.004273 | 6.3e-05 | 0.001127 | 75.0645 | 1.3422 | 2.5475 | 100.21 | 20.73 | 37.779 | 68.9537 | 88.2073 | 63.6395 | 0.009232 | 1.0483 | 0.024791 | 23.6499 | 0.031271 | 0.13856 | 7e-06 | 0.000591 | 0.000591 | 9.5409
2019-03-29 13:49 | 1553881725.534 | -262 | 0.20951 | 0.000448 | -1e-06 | 99.9954 | 99.6912 | 124.9249 | 21.1156 | 0.001924 | 0.001786 | -151.7994 | 0.23268 | 8.4022 | 0.002872 | 0.004407 | 0.004276 | 5.9e-05 | 0.001127 | 75.0624 | 1.3494 | 2.5474 | 100.21 | 20.73 | 37.779 | 67.981 | 90.7722 | 64.6639 | 0.009232 | 1.0451 | 0.02475 | 23.6825 | 0.02725 | 0.13842 | 1e-05 | 0.000784 | 0.000784 | 7.1828
2019-03-29 13:49 | 1553881726.533 | -261 | 0.20948 | 0.000444 | -1e-06 | 99.9936 | 99.7639 | 125.2726 | 21.1148 | 0.001948 | 0.001728 | -151.8024 | 0.23336 | 8.403 | 0.002796 | 0.003944 | 0.003714 | 6e-05 | 0.001123 | 75.0428 | 1.3526 | 2.5475 | 100.21 | 20.73 | 37.77 | 68.5529 | 88.9297 | 63.2664 | 0.00923 | 1.0494 | 0.024836 | 23.6663 | 0.020569 | 0.13882 | 1.6e-05 | 0.001248 | 0.001248 | 4.5271
2019-03-29 13:49 | 1553881727.5331 | -260 | 0.20952 | 0.000449 | -1e-06 | 99.9427 | 99.6878 | 125.6258 | 21.116 | 0.001896 | 0.001811 | -151.8083 | 0.23327 | 8.3994 | 0.002794 | 0.004255 | 0.004027 | 6.2e-05 | 0.00112 | 75.0107 | 1.3227 | 2.5472 | 100.21 | 20.73 | 37.77 | 67.2957 | 88.8957 | 65.2598 | 0.00923 | 1.0432 | 0.024798 | 23.7698 | 0.022842 | 0.13877 | 9e-06 | 0.000685 | 0.000685 | 8.2413
2019-03-29 13:49 | 1553881728.5346 | -259 | 0.20947 | 0.000447 | -1e-06 | 99.8987 | 99.7657 | 125.9286 | 21.1152 | 0.001946 | 0.001842 | -151.7943 | 0.23319 | 8.4029 | 0.00289 | 0.004497 | 0.0044 | 6.1e-05 | 0.00112 | 75.039 | 1.3369 | 2.5472 | 100.21 | 20.73 | 37.769 | 68.5064 | 91.2103 | 66.0021 | 0.009229 | 1.041 | 0.0248 | 23.8245 | 0.011924 | 0.13872 | 1.2e-05 | 0.000934 | 0.000934 | 6.0449
2019-03-29 13:49 | 1553881729.5332 | -258 | 0.2095 | 0.00045 | -1e-06 | 99.9305 | 99.6493 | 126.2352 | 21.1153 | 0.001935 | 0.001854 | -151.791 | 0.2325 | 8.4036 | 0.002865 | 0.00393 | 0.003744 | 6.4e-05 | 0.001116 | 75.0116 | 1.3419 | 2.5473 | 100.21 | 20.73 | 37.769 | 68.249 | 90.6009 | 66.2973 | 0.009229 | 1.0401 | 0.02482 | 23.8638 | 0.025205 | 0.13832 | 9e-06 | 0.000719 | 0.000719 | 7.8342
2019-03-29 13:49 | 1553881730.5342 | -257 | 0.20956 | 0.000445 | -1e-06 | 99.9092 | 99.553 | 126.633 | 21.1156 | 0.001964 | 0.001779 | -151.7952 | 0.23312 | 8.4024 | 0.002867 | 0.004985 | 0.004711 | 6.1e-05 | 0.001116 | 75.0658 | 1.3519 | 2.5472 | 100.21 | 20.73 | 37.765 | 68.9428 | 90.6496 | 64.4991 | 0.009228 | 1.0456 | 0.024925 | 23.838 | 0.031945 | 0.13868 | 1.4e-05 | 0.001131 | 0.001131 | 4.9912
2019-03-29 13:49 | 1553881731.5332 | -256 | 0.20948 | 0.000449 | -1e-06 | 99.8658 | 99.5605 | 126.8351 | 21.1152 | 0.00194 | 0.001787 | -151.7953 | 0.23205 | 8.4039 | 0.002909 | 0.004334 | 0.004065 | 6.2e-05 | 0.001118 | 75.0735 | 1.3623 | 2.5471 | 100.21 | 20.73 | 37.765 | 68.3643 | 91.6708 | 64.6955 | 0.009228 | 1.045 | 0.024938 | 23.864 | 0.027384 | 0.13805 | 1e-05 | 0.000793 | 0.000793 | 7.0873
2019-03-29 13:49 | 1553881732.5336 | -255 | 0.20958 | 0.000451 | -1e-06 | 99.8607 | 99.6687 | 126.849 | 21.1158 | 0.001886 | 0.001807 | -151.7944 | 0.23339 | 8.4027 | 0.002779 | 0.004421 | 0.004308 | 6.5e-05 | 0.001124 | 75.0337 | 1.3439 | 2.5473 | 100.21 | 20.74 | 37.761 | 67.0716 | 88.5245 | 65.1552 | 0.009233 | 1.0436 | 0.024922 | 23.8815 | 0.017214 | 0.13884 | 7e-06 | 0.000589 | 0.000589 | 9.5894
2019-03-29 13:49 | 1553881733.5333 | -254 | 0.20954 | 0.000455 | -1e-06 | 99.9023 | 99.68 | 126.7067 | 21.1157 | 0.001883 | 0.001758 | -151.7803 | 0.23412 | 8.401 | 0.002908 | 0.003988 | 0.003854 | 5.9e-05 | 0.001128 | 75.0351 | 1.3037 | 2.5473 | 100.21 | 20.74 | 37.758 | 67.0013 | 91.6494 | 63.9833 | 0.009232 | 1.0472 | 0.024951 | 23.8267 | 0.01993 | 0.13928 | 2e-06 | 0.000212 | 0.000212 | 26.6913
2019-03-29 13:49 | 1553881734.5324 | -253 | 0.20942 | 0.000451 | -1e-06 | 99.8547 | 99.5128 | 126.2112 | 21.1159 | 0.001913 | 0.001778 | -151.7839 | 0.23334 | 8.4022 | 0.002908 | 0.004455 | 0.004223 | 6e-05 | 0.00113 | 74.9642 | 1.2782 | 2.5473 | 100.21 | 20.74 | 37.756 | 67.7168 | 91.6369 | 64.476 | 0.009232 | 1.0457 | 0.024884 | 23.7975 | 0.030686 | 0.13882 | 6e-06 | 0.000474 | 0.000474 | 11.9203
2019-03-29 13:49 | 1553881735.5339 | -252 | 0.20948 | 0.000449 | -1e-06 | 99.7871 | 99.3384 | 126.3279 | 21.1154 | 0.001948 | 0.001795 | -151.7749 | 0.23226 | 8.4028 | 0.002831 | 0.004863 | 0.004627 | 6.1e-05 | 0.001136 | 75.0667 | 1.2965 | 2.5473 | 100.21 | 20.74 | 37.753 | 68.5541 | 89.7716 | 64.8884 | 0.009231 | 1.0444 | 0.024881 | 23.823 | 0.040314 | 0.13817 | 1e-05 | 0.000785 | 0.000785 | 7.1605
2019-03-29 13:49 | 1553881736.5325 | -251 | 0.20961 | 0.000447 | -1e-06 | 99.7979 | 99.5783 | 126.2063 | 21.1172 | 0.001902 | 0.001762 | -151.7675 | 0.23346 | 8.4031 | 0.002849 | 0.004596 | 0.004463 | 6e-05 | 0.001141 | 75.0258 | 1.3119 | 2.5473 | 100.21 | 20.74 | 37.753 | 67.463 | 90.2259 | 64.0728 | 0.009231 | 1.0469 | 0.024899 | 23.7828 | 0.019702 | 0.13889 | 1.1e-05 | 0.000881 | 0.000881 | 6.416
2019-03-29 13:49 | 1553881737.5333 | -250 | 0.20952 | 0.000446 | -1e-06 | 99.8449 | 99.5629 | 126.1924 | 21.1167 | 0.00192 | 0.001825 | -151.7764 | 0.23363 | 8.4034 | 0.002849 | 0.004962 | 0.004763 | 5.8e-05 | 0.001135 | 75.1087 | 1.3197 | 2.5473 | 100.21 | 20.74 | 37.753 | 67.8837 | 90.2286 | 65.6085 | 0.009231 | 1.0422 | 0.024841 | 23.8356 | 0.025293 | 0.13899 | 1.3e-05 | 0.001039 | 0.001039 | 5.4462
2019-03-29 13:49 | 1553881738.5335 | -249 | 0.20949 | 0.000445 | -1e-06 | 99.8116 | 99.5493 | 126.1199 | 21.1163 | 0.001912 | 0.001845 | -151.7705 | 0.23292 | 8.4028 | 0.002889 | 0.004379 | 0.004181 | 5.9e-05 | 0.001135 | 75.1191 | 1.326 | 2.5472 | 100.21 | 20.74 | 37.752 | 67.6946 | 91.1841 | 66.0765 | 0.009231 | 1.0407 | 0.024817 | 23.8452 | 0.023542 | 0.13856 | 1.4e-05 | 0.001072 | 0.001072 | 5.2585
2019-03-29 13:49 | 1553881739.5335 | -248 | 0.2095 | 0.000451 | -1e-06 | 99.8526 | 99.3271 | 126.1169 | 21.1168 | 0.001999 | 0.00178 | -151.7845 | 0.23367 | 8.4031 | 0.002842 | 0.004227 | 0.004056 | 6.1e-05 | 0.001136 | 75.0746 | 1.3591 | 2.5474 | 100.21 | 20.74 | 37.753 | 69.7907 | 90.0396 | 64.5096 | 0.009231 | 1.0456 | 0.024874 | 23.7898 | 0.047202 | 0.13901 | 7e-06 | 0.000555 | 0.000555 | 10.1854
2019-03-29 13:49 | 1553881740.5324 | -247 | 0.20952 | 0.000453 | -1e-06 | 99.7988 | 99.5622 | 126.088 | 21.1172 | 0.001882 | 0.001781 | -151.7966 | 0.23353 | 8.4038 | 0.002838 | 0.003764 | 0.003626 | 5.8e-05 | 0.001131 | 75.0532 | 1.344 | 2.5473 | 100.21 | 20.74 | 37.751 | 66.9645 | 89.9422 | 64.5482 | 0.009231 | 1.0454 | 0.024869 | 23.7884 | 0.021224 | 0.13893 | 4e-06 | 0.000383 | 0.000383 | 14.7549
2019-03-29 13:49 | 1553881741.5334 | -246 | 0.20947 | 0.000448 | -1e-06 | 99.7835 | 99.5844 | 126.107 | 21.1171 | 0.001924 | 0.001818 | -151.7938 | 0.23317 | 8.4037 | 0.00276 | 0.004203 | 0.003991 | 6e-05 | 0.001124 | 75.0604 | 1.374 | 2.5474 | 100.21 | 20.74 | 37.751 | 67.9731 | 88.0688 | 65.4294 | 0.009231 | 1.0427 | 0.024839 | 23.8212 | 0.017862 | 0.13871 | 1e-05 | 0.000821 | 0.000821 | 6.8804
2019-03-29 13:49 | 1553881742.5336 | -245 | 0.20949 | 0.000447 | -1e-06 | 99.6548 | 99.5837 | 125.7831 | 21.118 | 0.001942 | 0.001809 | -151.8126 | 0.23219 | 8.4019 | 0.002774 | 0.004404 | 0.004211 | 5.9e-05 | 0.001125 | 75.0267 | 1.3523 | 2.5474 | 100.21 | 20.74 | 37.752 | 68.4156 | 88.4195 | 65.215 | 0.009231 | 1.0434 | 0.024815 | 23.7831 | 0.006379 | 0.13813 | 1.2e-05 | 0.000993 | 0.000993 | 5.6639
2019-03-29 13:49 | 1553881743.5336 | -244 | 0.20952 | 0.000448 | -1e-06 | 99.767 | 99.4312 | 125.4465 | 21.1178 | 0.001927 | 0.001861 | -151.8255 | 0.23213 | 8.4019 | 0.00286 | 0.004003 | 0.003782 | 6e-05 | 0.001116 | 75.0258 | 1.3218 | 2.5472 | 100.21 | 20.74 | 37.752 | 68.0572 | 90.4922 | 66.4618 | 0.009231 | 1.0396 | 0.024736 | 23.7949 | 0.030161 | 0.1381 | 1e-05 | 0.000808 | 0.000808 | 6.9556
2019-03-29 13:49 | 1553881744.5336 | -243 | 0.2095 | 0.00045 | 0 | 99.7625 | 99.4486 | 125.7415 | 21.1185 | 0.001944 | 0.001821 | -151.8155 | 0.23301 | 8.4027 | 0.002798 | 0.003057 | 0.002915 | 5.7e-05 | 0.001113 | 75.1013 | 1.344 | 2.5473 | 100.21 | 20.74 | 37.752 | 68.4745 | 88.9886 | 65.5066 | 0.009231 | 1.0425 | 0.0248 | 23.7894 | 0.028193 | 0.13862 | 8e-06 | 0.000599 | 0.000599 | 9.4155
2019-03-29 13:49 | 1553881745.533 | -242 | 0.20948 | 0.000447 | -1e-06 | 99.7709 | 99.4268 | 125.7676 | 21.118 | 0.001909 | 0.001758 | -151.8289 | 0.23351 | 8.4013 | 0.002837 | 0.004163 | 0.003942 | 6.1e-05 | 0.001116 | 75.0672 | 1.3024 | 2.5473 | 100.21 | 20.74 | 37.753 | 67.6323 | 89.9274 | 63.9925 | 0.009231 | 1.0472 | 0.024858 | 23.7386 | 0.030909 | 0.13892 | 1.1e-05 | 0.000887 | 0.000887 | 6.3763
2019-03-29 13:49 | 1553881746.5343 | -241 | 0.20948 | 0.000451 | -1e-06 | 99.6662 | 99.4145 | 125.623 | 21.1168 | 0.001921 | 0.001811 | -151.8246 | 0.23334 | 8.402 | 0.002898 | 0.003928 | 0.003802 | 5.9e-05 | 0.001119 | 75.1218 | 1.3684 | 2.5473 | 100.21 | 20.74 | 37.752 | 67.9064 | 91.3978 | 65.2668 | 0.009231 | 1.0432 | 0.024797 | 23.7697 | 0.022617 | 0.13882 | 7e-06 | 0.000547 | 0.000547 | 10.3236
2019-03-29 13:49 | 1553881747.5327 | -240 | 0.2095 | 0.000447 | -1e-06 | 99.729 | 99.3742 | 125.3071 | 21.1176 | 0.001961 | 0.001811 | -151.8096 | 0.23331 | 8.4011 | 0.002812 | 0.004175 | 0.00406 | 6e-05 | 0.001124 | 75.0482 | 1.386 | 2.5474 | 100.21 | 20.74 | 37.751 | 68.8611 | 89.3247 | 65.2536 | 0.009231 | 1.0433 | 0.024767 | 23.7394 | 0.031881 | 0.1388 | 1.2e-05 | 0.000949 | 0.000949 | 5.9531
2019-03-29 13:49 | 1553881748.5331 | -239 | 0.20947 | 0.000448 | -1e-06 | 99.7224 | 99.4902 | 125.4481 | 21.1188 | 0.001949 | 0.001809 | -151.8235 | 0.23297 | 8.4014 | 0.002816 | 0.004323 | 0.004145 | 5.9e-05 | 0.001129 | 75.0651 | 1.3258 | 2.5473 | 100.21 | 20.74 | 37.748 | 68.5933 | 89.4157 | 65.219 | 0.00923 | 1.0434 | 0.024782 | 23.7515 | 0.020852 | 0.13859 | 1e-05 | 0.000783 | 0.000783 | 7.2002
2019-03-29 13:49 | 1553881749.5318 | -238 | 0.20949 | 0.00045 | -1e-06 | 99.741 | 99.479 | 125.2014 | 21.1182 | 0.001936 | 0.001823 | -151.807 | 0.23298 | 8.4007 | 0.002754 | 0.004478 | 0.004272 | 6.2e-05 | 0.001127 | 75.1175 | 1.3431 | 2.5473 | 100.21 | 20.74 | 37.747 | 68.2598 | 87.9378 | 65.5633 | 0.00923 | 1.0423 | 0.024745 | 23.7402 | 0.023529 | 0.1386 | 9e-06 | 0.0007 | 0.0007 | 8.0628
2019-03-29 13:49 | 1553881750.5328 | -237 | 0.20949 | 0.00045 | -1e-06 | 99.7275 | 99.4902 | 124.9296 | 21.1187 | 0.001953 | 0.001761 | -151.785 | 0.23239 | 8.4042 | 0.002848 | 0.003731 | 0.003647 | 6e-05 | 0.001127 | 75.0781 | 1.3638 | 2.5474 | 100.21 | 20.74 | 37.743 | 68.672 | 90.1985 | 64.0735 | 0.009229 | 1.0469 | 0.024772 | 23.6622 | 0.021308 | 0.13825 | 8e-06 | 0.000658 | 0.000658 | 8.5483
2019-03-29 13:49 | 1553881751.5318 | -236 | 0.20951 | 0.000448 | -1e-06 | 99.6934 | 99.3861 | 125.0146 | 21.1182 | 0.001942 | 0.001806 | -151.7934 | 0.2344 | 8.4034 | 0.002792 | 0.004364 | 0.004202 | 6.1e-05 | 0.001128 | 75.0794 | 1.3368 | 2.5473 | 100.21 | 20.74 | 37.743 | 68.4117 | 88.8485 | 65.1387 | 0.009229 | 1.0436 | 0.024742 | 23.7076 | 0.02761 | 0.13944 | 1e-05 | 0.000802 | 0.000802 | 7.0724
2019-03-29 13:49 | 1553881752.5333 | -235 | 0.20952 | 0.000451 | -1e-06 | 99.6695 | 99.494 | 125.3568 | 21.1184 | 0.00201 | 0.001787 | -151.7599 | 0.23365 | 8.4023 | 0.002798 | 0.004581 | 0.004312 | 6.4e-05 | 0.001132 | 75.0399 | 1.3568 | 2.5474 | 100.21 | 20.74 | 37.738 | 70.0449 | 88.9996 | 64.6854 | 0.009227 | 1.045 | 0.024792 | 23.7241 | 0.015761 | 0.139 | 7e-06 | 0.000569 | 0.000569 | 9.9461
2019-03-29 13:49 | 1553881753.5319 | -234 | 0.2095 | 0.000449 | -1e-06 | 99.7051 | 99.4128 | 125.6108 | 21.1191 | 0.001944 | 0.001854 | -151.7414 | 0.2336 | 8.4011 | 0.00284 | 0.003993 | 0.003889 | 6.1e-05 | 0.00114 | 75.0511 | 1.3302 | 2.5473 | 100.21 | 20.74 | 37.738 | 68.4549 | 90.0081 | 66.2923 | 0.009227 | 1.0401 | 0.024759 | 23.8046 | 0.026262 | 0.13897 | 1e-05 | 0.000778 | 0.000778 | 7.272
2019-03-29 13:49 | 1553881754.5319 | -233 | 0.20948 | 0.000452 | -1e-06 | 99.6907 | 99.4913 | 125.8213 | 21.1195 | 0.001977 | 0.001793 | -151.7494 | 0.2331 | 8.4027 | 0.002778 | 0.004424 | 0.004182 | 5.9e-05 | 0.001144 | 75.0582 | 1.3434 | 2.5473 | 100.21 | 20.74 | 37.736 | 69.2639 | 88.5062 | 64.8378 | 0.009227 | 1.0446 | 0.024833 | 23.7734 | 0.017913 | 0.13867 | 5e-06 | 0.000409 | 0.000409 | 13.7952
2019-03-29 13:49 | 1553881755.5333 | -232 | 0.20948 | 0.000448 | -1e-06 | 99.7341 | 99.4735 | 125.7657 | 21.1191 | 0.001913 | 0.001805 | -151.7355 | 0.23273 | 8.4029 | 0.002856 | 0.004975 | 0.004744 | 6.5e-05 | 0.001146 | 75.0807 | 1.3533 | 2.5473 | 100.21 | 20.74 | 37.736 | 67.7098 | 90.3949 | 65.1285 | 0.009227 | 1.0437 | 0.024816 | 23.7784 | 0.023402 | 0.13845 | 1e-05 | 0.000791 | 0.000791 | 7.1266
2019-03-29 13:49 | 1553881756.533 | -231 | 0.20949 | 0.000449 | -1e-06 | 99.6356 | 99.4759 | 125.7779 | 21.1192 | 0.001971 | 0.001765 | -151.7449 | 0.23342 | 8.4016 | 0.002689 | 0.004533 | 0.004315 | 6.1e-05 | 0.001149 | 75.0418 | 1.3687 | 2.5475 | 100.21 | 20.74 | 37.73 | 69.1062 | 86.3803 | 64.1498 | 0.009225 | 1.0467 | 0.024854 | 23.7451 | 0.014351 | 0.13886 | 9e-06 | 0.000721 | 0.000721 | 7.8355
2019-03-29 13:49 | 1553881757.5331 | -230 | 0.20946 | 0.000453 | -1e-06 | 99.7563 | 99.4138 | 125.9723 | 21.119 | 0.001942 | 0.001796 | -151.7415 | 0.2331 | 8.4027 | 0.002761 | 0.004 | 0.003867 | 6.1e-05 | 0.001149 | 75.0444 | 1.3158 | 2.5473 | 100.21 | 20.74 | 37.73 | 68.4103 | 88.0909 | 64.8999 | 0.009225 | 1.0444 | 0.024845 | 23.7899 | 0.030761 | 0.13867 | 5e-06 | 0.000446 | 0.000446 | 12.6571
2019-03-29 13:49 | 1553881758.5321 | -229 | 0.20948 | 0.000454 | -1e-06 | 99.7156 | 99.4237 | 125.8465 | 21.1192 | 0.001974 | 0.001793 | -151.7374 | 0.2339 | 8.4016 | 0.002741 | 0.004013 | 0.003894 | 6.1e-05 | 0.001141 | 75.0753 | 1.3517 | 2.5474 | 100.21 | 20.74 | 37.728 | 69.1966 | 87.611 | 64.8399 | 0.009225 | 1.0445 | 0.024835 | 23.7759 | 0.026219 | 0.13915 | 3e-06 | 0.000312 | 0.000312 | 18.1337
2019-03-29 13:49 | 1553881759.5331 | -228 | 0.20947 | 0.000453 | -1e-06 | 99.7218 | 99.4349 | 125.903 | 21.1191 | 0.001962 | 0.001793 | -151.7509 | 0.23321 | 8.4031 | 0.002759 | 0.00379 | 0.003647 | 5.6e-05 | 0.001141 | 75.0556 | 1.3392 | 2.5473 | 100.21 | 20.74 | 37.727 | 68.9017 | 88.0503 | 64.8218 | 0.009225 | 1.0446 | 0.024841 | 23.7806 | 0.02577 | 0.13873 | 4e-06 | 0.000354 | 0.000354 | 15.9295
2019-03-29 13:49 | 1553881760.5333 | -227 | 0.20951 | 0.000453 | -1e-06 | 99.797 | 99.6404 | 125.8176 | 21.1207 | 0.00192 | 0.001754 | -151.7638 | 0.23409 | 8.4018 | 0.002695 | 0.003549 | 0.003316 | 5.7e-05 | 0.001145 | 75.0873 | 1.3511 | 2.5473 | 100.21 | 20.74 | 37.722 | 67.8987 | 86.5038 | 63.8859 | 0.009223 | 1.0475 | 0.024867 | 23.7396 | 0.014044 | 0.13926 | 4e-06 | 0.000413 | 0.000413 | 13.7386
2019-03-29 13:49 | 1553881761.5321 | -226 | 0.20947 | 0.000454 | -1e-06 | 99.8182 | 99.5356 | 125.4089 | 21.1203 | 0.001962 | 0.0018 | -151.7485 | 0.23337 | 8.4047 | 0.002748 | 0.00401 | 0.0038 | 5.7e-05 | 0.001154 | 75.0644 | 1.3472 | 2.5475 | 100.21 | 20.74 | 37.722 | 68.8887 | 87.7853 | 65.0002 | 0.009223 | 1.044 | 0.024786 | 23.7401 | 0.025359 | 0.13883 | 4e-06 | 0.000348 | 0.000348 | 16.2418
2019-03-29 13:49 | 1553881762.5333 | -225 | 0.20948 | 0.000452 | -1e-06 | 99.7467 | 99.5001 | 125.0963 | 21.1198 | 0.001926 | 0.00178 | -151.755 | 0.23315 | 8.4024 | 0.002778 | 0.004012 | 0.003939 | 6e-05 | 0.001156 | 75.0276 | 1.3475 | 2.5474 | 100.21 | 20.74 | 37.72 | 68.024 | 88.5159 | 64.5087 | 0.009223 | 1.0456 | 0.024773 | 23.6933 | 0.022136 | 0.1387 | 5e-06 | 0.00044 | 0.00044 | 12.832
2019-03-29 13:49 | 1553881763.5322 | -224 | 0.20949 | 0.000451 | -1e-06 | 99.7287 | 99.5196 | 125.3399 | 21.1202 | 0.001973 | 0.0018 | -151.7649 | 0.23293 | 8.4033 | 0.002722 | 0.003811 | 0.003694 | 5.5e-05 | 0.001156 | 75.0262 | 1.3457 | 2.5474 | 100.21 | 20.74 | 37.718 | 69.168 | 87.1551 | 64.997 | 0.009222 | 1.0441 | 0.024779 | 23.7335 | 0.018777 | 0.13857 | 6e-06 | 0.000476 | 0.000476 | 11.8454
2019-03-29 13:49 | 1553881764.5327 | -223 | 0.20947 | 0.000449 | -1e-06 | 99.7931 | 99.4083 | 125.0568 | 21.1208 | 0.001931 | 0.001799 | -151.7607 | 0.23328 | 8.402 | 0.002707 | 0.004682 | 0.004464 | 5.6e-05 | 0.001158 | 75.0627 | 1.3533 | 2.5473 | 100.21 | 20.74 | 37.715 | 68.1452 | 86.8051 | 64.9702 | 0.009222 | 1.0441 | 0.024752 | 23.7057 | 0.034553 | 0.13878 | 9e-06 | 0.000673 | 0.000673 | 8.3912
2019-03-29 13:49 | 1553881765.5323 | -222 | 0.20951 | 0.000451 | -1e-06 | 99.8146 | 99.337 | 125.2472 | 21.1204 | 0.001908 | 0.001748 | -151.7642 | 0.23364 | 8.4031 | 0.0028 | 0.004331 | 0.004216 | 5.9e-05 | 0.001159 | 75.0399 | 1.386 | 2.5474 | 100.21 | 20.74 | 37.713 | 67.6019 | 89.0538 | 63.7606 | 0.009221 | 1.0479 | 0.024815 | 23.6813 | 0.042905 | 0.13899 | 8e-06 | 0.000645 | 0.000645 | 8.7676
2019-03-29 13:49 | 1553881766.5313 | -221 | 0.20948 | 0.00045 | -1e-06 | 99.8419 | 99.5329 | 125.7167 | 21.1208 | 0.001924 | 0.00181 | -151.7511 | 0.23334 | 8.4008 | 0.002801 | 0.004121 | 0.004003 | 5.7e-05 | 0.001161 | 75.0245 | 1.3598 | 2.5473 | 100.21 | 20.74 | 37.71 | 67.9954 | 89.0629 | 65.2379 | 0.009221 | 1.0433 | 0.024808 | 23.7776 | 0.027723 | 0.13881 | 7e-06 | 0.000593 | 0.000593 | 9.5221
2019-03-29 13:49 | 1553881767.5317 | -220 | 0.20951 | 0.000451 | -1e-06 | 99.814 | 99.4042 | 125.9835 | 21.1212 | 0.001994 | 0.001806 | -151.7564 | 0.23401 | 8.4005 | 0.00273 | 0.003996 | 0.003703 | 6e-05 | 0.001163 | 75.0051 | 1.3592 | 2.5473 | 100.21 | 20.74 | 37.709 | 69.6586 | 87.3691 | 65.14 | 0.00922 | 1.0436 | 0.024837 | 23.7994 | 0.036799 | 0.13921 | 6e-06 | 0.000505 | 0.000505 | 11.2273
2019-03-29 13:49 | 1553881768.5313 | -219 | 0.20951 | 0.00045 | -1e-06 | 99.7865 | 99.4732 | 125.9533 | 21.1214 | 0.001909 | 0.001816 | -151.7525 | 0.23346 | 8.4007 | 0.002759 | 0.004078 | 0.003996 | 5.7e-05 | 0.00116 | 75.0437 | 1.3561 | 2.5473 | 100.21 | 20.74 | 37.707 | 67.6137 | 88.0518 | 65.3901 | 0.00922 | 1.0428 | 0.024825 | 23.8053 | 0.02813 | 0.13889 | 7e-06 | 0.000581 | 0.000581 | 9.7268
2019-03-29 13:49 | 1553881769.5325 | -218 | 0.20952 | 0.000449 | -1e-06 | 99.8341 | 99.6912 | 125.9307 | 21.1223 | 0.001934 | 0.001807 | -151.7618 | 0.23353 | 8.4002 | 0.00274 | 0.004077 | 0.003864 | 6.1e-05 | 0.001156 | 75.02 | 1.3667 | 2.5474 | 100.21 | 20.74 | 37.706 | 68.2348 | 87.6084 | 65.1809 | 0.00922 | 1.0435 | 0.024831 | 23.7958 | 0.012807 | 0.13892 | 1e-05 | 0.000773 | 0.000773 | 7.3132
2019-03-29 13:49 | 1553881770.5328 | -217 | 0.20952 | 0.00045 | -1e-06 | 99.8158 | 99.3803 | 125.8854 | 21.1209 | 0.001949 | 0.001816 | -151.761 | 0.23318 | 8.4011 | 0.002765 | 0.003938 | 0.003737 | 5.9e-05 | 0.001151 | 74.991 | 1.3418 | 2.5473 | 100.21 | 20.74 | 37.704 | 68.5751 | 88.2052 | 65.3834 | 0.009219 | 1.0429 | 0.024819 | 23.7987 | 0.03911 | 0.13872 | 7e-06 | 0.000596 | 0.000596 | 9.4724
2019-03-29 13:50 | 1553881771.5326 | -216 | 0.20949 | 0.000449 | -1e-06 | 99.7991 | 99.6291 | 125.9561 | 21.1209 | 0.001973 | 0.001832 | -151.7755 | 0.23284 | 8.4023 | 0.002862 | 0.00416 | 0.003914 | 6.1e-05 | 0.001156 | 75.0112 | 1.3091 | 2.5474 | 100.21 | 20.74 | 37.704 | 69.1562 | 90.5288 | 65.7697 | 0.009219 | 1.0417 | 0.024812 | 23.8189 | 0.015243 | 0.13852 | 9e-06 | 0.000738 | 0.000738 | 7.6419
2019-03-29 13:50 | 1553881772.5315 | -215 | 0.20952 | 0.000451 | -1e-06 | 99.8272 | 99.5578 | 125.9199 | 21.1217 | 0.001999 | 0.001775 | -151.7733 | 0.23321 | 8.4021 | 0.0028 | 0.003781 | 0.003681 | 5.8e-05 | 0.00116 | 75.0182 | 1.3193 | 2.5473 | 100.21 | 20.74 | 37.704 | 69.7917 | 89.0318 | 64.402 | 0.009219 | 1.0459 | 0.024858 | 23.7674 | 0.024171 | 0.13874 | 5e-06 | 0.000443 | 0.000443 | 12.7434
2019-03-29 13:50 | 1553881773.5315 | -214 | 0.2095 | 0.00045 | -1e-06 | 99.8464 | 99.5131 | 125.4444 | 21.1219 | 0.001881 | 0.001796 | -151.775 | 0.23311 | 8.403 | 0.002785 | 0.004178 | 0.003988 | 6e-05 | 0.00116 | 74.9931 | 1.3376 | 2.5473 | 100.21 | 20.74 | 37.704 | 66.9535 | 88.6874 | 64.9191 | 0.009219 | 1.0443 | 0.024792 | 23.7406 | 0.029904 | 0.13868 | 8e-06 | 0.000634 | 0.000634 | 8.9025
2019-03-29 13:50 | 1553881774.532 | -213 | 0.20951 | 0.000449 | -1e-06 | 99.8056 | 99.5783 | 125.4573 | 21.1215 | 0.001932 | 0.001759 | -151.7861 | 0.23386 | 8.4027 | 0.002826 | 0.004323 | 0.004081 | 6.2e-05 | 0.001161 | 75.0412 | 1.3289 | 2.5472 | 100.21 | 20.74 | 37.705 | 68.1677 | 89.6634 | 64.0108 | 0.009219 | 1.0471 | 0.024827 | 23.7099 | 0.0204 | 0.13912 | 9e-06 | 0.000728 | 0.000728 | 7.776
2019-03-29 13:50 | 1553881775.5327 | -212 | 0.2095 | 0.000447 | -1e-06 | 99.8862 | 99.4206 | 125.2622 | 21.1217 | 0.001874 | 0.001759 | -151.7775 | 0.23234 | 8.4016 | 0.002767 | 0.003772 | 0.003624 | 5.5e-05 | 0.001162 | 75.0342 | 1.3137 | 2.5473 | 100.21 | 20.74 | 37.706 | 66.7832 | 88.2452 | 64.0062 | 0.00922 | 1.0471 | 0.024808 | 23.6913 | 0.041787 | 0.13822 | 1.1e-05 | 0.000876 | 0.000876 | 6.424
2019-03-29 13:50 | 1553881776.5322 | -211 | 0.20951 | 0.000447 | -1e-06 | 99.9026 | 99.7114 | 125.1528 | 21.1216 | 0.001942 | 0.001807 | -151.7866 | 0.23306 | 8.401 | 0.002893 | 0.003904 | 0.003777 | 5.8e-05 | 0.001158 | 75.0546 | 1.3404 | 2.5474 | 100.21 | 20.74 | 37.709 | 68.4268 | 91.2948 | 65.1615 | 0.00922 | 1.0436 | 0.024755 | 23.7215 | 0.017139 | 0.13865 | 1.1e-05 | 0.000866 | 0.000866 | 6.5184
2019-03-29 13:50 | 1553881777.5317 | -210 | 0.20956 | 0.000452 | -1e-06 | 99.929 | 99.6407 | 125.2629 | 21.121 | 0.001968 | 0.001801 | -151.7743 | 0.23345 | 8.4039 | 0.002847 | 0.004213 | 0.004048 | 6.1e-05 | 0.001157 | 75.048 | 1.3362 | 2.5473 | 100.21 | 20.74 | 37.71 | 69.0308 | 90.1726 | 65.0188 | 0.009221 | 1.044 | 0.024771 | 23.7269 | 0.025837 | 0.13888 | 5e-06 | 0.00044 | 0.00044 | 12.8326
2019-03-29 13:50 | 1553881778.5327 | -209 | 0.20949 | 0.000449 | -1e-06 | 99.9335 | 99.593 | 125.3556 | 21.1214 | 0.001951 | 0.001792 | -151.7733 | 0.2327 | 8.4035 | 0.002757 | 0.004225 | 0.004037 | 5.9e-05 | 0.001149 | 75.0595 | 1.3358 | 2.5474 | 100.21 | 20.74 | 37.712 | 68.6442 | 87.9972 | 64.8014 | 0.009221 | 1.0447 | 0.024788 | 23.7281 | 0.030529 | 0.13843 | 9e-06 | 0.000697 | 0.000697 | 8.0849
2019-03-29 13:50 | 1553881779.5307 | -208 | 0.2095 | 0.000451 | -1e-06 | 99.9814 | 99.6254 | 125.3815 | 21.1217 | 0.001954 | 0.001802 | -151.7571 | 0.23329 | 8.4008 | 0.002851 | 0.00434 | 0.004099 | 6e-05 | 0.001145 | 75.0638 | 1.3499 | 2.5473 | 100.21 | 20.74 | 37.714 | 68.7006 | 90.2662 | 65.0538 | 0.009222 | 1.0439 | 0.024781 | 23.7394 | 0.031903 | 0.13878 | 7e-06 | 0.000615 | 0.000615 | 9.1833
2019-03-29 13:50 | 1553881780.5321 | -207 | 0.20957 | 0.000447 | -1e-06 | 99.9538 | 99.6803 | 125.4901 | 21.1221 | 0.001912 | 0.001815 | -151.7583 | 0.23394 | 8.4022 | 0.002883 | 0.004005 | 0.003829 | 5.9e-05 | 0.001145 | 75.0383 | 1.3346 | 2.5473 | 100.21 | 20.74 | 37.717 | 67.7089 | 91.0404 | 65.354 | 0.009222 | 1.043 | 0.024781 | 23.7602 | 0.024507 | 0.13917 | 1.1e-05 | 0.000895 | 0.000895 | 6.3271
2019-03-29 13:50 | 1553881781.5308 | -206 | 0.20954 | 0.000452 | -1e-06 | 100.0301 | 99.7025 | 125.5984 | 21.1213 | 0.001939 | 0.001795 | -151.7518 | 0.23314 | 8.4037 | 0.00289 | 0.003469 | 0.003389 | 5.9e-05 | 0.00114 | 75.0622 | 1.3454 | 2.5474 | 100.21 | 20.74 | 37.717 | 68.3531 | 91.1991 | 64.8778 | 0.009222 | 1.0444 | 0.024809 | 23.7537 | 0.029345 | 0.13869 | 5e-06 | 0.000429 | 0.000429 | 13.1535
2019-03-29 13:50 | 1553881782.5308 | -205 | 0.20944 | 0.000452 | -1e-06 | 100.0017 | 99.6906 | 125.4678 | 21.1211 | 0.001918 | 0.001795 | -151.7529 | 0.23286 | 8.4019 | 0.002837 | 0.004114 | 0.003801 | 5.9e-05 | 0.00114 | 75.0387 | 1.4112 | 2.5473 | 100.21 | 20.74 | 37.719 | 67.839 | 89.9227 | 64.8846 | 0.009223 | 1.0444 | 0.024796 | 23.7416 | 0.027873 | 0.13853 | 5e-06 | 0.000464 | 0.000464 | 12.1602
2019-03-29 13:50 | 1553881783.5313 | -204 | 0.20947 | 0.000449 | -1e-06 | 100.0104 | 99.782 | 125.9601 | 21.1222 | 0.00191 | 0.001821 | -151.75 | 0.23264 | 8.4039 | 0.002855 | 0.003841 | 0.00369 | 6.1e-05 | 0.001134 | 75.0434 | 1.4083 | 2.5475 | 100.21 | 20.74 | 37.719 | 67.6544 | 90.3746 | 65.5172 | 0.009223 | 1.0425 | 0.024821 | 23.8104 | 0.020447 | 0.1384 | 8e-06 | 0.000642 | 0.000642 | 8.7744
2019-03-29 13:50 | 1553881784.5316 | -203 | 0.20939 | 0.000449 | -1e-06 | 100.061 | 99.7609 | 126.1335 | 21.1219 | 0.001936 | 0.001812 | -151.736 | 0.23353 | 8.4019 | 0.002901 | 0.004474 | 0.004338 | 5.7e-05 | 0.001132 | 75.0613 | 1.3955 | 2.5475 | 100.21 | 20.74 | 37.721 | 68.2852 | 91.485 | 65.2975 | 0.009223 | 1.0431 | 0.024846 | 23.8191 | 0.026867 | 0.13892 | 9e-06 | 0.000767 | 0.000767 | 7.3695
2019-03-29 13:50 | 1553881785.532 | -202 | 0.20946 | 0.00045 | -1e-06 | 100.1406 | 99.7933 | 126.0854 | 21.1217 | 0.00192 | 0.001755 | -151.7271 | 0.23357 | 8.4019 | 0.002817 | 0.004282 | 0.00415 | 6.2e-05 | 0.001131 | 75.0218 | 1.3004 | 2.5474 | 100.21 | 20.74 | 37.721 | 67.8785 | 89.4481 | 63.9226 | 0.009223 | 1.0474 | 0.024892 | 23.7661 | 0.031075 | 0.13895 | 8e-06 | 0.000668 | 0.000668 | 8.4642
2019-03-29 13:50 | 1553881786.531 | -201 | 0.2095 | 0.00045 | -1e-06 | 100.1376 | 99.8861 | 125.634 | 21.1224 | 0.001915 | 0.001796 | -151.7264 | 0.23306 | 8.4025 | 0.002862 | 0.004272 | 0.004104 | 5.6e-05 | 0.001132 | 74.9952 | 1.3331 | 2.5473 | 100.21 | 20.74 | 37.718 | 67.7693 | 90.5256 | 64.8993 | 0.009222 | 1.0444 | 0.024812 | 23.7579 | 0.022491 | 0.13865 | 9e-06 | 0.00075 | 0.00075 | 7.521
2019-03-29 13:50 | 1553881787.5317 | -200 | 0.20948 | 0.000443 | 0 | 100.1229 | 99.9564 | 125.132 | 21.1221 | 0.002 | 0.001796 | -151.7347 | 0.23366 | 8.4036 | 0.002886 | 0.004183 | 0.003957 | 6.2e-05 | 0.001133 | 74.9818 | 1.3222 | 2.5474 | 100.21 | 20.74 | 37.717 | 69.8084 | 91.1052 | 64.9005 | 0.009222 | 1.0444 | 0.024762 | 23.7104 | 0.014888 | 0.13901 | 1.6e-05 | 0.001199 | 0.001199 | 4.7202
2019-03-29 13:50 | 1553881788.5317 | -199 | 0.20947 | 0.000448 | -1e-06 | 100.2011 | 99.9148 | 125.2019 | 21.1229 | 0.00185 | 0.001858 | -151.7419 | 0.23347 | 8.4011 | 0.002825 | 0.004798 | 0.00464 | 6.1e-05 | 0.001129 | 75.0134 | 1.3266 | 2.5474 | 100.21 | 20.74 | 37.717 | 66.2055 | 89.6556 | 66.3888 | 0.009222 | 1.0398 | 0.024715 | 23.7692 | 0.025591 | 0.13889 | 1e-05 | 0.000786 | 0.000786 | 7.1912
2019-03-29 13:50 | 1553881789.5322 | -198 | 0.20953 | 0.000454 | -1e-06 | 100.29 | 100.0957 | 125.4101 | 21.1234 | 0.001873 | 0.001803 | -151.7543 | 0.23341 | 8.4019 | 0.002867 | 0.004558 | 0.00434 | 6.3e-05 | 0.001133 | 75.0877 | 1.3126 | 2.5473 | 100.21 | 20.74 | 37.714 | 66.7522 | 90.6642 | 65.0805 | 0.009222 | 1.0438 | 0.024783 | 23.7431 | 0.017345 | 0.13886 | 3e-06 | 0.000272 | 0.000272 | 20.7503
2019-03-29 13:50 | 1553881790.5312 | -197 | 0.20944 | 0.00045 | -1e-06 | 100.2214 | 100.0633 | 125.3338 | 21.123 | 0.001934 | 0.001805 | -151.7718 | 0.23365 | 8.402 | 0.002838 | 0.004274 | 0.004169 | 6.1e-05 | 0.001131 | 75.0654 | 1.3299 | 2.5472 | 100.21 | 20.74 | 37.714 | 68.2151 | 89.9555 | 65.1292 | 0.009222 | 1.0437 | 0.024774 | 23.7375 | 0.014127 | 0.139 | 7e-06 | 0.000591 | 0.000591 | 9.5752
2019-03-29 13:50 | 1553881791.5316 | -196 | 0.20946 | 0.000449 | -1e-06 | 100.3184 | 99.994 | 125.3481 | 21.1231 | 0.001919 | 0.001805 | -151.7749 | 0.23317 | 8.4027 | 0.002832 | 0.004376 | 0.004172 | 6.1e-05 | 0.001136 | 75.095 | 1.3395 | 2.5473 | 100.21 | 20.74 | 37.71 | 67.876 | 89.8205 | 65.1324 | 0.009221 | 1.0436 | 0.024775 | 23.739 | 0.028974 | 0.13872 | 9e-06 | 0.000693 | 0.000693 | 8.1486
2019-03-29 13:50 | 1553881792.5312 | -195 | 0.20946 | 0.000455 | -1e-06 | 100.3887 | 100.083 | 125.615 | 21.1235 | 0.001924 | 0.001852 | -151.7897 | 0.23336 | 8.4012 | 0.002821 | 0.00435 | 0.004176 | 6.1e-05 | 0.001146 | 75.0765 | 1.3417 | 2.5473 | 100.21 | 20.74 | 37.71 | 67.9793 | 89.5539 | 66.2591 | 0.009221 | 1.0402 | 0.02476 | 23.8038 | 0.027279 | 0.13883 | 2e-06 | 0.000172 | 0.000172 | 32.7929
2019-03-29 13:50 | 1553881793.5324 | -194 | 0.20945 | 0.000456 | -1e-06 | 100.3714 | 100.0906 | 125.6854 | 21.1232 | 0.001933 | 0.001779 | -151.7976 | 0.2334 | 8.4027 | 0.002763 | 0.00397 | 0.003775 | 5.9e-05 | 0.001156 | 75.0841 | 1.3575 | 2.5473 | 100.21 | 20.74 | 37.708 | 68.2045 | 88.1514 | 64.5096 | 0.00922 | 1.0456 | 0.024831 | 23.749 | 0.025061 | 0.13885 | 0 | 9.5e-05 | 9.5e-05 | 59.5789
2019-03-29 13:50 | 1553881794.5304 | -193 | 0.20942 | 0.000459 | -1e-06 | 100.4771 | 100.1851 | 125.8056 | 21.1225 | 0.00189 | 0.001752 | -151.8027 | 0.23416 | 8.4029 | 0.00281 | 0.004368 | 0.004214 | 5.8e-05 | 0.001157 | 75.0462 | 1.3631 | 2.5473 | 100.21 | 20.74 | 37.706 | 67.1726 | 89.287 | 63.8431 | 0.00922 | 1.0476 | 0.024868 | 23.7369 | 0.026028 | 0.1393 | -4e-06 | -0.000287 | -0.000287 | -19.7245
2019-03-29 13:50 | 1553881795.5314 | -192 | 0.20946 | 0.000461 | -1e-06 | 100.4432 | 100.4011 | 125.8058 | 21.1238 | 0.001949 | 0.00183 | -151.8217 | 0.23358 | 8.4032 | 0.002759 | 0.004369 | 0.004183 | 6.1e-05 | 0.001161 | 75.0331 | 1.4146 | 2.5473 | 100.21 | 20.74 | 37.705 | 68.5888 | 88.0504 | 65.7319 | 0.009219 | 1.0418 | 0.024798 | 23.8034 | 0.00375 | 0.13896 | 0.000211 | 0.015719 | 0.015719 | 0.35979
2019-03-29 13:50 | 1553881796.5319 | -191 | 0.20946 | 0.000467 | -1e-06 | 100.5294 | 100.4308 | 125.7512 | 21.1227 | 0.001908 | 0.001808 | -151.8284 | 0.23327 | 8.4022 | 0.002737 | 0.003849 | 0.003653 | 5.7e-05 | 0.001165 | 75.0347 | 1.3608 | 2.5474 | 100.21 | 20.74 | 37.705 | 67.5941 | 87.5183 | 65.2071 | 0.009219 | 1.0434 | 0.024812 | 23.7798 | 0.008778 | 0.13877 | 0.000163 | 0.012207 | 0.012207 | 0.4627
2019-03-29 13:50 | 1553881797.5304 | -190 | 0.20947 | 0.000465 | -1e-06 | 100.5743 | 100.4694 | 126.0307 | 21.123 | 0.001872 | 0.001782 | -151.8223 | 0.23381 | 8.3991 | 0.002728 | 0.00446 | 0.004263 | 5.6e-05 | 0.001163 | 75.0995 | 1.3445 | 2.5474 | 100.21 | 20.74 | 37.704 | 66.7477 | 87.3037 | 64.5734 | 0.009219 | 1.0454 | 0.024863 | 23.7839 | 0.009338 | 0.13909 | 0.000148 | 0.011067 | 0.011067 | 0.51153
2019-03-29 13:50 | 1553881798.5315 | -189 | 0.20946 | 0.000462 | -1e-06 | 100.5752 | 100.4445 | 126.1147 | 21.123 | 0.001939 | 0.001784 | -151.8125 | 0.23331 | 8.403 | 0.002805 | 0.004082 | 0.003969 | 5.5e-05 | 0.001158 | 75.0315 | 1.3478 | 2.5472 | 100.21 | 20.74 | 37.704 | 68.344 | 89.1548 | 64.6078 | 0.009219 | 1.0453 | 0.02487 | 23.793 | 0.011636 | 0.13879 | 0.000206 | 0.015463 | 0.015463 | 0.36532
2019-03-29 13:50 | 1553881799.5313 | -188 | 0.20947 | 0.000467 | 0 | 100.5941 | 100.4114 | 125.9312 | 21.123 | 0.001939 | 0.001793 | -151.7941 | 0.23332 | 8.4023 | 0.002747 | 0.004075 | 0.003817 | 5.5e-05 | 0.001151 | 75.1101 | 1.3434 | 2.5473 | 100.21 | 20.74 | 37.702 | 68.3486 | 87.7716 | 64.8343 | 0.009219 | 1.0446 | 0.024844 | 23.7837 | 0.016262 | 0.13881 | 0.00011 | 0.00822 | 0.00822 | 0.68727
2019-03-29 13:50 | 1553881800.5305 | -187 | 0.20946 | 0.000468 | -1e-06 | 100.5806 | 100.525 | 126.0171 | 21.1235 | 0.001898 | 0.001752 | -151.7679 | 0.23173 | 8.4033 | 0.002816 | 0.004181 | 0.004116 | 5.5e-05 | 0.001149 | 75.132 | 1.3285 | 2.5473 | 100.21 | 20.75 | 37.7 | 67.3537 | 89.4421 | 63.8548 | 0.009224 | 1.0476 | 0.024888 | 23.7573 | 0.004946 | 0.13786 | 0.000151 | 0.011337 | 0.011337 | 0.4949
2019-03-29 13:50 | 1553881801.5315 | -186 | 0.20947 | 0.000466 | 0 | 100.7036 | 100.5711 | 125.776 | 21.124 | 0.001999 | 0.00177 | -151.7792 | 0.23334 | 8.4 | 0.002724 | 0.004097 | 0.003974 | 5.8e-05 | 0.001148 | 75.1427 | 1.3148 | 2.5472 | 100.21 | 20.75 | 37.7 | 69.7869 | 87.2064 | 64.2723 | 0.009224 | 1.0463 | 0.024849 | 23.7492 | 0.011781 | 0.13881 | 0.000116 | 0.008692 | 0.008692 | 0.64999
2019-03-29 13:50 | 1553881802.5307 | -185 | 0.20947 | 0.000467 | -1e-06 | 100.7096 | 100.5305 | 125.7554 | 21.1243 | 0.001916 | 0.001813 | -151.7813 | 0.23284 | 8.4019 | 0.002824 | 0.003951 | 0.003697 | 5.8e-05 | 0.001148 | 75.1212 | 1.3034 | 2.5473 | 100.21 | 20.75 | 37.698 | 67.7969 | 89.6288 | 65.3241 | 0.009223 | 1.0431 | 0.024808 | 23.7843 | 0.015926 | 0.13851 | 9.6e-05 | 0.007161 | 0.007161 | 0.78724
2019-03-29 13:50 | 1553881803.5306 | -184 | 0.20949 | 0.000467 | -1e-06 | 100.7674 | 100.6622 | 125.2028 | 21.124 | 0.001907 | 0.0018 | -151.7718 | 0.23273 | 8.4017 | 0.002749 | 0.00457 | 0.004435 | 5.8e-05 | 0.001153 | 75.0889 | 1.3316 | 2.5474 | 100.21 | 20.75 | 37.698 | 67.5853 | 87.8154 | 65.0015 | 0.009223 | 1.044 | 0.024765 | 23.7207 | 0.009342 | 0.13845 | 1.6e-05 | 0.001258 | 0.001258 | 4.4802
2019-03-29 13:50 | 1553881804.5306 | -183 | 0.20949 | 0.000461 | 0 | 100.803 | 100.5458 | 125.2427 | 21.1243 | 0.001948 | 0.001815 | -151.7692 | 0.23266 | 8.4017 | 0.002776 | 0.004179 | 0.003986 | 5.6e-05 | 0.001155 | 75.0972 | 1.3128 | 2.5473 | 100.21 | 20.75 | 37.698 | 68.5597 | 88.4796 | 65.3693 | 0.009223 | 1.0429 | 0.024756 | 23.7373 | 0.02285 | 0.13841 | 3e-05 | 0.002279 | 0.002279 | 2.4714
2019-03-29 13:50 | 1553881805.5306 | -182 | 0.20949 | 0.000463 | -1e-06 | 100.8398 | 100.6363 | 125.4871 | 21.1235 | 0.001892 | 0.001785 | -151.7673 | 0.23307 | 8.4022 | 0.002689 | 0.004551 | 0.004316 | 5.6e-05 | 0.001153 | 75.0891 | 1.3435 | 2.5472 | 100.21 | 20.75 | 37.699 | 67.2202 | 86.3705 | 64.6458 | 0.009224 | 1.0451 | 0.024807 | 23.7351 | 0.018074 | 0.13865 | 2.3e-05 | 0.001785 | 0.001785 | 3.1608
2019-03-29 13:50 | 1553881806.5321 | -181 | 0.20947 | 0.000465 | -1e-06 | 100.8066 | 100.6946 | 125.7432 | 21.1242 | 0.001909 | 0.00177 | -151.7651 | 0.23276 | 8.4005 | 0.002787 | 0.003883 | 0.00375 | 5.7e-05 | 0.001154 | 75.1284 | 1.3286 | 2.5473 | 100.21 | 20.75 | 37.7 | 67.6327 | 88.731 | 64.281 | 0.009224 | 1.0463 | 0.024845 | 23.7464 | 0.009941 | 0.13847 | 9.5e-05 | 0.00714 | 0.00714 | 0.78929
2019-03-29 13:50 | 1553881807.5308 | -180 | 0.20948 | 0.000458 | -1e-06 | 100.8291 | 100.7827 | 125.4531 | 21.1232 | 0.001929 | 0.001767 | -151.7699 | 0.23315 | 8.4016 | 0.002683 | 0.004564 | 0.004356 | 5.9e-05 | 0.00115 | 75.0956 | 1.3341 | 2.5473 | 100.21 | 20.75 | 37.7 | 68.0961 | 86.236 | 64.216 | 0.009224 | 1.0465 | 0.024819 | 23.7168 | 0.004115 | 0.1387 | 5.5e-05 | 0.00416 | 0.00416 | 1.357
2019-03-29 13:50 | 1553881808.5308 | -179 | 0.20949 | 0.000458 | 0 | 100.8539 | 100.7103 | 125.2538 | 21.1247 | 0.0019 | 0.0018 | -151.7698 | 0.23309 | 8.4043 | 0.002755 | 0.004181 | 0.004025 | 5.6e-05 | 0.001153 | 75.0409 | 1.3153 | 2.5475 | 100.21 | 20.75 | 37.7 | 67.4073 | 87.9576 | 65.006 | 0.009224 | 1.044 | 0.02477 | 23.7256 | 0.012742 | 0.13867 | -1e-05 | -0.000707 | -0.000707 | -7.9844
2019-03-29 13:50 | 1553881809.5308 | -178 | 0.20949 | 0.000458 | -1e-06 | 100.9548 | 100.7315 | 125.1613 | 21.1244 | 0.001927 | 0.001773 | -151.7815 | 0.23323 | 8.4026 | 0.002797 | 0.004075 | 0.00387 | 5.9e-05 | 0.001151 | 75.01 | 1.3208 | 2.5473 | 100.21 | 20.75 | 37.701 | 68.0504 | 88.9797 | 64.3445 | 0.009224 | 1.0461 | 0.024785 | 23.6937 | 0.019805 | 0.13875 | 2.8e-05 | 0.002116 | 0.002116 | 2.6688
2019-03-29 13:50 | 1553881810.5302 | -177 | 0.20949 | 0.000461 | 0 | 100.827 | 100.8643 | 125.1536 | 21.1243 | 0.001903 | 0.001755 | -151.7994 | 0.23292 | 8.4026 | 0.002665 | 0.004174 | 0.003948 | 5.5e-05 | 0.001144 | 75.0226 | 1.3749 | 2.5473 | 100.21 | 20.75 | 37.702 | 67.4855 | 85.7992 | 63.9134 | 0.009224 | 1.0474 | 0.0248 | 23.6778 | -0.003307 | 0.13856 | -1.2e-05 | -0.000838 | -0.000838 | -6.7335
2019-03-29 13:50 | 1553881811.5303 | -176 | 0.2095 | 0.000456 | -1e-06 | 100.9527 | 100.9277 | 125.3742 | 21.124 | 0.002015 | 0.001866 | -151.7947 | 0.23268 | 8.4023 | 0.002765 | 0.004355 | 0.004188 | 5.6e-05 | 0.001143 | 75.0889 | 1.3467 | 2.5473 | 100.21 | 20.75 | 37.705 | 70.1823 | 88.2099 | 66.5897 | 0.009225 | 1.0392 | 0.024724 | 23.7926 | 0.002211 | 0.13842 | -5.1e-05 | -0.00377 | -0.00377 | -1.4945
2019-03-29 13:50 | 1553881812.5299 | -175 | 0.20953 | 0.000456 | 0 | 101.0155 | 100.9895 | 125.5627 | 21.1238 | 0.001944 | 0.001837 | -151.7832 | 0.23247 | 8.4041 | 0.002774 | 0.004016 | 0.003814 | 5.7e-05 | 0.001139 | 75.1026 | 1.2588 | 2.5473 | 100.21 | 20.75 | 37.705 | 68.4602 | 88.4283 | 65.8863 | 0.009225 | 1.0413 | 0.024769 | 23.7858 | 0.002306 | 0.13829 | -0.00022 | -0.016377 | -0.016377 | -0.34368
2019-03-29 13:50 | 1553881813.5306 | -174 | 0.20953 | 0.000453 | -1e-06 | 101.0009 | 101.0615 | 125.715 | 21.1242 | 0.001917 | 0.001783 | -151.7864 | 0.2327 | 8.4018 | 0.002642 | 0.004362 | 0.004206 | 5.8e-05 | 0.001141 | 75.0592 | 1.3055 | 2.5472 | 100.21 | 20.75 | 37.71 | 67.8076 | 85.2543 | 64.5881 | 0.009226 | 1.0453 | 0.024831 | 23.7546 | -0.005369 | 0.13843 | -0.00022 | -0.016378 | -0.016378 | -0.34402
2019-03-29 13:50 | 1553881814.531 | -173 | 0.20952 | 0.000454 | -1e-06 | 101.0688 | 101.0243 | 125.5826 | 21.1238 | 0.001906 | 0.001794 | -151.7841 | 0.23256 | 8.402 | 0.002759 | 0.004257 | 0.003979 | 5.6e-05 | 0.001142 | 75.0439 | 1.3244 | 2.5474 | 100.21 | 20.75 | 37.71 | 67.5559 | 88.058 | 64.86 | 0.009226 | 1.0445 | 0.024808 | 23.7516 | 0.00394 | 0.13835 | -0.000191 | -0.014192 | -0.014192 | -0.39677
2019-03-29 13:50 | 1553881815.531 | -172 | 0.20954 | 0.000459 | -1e-06 | 101.1799 | 100.9499 | 125.4939 | 21.1246 | 0.001923 | 0.001767 | -151.7764 | 0.23258 | 8.403 | 0.00279 | 0.003941 | 0.003817 | 5.8e-05 | 0.001151 | 75.0645 | 1.3412 | 2.5474 | 100.21 | 20.75 | 37.715 | 67.9585 | 88.7986 | 64.2174 | 0.009227 | 1.0465 | 0.024823 | 23.7207 | 0.020352 | 0.13836 | -0.000312 | -0.023205 | -0.023205 | -0.24267
2019-03-29 13:50 | 1553881816.5302 | -171 | 0.20952 | 0.000451 | -1e-06 | 101.1215 | 101.0178 | 125.7333 | 21.1241 | 0.001949 | 0.001815 | -151.7856 | 0.23256 | 8.4025 | 0.002832 | 0.004466 | 0.004347 | 5.7e-05 | 0.001155 | 74.9956 | 1.3435 | 2.5473 | 100.21 | 20.75 | 37.715 | 68.5977 | 89.8169 | 65.3699 | 0.009227 | 1.0429 | 0.024804 | 23.7838 | 0.009175 | 0.13835 | -0.000183 | -0.013554 | -0.013554 | -0.41545
2019-03-29 13:50 | 1553881817.531 | -170 | 0.20952 | 0.000451 | -1e-06 | 101.2158 | 101.0957 | 126.0331 | 21.1242 | 0.001981 | 0.001823 | -151.8119 | 0.2328 | 8.4032 | 0.002825 | 0.003788 | 0.003476 | 5.6e-05 | 0.001153 | 75.0769 | 1.3462 | 2.5474 | 100.21 | 20.75 | 37.716 | 69.3613 | 89.6584 | 65.5543 | 0.009228 | 1.0423 | 0.024827 | 23.8186 | 0.010624 | 0.1385 | -0.000159 | -0.011796 | -0.011796 | -0.47785
2019-03-29 13:50 | 1553881818.5311 | -169 | 0.20954 | 0.000453 | -1e-06 | 101.2062 | 101.0567 | 125.784 | 21.1235 | 0.001942 | 0.0018 | -151.8215 | 0.23333 | 8.4037 | 0.002839 | 0.004379 | 0.004256 | 5.5e-05 | 0.001153 | 75.0954 | 1.3631 | 2.5474 | 100.21 | 20.75 | 37.716 | 68.4209 | 89.9929 | 65.0149 | 0.009228 | 1.044 | 0.024822 | 23.7761 | 0.013221 | 0.13881 | -0.00031 | -0.023082 | -0.023082 | -0.24476
2019-03-29 13:50 | 1553881819.5306 | -168 | 0.2095 | 0.000452 | 0 | 101.3272 | 101.1458 | 125.781 | 21.1243 | 0.001942 | 0.00179 | -151.8135 | 0.23328 | 8.4037 | 0.002796 | 0.004048 | 0.003895 | 5.7e-05 | 0.001155 | 75.0686 | 1.3726 | 2.5474 | 100.21 | 20.75 | 37.719 | 68.4123 | 88.9445 | 64.7628 | 0.009228 | 1.0448 | 0.024831 | 23.767 | 0.016021 | 0.13878 | -7.3e-05 | -0.005387 | -0.005387 | -1.0485
2019-03-29 13:50 | 1553881820.5308 | -167 | 0.20954 | 0.000451 | 0 | 101.2996 | 101.238 | 125.7509 | 21.1247 | 0.001943 | 0.00179 | -151.8098 | 0.23353 | 8.4021 | 0.002713 | 0.004398 | 0.004274 | 5.5e-05 | 0.00115 | 75.0116 | 1.3877 | 2.5473 | 100.21 | 20.75 | 37.719 | 68.4531 | 86.9456 | 64.7649 | 0.009228 | 1.0448 | 0.024828 | 23.7642 | 0.005445 | 0.13893 | -0.000287 | -0.02139 | -0.02139 | -0.26435
2019-03-29 13:50 | 1553881821.5312 | -166 | 0.20954 | 0.000449 | -1e-06 | 101.3652 | 101.1728 | 125.8035 | 21.124 | 0.00195 | 0.001803 | -151.7882 | 0.23398 | 8.4027 | 0.002772 | 0.00413 | 0.00392 | 5.3e-05 | 0.001143 | 75.0781 | 1.3851 | 2.5473 | 100.21 | 20.75 | 37.722 | 68.6039 | 88.3808 | 65.0787 | 0.009229 | 1.0438 | 0.024822 | 23.7802 | 0.016992 | 0.1392 | -0.000267 | -0.01989 | -0.01989 | -0.28483
2019-03-29 13:50 | 1553881822.5312 | -165 | 0.2095 | 0.000451 | -1e-06 | 101.3658 | 101.3407 | 125.7655 | 21.1246 | 0.00189 | 0.001815 | -151.7798 | 0.23377 | 8.4014 | 0.002729 | 0.003994 | 0.003775 | 5.2e-05 | 0.001136 | 75.0269 | 1.3921 | 2.5475 | 100.21 | 20.75 | 37.723 | 67.1737 | 87.339 | 65.3658 | 0.009229 | 1.0429 | 0.024808 | 23.7867 | 0.002212 | 0.13907 | -1.3e-05 | -0.000921 | -0.000921 | -6.1455
2019-03-29 13:50 | 1553881823.5313 | -164 | 0.2095 | 0.000451 | 0 | 101.4346 | 101.2052 | 125.7371 | 21.1251 | 0.001909 | 0.001813 | -151.7846 | 0.2339 | 8.4004 | 0.002709 | 0.003924 | 0.003791 | 5.3e-05 | 0.001133 | 75.0773 | 1.3923 | 2.5474 | 100.21 | 20.75 | 37.725 | 67.6315 | 86.8677 | 65.3193 | 0.00923 | 1.0431 | 0.024807 | 23.7824 | 0.020252 | 0.13915 | -6.1e-05 | -0.004506 | -0.004506 | -1.2568
2019-03-29 13:50 | 1553881824.5318 | -163 | 0.2095 | 0.000447 | 0 | 101.431 | 101.2506 | 125.5286 | 21.1249 | 0.001941 | 0.001775 | -151.8042 | 0.23277 | 8.4019 | 0.002796 | 0.003708 | 0.00345 | 5.3e-05 | 0.001142 | 75.1004 | 1.338 | 2.5473 | 100.21 | 20.75 | 37.729 | 68.3918 | 88.9521 | 64.4044 | 0.009231 | 1.0459 | 0.02482 | 23.7305 | 0.015924 | 0.13848 | -2.2e-05 | -0.001606 | -0.001606 | -3.5096
2019-03-29 13:50 | 1553881825.5303 | -162 | 0.20952 | 0.000451 | 0 | 101.4801 | 101.5287 | 125.4493 | 21.1255 | 0.001958 | 0.001795 | -151.8148 | 0.23395 | 8.401 | 0.002748 | 0.004225 | 0.004101 | 5.3e-05 | 0.001145 | 75.0669 | 1.3278 | 2.5473 | 100.21 | 20.75 | 37.731 | 68.8041 | 87.7911 | 64.8951 | 0.009231 | 1.0444 | 0.024794 | 23.7403 | -0.004286 | 0.13918 | -0.000174 | -0.012932 | -0.012932 | -0.43803
2019-03-29 13:50 | 1553881826.5308 | -161 | 0.20951 | 0.000455 | -1e-06 | 101.5178 | 101.5045 | 125.4486 | 21.1263 | 0.001915 | 0.001783 | -151.8058 | 0.23328 | 8.3999 | 0.002711 | 0.004201 | 0.004001 | 5.5e-05 | 0.001151 | 75.0667 | 1.317 | 2.5474 | 100.21 | 20.75 | 37.732 | 67.7776 | 86.9026 | 64.6009 | 0.009232 | 1.0453 | 0.024804 | 23.7299 | 0.001173 | 0.13878 | -9.6e-05 | -0.007124 | -0.007124 | -0.79292
2019-03-29 13:50 | 1553881827.5301 | -160 | 0.20958 | 0.000449 | -1e-06 | 101.5053 | 101.4441 | 125.7293 | 21.1257 | 0.00191 | 0.001775 | -151.784 | 0.23335 | 8.4009 | 0.002688 | 0.004296 | 0.004151 | 5.5e-05 | 0.001153 | 75.0543 | 1.2803 | 2.5473 | 100.21 | 20.75 | 37.733 | 67.664 | 86.3524 | 64.4076 | 0.009232 | 1.0459 | 0.024839 | 23.7496 | 0.00539 | 0.13882 | -0.000522 | -0.038913 | -0.038913 | -0.1452
2019-03-29 13:50 | 1553881828.5294 | -159 | 0.20959 | 0.000453 | -1e-06 | 101.4837 | 101.5902 | 125.8889 | 21.1256 | 0.001907 | 0.001806 | -151.7728 | 0.23319 | 8.4015 | 0.002772 | 0.003954 | 0.003795 | 5.6e-05 | 0.001155 | 75.0602 | 1.3157 | 2.5475 | 100.21 | 20.75 | 37.735 | 67.5875 | 88.3698 | 65.1404 | 0.009232 | 1.0436 | 0.024828 | 23.7904 | -0.00938 | 0.13872 | -0.000581 | -0.043289 | -0.043289 | -0.13043
2019-03-29 13:50 | 1553881829.5309 | -158 | 0.20949 | 0.000447 | 0 | 101.5894 | 101.5468 | 125.9474 | 21.1255 | 0.001945 | 0.00188 | -151.7886 | 0.23292 | 8.4037 | 0.00286 | 0.003929 | 0.003647 | 5.1e-05 | 0.001149 | 75.0536 | 1.3114 | 2.5473 | 100.21 | 20.75 | 37.74 | 68.489 | 90.4935 | 66.9413 | 0.009234 | 1.0381 | 0.024768 | 23.8592 | 0.003746 | 0.13856 | 0 | 3.8e-05 | 3.8e-05 | 148.6407
2019-03-29 13:50 | 1553881830.5298 | -157 | 0.20954 | 0.000446 | -1e-06 | 101.6322 | 101.6328 | 125.7812 | 21.126 | 0.001967 | 0.001759 | -151.7733 | 0.23206 | 8.4004 | 0.002741 | 0.003528 | 0.003412 | 5.3e-05 | 0.001145 | 75.0931 | 1.3061 | 2.5474 | 100.21 | 20.76 | 37.742 | 69.0293 | 87.6151 | 64.0096 | 0.00924 | 1.0471 | 0.024859 | 23.7405 | -5.9e-05 | 0.13805 | -0.000251 | -0.018659 | -0.018659 | -0.30113
2019-03-29 13:50 | 1553881831.5305 | -156 | 0.20954 | 0.000447 | 0 | 101.613 | 101.6178 | 125.8516 | 21.1263 | 0.001889 | 0.001769 | -151.7858 | 0.23324 | 8.4039 | 0.002837 | 0.003901 | 0.003697 | 5.3e-05 | 0.001141 | 75.0784 | 1.3387 | 2.5475 | 100.21 | 20.75 | 37.746 | 67.1481 | 89.9416 | 64.2707 | 0.009235 | 1.0463 | 0.024856 | 23.7563 | -0.000423 | 0.13875 | -0.000251 | -0.018775 | -0.018775 | -0.3008
2019-03-29 13:51 | 1553881832.5305 | -155 | 0.20946 | 0.000445 | -1e-06 | 101.6199 | 101.6939 | 125.5817 | 21.1271 | 0.001929 | 0.001836 | -151.7833 | 0.23317 | 8.403 | 0.002746 | 0.004495 | 0.004296 | 5.2e-05 | 0.001134 | 75.0979 | 1.3566 | 2.5473 | 100.21 | 20.76 | 37.75 | 68.1044 | 87.7435 | 65.8645 | 0.009242 | 1.0414 | 0.024771 | 23.7868 | -0.006514 | 0.13872 | 0.000196 | 0.014601 | 0.014601 | 0.38666
2019-03-29 13:51 | 1553881833.53 | -154 | 0.20949 | 0.000448 | -1e-06 | 101.6088 | 101.6731 | 125.9615 | 21.1269 | 0.001941 | 0.001823 | -151.7724 | 0.23273 | 8.4057 | 0.002853 | 0.004115 | 0.004004 | 5.2e-05 | 0.001126 | 75.0957 | 1.3694 | 2.5474 | 100.21 | 20.75 | 37.753 | 68.4076 | 90.3353 | 65.5607 | 0.009237 | 1.0423 | 0.02482 | 23.8121 | -0.005657 | 0.13845 | 5.9e-05 | 0.004482 | 0.004482 | 1.2572
2019-03-29 13:51 | 1553881834.531 | -153 | 0.20955 | 0.000448 | -1e-06 | 101.6699 | 101.655 | 126.2248 | 21.1261 | 0.001954 | 0.001806 | -151.7616 | 0.23294 | 8.4037 | 0.002848 | 0.00445 | 0.004257 | 5.6e-05 | 0.001119 | 75.125 | 1.3723 | 2.5474 | 100.21 | 20.76 | 37.758 | 68.7095 | 90.2056 | 65.1425 | 0.009244 | 1.0436 | 0.024861 | 23.8222 | 0.001308 | 0.13858 | -0.000326 | -0.024291 | -0.024291 | -0.23219
2019-03-29 13:51 | 1553881835.531 | -152 | 0.20962 | 0.000448 | -1e-06 | 101.6193 | 101.5967 | 126.1638 | 21.1274 | 0.001884 | 0.001764 | -151.7605 | 0.23404 | 8.4044 | 0.002841 | 0.003753 | 0.003554 | 5.6e-05 | 0.001117 | 75.143 | 1.3435 | 2.5474 | 100.21 | 20.76 | 37.76 | 67.022 | 90.0381 | 64.1491 | 0.009244 | 1.0467 | 0.024892 | 23.7815 | 0.001993 | 0.13923 | -0.000737 | -0.055014 | -0.055014 | -0.103
2019-03-29 13:51 | 1553881836.5296 | -151 | 0.20943 | 0.000447 | -1e-06 | 101.6908 | 101.4936 | 125.8911 | 21.1269 | 0.001896 | 0.001744 | -151.7475 | 0.23334 | 8.4035 | 0.002828 | 0.004314 | 0.004099 | 5.4e-05 | 0.001114 | 75.0699 | 1.3282 | 2.5474 | 100.21 | 20.76 | 37.761 | 67.3157 | 89.7157 | 63.656 | 0.009244 | 1.0482 | 0.024883 | 23.7384 | 0.017366 | 0.13881 | 0.000414 | 0.030995 | 0.030995 | 0.18228
2019-03-29 13:51 | 1553881837.5309 | -150 | 0.20947 | 0.000451 | -1e-06 | 101.6585 | 101.5806 | 125.3413 | 21.1264 | 0.001972 | 0.001771 | -151.7428 | 0.23364 | 8.3994 | 0.002842 | 0.003612 | 0.003493 | 5.7e-05 | 0.001117 | 75.1306 | 1.3234 | 2.5474 | 100.21 | 20.76 | 37.766 | 69.1504 | 90.0511 | 64.3137 | 0.009246 | 1.0462 | 0.024804 | 23.7096 | 0.006856 | 0.13899 | 0.000117 | 0.008744 | 0.008744 | 0.647
2019-03-29 13:51 | 1553881838.5304 | -149 | 0.20953 | 0.000443 | -1e-06 | 101.6271 | 101.7885 | 125.2322 | 21.1276 | 0.001958 | 0.001842 | -151.7585 | 0.23402 | 8.4012 | 0.002785 | 0.00473 | 0.004548 | 5.8e-05 | 0.001122 | 75.1295 | 1.2947 | 2.5473 | 100.21 | 20.76 | 37.767 | 68.8086 | 88.6766 | 66.0165 | 0.009246 | 1.0409 | 0.024731 | 23.759 | -0.014193 | 0.13922 | -0.000218 | -0.016128 | -0.016128 | -0.35133
2019-03-29 13:51 | 1553881839.5297 | -148 | 0.2095 | 0.000445 | 0 | 101.6908 | 101.6284 | 125.1432 | 21.1272 | 0.001929 | 0.00181 | -151.7483 | 0.2335 | 8.4017 | 0.002811 | 0.004125 | 0.00399 | 5.7e-05 | 0.001126 | 75.123 | 1.2955 | 2.5472 | 100.21 | 20.76 | 37.768 | 68.1101 | 89.3051 | 65.2411 | 0.009246 | 1.0433 | 0.024751 | 23.7234 | 0.005494 | 0.13891 | -4.1e-05 | -0.002995 | -0.002995 | -1.8877
2019-03-29 13:51 | 1553881840.531 | -147 | 0.2095 | 0.000451 | -1e-06 | 101.6726 | 101.6946 | 125.4153 | 21.1277 | 0.001957 | 0.001839 | -151.764 | 0.23351 | 8.4024 | 0.002802 | 0.004113 | 0.003908 | 5.7e-05 | 0.001135 | 75.1211 | 1.3307 | 2.5473 | 100.21 | 20.76 | 37.769 | 68.7923 | 89.1024 | 65.9478 | 0.009246 | 1.0411 | 0.024752 | 23.774 | -0.001938 | 0.13891 | -2.1e-05 | -0.001497 | -0.001497 | -3.7774
2019-03-29 13:51 | 1553881841.5296 | -146 | 0.20952 | 0.000447 | 0 | 101.7525 | 101.7267 | 125.4047 | 21.1272 | 0.00194 | 0.001836 | -151.7535 | 0.23306 | 8.402 | 0.002828 | 0.004721 | 0.004505 | 5.3e-05 | 0.001141 | 75.1033 | 1.34 | 2.5474 | 100.21 | 20.76 | 37.769 | 68.3846 | 89.7229 | 65.885 | 0.009246 | 1.0413 | 0.024753 | 23.7708 | 0.002269 | 0.13865 | -0.000168 | -0.012485 | -0.012485 | -0.45198
2019-03-29 13:51 | 1553881842.5297 | -145 | 0.20952 | 0.000449 | -1e-06 | 101.7244 | 101.7165 | 125.4275 | 21.1274 | 0.001948 | 0.001793 | -151.7618 | 0.23362 | 8.4034 | 0.002817 | 0.003888 | 0.003734 | 5.7e-05 | 0.001147 | 75.115 | 1.3349 | 2.5474 | 100.21 | 20.76 | 37.769 | 68.5596 | 89.4613 | 64.8431 | 0.009246 | 1.0445 | 0.024793 | 23.7364 | 0.000696 | 0.13898 | -0.000138 | -0.010219 | -0.010219 | -0.55354
2019-03-29 13:51 | 1553881843.5288 | -144 | 0.2095 | 0.000449 | -1e-06 | 101.7301 | 101.5284 | 125.0977 | 21.1277 | 0.001927 | 0.001778 | -151.7679 | 0.23336 | 8.4048 | 0.002864 | 0.004392 | 0.004186 | 6e-05 | 0.001145 | 75.1058 | 1.331 | 2.5474 | 100.21 | 20.75 | 37.768 | 68.0716 | 90.5838 | 64.4798 | 0.00924 | 1.0457 | 0.024774 | 23.6924 | 0.017747 | 0.13883 | -3.1e-05 | -0.002249 | -0.002249 | -2.5126
2019-03-29 13:51 | 1553881844.5294 | -143 | 0.20952 | 0.000451 | 0 | 101.6837 | 101.6898 | 124.974 | 21.1279 | 0.001952 | 0.001848 | -151.7685 | 0.23356 | 8.3997 | 0.002803 | 0.004192 | 0.004046 | 5.6e-05 | 0.001136 | 75.0543 | 1.3103 | 2.5475 | 100.21 | 20.75 | 37.767 | 68.6598 | 89.1213 | 66.1642 | 0.00924 | 1.0405 | 0.0247 | 23.7397 | -0.000544 | 0.13895 | -0.000155 | -0.011463 | -0.011463 | -0.49333
2019-03-29 13:51 | 1553881845.5298 | -142 | 0.2095 | 0.000455 | -1e-06 | 101.684 | 101.6601 | 125.0991 | 21.1276 | 0.001931 | 0.00184 | -151.7537 | 0.23327 | 8.4028 | 0.002825 | 0.004971 | 0.004748 | 5.8e-05 | 0.00113 | 75.0471 | 1.3404 | 2.5474 | 100.21 | 20.76 | 37.766 | 68.1588 | 89.6484 | 65.9633 | 0.009246 | 1.0411 | 0.02472 | 23.7445 | 0.002095 | 0.13877 | -3.4e-05 | -0.002488 | -0.002488 | -2.2697
2019-03-29 13:51 | 1553881846.5291 | -141 | 0.20951 | 0.000452 | 0 | 101.728 | 101.7011 | 125.2108 | 21.1276 | 0.001912 | 0.001796 | -151.7549 | 0.23327 | 8.4003 | 0.002833 | 0.004302 | 0.004184 | 5.8e-05 | 0.00113 | 75.0696 | 1.3539 | 2.5472 | 100.21 | 20.76 | 37.766 | 67.7125 | 89.8367 | 64.9019 | 0.009246 | 1.0444 | 0.02477 | 23.7179 | 0.002362 | 0.13877 | -9.4e-05 | -0.006931 | -0.006931 | -0.81486
2019-03-29 13:51 | 1553881847.529 | -140 | 0.2095 | 0.000448 | 0 | 101.7169 | 101.7301 | 125.481 | 21.1272 | 0.001938 | 0.001794 | -151.7641 | 0.23276 | 8.4046 | 0.002807 | 0.004513 | 0.004281 | 5.8e-05 | 0.001127 | 75.0728 | 1.3252 | 2.5471 | 100.21 | 20.75 | 37.764 | 68.3387 | 89.2087 | 64.8532 | 0.009239 | 1.0445 | 0.024798 | 23.7418 | -0.001163 | 0.13847 | -2.4e-05 | -0.001768 | -0.001768 | -3.1869
2019-03-29 13:51 | 1553881848.529 | -139 | 0.20951 | 0.000449 | -1e-06 | 101.6019 | 101.4898 | 125.8033 | 21.1274 | 0.001892 | 0.001778 | -151.7606 | 0.23255 | 8.4042 | 0.002848 | 0.003823 | 0.00362 | 5.6e-05 | 0.001124 | 75.0814 | 1.3557 | 2.5473 | 100.21 | 20.76 | 37.762 | 67.2319 | 90.2132 | 64.4684 | 0.009245 | 1.0457 | 0.024844 | 23.7587 | 0.009874 | 0.13834 | -7.2e-05 | -0.00529 | -0.00529 | -1.0645
2019-03-29 13:51 | 1553881849.5299 | -138 | 0.20949 | 0.000449 | 0 | 101.6151 | 101.625 | 126.0953 | 21.1279 | 0.001899 | 0.001802 | -151.7641 | 0.23349 | 8.4026 | 0.002797 | 0.004115 | 0.00397 | 5.2e-05 | 0.00113 | 75.0419 | 1.3302 | 2.5472 | 100.21 | 20.76 | 37.762 | 67.3885 | 88.972 | 65.0514 | 0.009245 | 1.0439 | 0.024852 | 23.8068 | -0.00087 | 0.13891 | 6.4e-05 | 0.004791 | 0.004791 | 1.18
2019-03-29 13:51 | 1553881850.5295 | -137 | 0.2095 | 0.000451 | -1e-06 | 101.6253 | 101.7939 | 126.003 | 21.1279 | 0.001995 | 0.001847 | -151.7655 | 0.23298 | 8.4015 | 0.002874 | 0.004272 | 0.004019 | 5.9e-05 | 0.001131 | 75.0764 | 1.3483 | 2.5473 | 100.21 | 20.76 | 37.761 | 69.708 | 90.8241 | 66.1471 | 0.009244 | 1.0405 | 0.024802 | 23.8366 | -0.014831 | 0.1386 | -2.9e-05 | -0.002115 | -0.002115 | -2.6678
2019-03-29 13:51 | 1553881851.5295 | -136 | 0.20952 | 0.000454 | 0 | 101.6295 | 101.6448 | 125.9098 | 21.128 | 0.001905 | 0.001845 | -151.7739 | 0.23297 | 8.4016 | 0.002911 | 0.004761 | 0.004566 | 5.6e-05 | 0.001137 | 75.1187 | 1.3871 | 2.5473 | 100.21 | 20.76 | 37.757 | 67.5321 | 91.7121 | 66.0839 | 0.009243 | 1.0407 | 0.024796 | 23.8256 | -0.001347 | 0.13859 | -0.000158 | -0.011799 | -0.011799 | -0.47807
2019-03-29 13:51 | 1553881852.5301 | -135 | 0.20951 | 0.000452 | 0 | 101.5852 | 101.7137 | 126.0474 | 21.1275 | 0.001935 | 0.001868 | -151.7707 | 0.23366 | 8.4031 | 0.002825 | 0.004565 | 0.004385 | 5.6e-05 | 0.001139 | 75.0347 | 1.3586 | 2.5473 | 100.21 | 20.76 | 37.757 | 68.2656 | 89.6466 | 66.6539 | 0.009243 | 1.039 | 0.024788 | 23.8586 | -0.011312 | 0.13901 | -0.000108 | -0.008023 | -0.008023 | -0.70519
2019-03-29 13:51 | 1553881853.5305 | -134 | 0.2095 | 0.000458 | -1e-06 | 101.5211 | 101.3738 | 126.0349 | 21.1282 | 0.001899 | 0.001811 | -151.7681 | 0.23303 | 8.405 | 0.002808 | 0.004075 | 0.003979 | 5.6e-05 | 0.001141 | 75.0764 | 1.3408 | 2.5473 | 100.21 | 20.76 | 37.756 | 67.3794 | 89.2334 | 65.2695 | 0.009243 | 1.0432 | 0.024838 | 23.8088 | 0.012989 | 0.13863 | -5.5e-05 | -0.00402 | -0.00402 | -1.4036
2019-03-29 13:51 | 1553881854.5302 | -133 | 0.20952 | 0.000454 | -1e-06 | 101.5654 | 101.5079 | 125.9838 | 21.1272 | 0.00195 | 0.001782 | -151.7606 | 0.23355 | 8.4032 | 0.00285 | 0.004054 | 0.00382 | 5.5e-05 | 0.001142 | 75.0726 | 1.3007 | 2.5471 | 100.21 | 20.76 | 37.755 | 68.6063 | 90.241 | 64.5636 | 0.009243 | 1.0454 | 0.024859 | 23.7791 | 0.005064 | 0.13894 | -0.000141 | -0.010466 | -0.010466 | -0.54029
2019-03-29 13:51 | 1553881855.5292 | -132 | 0.20952 | 0.000455 | 0 | 101.3921 | 101.4902 | 126.0403 | 21.1281 | 0.001996 | 0.001781 | -151.7741 | 0.23326 | 8.401 | 0.00283 | 0.003737 | 0.003622 | 5.7e-05 | 0.001139 | 75.0346 | 1.3491 | 2.5474 | 100.21 | 20.76 | 37.752 | 69.7176 | 89.7669 | 64.5557 | 0.009242 | 1.0454 | 0.024865 | 23.7842 | -0.008648 | 0.13877 | -0.000148 | -0.011053 | -0.011053 | -0.51098
2019-03-29 13:51 | 1553881856.5293 | -131 | 0.20949 | 0.000454 | 0 | 101.4241 | 101.4257 | 125.8948 | 21.1281 | 0.001967 | 0.001804 | -151.7736 | 0.23278 | 8.4032 | 0.002936 | 0.004449 | 0.004228 | 5.6e-05 | 0.001143 | 75.0672 | 1.3931 | 2.5474 | 100.21 | 20.76 | 37.749 | 69.022 | 92.33 | 65.0951 | 0.009241 | 1.0438 | 0.02483 | 23.7894 | -0.000136 | 0.13848 | 3e-05 | 0.002295 | 0.002295 | 2.4559
2019-03-29 13:51 | 1553881857.5297 | -130 | 0.20953 | 0.000451 | 0 | 101.3966 | 101.4827 | 125.8854 | 21.1284 | 0.001923 | 0.001832 | -151.7659 | 0.23284 | 8.4041 | 0.002863 | 0.004551 | 0.004386 | 5.5e-05 | 0.001148 | 75.1207 | 1.354 | 2.5474 | 100.21 | 20.76 | 37.748 | 67.9763 | 90.568 | 65.7738 | 0.009241 | 1.0417 | 0.024805 | 23.8124 | -0.00759 | 0.13851 | -0.000214 | -0.015964 | -0.015964 | -0.35315
2019-03-29 13:51 | 1553881858.5284 | -129 | 0.2095 | 0.000452 | -1e-06 | 101.3487 | 101.2247 | 125.6315 | 21.1279 | 0.001946 | 0.00187 | -151.7551 | 0.23263 | 8.4031 | 0.002819 | 0.004805 | 0.00449 | 5.8e-05 | 0.001148 | 75.0895 | 1.3433 | 2.5473 | 100.21 | 20.76 | 37.747 | 68.5312 | 89.5049 | 66.6909 | 0.009241 | 1.0389 | 0.024746 | 23.8205 | 0.010954 | 0.13839 | -5e-06 | -0.000361 | -0.000361 | -15.6051
2019-03-29 13:51 | 1553881859.5294 | -128 | 0.20952 | 0.000455 | 0 | 101.2724 | 101.2107 | 125.343 | 21.1282 | 0.001909 | 0.001805 | -151.7636 | 0.2323 | 8.4031 | 0.002813 | 0.003933 | 0.003781 | 5.6e-05 | 0.001154 | 75.0957 | 1.3623 | 2.5475 | 100.21 | 20.76 | 37.746 | 67.6196 | 89.3646 | 65.1273 | 0.009241 | 1.0437 | 0.024775 | 23.7383 | 0.005452 | 0.13819 | -0.000156 | -0.011591 | -0.011591 | -0.48526
2019-03-29 13:51 | 1553881860.5286 | -127 | 0.20954 | 0.000453 | 0 | 101.2014 | 101.3158 | 125.1099 | 21.1288 | 0.001949 | 0.001815 | -151.7527 | 0.23307 | 8.4028 | 0.002857 | 0.004012 | 0.003827 | 5.2e-05 | 0.001157 | 75.1045 | 1.321 | 2.5473 | 100.21 | 20.76 | 37.745 | 68.5859 | 90.4246 | 65.366 | 0.00924 | 1.0429 | 0.024743 | 23.7246 | -0.010101 | 0.13865 | -0.000246 | -0.018303 | -0.018303 | -0.30832
2019-03-29 13:51 | 1553881861.5288 | -126 | 0.20953 | 0.000455 | -1e-06 | 101.132 | 101.0704 | 124.9477 | 21.1291 | 0.001937 | 0.001819 | -151.7588 | 0.23294 | 8.402 | 0.002798 | 0.004484 | 0.004283 | 5.8e-05 | 0.001146 | 75.1081 | 1.3603 | 2.5474 | 100.21 | 20.76 | 37.745 | 68.3124 | 88.9951 | 65.4762 | 0.00924 | 1.0426 | 0.024723 | 23.7131 | 0.005449 | 0.13858 | -0.00019 | -0.014058 | -0.014058 | -0.40119
2019-03-29 13:51 | 1553881862.5285 | -125 | 0.20954 | 0.000456 | -1e-06 | 101.1021 | 100.9663 | 124.8414 | 21.1291 | 0.001936 | 0.001793 | -151.7748 | 0.23323 | 8.4039 | 0.002811 | 0.003876 | 0.003716 | 6e-05 | 0.001144 | 75.0862 | 1.3387 | 2.5474 | 100.21 | 20.76 | 37.744 | 68.2843 | 89.3272 | 64.8496 | 0.00924 | 1.0445 | 0.024735 | 23.6811 | 0.012018 | 0.13875 | -0.000255 | -0.018845 | -0.018845 | -0.29966
2019-03-29 13:51 | 1553881863.5294 | -124 | 0.20954 | 0.000455 | -1e-06 | 101.0703 | 100.9349 | 125.066 | 21.128 | 0.001892 | 0.001797 | -151.7855 | 0.23332 | 8.4028 | 0.00278 | 0.003886 | 0.003707 | 5.4e-05 | 0.001141 | 75.1313 | 1.3594 | 2.5475 | 100.21 | 20.76 | 37.744 | 67.2207 | 88.5573 | 64.9368 | 0.00924 | 1.0442 | 0.024754 | 23.7054 | 0.011993 | 0.1388 | -0.000248 | -0.018395 | -0.018395 | -0.30711
2019-03-29 13:51 | 1553881864.5285 | -123 | 0.20949 | 0.000453 | -2e-06 | 100.9692 | 100.9926 | 125.0003 | 21.129 | 0.001939 | 0.001836 | -151.791 | 0.23308 | 8.402 | 0.002774 | 0.00369 | 0.003486 | 5.6e-05 | 0.001139 | 75.0701 | 1.3853 | 2.5475 | 100.21 | 20.76 | 37.744 | 68.345 | 88.4358 | 65.8781 | 0.00924 | 1.0413 | 0.024713 | 23.7322 | -0.002076 | 0.13866 | 4e-05 | 0.003051 | 0.003051 | 1.85
2019-03-29 13:51 | 1553881865.5285 | -122 | 0.20954 | 0.000447 | -1e-06 | 100.9308 | 100.869 | 124.7087 | 21.1294 | 0.001913 | 0.00188 | -151.7801 | 0.23317 | 8.4024 | 0.002811 | 0.004371 | 0.004178 | 5.7e-05 | 0.001138 | 75.0661 | 1.3391 | 2.5473 | 100.21 | 20.76 | 37.744 | 67.7222 | 89.3195 | 66.9419 | 0.00924 | 1.0381 | 0.024646 | 23.7416 | 0.005479 | 0.13872 | -0.000256 | -0.018847 | -0.018847 | -0.29956
2019-03-29 13:51 | 1553881866.5295 | -121 | 0.20952 | 0.000451 | -1e-06 | 100.8383 | 100.6196 | 124.7818 | 21.1293 | 0.001969 | 0.001834 | -151.7685 | 0.23273 | 8.4042 | 0.002785 | 0.003721 | 0.003541 | 5.7e-05 | 0.001142 | 75.0182 | 1.33 | 2.5474 | 100.21 | 20.76 | 37.745 | 69.0785 | 88.689 | 65.8214 | 0.00924 | 1.0415 | 0.024694 | 23.7094 | 0.019428 | 0.13845 | -0.00016 | -0.011827 | -0.011827 | -0.47643
2019-03-29 13:51 | 1553881867.5296 | -120 | 0.20953 | 0.000449 | -1e-06 | 100.7653 | 100.753 | 124.4963 | 21.13 | 0.001885 | 0.001821 | -151.7649 | 0.23344 | 8.402 | 0.002745 | 0.003768 | 0.003595 | 6.2e-05 | 0.001148 | 75.1028 | 1.3343 | 2.5473 | 100.21 | 20.76 | 37.746 | 67.0441 | 87.7158 | 65.5084 | 0.009241 | 1.0425 | 0.024677 | 23.6714 | 0.001093 | 0.13887 | -0.000227 | -0.016783 | -0.016783 | -0.33678
2019-03-29 13:51 | 1553881868.5286 | -119 | 0.20953 | 0.00045 | 0 | 100.7114 | 100.4356 | 124.7386 | 21.1286 | 0.001937 | 0.001854 | -151.7597 | 0.23204 | 8.4014 | 0.002855 | 0.004738 | 0.004575 | 5.9e-05 | 0.00115 | 75.1128 | 1.3052 | 2.5473 | 100.21 | 20.77 | 37.746 | 68.3049 | 90.3823 | 66.2998 | 0.009246 | 1.0401 | 0.024672 | 23.722 | 0.024531 | 0.13804 | -0.000237 | -0.017568 | -0.017568 | -0.31981
2019-03-29 13:51 | 1553881869.5287 | -118 | 0.20948 | 0.000452 | 0 | 100.6282 | 100.5718 | 124.731 | 21.1295 | 0.001917 | 0.001814 | -151.7465 | 0.23238 | 8.4027 | 0.00277 | 0.004468 | 0.004322 | 6.2e-05 | 0.001148 | 75.0968 | 1.3506 | 2.5473 | 100.21 | 20.77 | 37.746 | 67.8339 | 88.337 | 65.3518 | 0.009246 | 1.043 | 0.024706 | 23.6882 | 0.005018 | 0.13824 | 7.1e-05 | 0.005267 | 0.005267 | 1.0682
2019-03-29 13:51 | 1553881870.5305 | -117 | 0.20949 | 0.00045 | 0 | 100.5728 | 100.3868 | 124.7625 | 21.13 | 0.001936 | 0.001828 | -151.7353 | 0.23235 | 8.4017 | 0.002751 | 0.004 | 0.003822 | 6.4e-05 | 0.001145 | 75.0859 | 1.3335 | 2.5474 | 100.21 | 20.77 | 37.745 | 68.2828 | 87.8706 | 65.684 | 0.009246 | 1.0419 | 0.024697 | 23.7028 | 0.016561 | 0.13822 | 9e-06 | 0.000686 | 0.000686 | 8.203
2019-03-29 13:51 | 1553881871.5286 | -116 | 0.2095 | 0.000451 | -1e-06 | 100.4552 | 100.1353 | 124.7083 | 21.1296 | 0.001897 | 0.001786 | -151.7297 | 0.23277 | 8.4025 | 0.00274 | 0.004311 | 0.004035 | 6.1e-05 | 0.001145 | 75.1251 | 1.3304 | 2.5474 | 100.21 | 20.77 | 37.745 | 67.3426 | 87.5975 | 64.672 | 0.009246 | 1.0451 | 0.024729 | 23.6622 | 0.028532 | 0.13848 | -1.9e-05 | -0.00135 | -0.00135 | -4.1751
2019-03-29 13:51 | 1553881872.5297 | -115 | 0.20953 | 0.000445 | 0 | 100.3765 | 100.2049 | 124.816 | 21.1298 | 0.001948 | 0.001753 | -151.7196 | 0.23288 | 8.4015 | 0.002823 | 0.004451 | 0.004289 | 5.9e-05 | 0.001147 | 75.1013 | 1.3416 | 2.5475 | 100.21 | 20.77 | 37.743 | 68.5799 | 89.6052 | 63.8814 | 0.009246 | 1.0475 | 0.024768 | 23.6447 | 0.015303 | 0.13854 | -0.000178 | -0.013239 | -0.013239 | -0.42591
2019-03-29 13:51 | 1553881873.5289 | -114 | 0.20953 | 0.000449 | 0 | 100.3124 | 100.0636 | 124.5717 | 21.1293 | 0.001967 | 0.001783 | -151.712 | 0.23276 | 8.4017 | 0.002736 | 0.003741 | 0.003635 | 5.7e-05 | 0.001146 | 74.9895 | 1.3544 | 2.5474 | 100.21 | 20.77 | 37.743 | 69.0263 | 87.5168 | 64.6084 | 0.009246 | 1.0453 | 0.024717 | 23.647 | 0.022214 | 0.13847 | -0.000208 | -0.015434 | -0.015434 | -0.36515
2019-03-29 13:51 | 1553881874.5278 | -113 | 0.20953 | 0.000446 | 0 | 100.1936 | 99.8919 | 124.7095 | 21.1284 | 0.001984 | 0.001794 | -151.7092 | 0.23241 | 8.4021 | 0.002799 | 0.004112 | 0.003929 | 5.9e-05 | 0.00115 | 75.0593 | 1.3393 | 2.5474 | 100.21 | 20.77 | 37.743 | 69.4254 | 89.0366 | 64.8722 | 0.009246 | 1.0444 | 0.024721 | 23.6693 | 0.02697 | 0.13826 | -0.000219 | -0.016233 | -0.016233 | -0.34664
2019-03-29 13:51 | 1553881875.5288 | -112 | 0.20953 | 0.000449 | -1e-06 | 100.0768 | 99.7704 | 124.7773 | 21.1285 | 0.001947 | 0.00188 | -151.7289 | 0.23255 | 8.4033 | 0.002752 | 0.00385 | 0.003622 | 5.9e-05 | 0.001147 | 75.0249 | 1.369 | 2.5474 | 100.21 | 20.77 | 37.743 | 68.5505 | 87.8885 | 66.9459 | 0.009246 | 1.0381 | 0.024653 | 23.7483 | 0.027428 | 0.13834 | -0.000205 | -0.015141 | -0.015141 | -0.37189
2019-03-29 13:51 | 1553881876.5289 | -111 | 0.20954 | 0.000448 | -1e-06 | 100.0739 | 99.782 | 125.0545 | 21.1287 | 0.001936 | 0.001815 | -151.7508 | 0.23225 | 8.4033 | 0.002689 | 0.003971 | 0.003813 | 5.7e-05 | 0.001152 | 75.0091 | 1.3717 | 2.5475 | 100.21 | 20.77 | 37.744 | 68.2918 | 86.3732 | 65.3823 | 0.009246 | 1.0429 | 0.024737 | 23.7199 | 0.026121 | 0.13817 | -0.000257 | -0.019045 | -0.019045 | -0.29526
2019-03-29 13:51 | 1553881877.5289 | -110 | 0.20953 | 0.00045 | -1e-06 | 100.0116 | 99.7967 | 125.4345 | 21.128 | 0.001912 | 0.00185 | -151.7524 | 0.23416 | 8.4013 | 0.002719 | 0.004057 | 0.003947 | 5.8e-05 | 0.001153 | 74.9998 | 1.3526 | 2.5473 | 100.21 | 20.77 | 37.744 | 67.7084 | 87.0908 | 66.2228 | 0.009246 | 1.0403 | 0.024744 | 23.7854 | 0.019238 | 0.1393 | -0.000177 | -0.013102 | -0.013102 | -0.43275
2019-03-29 13:51 | 1553881878.5292 | -109 | 0.20951 | 0.000456 | -1e-06 | 99.8068 | 99.5578 | 125.3099 | 21.1281 | 0.001876 | 0.001818 | -151.7644 | 0.23314 | 8.4022 | 0.002714 | 0.003903 | 0.003617 | 6e-05 | 0.001148 | 75.0007 | 1.321 | 2.5472 | 100.21 | 20.77 | 37.742 | 66.8291 | 86.9871 | 65.4512 | 0.009245 | 1.0427 | 0.02476 | 23.7466 | 0.022347 | 0.1387 | -0.000108 | -0.007952 | -0.007952 | -0.70986
2019-03-29 13:51 | 1553881879.5289 | -108 | 0.20952 | 0.00045 | -1e-06 | 99.7048 | 99.5066 | 125.3744 | 21.1281 | 0.001907 | 0.001768 | -151.7654 | 0.23368 | 8.4034 | 0.002659 | 0.004257 | 0.004036 | 6.1e-05 | 0.001146 | 75.0152 | 1.3897 | 2.5475 | 100.21 | 20.77 | 37.742 | 67.5798 | 85.6586 | 64.2497 | 0.009245 | 1.0464 | 0.02481 | 23.7105 | 0.017795 | 0.13901 | -0.00017 | -0.01261 | -0.01261 | -0.44868
2019-03-29 13:51 | 1553881880.529 | -107 | 0.20954 | 0.000447 | -1e-06 | 99.6159 | 99.0742 | 125.4357 | 21.1285 | 0.001965 | 0.00177 | -151.7955 | 0.23404 | 8.4003 | 0.002661 | 0.004426 | 0.004225 | 5.9e-05 | 0.001145 | 75.104 | 1.3617 | 2.5475 | 100.21 | 20.77 | 37.744 | 68.9716 | 85.7125 | 64.282 | 0.009246 | 1.0463 | 0.024815 | 23.7174 | 0.04877 | 0.13923 | -0.000258 | -0.019152 | -0.019152 | -0.29587
2019-03-29 13:51 | 1553881881.5289 | -106 | 0.20955 | 0.000449 | -1e-06 | 99.5731 | 99.1909 | 125.4899 | 21.1288 | 0.001956 | 0.001838 | -151.7881 | 0.23353 | 8.4027 | 0.002676 | 0.004297 | 0.004046 | 6.2e-05 | 0.00114 | 75.0363 | 1.3225 | 2.5474 | 100.21 | 20.77 | 37.744 | 68.752 | 86.054 | 65.9147 | 0.009246 | 1.0412 | 0.02476 | 23.7799 | 0.034394 | 0.13893 | -0.000299 | -0.022219 | -0.022219 | -0.25448
2019-03-29 13:51 | 1553881882.5285 | -105 | 0.20952 | 0.000448 | -1e-06 | 99.348 | 99.1943 | 125.2864 | 21.1301 | 0.001985 | 0.001846 | -151.7798 | 0.23314 | 8.4029 | 0.00262 | 0.003805 | 0.003696 | 6e-05 | 0.001137 | 75.0629 | 1.3399 | 2.5473 | 100.21 | 20.77 | 37.745 | 69.4497 | 84.7083 | 66.1257 | 0.009246 | 1.0406 | 0.024733 | 23.768 | 0.013843 | 0.1387 | -0.000159 | -0.011725 | -0.011725 | -0.48144
2019-03-29 13:51 | 1553881883.5281 | -104 | 0.20952 | 0.000454 | -1e-06 | 99.2501 | 98.9964 | 125.2897 | 21.1297 | 0.001961 | 0.00185 | -151.7826 | 0.23394 | 8.4039 | 0.002657 | 0.004133 | 0.003929 | 5.8e-05 | 0.001133 | 75.1166 | 1.3478 | 2.5471 | 100.21 | 20.77 | 37.745 | 68.8777 | 85.6132 | 66.2237 | 0.009246 | 1.0403 | 0.024729 | 23.7717 | 0.022893 | 0.13917 | -0.000143 | -0.01055 | -0.01055 | -0.53688
2019-03-29 13:51 | 1553881884.5281 | -103 | 0.20955 | 0.000446 | -1e-06 | 99.1136 | 98.7333 | 125.358 | 21.1294 | 0.001967 | 0.00184 | -151.7801 | 0.23223 | 8.4025 | 0.00262 | 0.004982 | 0.004865 | 6.1e-05 | 0.001128 | 75.1487 | 1.3296 | 2.5472 | 100.21 | 20.77 | 37.745 | 69.0192 | 84.7092 | 65.9814 | 0.009246 | 1.041 | 0.024745 | 23.7697 | 0.03439 | 0.13816 | -0.000317 | -0.023537 | -0.023537 | -0.23889
2019-03-29 13:51 | 1553881885.5291 | -102 | 0.20949 | 0.000449 | -1e-06 | 98.9579 | 98.5947 | 125.5786 | 21.1298 | 0.00194 | 0.001827 | -151.7741 | 0.23383 | 8.4016 | 0.002636 | 0.004647 | 0.004439 | 6e-05 | 0.001132 | 75.1178 | 1.3152 | 2.5472 | 100.21 | 20.77 | 37.745 | 68.3825 | 85.0983 | 65.6611 | 0.009246 | 1.042 | 0.024778 | 23.7794 | 0.032894 | 0.13911 | 1.4e-05 | 0.001044 | 0.001044 | 5.423
2019-03-29 13:51 | 1553881886.5296 | -101 | 0.2095 | 0.000453 | -1e-06 | 98.8376 | 98.4845 | 125.4383 | 21.129 | 0.001933 | 0.001825 | -151.7751 | 0.23324 | 8.404 | 0.002627 | 0.003818 | 0.003654 | 6.2e-05 | 0.001129 | 75.0421 | 1.3262 | 2.5474 | 100.21 | 20.77 | 37.746 | 68.2046 | 84.8907 | 65.601 | 0.009246 | 1.0422 | 0.024767 | 23.764 | 0.032017 | 0.13875 | -4.2e-05 | -0.003073 | -0.003073 | -1.8376
2019-03-29 13:51 | 1553881887.5293 | -100 | 0.20952 | 0.000452 | -1e-06 | 98.7583 | 98.4097 | 125.2726 | 21.1295 | 0.001929 | 0.001807 | -151.779 | 0.23302 | 8.4024 | 0.002634 | 0.003876 | 0.003671 | 6.2e-05 | 0.001133 | 75.071 | 1.3637 | 2.5473 | 100.21 | 20.77 | 37.747 | 68.1154 | 85.0667 | 65.1691 | 0.009247 | 1.0435 | 0.024766 | 23.7331 | 0.031626 | 0.13862 | -0.000131 | -0.009703 | -0.009703 | -0.58146
2019-03-29 13:51 | 1553881888.5273 | -99 | 0.2095 | 0.000455 | -1e-06 | 98.5012 | 98.0125 | 125.3575 | 21.1297 | 0.001947 | 0.001826 | -151.7743 | 0.23323 | 8.4015 | 0.002639 | 0.004046 | 0.003906 | 6.1e-05 | 0.001136 | 75.0823 | 1.3459 | 2.5474 | 100.21 | 20.78 | 37.749 | 68.5438 | 85.1809 | 65.6416 | 0.009253 | 1.0421 | 0.024757 | 23.7578 | 0.044488 | 0.13875 | -1.5e-05 | -0.001065 | -0.001065 | -5.3041
2019-03-29 13:51 | 1553881889.5276 | -98 | 0.2095 | 0.000453 | -1e-06 | 98.429 | 98.0377 | 125.5516 | 21.1305 | 0.001926 | 0.001832 | -151.7715 | 0.23345 | 8.4017 | 0.002633 | 0.004722 | 0.004481 | 6.2e-05 | 0.001138 | 75.0713 | 1.3445 | 2.5475 | 100.21 | 20.77 | 37.751 | 68.0323 | 85.035 | 65.7821 | 0.009248 | 1.0416 | 0.024771 | 23.7811 | 0.035631 | 0.13888 | -5.3e-05 | -0.003908 | -0.003908 | -1.4465
2019-03-29 13:51 | 1553881890.5283 | -97 | 0.20951 | 0.000449 | -1e-06 | 98.2578 | 97.9418 | 125.4521 | 21.1299 | 0.001945 | 0.001822 | -151.7556 | 0.23186 | 8.4032 | 0.002723 | 0.004445 | 0.004332 | 6.4e-05 | 0.001143 | 75.1525 | 1.3547 | 2.5474 | 100.21 | 20.77 | 37.754 | 68.4911 | 87.2042 | 65.5357 | 0.009248 | 1.0424 | 0.024771 | 23.763 | 0.028812 | 0.13793 | -7.8e-05 | -0.005752 | -0.005752 | -0.976
2019-03-29 13:52 | 1553881891.5274 | -96 | 0.20954 | 0.000452 | -1e-06 | 98.0767 | 97.5234 | 125.6103 | 21.13 | 0.001909 | 0.00184 | -151.7582 | 0.23332 | 8.4015 | 0.002728 | 0.004563 | 0.004389 | 6.4e-05 | 0.001142 | 75.1822 | 1.3245 | 2.5473 | 100.21 | 20.78 | 37.758 | 67.6214 | 87.3115 | 65.9767 | 0.009255 | 1.041 | 0.02477 | 23.7935 | 0.050606 | 0.1388 | -0.00026 | -0.019249 | -0.019249 | -0.29348
2019-03-29 13:52 | 1553881892.5274 | -95 | 0.20956 | 0.000447 | -1e-06 | 97.9959 | 97.3469 | 125.9828 | 21.13 | 0.001945 | 0.001825 | -151.7567 | 0.23362 | 8.4027 | 0.002692 | 0.003938 | 0.003799 | 6.4e-05 | 0.001142 | 75.1769 | 1.3402 | 2.5475 | 100.21 | 20.78 | 37.763 | 68.5027 | 86.444 | 65.6176 | 0.009256 | 1.0421 | 0.02482 | 23.8161 | 0.05943 | 0.13898 | -0.000397 | -0.029541 | -0.029541 | -0.19148
2019-03-29 13:52 | 1553881893.5287 | -94 | 0.20954 | 0.000446 | -1e-06 | 97.845 | 97.1801 | 125.9427 | 21.1303 | 0.001941 | 0.001816 | -151.7536 | 0.23386 | 8.4037 | 0.002689 | 0.003989 | 0.003788 | 6.7e-05 | 0.001138 | 75.146 | 1.3339 | 2.5474 | 100.21 | 20.78 | 37.767 | 68.3908 | 86.3733 | 65.3939 | 0.009257 | 1.0428 | 0.024824 | 23.8044 | 0.060997 | 0.13913 | -0.000272 | -0.020184 | -0.020184 | -0.28054
2019-03-29 13:52 | 1553881894.5278 | -93 | 0.20953 | 0.000449 | -1e-06 | 97.6229 | 96.8664 | 125.884 | 21.1303 | 0.001916 | 0.00183 | -151.7509 | 0.23315 | 8.4009 | 0.002749 | 0.004609 | 0.004377 | 6.6e-05 | 0.001138 | 75.1677 | 1.327 | 2.5474 | 100.21 | 20.78 | 37.772 | 67.7926 | 87.8347 | 65.7225 | 0.009259 | 1.0418 | 0.024806 | 23.8104 | 0.069585 | 0.1387 | -0.000218 | -0.016192 | -0.016192 | -0.34863
2019-03-29 13:52 | 1553881895.5285 | -92 | 0.20952 | 0.000448 | -1e-06 | 97.3661 | 96.8988 | 125.7584 | 21.1296 | 0.001923 | 0.001809 | -151.7513 | 0.23356 | 8.4018 | 0.00271 | 0.003814 | 0.003741 | 6.6e-05 | 0.001135 | 75.2278 | 1.3626 | 2.5474 | 100.21 | 20.78 | 37.785 | 67.9751 | 86.8957 | 65.2371 | 0.009262 | 1.0433 | 0.024812 | 23.7815 | 0.043029 | 0.13895 | -0.000132 | -0.00974 | -0.00974 | -0.58061
2019-03-29 13:52 | 1553881896.5285 | -91 | 0.20952 | 0.000447 | -1e-06 | 97.3044 | 96.6087 | 125.6481 | 21.1303 | 0.001928 | 0.001789 | -151.7487 | 0.23315 | 8.4018 | 0.002683 | 0.003781 | 0.003618 | 6.5e-05 | 0.001138 | 75.2237 | 1.344 | 2.5475 | 100.21 | 20.78 | 37.791 | 68.0943 | 86.239 | 64.753 | 0.009263 | 1.0448 | 0.024819 | 23.7541 | 0.064181 | 0.1387 | -0.000111 | -0.008243 | -0.008243 | -0.68483
2019-03-29 13:52 | 1553881897.528 | -90 | 0.20953 | 0.000451 | -1e-06 | 97.115 | 96.4801 | 125.6185 | 21.1303 | 0.001899 | 0.00183 | -151.7397 | 0.23327 | 8.4025 | 0.002678 | 0.003435 | 0.003238 | 6.7e-05 | 0.001138 | 75.1505 | 1.3331 | 2.5474 | 100.21 | 20.78 | 37.807 | 67.4033 | 86.1226 | 65.7264 | 0.009267 | 1.0418 | 0.02478 | 23.7855 | 0.058667 | 0.13877 | -0.000178 | -0.013167 | -0.013167 | -0.42896
2019-03-29 13:52 | 1553881898.5276 | -89 | 0.20954 | 0.000446 | -1e-06 | 96.9629 | 96.3838 | 125.6859 | 21.1307 | 0.001925 | 0.001777 | -151.7376 | 0.23354 | 8.404 | 0.002769 | 0.004453 | 0.004268 | 6.3e-05 | 0.001141 | 75.0952 | 1.3458 | 2.5474 | 100.21 | 20.78 | 37.815 | 68.0089 | 88.3094 | 64.4609 | 0.009269 | 1.0457 | 0.024833 | 23.7474 | 0.053578 | 0.13893 | -0.000259 | -0.019278 | -0.019278 | -0.29332
2019-03-29 13:52 | 1553881899.5276 | -88 | 0.20952 | 0.000448 | -1e-06 | 96.789 | 96.1255 | 125.6685 | 21.1302 | 0.00194 | 0.001785 | -151.7398 | 0.23297 | 8.4033 | 0.002748 | 0.00414 | 0.004008 | 6.5e-05 | 0.001145 | 65.4701 | 1.3286 | 2.5472 | 100.21 | 20.78 | 37.824 | 68.3757 | 87.798 | 64.6593 | 0.009271 | 1.0451 | 0.024824 | 23.7527 | 0.061529 | 0.13859 | -0.000137 | -0.010109 | -0.010109 | -0.55801
2019-03-29 13:52 | 1553881900.529 | -87 | 0.20952 | 0.000448 | -1e-06 | 96.6276 | 95.8944 | 125.8737 | 21.1299 | 0.001897 | 0.001877 | -151.7475 | 0.23328 | 8.3999 | 0.002785 | 0.003907 | 0.003687 | 6.6e-05 | 0.001148 | 54.6493 | 1.3024 | 2.5473 | 100.21 | 20.78 | 37.851 | 67.3397 | 88.6966 | 66.8552 | 0.009278 | 1.0384 | 0.024764 | 23.8492 | 0.068132 | 0.13878 | -0.000151 | -0.011161 | -0.011161 | -0.50609
2019-03-29 13:52 | 1553881901.5291 | -86 | 0.20953 | 0.000445 | -1e-06 | 96.4388 | 95.7815 | 125.8331 | 21.1304 | 0.001946 | 0.001821 | -151.7506 | 0.23351 | 8.403 | 0.002759 | 0.004176 | 0.00401 | 6.3e-05 | 0.001149 | 40.2825 | 1.3411 | 2.5474 | 100.21 | 20.78 | 37.851 | 68.5218 | 88.0529 | 65.5095 | 0.009278 | 1.0425 | 0.024809 | 23.7981 | 0.061174 | 0.13891 | -0.000185 | -0.013734 | -0.013734 | -0.41166
2019-03-29 13:52 | 1553881902.5277 | -85 | 0.20953 | 0.000444 | -1e-06 | 96.2702 | 95.619 | 125.5343 | 21.13 | 0.001972 | 0.001818 | -151.7587 | 0.23215 | 8.4017 | 0.002827 | 0.003639 | 0.003533 | 6.4e-05 | 0.001149 | 17.8728 | 1.3406 | 2.5474 | 100.21 | 20.78 | 37.861 | 69.1563 | 89.6946 | 65.4507 | 0.00928 | 1.0427 | 0.024782 | 23.7678 | 0.060713 | 0.13811 | -0.000184 | -0.01368 | -0.01368 | -0.4109
2019-03-29 13:52 | 1553881903.5267 | -84 | 0.20953 | 0.000445 | -1e-06 | 96.0496 | 95.4296 | 125.5366 | 21.1297 | 0.001877 | 0.001781 | -151.7708 | 0.23319 | 8.404 | 0.00286 | 0.004292 | 0.004096 | 6.5e-05 | 0.001143 | 5.8877 | 1.3392 | 2.5474 | 100.21 | 20.78 | 37.882 | 66.855 | 90.5021 | 64.5621 | 0.009285 | 1.0454 | 0.024815 | 23.7368 | 0.057929 | 0.13872 | -0.000212 | -0.015727 | -0.015727 | -0.35901
2019-03-29 13:52 | 1553881904.5281 | -83 | 0.20953 | 0.000445 | -1e-06 | 95.8084 | 94.9579 | 125.4925 | 21.13 | 0.001967 | 0.001775 | -151.7732 | 0.23227 | 8.4027 | 0.002759 | 0.003667 | 0.003515 | 6.4e-05 | 0.001138 | 1.2492 | 1.3336 | 2.5476 | 100.21 | 20.78 | 37.893 | 69.0302 | 88.0619 | 64.4094 | 0.009288 | 1.0459 | 0.024816 | 23.7273 | 0.079752 | 0.13818 | -0.000195 | -0.014453 | -0.014453 | -0.3891
2019-03-29 13:52 | 1553881905.5268 | -82 | 0.20952 | 0.000445 | -1e-06 | 95.6234 | 94.734 | 125.4831 | 21.1295 | 0.001955 | 0.001835 | -151.7501 | 0.23274 | 8.4009 | 0.002782 | 0.004469 | 0.004297 | 6.5e-05 | 0.001136 | 0.74982 | 1.3659 | 2.5474 | 100.21 | 20.78 | 37.917 | 68.7472 | 88.6197 | 65.8622 | 0.009294 | 1.0414 | 0.024762 | 23.7774 | 0.083579 | 0.13846 | -0.000147 | -0.010914 | -0.010914 | -0.51631
2019-03-29 13:52 | 1553881906.5278 | -81 | 0.20952 | 0.000443 | -1e-06 | 95.3881 | 94.59 | 125.7361 | 21.1307 | 0.001949 | 0.001842 | -151.7443 | 0.23313 | 8.4021 | 0.002785 | 0.004224 | 0.004 | 6.3e-05 | 0.001135 | 0.67146 | 1.3646 | 2.5475 | 100.21 | 20.79 | 37.929 | 68.5854 | 88.689 | 66.0335 | 0.009303 | 1.0409 | 0.02478 | 23.8074 | 0.075153 | 0.13869 | -0.000129 | -0.009553 | -0.009553 | -0.59089
2019-03-29 13:52 | 1553881907.5285 | -80 | 0.20953 | 0.000444 | -1e-06 | 95.1627 | 94.1781 | 125.4998 | 21.1299 | 0.001951 | 0.001784 | -151.7622 | 0.23364 | 8.4009 | 0.002765 | 0.004393 | 0.004246 | 6.3e-05 | 0.001132 | 0.48642 | 1.3532 | 2.5474 | 100.21 | 20.79 | 37.94 | 68.6471 | 88.2007 | 64.6333 | 0.009305 | 1.0452 | 0.024808 | 23.7358 | 0.09303 | 0.13899 | -0.000173 | -0.012851 | -0.012851 | -0.4402
2019-03-29 13:52 | 1553881908.5273 | -79 | 0.20952 | 0.000442 | -1e-06 | 94.826 | 93.9975 | 125.5484 | 21.1309 | 0.001888 | 0.001773 | -151.7878 | 0.23319 | 8.4016 | 0.002757 | 0.004599 | 0.004488 | 6.1e-05 | 0.001134 | 0.40767 | 1.3609 | 2.5475 | 100.21 | 20.78 | 37.962 | 67.1402 | 88.012 | 64.3525 | 0.009305 | 1.0461 | 0.024823 | 23.7306 | 0.078486 | 0.13872 | -0.000155 | -0.011433 | -0.011433 | -0.49384
2019-03-29 13:52 | 1553881909.5269 | -78 | 0.20957 | 0.000447 | -1e-06 | 94.5985 | 93.6402 | 125.431 | 21.1307 | 0.001935 | 0.001808 | -151.7977 | 0.23302 | 8.4021 | 0.002801 | 0.004617 | 0.004441 | 6.3e-05 | 0.001134 | 0.39511 | 1.3797 | 2.5475 | 100.21 | 20.78 | 37.971 | 68.2493 | 89.0778 | 65.1992 | 0.009307 | 1.0434 | 0.024781 | 23.7492 | 0.091069 | 0.13862 | -0.000447 | -0.033239 | -0.033239 | -0.16974
2019-03-29 13:52 | 1553881910.528 | -77 | 0.20958 | 0.000442 | -1e-06 | 94.3692 | 93.3713 | 125.3592 | 21.1319 | 0.001957 | 0.001827 | -151.7924 | 0.23318 | 8.4021 | 0.002749 | 0.004318 | 0.004112 | 6.2e-05 | 0.001132 | 0.37266 | 1.3813 | 2.5475 | 100.21 | 20.79 | 37.988 | 68.7995 | 87.8185 | 65.657 | 0.009317 | 1.042 | 0.024757 | 23.7584 | 0.095089 | 0.13872 | -0.000509 | -0.037785 | -0.037785 | -0.14942
2019-03-29 13:52 | 1553881911.528 | -76 | 0.20952 | 0.000447 | -1e-06 | 94.0345 | 93.0767 | 125.3664 | 21.1314 | 0.001937 | 0.001842 | -151.798 | 0.23336 | 8.4017 | 0.002715 | 0.003787 | 0.003686 | 6.1e-05 | 0.001127 | 0.38361 | 1.3409 | 2.5475 | 100.21 | 20.78 | 37.996 | 68.317 | 87.0154 | 66.0195 | 0.009313 | 1.0409 | 0.024744 | 23.7718 | 0.091572 | 0.13883 | -0.000135 | -0.009967 | -0.009967 | -0.56693
2019-03-29 13:52 | 1553881912.5269 | -75 | 0.20953 | 0.000442 | -1e-06 | 93.7253 | 92.5962 | 125.726 | 21.1306 | 0.001935 | 0.001854 | -151.7969 | 0.23296 | 8.4029 | 0.002703 | 0.003941 | 0.003712 | 6.2e-05 | 0.001125 | 0.36801 | 1.3649 | 2.5474 | 100.21 | 20.78 | 38.005 | 68.2646 | 86.7155 | 66.2996 | 0.009316 | 1.0401 | 0.02477 | 23.8157 | 0.10841 | 0.13859 | -0.00016 | -0.01179 | -0.01179 | -0.47841
2019-03-29 13:52 | 1553881913.5287 | -74 | 0.20954 | 0.000446 | -1e-06 | 93.3482 | 92.2638 | 125.9969 | 21.1309 | 0.001979 | 0.001826 | -151.7766 | 0.23347 | 8.4023 | 0.00269 | 0.003829 | 0.003659 | 6.2e-05 | 0.001127 | 0.35889 | 1.3448 | 2.5474 | 100.21 | 20.78 | 38.018 | 69.3107 | 86.4066 | 65.6483 | 0.009319 | 1.0421 | 0.02482 | 23.8185 | 0.10451 | 0.13889 | -0.000238 | -0.017669 | -0.017669 | -0.31994
2019-03-29 13:52 | 1553881914.5281 | -73 | 0.20954 | 0.000446 | -1e-06 | 92.9869 | 91.7717 | 126.0422 | 21.1313 | 0.00192 | 0.001847 | -151.781 | 0.23279 | 8.4046 | 0.002815 | 0.003807 | 0.003599 | 6.1e-05 | 0.001126 | 0.35723 | 1.3387 | 2.5473 | 100.21 | 20.78 | 38.022 | 67.9042 | 89.4021 | 66.1332 | 0.00932 | 1.0406 | 0.024807 | 23.8398 | 0.11767 | 0.13849 | -0.000235 | -0.017409 | -0.017409 | -0.32377
2019-03-29 13:52 | 1553881915.527 | -72 | 0.2095 | 0.000443 | 0 | 92.6274 | 91.4488 | 126.2793 | 21.1308 | 0.001963 | 0.001847 | -151.797 | 0.23331 | 8.401 | 0.002746 | 0.004436 | 0.004271 | 6.3e-05 | 0.001129 | 0.36121 | 1.3524 | 2.5473 | 100.21 | 20.78 | 38.03 | 68.9419 | 87.7628 | 66.1437 | 0.009322 | 1.0405 | 0.02483 | 23.8626 | 0.11454 | 0.13879 | -3e-05 | -0.002185 | -0.002185 | -2.5855
2019-03-29 13:52 | 1553881916.5285 | -71 | 0.20953 | 0.000445 | -1e-06 | 92.1853 | 91.0939 | 125.9429 | 21.1313 | 0.001937 | 0.001801 | -151.7944 | 0.23243 | 8.4031 | 0.002793 | 0.004438 | 0.004277 | 6.7e-05 | 0.001129 | 0.34634 | 1.3357 | 2.5475 | 100.21 | 20.78 | 38.03 | 68.3051 | 88.8723 | 65.0429 | 0.009322 | 1.0439 | 0.024837 | 23.7921 | 0.10653 | 0.13827 | -0.000198 | -0.014678 | -0.014678 | -0.38341
2019-03-29 13:52 | 1553881917.5281 | -70 | 0.20952 | 0.000444 | -1e-06 | 91.894 | 90.3894 | 126.3732 | 21.1316 | 0.001946 | 0.001805 | -151.7802 | 0.23353 | 8.403 | 0.002736 | 0.00362 | 0.003425 | 6.4e-05 | 0.001132 | 0.35513 | 1.3532 | 2.5474 | 100.21 | 20.78 | 38.031 | 68.516 | 87.5065 | 65.1322 | 0.009322 | 1.0436 | 0.024876 | 23.8359 | 0.14766 | 0.13892 | -0.000157 | -0.011696 | -0.011696 | -0.48342
2019-03-29 13:52 | 1553881918.5271 | -69 | 0.20951 | 0.000445 | -1e-06 | 91.4468 | 90.0024 | 126.4006 | 21.1307 | 0.00192 | 0.001825 | -151.7683 | 0.23299 | 8.403 | 0.002727 | 0.003823 | 0.003673 | 6.3e-05 | 0.001134 | 0.34827 | 1.4117 | 2.5474 | 100.21 | 20.78 | 38.031 | 67.8942 | 87.2854 | 65.6152 | 0.009322 | 1.0422 | 0.024861 | 23.8555 | 0.14241 | 0.1386 | -6.8e-05 | -0.00499 | -0.00499 | -1.1304
2019-03-29 13:52 | 1553881919.5276 | -68 | 0.20953 | 0.000449 | -1e-06 | 91.0086 | 89.5888 | 126.6133 | 21.1307 | 0.001954 | 0.001777 | -151.7663 | 0.23416 | 8.4015 | 0.002696 | 0.004041 | 0.003861 | 6.6e-05 | 0.001139 | 0.33516 | 1.394 | 2.5475 | 100.21 | 20.78 | 38.032 | 68.715 | 86.5369 | 64.4678 | 0.009322 | 1.0457 | 0.024924 | 23.8351 | 0.14064 | 0.1393 | -0.000186 | -0.013874 | -0.013874 | -0.40866
2019-03-29 13:52 | 1553881920.5267 | -67 | 0.20951 | 0.000448 | -1e-06 | 90.6482 | 89.2096 | 126.9241 | 21.1315 | 0.001973 | 0.001838 | -151.7706 | 0.23386 | 8.4006 | 0.002767 | 0.003925 | 0.003743 | 6.3e-05 | 0.001149 | 0.33212 | 1.3826 | 2.5474 | 100.21 | 20.78 | 38.032 | 69.1679 | 88.2492 | 65.9289 | 0.009322 | 1.0412 | 0.024901 | 23.9159 | 0.14309 | 0.13913 | -9.5e-05 | -0.007029 | -0.007029 | -0.80555
2019-03-29 13:52 | 1553881921.5273 | -66 | 0.20953 | 0.00045 | -1e-06 | 90.4758 | 89.0523 | 126.9701 | 21.1304 | 0.002005 | 0.001839 | -151.7574 | 0.23381 | 8.4028 | 0.002744 | 0.004647 | 0.004472 | 6.3e-05 | 0.001149 | 0.31564 | 1.355 | 2.5473 | 100.21 | 20.78 | 38.036 | 69.9461 | 87.703 | 65.9529 | 0.009323 | 1.0411 | 0.024905 | 23.9211 | 0.14185 | 0.13909 | -0.000225 | -0.016741 | -0.016741 | -0.33815
2019-03-29 13:52 | 1553881922.5263 | -65 | 0.20953 | 0.000449 | -1e-06 | 90.3531 | 88.9052 | 126.8783 | 21.1316 | 0.001932 | 0.001827 | -151.7472 | 0.23328 | 8.4015 | 0.002706 | 0.0035 | 0.003396 | 6.5e-05 | 0.001151 | 0.33196 | 1.3134 | 2.5474 | 100.21 | 20.78 | 38.036 | 68.1783 | 86.7899 | 65.6528 | 0.009323 | 1.042 | 0.024907 | 23.9018 | 0.1445 | 0.13878 | -0.000177 | -0.013198 | -0.013198 | -0.42797
2019-03-29 13:52 | 1553881923.5278 | -64 | 0.20951 | 0.000452 | -1e-06 | 90.5189 | 89.0618 | 126.963 | 21.1314 | 0.001879 | 0.001816 | -151.7472 | 0.23276 | 8.4027 | 0.002699 | 0.003855 | 0.003638 | 6.4e-05 | 0.001158 | 0.33035 | 1.333 | 2.5475 | 100.21 | 20.78 | 38.039 | 66.9223 | 86.6154 | 65.4075 | 0.009324 | 1.0428 | 0.024924 | 23.9011 | 0.14515 | 0.13847 | -8e-05 | -0.005928 | -0.005928 | -0.95061
2019-03-29 13:52 | 1553881924.5264 | -63 | 0.20953 | 0.000448 | -1e-06 | 90.8625 | 89.399 | 127.2379 | 21.1312 | 0.001934 | 0.00181 | -151.7468 | 0.23345 | 8.4024 | 0.002693 | 0.003544 | 0.003367 | 6.7e-05 | 0.001153 | 0.35419 | 1.3627 | 2.5475 | 100.21 | 20.78 | 38.039 | 68.231 | 86.4716 | 65.2425 | 0.009324 | 1.0433 | 0.024957 | 23.9212 | 0.14524 | 0.13888 | -0.000212 | -0.015865 | -0.015865 | -0.35629
2019-03-29 13:52 | 1553881925.5274 | -62 | 0.20954 | 0.000449 | 0 | 91.1837 | 89.8953 | 127.2818 | 21.1316 | 0.001957 | 0.0018 | -151.7456 | 0.23319 | 8.3996 | 0.002674 | 0.003934 | 0.003763 | 6.7e-05 | 0.001154 | 0.34108 | 1.3566 | 2.5475 | 100.21 | 20.78 | 38.044 | 68.7937 | 86.0257 | 65.0206 | 0.009325 | 1.044 | 0.024969 | 23.9175 | 0.12729 | 0.13873 | -0.000242 | -0.018128 | -0.018128 | -0.31145
2019-03-29 13:52 | 1553881926.5278 | -61 | 0.20953 | 0.000447 | -1e-06 | 91.6061 | 90.4099 | 127.4818 | 21.1317 | 0.001969 | 0.001794 | -151.7508 | 0.23364 | 8.4027 | 0.002677 | 0.003957 | 0.003731 | 6.6e-05 | 0.001153 | 0.33085 | 1.2916 | 2.5473 | 100.21 | 20.78 | 38.044 | 69.0777 | 86.1007 | 64.8598 | 0.009325 | 1.0445 | 0.024995 | 23.9305 | 0.11756 | 0.13899 | -0.000195 | -0.014628 | -0.014628 | -0.38674
2019-03-29 13:52 | 1553881927.5265 | -60 | 0.20947 | 0.000452 | -1e-06 | 92.0266 | 90.7526 | 127.5865 | 21.1314 | 0.001972 | 0.001825 | -151.7597 | 0.23265 | 8.4021 | 0.002719 | 0.003978 | 0.003819 | 6.8e-05 | 0.00115 | 0.34971 | 1.3672 | 2.5475 | 100.21 | 20.78 | 38.046 | 69.1552 | 87.1097 | 65.6144 | 0.009326 | 1.0422 | 0.024977 | 23.9671 | 0.1247 | 0.1384 | 0.00017 | 0.012808 | 0.012808 | 0.43981
2019-03-29 13:52 | 1553881928.5268 | -59 | 0.20952 | 0.000454 | -1e-06 | 92.4595 | 91.1983 | 127.34 | 21.1316 | 0.001965 | 0.001788 | -151.7521 | 0.23289 | 8.4024 | 0.00268 | 0.004616 | 0.004373 | 7e-05 | 0.001148 | 0.36248 | 1.3747 | 2.5475 | 100.21 | 20.78 | 38.051 | 68.981 | 86.1607 | 64.7334 | 0.009327 | 1.0449 | 0.024986 | 23.9128 | 0.12284 | 0.13855 | -0.000144 | -0.010723 | -0.010723 | -0.52586
2019-03-29 13:52 | 1553881929.5265 | -58 | 0.20954 | 0.000453 | -1e-06 | 92.8743 | 91.674 | 127.4576 | 21.1317 | 0.001888 | 0.001805 | -151.7578 | 0.23394 | 8.4015 | 0.002649 | 0.003799 | 0.003801 | 6.7e-05 | 0.001146 | 0.31365 | 1.3438 | 2.5473 | 100.21 | 20.78 | 38.053 | 67.1238 | 85.4256 | 65.1223 | 0.009327 | 1.0437 | 0.024983 | 23.9376 | 0.11635 | 0.13917 | -0.000274 | -0.020529 | -0.020529 | -0.27592
2019-03-29 13:52 | 1553881930.527 | -57 | 0.20952 | 0.000446 | -1e-06 | 93.2709 | 92.0048 | 127.543 | 21.1312 | 0.001933 | 0.001783 | -151.7515 | 0.2339 | 8.4017 | 0.002676 | 0.003969 | 0.003746 | 6.6e-05 | 0.001151 | 0.30646 | 1.3746 | 2.5474 | 100.21 | 20.79 | 38.062 | 68.2029 | 86.0643 | 64.6088 | 0.009335 | 1.0453 | 0.02501 | 23.9274 | 0.12226 | 0.13915 | -0.000111 | -0.008319 | -0.008319 | -0.68079
2019-03-29 13:52 | 1553881931.5276 | -56 | 0.20951 | 0.000452 | -1e-06 | 93.7451 | 92.7529 | 127.5736 | 21.1307 | 0.001947 | 0.001838 | -151.7632 | 0.23282 | 8.4032 | 0.002695 | 0.003659 | 0.00355 | 6.7e-05 | 0.001153 | 0.31581 | 1.3956 | 2.5474 | 100.21 | 20.79 | 38.062 | 68.5503 | 86.5173 | 65.936 | 0.009335 | 1.0412 | 0.024964 | 23.9772 | 0.095176 | 0.13851 | -5.9e-05 | -0.004341 | -0.004341 | -1.2986
2019-03-29 13:52 | 1553881932.5267 | -55 | 0.20953 | 0.000453 | -1e-06 | 94.0989 | 93.0959 | 127.3175 | 21.1312 | 0.001909 | 0.001812 | -151.7576 | 0.23327 | 8.4028 | 0.002724 | 0.004449 | 0.004231 | 7e-05 | 0.001148 | 0.32659 | 1.378 | 2.5474 | 100.21 | 20.78 | 38.067 | 67.6288 | 87.2298 | 65.3082 | 0.009331 | 1.0431 | 0.024962 | 23.931 | 0.095855 | 0.13877 | -0.000166 | -0.012375 | -0.012375 | -0.45641
2019-03-29 13:52 | 1553881933.528 | -54 | 0.20956 | 0.000449 | -1e-06 | 94.4799 | 93.3634 | 126.9994 | 21.1314 | 0.001843 | 0.00177 | -151.7461 | 0.23292 | 8.4031 | 0.002717 | 0.003844 | 0.003567 | 6.9e-05 | 0.001141 | 0.33915 | 1.3335 | 2.5474 | 100.21 | 20.78 | 38.067 | 66.0459 | 87.0474 | 64.2953 | 0.009331 | 1.0462 | 0.024969 | 23.8653 | 0.10633 | 0.13856 | -0.000362 | -0.027081 | -0.027081 | -0.20824
2019-03-29 13:52 | 1553881934.5268 | -53 | 0.20952 | 0.000448 | -1e-06 | 94.846 | 93.919 | 126.4396 | 21.131 | 0.001962 | 0.001779 | -151.7521 | 0.23238 | 8.4029 | 0.002663 | 0.003649 | 0.003471 | 7e-05 | 0.001135 | 0.32316 | 1.3261 | 2.5475 | 100.21 | 20.78 | 38.075 | 68.9039 | 85.7549 | 64.515 | 0.009333 | 1.0455 | 0.024905 | 23.8204 | 0.08785 | 0.13824 | -0.000152 | -0.011345 | -0.011345 | -0.49594
2019-03-29 13:52 | 1553881935.5267 | -52 | 0.20953 | 0.000453 | -1e-06 | 95.2046 | 94.3743 | 126.0953 | 21.1312 | 0.001984 | 0.001797 | -151.7614 | 0.2327 | 8.4027 | 0.002666 | 0.003444 | 0.003257 | 7.2e-05 | 0.00113 | 0.32565 | 1.344 | 2.5475 | 100.21 | 20.78 | 38.075 | 69.4344 | 85.8394 | 64.95 | 0.009333 | 1.0442 | 0.024855 | 23.8032 | 0.07835 | 0.13843 | -0.000183 | -0.013582 | -0.013582 | -0.41482
2019-03-29 13:52 | 1553881936.5258 | -51 | 0.20955 | 0.000451 | -1e-06 | 95.5411 | 94.6033 | 125.8509 | 21.1308 | 0.001952 | 0.001809 | -151.7671 | 0.23226 | 8.4035 | 0.002698 | 0.00347 | 0.003282 | 7.1e-05 | 0.00113 | 0.30055 | 1.3407 | 2.5474 | 100.21 | 20.79 | 38.083 | 68.6612 | 86.5947 | 65.2181 | 0.009341 | 1.0434 | 0.024821 | 23.7896 | 0.088226 | 0.13817 | -0.000287 | -0.021308 | -0.021308 | -0.26392
2019-03-29 13:52 | 1553881937.5269 | -50 | 0.20953 | 0.000455 | -1e-06 | 95.8353 | 94.9709 | 125.488 | 21.1322 | 0.001967 | 0.001783 | -151.7589 | 0.23247 | 8.4028 | 0.002623 | 0.004464 | 0.004216 | 7e-05 | 0.001131 | 0.30127 | 1.329 | 2.5474 | 100.21 | 20.79 | 38.083 | 69.025 | 84.7829 | 64.6074 | 0.009341 | 1.0453 | 0.024808 | 23.7338 | 0.081045 | 0.13829 | -0.000184 | -0.013705 | -0.013705 | -0.4107
2019-03-29 13:52 | 1553881938.5268 | -49 | 0.20952 | 0.000451 | -1e-06 | 96.0853 | 95.3439 | 125.5007 | 21.1311 | 0.001857 | 0.001811 | -151.7576 | 0.23245 | 8.4024 | 0.002621 | 0.004451 | 0.004346 | 7.1e-05 | 0.001137 | 0.32372 | 1.3563 | 2.5474 | 100.21 | 20.79 | 38.086 | 66.3929 | 84.739 | 65.2682 | 0.009341 | 1.0432 | 0.024785 | 23.7582 | 0.069279 | 0.13828 | -0.000125 | -0.009244 | -0.009244 | -0.60886
2019-03-29 13:52 | 1553881939.5269 | -48 | 0.20952 | 0.000451 | 0 | 96.3298 | 95.5886 | 125.143 | 21.1316 | 0.001957 | 0.001806 | -151.7501 | 0.23239 | 8.4027 | 0.002625 | 0.004075 | 0.003866 | 7.2e-05 | 0.001144 | 0.32731 | 1.3238 | 2.5474 | 100.21 | 20.79 | 38.091 | 68.7961 | 84.8432 | 65.1626 | 0.009342 | 1.0435 | 0.024754 | 23.7206 | 0.069088 | 0.13825 | -0.000106 | -0.007876 | -0.007876 | -0.71441
2019-03-29 13:52 | 1553881940.527 | -47 | 0.20949 | 0.00045 | -1e-06 | 96.6157 | 95.8442 | 124.6071 | 21.1314 | 0.001909 | 0.00183 | -151.747 | 0.23308 | 8.4018 | 0.002695 | 0.004317 | 0.004119 | 7.5e-05 | 0.001142 | 0.30663 | 1.3221 | 2.5474 | 100.21 | 20.79 | 38.094 | 67.6453 | 86.5219 | 65.7253 | 0.009343 | 1.0418 | 0.02468 | 23.6895 | 0.071704 | 0.13866 | 3.3e-05 | 0.00245 | 0.00245 | 2.3036
2019-03-29 13:52 | 1553881941.5273 | -46 | 0.20953 | 0.000451 | -1e-06 | 96.868 | 96.1378 | 124.3822 | 21.1317 | 0.001929 | 0.001765 | -151.7547 | 0.23242 | 8.4005 | 0.002691 | 0.003905 | 0.003763 | 7.4e-05 | 0.001143 | 0.33212 | 1.3364 | 2.5472 | 100.21 | 20.79 | 38.098 | 68.1089 | 86.4214 | 64.1815 | 0.009344 | 1.0466 | 0.024714 | 23.6141 | 0.067685 | 0.13827 | -0.000167 | -0.012364 | -0.012364 | -0.45514
2019-03-29 13:52 | 1553881942.526 | -45 | 0.20954 | 0.000446 | -1e-06 | 97.035 | 96.5248 | 124.4207 | 21.1315 | 0.001965 | 0.001818 | -151.7691 | 0.23403 | 8.4032 | 0.002709 | 0.004243 | 0.004016 | 7.3e-05 | 0.001149 | 0.28799 | 1.344 | 2.5473 | 100.21 | 20.79 | 38.099 | 68.983 | 86.8639 | 65.4434 | 0.009344 | 1.0427 | 0.024672 | 23.6619 | 0.047159 | 0.13923 | -0.000266 | -0.019638 | -0.019638 | -0.28855
2019-03-29 13:52 | 1553881943.526 | -44 | 0.20955 | 0.00045 | -1e-06 | 97.338 | 96.7582 | 124.2736 | 21.1331 | 0.001965 | 0.001833 | -151.7953 | 0.23269 | 8.3996 | 0.002699 | 0.00375 | 0.003568 | 7e-05 | 0.001157 | 0.31996 | 1.329 | 2.5474 | 100.21 | 20.79 | 38.103 | 68.9762 | 86.6149 | 65.8174 | 0.009345 | 1.0415 | 0.024644 | 23.661 | 0.053433 | 0.13843 | -0.000278 | -0.020566 | -0.020566 | -0.27395
2019-03-29 13:52 | 1553881944.5271 | -43 | 0.20953 | 0.00045 | -1e-06 | 97.5571 | 97.0377 | 124.2344 | 21.1316 | 0.001988 | 0.001826 | -151.7875 | 0.23384 | 8.401 | 0.00268 | 0.004147 | 0.004066 | 6.7e-05 | 0.001162 | 0.28655 | 1.3597 | 2.5475 | 100.21 | 20.79 | 38.103 | 69.526 | 86.1589 | 65.6292 | 0.009345 | 1.0421 | 0.024647 | 23.6507 | 0.047743 | 0.13911 | -0.000193 | -0.01423 | -0.01423 | -0.39787
2019-03-29 13:52 | 1553881945.5275 | -42 | 0.20952 | 0.000451 | 0 | 97.7346 | 97.1493 | 124.2006 | 21.1317 | 0.001916 | 0.001799 | -151.7644 | 0.23238 | 8.4041 | 0.002706 | 0.003774 | 0.003565 | 7e-05 | 0.001161 | 0.31238 | 1.3783 | 2.5474 | 100.21 | 20.79 | 38.102 | 67.8009 | 86.7829 | 64.9863 | 0.009345 | 1.0441 | 0.024667 | 23.625 | 0.053721 | 0.13824 | -0.00016 | -0.011815 | -0.011815 | -0.4762
2019-03-29 13:52 | 1553881946.5275 | -41 | 0.20949 | 0.000451 | -1e-06 | 97.9184 | 97.378 | 124.2301 | 21.1322 | 0.001991 | 0.001855 | -151.7707 | 0.23355 | 8.4018 | 0.002716 | 0.003921 | 0.003734 | 7.1e-05 | 0.001164 | 0.32173 | 1.3431 | 2.5474 | 100.21 | 20.79 | 38.102 | 69.599 | 87.0315 | 66.3408 | 0.009345 | 1.0399 | 0.02462 | 23.6751 | 0.049498 | 0.13894 | 5.6e-05 | 0.004194 | 0.004194 | 1.3483
2019-03-29 13:52 | 1553881947.5262 | -40 | 0.20952 | 0.000454 | 0 | 98.121 | 97.7104 | 124.1919 | 21.132 | 0.001945 | 0.001825 | -151.7903 | 0.23264 | 8.4028 | 0.002699 | 0.004061 | 0.003882 | 6.8e-05 | 0.00117 | 0.30862 | 1.3665 | 2.5476 | 100.21 | 20.79 | 38.103 | 68.5119 | 86.6286 | 65.6112 | 0.009345 | 1.0422 | 0.024643 | 23.646 | 0.037508 | 0.13839 | -0.00013 | -0.00958 | -0.00958 | -0.58796
2019-03-29 13:52 | 1553881948.5263 | -39 | 0.2095 | 0.000446 | -1e-06 | 98.2791 | 97.7616 | 124.3477 | 21.1325 | 0.001971 | 0.001843 | -151.7903 | 0.23281 | 8.4004 | 0.002675 | 0.003432 | 0.003183 | 7.1e-05 | 0.001171 | 0.31526 | 1.3416 | 2.5474 | 100.21 | 20.79 | 38.103 | 69.1222 | 86.0405 | 66.0587 | 0.009345 | 1.0408 | 0.024642 | 23.6764 | 0.047219 | 0.1385 | 1.1e-05 | 0.000847 | 0.000847 | 6.6556
2019-03-29 13:52 | 1553881949.5272 | -38 | 0.20954 | 0.000451 | -1e-06 | 98.5066 | 98.0186 | 124.325 | 21.1322 | 0.001992 | 0.001789 | -151.797 | 0.23367 | 8.4028 | 0.002674 | 0.004437 | 0.004297 | 6.9e-05 | 0.001172 | 0.31669 | 1.3612 | 2.5474 | 100.21 | 20.79 | 38.099 | 69.6396 | 86.0145 | 64.759 | 0.009344 | 1.0448 | 0.024687 | 23.6289 | 0.044417 | 0.13901 | -0.000245 | -0.01811 | -0.01811 | -0.31241
2019-03-29 13:52 | 1553881950.5256 | -37 | 0.20951 | 0.000446 | -1e-06 | 98.6203 | 98.2886 | 124.5137 | 21.1318 | 0.00194 | 0.001842 | -151.8 | 0.23269 | 8.4018 | 0.002667 | 0.004145 | 0.003949 | 6.9e-05 | 0.001167 | 0.33339 | 1.3503 | 2.5474 | 100.21 | 20.79 | 38.097 | 68.3703 | 85.8517 | 66.0251 | 0.009344 | 1.0409 | 0.02466 | 23.6911 | 0.030138 | 0.13843 | -7.2e-05 | -0.005297 | -0.005297 | -1.0637
2019-03-29 13:52 | 1553881951.5263 | -36 | 0.20952 | 0.000451 | -1e-06 | 98.6948 | 98.3838 | 124.8029 | 21.1315 | 0.001942 | 0.001808 | -151.7846 | 0.23342 | 8.4037 | 0.002606 | 0.003283 | 0.003111 | 7e-05 | 0.001161 | 0.29717 | 1.3426 | 2.5474 | 100.21 | 20.79 | 38.092 | 68.4207 | 84.3928 | 65.2086 | 0.009343 | 1.0434 | 0.024718 | 23.69 | 0.028235 | 0.13886 | -0.000136 | -0.010038 | -0.010038 | -0.56302
2019-03-29 13:53 | 1553881952.5263 | -35 | 0.20952 | 0.00045 | 0 | 98.9636 | 98.7872 | 124.7925 | 21.1316 | 0.001992 | 0.001749 | -151.7925 | 0.23292 | 8.4001 | 0.00261 | 0.004296 | 0.00404 | 6.8e-05 | 0.001158 | 0.323 | 1.3647 | 2.5474 | 100.21 | 20.79 | 38.087 | 69.6335 | 84.4906 | 63.7831 | 0.009342 | 1.0478 | 0.024769 | 23.639 | 0.015962 | 0.13856 | -0.000122 | -0.009052 | -0.009052 | -0.623
2019-03-29 13:53 | 1553881953.5252 | -34 | 0.20954 | 0.000453 | -1e-06 | 99.0615 | 98.8664 | 125.1467 | 21.1318 | 0.001892 | 0.001843 | -151.7868 | 0.23279 | 8.4021 | 0.002701 | 0.003877 | 0.003772 | 6.9e-05 | 0.001152 | 0.3011 | 1.3358 | 2.5475 | 100.21 | 20.79 | 38.083 | 67.23 | 86.676 | 66.037 | 0.009341 | 1.0409 | 0.024722 | 23.7516 | 0.017638 | 0.13849 | -0.00027 | -0.019944 | -0.019944 | -0.28261
2019-03-29 13:53 | 1553881954.5263 | -33 | 0.20948 | 0.000452 | 0 | 99.2208 | 98.9773 | 125.3967 | 21.1315 | 0.00195 | 0.001776 | -151.7844 | 0.23294 | 8.3997 | 0.002646 | 0.004186 | 0.003957 | 6.9e-05 | 0.001145 | 0.29159 | 1.3728 | 2.5475 | 100.21 | 20.79 | 38.064 | 68.6282 | 85.3504 | 64.4359 | 0.009336 | 1.0458 | 0.024805 | 23.7191 | 0.021976 | 0.13858 | 0.000107 | 0.008034 | 0.008034 | 0.70201
2019-03-29 13:53 | 1553881955.5264 | -32 | 0.20952 | 0.000452 | -1e-06 | 99.418 | 99.3704 | 125.5744 | 21.1325 | 0.001947 | 0.001806 | -151.799 | 0.23337 | 8.4018 | 0.00267 | 0.003822 | 0.003626 | 6.6e-05 | 0.001146 | 0.31902 | 1.3587 | 2.5475 | 100.21 | 20.79 | 38.064 | 68.5419 | 85.9233 | 65.155 | 0.009336 | 1.0436 | 0.024797 | 23.7612 | 0.004281 | 0.13883 | -0.000116 | -0.008579 | -0.008579 | -0.65866
2019-03-29 13:53 | 1553881956.5257 | -31 | 0.20956 | 0.000447 | -1e-06 | 99.5369 | 99.2223 | 125.6547 | 21.1331 | 0.001951 | 0.001813 | -151.7915 | 0.23317 | 8.4025 | 0.002657 | 0.003886 | 0.003774 | 6.9e-05 | 0.001148 | 0.34722 | 1.3671 | 2.5475 | 100.21 | 20.79 | 38.056 | 68.6353 | 85.6211 | 65.3187 | 0.009334 | 1.0431 | 0.024798 | 23.7746 | 0.028309 | 0.13871 | -0.000358 | -0.026599 | -0.026599 | -0.21224
2019-03-29 13:53 | 1553881957.5255 | -30 | 0.20953 | 0.000449 | -1e-06 | 99.6817 | 99.3653 | 125.5847 | 21.132 | 0.001963 | 0.001801 | -151.7698 | 0.23379 | 8.4031 | 0.002729 | 0.00438 | 0.004123 | 7.1e-05 | 0.001145 | 0.33323 | 1.3967 | 2.5476 | 100.21 | 20.8 | 38.039 | 68.9213 | 87.3538 | 65.0294 | 0.009335 | 1.044 | 0.024802 | 23.7578 | 0.028437 | 0.13908 | -0.000156 | -0.01154 | -0.01154 | -0.49053
2019-03-29 13:53 | 1553881958.5265 | -29 | 0.20954 | 0.000447 | -1e-06 | 99.8266 | 99.4752 | 125.3843 | 21.1315 | 0.00192 | 0.001807 | -151.7601 | 0.23464 | 8.4005 | 0.002679 | 0.004035 | 0.003956 | 6.8e-05 | 0.001146 | 0.32731 | 1.3632 | 2.5474 | 100.21 | 20.8 | 38.029 | 67.8966 | 86.1512 | 65.1788 | 0.009333 | 1.0435 | 0.024777 | 23.7441 | 0.03154 | 0.13959 | -0.000223 | -0.016558 | -0.016558 | -0.34312
2019-03-29 13:53 | 1553881959.5265 | -28 | 0.20951 | 0.000451 | -1e-06 | 99.9864 | 99.7445 | 125.726 | 21.1316 | 0.001969 | 0.001822 | -151.7423 | 0.23269 | 8.4018 | 0.002688 | 0.003634 | 0.003346 | 6.9e-05 | 0.001145 | 0.32565 | 1.3912 | 2.5475 | 100.21 | 20.8 | 38.008 | 69.0877 | 86.3596 | 65.5311 | 0.009328 | 1.0424 | 0.024798 | 23.7888 | 0.021671 | 0.13843 | -5.5e-05 | -0.004047 | -0.004047 | -1.392
2019-03-29 13:53 | 1553881960.5974 | -27 | 0.20954 | 0.000448 | -1e-06 | 100.1029 | 99.8387 | 126.0406 | 21.1322 | 0.002003 | 0.001836 | -151.7366 | 0.23242 | 8.402 | 0.002666 | 0.00395 | 0.003772 | 6.7e-05 | 0.001136 | 0.30038 | 1.3229 | 2.5473 | 100.21 | 20.8 | 37.999 | 69.8908 | 85.8403 | 65.8718 | 0.009326 | 1.0414 | 0.024816 | 23.8305 | 0.023638 | 0.13827 | -0.000216 | -0.016047 | -0.016047 | -0.35069
2019-03-29 13:53 | 1553881961.5418 | -26 | 0.20954 | 0.000449 | -1e-06 | 100.2813 | 100.0534 | 125.999 | 21.1318 | 0.001971 | 0.001835 | -151.748 | 0.23245 | 8.4057 | 0.002711 | 0.004045 | 0.003875 | 6.9e-05 | 0.001132 | 0.31166 | 1.3605 | 2.5473 | 100.21 | 20.8 | 37.988 | 69.1277 | 86.9048 | 65.8512 | 0.009323 | 1.0414 | 0.024813 | 23.8258 | 0.020353 | 0.13829 | -0.000261 | -0.019432 | -0.019432 | -0.28964
2019-03-29 13:53 | 1553881962.5418 | -25 | 0.20952 | 0.00045 | 0 | 100.3349 | 100.2656 | 126.2213 | 21.1316 | 0.0019 | 0.00187 | -151.7426 | 0.23248 | 8.4039 | 0.002684 | 0.003558 | 0.003438 | 6.8e-05 | 0.001126 | 0.31238 | 1.3513 | 2.5473 | 100.21 | 20.8 | 37.965 | 67.414 | 86.2566 | 66.7063 | 0.009317 | 1.0388 | 0.024804 | 23.8769 | 0.006174 | 0.13831 | -0.000119 | -0.008819 | -0.008819 | -0.63828
2019-03-29 13:53 | 1553881963.5408 | -24 | 0.20953 | 0.000454 | 0 | 100.4893 | 100.4987 | 126.1969 | 21.1322 | 0.001907 | 0.00186 | -151.7393 | 0.23421 | 8.4015 | 0.002666 | 0.003276 | 0.003179 | 6.4e-05 | 0.001118 | 0.32892 | 1.3196 | 2.5474 | 100.21 | 20.8 | 37.955 | 67.5759 | 85.834 | 66.45 | 0.009315 | 1.0396 | 0.02481 | 23.8656 | -0.000838 | 0.13933 | -0.000181 | -0.013455 | -0.013455 | -0.42145
2019-03-29 13:53 | 1553881964.5408 | -23 | 0.2095 | 0.000448 | -1e-06 | 100.5857 | 100.3803 | 126.3427 | 21.1328 | 0.00194 | 0.001857 | -151.7496 | 0.23295 | 8.402 | 0.002702 | 0.003844 | 0.003653 | 6.6e-05 | 0.001116 | 0.32604 | 1.359 | 2.5474 | 100.21 | 20.8 | 37.938 | 68.3775 | 86.6919 | 66.3938 | 0.009311 | 1.0398 | 0.024827 | 23.8774 | 0.018284 | 0.13858 | -9e-06 | -0.00059 | -0.00059 | -9.5586
2019-03-29 13:53 | 1553881965.5401 | -22 | 0.20952 | 0.000451 | -1e-06 | 100.6324 | 100.5663 | 126.5426 | 21.1331 | 0.001868 | 0.001806 | -151.7345 | 0.23372 | 8.4043 | 0.002631 | 0.003948 | 0.003746 | 6.8e-05 | 0.001114 | 0.30469 | 1.3335 | 2.5473 | 100.21 | 20.8 | 37.921 | 66.6382 | 84.9795 | 65.1664 | 0.009307 | 1.0435 | 0.024892 | 23.853 | 0.005876 | 0.13904 | -8.9e-05 | -0.006592 | -0.006592 | -0.85845
2019-03-29 13:53 | 1553881966.541 | -21 | 0.20952 | 0.000452 | -1e-06 | 100.6794 | 100.7138 | 126.511 | 21.1326 | 0.00189 | 0.001809 | -151.7353 | 0.23269 | 8.4027 | 0.002677 | 0.004168 | 0.004019 | 6.7e-05 | 0.001119 | 0.33627 | 1.3472 | 2.5475 | 100.21 | 20.8 | 37.921 | 67.1699 | 86.0859 | 65.2323 | 0.009307 | 1.0433 | 0.024886 | 23.8524 | -0.003051 | 0.13843 | -0.000141 | -0.010461 | -0.010461 | -0.53858
2019-03-29 13:53 | 1553881967.54 | -20 | 0.20952 | 0.000455 | -1e-06 | 100.7922 | 100.6093 | 127.0069 | 21.1321 | 0.001961 | 0.001778 | -151.7422 | 0.23328 | 8.4037 | 0.002719 | 0.003412 | 0.003306 | 6.9e-05 | 0.001116 | 0.31022 | 1.3856 | 2.5474 | 100.21 | 20.8 | 37.907 | 68.8783 | 87.0972 | 64.473 | 0.009303 | 1.0457 | 0.024963 | 23.8723 | 0.016248 | 0.13878 | -0.000104 | -0.007742 | -0.007742 | -0.72958
2019-03-29 13:53 | 1553881968.5415 | -19 | 0.2095 | 0.000452 | -1e-06 | 100.8713 | 100.5663 | 127.3215 | 21.1326 | 0.001945 | 0.001829 | -151.7492 | 0.23375 | 8.4013 | 0.002683 | 0.004081 | 0.003819 | 6.9e-05 | 0.001118 | 0.28854 | 1.3688 | 2.5474 | 100.21 | 20.8 | 37.907 | 68.5014 | 86.2365 | 65.7046 | 0.009303 | 1.0419 | 0.024948 | 23.9454 | 0.027082 | 0.13906 | 1.1e-05 | 0.00089 | 0.00089 | 6.3614
2019-03-29 13:53 | 1553881969.5414 | -18 | 0.20951 | 0.000451 | -1e-06 | 100.9374 | 100.7018 | 127.4149 | 21.1331 | 0.001962 | 0.001791 | -151.7662 | 0.23336 | 8.4014 | 0.002639 | 0.004195 | 0.004018 | 6.7e-05 | 0.001124 | 0.29999 | 1.3744 | 2.5474 | 100.21 | 20.8 | 37.903 | 68.9185 | 85.1756 | 64.8034 | 0.009302 | 1.0447 | 0.024991 | 23.9223 | 0.020904 | 0.13883 | -4.3e-05 | -0.00316 | -0.00316 | -1.7882
2019-03-29 13:53 | 1553881970.5404 | -17 | 0.20954 | 0.000453 | 0 | 100.9922 | 100.9315 | 127.5632 | 21.133 | 0.001854 | 0.001814 | -151.7828 | 0.23408 | 8.4012 | 0.002729 | 0.004058 | 0.003847 | 6.7e-05 | 0.00113 | 0.31653 | 1.3421 | 2.5475 | 100.21 | 20.8 | 37.899 | 66.3103 | 87.3565 | 65.3583 | 0.009301 | 1.0429 | 0.024985 | 23.9558 | 0.005378 | 0.13925 | -0.000236 | -0.017665 | -0.017665 | -0.32084
2019-03-29 13:53 | 1553881971.541 | -16 | 0.20952 | 0.000451 | -1e-06 | 101.047 | 100.9393 | 127.5905 | 21.1329 | 0.001949 | 0.001785 | -151.785 | 0.23391 | 8.4007 | 0.002722 | 0.003427 | 0.003266 | 6.9e-05 | 0.001129 | 0.3178 | 1.3813 | 2.5474 | 100.21 | 20.8 | 37.898 | 68.607 | 87.1791 | 64.6581 | 0.009301 | 1.0451 | 0.025013 | 23.9336 | 0.009533 | 0.13915 | -9.6e-05 | -0.007159 | -0.007159 | -0.79111
2019-03-29 13:53 | 1553881972.5402 | -15 | 0.20951 | 0.000456 | 0 | 101.0464 | 100.883 | 127.5714 | 21.1324 | 0.001927 | 0.001794 | -151.7881 | 0.23386 | 8.4041 | 0.00268 | 0.00393 | 0.003744 | 6.5e-05 | 0.001132 | 0.32228 | 1.3953 | 2.5475 | 100.21 | 20.81 | 37.897 | 68.066 | 86.1708 | 64.876 | 0.009306 | 1.0444 | 0.025003 | 23.9395 | 0.014471 | 0.13912 | -5.4e-05 | -0.004062 | -0.004062 | -1.394
2019-03-29 13:53 | 1553881973.5415 | -14 | 0.20951 | 0.000451 | -1e-06 | 101.1095 | 100.997 | 127.3628 | 21.1334 | 0.001944 | 0.001768 | -151.7704 | 0.23308 | 8.4044 | 0.002693 | 0.004216 | 0.00407 | 6.7e-05 | 0.001135 | 0.33052 | 1.4184 | 2.5475 | 100.21 | 20.81 | 37.898 | 68.4865 | 86.4747 | 64.2416 | 0.009307 | 1.0464 | 0.025006 | 23.8975 | 0.009959 | 0.13866 | -3.6e-05 | -0.002606 | -0.002606 | -2.1655
2019-03-29 13:53 | 1553881974.5406 | -13 | 0.20951 | 0.00045 | -1e-06 | 101.105 | 100.9127 | 127.5247 | 21.1329 | 0.001914 | 0.001753 | -151.7907 | 0.23356 | 8.4046 | 0.002713 | 0.003832 | 0.003657 | 7e-05 | 0.001139 | 0.33339 | 1.3604 | 2.5474 | 100.21 | 20.81 | 37.9 | 67.7528 | 86.9636 | 63.8765 | 0.009307 | 1.0475 | 0.025036 | 23.8997 | 0.017031 | 0.13895 | -5.8e-05 | -0.004312 | -0.004312 | -1.3116
2019-03-29 13:53 | 1553881975.5402 | -12 | 0.20952 | 0.000451 | -1e-06 | 101.1715 | 101.1076 | 127.5374 | 21.134 | 0.00189 | 0.001809 | -150.658 | 0.23294 | 8.4037 | 0.002708 | 0.003485 | 0.00333 | 6.5e-05 | 0.001145 | 0.31238 | 1.3731 | 2.5473 | 100.21 | 20.81 | 37.902 | 67.1727 | 86.8492 | 65.239 | 0.009308 | 1.0433 | 0.024986 | 23.9492 | 0.005651 | 0.13858 | -8.6e-05 | -0.006401 | -0.006401 | -0.88111
2019-03-29 13:53 | 1553881976.5413 | -11 | 0.20952 | 0.000448 | -1e-06 | 101.191 | 101.1486 | 127.6745 | 21.1332 | 0.00191 | 0.001723 | -111.39 | 0.23377 | 8.4009 | 0.002734 | 0.004309 | 0.004111 | 7.1e-05 | 0.00115 | 0.3095 | 1.3332 | 2.5474 | 100.21 | 20.81 | 37.905 | 67.6589 | 87.4667 | 63.1608 | 0.009308 | 1.0498 | 0.025077 | 23.8884 | 0.003749 | 0.13907 | -0.000105 | -0.007826 | -0.007826 | -0.72324
2019-03-29 13:53 | 1553881977.5412 | -10 | 0.20954 | 0.000455 | 0 | 101.179 | 100.9991 | 127.6531 | 21.1338 | 0.001865 | 0.001777 | -69.4954 | 0.2331 | 8.4018 | 0.002714 | 0.003985 | 0.003836 | 6.7e-05 | 0.001148 | 0.3298 | 1.3827 | 2.5474 | 100.21 | 20.81 | 37.915 | 66.5694 | 86.9784 | 64.4684 | 0.009311 | 1.0457 | 0.025026 | 23.9328 | 0.015919 | 0.13867 | -0.000271 | -0.020329 | -0.020329 | -0.27762
2019-03-29 13:53 | 1553881978.5413 | -9 | 0.20949 | 0.000455 | -1e-06 | 101.2014 | 101.1352 | 127.4487 | 21.1336 | 0.001894 | 0.001804 | -27.6179 | 0.23285 | 8.401 | 0.002694 | 0.00394 | 0.003729 | 6.7e-05 | 0.001147 | 0.32477 | 1.4076 | 2.5476 | 100.21 | 20.81 | 37.92 | 67.2637 | 86.499 | 65.0998 | 0.009312 | 1.0437 | 0.024983 | 23.9359 | 0.005852 | 0.13852 | 4.3e-05 | 0.003233 | 0.003233 | 1.7436
2019-03-29 13:53 | 1553881979.5413 | -8 | 0.20946 | 0.000449 | -1e-06 | 101.2212 | 101.0666 | 127.2032 | 21.1328 | 0.001875 | 0.001714 | 14.2817 | 0.23391 | 8.4014 | 0.00273 | 0.004472 | 0.004308 | 6.6e-05 | 0.001137 | 0.3183 | 1.3332 | 2.5473 | 100.21 | 20.81 | 37.925 | 66.8235 | 87.357 | 62.9415 | 0.009313 | 1.0504 | 0.025039 | 23.8364 | 0.013667 | 0.13915 | 0.000228 | 0.017237 | 0.017237 | 0.32857
2019-03-29 13:53 | 1553881980.5413 | -7 | 0.20948 | 0.000452 | 0 | 101.238 | 101.1451 | 126.8821 | 21.1336 | 0.001849 | 0.001731 | 56.1974 | 0.23423 | 8.4009 | 0.002708 | 0.004422 | 0.004318 | 6.8e-05 | 0.001133 | 0.29734 | 1.3602 | 2.5473 | 100.21 | 20.81 | 37.933 | 66.186 | 86.8488 | 63.3423 | 0.009315 | 1.0492 | 0.024992 | 23.8205 | 0.008204 | 0.13935 | 0.000128 | 0.009577 | 0.009577 | 0.5922
2019-03-29 13:53 | 1553881981.5404 | -6 | 0.20947 | 0.000456 | -1e-06 | 101.2062 | 101.184 | 126.5645 | 21.1338 | 0.00191 | 0.001767 | 98.1153 | 0.26347 | 8.403 | 0.00271 | 0.00429 | 0.004161 | 7e-05 | 0.001135 | 0.29192 | 1.3343 | 2.5473 | 100.21 | 20.81 | 37.956 | 67.6695 | 86.8935 | 64.2131 | 0.009321 | 1.0465 | 0.024929 | 23.8215 | 0.00196 | 0.15674 | 0.00016 | 0.012046 | 0.012046 | 0.52957
2019-03-29 13:53 | 1553881982.5404 | -5 | 0.20951 | 0.000456 | -1e-06 | 101.2326 | 101.0854 | 126.5002 | 21.1337 | 0.001795 | 0.001709 | 267.1637 | 0.26572 | 8.4009 | 0.002679 | 0.004467 | 0.004455 | 6.7e-05 | 0.001135 | 0.3131 | 1.3686 | 2.5475 | 100.21 | 20.81 | 37.97 | 64.8985 | 86.1338 | 62.8128 | 0.009324 | 1.0508 | 0.024974 | 23.7659 | 0.013011 | 0.15808 | -6.2e-05 | -0.00458 | -0.00458 | -1.4046
2019-03-29 13:53 | 1553881983.5394 | -4 | 0.20949 | 0.000457 | 0 | 101.2607 | 101.1875 | 126.6722 | 21.1329 | 0.001866 | 0.001727 | 267.0627 | 0.26584 | 8.4028 | 0.002723 | 0.004274 | 0.004247 | 7e-05 | 0.001131 | 0.29878 | 1.3834 | 2.5474 | 100.21 | 20.81 | 37.986 | 66.5952 | 87.2096 | 63.2658 | 0.009328 | 1.0494 | 0.024975 | 23.7981 | 0.006472 | 0.15815 | 3e-05 | 0.002244 | 0.002244 | 2.8677
2019-03-29 13:53 | 1553881984.5394 | -3 | 0.20951 | 0.000454 | 0 | 101.2442 | 101.2387 | 126.6154 | 21.1335 | 0.001869 | 0.001722 | 267.142 | 0.266 | 8.4033 | 0.002779 | 0.00434 | 0.004244 | 6.5e-05 | 0.00113 | 0.30685 | 1.344 | 2.5473 | 100.21 | 20.81 | 38.022 | 66.6668 | 88.5456 | 63.143 | 0.009337 | 1.0498 | 0.024973 | 23.7884 | 0.000494 | 0.15824 | -6.3e-05 | -0.004686 | -0.004686 | -1.3744
2019-03-29 13:53 | 1553881985.5395 | -2 | 0.20951 | 0.000454 | -1e-06 | 101.2338 | 101.0632 | 126.3286 | 21.1332 | 0.001903 | 0.001763 | 267.2708 | 0.26552 | 8.4031 | 0.002851 | 0.003674 | 0.00368 | 6.6e-05 | 0.001127 | 0.29717 | 1.3858 | 2.5474 | 100.21 | 20.81 | 38.042 | 67.4847 | 90.2829 | 64.1133 | 0.009342 | 1.0468 | 0.024909 | 23.7957 | 0.01508 | 0.15796 | -5.2e-05 | -0.00385 | -0.00385 | -1.6699
2019-03-29 13:53 | 1553881986.5395 | -1 | 0.2095 | 0.000456 | -1e-06 | 101.3092 | 101.1492 | 126.1915 | 21.1339 | 0.00184 | 0.001739 | 267.3372 | 0.26646 | 8.4006 | 0.002799 | 0.004199 | 0.004106 | 6.6e-05 | 0.001127 | 0.30397 | 1.3382 | 2.5474 | 100.21 | 20.81 | 38.061 | 65.9673 | 89.0357 | 63.5362 | 0.009347 | 1.0486 | 0.024917 | 23.7625 | 0.014134 | 0.15852 | 2.3e-05 | 0.001792 | 0.001792 | 3.6011
2019-03-29 13:53 | 1553881987.5405 | 0 | 0.20946 | 0.000454 | -1e-06 | 100.1586 | 101.0591 | 125.0181 | 21.1238 | 0.001897 | 0.001733 | 267.2363 | 0.26612 | 8.3996 | 0.002777 | 0.005164 | 0.004866 | 6.8e-05 | 0.001129 | 0.25409 | 1.4012 | 2.7157 | 100.21 | 20.81 | 38.101 | 67.3502 | 88.4846 | 63.3969 | 0.009356 | 1.049 | 0.024806 | 23.6468 | -0.080065 | 0.15831 | 0.000261 | 0.019492 | 0.019492 | 0.33056
2019-03-29 13:53 | 1553881988.5406 | 1 | 0.20949 | 0.000451 | 0 | 96.2089 | 99.7779 | 124.0212 | 21.1214 | 0.001844 | 0.001769 | 267.4351 | 0.26652 | 8.4031 | 0.002866 | 0.004885 | 0.024577 | 6.1e-05 | 0.001132 | 0.28279 | 1.4041 | 2.5483 | 100.21 | 20.81 | 38.12 | 66.0713 | 90.6309 | 64.2644 | 0.009361 | 1.0463 | 0.024675 | 23.5827 | -0.32581 | 0.15855 | 6.1e-05 | 0.004521 | 0.004521 | 1.4274
2019-03-29 13:53 | 1553881989.5396 | 2 | 0.20949 | 0.000459 | 1e-06 | 94.7353 | 99.0223 | 122.654 | 21.1295 | 0.001843 | 0.001783 | 267.2074 | 0.26439 | 8.4032 | 0.002838 | 0.041019 | 0.073713 | 6.5e-05 | 0.001131 | 0.69568 | 1.3911 | 2.5477 | 100.21 | 20.81 | 38.152 | 66.0534 | 89.9761 | 64.604 | 0.009369 | 1.0453 | 0.024526 | 23.4642 | -0.39588 | 0.15729 | 3.1e-05 | 0.002232 | 0.002232 | 2.8682
2019-03-29 13:53 | 1553881990.5406 | 3 | 0.20947 | 0.000457 | 3e-06 | 95.3752 | 99.3981 | 120.8588 | 21.1228 | 0.001882 | 0.001891 | 267.8146 | 0.26617 | 8.4027 | 0.002821 | 0.006948 | 0.009316 | 9e-05 | 0.001126 | 0.33373 | 1.3712 | 2.5475 | 100.21 | 20.81 | 38.165 | 66.9692 | 89.5522 | 67.1861 | 0.009372 | 1.0373 | 0.024254 | 23.3807 | -0.36953 | 0.15835 | 0.000153 | 0.010944 | 0.010944 | 0.5889
2019-03-29 13:53 | 1553881991.5397 | 4 | 0.20944 | 0.000462 | 5e-06 | 98.2075 | 99.995 | 118.9491 | 21.1217 | 0.001831 | 0.001839 | 267.3184 | 0.26595 | 8.4005 | 0.002871 | 0.00499 | 0.00529 | 7e-05 | 0.001113 | 0.33356 | 1.4095 | 2.5473 | 100.21 | 20.81 | 38.178 | 65.7485 | 90.7499 | 65.9482 | 0.009375 | 1.0411 | 0.024105 | 23.153 | -0.16134 | 0.15821 | 0.000351 | 0.025091 | 0.025091 | 0.25663
2019-03-29 13:53 | 1553881992.5397 | 5 | 0.20928 | 0.000463 | 9e-06 | 99.6404 | 100.1141 | 116.5181 | 21.127 | 0.001942 | 0.001935 | 267.378 | 0.26587 | 8.4015 | 0.002805 | 0.003778 | 0.003851 | 6.4e-05 | 0.001119 | 0.31653 | 1.3657 | 1.7444 | 100.21 | 20.82 | 38.194 | 68.4244 | 89.1602 | 68.2465 | 0.009385 | 1.0341 | 0.023777 | 22.9927 | -0.042421 | 0.15817 | 0.001332 | 0.094564 | 0.094564 | 0.068074
2019-03-29 13:53 | 1553881993.5397 | 6 | 0.20856 | 0.000473 | 1.6e-05 | 100.3744 | 99.5882 | 114.3819 | 21.1322 | 0.001982 | 0.002 | 267.1781 | 0.2664 | 8.4023 | 0.002818 | 0.004699 | 0.004597 | 6.2e-05 | 0.001124 | 0.26897 | 1.34 | 2.4181 | 100.21 | 20.82 | 38.194 | 69.38 | 89.4847 | 69.8251 | 0.009385 | 1.0294 | 0.023504 | 22.8336 | 0.070335 | 0.15848 | 0.005604 | 0.39437 | 0.39437 | 0.016355
2019-03-29 13:53 | 1553881994.5397 | 7 | 0.20758 | 0.000526 | 2.6e-05 | 100.8374 | 99.8076 | 112.0032 | 21.1326 | 0.002071 | 0.002093 | 267.1335 | 0.26624 | 8.3995 | 0.002749 | 0.004504 | 0.004461 | 6.1e-05 | 0.001119 | 0.28152 | 1.3565 | 2.5474 | 100.21 | 20.82 | 38.188 | 71.5183 | 87.8369 | 72.047 | 0.009384 | 1.0227 | 0.023183 | 22.668 | 0.091817 | 0.15839 | 0.011449 | 0.7946 | 0.7946 | 0.008113
2019-03-29 13:53 | 1553881995.5408 | 8 | 0.20577 | 0.000647 | 4e-05 | 101.0233 | 97.6111 | 110.0193 | 21.1338 | 0.00215 | 0.002182 | 267.0083 | 0.26517 | 8.4023 | 0.002809 | 0.004442 | 0.004322 | 6.4e-05 | 0.00112 | 0.31708 | 1.3639 | 2.5474 | 100.21 | 20.82 | 38.188 | 73.4256 | 89.2691 | 74.2067 | 0.009384 | 1.0164 | 0.022906 | 22.5365 | 0.30734 | 0.15775 | 0.022141 | 1.5172 | 1.5172 | 0.004232
2019-03-29 13:53 | 1553881996.5398 | 9 | 0.20416 | 0.001136 | 5.6e-05 | 101.02 | 93.3662 | 108.0988 | 21.1331 | 0.002238 | 0.002226 | 266.927 | 0.26573 | 8.4043 | 0.002797 | 0.004389 | 0.004376 | 6.1e-05 | 0.001121 | 0.27688 | 1.3557 | 2.5474 | 100.21 | 20.82 | 38.182 | 75.5459 | 88.9834 | 75.2521 | 0.009382 | 1.0133 | 0.022671 | 22.3725 | 0.70474 | 0.15809 | 0.031133 | 2.1095 | 2.1095 | 0.00305
2019-03-29 13:53 | 1553881997.5389 | 10 | 0.20327 | 0.001803 | 7.3e-05 | 101.0227 | 87.5421 | 107.0939 | 21.1337 | 0.002234 | 0.002328 | 266.8257 | 0.26586 | 8.3991 | 0.002717 | 0.004162 | 0.004106 | 6.3e-05 | 0.001121 | 0.27721 | 1.3837 | 2.5474 | 100.21 | 20.82 | 38.167 | 75.4427 | 87.0504 | 77.7112 | 0.009378 | 1.0062 | 0.022486 | 22.3467 | 1.2811 | 0.15816 | 0.035642 | 2.3937 | 2.3937 | 0.002689
2019-03-29 13:53 | 1553881998.5399 | 11 | 0.20176 | 0.002588 | 8.8e-05 | 101.0398 | 80.6887 | 106.0475 | 21.1339 | 0.002452 | 0.002376 | 266.7246 | 0.26552 | 8.4024 | 0.002893 | 0.004505 | 0.004362 | 6.4e-05 | 0.001126 | 0.30718 | 1.3548 | 2.5474 | 100.21 | 20.82 | 38.16 | 80.6833 | 91.2841 | 78.861 | 0.009377 | 1.0029 | 0.022339 | 22.2737 | 2.0118 | 0.15796 | 0.043606 | 2.9071 | 2.9071 | 0.002211
2019-03-29 13:53 | 1553881999.5399 | 12 | 0.20077 | 0.003469 | 0.0001 | 101.1176 | 74.9953 | 105.3614 | 21.1345 | 0.002419 | 0.002388 | 266.6168 | 0.26523 | 8.3991 | 0.002932 | 0.004404 | 0.004395 | 6.1e-05 | 0.001118 | 0.32997 | 1.395 | 2.5474 | 100.21 | 20.82 | 38.135 | 79.9017 | 92.2456 | 79.1378 | 0.009371 | 1.0022 | 0.022258 | 22.2103 | 2.6732 | 0.15779 | 0.048428 | 3.2152 | 3.2152 | 0.001997
2019-03-29 13:53 | 1553882000.54 | 13 | 0.20016 | 0.004375 | 0.000109 | 101.0446 | 71.8708 | 105.1948 | 21.1346 | 0.002455 | 0.002459 | 266.5048 | 0.26595 | 8.4021 | 0.002972 | 0.004399 | 0.00427 | 6.6e-05 | 0.001112 | 0.31974 | 1.3306 | 2.5474 | 100.21 | 20.82 | 38.135 | 80.7636 | 93.2054 | 80.8462 | 0.009371 | 0.99731 | 0.022187 | 22.2464 | 3.0473 | 0.15822 | 0.050983 | 3.3728 | 3.3728 | 0.001909
2019-03-29 13:53 | 1553882001.5399 | 14 | 0.19999 | 0.00512 | 0.000115 | 101.0515 | 69.9507 | 105.3617 | 21.1337 | 0.002567 | 0.002508 | 266.3995 | 0.26472 | 8.4028 | 0.002958 | 0.004708 | 0.004621 | 6.1e-05 | 0.001108 | 0.28622 | 1.3148 | 2.5474 | 100.21 | 20.82 | 38.117 | 83.4519 | 92.8728 | 82.0251 | 0.009366 | 0.994 | 0.022167 | 22.3011 | 3.2902 | 0.15748 | 0.051099 | 3.3773 | 3.3773 | 0.001898
2019-03-29 13:53 | 1553882002.54 | 15 | 0.19999 | 0.005543 | 0.000119 | 101.0045 | 70.594 | 105.5229 | 21.1346 | 0.002539 | 0.002519 | 266.2928 | 0.26558 | 8.3993 | 0.003072 | 0.004626 | 0.004615 | 6.6e-05 | 0.001106 | 0.30021 | 1.3834 | 2.5475 | 100.21 | 20.82 | 38.117 | 82.7747 | 95.6067 | 82.3052 | 0.009366 | 0.99322 | 0.022176 | 22.327 | 3.2041 | 0.158 | 0.05056 | 3.3428 | 3.3428 | 0.001924
2019-03-29 13:53 | 1553882003.54 | 16 | 0.19993 | 0.005772 | 0.00012 | 100.9632 | 71.9182 | 105.1312 | 21.1335 | 0.00268 | 0.002612 | 266.1528 | 0.26556 | 8.4016 | 0.003061 | 0.004716 | 0.004601 | 6.5e-05 | 0.001104 | 0.29192 | 1.3278 | 2.5475 | 100.21 | 20.82 | 38.112 | 86.1764 | 95.337 | 84.5415 | 0.009365 | 0.98701 | 0.022065 | 22.3555 | 3.0342 | 0.15798 | 0.05067 | 3.3333 | 3.3333 | 0.001929
2019-03-29 13:53 | 1553882004.5391 | 17 | 0.19984 | 0.005987 | 0.000121 | 100.9461 | 72.4684 | 104.826 | 21.1341 | 0.002714 | 0.002705 | 266.0169 | 0.26611 | 8.404 | 0.003012 | 0.005428 | 0.005284 | 6.5e-05 | 0.001104 | 0.29645 | 1.355 | 2.5473 | 100.21 | 20.82 | 38.108 | 86.9819 | 94.1633 | 86.7767 | 0.009364 | 0.98088 | 0.021965 | 22.3926 | 2.9645 | 0.15831 | 0.050926 | 3.3347 | 3.3347 | 0.001932
2019-03-29 13:53 | 1553882005.5401 | 18 | 0.19983 | 0.006186 | 0.000121 | 100.9419 | 73.3039 | 104.6037 | 21.1332 | 0.002767 | 0.002703 | 265.8988 | 0.26582 | 8.4038 | 0.003066 | 0.004542 | 0.004511 | 6.6e-05 | 0.00111 | 0.30718 | 1.3578 | 2.5475 | 100.21 | 20.82 | 38.104 | 88.2679 | 95.4629 | 86.7204 | 0.009363 | 0.98103 | 0.021943 | 22.3671 | 2.8616 | 0.15813 | 0.050751 | 3.32 | 3.32 | 0.001939
2019-03-29 13:53 | 1553882006.5392 | 19 | 0.19974 | 0.006348 | 0.000123 | 100.8937 | 74.4547 | 104.3355 | 21.1343 | 0.002831 | 0.002755 | 265.7754 | 0.26659 | 8.4034 | 0.003048 | 0.004752 | 0.004675 | 6.6e-05 | 0.001117 | 0.28008 | 1.361 | 2.5474 | 100.21 | 20.82 | 38.102 | 89.7917 | 95.0325 | 87.9746 | 0.009362 | 0.97763 | 0.021877 | 22.3773 | 2.718 | 0.1586 | 0.051085 | 3.3316 | 3.3316 | 0.001937
2019-03-29 13:53 | 1553882007.5402 | 20 | 0.19942 | 0.006489 | 0.000124 | 100.8832 | 75.013 | 103.9822 | 21.1346 | 0.002875 | 0.002795 | 265.6568 | 0.26581 | 8.4025 | 0.003098 | 0.004932 | 0.004912 | 6.5e-05 | 0.001114 | 0.30127 | 1.3556 | 2.5474 | 100.21 | 20.82 | 38.1 | 90.8635 | 96.2473 | 88.9263 | 0.009362 | 0.97506 | 0.021811 | 22.3688 | 2.6503 | 0.15813 | 0.052797 | 3.4325 | 3.4325 | 0.001875
2019-03-29 13:53 | 1553882008.5392 | 21 | 0.19922 | 0.006628 | 0.000123 | 100.8542 | 75.0229 | 103.5754 | 21.1349 | 0.002939 | 0.002823 | 265.5406 | 0.26563 | 8.4031 | 0.003079 | 0.005396 | 0.005281 | 6.5e-05 | 0.001116 | 0.30055 | 1.3493 | 2.5473 | 100.21 | 20.83 | 38.1 | 92.3954 | 95.7807 | 89.6109 | 0.009368 | 0.97322 | 0.021748 | 22.3461 | 2.6465 | 0.15802 | 0.053861 | 3.4913 | 3.4913 | 0.001842
2019-03-29 13:53 | 1553882009.5392 | 22 | 0.19877 | 0.006759 | 0.000124 | 100.8312 | 73.6991 | 102.8966 | 21.1346 | 0.00299 | 0.002942 | 265.4085 | 0.26495 | 8.4054 | 0.003126 | 0.004787 | 0.004808 | 6.6e-05 | 0.001113 | 0.30917 | 1.3277 | 2.5474 | 100.21 | 20.83 | 38.1 | 93.646 | 96.9248 | 92.4677 | 0.009368 | 0.96561 | 0.021591 | 22.3603 | 2.8037 | 0.15762 | 0.056375 | 3.6273 | 3.6273 | 0.001769
2019-03-29 13:53 | 1553882010.5393 | 23 | 0.19857 | 0.006904 | 0.000124 | 100.8177 | 73.2926 | 102.303 | 21.1341 | 0.003062 | 0.002934 | 265.2828 | 0.26465 | 8.4012 | 0.003194 | 0.004249 | 0.004077 | 6.2e-05 | 0.001114 | 0.25713 | 1.3607 | 2.5474 | 100.21 | 20.83 | 38.1 | 95.3745 | 98.568 | 92.2843 | 0.009368 | 0.9661 | 0.021534 | 22.2901 | 2.852 | 0.15744 | 0.057384 | 3.6822 | 3.6822 | 0.00174
2019-03-29 13:54 | 1553882011.5393 | 24 | 0.1984 | 0.007062 | 0.000125 | 100.7818 | 74.867 | 101.6913 | 21.1347 | 0.003106 | 0.003056 | 265.1435 | 0.26514 | 8.4017 | 0.00324 | 0.005052 | 0.005048 | 6.4e-05 | 0.001113 | 0.28584 | 1.327 | 2.5474 | 100.21 | 20.83 | 38.099 | 96.4222 | 99.6714 | 95.2146 | 0.009368 | 0.95841 | 0.021384 | 22.3123 | 2.6587 | 0.15773 | 0.058212 | 3.709 | 3.709 | 0.001731
2019-03-29 13:54 | 1553882012.5384 | 25 | 0.1981 | 0.007225 | 0.000126 | 100.7479 | 75.5936 | 101.3113 | 21.1345 | 0.003144 | 0.003071 | 264.997 | 0.26472 | 8.4014 | 0.003165 | 0.004669 | 0.004647 | 6.4e-05 | 0.001112 | 0.30287 | 1.3446 | 2.5475 | 100.21 | 20.83 | 38.099 | 97.3386 | 97.8698 | 95.5901 | 0.009368 | 0.95744 | 0.021334 | 22.2819 | 2.5693 | 0.15749 | 0.059789 | 3.8 | 3.8 | 0.001687
2019-03-29 13:54 | 1553882013.5384 | 26 | 0.19785 | 0.0074 | 0.000126 | 100.7186 | 75.5308 | 101.3739 | 21.1343 | 0.003218 | 0.003166 | 264.8695 | 0.2655 | 8.4048 | 0.003241 | 0.004927 | 0.004919 | 6.5e-05 | 0.001105 | 0.28512 | 1.3315 | 2.5475 | 100.21 | 20.83 | 38.099 | 99.1376 | 99.6828 | 97.8822 | 0.009368 | 0.95152 | 0.021274 | 22.358 | 2.5742 | 0.15795 | 0.061068 | 3.87 | 3.87 | 0.001661
2019-03-29 13:54 | 1553882014.5397 | 27 | 0.19758 | 0.007573 | 0.000126 | 100.7363 | 75.1837 | 101.2941 | 21.1345 | 0.003317 | 0.003204 | 264.7469 | 0.26495 | 8.4012 | 0.003195 | 0.00521 | 0.005062 | 6.5e-05 | 0.001108 | 0.28191 | 1.345 | 2.5474 | 100.21 | 20.83 | 38.099 | 101.5373 | 98.5904 | 98.8067 | 0.009368 | 0.94916 | 0.021239 | 22.377 | 2.6169 | 0.15762 | 0.062427 | 3.9492 | 3.9492 | 0.001624
2019-03-29 13:54 | 1553882015.5385 | 28 | 0.19727 | 0.007763 | 0.000127 | 100.7057 | 74.9513 | 101.0115 | 21.1343 | 0.003322 | 0.00323 | 264.629 | 0.26437 | 8.4036 | 0.003307 | 0.00459 | 0.0045 | 6.1e-05 | 0.0011 | 0.28639 | 1.3355 | 2.5474 | 100.21 | 20.83 | 38.1 | 101.6543 | 101.3025 | 99.4396 | 0.009368 | 0.94754 | 0.021192 | 22.3648 | 2.6419 | 0.15728 | 0.064054 | 4.0425 | 4.0425 | 0.001583
2019-03-29 13:54 | 1553882016.5394 | 29 | 0.1972 | 0.007959 | 0.000127 | 100.654 | 75.5185 | 100.8838 | 21.1353 | 0.003408 | 0.003328 | 264.4923 | 0.26442 | 8.4016 | 0.003298 | 0.004898 | 0.00486 | 6.6e-05 | 0.001105 | 0.31454 | 1.3366 | 2.5475 | 100.21 | 20.83 | 38.101 | 103.7393 | 101.0757 | 101.7924 | 0.009368 | 0.9416 | 0.021112 | 22.4211 | 2.5699 | 0.1573 | 0.064262 | 4.0403 | 4.0403 | 0.001585
2019-03-29 13:54 | 1553882017.5385 | 30 | 0.19709 | 0.008156 | 0.000127 | 100.6953 | 75.9072 | 100.6766 | 21.1355 | 0.003462 | 0.00332 | 264.3382 | 0.26558 | 8.4014 | 0.003339 | 0.004894 | 0.004847 | 6.5e-05 | 0.001108 | 0.27865 | 1.3375 | 2.5474 | 100.21 | 20.83 | 38.102 | 105.0386 | 102.059 | 101.5989 | 0.009368 | 0.94208 | 0.021095 | 22.3923 | 2.5276 | 0.15799 | 0.064676 | 4.0631 | 4.0631 | 0.001583
2019-03-29 13:54 | 1553882018.5385 | 31 | 0.19689 | 0.008348 | 0.000126 | 100.5977 | 76.0188 | 100.5726 | 21.1342 | 0.00355 | 0.003377 | 264.1945 | 0.26532 | 8.4038 | 0.003303 | 0.005091 | 0.005066 | 6.4e-05 | 0.001106 | 0.29999 | 1.3549 | 2.5473 | 100.21 | 20.83 | 38.102 | 107.1733 | 101.1939 | 102.9816 | 0.009368 | 0.93862 | 0.021046 | 22.422 | 2.5058 | 0.15784 | 0.065625 | 4.1127 | 4.1127 | 0.001562
2019-03-29 13:54 | 1553882019.5396 | 32 | 0.19684 | 0.008522 | 0.000124 | 100.545 | 75.8157 | 100.8909 | 21.1345 | 0.003561 | 0.003383 | 264.0702 | 0.26524 | 8.4002 | 0.003382 | 0.004453 | 0.004408 | 6.3e-05 | 0.001102 | 0.28102 | 1.3572 | 2.5475 | 100.21 | 20.83 | 38.103 | 107.4484 | 103.1037 | 103.1429 | 0.009369 | 0.93822 | 0.021075 | 22.4622 | 2.525 | 0.15779 | 0.06575 | 4.1262 | 4.1262 | 0.001556
2019-03-29 13:54 | 1553882020.5386 | 33 | 0.19686 | 0.008689 | 0.000123 | 100.5731 | 75.6355 | 100.9815 | 21.1349 | 0.003539 | 0.00345 | 263.9183 | 0.26473 | 8.4047 | 0.003373 | 0.005108 | 0.004958 | 6.6e-05 | 0.001099 | 0.30309 | 1.3629 | 2.5473 | 100.2 | 20.83 | 38.108 | 106.9103 | 102.8801 | 104.7471 | 0.009371 | 0.93424 | 0.021039 | 22.5202 | 2.5488 | 0.15749 | 0.065406 | 4.0979 | 4.0979 | 0.001564
2019-03-29 13:54 | 1553882021.5387 | 34 | 0.19688 | 0.008846 | 0.000122 | 100.554 | 75.9523 | 100.9679 | 21.1344 | 0.003685 | 0.003447 | 263.7875 | 0.2661 | 8.4013 | 0.003296 | 0.004411 | 0.004356 | 6.7e-05 | 0.001094 | 0.30071 | 1.3613 | 2.5476 | 100.2 | 20.83 | 38.108 | 110.4568 | 101.0184 | 104.6843 | 0.009371 | 0.93439 | 0.021039 | 22.5168 | 2.5097 | 0.1583 | 0.065102 | 4.0791 | 4.0791 | 0.001579
2019-03-29 13:54 | 1553882022.5397 | 35 | 0.19696 | 0.008977 | 0.000121 | 100.5064 | 75.9707 | 101.1559 | 21.1351 | 0.003672 | 0.003498 | 263.657 | 0.26534 | 8.4016 | 0.003388 | 0.004717 | 0.004651 | 6.4e-05 | 0.001097 | 0.31758 | 1.3344 | 2.5474 | 100.2 | 20.83 | 38.114 | 110.1374 | 103.2566 | 105.9296 | 0.009372 | 0.93132 | 0.021024 | 22.5748 | 2.5033 | 0.15785 | 0.064489 | 4.0381 | 4.0381 | 0.001591
2019-03-29 13:54 | 1553882023.5397 | 36 | 0.19712 | 0.009088 | 0.000118 | 100.4363 | 76.5485 | 101.2596 | 21.1351 | 0.003704 | 0.003577 | 263.5284 | 0.26506 | 8.4039 | 0.00342 | 0.004864 | 0.004779 | 6.5e-05 | 0.001092 | 0.28351 | 1.3152 | 2.5474 | 100.2 | 20.83 | 38.114 | 110.9255 | 104.0278 | 107.8396 | 0.009372 | 0.92665 | 0.020982 | 22.6432 | 2.4293 | 0.15768 | 0.063391 | 3.9618 | 3.9618 | 0.00162
2019-03-29 13:54 | 1553882024.5402 | 37 | 0.19732 | 0.009173 | 0.000117 | 100.4528 | 76.5925 | 101.8819 | 21.1346 | 0.003791 | 0.00362 | 263.4042 | 0.26556 | 8.4033 | 0.003347 | 0.005172 | 0.005122 | 6.5e-05 | 0.00109 | 0.27848 | 1.2842 | 2.5472 | 100.2 | 20.83 | 38.121 | 113.0401 | 102.271 | 108.8894 | 0.009374 | 0.92411 | 0.021018 | 22.744 | 2.4257 | 0.15798 | 0.062107 | 3.8886 | 3.8886 | 0.001653
2019-03-29 13:54 | 1553882025.5391 | 38 | 0.1974 | 0.009235 | 0.000114 | 100.3971 | 77.771 | 102.4342 | 21.1349 | 0.003807 | 0.003544 | 263.2615 | 0.26428 | 8.4006 | 0.003438 | 0.004634 | 0.004558 | 6.7e-05 | 0.001082 | 0.28926 | 1.2934 | 2.5473 | 100.2 | 20.83 | 38.121 | 113.4298 | 104.4591 | 107.0424 | 0.009374 | 0.9286 | 0.021126 | 22.7504 | 2.2841 | 0.15722 | 0.06156 | 3.8745 | 3.8745 | 0.001652
2019-03-29 13:54 | 1553882026.5388 | 39 | 0.19762 | 0.009262 | 0.000112 | 100.3271 | 78.3788 | 102.8159 | 21.1352 | 0.003804 | 0.003556 | 263.128 | 0.26469 | 8.4009 | 0.003396 | 0.004595 | 0.00457 | 6.4e-05 | 0.001076 | 0.28263 | 1.3355 | 2.5475 | 100.2 | 20.83 | 38.126 | 113.3498 | 103.4531 | 107.3262 | 0.009375 | 0.9279 | 0.021157 | 22.8012 | 2.2082 | 0.15746 | 0.060205 | 3.7954 | 3.7954 | 0.001689
2019-03-29 13:54 | 1553882027.5389 | 40 | 0.19774 | 0.009264 | 0.00011 | 100.3178 | 78.6887 | 103.1466 | 21.1345 | 0.003838 | 0.003564 | 262.9969 | 0.26473 | 8.4024 | 0.00345 | 0.005082 | 0.00505 | 6.4e-05 | 0.001075 | 0.29214 | 1.325 | 2.5474 | 100.2 | 20.83 | 38.126 | 114.1812 | 104.7513 | 107.5137 | 0.009375 | 0.92745 | 0.021186 | 22.8435 | 2.1721 | 0.15749 | 0.059523 | 3.7578 | 3.7578 | 0.001706
2019-03-29 13:54 | 1553882028.5378 | 41 | 0.19788 | 0.009242 | 0.000107 | 100.2705 | 79.1573 | 103.6564 | 21.1349 | 0.003893 | 0.003657 | 262.8708 | 0.2652 | 8.403 | 0.003499 | 0.004728 | 0.004626 | 6.4e-05 | 0.001076 | 0.30254 | 1.3507 | 2.5475 | 100.21 | 20.83 | 38.128 | 115.5237 | 105.9497 | 109.7886 | 0.009375 | 0.92194 | 0.021175 | 22.9682 | 2.1148 | 0.15777 | 0.058735 | 3.7065 | 3.7065 | 0.001732
2019-03-29 13:54 | 1553882029.5378 | 42 | 0.19803 | 0.009207 | 0.000103 | 100.2424 | 79.3894 | 104.0648 | 21.1348 | 0.003955 | 0.003615 | 262.7558 | 0.26627 | 8.4026 | 0.003426 | 0.004937 | 0.00485 | 6.3e-05 | 0.001073 | 0.27632 | 1.3621 | 2.5475 | 100.21 | 20.83 | 38.129 | 117.0298 | 104.1735 | 108.7667 | 0.009375 | 0.9244 | 0.021245 | 22.9827 | 2.0861 | 0.15841 | 0.057885 | 3.6654 | 3.6654 | 0.001759
2019-03-29 13:54 | 1553882030.5379 | 43 | 0.19832 | 0.00916 | 0.0001 | 100.2298 | 80.0064 | 104.3991 | 21.1355 | 0.003921 | 0.003676 | 262.6318 | 0.26491 | 8.4027 | 0.003384 | 0.004792 | 0.004753 | 6.5e-05 | 0.001073 | 0.30038 | 1.3223 | 2.5474 | 100.21 | 20.83 | 38.133 | 116.2009 | 103.1635 | 110.2442 | 0.009376 | 0.92084 | 0.021238 | 23.064 | 2.0157 | 0.1576 | 0.056199 | 3.5581 | 3.5581 | 0.001803
2019-03-29 13:54 | 1553882031.5389 | 44 | 0.19874 | 0.009111 | 9.6e-05 | 100.2097 | 80.6143 | 105.0755 | 21.1359 | 0.003891 | 0.003642 | 262.4997 | 0.26552 | 8.4038 | 0.003302 | 0.004259 | 0.004236 | 6.1e-05 | 0.001069 | 0.29286 | 1.3256 | 2.5474 | 100.21 | 20.84 | 38.133 | 115.4725 | 101.1602 | 109.4102 | 0.009382 | 0.92285 | 0.02133 | 23.1134 | 1.9462 | 0.15796 | 0.0538 | 3.4218 | 3.4218 | 0.001879
2019-03-29 13:54 | 1553882032.539 | 45 | 0.19908 | 0.009038 | 9.1e-05 | 100.1726 | 81.3497 | 105.6344 | 21.1351 | 0.003952 | 0.003671 | 262.3761 | 0.26567 | 8.4044 | 0.003398 | 0.005247 | 0.005172 | 6.5e-05 | 0.001074 | 0.27024 | 1.3143 | 2.5473 | 100.21 | 20.84 | 38.135 | 116.9644 | 103.4877 | 110.1215 | 0.009382 | 0.92114 | 0.021367 | 23.1964 | 1.8617 | 0.15805 | 0.051876 | 3.3058 | 3.3058 | 0.001946
2019-03-29 13:54 | 1553882033.539 | 46 | 0.19965 | 0.008957 | 8.7e-05 | 100.1182 | 82.8135 | 106.2409 | 21.1351 | 0.00399 | 0.003652 | 262.263 | 0.26471 | 8.4018 | 0.003473 | 0.004785 | 0.004685 | 6.5e-05 | 0.001079 | 0.31957 | 1.3546 | 2.5475 | 100.2 | 20.84 | 38.138 | 117.8855 | 105.3087 | 109.6546 | 0.009384 | 0.92226 | 0.021441 | 23.2487 | 1.6973 | 0.15747 | 0.048555 | 3.106 | 3.106 | 0.002064
2019-03-29 13:54 | 1553882034.538 | 47 | 0.19986 | 0.008844 | 8.3e-05 | 100.1355 | 84.0053 | 106.5848 | 21.1366 | 0.003998 | 0.003613 | 262.1446 | 0.26554 | 8.4014 | 0.003448 | 0.004647 | 0.004602 | 6.5e-05 | 0.001082 | 0.30254 | 1.3951 | 2.5474 | 100.2 | 20.84 | 38.139 | 118.0906 | 104.7165 | 108.7236 | 0.009384 | 0.92451 | 0.021502 | 23.2579 | 1.5711 | 0.15797 | 0.047442 | 3.0438 | 3.0438 | 0.002112
2019-03-29 13:54 | 1553882035.5381 | 48 | 0.20016 | 0.008713 | 8e-05 | 100.0209 | 85.3588 | 107.3113 | 21.1364 | 0.00395 | 0.00371 | 262.0295 | 0.26556 | 8.4022 | 0.003477 | 0.004742 | 0.004577 | 6.2e-05 | 0.001087 | 0.27815 | 1.3703 | 2.5473 | 100.21 | 20.84 | 38.14 | 116.9174 | 105.4029 | 111.0617 | 0.009383 | 0.91888 | 0.02151 | 23.4084 | 1.4178 | 0.15798 | 0.045812 | 2.9408 | 2.9408 | 0.002186
2019-03-29 13:54 | 1553882036.5395 | 49 | 0.20079 | 0.008545 | 7.6e-05 | 100.0676 | 86.2414 | 108.0354 | 21.1361 | 0.003953 | 0.003606 | 261.9253 | 0.26521 | 8.4022 | 0.003435 | 0.004893 | 0.004843 | 6.2e-05 | 0.00109 | 0.30093 | 1.3858 | 2.5474 | 100.21 | 20.84 | 38.143 | 117.0011 | 104.3976 | 108.5521 | 0.009384 | 0.92492 | 0.021653 | 23.4104 | 1.33 | 0.15777 | 0.042259 | 2.7317 | 2.7317 | 0.002351
2019-03-29 13:54 | 1553882037.5382 | 50 | 0.20112 | 0.008355 | 7.2e-05 | 100.0771 | 87.1052 | 109.0806 | 21.137 | 0.003905 | 0.003553 | 261.8341 | 0.2657 | 8.4009 | 0.003412 | 0.005135 | 0.005074 | 6.2e-05 | 0.001088 | 0.27273 | 1.3723 | 2.5474 | 100.21 | 20.84 | 38.143 | 115.8131 | 103.8394 | 107.2593 | 0.009384 | 0.92807 | 0.021794 | 23.4835 | 1.2417 | 0.15807 | 0.040508 | 2.6361 | 2.6361 | 0.00244
2019-03-29 13:54 | 1553882038.5382 | 51 | 0.20172 | 0.008145 | 6.9e-05 | 99.9409 | 87.8813 | 109.9977 | 21.1366 | 0.003991 | 0.003538 | 261.7422 | 0.26514 | 8.401 | 0.003402 | 0.004981 | 0.004922 | 6.4e-05 | 0.00108 | 0.28064 | 1.3467 | 2.5474 | 100.21 | 20.84 | 38.142 | 117.9202 | 103.6023 | 106.8828 | 0.009384 | 0.92899 | 0.021897 | 23.5704 | 1.1502 | 0.15773 | 0.037158 | 2.4301 | 2.4301 | 0.002642
2019-03-29 13:54 | 1553882039.5392 | 52 | 0.20233 | 0.007908 | 6.7e-05 | 99.9565 | 88.041 | 110.8373 | 21.1361 | 0.00386 | 0.003584 | 261.6607 | 0.26487 | 8.3993 | 0.003353 | 0.004966 | 0.004892 | 6.6e-05 | 0.001068 | 0.29247 | 1.3193 | 2.5473 | 100.21 | 20.83 | 38.141 | 114.7306 | 102.4066 | 108.0122 | 0.009378 | 0.92623 | 0.021947 | 23.6953 | 1.1354 | 0.15757 | 0.033819 | 2.2175 | 2.2175 | 0.002892
2019-03-29 13:54 | 1553882040.5382 | 53 | 0.20273 | 0.007572 | 6.4e-05 | 99.9353 | 88.6441 | 111.7217 | 21.1359 | 0.003943 | 0.003439 | 261.5915 | 0.26537 | 8.4011 | 0.00337 | 0.005016 | 0.00496 | 6.4e-05 | 0.001063 | 0.29103 | 1.3587 | 2.5473 | 100.21 | 20.84 | 38.141 | 116.7444 | 102.8181 | 104.4868 | 0.009384 | 0.93488 | 0.022137 | 23.6794 | 1.0724 | 0.15787 | 0.031824 | 2.1051 | 2.1051 | 0.003052
2019-03-29 13:54 | 1553882041.5393 | 54 | 0.20297 | 0.007265 | 6.2e-05 | 99.9427 | 88.9495 | 112.3726 | 21.1363 | 0.003904 | 0.003493 | 261.5276 | 0.2655 | 8.3993 | 0.003335 | 0.004867 | 0.004797 | 6.3e-05 | 0.001063 | 0.27865 | 1.3364 | 2.5474 | 100.21 | 20.83 | 38.138 | 115.7892 | 101.9694 | 105.793 | 0.009377 | 0.93166 | 0.022163 | 23.7893 | 1.0423 | 0.15795 | 0.030792 | 2.0394 | 2.0394 | 0.003152
2019-03-29 13:54 | 1553882042.5399 | 55 | 0.20328 | 0.006927 | 6e-05 | 99.8715 | 89.5151 | 113.1188 | 21.1364 | 0.003893 | 0.003498 | 261.4505 | 0.26532 | 8.4047 | 0.003398 | 0.004542 | 0.004456 | 6.4e-05 | 0.001066 | 0.31636 | 1.3238 | 2.5475 | 100.21 | 20.83 | 38.136 | 115.5204 | 103.4883 | 105.9309 | 0.009377 | 0.93132 | 0.022233 | 23.8725 | 0.97923 | 0.15784 | 0.02933 | 1.9489 | 1.9489 | 0.003296
2019-03-29 13:54 | 1553882043.5374 | 56 | 0.20352 | 0.006644 | 5.8e-05 | 99.7922 | 89.6809 | 113.8157 | 21.1366 | 0.00389 | 0.003508 | 261.3842 | 0.26533 | 8.4037 | 0.003315 | 0.00522 | 0.005146 | 6.4e-05 | 0.001062 | 0.29645 | 1.3347 | 2.5474 | 100.21 | 20.83 | 38.132 | 115.4567 | 101.4908 | 106.1522 | 0.009376 | 0.93078 | 0.022295 | 23.9529 | 0.95556 | 0.15785 | 0.02826 | 1.8833 | 1.8833 | 0.003411
2019-03-29 13:54 | 1553882044.5384 | 57 | 0.20377 | 0.006415 | 5.7e-05 | 99.8817 | 89.3045 | 114.2443 | 21.1371 | 0.003874 | 0.003506 | 261.3201 | 0.26442 | 8.4045 | 0.003373 | 0.004789 | 0.00482 | 6.1e-05 | 0.001069 | 0.30574 | 1.3421 | 2.5474 | 100.21 | 20.83 | 38.129 | 115.0614 | 102.9011 | 106.1134 | 0.009375 | 0.93087 | 0.022338 | 23.9967 | 1.0012 | 0.1573 | 0.027005 | 1.8032 | 1.8032 | 0.00355
2019-03-29 13:54 | 1553882045.5384 | 58 | 0.204 | 0.006191 | 5.6e-05 | 99.8167 | 89.9038 | 114.2283 | 21.1371 | 0.003879 | 0.003454 | 261.2508 | 0.26457 | 8.4021 | 0.003339 | 0.00499 | 0.004887 | 6.5e-05 | 0.001075 | 0.2959 | 1.3635 | 2.5473 | 100.2 | 20.83 | 38.126 | 115.196 | 102.0723 | 104.8554 | 0.009375 | 0.93397 | 0.022373 | 23.9552 | 0.93556 | 0.15739 | 0.02589 | 1.7317 | 1.7317 | 0.003699
2019-03-29 13:54 | 1553882046.5374 | 59 | 0.20438 | 0.005971 | 5.5e-05 | 99.8634 | 90.0604 | 114.4023 | 21.1371 | 0.003854 | 0.00342 | 261.1778 | 0.26558 | 8.4041 | 0.003353 | 0.004419 | 0.004363 | 6.6e-05 | 0.001077 | 0.30879 | 1.3467 | 2.5476 | 100.2 | 20.83 | 38.122 | 114.5854 | 102.407 | 104.0343 | 0.009374 | 0.936 | 0.022415 | 23.9474 | 0.92417 | 0.158 | 0.023896 | 1.6014 | 1.6014 | 0.004016
2019-03-29 13:54 | 1553882047.5374 | 60 | 0.20461 | 0.005756 | 5.4e-05 | 99.8392 | 90.4352 | 114.414 | 21.137 | 0.003823 | 0.003397 | 261.1194 | 0.26545 | 8.3998 | 0.003317 | 0.004391 | 0.004324 | 6.8e-05 | 0.001081 | 0.33301 | 1.3745 | 2.5474 | 100.2 | 20.83 | 38.119 | 113.8312 | 101.5335 | 103.4693 | 0.009373 | 0.93741 | 0.022433 | 23.9307 | 0.88486 | 0.15791 | 0.022795 | 1.5289 | 1.5289 | 0.004204
2019-03-29 13:54 | 1553882048.5385 | 61 | 0.20462 | 0.005553 | 5.3e-05 | 99.8062 | 90.2441 | 114.4581 | 21.1371 | 0.003809 | 0.00342 | 261.0567 | 0.26607 | 8.4008 | 0.003175 | 0.005321 | 0.005241 | 6.5e-05 | 0.001081 | 0.32045 | 1.3521 | 2.5474 | 100.2 | 20.83 | 38.113 | 113.4764 | 98.0966 | 104.0234 | 0.009372 | 0.93603 | 0.022421 | 23.9529 | 0.90083 | 0.15829 | 0.022949 | 1.5386 | 1.5386 | 0.004187
2019-03-29 13:54 | 1553882049.5375 | 62 | 0.20481 | 0.00535 | 5.1e-05 | 99.7847 | 90.4553 | 114.8473 | 21.1371 | 0.003854 | 0.003341 | 261.0009 | 0.26518 | 8.4024 | 0.00329 | 0.004137 | 0.00408 | 6.3e-05 | 0.001085 | 0.29551 | 1.3723 | 2.5475 | 100.2 | 20.83 | 38.111 | 114.572 | 100.8924 | 102.1196 | 0.009371 | 0.94078 | 0.022516 | 23.9329 | 0.87799 | 0.15776 | 0.022112 | 1.4888 | 1.4888 | 0.004313
2019-03-29 13:54 | 1553882050.5375 | 63 | 0.20497 | 0.005156 | 5.1e-05 | 99.7605 | 90.357 | 115.1158 | 21.1372 | 0.003754 | 0.003304 | 260.9424 | 0.26536 | 8.4007 | 0.003225 | 0.005148 | 0.005009 | 6.4e-05 | 0.001076 | 0.30359 | 1.356 | 2.5474 | 100.2 | 20.83 | 38.11 | 112.1447 | 99.3216 | 101.2252 | 0.009371 | 0.94303 | 0.022569 | 23.9323 | 0.88554 | 0.15786 | 0.021397 | 1.4441 | 1.4441 | 0.004449
2019-03-29 13:54 | 1553882051.5386 | 64 | 0.20525 | 0.004968 | 5e-05 | 99.7506 | 90.7454 | 115.3513 | 21.137 | 0.003701 | 0.003344 | 260.8658 | 0.26573 | 8.4016 | 0.003178 | 0.004249 | 0.004183 | 6.7e-05 | 0.001075 | 0.29839 | 1.3259 | 2.5473 | 100.2 | 20.83 | 38.108 | 110.8539 | 98.1789 | 102.1853 | 0.009371 | 0.94061 | 0.022563 | 23.9875 | 0.84629 | 0.15808 | 0.019928 | 1.3447 | 1.3447 | 0.004785
2019-03-29 13:54 | 1553882052.5376 | 65 | 0.20548 | 0.004793 | 4.9e-05 | 99.7341 | 91.2966 | 115.6829 | 21.1368 | 0.003694 | 0.00329 | 260.8268 | 0.26557 | 8.4031 | 0.003223 | 0.005497 | 0.005437 | 6.1e-05 | 0.001076 | 0.30143 | 1.3197 | 2.5474 | 100.2 | 20.83 | 38.107 | 110.6793 | 99.2607 | 100.8866 | 0.00937 | 0.94388 | 0.022635 | 23.9804 | 0.79065 | 0.15799 | 0.018748 | 1.2692 | 1.2692 | 0.005066
2019-03-29 13:54 | 1553882053.5387 | 66 | 0.20569 | 0.004614 | 4.7e-05 | 99.7832 | 91.6409 | 116.0092 | 21.137 | 0.003736 | 0.00322 | 260.7621 | 0.26531 | 8.403 | 0.003202 | 0.004751 | 0.004742 | 6.5e-05 | 0.001071 | 0.28871 | 1.3335 | 2.5474 | 100.2 | 20.83 | 38.105 | 111.7009 | 98.7521 | 99.184 | 0.00937 | 0.9482 | 0.022718 | 23.9594 | 0.76138 | 0.15783 | 0.017687 | 1.2017 | 1.2017 | 0.005345
2019-03-29 13:54 | 1553882054.5376 | 67 | 0.20588 | 0.004445 | 4.6e-05 | 99.7608 | 91.7969 | 116.5085 | 21.1375 | 0.003711 | 0.003297 | 260.7234 | 0.26523 | 8.4018 | 0.003187 | 0.00517 | 0.005018 | 6.3e-05 | 0.001069 | 0.30127 | 1.2989 | 2.5474 | 100.2 | 20.83 | 38.102 | 111.0958 | 98.4029 | 101.0442 | 0.009369 | 0.94348 | 0.02271 | 24.0708 | 0.74415 | 0.15779 | 0.016801 | 1.1412 | 1.1412 | 0.005627
2019-03-29 13:54 | 1553882055.5377 | 68 | 0.20603 | 0.004275 | 4.5e-05 | 99.6841 | 92.0228 | 116.8415 | 21.1374 | 0.003707 | 0.003254 | 260.6738 | 0.26449 | 8.4019 | 0.00313 | 0.004648 | 0.004616 | 6.7e-05 | 0.001068 | 0.30757 | 1.3705 | 2.5474 | 100.2 | 20.83 | 38.099 | 110.9892 | 97.021 | 100.012 | 0.009368 | 0.94609 | 0.022774 | 24.0719 | 0.71529 | 0.15735 | 0.016114 | 1.0977 | 1.0977 | 0.005834
2019-03-29 13:54 | 1553882056.5377 | 69 | 0.20615 | 0.004115 | 4.4e-05 | 99.7521 | 92.1757 | 117.0079 | 21.1364 | 0.003691 | 0.003223 | 260.6111 | 0.26474 | 8.4033 | 0.003135 | 0.004939 | 0.004841 | 6.3e-05 | 0.001065 | 0.29999 | 1.3794 | 2.5475 | 100.2 | 20.83 | 38.094 | 110.6055 | 97.1249 | 99.2494 | 0.009367 | 0.94803 | 0.022814 | 24.0645 | 0.70654 | 0.1575 | 0.015583 | 1.0634 | 1.0634 | 0.006028
2019-03-29 13:54 | 1553882057.5387 | 70 | 0.20623 | 0.00397 | 4.4e-05 | 99.7275 | 92.4566 | 116.8516 | 21.1359 | 0.0036 | 0.003168 | 260.5572 | 0.26503 | 8.4019 | 0.003083 | 0.005229 | 0.00517 | 6.3e-05 | 0.001056 | 0.30453 | 1.4003 | 2.5473 | 100.2 | 20.83 | 38.09 | 108.3943 | 95.8727 | 97.9292 | 0.009366 | 0.9514 | 0.022839 | 24.0057 | 0.67712 | 0.15767 | 0.015275 | 1.0435 | 1.0435 | 0.006149
2019-03-29 13:54 | 1553882058.5388 | 71 | 0.20633 | 0.003833 | 4.2e-05 | 99.671 | 92.3699 | 117.0305 | 21.1364 | 0.00365 | 0.003203 | 260.5299 | 0.26575 | 8.4017 | 0.003055 | 0.004454 | 0.004434 | 6.1e-05 | 0.001058 | 0.31509 | 1.3856 | 2.5476 | 100.2 | 20.83 | 38.086 | 109.6117 | 95.21 | 98.7788 | 0.009365 | 0.94923 | 0.02283 | 24.0516 | 0.68044 | 0.1581 | 0.014865 | 1.0152 | 1.0152 | 0.006338
2019-03-29 13:54 | 1553882059.5368 | 72 | 0.20638 | 0.003705 | 4.2e-05 | 99.7227 | 92.5102 | 116.5643 | 21.1375 | 0.003616 | 0.003178 | 260.4675 | 0.26532 | 8.4016 | 0.003096 | 0.004733 | 0.004632 | 6.3e-05 | 0.001061 | 0.32621 | 1.3308 | 2.5475 | 100.2 | 20.84 | 38.078 | 108.7947 | 96.1859 | 98.17 | 0.009369 | 0.95078 | 0.022804 | 23.984 | 0.67151 | 0.15784 | 0.014705 | 1.0031 | 1.0031 | 0.006404
2019-03-29 13:54 | 1553882060.5369 | 73 | 0.20643 | 0.003594 | 4.1e-05 | 99.7006 | 92.9866 | 116.5641 | 21.1371 | 0.003566 | 0.003151 | 260.4133 | 0.26508 | 8.4032 | 0.00302 | 0.004803 | 0.004799 | 6.4e-05 | 0.001059 | 0.29767 | 1.3645 | 2.5476 | 100.2 | 20.84 | 38.076 | 107.58 | 94.3589 | 97.5102 | 0.009369 | 0.95248 | 0.022824 | 23.9626 | 0.62357 | 0.15769 | 0.014523 | 0.9916 | 0.9916 | 0.006472
2019-03-29 13:54 | 1553882061.5369 | 74 | 0.20645 | 0.003487 | 4.1e-05 | 99.6817 | 92.9699 | 116.4986 | 21.1371 | 0.003576 | 0.003222 | 260.383 | 0.26652 | 8.4022 | 0.002997 | 0.004698 | 0.004595 | 6.5e-05 | 0.001062 | 0.28816 | 1.391 | 2.5474 | 100.2 | 20.83 | 38.067 | 107.8132 | 93.805 | 99.2344 | 0.009361 | 0.94807 | 0.022765 | 24.0115 | 0.62349 | 0.15855 | 0.014573 | 0.99243 | 0.99243 | 0.006502
2019-03-29 13:54 | 1553882062.537 | 75 | 0.20647 | 0.003399 | 4.1e-05 | 99.6985 | 92.8716 | 116.566 | 21.1361 | 0.003517 | 0.00312 | 260.3417 | 0.26573 | 8.4019 | 0.002947 | 0.004316 | 0.004242 | 6.3e-05 | 0.001068 | 0.29645 | 1.3619 | 2.5475 | 100.2 | 20.83 | 38.063 | 106.3831 | 92.5896 | 96.776 | 0.00936 | 0.95437 | 0.022847 | 23.9391 | 0.63446 | 0.15809 | 0.014535 | 0.9934 | 0.9934 | 0.006477
2019-03-29 13:54 | 1553882063.538 | 76 | 0.20651 | 0.003319 | 4.2e-05 | 99.6572 | 92.7959 | 116.4993 | 21.1376 | 0.003482 | 0.003118 | 260.3174 | 0.26444 | 8.4025 | 0.00299 | 0.004156 | 0.004074 | 6.4e-05 | 0.001059 | 0.28313 | 1.3934 | 2.5475 | 100.2 | 20.84 | 38.059 | 105.5306 | 93.6416 | 96.7319 | 0.009364 | 0.95448 | 0.022841 | 23.9308 | 0.63805 | 0.15731 | 0.014392 | 0.98339 | 0.98339 | 0.006511
2019-03-29 13:54 | 1553882064.5381 | 77 | 0.20649 | 0.003247 | 4.2e-05 | 99.6269 | 92.9211 | 116.2812 | 21.137 | 0.00351 | 0.003106 | 260.2777 | 0.26535 | 8.4023 | 0.002965 | 0.004205 | 0.004122 | 6.3e-05 | 0.00106 | 0.32068 | 1.344 | 2.5474 | 100.2 | 20.84 | 38.05 | 106.2059 | 93.0302 | 96.4251 | 0.009362 | 0.95527 | 0.02283 | 23.8985 | 0.62327 | 0.15786 | 0.014637 | 0.99959 | 0.99959 | 0.006427
2019-03-29 13:54 | 1553882065.538 | 78 | 0.20651 | 0.00319 | 4.2e-05 | 99.6102 | 92.7061 | 116.5847 | 21.1372 | 0.003475 | 0.003083 | 260.2365 | 0.26465 | 8.4037 | 0.002984 | 0.004478 | 0.004424 | 6.1e-05 | 0.001064 | 0.32173 | 1.341 | 2.5474 | 100.2 | 20.84 | 38.046 | 105.3707 | 93.4857 | 95.8786 | 0.009361 | 0.95669 | 0.022876 | 23.9119 | 0.64249 | 0.15744 | 0.014543 | 0.9952 | 0.9952 | 0.006439
2019-03-29 13:54 | 1553882066.5371 | 79 | 0.20653 | 0.003137 | 4.2e-05 | 99.6383 | 92.7102 | 116.7859 | 21.1369 | 0.003498 | 0.003058 | 260.2116 | 0.26508 | 8.4002 | 0.002895 | 0.004641 | 0.004562 | 6.2e-05 | 0.001061 | 0.30685 | 1.3396 | 2.5473 | 100.2 | 20.84 | 38.036 | 105.932 | 91.3456 | 95.2879 | 0.009359 | 0.95822 | 0.022914 | 23.9134 | 0.64462 | 0.1577 | 0.014485 | 0.99284 | 0.99284 | 0.006465
2019-03-29 13:54 | 1553882067.5371 | 80 | 0.20653 | 0.003084 | 4.1e-05 | 99.6701 | 92.3884 | 117.2976 | 21.1367 | 0.003462 | 0.003087 | 260.1866 | 0.26415 | 8.4031 | 0.002881 | 0.004461 | 0.004368 | 6.2e-05 | 0.00106 | 0.32836 | 1.2871 | 2.5474 | 100.2 | 20.84 | 38.036 | 105.0537 | 91.0092 | 95.9813 | 0.009359 | 0.95642 | 0.022943 | 23.9883 | 0.67857 | 0.15714 | 0.014581 | 1.0007 | 1.0007 | 0.006391
2019-03-29 13:54 | 1553882068.5381 | 81 | 0.20657 | 0.003047 | 4.2e-05 | 99.7072 | 92.5256 | 117.4243 | 21.1369 | 0.003456 | 0.003047 | 260.146 | 0.2658 | 8.4049 | 0.002883 | 0.004455 | 0.004441 | 6.1e-05 | 0.001055 | 0.28031 | 1.3014 | 2.5474 | 100.2 | 20.84 | 38.033 | 104.9092 | 91.0478 | 95.0005 | 0.009358 | 0.95897 | 0.022986 | 23.9693 | 0.66863 | 0.15812 | 0.014365 | 0.98768 | 0.98768 | 0.006516
2019-03-29 13:54 | 1553882069.5371 | 82 | 0.20653 | 0.003015 | 4.3e-05 | 99.6593 | 92.6518 | 117.7201 | 21.1376 | 0.003398 | 0.003065 | 260.1151 | 0.26499 | 8.4013 | 0.0029 | 0.004244 | 0.004211 | 6.3e-05 | 0.001055 | 0.29502 | 1.3019 | 2.5471 | 100.2 | 20.84 | 38.027 | 103.4933 | 91.4721 | 95.4453 | 0.009357 | 0.95781 | 0.023001 | 24.014 | 0.65213 | 0.15764 | 0.014642 | 1.0074 | 1.0074 | 0.006369
2019-03-29 13:54 | 1553882070.5372 | 83 | 0.20655 | 0.002981 | 4.3e-05 | 99.6605 | 92.7204 | 117.8487 | 21.1363 | 0.00344 | 0.003039 | 260.0731 | 0.26507 | 8.4016 | 0.002797 | 0.003984 | 0.003909 | 6.2e-05 | 0.001055 | 0.29175 | 1.3186 | 2.5473 | 100.2 | 20.84 | 38.025 | 104.5211 | 88.9966 | 94.82 | 0.009356 | 0.95944 | 0.023033 | 24.0067 | 0.64562 | 0.15769 | 0.014558 | 1.003 | 1.003 | 0.006399
2019-03-29 13:54 | 1553882071.5382 | 84 | 0.20655 | 0.002952 | 4.3e-05 | 99.6252 | 92.7409 | 118.0841 | 21.1363 | 0.003434 | 0.003018 | 260.0386 | 0.26515 | 8.401 | 0.002888 | 0.004537 | 0.004394 | 6.4e-05 | 0.001056 | 0.32117 | 1.3488 | 2.5473 | 100.2 | 20.84 | 38.021 | 104.3624 | 91.166 | 94.3112 | 0.009355 | 0.96077 | 0.023072 | 24.014 | 0.64047 | 0.15774 | 0.014605 | 1.0079 | 1.0079 | 0.00637
2019-03-29 13:55 | 1553882072.5382 | 85 | 0.20658 | 0.002928 | 4.4e-05 | 99.7329 | 92.6938 | 117.813 | 21.1379 | 0.00346 | 0.003001 | 259.9933 | 0.26493 | 8.4028 | 0.002793 | 0.003885 | 0.003868 | 6.6e-05 | 0.001046 | 0.32748 | 1.3552 | 2.5475 | 100.2 | 20.84 | 38.018 | 105.0009 | 88.8801 | 93.9002 | 0.009354 | 0.96184 | 0.023058 | 23.9731 | 0.65469 | 0.1576 | 0.014481 | 0.99874 | 0.99874 | 0.006423
2019-03-29 13:55 | 1553882073.5383 | 86 | 0.20655 | 0.002908 | 4.4e-05 | 99.7194 | 92.6819 | 117.4985 | 21.1374 | 0.003447 | 0.002978 | 259.9479 | 0.2647 | 8.4028 | 0.002799 | 0.003507 | 0.003459 | 6.4e-05 | 0.001054 | 0.32244 | 1.3999 | 2.5475 | 100.2 | 20.84 | 38.008 | 104.6865 | 89.0439 | 93.3504 | 0.009352 | 0.96329 | 0.023045 | 23.9231 | 0.65463 | 0.15747 | 0.014677 | 1.0117 | 1.0117 | 0.006335
2019-03-29 13:55 | 1553882074.5363 | 87 | 0.20657 | 0.00289 | 4.5e-05 | 99.6967 | 92.6867 | 117.368 | 21.1368 | 0.003368 | 0.003007 | 259.8863 | 0.26543 | 8.4009 | 0.00284 | 0.004499 | 0.004362 | 6.2e-05 | 0.001054 | 0.29983 | 1.3782 | 2.5475 | 100.2 | 20.84 | 38.004 | 102.7815 | 90.0261 | 94.0432 | 0.009351 | 0.96147 | 0.02301 | 23.9324 | 0.65213 | 0.1579 | 0.014583 | 1.0036 | 1.0036 | 0.006403
2019-03-29 13:55 | 1553882075.5374 | 88 | 0.20658 | 0.00287 | 4.4e-05 | 99.7452 | 92.6819 | 117.0777 | 21.1375 | 0.003448 | 0.002997 | 259.8513 | 0.26582 | 8.3989 | 0.002719 | 0.004097 | 0.004096 | 6.5e-05 | 0.001057 | 0.30702 | 1.3486 | 2.5474 | 100.2 | 20.84 | 38 | 104.7157 | 87.1142 | 93.8022 | 0.00935 | 0.9621 | 0.022989 | 23.8949 | 0.65694 | 0.15813 | 0.014547 | 1.0003 | 1.0003 | 0.006434
2019-03-29 13:55 | 1553882076.5364 | 89 | 0.20657 | 0.002857 | 4.3e-05 | 99.741 | 92.6593 | 116.6603 | 21.1379 | 0.003363 | 0.002966 | 259.8144 | 0.2643 | 8.4044 | 0.002777 | 0.004064 | 0.003953 | 6.3e-05 | 0.001057 | 0.2907 | 1.3491 | 2.5474 | 100.2 | 20.84 | 37.99 | 102.6527 | 88.4998 | 93.0553 | 0.009347 | 0.96406 | 0.022972 | 23.828 | 0.65874 | 0.15723 | 0.014607 | 1.0037 | 1.0037 | 0.006376
2019-03-29 13:55 | 1553882077.5364 | 90 | 0.20659 | 0.002844 | 4.3e-05 | 99.7593 | 92.8791 | 116.4608 | 21.1374 | 0.003365 | 0.00299 | 259.7568 | 0.26525 | 8.4023 | 0.002754 | 0.004133 | 0.004076 | 6.4e-05 | 0.001054 | 0.3063 | 1.3513 | 2.5474 | 100.2 | 20.84 | 37.984 | 102.7034 | 87.9425 | 93.6306 | 0.009346 | 0.96255 | 0.022934 | 23.8263 | 0.63918 | 0.1578 | 0.014532 | 0.99692 | 0.99692 | 0.006442
2019-03-29 13:55 | 1553882078.5375 | 91 | 0.20661 | 0.002829 | 4.3e-05 | 99.7554 | 92.9286 | 116.2615 | 21.1374 | 0.003355 | 0.002952 | 259.7153 | 0.26484 | 8.4016 | 0.002696 | 0.004448 | 0.004332 | 6.6e-05 | 0.001058 | 0.31653 | 1.347 | 2.5474 | 100.2 | 20.84 | 37.973 | 102.4614 | 86.5457 | 92.7199 | 0.009343 | 0.96495 | 0.022943 | 23.7764 | 0.63407 | 0.15755 | 0.014407 | 0.98873 | 0.98873 | 0.006486
2019-03-29 13:55 | 1553882079.5366 | 92 | 0.2066 | 0.002814 | 4.3e-05 | 99.8056 | 93.13 | 116.3632 | 21.1379 | 0.003321 | 0.002963 | 259.6624 | 0.26548 | 8.3989 | 0.002691 | 0.004212 | 0.004176 | 6.4e-05 | 0.00107 | 0.2912 | 1.3577 | 2.5474 | 100.2 | 20.84 | 37.967 | 101.6333 | 86.4461 | 92.9761 | 0.009342 | 0.96427 | 0.022945 | 23.7951 | 0.61922 | 0.15793 | 0.014467 | 0.99288 | 0.99288 | 0.006474
2019-03-29 13:55 | 1553882080.5375 | 93 | 0.20659 | 0.002802 | 4.4e-05 | 99.8033 | 93.2904 | 116.3737 | 21.1382 | 0.003307 | 0.002898 | 259.6229 | 0.26491 | 8.4039 | 0.002687 | 0.003933 | 0.003873 | 6.3e-05 | 0.001075 | 0.30469 | 1.3404 | 2.5473 | 100.2 | 20.84 | 37.961 | 101.2913 | 86.3473 | 91.4274 | 0.00934 | 0.96837 | 0.022995 | 23.7458 | 0.60361 | 0.1576 | 0.014583 | 1.003 | 1.003 | 0.006395
2019-03-29 13:55 | 1553882081.5375 | 94 | 0.20659 | 0.002786 | 4.5e-05 | 99.8595 | 93.1102 | 116.2751 | 21.1372 | 0.003304 | 0.002947 | 259.5665 | 0.26493 | 8.3995 | 0.002724 | 0.004356 | 0.00429 | 6.4e-05 | 0.001074 | 0.3246 | 1.3554 | 2.5474 | 100.2 | 20.84 | 37.951 | 101.2276 | 87.2277 | 92.5883 | 0.009338 | 0.9653 | 0.022948 | 23.7735 | 0.62595 | 0.15761 | 0.014584 | 1.0011 | 1.0011 | 0.006408
2019-03-29 13:55 | 1553882082.5375 | 95 | 0.2066 | 0.002781 | 4.4e-05 | 99.8868 | 93.3911 | 116.6655 | 21.1374 | 0.003323 | 0.002928 | 259.5299 | 0.26507 | 8.4022 | 0.00268 | 0.004023 | 0.003936 | 6.4e-05 | 0.001072 | 0.31653 | 1.3629 | 2.5476 | 100.2 | 20.84 | 37.946 | 101.6769 | 86.1718 | 92.1373 | 0.009337 | 0.96649 | 0.023001 | 23.7986 | 0.60145 | 0.15769 | 0.014555 | 1.0014 | 1.0014 | 0.006409
2019-03-29 13:55 | 1553882083.5376 | 96 | 0.20663 | 0.002773 | 4.4e-05 | 99.9065 | 93.5034 | 117.1631 | 21.1377 | 0.003322 | 0.002926 | 259.4928 | 0.26533 | 8.4013 | 0.002704 | 0.00415 | 0.00409 | 6.6e-05 | 0.001062 | 0.29999 | 1.317 | 2.5473 | 100.2 | 20.85 | 37.937 | 101.6551 | 86.7373 | 92.0939 | 0.00934 | 0.9666 | 0.023051 | 23.8479 | 0.59247 | 0.15785 | 0.01438 | 0.99146 | 0.99146 | 0.00648
2019-03-29 13:55 | 1553882084.5372 | 97 | 0.20662 | 0.002766 | 4.5e-05 | 99.8529 | 93.3788 | 117.5194 | 21.1376 | 0.003293 | 0.002896 | 259.4503 | 0.26453 | 8.4029 | 0.00278 | 0.003908 | 0.003827 | 6.4e-05 | 0.001066 | 0.28584 | 1.3212 | 2.5474 | 100.2 | 20.85 | 37.93 | 100.9517 | 88.577 | 91.3745 | 0.009338 | 0.96851 | 0.023109 | 23.8606 | 0.59959 | 0.15737 | 0.014441 | 0.99811 | 0.99811 | 0.006417
2019-03-29 13:55 | 1553882085.5366 | 98 | 0.20664 | 0.002759 | 4.6e-05 | 99.9655 | 93.5088 | 117.9161 | 21.139 | 0.00325 | 0.002897 | 259.4035 | 0.26478 | 8.4025 | 0.002749 | 0.004742 | 0.004603 | 6.5e-05 | 0.001066 | 0.28871 | 1.351 | 2.5474 | 100.2 | 20.85 | 37.93 | 99.9202 | 87.8348 | 91.3874 | 0.009338 | 0.96847 | 0.023148 | 23.9013 | 0.59722 | 0.15752 | 0.014293 | 0.9895 | 0.9895 | 0.006479
2019-03-29 13:55 | 1553882086.5367 | 99 | 0.2066 | 0.002752 | 4.6e-05 | 99.9475 | 93.5692 | 118.5691 | 21.1386 | 0.003238 | 0.00283 | 259.3663 | 0.26421 | 8.4006 | 0.002722 | 0.003738 | 0.003705 | 6.1e-05 | 0.001067 | 0.29767 | 1.3475 | 2.5473 | 100.2 | 20.85 | 37.919 | 99.6147 | 87.1752 | 89.7764 | 0.009336 | 0.97277 | 0.023263 | 23.9144 | 0.58984 | 0.15718 | 0.014566 | 1.0134 | 1.0134 | 0.006312
2019-03-29 13:55 | 1553882087.5377 | 100 | 0.20663 | 0.002747 | 4.6e-05 | 99.9667 | 93.7839 | 118.8078 | 21.1381 | 0.003246 | 0.002909 | 259.3321 | 0.2643 | 8.4014 | 0.002653 | 0.00465 | 0.004636 | 6.5e-05 | 0.001065 | 0.30215 | 1.3448 | 2.5474 | 100.2 | 20.85 | 37.919 | 99.8166 | 85.5143 | 91.6759 | 0.009336 | 0.96771 | 0.023226 | 24.001 | 0.57105 | 0.15723 | 0.014407 | 1.0008 | 1.0008 | 0.006395
2019-03-29 13:55 | 1553882088.5368 | 101 | 0.20664 | 0.002738 | 4.7e-05 | 100.02 | 93.773 | 119.1259 | 21.1383 | 0.003267 | 0.002884 | 259.2924 | 0.26458 | 8.401 | 0.002676 | 0.004913 | 0.004846 | 6.3e-05 | 0.001064 | 0.30469 | 1.325 | 2.5474 | 100.2 | 20.85 | 37.914 | 100.3159 | 86.0719 | 91.0725 | 0.009334 | 0.96931 | 0.023276 | 24.0132 | 0.57686 | 0.1574 | 0.01437 | 1.0003 | 1.0003 | 0.006404
2019-03-29 13:55 | 1553882089.5368 | 102 | 0.20664 | 0.002731 | 4.7e-05 | 100.0753 | 93.9463 | 119.0341 | 21.1378 | 0.003276 | 0.002806 | 259.2547 | 0.26511 | 8.4027 | 0.002655 | 0.004621 | 0.004473 | 6.5e-05 | 0.001066 | 0.28943 | 1.3787 | 2.5474 | 100.2 | 20.85 | 37.913 | 100.5344 | 85.5586 | 89.2121 | 0.009334 | 0.97429 | 0.023327 | 23.9426 | 0.56529 | 0.15771 | 0.014372 | 1.0026 | 1.0026 | 0.006402
2019-03-29 13:55 | 1553882090.5359 | 103 | 0.20666 | 0.002721 | 4.7e-05 | 100.0888 | 94.0225 | 119.2054 | 21.1372 | 0.003257 | 0.002854 | 259.2005 | 0.26377 | 8.4016 | 0.002717 | 0.004798 | 0.004768 | 6.5e-05 | 0.001063 | 0.30735 | 1.3527 | 2.5474 | 100.2 | 20.85 | 37.912 | 100.0723 | 87.0614 | 90.368 | 0.009334 | 0.97119 | 0.023307 | 23.998 | 0.55925 | 0.15692 | 0.01421 | 0.99044 | 0.99044 | 0.006448
2019-03-29 13:55 | 1553882091.5368 | 104 | 0.20669 | 0.002712 | 4.7e-05 | 100.128 | 93.7426 | 119.2801 | 21.1377 | 0.003232 | 0.002868 | 259.1471 | 0.26444 | 8.4015 | 0.00262 | 0.004598 | 0.004525 | 6.6e-05 | 0.001066 | 0.30469 | 1.3058 | 2.5474 | 100.2 | 20.84 | 37.91 | 99.4852 | 84.7266 | 90.6905 | 0.009328 | 0.97033 | 0.023304 | 24.0161 | 0.58942 | 0.15731 | 0.014091 | 0.98198 | 0.98198 | 0.00652
2019-03-29 13:55 | 1553882092.5369 | 105 | 0.20669 | 0.002709 | 4.6e-05 | 100.0936 | 93.9951 | 119.0148 | 21.1375 | 0.003323 | 0.002878 | 259.1092 | 0.26437 | 8.4014 | 0.002673 | 0.004417 | 0.004373 | 6.4e-05 | 0.001068 | 0.30093 | 1.3758 | 2.5475 | 100.2 | 20.84 | 37.91 | 101.6896 | 85.9984 | 90.9272 | 0.009328 | 0.9697 | 0.02327 | 23.9972 | 0.56228 | 0.15727 | 0.014092 | 0.98068 | 0.98068 | 0.006527
2019-03-29 13:55 | 1553882093.5362 | 106 | 0.20669 | 0.002699 | 4.6e-05 | 100.1663 | 94.2293 | 118.7501 | 21.1376 | 0.003282 | 0.002883 | 259.0669 | 0.2653 | 8.4031 | 0.002706 | 0.005115 | 0.005046 | 6.7e-05 | 0.001059 | 0.29407 | 1.34 | 2.5474 | 100.2 | 20.84 | 37.91 | 100.695 | 86.7932 | 91.061 | 0.009328 | 0.96934 | 0.02324 | 23.9749 | 0.54652 | 0.15783 | 0.014065 | 0.97755 | 0.97755 | 0.006571
2019-03-29 13:55 | 1553882094.536 | 107 | 0.20669 | 0.002687 | 4.5e-05 | 100.2029 | 94.3678 | 118.4179 | 21.1379 | 0.003196 | 0.002827 | 259.0467 | 0.26411 | 8.4024 | 0.002738 | 0.004181 | 0.004075 | 6.5e-05 | 0.001053 | 0.30237 | 1.3328 | 2.5473 | 100.2 | 20.84 | 37.91 | 98.6039 | 87.5596 | 89.7194 | 0.009328 | 0.97293 | 0.02325 | 23.8972 | 0.53664 | 0.15712 | 0.014101 | 0.98055 | 0.98055 | 0.006521
2019-03-29 13:55 | 1553882095.537 | 108 | 0.20669 | 0.002682 | 4.5e-05 | 100.1789 | 94.4951 | 117.9797 | 21.1383 | 0.003213 | 0.002819 | 259.0046 | 0.26549 | 8.4027 | 0.00275 | 0.003671 | 0.003615 | 6.6e-05 | 0.00105 | 0.28528 | 1.3092 | 2.5474 | 100.2 | 20.84 | 37.91 | 99.019 | 87.8591 | 89.5276 | 0.009328 | 0.97344 | 0.023213 | 23.8467 | 0.52245 | 0.15794 | 0.014091 | 0.97835 | 0.97835 | 0.00657
2019-03-29 13:55 | 1553882096.5375 | 109 | 0.20673 | 0.002671 | 4.3e-05 | 100.2412 | 94.6433 | 117.787 | 21.138 | 0.003202 | 0.002862 | 258.9664 | 0.26523 | 8.4032 | 0.002662 | 0.004235 | 0.004205 | 6.3e-05 | 0.001052 | 0.27472 | 1.3155 | 2.5474 | 100.2 | 20.84 | 37.908 | 98.7562 | 85.7304 | 90.5578 | 0.009327 | 0.97068 | 0.023162 | 23.861 | 0.51399 | 0.15779 | 0.013908 | 0.96357 | 0.96357 | 0.006665
2019-03-29 13:55 | 1553882097.5371 | 110 | 0.20671 | 0.002657 | 4.3e-05 | 100.1843 | 95.0013 | 117.6084 | 21.1386 | 0.003221 | 0.002804 | 258.9333 | 0.26571 | 8.4028 | 0.002714 | 0.004296 | 0.004186 | 6.4e-05 | 0.001047 | 0.31725 | 1.3457 | 2.5476 | 100.2 | 20.84 | 37.906 | 99.2107 | 86.981 | 89.1488 | 0.009327 | 0.97446 | 0.023189 | 23.7967 | 0.47515 | 0.15807 | 0.014045 | 0.9742 | 0.9742 | 0.006604
2019-03-29 13:55 | 1553882098.5361 | 111 | 0.20678 | 0.002648 | 4.2e-05 | 100.2915 | 95.2337 | 117.242 | 21.1388 | 0.003122 | 0.002863 | 258.8971 | 0.265 | 8.4019 | 0.002715 | 0.004583 | 0.004492 | 6.4e-05 | 0.001044 | 0.31852 | 1.3211 | 2.5475 | 100.2 | 20.84 | 37.905 | 96.835 | 87.0153 | 90.5806 | 0.009327 | 0.97062 | 0.023107 | 23.8065 | 0.46285 | 0.15765 | 0.013609 | 0.94069 | 0.94069 | 0.006821
2019-03-29 13:55 | 1553882099.5371 | 112 | 0.20679 | 0.002633 | 4.1e-05 | 100.257 | 95.2828 | 117.0267 | 21.1376 | 0.00321 | 0.002812 | 258.8598 | 0.26478 | 8.4012 | 0.002744 | 0.004503 | 0.004474 | 6.6e-05 | 0.001037 | 0.3282 | 1.3123 | 2.5475 | 100.2 | 20.84 | 37.901 | 98.9509 | 87.7044 | 89.3441 | 0.009326 | 0.97393 | 0.023125 | 23.7442 | 0.45517 | 0.15752 | 0.013601 | 0.94086 | 0.94086 | 0.006814
2019-03-29 13:55 | 1553882100.5362 | 113 | 0.20679 | 0.002629 | 4e-05 | 100.2927 | 95.5402 | 116.6272 | 21.1384 | 0.003157 | 0.002794 | 258.8275 | 0.26515 | 8.4001 | 0.002702 | 0.004639 | 0.004647 | 6.3e-05 | 0.001032 | 0.32278 | 1.3197 | 2.5474 | 100.2 | 20.84 | 37.899 | 97.6753 | 86.691 | 88.9112 | 0.009325 | 0.9751 | 0.0231 | 23.6894 | 0.43422 | 0.15774 | 0.013561 | 0.93709 | 0.93709 | 0.006851
2019-03-29 13:55 | 1553882101.5366 | 114 | 0.20684 | 0.002612 | 3.9e-05 | 100.3112 | 95.6378 | 116.634 | 21.1378 | 0.003178 | 0.002852 | 258.8096 | 0.26495 | 8.4026 | 0.002673 | 0.004462 | 0.004244 | 6.5e-05 | 0.001027 | 0.31974 | 1.3926 | 2.5474 | 100.2 | 20.84 | 37.893 | 98.1632 | 86.0057 | 90.3014 | 0.009324 | 0.97137 | 0.023056 | 23.7356 | 0.42674 | 0.15762 | 0.013287 | 0.91654 | 0.91654 | 0.006999
2019-03-29 13:55 | 1553882102.5362 | 115 | 0.20688 | 0.0026 | 3.8e-05 | 100.3732 | 95.6088 | 116.4556 | 21.1382 | 0.003119 | 0.002848 | 258.7564 | 0.26495 | 8.4027 | 0.002657 | 0.004495 | 0.004412 | 5.9e-05 | 0.001027 | 0.29463 | 1.3731 | 2.5473 | 100.2 | 20.84 | 37.893 | 96.7419 | 85.6065 | 90.2049 | 0.009324 | 0.97163 | 0.023041 | 23.7143 | 0.43498 | 0.15762 | 0.013098 | 0.90296 | 0.90296 | 0.007105
2019-03-29 13:55 | 1553882103.5362 | 116 | 0.20689 | 0.002586 | 3.7e-05 | 100.4124 | 95.9634 | 116.2803 | 21.1385 | 0.003216 | 0.002798 | 258.7153 | 0.2653 | 8.4029 | 0.002686 | 0.004578 | 0.004581 | 5.8e-05 | 0.001029 | 0.30558 | 1.3988 | 2.5476 | 100.2 | 20.84 | 37.884 | 99.0837 | 86.3055 | 89.0129 | 0.009321 | 0.97483 | 0.023062 | 23.6575 | 0.40536 | 0.15782 | 0.013044 | 0.9001 | 0.9001 | 0.007136
2019-03-29 13:55 | 1553882104.5375 | 117 | 0.2069 | 0.002572 | 3.6e-05 | 100.428 | 96.1576 | 116.5192 | 21.1386 | 0.003135 | 0.002805 | 258.6619 | 0.26587 | 8.404 | 0.002607 | 0.004704 | 0.004637 | 6.1e-05 | 0.001038 | 0.32084 | 1.4087 | 2.5475 | 100.2 | 20.84 | 37.884 | 97.1287 | 84.421 | 89.17 | 0.009321 | 0.9744 | 0.023081 | 23.6869 | 0.38866 | 0.15817 | 0.012974 | 0.89599 | 0.89599 | 0.007185
2019-03-29 13:55 | 1553882105.5373 | 118 | 0.20696 | 0.002556 | 3.6e-05 | 100.4648 | 96.2446 | 116.6629 | 21.1379 | 0.003156 | 0.002799 | 258.6286 | 0.26448 | 8.4015 | 0.002747 | 0.004454 | 0.004426 | 5.8e-05 | 0.001037 | 0.30663 | 1.3935 | 2.5475 | 100.2 | 20.85 | 37.878 | 97.6553 | 87.7906 | 89.0403 | 0.009326 | 0.97475 | 0.023099 | 23.6973 | 0.38385 | 0.15734 | 0.012684 | 0.87666 | 0.87666 | 0.007305
2019-03-29 13:55 | 1553882106.5354 | 119 | 0.20689 | 0.002544 | 3.6e-05 | 100.4585 | 96.4753 | 116.4418 | 21.1379 | 0.003151 | 0.002786 | 258.5815 | 0.26465 | 8.4034 | 0.002705 | 0.004131 | 0.003981 | 5.7e-05 | 0.001038 | 0.29717 | 1.3859 | 2.5475 | 100.2 | 20.85 | 37.878 | 97.5197 | 86.7632 | 88.727 | 0.009326 | 0.97559 | 0.023087 | 23.6646 | 0.36188 | 0.15744 | 0.013101 | 0.90506 | 0.90506 | 0.00708
2019-03-29 13:55 | 1553882107.5367 | 120 | 0.20694 | 0.002536 | 3.6e-05 | 100.5396 | 96.5432 | 116.4094 | 21.1388 | 0.003151 | 0.002756 | 258.5359 | 0.26562 | 8.4039 | 0.002675 | 0.003744 | 0.00368 | 5.6e-05 | 0.001043 | 0.30055 | 1.3384 | 2.5474 | 100.2 | 20.85 | 37.874 | 97.5254 | 86.0591 | 87.9908 | 0.009325 | 0.97758 | 0.023107 | 23.6372 | 0.3628 | 0.15802 | 0.012804 | 0.88528 | 0.88528 | 0.007265
2019-03-29 13:55 | 1553882108.5357 | 121 | 0.20694 | 0.002517 | 3.5e-05 | 100.6348 | 96.9063 | 116.7289 | 21.1375 | 0.003106 | 0.002784 | 258.4891 | 0.26571 | 8.402 | 0.002702 | 0.004617 | 0.004581 | 5.8e-05 | 0.001042 | 0.32156 | 1.3203 | 2.5474 | 100.2 | 20.85 | 37.871 | 96.4249 | 86.694 | 88.6729 | 0.009324 | 0.97574 | 0.023117 | 23.692 | 0.33769 | 0.15807 | 0.012862 | 0.88971 | 0.88971 | 0.007231
2019-03-29 13:55 | 1553882109.5365 | 122 | 0.20697 | 0.002504 | 3.5e-05 | 100.6549 | 96.9609 | 117.0004 | 21.138 | 0.003131 | 0.002792 | 258.4506 | 0.26455 | 8.4031 | 0.002683 | 0.004306 | 0.004275 | 5.6e-05 | 0.001045 | 0.28031 | 1.3312 | 2.5474 | 100.2 | 20.85 | 37.868 | 97.0301 | 86.2549 | 88.8593 | 0.009323 | 0.97524 | 0.023138 | 23.7256 | 0.33443 | 0.15738 | 0.012656 | 0.87628 | 0.87628 | 0.00731
2019-03-29 13:55 | 1553882110.5355 | 123 | 0.207 | 0.002495 | 3.6e-05 | 100.6261 | 97.079 | 117.165 | 21.1379 | 0.00308 | 0.002752 | 258.4063 | 0.2653 | 8.4028 | 0.002732 | 0.003819 | 0.00378 | 5.6e-05 | 0.001044 | 0.29407 | 1.3435 | 2.5474 | 100.2 | 20.85 | 37.867 | 95.7979 | 87.4113 | 87.8984 | 0.009323 | 0.97783 | 0.023185 | 23.7108 | 0.32099 | 0.15783 | 0.012499 | 0.86713 | 0.86713 | 0.007408
2019-03-29 13:55 | 1553882111.5365 | 124 | 0.207 | 0.002475 | 3.6e-05 | 100.6402 | 97.092 | 117.2725 | 21.1385 | 0.003095 | 0.002711 | 258.345 | 0.26532 | 8.4022 | 0.002679 | 0.004148 | 0.004023 | 5.6e-05 | 0.001044 | 0.33107 | 1.3221 | 2.5474 | 100.2 | 20.85 | 37.867 | 96.1712 | 86.1368 | 86.9176 | 0.009323 | 0.9805 | 0.023227 | 23.6894 | 0.32104 | 0.15784 | 0.012518 | 0.87001 | 0.87001 | 0.007384
2019-03-29 13:55 | 1553882112.5356 | 125 | 0.20703 | 0.002461 | 3.6e-05 | 100.6962 | 97.3005 | 117.5156 | 21.1377 | 0.003123 | 0.002811 | 258.3174 | 0.26467 | 8.4026 | 0.002673 | 0.003768 | 0.003602 | 5.5e-05 | 0.001044 | 0.29535 | 1.3408 | 2.5472 | 100.2 | 20.85 | 37.865 | 96.8371 | 86.0004 | 89.3143 | 0.009322 | 0.97401 | 0.023174 | 23.7927 | 0.30683 | 0.15745 | 0.012345 | 0.85603 | 0.85603 | 0.007486
2019-03-29 13:55 | 1553882113.5356 | 126 | 0.20705 | 0.002448 | 3.6e-05 | 100.7007 | 97.2951 | 117.6342 | 21.1371 | 0.003124 | 0.002767 | 258.2853 | 0.26458 | 8.4025 | 0.00263 | 0.004414 | 0.004407 | 5.7e-05 | 0.001038 | 0.2896 | 1.429 | 2.5475 | 100.2 | 20.85 | 37.863 | 96.8756 | 84.968 | 88.2662 | 0.009322 | 0.97684 | 0.02322 | 23.7703 | 0.30773 | 0.1574 | 0.012224 | 0.84924 | 0.84924 | 0.007543
2019-03-29 13:55 | 1553882114.5369 | 127 | 0.20709 | 0.002438 | 3.6e-05 | 100.7438 | 97.6845 | 117.7666 | 21.1378 | 0.003122 | 0.002758 | 258.2639 | 0.26489 | 8.4027 | 0.002547 | 0.003786 | 0.003762 | 5.4e-05 | 0.001036 | 0.34667 | 1.3881 | 2.5476 | 100.2 | 20.85 | 37.861 | 96.8232 | 82.9777 | 88.0539 | 0.009321 | 0.97741 | 0.02324 | 23.7767 | 0.27583 | 0.15758 | 0.012005 | 0.83476 | 0.83476 | 0.007683
2019-03-29 13:55 | 1553882115.5367 | 128 | 0.20707 | 0.002424 | 3.5e-05 | 100.8351 | 97.9142 | 117.8889 | 21.1388 | 0.003088 | 0.002727 | 258.2145 | 0.26458 | 8.4006 | 0.00259 | 0.004868 | 0.004739 | 5.5e-05 | 0.001042 | 0.32532 | 1.3479 | 2.5474 | 100.2 | 20.85 | 37.861 | 96.0008 | 84.0005 | 87.3051 | 0.009321 | 0.97944 | 0.023276 | 23.7643 | 0.26292 | 0.1574 | 0.012152 | 0.84634 | 0.84634 | 0.007569
2019-03-29 13:55 | 1553882116.5361 | 129 | 0.20706 | 0.002409 | 3.5e-05 | 100.7671 | 97.9282 | 118.0877 | 21.1391 | 0.00303 | 0.002743 | 258.17 | 0.2645 | 8.4033 | 0.002647 | 0.003885 | 0.003809 | 5.6e-05 | 0.00104 | 0.31437 | 1.3651 | 2.5474 | 100.2 | 20.85 | 37.861 | 94.6051 | 85.3883 | 87.6895 | 0.009321 | 0.9784 | 0.023283 | 23.7971 | 0.25562 | 0.15735 | 0.012213 | 0.85085 | 0.85085 | 0.007527
2019-03-29 13:55 | 1553882117.5368 | 130 | 0.2071 | 0.002392 | 3.5e-05 | 100.7896 | 97.9845 | 118.1999 | 21.139 | 0.003041 | 0.00274 | 258.132 | 0.26477 | 8.403 | 0.002643 | 0.004088 | 0.004049 | 5.6e-05 | 0.001042 | 0.29839 | 1.3846 | 2.5474 | 100.2 | 20.84 | 37.86 | 94.8711 | 85.2905 | 87.6215 | 0.009315 | 0.97858 | 0.023296 | 23.8061 | 0.25247 | 0.15751 | 0.01205 | 0.84001 | 0.84001 | 0.007632
2019-03-29 13:55 | 1553882118.5357 | 131 | 0.20711 | 0.002381 | 3.6e-05 | 100.8632 | 98.207 | 118.4977 | 21.139 | 0.003074 | 0.002754 | 258.1059 | 0.26454 | 8.4026 | 0.00262 | 0.004295 | 0.004255 | 5.7e-05 | 0.001042 | 0.29374 | 1.3331 | 2.5473 | 100.2 | 20.84 | 37.859 | 95.6564 | 84.7217 | 87.9409 | 0.009315 | 0.97772 | 0.023315 | 23.8466 | 0.23871 | 0.15737 | 0.011996 | 0.83683 | 0.83683 | 0.007654
2019-03-29 13:55 | 1553882119.5362 | 132 | 0.20713 | 0.002372 | 3.5e-05 | 100.9305 | 98.424 | 118.6625 | 21.1389 | 0.003034 | 0.002706 | 258.0689 | 0.26503 | 8.404 | 0.002623 | 0.004173 | 0.004107 | 5.7e-05 | 0.001041 | 0.31454 | 1.355 | 2.5474 | 100.2 | 20.85 | 37.858 | 94.7082 | 84.8113 | 86.7984 | 0.009321 | 0.98082 | 0.023369 | 23.8254 | 0.22493 | 0.15767 | 0.011847 | 0.82832 | 0.82832 | 0.007747
2019-03-29 13:55 | 1553882120.5359 | 133 | 0.20712 | 0.002356 | 3.5e-05 | 100.8761 | 98.5756 | 118.6099 | 21.1395 | 0.003069 | 0.002714 | 258.0309 | 0.26553 | 8.4006 | 0.002587 | 0.003618 | 0.003542 | 6e-05 | 0.001039 | 0.28639 | 1.34 | 2.5474 | 100.2 | 20.84 | 37.857 | 95.5448 | 83.941 | 86.9895 | 0.009315 | 0.9803 | 0.023357 | 23.8265 | 0.20634 | 0.15796 | 0.011935 | 0.83414 | 0.83414 | 0.007707
2019-03-29 13:55 | 1553882121.5348 | 134 | 0.20712 | 0.002353 | 3.5e-05 | 100.9177 | 98.4984 | 118.3884 | 21.1393 | 0.003071 | 0.002732 | 258.0125 | 0.26491 | 8.4039 | 0.002586 | 0.003885 | 0.003883 | 5.8e-05 | 0.001037 | 0.32731 | 1.3293 | 2.5474 | 100.2 | 20.85 | 37.857 | 95.591 | 83.9082 | 87.425 | 0.00932 | 0.97912 | 0.023321 | 23.8186 | 0.21703 | 0.1576 | 0.011969 | 0.83519 | 0.83519 | 0.00768
2019-03-29 13:55 | 1553882122.535 | 135 | 0.20716 | 0.002344 | 3.4e-05 | 100.9485 | 98.5008 | 118.3592 | 21.1398 | 0.00307 | 0.002755 | 257.9883 | 0.265 | 8.4026 | 0.002561 | 0.004574 | 0.004441 | 6.1e-05 | 0.001033 | 0.31454 | 1.3616 | 2.5473 | 100.2 | 20.85 | 37.857 | 95.5741 | 83.3039 | 87.9737 | 0.00932 | 0.97763 | 0.023301 | 23.8338 | 0.21955 | 0.15765 | 0.011736 | 0.81829 | 0.81829 | 0.007841
2019-03-29 13:55 | 1553882123.5364 | 136 | 0.20721 | 0.00233 | 3.4e-05 | 100.9479 | 98.7602 | 118.6827 | 21.1399 | 0.003079 | 0.002701 | 257.9602 | 0.26462 | 8.4013 | 0.002577 | 0.004165 | 0.004151 | 5.9e-05 | 0.001032 | 0.30359 | 1.4052 | 2.5475 | 100.2 | 20.85 | 37.857 | 95.7793 | 83.6981 | 86.6675 | 0.00932 | 0.98118 | 0.023375 | 23.8231 | 0.19597 | 0.15742 | 0.011432 | 0.79964 | 0.79964 | 0.008013
2019-03-29 13:55 | 1553882124.535 | 137 | 0.20721 | 0.002318 | 3.3e-05 | 101.0338 | 98.8414 | 118.7679 | 21.139 | 0.003025 | 0.00274 | 257.8857 | 0.26595 | 8.4039 | 0.002567 | 0.004381 | 0.004226 | 6.1e-05 | 0.001037 | 0.29502 | 1.3584 | 2.5473 | 100.2 | 20.85 | 37.858 | 94.4861 | 83.4467 | 87.6254 | 0.009321 | 0.97857 | 0.023352 | 23.8634 | 0.19623 | 0.15821 | 0.011491 | 0.80301 | 0.80301 | 0.008019
2019-03-29 13:55 | 1553882125.536 | 138 | 0.20722 | 0.002305 | 3.3e-05 | 101.0057 | 98.7626 | 118.5811 | 21.1389 | 0.00306 | 0.002737 | 257.8635 | 0.26487 | 8.4031 | 0.002652 | 0.004103 | 0.004104 | 6.3e-05 | 0.001034 | 0.29966 | 1.3435 | 2.5474 | 100.2 | 20.85 | 37.859 | 95.3213 | 85.5043 | 87.5331 | 0.009321 | 0.97882 | 0.023337 | 23.8415 | 0.20087 | 0.15757 | 0.011443 | 0.79917 | 0.79917 | 0.008025
2019-03-29 13:55 | 1553882126.5354 | 139 | 0.20725 | 0.002296 | 3.3e-05 | 101.0548 | 98.9053 | 118.2297 | 21.1393 | 0.002979 | 0.002715 | 257.8376 | 0.26425 | 8.4017 | 0.002642 | 0.004218 | 0.004055 | 6.1e-05 | 0.001041 | 0.32029 | 1.3411 | 2.5474 | 100.2 | 20.85 | 37.86 | 93.3648 | 85.2511 | 87.005 | 0.009321 | 0.98026 | 0.023319 | 23.7888 | 0.19231 | 0.1572 | 0.011238 | 0.78423 | 0.78423 | 0.008159
2019-03-29 13:55 | 1553882127.536 | 140 | 0.20725 | 0.002281 | 3.2e-05 | 101.1251 | 99.0302 | 118.2316 | 21.1395 | 0.002996 | 0.002729 | 257.8041 | 0.26608 | 8.4007 | 0.002571 | 0.003871 | 0.003834 | 6.2e-05 | 0.001036 | 0.30989 | 1.345 | 2.5474 | 100.2 | 20.85 | 37.861 | 93.7727 | 83.5416 | 87.3534 | 0.009321 | 0.97931 | 0.023308 | 23.8005 | 0.18724 | 0.15829 | 0.011292 | 0.78765 | 0.78765 | 0.00818
2019-03-29 13:55 | 1553882128.5351 | 141 | 0.20726 | 0.002277 | 3.2e-05 | 101.1314 | 99.3923 | 117.8659 | 21.1399 | 0.002961 | 0.002705 | 257.7665 | 0.26427 | 8.4051 | 0.002666 | 0.004046 | 0.003942 | 6.1e-05 | 0.001038 | 0.27936 | 1.3349 | 2.5475 | 100.2 | 20.85 | 37.862 | 92.9414 | 85.8244 | 86.7704 | 0.009322 | 0.9809 | 0.023291 | 23.7444 | 0.15515 | 0.15721 | 0.011204 | 0.78092 | 0.78092 | 0.008194
2019-03-29 13:55 | 1553882129.5351 | 142 | 0.20726 | 0.002265 | 3.1e-05 | 101.1218 | 99.6718 | 117.7823 | 21.1395 | 0.002966 | 0.002714 | 257.7385 | 0.26526 | 8.4028 | 0.002536 | 0.004209 | 0.004097 | 6.2e-05 | 0.00104 | 0.3147 | 1.3342 | 2.5474 | 100.2 | 20.85 | 37.862 | 93.0482 | 82.7142 | 86.9903 | 0.009322 | 0.9803 | 0.023275 | 23.7432 | 0.12919 | 0.1578 | 0.011243 | 0.78317 | 0.78317 | 0.008201
2019-03-29 13:55 | 1553882130.5354 | 143 | 0.20731 | 0.002253 | 3.1e-05 | 101.1535 | 99.6779 | 118.0285 | 21.1395 | 0.003013 | 0.002687 | 257.7089 | 0.26595 | 8.4029 | 0.00256 | 0.003905 | 0.003823 | 6.1e-05 | 0.001033 | 0.30807 | 1.349 | 2.5474 | 100.2 | 20.85 | 37.862 | 94.1835 | 83.2765 | 86.3464 | 0.009322 | 0.98206 | 0.023321 | 23.7468 | 0.13144 | 0.15821 | 0.010937 | 0.76335 | 0.76335 | 0.008436
2019-03-29 13:56 | 1553882131.5352 | 144 | 0.20729 | 0.002242 | 3.1e-05 | 101.1622 | 99.9523 | 118.3315 | 21.1393 | 0.002997 | 0.002724 | 257.6648 | 0.26463 | 8.4025 | 0.002547 | 0.00441 | 0.004349 | 6e-05 | 0.001034 | 0.28335 | 1.3557 | 2.5474 | 100.2 | 20.85 | 37.86 | 93.7991 | 82.969 | 87.2353 | 0.009321 | 0.97963 | 0.023322 | 23.8066 | 0.10762 | 0.15743 | 0.011099 | 0.7747 | 0.7747 | 0.008271
2019-03-29 13:56 | 1553882132.5352 | 145 | 0.20733 | 0.002237 | 3.1e-05 | 101.1006 | 99.8796 | 118.1806 | 21.1398 | 0.003002 | 0.002701 | 257.6269 | 0.26599 | 8.4 | 0.002586 | 0.003781 | 0.003755 | 5.6e-05 | 0.001028 | 0.31636 | 1.3548 | 2.5474 | 100.2 | 20.85 | 37.859 | 93.9341 | 83.9016 | 86.6673 | 0.009321 | 0.98118 | 0.023325 | 23.7727 | 0.10867 | 0.15823 | 0.010872 | 0.75897 | 0.75897 | 0.008485
2019-03-29 13:56 | 1553882133.5362 | 146 | 0.20738 | 0.002222 | 3.1e-05 | 101.1895 | 100.0858 | 117.9999 | 21.1396 | 0.003017 | 0.002676 | 257.5992 | 0.26537 | 8.4028 | 0.0025 | 0.003849 | 0.003818 | 6e-05 | 0.001027 | 0.32117 | 1.3318 | 2.5473 | 100.2 | 20.85 | 37.856 | 94.2823 | 81.832 | 86.0763 | 0.00932 | 0.98279 | 0.023327 | 23.735 | 0.098095 | 0.15787 | 0.010582 | 0.73875 | 0.73875 | 0.008697
2019-03-29 13:56 | 1553882134.5367 | 147 | 0.20738 | 0.002212 | 3e-05 | 101.2403 | 100.1622 | 118.0661 | 21.1398 | 0.003021 | 0.002657 | 257.5882 | 0.26461 | 8.4035 | 0.002551 | 0.003735 | 0.003716 | 5.6e-05 | 0.001027 | 0.29734 | 1.3497 | 2.5475 | 100.2 | 20.85 | 37.853 | 94.392 | 83.0763 | 85.6081 | 0.009319 | 0.98408 | 0.023348 | 23.7262 | 0.095762 | 0.15741 | 0.010593 | 0.74025 | 0.74025 | 0.008655
2019-03-29 13:56 | 1553882135.5354 | 148 | 0.20737 | 0.002195 | 2.9e-05 | 101.2326 | 100.1428 | 118.2379 | 21.1392 | 0.003 | 0.002705 | 257.5571 | 0.26499 | 8.4018 | 0.002581 | 0.004283 | 0.004207 | 5.7e-05 | 0.001022 | 0.30685 | 1.3444 | 2.5473 | 100.2 | 20.85 | 37.847 | 93.8782 | 83.7917 | 86.7761 | 0.009318 | 0.98088 | 0.023327 | 23.782 | 0.096812 | 0.15764 | 0.010635 | 0.74255 | 0.74255 | 0.008641
2019-03-29 13:56 | 1553882136.5354 | 149 | 0.20741 | 0.002186 | 2.9e-05 | 101.284 | 100.2138 | 118.4886 | 21.1404 | 0.003052 | 0.002667 | 257.526 | 0.26511 | 8.4019 | 0.002571 | 0.003868 | 0.003754 | 5.4e-05 | 0.001023 | 0.29839 | 1.3315 | 2.5473 | 100.2 | 20.85 | 37.844 | 95.1379 | 83.5527 | 85.8618 | 0.009317 | 0.98338 | 0.023382 | 23.777 | 0.095021 | 0.15772 | 0.010447 | 0.7311 | 0.7311 | 0.00878
2019-03-29 13:56 | 1553882137.5357 | 150 | 0.20741 | 0.002168 | 2.9e-05 | 101.214 | 100.2984 | 118.2858 | 21.1397 | 0.002957 | 0.002699 | 257.5043 | 0.26472 | 8.4028 | 0.002522 | 0.003964 | 0.003831 | 5.8e-05 | 0.001018 | 0.29574 | 1.3386 | 2.5474 | 100.2 | 20.84 | 37.83 | 92.8419 | 82.3672 | 86.6369 | 0.009308 | 0.98126 | 0.023337 | 23.7822 | 0.081282 | 0.15748 | 0.010451 | 0.73001 | 0.73001 | 0.00878
2019-03-29 13:56 | 1553882138.5344 | 151 | 0.20741 | 0.00216 | 2.8e-05 | 101.3577 | 100.3325 | 118.2248 | 21.1396 | 0.002968 | 0.002633 | 257.4721 | 0.2641 | 8.4033 | 0.002555 | 0.004664 | 0.004686 | 5.3e-05 | 0.00102 | 0.29214 | 1.2997 | 2.5473 | 100.2 | 20.84 | 37.83 | 93.1054 | 83.1696 | 85.0376 | 0.009308 | 0.98564 | 0.023383 | 23.7232 | 0.090928 | 0.15711 | 0.010493 | 0.73443 | 0.73443 | 0.008707
2019-03-29 13:56 | 1553882139.5345 | 152 | 0.20745 | 0.00215 | 2.9e-05 | 101.282 | 100.5237 | 118.382 | 21.1403 | 0.002983 | 0.002728 | 257.4156 | 0.26517 | 8.4021 | 0.002512 | 0.003611 | 0.0035 | 5.2e-05 | 0.001024 | 0.29574 | 1.2946 | 2.5474 | 100.2 | 20.84 | 37.818 | 93.4825 | 82.1236 | 87.3278 | 0.009305 | 0.97938 | 0.023324 | 23.8147 | 0.06722 | 0.15775 | 0.010214 | 0.71306 | 0.71306 | 0.009004
2019-03-29 13:56 | 1553882140.5354 | 153 | 0.20746 | 0.00214 | 2.8e-05 | 101.3894 | 100.4786 | 118.4597 | 21.14 | 0.002989 | 0.002649 | 257.3847 | 0.26477 | 8.4009 | 0.002505 | 0.003705 | 0.003659 | 5.6e-05 | 0.001029 | 0.26897 | 1.3501 | 2.5475 | 100.2 | 20.84 | 37.818 | 93.6223 | 81.954 | 85.4281 | 0.009305 | 0.98457 | 0.023393 | 23.7597 | 0.080715 | 0.15751 | 0.010177 | 0.7126 | 0.7126 | 0.008996
2019-03-29 13:56 | 1553882141.5355 | 154 | 0.20747 | 0.002128 | 3e-05 | 101.3532 | 100.623 | 118.5569 | 21.1396 | 0.002944 | 0.002673 | 257.3511 | 0.26565 | 8.4019 | 0.002509 | 0.003634 | 0.003607 | 5.6e-05 | 0.001026 | 0.30541 | 1.3481 | 2.5474 | 100.2 | 20.84 | 37.809 | 92.5338 | 82.0551 | 86.0085 | 0.009303 | 0.98298 | 0.023384 | 23.7887 | 0.064677 | 0.15804 | 0.01015 | 0.71035 | 0.71035 | 0.009055
2019-03-29 13:56 | 1553882142.5346 | 155 | 0.20748 | 0.002117 | 3e-05 | 101.3801 | 100.6209 | 118.6834 | 21.1396 | 0.002948 | 0.002648 | 257.3141 | 0.26458 | 8.4035 | 0.002562 | 0.003958 | 0.003917 | 5.5e-05 | 0.001028 | 0.30718 | 1.3403 | 2.5474 | 100.2 | 20.84 | 37.809 | 92.6376 | 83.3436 | 85.3972 | 0.009303 | 0.98465 | 0.023416 | 23.7811 | 0.067236 | 0.1574 | 0.010091 | 0.70716 | 0.70716 | 0.009059
2019-03-29 13:56 | 1553882143.5345 | 156 | 0.20752 | 0.002105 | 3.1e-05 | 101.3993 | 100.7725 | 118.6179 | 21.14 | 0.002966 | 0.002658 | 257.2771 | 0.26448 | 8.4023 | 0.002464 | 0.003505 | 0.003415 | 5.5e-05 | 0.001031 | 0.30702 | 1.3248 | 2.5475 | 100.2 | 20.85 | 37.802 | 93.0653 | 80.9762 | 85.642 | 0.009307 | 0.98398 | 0.023402 | 23.7827 | 0.055466 | 0.15734 | 0.009873 | 0.69137 | 0.69137 | 0.009263
2019-03-29 13:56 | 1553882144.5357 | 157 | 0.20752 | 0.002102 | 3.1e-05 | 101.4152 | 100.7151 | 118.1733 | 21.1401 | 0.00294 | 0.002634 | 257.2409 | 0.26526 | 8.4033 | 0.00253 | 0.003755 | 0.003767 | 5.3e-05 | 0.001032 | 0.29734 | 1.3668 | 2.5474 | 100.2 | 20.85 | 37.798 | 92.4238 | 82.5774 | 85.077 | 0.009306 | 0.98554 | 0.023376 | 23.7194 | 0.061956 | 0.15781 | 0.009844 | 0.68861 | 0.68861 | 0.009327
2019-03-29 13:56 | 1553882145.536 | 158 | 0.20747 | 0.002093 | 3.2e-05 | 101.4325 | 100.9059 | 118.1163 | 21.14 | 0.002919 | 0.002624 | 257.1974 | 0.26554 | 8.4026 | 0.002573 | 0.004541 | 0.004384 | 5.7e-05 | 0.001035 | 0.3131 | 1.3486 | 2.5475 | 100.2 | 20.85 | 37.79 | 91.9373 | 83.5898 | 84.8273 | 0.009304 | 0.98622 | 0.023379 | 23.7054 | 0.04656 | 0.15797 | 0.010174 | 0.71177 | 0.71177 | 0.009033
2019-03-29 13:56 | 1553882146.536 | 159 | 0.20754 | 0.002083 | 3.1e-05 | 101.4056 | 100.7684 | 118.3719 | 21.1393 | 0.002971 | 0.00263 | 257.1518 | 0.26458 | 8.4019 | 0.002562 | 0.004246 | 0.004216 | 5.8e-05 | 0.001039 | 0.31022 | 1.3692 | 2.5474 | 100.2 | 20.85 | 37.788 | 93.1772 | 83.3354 | 84.9622 | 0.009303 | 0.98585 | 0.0234 | 23.7355 | 0.056384 | 0.1574 | 0.009752 | 0.68279 | 0.68279 | 0.009382
2019-03-29 13:56 | 1553882147.5357 | 160 | 0.2075 | 0.002074 | 3.1e-05 | 101.478 | 100.7677 | 118.7928 | 21.1396 | 0.002945 | 0.00264 | 257.1407 | 0.26589 | 8.4017 | 0.002501 | 0.00358 | 0.003605 | 5.7e-05 | 0.001043 | 0.30525 | 1.377 | 2.5475 | 100.2 | 20.85 | 37.785 | 92.5596 | 81.8765 | 85.2201 | 0.009303 | 0.98514 | 0.023433 | 23.7862 | 0.062832 | 0.15818 | 0.010034 | 0.70359 | 0.70359 | 0.00915
2019-03-29 13:56 | 1553882148.5348 | 161 | 0.20753 | 0.002065 | 3.1e-05 | 101.4861 | 100.9622 | 119.1362 | 21.1403 | 0.002954 | 0.002593 | 257.112 | 0.26523 | 8.4024 | 0.002538 | 0.004577 | 0.004431 | 5.7e-05 | 0.001046 | 0.31454 | 1.411 | 2.5473 | 100.2 | 20.85 | 37.778 | 92.7714 | 82.7545 | 84.0777 | 0.009301 | 0.98829 | 0.023504 | 23.7825 | 0.046293 | 0.15779 | 0.009834 | 0.69166 | 0.69166 | 0.009285
2019-03-29 13:56 | 1553882149.536 | 162 | 0.20755 | 0.002054 | 3e-05 | 101.5199 | 100.9028 | 119.1104 | 21.1389 | 0.002922 | 0.002652 | 257.1058 | 0.26427 | 8.403 | 0.002554 | 0.003931 | 0.003953 | 5.8e-05 | 0.001044 | 0.31078 | 1.363 | 2.5474 | 100.2 | 20.85 | 37.774 | 92.0004 | 83.1526 | 85.5012 | 0.0093 | 0.98437 | 0.023455 | 23.8273 | 0.054536 | 0.15721 | 0.009735 | 0.6833 | 0.6833 | 0.009364
2019-03-29 13:56 | 1553882150.5358 | 163 | 0.20756 | 0.002053 | 3e-05 | 101.5065 | 101.08 | 119.1834 | 21.1408 | 0.002953 | 0.002639 | 257.0687 | 0.26503 | 8.4022 | 0.002531 | 0.003602 | 0.003543 | 5.9e-05 | 0.001043 | 0.30397 | 1.3606 | 2.5475 | 100.2 | 20.85 | 37.768 | 92.7514 | 82.5897 | 85.1877 | 0.009299 | 0.98523 | 0.023472 | 23.8242 | 0.037661 | 0.15767 | 0.009707 | 0.68184 | 0.68184 | 0.009411
2019-03-29 13:56 | 1553882151.5357 | 164 | 0.20757 | 0.002038 | 2.9e-05 | 101.6136 | 101.2462 | 119.1815 | 21.1403 | 0.002884 | 0.002665 | 257.0548 | 0.26538 | 8.4018 | 0.002569 | 0.004753 | 0.004621 | 6.1e-05 | 0.001034 | 0.30773 | 1.3445 | 2.5474 | 100.2 | 20.85 | 37.765 | 91.0769 | 83.5081 | 85.8116 | 0.009298 | 0.98352 | 0.023452 | 23.8447 | 0.032404 | 0.15788 | 0.009634 | 0.67615 | 0.67615 | 0.009503
2019-03-29 13:56 | 1553882152.5353 | 165 | 0.20758 | 0.002032 | 2.9e-05 | 101.5391 | 101.111 | 119.2665 | 21.1404 | 0.002938 | 0.002651 | 257.0293 | 0.26491 | 8.4027 | 0.002535 | 0.004362 | 0.004274 | 6.3e-05 | 0.001029 | 0.31924 | 1.3575 | 2.5475 | 100.2 | 20.85 | 37.762 | 92.3843 | 82.6938 | 85.4783 | 0.009297 | 0.98443 | 0.023471 | 23.8421 | 0.037788 | 0.1576 | 0.009584 | 0.67324 | 0.67324 | 0.009527
2019-03-29 13:56 | 1553882153.5339 | 166 | 0.20761 | 0.002027 | 2.9e-05 | 101.5711 | 101.2226 | 119.3115 | 21.1402 | 0.002893 | 0.002642 | 257.0098 | 0.26407 | 8.4028 | 0.00254 | 0.004507 | 0.004493 | 6.4e-05 | 0.001032 | 0.30342 | 1.344 | 2.5473 | 100.2 | 20.85 | 37.755 | 91.3035 | 82.817 | 85.266 | 0.009295 | 0.98502 | 0.023482 | 23.8396 | 0.030742 | 0.1571 | 0.009405 | 0.66099 | 0.66099 | 0.009673
2019-03-29 13:56 | 1553882154.5349 | 167 | 0.20764 | 0.00202 | 2.9e-05 | 101.6199 | 101.2192 | 119.5153 | 21.1393 | 0.002922 | 0.002633 | 256.9731 | 0.2642 | 8.4027 | 0.002499 | 0.004223 | 0.004059 | 6e-05 | 0.001031 | 0.33251 | 1.3515 | 2.5474 | 100.2 | 20.85 | 37.752 | 91.9875 | 81.8192 | 85.052 | 0.009295 | 0.9856 | 0.023509 | 23.8528 | 0.035339 | 0.15717 | 0.009228 | 0.64928 | 0.64928 | 0.009852
2019-03-29 13:56 | 1553882155.5356 | 168 | 0.20765 | 0.002008 | 2.8e-05 | 101.5831 | 101.3137 | 119.7307 | 21.1411 | 0.00293 | 0.002587 | 256.9479 | 0.26542 | 8.4025 | 0.002525 | 0.004287 | 0.004209 | 6.3e-05 | 0.001028 | 0.31813 | 1.3339 | 2.5474 | 100.2 | 20.85 | 37.747 | 92.1903 | 82.4534 | 83.9294 | 0.009293 | 0.9887 | 0.023568 | 23.8369 | 0.023748 | 0.1579 | 0.009216 | 0.65007 | 0.65007 | 0.009886
2019-03-29 13:56 | 1553882156.535 | 169 | 0.20765 | 0.002002 | 2.8e-05 | 101.6352 | 101.3346 | 119.8361 | 21.1401 | 0.002907 | 0.002639 | 256.929 | 0.26452 | 8.4016 | 0.002502 | 0.004267 | 0.004299 | 6.3e-05 | 0.001029 | 0.33339 | 1.3606 | 2.5473 | 100.2 | 20.85 | 37.743 | 91.6441 | 81.9019 | 85.18 | 0.009292 | 0.98525 | 0.023537 | 23.8891 | 0.026495 | 0.15737 | 0.009227 | 0.65 | 0.65 | 0.009853
2019-03-29 13:56 | 1553882157.534 | 170 | 0.20766 | 0.001992 | 2.7e-05 | 101.6456 | 101.5328 | 119.7936 | 21.1394 | 0.002922 | 0.002607 | 256.8756 | 0.26381 | 8.4046 | 0.002568 | 0.004356 | 0.004122 | 6.2e-05 | 0.001026 | 0.30127 | 1.375 | 2.5475 | 100.2 | 20.85 | 37.74 | 92.0026 | 83.4697 | 84.4193 | 0.009292 | 0.98735 | 0.023558 | 23.8595 | 0.009932 | 0.15694 | 0.009164 | 0.64618 | 0.64618 | 0.009885
2019-03-29 13:56 | 1553882158.5344 | 171 | 0.20768 | 0.001981 | 2.7e-05 | 101.6576 | 101.3346 | 119.5599 | 21.1407 | 0.0029 | 0.00264 | 256.8514 | 0.2651 | 8.4039 | 0.002472 | 0.00383 | 0.003803 | 6.3e-05 | 0.001019 | 0.31293 | 1.4317 | 2.5475 | 100.2 | 20.85 | 37.736 | 91.4799 | 81.1757 | 85.215 | 0.009291 | 0.98516 | 0.023508 | 23.8627 | 0.02847 | 0.15771 | 0.00904 | 0.63607 | 0.63607 | 0.010091
2019-03-29 13:56 | 1553882159.5354 | 172 | 0.20768 | 0.001971 | 2.7e-05 | 101.6334 | 101.4298 | 119.5101 | 21.1401 | 0.002888 | 0.002621 | 256.8199 | 0.26564 | 8.4001 | 0.002481 | 0.004165 | 0.004059 | 6.4e-05 | 0.001025 | 0.31598 | 1.3501 | 2.5473 | 100.2 | 20.85 | 37.735 | 91.1837 | 81.3871 | 84.7468 | 0.00929 | 0.98645 | 0.023519 | 23.8421 | 0.017936 | 0.15803 | 0.009102 | 0.64073 | 0.64073 | 0.010038
2019-03-29 13:56 | 1553882160.5351 | 173 | 0.20769 | 0.001959 | 2.7e-05 | 101.6301 | 101.4813 | 119.6141 | 21.1391 | 0.002857 | 0.002595 | 256.8007 | 0.26454 | 8.4027 | 0.002405 | 0.003783 | 0.003722 | 6.4e-05 | 0.001025 | 0.27433 | 1.3428 | 2.5474 | 100.2 | 20.85 | 37.733 | 90.4277 | 79.5695 | 84.1225 | 0.00929 | 0.98817 | 0.02355 | 23.8317 | 0.013103 | 0.15737 | 0.009056 | 0.63838 | 0.63838 | 0.010033
2019-03-29 13:56 | 1553882161.5342 | 174 | 0.2077 | 0.00195 | 2.6e-05 | 101.6103 | 101.3281 | 119.4904 | 21.1398 | 0.002908 | 0.002589 | 256.7725 | 0.26545 | 8.4001 | 0.002461 | 0.003851 | 0.003821 | 6.5e-05 | 0.001026 | 0.31349 | 1.3508 | 2.5474 | 100.2 | 20.85 | 37.733 | 91.6643 | 80.9097 | 83.9726 | 0.00929 | 0.98858 | 0.023542 | 23.8144 | 0.024881 | 0.15792 | 0.008966 | 0.63188 | 0.63188 | 0.010172
2019-03-29 13:56 | 1553882162.5352 | 175 | 0.20772 | 0.001937 | 2.6e-05 | 101.6316 | 101.413 | 119.4451 | 21.1399 | 0.00285 | 0.002586 | 256.7466 | 0.26483 | 8.402 | 0.002435 | 0.004016 | 0.003947 | 6.4e-05 | 0.001028 | 0.29679 | 1.389 | 2.5474 | 100.2 | 20.86 | 37.729 | 90.2574 | 80.2813 | 83.904 | 0.009295 | 0.98877 | 0.02354 | 23.8076 | 0.019253 | 0.15755 | 0.008857 | 0.62409 | 0.62409 | 0.010275
2019-03-29 13:56 | 1553882163.5343 | 176 | 0.20769 | 0.001933 | 2.6e-05 | 101.6522 | 101.4243 | 119.5721 | 21.1405 | 0.002813 | 0.002594 | 256.7111 | 0.26475 | 8.4034 | 0.002445 | 0.004894 | 0.004824 | 6.5e-05 | 0.001026 | 0.30934 | 1.3642 | 2.5473 | 100.2 | 20.86 | 37.727 | 89.371 | 80.5165 | 84.0967 | 0.009294 | 0.98824 | 0.023546 | 23.8266 | 0.020077 | 0.1575 | 0.009043 | 0.63738 | 0.63738 | 0.010057
2019-03-29 13:56 | 1553882164.5344 | 177 | 0.2077 | 0.001923 | 2.6e-05 | 101.6864 | 101.3755 | 119.6491 | 21.1406 | 0.002902 | 0.002599 | 256.6772 | 0.26483 | 8.4011 | 0.002462 | 0.00436 | 0.00429 | 6.6e-05 | 0.001021 | 0.33251 | 1.3341 | 2.5475 | 100.2 | 20.86 | 37.726 | 91.515 | 80.9314 | 84.2168 | 0.009294 | 0.98791 | 0.02355 | 23.8383 | 0.027384 | 0.15755 | 0.009026 | 0.6363 | 0.6363 | 0.010077
2019-03-29 13:56 | 1553882165.5353 | 178 | 0.20772 | 0.001914 | 2.7e-05 | 101.6307 | 101.4956 | 119.3742 | 21.1401 | 0.002867 | 0.002599 | 256.6422 | 0.2646 | 8.405 | 0.002538 | 0.00409 | 0.004001 | 6e-05 | 0.001016 | 0.29839 | 1.3467 | 2.5475 | 100.2 | 20.86 | 37.725 | 90.6662 | 82.7467 | 84.2218 | 0.009294 | 0.98789 | 0.023523 | 23.8111 | 0.011892 | 0.15741 | 0.008888 | 0.62579 | 0.62579 | 0.010238
2019-03-29 13:56 | 1553882166.5343 | 179 | 0.20775 | 0.0019 | 2.7e-05 | 101.6801 | 101.6031 | 119.0153 | 21.1406 | 0.00284 | 0.002568 | 256.6192 | 0.26508 | 8.4007 | 0.00243 | 0.004844 | 0.00478 | 6.5e-05 | 0.001019 | 0.32139 | 1.3657 | 2.5474 | 100.2 | 20.86 | 37.724 | 90.0298 | 80.1549 | 83.4827 | 0.009293 | 0.98994 | 0.023512 | 23.7507 | 0.006769 | 0.1577 | 0.008774 | 0.61741 | 0.61741 | 0.010395
2019-03-29 13:56 | 1553882167.5344 | 180 | 0.20773 | 0.001899 | 2.8e-05 | 101.6989 | 101.6175 | 118.928 | 21.1406 | 0.002834 | 0.002598 | 256.5913 | 0.26483 | 8.4018 | 0.002505 | 0.003529 | 0.003547 | 6.4e-05 | 0.001018 | 0.34075 | 1.3457 | 2.5474 | 100.2 | 20.86 | 37.72 | 89.8869 | 81.9733 | 84.1997 | 0.009292 | 0.98796 | 0.02348 | 23.7658 | 0.007166 | 0.15755 | 0.008862 | 0.62272 | 0.62272 | 0.010297
2019-03-29 13:56 | 1553882168.5345 | 181 | 0.20774 | 0.00189 | 2.8e-05 | 101.6768 | 101.4898 | 119.1472 | 21.1409 | 0.002904 | 0.002602 | 256.5382 | 0.26478 | 8.4036 | 0.002393 | 0.004554 | 0.004408 | 6e-05 | 0.001021 | 0.29607 | 1.3312 | 2.5474 | 100.2 | 20.86 | 37.719 | 91.57 | 79.275 | 84.2967 | 0.009292 | 0.98769 | 0.023498 | 23.7909 | 0.01646 | 0.15752 | 0.008851 | 0.62248 | 0.62248 | 0.010299
2019-03-29 13:56 | 1553882169.5333 | 182 | 0.20776 | 0.001881 | 2.8e-05 | 101.6268 | 101.3878 | 119.0918 | 21.141 | 0.00285 | 0.002558 | 256.5097 | 0.26444 | 8.4022 | 0.002508 | 0.004412 | 0.004386 | 6.3e-05 | 0.001017 | 0.31382 | 1.3676 | 2.5475 | 100.2 | 20.86 | 37.717 | 90.2561 | 82.0377 | 83.2405 | 0.009292 | 0.99061 | 0.023527 | 23.7502 | 0.021059 | 0.15732 | 0.008739 | 0.6153 | 0.6153 | 0.010406
2019-03-29 13:56 | 1553882170.5334 | 183 | 0.20777 | 0.001879 | 2.8e-05 | 101.6872 | 101.57 | 118.8562 | 21.1392 | 0.002789 | 0.002618 | 256.468 | 0.26512 | 8.4031 | 0.00242 | 0.003988 | 0.003855 | 6.5e-05 | 0.001015 | 0.33339 | 1.3167 | 2.5474 | 100.2 | 20.87 | 37.712 | 88.801 | 79.9312 | 84.6913 | 0.009296 | 0.9866 | 0.023456 | 23.775 | 0.010315 | 0.15772 | 0.008659 | 0.60787 | 0.60787 | 0.01056
2019-03-29 13:56 | 1553882171.5334 | 184 | 0.20782 | 0.00188 | 2.8e-05 | 101.7127 | 101.4223 | 118.977 | 21.1401 | 0.002858 | 0.002578 | 256.4358 | 0.26484 | 8.4023 | 0.002516 | 0.003964 | 0.003941 | 6.3e-05 | 0.001008 | 0.31852 | 1.4066 | 2.5474 | 100.2 | 20.87 | 37.71 | 90.4491 | 82.2209 | 83.7077 | 0.009296 | 0.98932 | 0.023501 | 23.7543 | 0.025577 | 0.15755 | 0.008378 | 0.58916 | 0.58916 | 0.010884
2019-03-29 13:56 | 1553882172.5345 | 185 | 0.20783 | 0.001875 | 2.8e-05 | 101.7016 | 101.6089 | 119.3714 | 21.1408 | 0.00283 | 0.002607 | 256.4096 | 0.26542 | 8.4025 | 0.00242 | 0.003386 | 0.003362 | 6.4e-05 | 0.001002 | 0.31797 | 1.3347 | 2.5475 | 100.2 | 20.87 | 37.707 | 89.7949 | 79.9306 | 84.4221 | 0.009295 | 0.98734 | 0.023516 | 23.8175 | 0.008154 | 0.1579 | 0.008321 | 0.58552 | 0.58552 | 0.010976
2019-03-29 13:56 | 1553882173.5334 | 186 | 0.20781 | 0.001872 | 2.7e-05 | 101.6573 | 101.5294 | 119.484 | 21.1405 | 0.002849 | 0.002531 | 256.3695 | 0.26444 | 8.4045 | 0.002509 | 0.003968 | 0.003865 | 6.3e-05 | 0.001 | 0.3147 | 1.4106 | 2.5476 | 100.2 | 20.87 | 37.701 | 90.2313 | 82.0629 | 82.5817 | 0.009294 | 0.99245 | 0.023588 | 23.7673 | 0.011259 | 0.15731 | 0.008403 | 0.59318 | 0.59318 | 0.010794
2019-03-29 13:56 | 1553882174.5335 | 187 | 0.20781 | 0.001866 | 2.7e-05 | 101.6977 | 101.6864 | 119.5488 | 21.141 | 0.002856 | 0.002579 | 256.3317 | 0.26558 | 8.4038 | 0.002473 | 0.003884 | 0.00379 | 6.3e-05 | 0.000996 | 0.32477 | 1.3507 | 2.5473 | 100.2 | 20.87 | 37.693 | 90.4076 | 81.1868 | 83.7374 | 0.009292 | 0.98923 | 0.023556 | 23.8123 | 0.000994 | 0.15799 | 0.008407 | 0.59266 | 0.59266 | 0.01085
2019-03-29 13:56 | 1553882175.5346 | 188 | 0.20786 | 0.001863 | 2.7e-05 | 101.7253 | 101.4117 | 119.5979 | 21.1414 | 0.00283 | 0.002522 | 256.3024 | 0.26537 | 8.4011 | 0.002471 | 0.00342 | 0.003325 | 6.5e-05 | 0.000993 | 0.30901 | 1.3519 | 2.5475 | 100.2 | 20.87 | 37.693 | 89.7906 | 81.1367 | 82.3712 | 0.009292 | 0.99304 | 0.023606 | 23.7716 | 0.027615 | 0.15787 | 0.008143 | 0.57527 | 0.57527 | 0.011169
2019-03-29 13:56 | 1553882176.5337 | 189 | 0.20786 | 0.001853 | 2.7e-05 | 101.7199 | 101.8335 | 119.7934 | 21.1412 | 0.002772 | 0.002531 | 256.2744 | 0.26513 | 8.4003 | 0.002392 | 0.003787 | 0.003719 | 6.4e-05 | 0.000989 | 0.28313 | 1.3493 | 2.5475 | 100.2 | 20.87 | 37.689 | 88.3961 | 79.2488 | 82.6011 | 0.009291 | 0.99239 | 0.023618 | 23.7987 | -0.009987 | 0.15773 | 0.008148 | 0.57585 | 0.57585 | 0.011148
2019-03-29 13:56 | 1553882177.534 | 190 | 0.20791 | 0.001843 | 2.7e-05 | 101.7208 | 101.681 | 119.9702 | 21.1396 | 0.002835 | 0.002511 | 256.2447 | 0.26458 | 8.4012 | 0.002413 | 0.00398 | 0.003946 | 6.5e-05 | 0.000996 | 0.31885 | 1.336 | 2.5474 | 100.2 | 20.87 | 37.687 | 89.9069 | 79.7454 | 82.1166 | 0.00929 | 0.99375 | 0.023651 | 23.8 | 0.003501 | 0.1574 | 0.007876 | 0.5574 | 0.5574 | 0.011493
2019-03-29 13:56 | 1553882178.5347 | 191 | 0.20789 | 0.001834 | 2.6e-05 | 101.7363 | 101.4554 | 119.9826 | 21.1411 | 0.00283 | 0.002556 | 256.2223 | 0.26539 | 8.4025 | 0.002445 | 0.003912 | 0.003853 | 6.5e-05 | 0.001001 | 0.31404 | 1.3836 | 2.5473 | 100.2 | 20.87 | 37.681 | 89.7872 | 80.5326 | 83.1998 | 0.009289 | 0.99073 | 0.023616 | 23.8375 | 0.024737 | 0.15788 | 0.008004 | 0.56571 | 0.56571 | 0.011359
2019-03-29 13:56 | 1553882179.5337 | 192 | 0.20789 | 0.001828 | 2.6e-05 | 101.6534 | 101.3861 | 119.8889 | 21.1405 | 0.002844 | 0.00254 | 256.1759 | 0.26517 | 8.4039 | 0.002442 | 0.004002 | 0.003853 | 6.2e-05 | 0.001003 | 0.32803 | 1.3266 | 2.5472 | 100.2 | 20.88 | 37.678 | 90.1321 | 80.4449 | 82.8017 | 0.009294 | 0.99184 | 0.02362 | 23.8149 | 0.023554 | 0.15775 | 0.008027 | 0.56744 | 0.56744 | 0.011315
2019-03-29 13:56 | 1553882180.5337 | 193 | 0.20787 | 0.001823 | 2.5e-05 | 101.6887 | 101.582 | 119.9281 | 21.141 | 0.002789 | 0.002603 | 256.1542 | 0.26447 | 8.4031 | 0.002392 | 0.003825 | 0.00368 | 6.5e-05 | 0.001002 | 0.28191 | 1.3439 | 2.5475 | 100.2 | 20.88 | 37.672 | 88.809 | 79.2539 | 84.3326 | 0.009292 | 0.98759 | 0.023574 | 23.87 | 0.009395 | 0.15733 | 0.008136 | 0.5741 | 0.5741 | 0.011154
2019-03-29 13:56 | 1553882181.5348 | 194 | 0.20791 | 0.001814 | 2.6e-05 | 101.6651 | 101.5072 | 119.7037 | 21.1411 | 0.002774 | 0.002546 | 256.1289 | 0.26394 | 8.4041 | 0.002393 | 0.004182 | 0.004151 | 6.4e-05 | 0.001006 | 0.28871 | 1.3636 | 2.5474 | 100.2 | 20.88 | 37.669 | 88.4405 | 79.2727 | 82.9496 | 0.009291 | 0.99142 | 0.023597 | 23.8015 | 0.013899 | 0.15702 | 0.007912 | 0.55873 | 0.55873 | 0.011438
2019-03-29 13:56 | 1553882182.5348 | 195 | 0.20786 | 0.001809 | 2.6e-05 | 101.7459 | 101.6352 | 119.4411 | 21.1398 | 0.002829 | 0.002551 | 256.0777 | 0.2643 | 8.4013 | 0.002368 | 0.003886 | 0.003816 | 6.4e-05 | 0.001008 | 0.30807 | 1.3589 | 2.5475 | 100.2 | 20.88 | 37.666 | 89.7546 | 78.6798 | 83.0582 | 0.009291 | 0.99112 | 0.023568 | 23.779 | 0.009736 | 0.15723 | 0.008195 | 0.57808 | 0.57808 | 0.01107
2019-03-29 13:56 | 1553882183.5348 | 196 | 0.20786 | 0.001806 | 2.6e-05 | 101.6525 | 101.5363 | 119.6152 | 21.1405 | 0.002854 | 0.002529 | 256.0376 | 0.26388 | 8.4032 | 0.002407 | 0.004194 | 0.00414 | 6.4e-05 | 0.001007 | 0.28335 | 1.3598 | 2.5475 | 100.2 | 20.88 | 37.661 | 90.3618 | 79.6004 | 82.5411 | 0.009289 | 0.99256 | 0.023602 | 23.779 | 0.010237 | 0.15698 | 0.008203 | 0.57944 | 0.57944 | 0.011026
2019-03-29 13:56 | 1553882184.5333 | 197 | 0.20785 | 0.001799 | 2.6e-05 | 101.6277 | 101.6926 | 119.7789 | 21.1399 | 0.002825 | 0.002557 | 256.0121 | 0.26442 | 8.4019 | 0.002444 | 0.0035 | 0.003436 | 6.2e-05 | 0.001004 | 0.31725 | 1.3439 | 2.5474 | 100.2 | 20.88 | 37.658 | 89.6697 | 80.4957 | 83.2206 | 0.009289 | 0.99067 | 0.023596 | 23.818 | -0.005709 | 0.1573 | 0.008271 | 0.5841 | 0.5841 | 0.010961
2019-03-29 13:56 | 1553882185.533 | 198 | 0.20785 | 0.001795 | 2.7e-05 | 101.6555 | 101.7424 | 120.0263 | 21.1399 | 0.00276 | 0.002574 | 255.9794 | 0.26509 | 8.4034 | 0.002394 | 0.003763 | 0.003693 | 5.7e-05 | 0.001009 | 0.30071 | 1.3943 | 2.5474 | 100.2 | 20.88 | 37.652 | 88.0894 | 79.299 | 83.621 | 0.009287 | 0.98956 | 0.023607 | 23.856 | -0.007641 | 0.1577 | 0.008289 | 0.58562 | 0.58562 | 0.01096
2019-03-29 13:56 | 1553882186.5329 | 199 | 0.20786 | 0.001792 | 2.6e-05 | 101.6968 | 101.7574 | 120.1467 | 21.1402 | 0.002797 | 0.002561 | 255.943 | 0.26482 | 8.4028 | 0.002408 | 0.004138 | 0.004109 | 5.9e-05 | 0.00101 | 0.27544 | 1.3783 | 2.5474 | 100.2 | 20.88 | 37.649 | 88.9965 | 79.6344 | 83.3146 | 0.009286 | 0.99041 | 0.023629 | 23.8577 | -0.005327 | 0.15754 | 0.008208 | 0.58043 | 0.58043 | 0.011047
2019-03-29 13:56 | 1553882187.533 | 200 | 0.20793 | 0.001783 | 2.6e-05 | 101.7139 | 101.7124 | 120.115 | 21.1401 | 0.002838 | 0.002574 | 255.9168 | 0.26617 | 8.4012 | 0.002439 | 0.00329 | 0.003215 | 6e-05 | 0.001008 | 0.30525 | 1.3822 | 2.5473 | 100.2 | 20.88 | 37.648 | 89.9732 | 80.3758 | 83.6337 | 0.009286 | 0.98952 | 0.023615 | 23.8652 | 0.000135 | 0.15834 | 0.007827 | 0.55313 | 0.55313 | 0.011651
2019-03-29 13:56 | 1553882188.534 | 201 | 0.20792 | 0.001783 | 2.6e-05 | 101.6902 | 101.6683 | 119.7904 | 21.1393 | 0.002796 | 0.002554 | 255.8874 | 0.26397 | 8.4054 | 0.002448 | 0.004124 | 0.00406 | 5.9e-05 | 0.001011 | 0.31365 | 1.3728 | 2.5475 | 100.2 | 20.89 | 37.645 | 88.973 | 80.5997 | 83.1303 | 0.009291 | 0.99092 | 0.0236 | 23.8161 | 0.001927 | 0.15704 | 0.007889 | 0.55712 | 0.55712 | 0.011472
2019-03-29 13:56 | 1553882189.5343 | 202 | 0.20791 | 0.001779 | 2.7e-05 | 101.6744 | 101.5117 | 119.6878 | 21.1395 | 0.002797 | 0.00255 | 255.8491 | 0.2645 | 8.4027 | 0.002478 | 0.004344 | 0.0042 | 5.5e-05 | 0.001017 | 0.28943 | 1.3613 | 2.5475 | 100.2 | 20.89 | 37.643 | 88.9951 | 81.3094 | 83.0423 | 0.009291 | 0.99116 | 0.023593 | 23.803 | 0.014324 | 0.15735 | 0.007926 | 0.55954 | 0.55954 | 0.011445
2019-03-29 13:56 | 1553882190.5345 | 203 | 0.20796 | 0.001772 | 2.7e-05 | 101.7543 | 101.4854 | 119.6732 | 21.139 | 0.002797 | 0.002558 | 255.7951 | 0.26345 | 8.4024 | 0.002497 | 0.004136 | 0.004124 | 5.9e-05 | 0.001021 | 0.2896 | 1.3257 | 2.5475 | 100.2 | 20.89 | 37.643 | 88.9845 | 81.7706 | 83.2401 | 0.009291 | 0.99062 | 0.023585 | 23.8081 | 0.023668 | 0.15673 | 0.007662 | 0.54073 | 0.54073 | 0.011796
2019-03-29 13:56 | 1553882191.5341 | 204 | 0.20793 | 0.001767 | 2.7e-05 | 101.7363 | 101.6851 | 119.7145 | 21.1394 | 0.002759 | 0.002486 | 255.7487 | 0.26377 | 8.3998 | 0.002407 | 0.004199 | 0.004144 | 5.6e-05 | 0.001013 | 0.29535 | 1.3068 | 2.5474 | 100.2 | 20.89 | 37.641 | 88.0718 | 79.6205 | 81.5121 | 0.00929 | 0.99544 | 0.023646 | 23.7544 | 0.00451 | 0.15692 | 0.007825 | 0.55364 | 0.55364 | 0.011535
2019-03-29 13:57 | 1553882192.5331 | 205 | 0.20796 | 0.00176 | 2.7e-05 | 101.6447 | 101.6885 | 119.6063 | 21.1392 | 0.002753 | 0.002534 | 255.7303 | 0.26331 | 8.4033 | 0.002448 | 0.00476 | 0.004688 | 5.6e-05 | 0.001013 | 0.29214 | 1.331 | 2.5474 | 100.2 | 20.89 | 37.639 | 87.9254 | 80.5957 | 82.6537 | 0.00929 | 0.99225 | 0.023598 | 23.7819 | -0.003847 | 0.15664 | 0.00765 | 0.54017 | 0.54017 | 0.011802
2019-03-29 13:57 | 1553882193.5345 | 206 | 0.2079 | 0.001759 | 2.7e-05 | 101.7196 | 101.6905 | 119.319 | 21.1397 | 0.002758 | 0.00253 | 255.694 | 0.26443 | 8.4024 | 0.002419 | 0.004388 | 0.004313 | 5.7e-05 | 0.001012 | 0.28816 | 1.3238 | 2.5474 | 100.2 | 20.88 | 37.633 | 88.0456 | 79.9046 | 82.5684 | 0.009282 | 0.99249 | 0.023572 | 23.7505 | 0.002556 | 0.15731 | 0.008021 | 0.56581 | 0.56581 | 0.011316
2019-03-29 13:57 | 1553882194.5345 | 207 | 0.20791 | 0.001757 | 2.6e-05 | 101.695 | 101.7813 | 119.2089 | 21.1398 | 0.002721 | 0.002521 | 255.6607 | 0.26482 | 8.4007 | 0.002421 | 0.004537 | 0.00447 | 5.5e-05 | 0.001016 | 0.28102 | 1.3202 | 2.5472 | 100.2 | 20.88 | 37.633 | 87.1537 | 79.9443 | 82.3555 | 0.009282 | 0.99308 | 0.023568 | 23.7324 | -0.007585 | 0.15754 | 0.007958 | 0.5613 | 0.5613 | 0.011423
2019-03-29 13:57 | 1553882195.5332 | 208 | 0.20797 | 0.001751 | 2.7e-05 | 101.6717 | 101.6008 | 119.3446 | 21.1403 | 0.002769 | 0.002536 | 255.6062 | 0.26381 | 8.4011 | 0.002391 | 0.004057 | 0.004005 | 5.6e-05 | 0.001016 | 0.30541 | 1.3443 | 2.5471 | 100.2 | 20.88 | 37.623 | 88.31 | 79.2181 | 82.7077 | 0.00928 | 0.9921 | 0.02357 | 23.7576 | 0.006242 | 0.15694 | 0.007645 | 0.53917 | 0.53917 | 0.011847
2019-03-29 13:57 | 1553882196.5332 | 209 | 0.20799 | 0.001743 | 2.7e-05 | 101.7031 | 101.5892 | 119.3028 | 21.1395 | 0.002719 | 0.00252 | 255.5839 | 0.26405 | 8.4016 | 0.002435 | 0.004042 | 0.003984 | 5.5e-05 | 0.001011 | 0.30596 | 1.3285 | 2.5472 | 100.2 | 20.88 | 37.623 | 87.102 | 80.2824 | 82.3257 | 0.00928 | 0.99316 | 0.023578 | 23.7407 | 0.010028 | 0.15708 | 0.007529 | 0.53117 | 0.53117 | 0.012036
2019-03-29 13:57 | 1553882197.5343 | 210 | 0.20793 | 0.001743 | 2.8e-05 | 101.6693 | 101.6618 | 119.0653 | 21.1395 | 0.002796 | 0.002521 | 255.5525 | 0.26439 | 8.4008 | 0.002493 | 0.003988 | 0.003912 | 5.6e-05 | 0.001012 | 0.3011 | 1.3689 | 2.5474 | 100.2 | 20.88 | 37.619 | 88.9662 | 81.6752 | 82.3552 | 0.009279 | 0.99308 | 0.023554 | 23.7181 | 0.000655 | 0.15729 | 0.007847 | 0.55301 | 0.55301 | 0.011576
2019-03-29 13:57 | 1553882198.5324 | 211 | 0.20793 | 0.001738 | 2.8e-05 | 101.6771 | 101.5731 | 119.1829 | 21.1394 | 0.002798 | 0.002517 | 255.5053 | 0.2642 | 8.4045 | 0.002481 | 0.003833 | 0.003878 | 5.5e-05 | 0.001014 | 0.28943 | 1.3459 | 2.5475 | 100.2 | 20.88 | 37.61 | 89.0065 | 81.3818 | 82.2536 | 0.009277 | 0.99337 | 0.023569 | 23.7264 | 0.00915 | 0.15717 | 0.007897 | 0.55691 | 0.55691 | 0.011486
2019-03-29 13:57 | 1553882199.5333 | 212 | 0.20799 | 0.00174 | 2.8e-05 | 101.6441 | 101.5189 | 119.37 | 21.1401 | 0.002752 | 0.002486 | 255.4655 | 0.2638 | 8.4018 | 0.00249 | 0.004278 | 0.00413 | 5.4e-05 | 0.00102 | 0.31741 | 1.3493 | 2.5473 | 100.2 | 20.88 | 37.605 | 87.895 | 81.5969 | 81.5174 | 0.009276 | 0.99543 | 0.023612 | 23.7204 | 0.011033 | 0.15693 | 0.007487 | 0.5289 | 0.5289 | 0.012076
2019-03-29 13:57 | 1553882200.5334 | 213 | 0.20794 | 0.001736 | 2.7e-05 | 101.6409 | 101.5543 | 119.5202 | 21.1401 | 0.002708 | 0.002485 | 255.4535 | 0.26476 | 8.4008 | 0.002447 | 0.004178 | 0.00412 | 5.3e-05 | 0.001019 | 0.31116 | 1.3221 | 2.5474 | 100.2 | 20.88 | 37.598 | 86.8347 | 80.5716 | 81.4777 | 0.009274 | 0.99554 | 0.023628 | 23.734 | 0.007616 | 0.1575 | 0.007838 | 0.55417 | 0.55417 | 0.011567
2019-03-29 13:57 | 1553882201.5336 | 214 | 0.20797 | 0.001738 | 2.8e-05 | 101.6468 | 101.6107 | 119.6657 | 21.1397 | 0.0028 | 0.002484 | 255.4148 | 0.26412 | 8.4018 | 0.0025 | 0.004111 | 0.004034 | 5.9e-05 | 0.001016 | 0.28783 | 1.3506 | 2.5474 | 100.2 | 20.88 | 37.585 | 89.0726 | 81.8532 | 81.4516 | 0.009271 | 0.99561 | 0.023643 | 23.7476 | 0.003185 | 0.15712 | 0.007648 | 0.54103 | 0.54103 | 0.01182
2019-03-29 13:57 | 1553882202.5339 | 215 | 0.20799 | 0.001734 | 2.8e-05 | 101.625 | 101.7134 | 120.1438 | 21.1395 | 0.00274 | 0.002532 | 255.3577 | 0.26398 | 8.4029 | 0.002545 | 0.004338 | 0.004266 | 5.8e-05 | 0.001022 | 0.28783 | 1.3619 | 2.5475 | 100.2 | 20.88 | 37.579 | 87.6056 | 82.9355 | 82.611 | 0.009269 | 0.99237 | 0.023652 | 23.8338 | -0.007777 | 0.15704 | 0.007539 | 0.5335 | 0.5335 | 0.011981
2019-03-29 13:57 | 1553882203.5325 | 216 | 0.20797 | 0.001729 | 2.8e-05 | 101.6385 | 101.5663 | 120.5516 | 21.1391 | 0.002766 | 0.002517 | 255.3133 | 0.26553 | 8.4048 | 0.002423 | 0.004119 | 0.004094 | 5.5e-05 | 0.001027 | 0.31365 | 1.3094 | 2.5474 | 100.2 | 20.88 | 37.566 | 88.2423 | 80.0037 | 82.2502 | 0.009266 | 0.99337 | 0.023704 | 23.8621 | 0.006353 | 0.15796 | 0.007673 | 0.54418 | 0.54418 | 0.011814
2019-03-29 13:57 | 1553882204.5325 | 217 | 0.20796 | 0.001728 | 2.8e-05 | 101.5983 | 101.4001 | 120.3518 | 21.1389 | 0.0027 | 0.002508 | 255.2846 | 0.26549 | 8.402 | 0.002427 | 0.003922 | 0.003889 | 5.6e-05 | 0.001026 | 0.32604 | 1.41 | 2.5474 | 100.2 | 20.88 | 37.559 | 86.6509 | 80.0948 | 82.0318 | 0.009264 | 0.99399 | 0.023692 | 23.835 | 0.017472 | 0.15794 | 0.007707 | 0.5463 | 0.5463 | 0.011767
2019-03-29 13:57 | 1553882205.5336 | 218 | 0.20798 | 0.001722 | 2.8e-05 | 101.5822 | 101.41 | 120.195 | 21.139 | 0.002724 | 0.002459 | 255.2818 | 0.26487 | 8.4042 | 0.00249 | 0.003844 | 0.003798 | 6.1e-05 | 0.001031 | 0.31349 | 1.3797 | 2.5474 | 100.2 | 20.88 | 37.552 | 87.2295 | 81.6028 | 80.8551 | 0.009262 | 0.99729 | 0.023716 | 23.78 | 0.015176 | 0.15757 | 0.007609 | 0.53992 | 0.53992 | 0.011878
2019-03-29 13:57 | 1553882206.5326 | 219 | 0.20803 | 0.001715 | 2.7e-05 | 101.6456 | 101.7008 | 120.1157 | 21.1391 | 0.002699 | 0.0025 | 255.2365 | 0.26558 | 8.401 | 0.002459 | 0.004518 | 0.004372 | 5.8e-05 | 0.001027 | 0.29286 | 1.3849 | 2.5474 | 100.2 | 20.88 | 37.541 | 86.6222 | 80.8684 | 81.8385 | 0.00926 | 0.99453 | 0.023675 | 23.8051 | -0.004849 | 0.158 | 0.00733 | 0.5192 | 0.5192 | 0.012385
2019-03-29 13:57 | 1553882207.5336 | 220 | 0.20801 | 0.001714 | 2.8e-05 | 101.5421 | 101.5427 | 120.2694 | 21.1392 | 0.002745 | 0.002533 | 255.1899 | 0.26495 | 8.4017 | 0.002507 | 0.003936 | 0.00382 | 5.7e-05 | 0.001034 | 0.29358 | 1.3795 | 2.5474 | 100.2 | 20.88 | 37.535 | 87.7358 | 82.0203 | 82.6333 | 0.009258 | 0.9923 | 0.023664 | 23.847 | -5.9e-05 | 0.15762 | 0.007418 | 0.52517 | 0.52517 | 0.012215
2019-03-29 13:57 | 1553882208.5327 | 221 | 0.20794 | 0.001706 | 2.8e-05 | 101.4942 | 101.3851 | 120.3884 | 21.1394 | 0.002764 | 0.002516 | 255.1605 | 0.26467 | 8.4037 | 0.002495 | 0.003348 | 0.003356 | 5.7e-05 | 0.001035 | 0.32516 | 1.3601 | 2.5475 | 100.2 | 20.88 | 37.531 | 88.1849 | 81.7203 | 82.228 | 0.009257 | 0.99344 | 0.023689 | 23.8452 | 0.009621 | 0.15745 | 0.007878 | 0.55842 | 0.55842 | 0.011476
2019-03-29 13:57 | 1553882209.5327 | 222 | 0.208 | 0.001703 | 2.9e-05 | 101.4672 | 101.4584 | 120.7771 | 21.1397 | 0.002731 | 0.002488 | 255.1385 | 0.26483 | 8.4022 | 0.002408 | 0.003691 | 0.003581 | 5.8e-05 | 0.001037 | 0.29518 | 1.3353 | 2.5474 | 100.2 | 20.88 | 37.52 | 87.402 | 79.6371 | 81.5663 | 0.009255 | 0.99529 | 0.023749 | 23.8614 | 0.000776 | 0.15755 | 0.007488 | 0.532 | 0.532 | 0.012053
2019-03-29 13:57 | 1553882210.5328 | 223 | 0.20802 | 0.001699 | 3e-05 | 101.4334 | 101.3943 | 120.7053 | 21.1402 | 0.002704 | 0.002464 | 255.0993 | 0.26522 | 8.4066 | 0.002534 | 0.003463 | 0.003388 | 5.6e-05 | 0.001041 | 0.29717 | 1.3211 | 2.5473 | 100.2 | 20.88 | 37.515 | 86.7623 | 82.6578 | 80.9764 | 0.009253 | 0.99695 | 0.023762 | 23.8345 | 0.003452 | 0.15778 | 0.007363 | 0.52329 | 0.52329 | 0.012272
2019-03-29 13:57 | 1553882211.5328 | 224 | 0.20803 | 0.001696 | 3e-05 | 101.4472 | 101.2782 | 120.4898 | 21.1394 | 0.002742 | 0.002454 | 255.0859 | 0.26477 | 8.4009 | 0.002421 | 0.003723 | 0.003628 | 5.6e-05 | 0.001039 | 0.30807 | 1.3268 | 2.5473 | 100.2 | 20.88 | 37.506 | 87.6742 | 79.937 | 80.7471 | 0.009251 | 0.99759 | 0.023748 | 23.8055 | 0.014908 | 0.15751 | 0.007331 | 0.52072 | 0.52072 | 0.012311
2019-03-29 13:57 | 1553882212.5332 | 225 | 0.20807 | 0.00169 | 3e-05 | 101.4116 | 101.2636 | 120.2032 | 21.1383 | 0.002731 | 0.002493 | 255.0776 | 0.26488 | 8.4023 | 0.002446 | 0.004241 | 0.004223 | 5.8e-05 | 0.001036 | 0.29966 | 1.3441 | 2.5474 | 100.2 | 20.88 | 37.503 | 87.4068 | 80.5487 | 81.6639 | 0.00925 | 0.99502 | 0.023689 | 23.808 | 0.013063 | 0.15758 | 0.007094 | 0.50257 | 0.50257 | 0.012761
2019-03-29 13:57 | 1553882213.5338 | 226 | 0.20806 | 0.001691 | 3e-05 | 101.425 | 101.1823 | 120.0626 | 21.1399 | 0.002705 | 0.002494 | 255.045 | 0.26406 | 8.4023 | 0.002515 | 0.003416 | 0.003446 | 5.8e-05 | 0.001039 | 0.30989 | 1.3636 | 2.5475 | 100.2 | 20.88 | 37.497 | 86.7642 | 82.2097 | 81.7114 | 0.009249 | 0.99488 | 0.023674 | 23.7956 | 0.021429 | 0.15709 | 0.007175 | 0.50799 | 0.50799 | 0.012586
2019-03-29 13:57 | 1553882214.5322 | 227 | 0.20802 | 0.001685 | 3e-05 | 101.2891 | 101.1902 | 120.2152 | 21.1386 | 0.002701 | 0.002517 | 255.0169 | 0.26504 | 8.4033 | 0.002452 | 0.003809 | 0.003708 | 6.1e-05 | 0.001041 | 0.29535 | 1.3517 | 2.5474 | 100.2 | 20.88 | 37.485 | 86.6717 | 80.6955 | 82.2521 | 0.009246 | 0.99337 | 0.023671 | 23.8289 | 0.008742 | 0.15767 | 0.00742 | 0.52535 | 0.52535 | 0.012215
2019-03-29 13:57 | 1553882215.5319 | 228 | 0.20803 | 0.001687 | 3e-05 | 101.2781 | 101.1687 | 120.046 | 21.1381 | 0.002744 | 0.002472 | 254.9762 | 0.26459 | 8.4037 | 0.002444 | 0.004506 | 0.00447 | 5.8e-05 | 0.001042 | 0.31094 | 1.3566 | 2.5477 | 100.2 | 20.88 | 37.48 | 87.7058 | 80.4932 | 81.1582 | 0.009245 | 0.99644 | 0.023691 | 23.7754 | 0.009665 | 0.1574 | 0.007308 | 0.51789 | 0.51789 | 0.01237
2019-03-29 13:57 | 1553882216.5329 | 229 | 0.20808 | 0.001685 | 2.9e-05 | 101.2595 | 101.0281 | 120.141 | 21.1385 | 0.002666 | 0.002535 | 254.9433 | 0.26466 | 8.4024 | 0.00247 | 0.004155 | 0.004116 | 5.7e-05 | 0.001037 | 0.274 | 1.3676 | 2.5475 | 100.2 | 20.88 | 37.47 | 85.8287 | 81.1267 | 82.6771 | 0.009242 | 0.99218 | 0.023649 | 23.8358 | 0.020466 | 0.15745 | 0.007026 | 0.49697 | 0.49697 | 0.012894
2019-03-29 13:57 | 1553882217.5329 | 230 | 0.2081 | 0.001678 | 2.9e-05 | 101.155 | 101.037 | 120.2269 | 21.14 | 0.002726 | 0.002458 | 254.9174 | 0.26467 | 8.4019 | 0.002553 | 0.003849 | 0.003791 | 5.9e-05 | 0.001031 | 0.30917 | 1.3174 | 2.5474 | 100.2 | 20.88 | 37.466 | 87.2891 | 83.1192 | 80.8466 | 0.009241 | 0.99731 | 0.023719 | 23.7829 | 0.010448 | 0.15745 | 0.006906 | 0.4899 | 0.4899 | 0.013081
2019-03-29 13:57 | 1553882218.5329 | 231 | 0.20802 | 0.001678 | 2.9e-05 | 101.1293 | 101.0421 | 120.3002 | 21.139 | 0.002651 | 0.002494 | 254.8696 | 0.26407 | 8.4011 | 0.002547 | 0.003527 | 0.003457 | 6.2e-05 | 0.001026 | 0.26322 | 1.386 | 2.5475 | 100.2 | 20.88 | 37.461 | 85.4699 | 82.9778 | 81.6886 | 0.00924 | 0.99495 | 0.023698 | 23.8184 | 0.007718 | 0.15709 | 0.007404 | 0.52484 | 0.52484 | 0.012182
2019-03-29 13:57 | 1553882219.5334 | 232 | 0.20809 | 0.001668 | 2.9e-05 | 101.1012 | 100.9462 | 119.8899 | 21.1385 | 0.00275 | 0.002483 | 254.8356 | 0.26528 | 8.4017 | 0.002438 | 0.003479 | 0.003436 | 5.6e-05 | 0.00102 | 0.3079 | 1.3739 | 2.5474 | 100.2 | 20.88 | 37.453 | 87.8678 | 80.354 | 81.4456 | 0.009238 | 0.99563 | 0.023666 | 23.7696 | 0.013723 | 0.15781 | 0.006983 | 0.49432 | 0.49432 | 0.012994
2019-03-29 13:57 | 1553882220.533 | 233 | 0.20808 | 0.001663 | 2.8e-05 | 101.0293 | 100.9069 | 119.7183 | 21.1394 | 0.002684 | 0.002454 | 254.8099 | 0.26489 | 8.4027 | 0.002455 | 0.004185 | 0.00409 | 5.9e-05 | 0.00101 | 0.30165 | 1.3231 | 2.5473 | 100.2 | 20.88 | 37.449 | 86.275 | 80.7557 | 80.7345 | 0.009237 | 0.99763 | 0.023672 | 23.7288 | 0.010843 | 0.15758 | 0.007072 | 0.50081 | 0.50081 | 0.012806
2019-03-29 13:57 | 1553882221.5323 | 234 | 0.20813 | 0.001658 | 2.9e-05 | 101.0123 | 101.0796 | 119.8502 | 21.138 | 0.002732 | 0.002542 | 254.7768 | 0.26461 | 8.4029 | 0.002394 | 0.004678 | 0.00456 | 5.7e-05 | 0.001009 | 0.33682 | 1.3342 | 2.5474 | 100.2 | 20.88 | 37.445 | 87.4166 | 79.2963 | 82.8576 | 0.009236 | 0.99168 | 0.023615 | 23.8129 | -0.005962 | 0.15742 | 0.006801 | 0.48038 | 0.48038 | 0.013337
2019-03-29 13:57 | 1553882222.5335 | 235 | 0.20813 | 0.001656 | 2.9e-05 | 100.9171 | 100.694 | 119.9033 | 21.1384 | 0.002752 | 0.002487 | 254.7278 | 0.26517 | 8.4011 | 0.002468 | 0.003294 | 0.003259 | 6.1e-05 | 0.001005 | 0.2959 | 1.3299 | 2.5474 | 100.2 | 20.88 | 37.439 | 87.9076 | 81.065 | 81.5188 | 0.009235 | 0.99542 | 0.023665 | 23.7734 | 0.019797 | 0.15775 | 0.006795 | 0.48096 | 0.48096 | 0.013349
2019-03-29 13:57 | 1553882223.5335 | 236 | 0.20811 | 0.00165 | 2.8e-05 | 100.8362 | 100.7428 | 120.1413 | 21.1393 | 0.002682 | 0.002498 | 254.705 | 0.26527 | 8.4016 | 0.002512 | 0.003536 | 0.003473 | 6.1e-05 | 0.001004 | 0.30989 | 1.336 | 2.5474 | 100.2 | 20.88 | 37.436 | 86.221 | 82.1249 | 81.7942 | 0.009234 | 0.99465 | 0.023679 | 23.8062 | 0.008297 | 0.15781 | 0.006893 | 0.48821 | 0.48821 | 0.013156
2019-03-29 13:57 | 1553882224.5332 | 237 | 0.20814 | 0.001638 | 2.8e-05 | 100.7186 | 100.6349 | 120.0547 | 21.1397 | 0.002712 | 0.002545 | 254.6553 | 0.26491 | 8.4031 | 0.002443 | 0.003982 | 0.003826 | 6.2e-05 | 0.001008 | 0.30342 | 1.3489 | 2.5474 | 100.2 | 20.88 | 37.431 | 86.9532 | 80.4784 | 82.9201 | 0.009233 | 0.9915 | 0.023633 | 23.8353 | 0.007436 | 0.1576 | 0.006736 | 0.47613 | 0.47613 | 0.013472
2019-03-29 13:57 | 1553882225.5332 | 238 | 0.20813 | 0.001631 | 2.9e-05 | 100.6758 | 100.452 | 119.965 | 21.1388 | 0.00265 | 0.002478 | 254.6273 | 0.26522 | 8.4032 | 0.002482 | 0.003822 | 0.003795 | 6.3e-05 | 0.001006 | 0.30038 | 1.3337 | 2.5474 | 100.2 | 20.88 | 37.429 | 85.4433 | 81.4108 | 81.3125 | 0.009232 | 0.996 | 0.023678 | 23.7726 | 0.019908 | 0.15778 | 0.006794 | 0.48113 | 0.48113 | 0.013347
2019-03-29 13:57 | 1553882226.5322 | 239 | 0.20815 | 0.001627 | 2.9e-05 | 100.5573 | 100.2981 | 119.8504 | 21.1395 | 0.002666 | 0.002502 | 254.5949 | 0.26471 | 8.4023 | 0.002458 | 0.003928 | 0.003804 | 5.8e-05 | 0.001005 | 0.33445 | 1.353 | 2.5474 | 100.2 | 20.88 | 37.427 | 85.8296 | 80.8306 | 81.8958 | 0.009232 | 0.99437 | 0.023647 | 23.7808 | 0.023086 | 0.15747 | 0.006717 | 0.47506 | 0.47506 | 0.013491
2019-03-29 13:57 | 1553882227.5333 | 240 | 0.20811 | 0.001618 | 2.9e-05 | 100.4007 | 100.1718 | 119.8448 | 21.1395 | 0.002663 | 0.002465 | 254.5418 | 0.26411 | 8.4036 | 0.002409 | 0.004273 | 0.00424 | 6e-05 | 0.00101 | 0.335 | 1.3221 | 2.5475 | 100.2 | 20.88 | 37.426 | 85.7709 | 79.6455 | 80.9997 | 0.009231 | 0.99688 | 0.023676 | 23.7502 | 0.020419 | 0.15712 | 0.006915 | 0.48969 | 0.48969 | 0.013059
2019-03-29 13:57 | 1553882228.5324 | 241 | 0.20811 | 0.001616 | 3e-05 | 100.3313 | 100.1377 | 119.5826 | 21.1388 | 0.002693 | 0.002482 | 254.5209 | 0.26487 | 8.4031 | 0.002404 | 0.004313 | 0.004276 | 5.7e-05 | 0.001009 | 0.30862 | 1.3427 | 2.5474 | 100.2 | 20.88 | 37.425 | 86.4963 | 79.5293 | 81.4055 | 0.009231 | 0.99574 | 0.023637 | 23.7378 | 0.017278 | 0.15757 | 0.006927 | 0.48966 | 0.48966 | 0.013097
2019-03-29 13:57 | 1553882229.5318 | 242 | 0.20816 | 0.00161 | 3e-05 | 100.1125 | 99.8121 | 119.6599 | 21.1397 | 0.002694 | 0.002414 | 254.4844 | 0.26464 | 8.4027 | 0.002319 | 0.003574 | 0.003458 | 6e-05 | 0.001017 | 0.31238 | 1.3827 | 2.5475 | 100.2 | 20.88 | 37.42 | 86.5097 | 77.5034 | 79.7789 | 0.00923 | 1.0003 | 0.023699 | 23.6909 | 0.026879 | 0.15743 | 0.006649 | 0.47119 | 0.47119 | 0.013599
2019-03-29 13:57 | 1553882230.5324 | 243 | 0.20816 | 0.001607 | 3e-05 | 100.0457 | 99.779 | 119.6235 | 21.1385 | 0.002724 | 0.002485 | 254.4502 | 0.26463 | 8.4007 | 0.002437 | 0.003604 | 0.00356 | 5.8e-05 | 0.001025 | 0.28313 | 1.3787 | 2.5474 | 100.2 | 20.88 | 37.418 | 87.23 | 80.3353 | 81.4918 | 0.009229 | 0.9955 | 0.023638 | 23.7447 | 0.023881 | 0.15743 | 0.00667 | 0.47145 | 0.47145 | 0.013591
2019-03-29 13:57 | 1553882231.5317 | 244 | 0.20818 | 0.001603 | 3e-05 | 99.938 | 99.7616 | 119.9126 | 21.1396 | 0.002707 | 0.002474 | 254.4207 | 0.26397 | 8.4041 | 0.002377 | 0.004282 | 0.004225 | 5.7e-05 | 0.001023 | 0.30702 | 1.3742 | 2.5474 | 100.2 | 20.88 | 37.409 | 86.8164 | 78.8909 | 81.2292 | 0.009227 | 0.99624 | 0.023675 | 23.7646 | 0.015802 | 0.15704 | 0.006554 | 0.46398 | 0.46398 | 0.013775
2019-03-29 13:57 | 1553882232.5325 | 245 | 0.20818 | 0.001602 | 3e-05 | 99.7479 | 99.5111 | 119.9004 | 21.1383 | 0.00266 | 0.002485 | 254.3804 | 0.26457 | 8.4018 | 0.002298 | 0.003622 | 0.003523 | 5.7e-05 | 0.001023 | 0.31349 | 1.3451 | 2.5475 | 100.2 | 20.88 | 37.409 | 85.6964 | 76.9828 | 81.4735 | 0.009227 | 0.99555 | 0.023666 | 23.7716 | 0.021262 | 0.15739 | 0.006568 | 0.46473 | 0.46473 | 0.013784
2019-03-29 13:57 | 1553882233.5314 | 246 | 0.20817 | 0.001599 | 3.1e-05 | 99.5566 | 99.1073 | 119.3643 | 21.1382 | 0.002683 | 0.002426 | 254.3363 | 0.26353 | 8.4028 | 0.002454 | 0.00355 | 0.003389 | 5.6e-05 | 0.001018 | 0.29574 | 1.3442 | 2.5474 | 100.2 | 20.88 | 37.406 | 86.2366 | 80.7334 | 80.073 | 0.009226 | 0.9995 | 0.02366 | 23.6715 | 0.040458 | 0.15677 | 0.006629 | 0.46895 | 0.46895 | 0.013606
2019-03-29 13:57 | 1553882234.5315 | 247 | 0.20814 | 0.001595 | 3.1e-05 | 99.3998 | 98.996 | 119.4676 | 21.1381 | 0.002653 | 0.002476 | 254.3074 | 0.26486 | 8.4016 | 0.002414 | 0.003561 | 0.003574 | 5.7e-05 | 0.001021 | 0.28766 | 1.307 | 2.5475 | 100.2 | 20.88 | 37.4 | 85.5206 | 79.772 | 81.2767 | 0.009225 | 0.9961 | 0.02363 | 23.722 | 0.036402 | 0.15756 | 0.006809 | 0.48108 | 0.48108 | 0.01333
2019-03-29 13:57 | 1553882235.5325 | 248 | 0.20818 | 0.001596 | 3.2e-05 | 99.2441 | 98.7677 | 119.784 | 21.1385 | 0.002631 | 0.002429 | 254.2834 | 0.26399 | 8.4005 | 0.002417 | 0.003685 | 0.003518 | 5.8e-05 | 0.001024 | 0.28119 | 1.348 | 2.5474 | 100.2 | 20.88 | 37.396 | 84.9942 | 79.859 | 80.1442 | 0.009224 | 0.9993 | 0.023699 | 23.7155 | 0.043038 | 0.15705 | 0.006537 | 0.46311 | 0.46311 | 0.013802
2019-03-29 13:57 | 1553882236.5325 | 249 | 0.20817 | 0.00159 | 3.2e-05 | 99.0262 | 98.5462 | 119.6512 | 21.1393 | 0.00269 | 0.002494 | 254.2571 | 0.26422 | 8.4032 | 0.002344 | 0.003897 | 0.003867 | 5.5e-05 | 0.001021 | 0.27505 | 1.3207 | 2.5474 | 100.2 | 20.88 | 37.393 | 86.4147 | 78.1024 | 81.6944 | 0.009223 | 0.99493 | 0.023634 | 23.7542 | 0.043458 | 0.15718 | 0.006636 | 0.46888 | 0.46888 | 0.013644
2019-03-29 13:57 | 1553882237.5322 | 250 | 0.20818 | 0.001589 | 3.3e-05 | 98.7679 | 98.3667 | 119.5265 | 21.1384 | 0.002685 | 0.002457 | 254.2311 | 0.26531 | 8.4009 | 0.002423 | 0.004378 | 0.004249 | 5.5e-05 | 0.001013 | 0.29535 | 1.3378 | 2.5473 | 100.2 | 20.89 | 37.384 | 86.2998 | 80.0037 | 80.8086 | 0.009227 | 0.99742 | 0.023651 | 23.7122 | 0.036403 | 0.15783 | 0.006538 | 0.46219 | 0.46219 | 0.013899
2019-03-29 13:57 | 1553882238.5316 | 251 | 0.20818 | 0.001589 | 3.2e-05 | 98.6137 | 98.1759 | 119.918 | 21.1387 | 0.002673 | 0.002472 | 254.193 | 0.26575 | 8.3992 | 0.002361 | 0.003989 | 0.003985 | 5.4e-05 | 0.001012 | 0.28854 | 1.3485 | 2.5474 | 100.2 | 20.89 | 37.379 | 86.0146 | 78.5111 | 81.1762 | 0.009226 | 0.99638 | 0.023677 | 23.7634 | 0.039796 | 0.15809 | 0.006579 | 0.46564 | 0.46564 | 0.013818
2019-03-29 13:57 | 1553882239.532 | 252 | 0.20819 | 0.001585 | 3.2e-05 | 98.3733 | 97.8206 | 120.0295 | 21.1384 | 0.002622 | 0.002431 | 254.1801 | 0.26562 | 8.4009 | 0.002399 | 0.003662 | 0.003622 | 5.6e-05 | 0.001013 | 0.30198 | 1.3802 | 2.5474 | 100.2 | 20.89 | 37.371 | 84.7718 | 79.4152 | 80.1847 | 0.009224 | 0.99918 | 0.023722 | 23.7411 | 0.050396 | 0.15801 | 0.006525 | 0.46269 | 0.46269 | 0.0139
2019-03-29 13:57 | 1553882240.5317 | 253 | 0.20817 | 0.001579 | 3.2e-05 | 98.192 | 97.5415 | 120.1746 | 21.1379 | 0.002623 | 0.002399 | 254.1517 | 0.26483 | 8.4036 | 0.002453 | 0.003847 | 0.0039 | 5.7e-05 | 0.00101 | 0.28766 | 1.3603 | 2.5474 | 100.2 | 20.89 | 37.367 | 84.7947 | 80.7098 | 79.4107 | 0.009223 | 1.0014 | 0.023762 | 23.7294 | 0.059451 | 0.15755 | 0.006634 | 0.47123 | 0.47123 | 0.013607
2019-03-29 13:57 | 1553882241.5317 | 254 | 0.20817 | 0.001576 | 3.1e-05 | 97.8971 | 97.3886 | 120.1345 | 21.1387 | 0.002632 | 0.002457 | 254.1105 | 0.26514 | 8.402 | 0.002417 | 0.003753 | 0.003554 | 5.7e-05 | 0.001008 | 0.26753 | 1.3432 | 2.5475 | 100.2 | 20.89 | 37.362 | 85.0155 | 79.8598 | 80.8211 | 0.009221 | 0.99738 | 0.023711 | 23.7729 | 0.046584 | 0.15773 | 0.006644 | 0.47101 | 0.47101 | 0.01363
2019-03-29 13:57 | 1553882242.5321 | 255 | 0.20819 | 0.001573 | 3.1e-05 | 97.6816 | 97.2749 | 120.122 | 21.1379 | 0.002682 | 0.002436 | 254.0925 | 0.26494 | 8.4037 | 0.00241 | 0.00487 | 0.004909 | 5.6e-05 | 0.001009 | 0.29983 | 1.3596 | 2.5474 | 100.2 | 20.89 | 37.352 | 86.212 | 79.6735 | 80.3128 | 0.009219 | 0.99882 | 0.023727 | 23.7546 | 0.037315 | 0.15761 | 0.006546 | 0.46435 | 0.46435 | 0.013815
2019-03-29 13:57 | 1553882243.5328 | 256 | 0.20818 | 0.00157 | 3.1e-05 | 97.4852 | 96.7067 | 120.1277 | 21.1393 | 0.002656 | 0.002437 | 254.053 | 0.26459 | 8.4014 | 0.002411 | 0.004466 | 0.004355 | 5.3e-05 | 0.001012 | 0.28224 | 1.3888 | 2.5474 | 100.2 | 20.89 | 37.347 | 85.6045 | 79.7107 | 80.3315 | 0.009218 | 0.99877 | 0.023726 | 23.7557 | 0.071721 | 0.1574 | 0.006594 | 0.46775 | 0.46775 | 0.013696
2019-03-29 13:57 | 1553882244.5318 | 257 | 0.20819 | 0.001564 | 3.1e-05 | 97.182 | 96.547 | 120.0403 | 21.1384 | 0.00264 | 0.002432 | 254.0236 | 0.2645 | 8.4032 | 0.002355 | 0.003905 | 0.003844 | 5.4e-05 | 0.001015 | 0.28766 | 1.3642 | 2.5474 | 100.2 | 20.89 | 37.343 | 85.2016 | 78.351 | 80.2128 | 0.009217 | 0.9991 | 0.023722 | 23.7431 | 0.058641 | 0.15735 | 0.006524 | 0.46266 | 0.46266 | 0.013842
2019-03-29 13:57 | 1553882245.5322 | 258 | 0.2082 | 0.001557 | 3.2e-05 | 96.868 | 96.148 | 120.0265 | 21.1384 | 0.002631 | 0.002448 | 253.9805 | 0.26457 | 8.4012 | 0.002411 | 0.003713 | 0.003598 | 5.4e-05 | 0.001008 | 0.28031 | 1.3529 | 2.5473 | 100.2 | 20.89 | 37.335 | 84.9934 | 79.7022 | 80.6003 | 0.009215 | 0.99801 | 0.023707 | 23.7548 | 0.066733 | 0.15739 | 0.006496 | 0.46038 | 0.46038 | 0.013914
2019-03-29 13:57 | 1553882246.5322 | 259 | 0.2082 | 0.001556 | 3.2e-05 | 96.6552 | 95.9053 | 120.13 | 21.1394 | 0.002662 | 0.002464 | 253.9452 | 0.265 | 8.4026 | 0.00228 | 0.003973 | 0.003849 | 5.5e-05 | 0.001004 | 0.2792 | 1.3206 | 2.5474 | 100.2 | 20.9 | 37.33 | 85.7356 | 76.5581 | 80.9826 | 0.009219 | 0.99693 | 0.023705 | 23.7778 | 0.069662 | 0.15765 | 0.006467 | 0.45829 | 0.45829 | 0.014
2019-03-29 13:57 | 1553882247.5309 | 260 | 0.20818 | 0.001553 | 3.2e-05 | 96.3349 | 95.5777 | 120.0556 | 21.1383 | 0.00269 | 0.002468 | 253.9151 | 0.26448 | 8.4023 | 0.002343 | 0.003457 | 0.003407 | 5.3e-05 | 0.001001 | 0.30431 | 1.3316 | 2.5474 | 100.2 | 20.89 | 37.321 | 86.402 | 78.0772 | 81.0705 | 0.009211 | 0.99668 | 0.023695 | 23.7734 | 0.070582 | 0.15734 | 0.006614 | 0.46851 | 0.46851 | 0.013668
2019-03-29 13:57 | 1553882248.5319 | 261 | 0.2082 | 0.001547 | 3.2e-05 | 96.0511 | 95.2603 | 120.1004 | 21.1389 | 0.002619 | 0.002465 | 253.8853 | 0.26538 | 8.402 | 0.002389 | 0.003454 | 0.003438 | 5.6e-05 | 0.001001 | 0.29927 | 1.3392 | 2.5474 | 100.2 | 20.9 | 37.316 | 84.7067 | 79.1764 | 80.9929 | 0.009216 | 0.9969 | 0.023702 | 23.7753 | 0.073949 | 0.15788 | 0.006479 | 0.45904 | 0.45904 | 0.013998
2019-03-29 13:57 | 1553882249.5309 | 262 | 0.20817 | 0.001546 | 3.2e-05 | 95.815 | 94.8306 | 120.1866 | 21.1389 | 0.002644 | 0.002388 | 253.8391 | 0.2643 | 8.4027 | 0.002433 | 0.003555 | 0.00339 | 5.5e-05 | 0.001 | 0.29463 | 1.3462 | 2.5474 | 100.2 | 20.9 | 37.314 | 85.3084 | 80.2296 | 79.1506 | 0.009215 | 1.0021 | 0.023772 | 23.7218 | 0.092365 | 0.15723 | 0.006674 | 0.4743 | 0.4743 | 0.013492
2019-03-29 13:57 | 1553882250.531 | 263 | 0.20817 | 0.001536 | 3.2e-05 | 95.5842 | 94.645 | 120.1922 | 21.1386 | 0.00255 | 0.002464 | 253.8139 | 0.26472 | 8.401 | 0.00238 | 0.003318 | 0.003261 | 6e-05 | 0.000995 | 0.30453 | 1.3063 | 2.5475 | 100.2 | 20.9 | 37.311 | 83.0501 | 78.9546 | 80.9864 | 0.009214 | 0.99692 | 0.023711 | 23.7841 | 0.088324 | 0.15749 | 0.006676 | 0.47321 | 0.47321 | 0.013545
2019-03-29 13:57 | 1553882251.5321 | 264 | 0.20814 | 0.001539 | 3.2e-05 | 95.2872 | 94.4692 | 119.8986 | 21.1389 | 0.00271 | 0.002438 | 253.7968 | 0.26491 | 8.402 | 0.002404 | 0.003711 | 0.003604 | 5.8e-05 | 0.000997 | 0.3183 | 1.3133 | 2.5474 | 100.2 | 20.9 | 37.305 | 86.8954 | 79.5342 | 80.3458 | 0.009213 | 0.99873 | 0.023703 | 23.7336 | 0.077121 | 0.1576 | 0.006848 | 0.48533 | 0.48533 | 0.013216
2019-03-29 13:58 | 1553882252.5311 | 265 | 0.20817 | 0.001539 | 3.2e-05 | 95.0409 | 94.0692 | 119.9676 | 21.1384 | 0.002654 | 0.002472 | 253.7558 | 0.26464 | 8.404 | 0.002429 | 0.004314 | 0.004241 | 5.9e-05 | 0.000994 | 0.30901 | 1.3812 | 2.5476 | 100.2 | 20.9 | 37.302 | 85.5438 | 80.1403 | 81.1603 | 0.009212 | 0.99643 | 0.023683 | 23.7677 | 0.091918 | 0.15743 | 0.006711 | 0.47519 | 0.47519 | 0.013484
2019-03-29 13:58 | 1553882253.5325 | 266 | 0.20814 | 0.001539 | 3.2e-05 | 94.7655 | 93.6999 | 120.214 | 21.1386 | 0.002644 | 0.002433 | 253.7273 | 0.26423 | 8.4035 | 0.002435 | 0.004155 | 0.004055 | 5.7e-05 | 0.000997 | 0.30486 | 1.3955 | 2.5475 | 100.2 | 20.9 | 37.3 | 85.2971 | 80.2863 | 80.2442 | 0.009212 | 0.99901 | 0.023738 | 23.7614 | 0.10114 | 0.15719 | 0.006862 | 0.48699 | 0.48699 | 0.013137
2019-03-29 13:58 | 1553882254.5311 | 267 | 0.20817 | 0.001538 | 3.2e-05 | 94.4725 | 93.5327 | 120.3891 | 21.1383 | 0.002623 | 0.002433 | 253.7067 | 0.26483 | 8.4027 | 0.00238 | 0.004025 | 0.003918 | 5.7e-05 | 0.001002 | 0.29463 | 1.3247 | 2.5475 | 100.2 | 20.9 | 37.297 | 84.8021 | 78.9685 | 80.2228 | 0.009211 | 0.99907 | 0.023756 | 23.7779 | 0.089421 | 0.15755 | 0.006678 | 0.47428 | 0.47428 | 0.01352
2019-03-29 13:58 | 1553882255.5322 | 268 | 0.20815 | 0.001536 | 3.3e-05 | 94.2177 | 93.1764 | 120.0943 | 21.1374 | 0.002633 | 0.00244 | 253.6896 | 0.26445 | 8.4011 | 0.002391 | 0.003736 | 0.003687 | 5.6e-05 | 0.000996 | 0.28385 | 1.3066 | 2.5474 | 100.2 | 20.9 | 37.294 | 85.0324 | 79.2317 | 80.3998 | 0.00921 | 0.99857 | 0.023721 | 23.7547 | 0.099409 | 0.15732 | 0.00682 | 0.48359 | 0.48359 | 0.013241
2019-03-29 13:58 | 1553882256.5322 | 269 | 0.20813 | 0.001537 | 3.3e-05 | 93.9624 | 92.873 | 120.2274 | 21.1385 | 0.002666 | 0.00243 | 253.6674 | 0.26508 | 8.4033 | 0.002348 | 0.003627 | 0.003592 | 6.1e-05 | 0.000993 | 0.2855 | 1.3775 | 2.5474 | 100.2 | 20.9 | 37.292 | 85.8333 | 78.1789 | 80.1591 | 0.00921 | 0.99925 | 0.023742 | 23.7598 | 0.10431 | 0.1577 | 0.006938 | 0.4925 | 0.4925 | 0.013032
2019-03-29 13:58 | 1553882257.5311 | 270 | 0.20816 | 0.001537 | 3.3e-05 | 93.6921 | 92.4798 | 120.4945 | 21.1379 | 0.002642 | 0.002448 | 253.6217 | 0.2651 | 8.4019 | 0.002326 | 0.004196 | 0.004192 | 5.7e-05 | 0.000991 | 0.27992 | 1.3411 | 2.8152 | 100.2 | 20.9 | 37.29 | 85.2676 | 77.6576 | 80.5841 | 0.009209 | 0.99805 | 0.023754 | 23.8005 | 0.11649 | 0.15771 | 0.006739 | 0.47854 | 0.47854 | 0.013413
2019-03-29 13:58 | 1553882258.5311 | 271 | 0.20816 | 0.001535 | 3.2e-05 | 93.4688 | 92.3034 | 120.1539 | 21.1375 | 0.002629 | 0.002444 | 253.5931 | 0.2655 | 8.4011 | 0.002414 | 0.003668 | 0.003605 | 6.1e-05 | 0.000988 | 0.31492 | 1.3266 | 2.5474 | 100.2 | 20.9 | 37.29 | 84.9441 | 79.7822 | 80.4895 | 0.009209 | 0.99832 | 0.023724 | 23.7637 | 0.11223 | 0.15795 | 0.006742 | 0.47815 | 0.47815 | 0.013445
2019-03-29 13:58 | 1553882259.5326 | 272 | 0.20815 | 0.001534 | 3.2e-05 | 93.1895 | 91.914 | 120.3739 | 21.1375 | 0.002652 | 0.002438 | 253.5598 | 0.26426 | 8.403 | 0.002462 | 0.004007 | 0.003976 | 5.9e-05 | 0.000987 | 0.30879 | 1.3376 | 2.5475 | 100.2 | 20.9 | 37.29 | 85.4984 | 80.9183 | 80.36 | 0.009209 | 0.99869 | 0.02375 | 23.781 | 0.12328 | 0.15721 | 0.006813 | 0.48374 | 0.48374 | 0.013227
2019-03-29 13:58 | 1553882260.5323 | 273 | 0.20816 | 0.001544 | 3.1e-05 | 92.904 | 91.5744 | 120.3884 | 21.1383 | 0.00264 | 0.002452 | 253.5305 | 0.26548 | 8.404 | 0.002457 | 0.004404 | 0.004302 | 5.8e-05 | 0.000982 | 0.31957 | 1.3459 | 2.5474 | 100.2 | 20.9 | 37.291 | 85.2151 | 80.8138 | 80.696 | 0.009209 | 0.99774 | 0.02374 | 23.7938 | 0.12893 | 0.15793 | 0.006777 | 0.48106 | 0.48106 | 0.013362
2019-03-29 13:58 | 1553882261.5993 | 274 | 0.20814 | 0.001541 | 3.1e-05 | 92.5358 | 91.3863 | 120.8593 | 21.1377 | 0.002618 | 0.002465 | 253.4792 | 0.2651 | 8.4033 | 0.002437 | 0.00364 | 0.003571 | 6e-05 | 0.000979 | 0.29214 | 1.3693 | 2.5473 | 100.2 | 20.9 | 37.292 | 84.686 | 80.3267 | 81.0112 | 0.00921 | 0.99685 | 0.023776 | 23.8509 | 0.1118 | 0.15771 | 0.006844 | 0.48659 | 0.48659 | 0.013191
2019-03-29 13:58 | 1553882262.5437 | 275 | 0.20814 | 0.001541 | 3.2e-05 | 92.1446 | 90.7672 | 121.1238 | 21.1381 | 0.002574 | 0.002419 | 253.4414 | 0.26557 | 8.4015 | 0.002457 | 0.004152 | 0.003999 | 5.8e-05 | 0.000982 | 0.31653 | 1.3378 | 2.5475 | 100.2 | 20.9 | 37.294 | 83.6324 | 80.8104 | 79.8967 | 0.00921 | 1 | 0.023839 | 23.8394 | 0.13471 | 0.15799 | 0.006839 | 0.48746 | 0.48746 | 0.013191
2019-03-29 13:58 | 1553882263.5451 | 276 | 0.20818 | 0.001547 | 3.2e-05 | 91.7755 | 90.2867 | 121.263 | 21.137 | 0.002615 | 0.002476 | 253.4242 | 0.26446 | 8.4027 | 0.002462 | 0.003783 | 0.003735 | 5.4e-05 | 0.000976 | 0.29574 | 1.3444 | 2.5474 | 100.2 | 20.9 | 37.295 | 84.6012 | 80.9209 | 81.2586 | 0.00921 | 0.99615 | 0.023807 | 23.899 | 0.14629 | 0.15733 | 0.006611 | 0.47051 | 0.47051 | 0.013609
2019-03-29 13:58 | 1553882264.5447 | 277 | 0.20818 | 0.001542 | 3.2e-05 | 91.3717 | 89.9174 | 121.0994 | 21.1369 | 0.002605 | 0.00241 | 253.398 | 0.26513 | 8.3987 | 0.002434 | 0.003911 | 0.003836 | 5.6e-05 | 0.000983 | 0.2855 | 1.349 | 2.5475 | 100.2 | 20.9 | 37.295 | 84.356 | 80.2457 | 79.6771 | 0.00921 | 1.0006 | 0.023844 | 23.8296 | 0.1435 | 0.15773 | 0.006623 | 0.47211 | 0.47211 | 0.013598
2019-03-29 13:58 | 1553882265.5438 | 278 | 0.2082 | 0.001541 | 3.2e-05 | 90.9975 | 89.4526 | 121.4468 | 21.139 | 0.00261 | 0.002439 | 253.3676 | 0.26425 | 8.402 | 0.002472 | 0.003865 | 0.003801 | 5.7e-05 | 0.000979 | 0.28567 | 1.3244 | 2.5475 | 100.2 | 20.9 | 37.295 | 84.4867 | 81.1608 | 80.385 | 0.00921 | 0.99862 | 0.023855 | 23.8876 | 0.15316 | 0.1572 | 0.006517 | 0.46478 | 0.46478 | 0.013766
2019-03-29 13:58 | 1553882266.5438 | 279 | 0.20817 | 0.001539 | 3.2e-05 | 90.562 | 89.1499 | 121.3977 | 21.1365 | 0.002588 | 0.002439 | 253.3241 | 0.2643 | 8.4033 | 0.002406 | 0.004052 | 0.003914 | 5.2e-05 | 0.00098 | 0.30127 | 1.4126 | 2.5474 | 100.2 | 20.9 | 37.295 | 83.9593 | 79.5746 | 80.3707 | 0.00921 | 0.99865 | 0.02385 | 23.8823 | 0.14057 | 0.15723 | 0.00671 | 0.47846 | 0.47846 | 0.013375
2019-03-29 13:58 | 1553882267.5448 | 280 | 0.20819 | 0.001538 | 3.2e-05 | 90.3144 | 88.8103 | 121.3489 | 21.1374 | 0.002599 | 0.002466 | 253.3014 | 0.26483 | 8.4038 | 0.002458 | 0.004152 | 0.004088 | 5.4e-05 | 0.000986 | 0.29734 | 1.3556 | 2.5474 | 100.2 | 20.9 | 37.295 | 84.2161 | 80.8271 | 81.0244 | 0.00921 | 0.99681 | 0.023823 | 23.8996 | 0.15022 | 0.15755 | 0.006584 | 0.46893 | 0.46893 | 0.013674
2019-03-29 13:58 | 1553882268.5448 | 281 | 0.20817 | 0.001538 | 3.2e-05 | 90.2675 | 88.6653 | 121.5879 | 21.1379 | 0.002586 | 0.002437 | 253.2802 | 0.26415 | 8.4029 | 0.002451 | 0.004449 | 0.004351 | 5.3e-05 | 0.000995 | 0.29319 | 1.3622 | 2.5474 | 100.2 | 20.9 | 37.294 | 83.9162 | 80.6553 | 80.3207 | 0.00921 | 0.9988 | 0.023871 | 23.8993 | 0.16019 | 0.15714 | 0.00667 | 0.47598 | 0.47598 | 0.013437
2019-03-29 13:58 | 1553882269.5438 | 282 | 0.20819 | 0.001528 | 3.2e-05 | 90.3103 | 88.5328 | 121.401 | 21.1378 | 0.002575 | 0.002409 | 253.2629 | 0.26477 | 8.4022 | 0.00244 | 0.004883 | 0.004808 | 5.4e-05 | 0.001004 | 0.2959 | 1.3813 | 2.5474 | 100.2 | 20.9 | 37.289 | 83.646 | 80.3932 | 79.6592 | 0.009209 | 1.0007 | 0.023875 | 23.8586 | 0.1778 | 0.15751 | 0.006569 | 0.46887 | 0.46887 | 0.013673
2019-03-29 13:58 | 1553882270.5453 | 283 | 0.20821 | 0.001527 | 3.2e-05 | 90.4985 | 89.0574 | 121.5409 | 21.1375 | 0.0026 | 0.002395 | 253.2295 | 0.26439 | 8.4014 | 0.002427 | 0.004142 | 0.004099 | 5.7e-05 | 0.001004 | 0.28296 | 1.3665 | 2.5475 | 100.2 | 20.9 | 37.289 | 84.2496 | 80.0855 | 79.3293 | 0.009209 | 1.0016 | 0.0239 | 23.8612 | 0.14358 | 0.15728 | 0.006445 | 0.46043 | 0.46043 | 0.013903
2019-03-29 13:58 | 1553882271.545 | 284 | 0.20821 | 0.001528 | 3.2e-05 | 90.797 | 89.2758 | 121.3862 | 21.1375 | 0.002621 | 0.002407 | 253.2081 | 0.26487 | 8.4023 | 0.002429 | 0.0042 | 0.00412 | 5.8e-05 | 0.001007 | 0.29607 | 1.3556 | 2.5474 | 100.2 | 20.9 | 37.285 | 84.7623 | 80.1432 | 79.6023 | 0.009208 | 1.0008 | 0.023875 | 23.8552 | 0.15112 | 0.15757 | 0.006464 | 0.46134 | 0.46134 | 0.013901
2019-03-29 13:58 | 1553882272.545 | 285 | 0.20817 | 0.001524 | 3.3e-05 | 91.1349 | 89.6229 | 121.4095 | 21.1371 | 0.002625 | 0.002449 | 253.174 | 0.26506 | 8.4007 | 0.002438 | 0.004306 | 0.004191 | 5.5e-05 | 0.00101 | 0.29175 | 1.3343 | 2.5473 | 100.2 | 20.9 | 37.285 | 84.8478 | 80.3511 | 80.6177 | 0.009208 | 0.99796 | 0.023843 | 23.8918 | 0.14964 | 0.15769 | 0.006736 | 0.48013 | 0.48013 | 0.013367
2019-03-29 13:58 | 1553882273.544 | 286 | 0.20824 | 0.001518 | 3.2e-05 | 91.4495 | 90.0499 | 121.3156 | 21.1375 | 0.002594 | 0.002402 | 253.141 | 0.26551 | 8.4027 | 0.002522 | 0.004218 | 0.004227 | 5.6e-05 | 0.001009 | 0.27903 | 1.3258 | 2.5473 | 100.2 | 20.9 | 37.283 | 84.0969 | 82.3814 | 79.4915 | 0.009207 | 1.0011 | 0.023872 | 23.8445 | 0.13796 | 0.15795 | 0.006289 | 0.44873 | 0.44873 | 0.014326
2019-03-29 13:58 | 1553882274.5451 | 287 | 0.20823 | 0.001518 | 3.2e-05 | 91.8371 | 90.486 | 121.3827 | 21.1374 | 0.00262 | 0.002401 | 253.1017 | 0.26454 | 8.4027 | 0.002418 | 0.004126 | 0.004007 | 5.7e-05 | 0.001013 | 0.29319 | 1.3047 | 2.5473 | 100.2 | 20.9 | 37.281 | 84.7247 | 79.8742 | 79.4525 | 0.009207 | 1.0013 | 0.02388 | 23.8498 | 0.13257 | 0.15737 | 0.006333 | 0.45201 | 0.45201 | 0.01417
2019-03-29 13:58 | 1553882275.5451 | 288 | 0.20819 | 0.001517 | 3.2e-05 | 92.0862 | 90.8178 | 121.5515 | 21.1371 | 0.002673 | 0.002393 | 253.0624 | 0.26435 | 8.403 | 0.002394 | 0.003931 | 0.003777 | 5.4e-05 | 0.001019 | 0.30182 | 1.3001 | 2.5474 | 100.2 | 20.9 | 37.277 | 86.0074 | 79.2861 | 79.2688 | 0.009206 | 1.0018 | 0.023903 | 23.8602 | 0.12406 | 0.15726 | 0.006579 | 0.47016 | 0.47016 | 0.013614
2019-03-29 13:58 | 1553882276.5441 | 289 | 0.20821 | 0.001513 | 3.1e-05 | 92.5268 | 91.0829 | 121.571 | 21.1374 | 0.002581 | 0.002368 | 253.0327 | 0.26465 | 8.4032 | 0.002338 | 0.004002 | 0.003968 | 5.4e-05 | 0.001017 | 0.31542 | 1.313 | 2.5474 | 100.2 | 20.9 | 37.276 | 83.7993 | 77.9603 | 78.6612 | 0.009206 | 1.0035 | 0.023925 | 23.8415 | 0.14068 | 0.15744 | 0.006488 | 0.46407 | 0.46407 | 0.013808
2019-03-29 13:58 | 1553882277.5441 | 290 | 0.20825 | 0.00151 | 3.2e-05 | 92.7423 | 91.4662 | 121.5919 | 21.1382 | 0.002634 | 0.002419 | 252.9984 | 0.26481 | 8.4022 | 0.002443 | 0.003287 | 0.003226 | 5.7e-05 | 0.001011 | 0.29806 | 1.3374 | 2.5474 | 100.2 | 20.9 | 37.274 | 85.0641 | 80.4784 | 79.9007 | 0.009205 | 0.99999 | 0.023885 | 23.8855 | 0.12393 | 0.15754 | 0.006272 | 0.44777 | 0.44777 | 0.014319
2019-03-29 13:58 | 1553882278.5442 | 291 | 0.20821 | 0.001509 | 3.2e-05 | 92.9641 | 91.7027 | 121.5074 | 21.1375 | 0.002608 | 0.002402 | 252.9598 | 0.26404 | 8.4011 | 0.00239 | 0.004014 | 0.003869 | 5.5e-05 | 0.001015 | 0.26178 | 1.3452 | 2.5474 | 100.2 | 20.9 | 37.275 | 84.4377 | 79.2015 | 79.4908 | 0.009206 | 1.0011 | 0.023891 | 23.8634 | 0.1222 | 0.15708 | 0.006476 | 0.46248 | 0.46248 | 0.013824
2019-03-29 13:58 | 1553882279.5432 | 292 | 0.20823 | 0.001503 | 3.2e-05 | 93.177 | 91.9041 | 121.26 | 21.137 | 0.002623 | 0.002364 | 252.9313 | 0.26462 | 8.4014 | 0.002402 | 0.00417 | 0.004159 | 5.5e-05 | 0.001019 | 0.29911 | 1.3092 | 2.5474 | 100.2 | 20.9 | 37.276 | 84.804 | 79.4973 | 78.5726 | 0.009206 | 1.0038 | 0.023898 | 23.808 | 0.12303 | 0.15742 | 0.006355 | 0.45393 | 0.45393 | 0.014115
2019-03-29 13:58 | 1553882280.5442 | 293 | 0.20821 | 0.001499 | 3.3e-05 | 93.4194 | 92.1682 | 120.9248 | 21.1361 | 0.00261 | 0.002341 | 252.8975 | 0.26498 | 8.4024 | 0.00244 | 0.003495 | 0.003462 | 5.7e-05 | 0.001021 | 0.29175 | 1.3099 | 2.5473 | 100.2 | 20.9 | 37.282 | 84.4817 | 80.3867 | 78.0198 | 0.009207 | 1.0053 | 0.023883 | 23.7564 | 0.12061 | 0.15764 | 0.006528 | 0.46602 | 0.46602 | 0.013767
2019-03-29 13:58 | 1553882281.5437 | 294 | 0.2082 | 0.001501 | 3.3e-05 | 93.6658 | 92.4713 | 120.7483 | 21.1359 | 0.002589 | 0.00244 | 252.8568 | 0.2647 | 8.4008 | 0.002439 | 0.00434 | 0.004232 | 5.8e-05 | 0.001022 | 0.31166 | 1.32 | 2.5473 | 100.2 | 20.9 | 37.284 | 83.9907 | 80.3765 | 80.3935 | 0.009208 | 0.99859 | 0.023786 | 23.8191 | 0.1148 | 0.15747 | 0.006538 | 0.46487 | 0.46487 | 0.013787
2019-03-29 13:58 | 1553882282.5444 | 295 | 0.20822 | 0.001492 | 3.3e-05 | 93.8136 | 92.7433 | 120.2678 | 21.1363 | 0.002623 | 0.002432 | 252.8184 | 0.26425 | 8.4042 | 0.002431 | 0.003797 | 0.003773 | 5.6e-05 | 0.001023 | 0.2943 | 1.3101 | 2.5473 | 100.2 | 20.9 | 37.289 | 84.7956 | 80.1781 | 80.2163 | 0.009209 | 0.99909 | 0.023744 | 23.7657 | 0.10264 | 0.1572 | 0.006436 | 0.45674 | 0.45674 | 0.014008
2019-03-29 13:58 | 1553882283.5443 | 296 | 0.20822 | 0.001497 | 3.2e-05 | 94.0717 | 92.8658 | 119.9321 | 21.1369 | 0.002561 | 0.002456 | 252.8012 | 0.26522 | 8.4017 | 0.002478 | 0.00414 | 0.004034 | 5.2e-05 | 0.001027 | 0.29806 | 1.3051 | 2.5474 | 100.2 | 20.9 | 37.293 | 83.3145 | 81.3137 | 80.7852 | 0.00921 | 0.99749 | 0.023692 | 23.7516 | 0.11539 | 0.15778 | 0.006427 | 0.45514 | 0.45514 | 0.014109
2019-03-29 13:58 | 1553882284.5433 | 297 | 0.20821 | 0.001496 | 3.2e-05 | 94.3895 | 93.433 | 119.941 | 21.1364 | 0.002587 | 0.002411 | 252.7871 | 0.26553 | 8.4014 | 0.002418 | 0.004093 | 0.004121 | 5.5e-05 | 0.001028 | 0.28871 | 1.3337 | 2.5473 | 100.2 | 20.9 | 37.296 | 83.9232 | 79.8614 | 79.704 | 0.009211 | 1.0005 | 0.023729 | 23.7162 | 0.091102 | 0.15796 | 0.006516 | 0.46219 | 0.46219 | 0.01391
2019-03-29 13:58 | 1553882285.5448 | 298 | 0.20822 | 0.001493 | 3.2e-05 | 94.65 | 93.7115 | 119.7117 | 21.1362 | 0.002576 | 0.002382 | 252.7467 | 0.26473 | 8.4049 | 0.002461 | 0.004147 | 0.004122 | 5.8e-05 | 0.001031 | 0.29911 | 1.3705 | 2.5475 | 100.2 | 20.91 | 37.303 | 83.6619 | 80.9132 | 79.0006 | 0.009218 | 1.0025 | 0.02373 | 23.6699 | 0.089125 | 0.15749 | 0.006461 | 0.45832 | 0.45832 | 0.013985
2019-03-29 13:58 | 1553882286.5434 | 299 | 0.20821 | 0.001492 | 3.1e-05 | 94.9008 | 94.0136 | 119.8533 | 21.1362 | 0.002566 | 0.002355 | 252.6984 | 0.26441 | 8.4032 | 0.002408 | 0.003685 | 0.003539 | 5.9e-05 | 0.001028 | 0.27417 | 1.3646 | 2.5474 | 100.2 | 20.9 | 37.306 | 83.4192 | 79.6188 | 78.3458 | 0.009213 | 1.0044 | 0.023766 | 23.6619 | 0.084016 | 0.1573 | 0.006539 | 0.46462 | 0.46462 | 0.013779
2019-03-29 13:58 | 1553882287.5445 | 300 | 0.20823 | 0.0015 | 3.1e-05 | 95.1011 | 94.117 | 119.8617 | 21.1358 | 0.002588 | 0.002346 | 252.6665 | 0.26415 | 8.402 | 0.002397 | 0.003258 | 0.003249 | 6e-05 | 0.001023 | 0.28241 | 1.3654 | 2.5473 | 100.2 | 20.9 | 37.309 | 83.95 | 79.3675 | 78.1481 | 0.009214 | 1.005 | 0.023774 | 23.656 | 0.093037 | 0.15714 | 0.00641 | 0.4556 | 0.4556 | 0.014038
2019-03-29 13:58 | 1553882288.5435 | 301 | 0.20824 | 0.00149 | 3.1e-05 | 95.3767 | 94.5617 | 119.9176 | 21.1357 | 0.002627 | 0.00239 | 252.634 | 0.26543 | 8.4001 | 0.002433 | 0.00384 | 0.003798 | 5.8e-05 | 0.00102 | 0.30359 | 1.3245 | 2.5474 | 100.2 | 20.91 | 37.309 | 84.8996 | 80.239 | 79.1966 | 0.00922 | 1.002 | 0.023744 | 23.6968 | 0.076765 | 0.15791 | 0.006365 | 0.4518 | 0.4518 | 0.014225
2019-03-29 13:58 | 1553882289.5439 | 302 | 0.20825 | 0.001488 | 3.1e-05 | 95.6955 | 94.732 | 120.092 | 21.1364 | 0.002602 | 0.002346 | 252.5985 | 0.2643 | 8.4045 | 0.002435 | 0.004332 | 0.004153 | 5.6e-05 | 0.001016 | 0.31382 | 1.3409 | 2.5473 | 100.2 | 20.91 | 37.308 | 84.294 | 80.2772 | 78.1463 | 0.009219 | 1.005 | 0.023797 | 23.6787 | 0.090517 | 0.15723 | 0.006313 | 0.44907 | 0.44907 | 0.01425
2019-03-29 13:58 | 1553882290.5446 | 303 | 0.20825 | 0.001496 | 3.1e-05 | 95.8069 | 94.8334 | 120.4633 | 21.1354 | 0.002588 | 0.002414 | 252.5835 | 0.26415 | 8.4013 | 0.002511 | 0.003791 | 0.003801 | 5.7e-05 | 0.001023 | 0.3131 | 1.2967 | 2.5473 | 100.2 | 20.91 | 37.304 | 83.9653 | 82.108 | 79.7654 | 0.009218 | 1.0004 | 0.023779 | 23.7699 | 0.091354 | 0.15714 | 0.006296 | 0.44755 | 0.44755 | 0.014291
2019-03-29 13:58 | 1553882291.5436 | 304 | 0.20826 | 0.001483 | 3.2e-05 | 96.0634 | 95.2651 | 120.6483 | 21.1361 | 0.00261 | 0.00242 | 252.5558 | 0.26471 | 8.4039 | 0.002453 | 0.003723 | 0.0036 | 5.8e-05 | 0.001025 | 0.2959 | 1.3121 | 2.5474 | 100.2 | 20.91 | 37.301 | 84.4919 | 80.7018 | 79.9064 | 0.009218 | 0.99997 | 0.023792 | 23.7928 | 0.074643 | 0.15748 | 0.006229 | 0.44298 | 0.44298 | 0.014469
2019-03-29 13:58 | 1553882292.544 | 305 | 0.20824 | 0.001483 | 3.2e-05 | 96.2484 | 95.5484 | 121.1515 | 21.1349 | 0.002541 | 0.002404 | 252.5408 | 0.26444 | 8.4037 | 0.002465 | 0.004828 | 0.00479 | 5.2e-05 | 0.001022 | 0.29358 | 1.3511 | 2.5475 | 100.2 | 20.91 | 37.296 | 82.8282 | 80.9945 | 79.5255 | 0.009216 | 1.001 | 0.023855 | 23.8296 | 0.065294 | 0.15731 | 0.006328 | 0.45118 | 0.45118 | 0.014191
2019-03-29 13:58 | 1553882293.5437 | 306 | 0.20826 | 0.001485 | 3.2e-05 | 96.466 | 95.7999 | 121.4275 | 21.1353 | 0.002543 | 0.002399 | 252.5058 | 0.26502 | 8.4022 | 0.002453 | 0.004112 | 0.004029 | 5.6e-05 | 0.00102 | 0.28351 | 1.366 | 2.5475 | 100.2 | 20.91 | 37.293 | 82.8707 | 80.7003 | 79.4144 | 0.009216 | 1.0014 | 0.023885 | 23.8529 | 0.061979 | 0.15766 | 0.006222 | 0.44418 | 0.44418 | 0.014447
2019-03-29 13:58 | 1553882294.5427 | 307 | 0.20825 | 0.001483 | 3.3e-05 | 96.6265 | 95.9234 | 121.5022 | 21.135 | 0.00251 | 0.002403 | 252.4978 | 0.26453 | 8.4047 | 0.002518 | 0.004183 | 0.004169 | 5.8e-05 | 0.001021 | 0.30685 | 1.3525 | 2.5474 | 100.2 | 20.91 | 37.291 | 82.0724 | 82.2837 | 79.5171 | 0.009215 | 1.0011 | 0.023889 | 23.8637 | 0.065315 | 0.15737 | 0.0063 | 0.44982 | 0.44982 | 0.014239
2019-03-29 13:58 | 1553882295.5437 | 308 | 0.20825 | 0.001486 | 3.3e-05 | 96.8105 | 95.9166 | 121.421 | 21.1353 | 0.002663 | 0.00241 | 252.4839 | 0.26493 | 8.4017 | 0.002504 | 0.004034 | 0.00394 | 5.6e-05 | 0.00102 | 0.30254 | 1.3466 | 2.5474 | 100.2 | 20.9 | 37.286 | 85.7505 | 81.9394 | 79.6711 | 0.009208 | 1.0006 | 0.023876 | 23.861 | 0.082977 | 0.1576 | 0.006266 | 0.44711 | 0.44711 | 0.014347
2019-03-29 13:58 | 1553882296.5427 | 309 | 0.20824 | 0.001483 | 3.2e-05 | 96.9848 | 96.3459 | 121.3222 | 21.1349 | 0.002589 | 0.002401 | 252.4649 | 0.26446 | 8.4033 | 0.002473 | 0.003827 | 0.003752 | 5.8e-05 | 0.001022 | 0.31238 | 1.3748 | 2.5474 | 100.2 | 20.91 | 37.284 | 83.9797 | 81.1919 | 79.4688 | 0.009213 | 1.0012 | 0.023873 | 23.8444 | 0.05911 | 0.15733 | 0.006338 | 0.45224 | 0.45224 | 0.014159
2019-03-29 13:58 | 1553882297.5431 | 310 | 0.20827 | 0.001483 | 3.3e-05 | 97.1186 | 96.3907 | 121.2355 | 21.1353 | 0.002481 | 0.002406 | 252.4509 | 0.26451 | 8.4035 | 0.00247 | 0.004253 | 0.00408 | 5.6e-05 | 0.001024 | 0.3063 | 1.3451 | 2.5475 | 100.2 | 20.9 | 37.282 | 81.3771 | 81.1182 | 79.5727 | 0.009207 | 1.0009 | 0.023861 | 23.8394 | 0.067292 | 0.15736 | 0.006165 | 0.43963 | 0.43963 | 0.014568
2019-03-29 13:58 | 1553882298.5438 | 311 | 0.2083 | 0.001484 | 3.2e-05 | 97.3023 | 96.6497 | 121.1381 | 21.1359 | 0.002551 | 0.002414 | 252.4471 | 0.26447 | 8.402 | 0.00245 | 0.003264 | 0.003301 | 5.7e-05 | 0.001027 | 0.31868 | 1.3763 | 2.5475 | 100.2 | 20.9 | 37.28 | 83.0679 | 80.6339 | 79.7769 | 0.009207 | 1.0003 | 0.023845 | 23.8367 | 0.060197 | 0.15733 | 0.005995 | 0.42721 | 0.42721 | 0.014989
2019-03-29 13:58 | 1553882299.5428 | 312 | 0.20826 | 0.001488 | 3.2e-05 | 97.4376 | 96.9627 | 121.1884 | 21.1348 | 0.002601 | 0.002416 | 252.4135 | 0.26481 | 8.4023 | 0.002354 | 0.004275 | 0.00414 | 6e-05 | 0.001027 | 0.27848 | 1.3785 | 2.5474 | 100.2 | 20.91 | 37.281 | 84.265 | 78.3187 | 79.8318 | 0.009213 | 1.0002 | 0.023848 | 23.8435 | 0.043709 | 0.15754 | 0.00621 | 0.44262 | 0.44262 | 0.014486
2019-03-29 13:58 | 1553882300.5439 | 313 | 0.2083 | 0.001482 | 3.2e-05 | 97.6142 | 97.0135 | 121.3344 | 21.1338 | 0.002565 | 0.002399 | 252.3657 | 0.26455 | 8.4033 | 0.002411 | 0.003728 | 0.003686 | 5.7e-05 | 0.001034 | 0.28567 | 1.3507 | 2.5475 | 100.2 | 20.91 | 37.285 | 83.4035 | 79.6883 | 79.4209 | 0.009214 | 1.0013 | 0.023876 | 23.844 | 0.055217 | 0.15738 | 0.006 | 0.42813 | 0.42813 | 0.014962
2019-03-29 13:58 | 1553882301.5443 | 314 | 0.20825 | 0.001484 | 3.2e-05 | 97.7891 | 97.1589 | 121.2959 | 21.1344 | 0.002599 | 0.002387 | 252.3386 | 0.26459 | 8.4023 | 0.002394 | 0.004042 | 0.003958 | 5.8e-05 | 0.001029 | 0.28871 | 1.3281 | 2.5474 | 100.2 | 20.91 | 37.287 | 84.2108 | 79.2943 | 79.1144 | 0.009214 | 1.0022 | 0.023883 | 23.8298 | 0.057826 | 0.1574 | 0.006285 | 0.44864 | 0.44864 | 0.01428
2019-03-29 13:58 | 1553882302.5429 | 315 | 0.20821 | 0.001482 | 3.2e-05 | 97.9022 | 97.3155 | 121.0802 | 21.1342 | 0.002608 | 0.002408 | 252.3159 | 0.26469 | 8.4022 | 0.002423 | 0.003521 | 0.003511 | 6e-05 | 0.001019 | 0.28744 | 1.3114 | 2.5473 | 100.2 | 20.91 | 37.289 | 84.4463 | 79.9841 | 79.6255 | 0.009215 | 1.0008 | 0.023844 | 23.8259 | 0.053759 | 0.15746 | 0.00655 | 0.4669 | 0.4669 | 0.013726
2019-03-29 13:58 | 1553882303.5429 | 316 | 0.20823 | 0.001477 | 3.2e-05 | 97.9959 | 97.4285 | 120.952 | 21.1351 | 0.00253 | 0.002427 | 252.2724 | 0.26506 | 8.4028 | 0.002409 | 0.003808 | 0.003605 | 5.5e-05 | 0.00102 | 0.28008 | 1.3105 | 2.5473 | 100.2 | 20.91 | 37.291 | 82.5716 | 79.6528 | 80.0731 | 0.009215 | 0.9995 | 0.023816 | 23.8284 | 0.051938 | 0.15769 | 0.006412 | 0.45649 | 0.45649 | 0.014059
2019-03-29 13:58 | 1553882304.5444 | 317 | 0.20826 | 0.00148 | 3.1e-05 | 98.1827 | 97.5947 | 120.6349 | 21.1341 | 0.00252 | 0.002377 | 252.2122 | 0.26483 | 8.4024 | 0.002449 | 0.003855 | 0.003863 | 5.3e-05 | 0.00102 | 0.27201 | 1.3675 | 2.5473 | 100.2 | 20.91 | 37.291 | 82.3291 | 80.6016 | 78.888 | 0.009215 | 1.0029 | 0.023825 | 23.7572 | 0.053725 | 0.15755 | 0.006233 | 0.44392 | 0.44392 | 0.014444
2019-03-29 13:58 | 1553882305.5442 | 318 | 0.20829 | 0.001471 | 3.2e-05 | 98.3332 | 97.734 | 120.5624 | 21.1331 | 0.002574 | 0.00238 | 252.1681 | 0.26436 | 8.4031 | 0.002465 | 0.004081 | 0.003991 | 5.2e-05 | 0.001023 | 0.29407 | 1.3441 | 2.5473 | 100.2 | 20.91 | 37.287 | 83.6134 | 81.0005 | 78.9624 | 0.009214 | 1.0027 | 0.023815 | 23.7526 | 0.054678 | 0.15727 | 0.006073 | 0.4323 | 0.4323 | 0.014806
2019-03-29 13:58 | 1553882306.5441 | 319 | 0.20824 | 0.001471 | 3.1e-05 | 98.4769 | 98.0974 | 120.4804 | 21.1349 | 0.002549 | 0.002364 | 252.144 | 0.26346 | 8.4036 | 0.002479 | 0.003953 | 0.003843 | 5.2e-05 | 0.001026 | 0.28263 | 1.3743 | 2.5474 | 100.2 | 20.91 | 37.287 | 83.0201 | 81.3448 | 78.5761 | 0.009214 | 1.0038 | 0.02382 | 23.7315 | 0.034535 | 0.15673 | 0.006345 | 0.45181 | 0.45181 | 0.014119
2019-03-29 13:58 | 1553882307.544 | 320 | 0.20826 | 0.001473 | 3.1e-05 | 98.687 | 98.2544 | 120.5377 | 21.1346 | 0.002551 | 0.002377 | 252.1091 | 0.26428 | 8.4023 | 0.002457 | 0.004126 | 0.004104 | 5.1e-05 | 0.00103 | 0.30525 | 1.3536 | 2.5475 | 100.2 | 20.91 | 37.281 | 83.0594 | 80.8018 | 78.8735 | 0.009213 | 1.0029 | 0.023816 | 23.7471 | 0.039297 | 0.15722 | 0.00625 | 0.44498 | 0.44498 | 0.01438
2019-03-29 13:58 | 1553882308.5422 | 321 | 0.20829 | 0.001469 | 3e-05 | 98.8613 | 98.3671 | 120.5541 | 21.1335 | 0.002538 | 0.002411 | 252.0814 | 0.26396 | 8.4019 | 0.002399 | 0.004342 | 0.004189 | 5.2e-05 | 0.001024 | 0.31166 | 1.3854 | 2.5476 | 100.2 | 20.91 | 37.281 | 82.7619 | 79.3987 | 79.6981 | 0.009213 | 1.0006 | 0.02379 | 23.7766 | 0.044826 | 0.15703 | 0.006093 | 0.43329 | 0.43329 | 0.01475
2019-03-29 13:58 | 1553882309.5421 | 322 | 0.20827 | 0.00146 | 3.1e-05 | 99.0088 | 98.6234 | 120.5678 | 21.1333 | 0.002607 | 0.002359 | 252.0725 | 0.26531 | 8.3991 | 0.002418 | 0.003568 | 0.003575 | 5.5e-05 | 0.001027 | 0.29645 | 1.409 | 2.8281 | 100.2 | 20.91 | 37.279 | 84.4179 | 79.8615 | 78.4505 | 0.009212 | 1.0041 | 0.023833 | 23.7358 | 0.034892 | 0.15783 | 0.006221 | 0.44322 | 0.44322 | 0.014494
2019-03-29 13:58 | 1553882310.5436 | 323 | 0.20828 | 0.001464 | 3.1e-05 | 99.1289 | 98.7841 | 120.3396 | 21.1338 | 0.002577 | 0.002417 | 252.038 | 0.26508 | 8.4013 | 0.002446 | 0.003561 | 0.003531 | 5.1e-05 | 0.00103 | 0.29231 | 1.3573 | 2.5475 | 100.2 | 20.91 | 37.276 | 83.6953 | 80.5304 | 79.8527 | 0.009211 | 1.0001 | 0.023763 | 23.7606 | 0.031162 | 0.1577 | 0.006146 | 0.43661 | 0.43661 | 0.0147
2019-03-29 13:58 | 1553882311.5423 | 324 | 0.20827 | 0.001459 | 3.1e-05 | 99.2645 | 98.878 | 120.5154 | 21.1336 | 0.002601 | 0.002357 | 252.0051 | 0.26452 | 8.4018 | 0.002535 | 0.003816 | 0.003625 | 5.2e-05 | 0.001031 | 0.30917 | 1.3317 | 2.5475 | 100.2 | 20.91 | 37.272 | 84.2583 | 82.681 | 78.3954 | 0.00921 | 1.0043 | 0.02383 | 23.7288 | 0.034895 | 0.15736 | 0.006235 | 0.44415 | 0.44415 | 0.01442
2019-03-29 13:59 | 1553882312.5429 | 325 | 0.20827 | 0.001457 | 3.2e-05 | 99.4485 | 99.1291 | 120.3565 | 21.1329 | 0.002603 | 0.002405 | 251.9798 | 0.265 | 8.4022 | 0.002496 | 0.004311 | 0.00436 | 4.9e-05 | 0.001034 | 0.28655 | 1.3198 | 2.5474 | 100.2 | 20.91 | 37.269 | 84.3178 | 81.7445 | 79.5628 | 0.00921 | 1.0009 | 0.023775 | 23.7525 | 0.028773 | 0.15765 | 0.006236 | 0.44316 | 0.44316 | 0.014479
2019-03-29 13:59 | 1553882313.5422 | 326 | 0.20829 | 0.001455 | 3.1e-05 | 99.6677 | 99.0605 | 120.3215 | 21.1344 | 0.002551 | 0.002405 | 251.9645 | 0.26527 | 8.4005 | 0.002419 | 0.003342 | 0.003283 | 5.1e-05 | 0.001029 | 0.27632 | 1.3639 | 2.5475 | 100.2 | 20.91 | 37.267 | 83.0561 | 79.8932 | 79.5632 | 0.009209 | 1.0009 | 0.023771 | 23.7491 | 0.054651 | 0.15781 | 0.006078 | 0.43187 | 0.43187 | 0.014872
2019-03-29 13:59 | 1553882314.5423 | 327 | 0.20828 | 0.001452 | 3.1e-05 | 99.694 | 99.3783 | 120.5642 | 21.1343 | 0.002556 | 0.002409 | 251.9352 | 0.26518 | 8.402 | 0.002444 | 0.00358 | 0.003496 | 5.2e-05 | 0.001027 | 0.28639 | 1.3659 | 2.5474 | 100.2 | 20.91 | 37.267 | 83.1731 | 80.4888 | 79.66 | 0.009209 | 1.0007 | 0.023792 | 23.7763 | 0.028371 | 0.15776 | 0.006142 | 0.43682 | 0.43682 | 0.014699
2019-03-29 13:59 | 1553882315.5423 | 328 | 0.20826 | 0.00145 | 3.1e-05 | 99.8646 | 99.4141 | 120.5525 | 21.1341 | 0.002523 | 0.00246 | 251.9113 | 0.2645 | 8.4032 | 0.002438 | 0.003588 | 0.003599 | 5.2e-05 | 0.001024 | 0.28998 | 1.3734 | 2.5474 | 100.2 | 20.91 | 37.268 | 82.3843 | 80.3513 | 80.8736 | 0.009209 | 0.99724 | 0.02375 | 23.816 | 0.04044 | 0.15735 | 0.006282 | 0.44602 | 0.44602 | 0.014359
2019-03-29 13:59 | 1553882316.5423 | 329 | 0.2083 | 0.001444 | 3.1e-05 | 100.0349 | 99.7175 | 120.3833 | 21.1336 | 0.002579 | 0.002399 | 251.8686 | 0.26464 | 8.4014 | 0.00235 | 0.003638 | 0.003522 | 5.4e-05 | 0.001028 | 0.3027 | 1.346 | 2.5474 | 100.2 | 20.91 | 37.268 | 83.7334 | 78.231 | 79.3988 | 0.009209 | 1.0014 | 0.023783 | 23.7496 | 0.028426 | 0.15743 | 0.006028 | 0.42856 | 0.42856 | 0.014952
2019-03-29 13:59 | 1553882317.5424 | 330 | 0.20829 | 0.001443 | 3.1e-05 | 100.0325 | 99.6124 | 120.2025 | 21.1336 | 0.002555 | 0.002386 | 251.8509 | 0.26381 | 8.4051 | 0.002346 | 0.003952 | 0.003915 | 5.4e-05 | 0.001029 | 0.31974 | 1.3548 | 2.5474 | 100.2 | 20.91 | 37.27 | 83.1693 | 78.132 | 79.0893 | 0.00921 | 1.0023 | 0.023776 | 23.7214 | 0.037646 | 0.15694 | 0.006089 | 0.43278 | 0.43278 | 0.014759
2019-03-29 13:59 | 1553882318.5424 | 331 | 0.20827 | 0.001443 | 3e-05 | 100.1547 | 99.8411 | 120.3842 | 21.1334 | 0.002543 | 0.002402 | 251.809 | 0.26439 | 8.403 | 0.002425 | 0.004607 | 0.004468 | 5.6e-05 | 0.001034 | 0.28512 | 1.3518 | 2.5475 | 100.2 | 20.91 | 37.274 | 82.8827 | 80.0426 | 79.4945 | 0.009211 | 1.0011 | 0.02378 | 23.7529 | 0.02805 | 0.15729 | 0.006207 | 0.44131 | 0.44131 | 0.014506
2019-03-29 13:59 | 1553882319.5438 | 332 | 0.2083 | 0.001442 | 3e-05 | 100.2109 | 99.9185 | 120.3844 | 21.1335 | 0.002604 | 0.00239 | 251.7788 | 0.26482 | 8.4023 | 0.002446 | 0.003598 | 0.003595 | 5.3e-05 | 0.001037 | 0.31454 | 1.3303 | 2.5473 | 100.2 | 20.91 | 37.277 | 84.3325 | 80.5416 | 79.1882 | 0.009212 | 1.002 | 0.02379 | 23.7426 | 0.026137 | 0.15754 | 0.00607 | 0.4317 | 0.4317 | 0.014853
2019-03-29 13:59 | 1553882320.5439 | 333 | 0.20829 | 0.001442 | 3e-05 | 100.3196 | 99.8107 | 120.4941 | 21.1335 | 0.002575 | 0.002397 | 251.7639 | 0.26432 | 8.4022 | 0.00248 | 0.00347 | 0.00341 | 5.8e-05 | 0.001037 | 0.30541 | 1.3223 | 2.5474 | 100.2 | 20.91 | 37.279 | 83.6325 | 81.3487 | 79.3636 | 0.009212 | 1.0015 | 0.023795 | 23.7594 | 0.045489 | 0.15724 | 0.00609 | 0.43324 | 0.43324 | 0.014772
2019-03-29 13:59 | 1553882321.5429 | 334 | 0.20829 | 0.001439 | 2.9e-05 | 100.3654 | 100.0216 | 120.2819 | 21.1335 | 0.002599 | 0.002367 | 251.7464 | 0.26469 | 8.4046 | 0.002383 | 0.003365 | 0.003303 | 5.7e-05 | 0.001034 | 0.32101 | 1.3156 | 2.5474 | 100.2 | 20.91 | 37.285 | 84.2199 | 79.0245 | 78.6418 | 0.009214 | 1.0036 | 0.023799 | 23.7141 | 0.03069 | 0.15746 | 0.006117 | 0.43521 | 0.43521 | 0.014726
2019-03-29 13:59 | 1553882322.5425 | 335 | 0.20827 | 0.001436 | 3e-05 | 100.3878 | 100.0612 | 120.2983 | 21.1334 | 0.002536 | 0.002355 | 251.7355 | 0.26476 | 8.4026 | 0.002519 | 0.003697 | 0.003689 | 5.5e-05 | 0.001032 | 0.29358 | 1.3282 | 2.5474 | 100.2 | 20.91 | 37.287 | 82.7128 | 82.2854 | 78.3611 | 0.009214 | 1.0044 | 0.02381 | 23.7063 | 0.029151 | 0.15751 | 0.006267 | 0.44614 | 0.44614 | 0.014369
2019-03-29 13:59 | 1553882323.5426 | 336 | 0.20831 | 0.001434 | 3e-05 | 100.4729 | 100.2148 | 120.4009 | 21.1341 | 0.002595 | 0.002403 | 251.7191 | 0.26433 | 8.4021 | 0.002504 | 0.004316 | 0.004201 | 5.5e-05 | 0.001027 | 0.28854 | 1.3514 | 2.5475 | 100.2 | 20.91 | 37.293 | 84.1232 | 81.9406 | 79.4969 | 0.009216 | 1.0011 | 0.023781 | 23.7547 | 0.023004 | 0.15725 | 0.005983 | 0.42533 | 0.42533 | 0.015047
2019-03-29 13:59 | 1553882324.5416 | 337 | 0.2083 | 0.001436 | 3e-05 | 100.5642 | 100.295 | 120.2924 | 21.1336 | 0.002537 | 0.002396 | 251.7147 | 0.26448 | 8.4027 | 0.002491 | 0.003604 | 0.003611 | 5.8e-05 | 0.001022 | 0.29015 | 1.3444 | 2.5475 | 100.2 | 20.91 | 37.295 | 82.7231 | 81.611 | 79.3317 | 0.009216 | 1.0016 | 0.023776 | 23.7384 | 0.023973 | 0.15734 | 0.006087 | 0.43265 | 0.43265 | 0.014802
2019-03-29 13:59 | 1553882325.5416 | 338 | 0.20829 | 0.001434 | 3e-05 | 100.6088 | 100.5066 | 120.6079 | 21.1333 | 0.002532 | 0.002377 | 251.6885 | 0.26478 | 8.4025 | 0.002425 | 0.00332 | 0.003228 | 5.6e-05 | 0.001012 | 0.28296 | 1.3163 | 2.5473 | 100.2 | 20.91 | 37.297 | 82.6143 | 80.0348 | 78.8785 | 0.009217 | 1.0029 | 0.023823 | 23.7542 | 0.009088 | 0.15752 | 0.00611 | 0.43512 | 0.43512 | 0.014733
2019-03-29 13:59 | 1553882326.5417 | 339 | 0.20827 | 0.001434 | 3e-05 | 100.7027 | 100.4175 | 120.7621 | 21.1343 | 0.002591 | 0.002355 | 251.7006 | 0.26508 | 8.4028 | 0.002422 | 0.003215 | 0.003165 | 6e-05 | 0.001005 | 0.28368 | 1.3336 | 2.5473 | 100.2 | 20.91 | 37.303 | 84.0204 | 79.9741 | 78.354 | 0.009218 | 1.0044 | 0.023856 | 23.7517 | 0.025371 | 0.15769 | 0.006238 | 0.44487 | 0.44487 | 0.014427
2019-03-29 13:59 | 1553882327.5426 | 340 | 0.2083 | 0.001432 | 3e-05 | 100.648 | 100.4981 | 120.6612 | 21.1338 | 0.002531 | 0.002372 | 251.6782 | 0.26416 | 8.4003 | 0.00238 | 0.004356 | 0.004163 | 5.9e-05 | 0.001001 | 0.31581 | 1.3406 | 2.5474 | 100.2 | 20.91 | 37.305 | 82.5953 | 78.9484 | 78.766 | 0.009219 | 1.0032 | 0.023832 | 23.7557 | 0.013333 | 0.15715 | 0.006082 | 0.43332 | 0.43332 | 0.01476
2019-03-29 13:59 | 1553882328.5427 | 341 | 0.20829 | 0.00143 | 3e-05 | 100.6878 | 100.4708 | 120.5053 | 21.1336 | 0.002577 | 0.002323 | 251.6638 | 0.26295 | 8.4008 | 0.002383 | 0.004048 | 0.003994 | 6.1e-05 | 0.000994 | 0.28512 | 1.3469 | 2.5473 | 100.2 | 20.91 | 37.307 | 83.681 | 79.0355 | 77.5895 | 0.009219 | 1.0066 | 0.023856 | 23.7006 | 0.0193 | 0.15643 | 0.006114 | 0.43601 | 0.43601 | 0.014602
2019-03-29 13:59 | 1553882329.5417 | 342 | 0.20829 | 0.001431 | 3.1e-05 | 100.7623 | 100.4086 | 120.5762 | 21.1345 | 0.002596 | 0.002455 | 251.6291 | 0.2637 | 8.4024 | 0.002415 | 0.004204 | 0.004198 | 6.1e-05 | 0.000993 | 0.27975 | 1.341 | 2.5472 | 100.2 | 20.91 | 37.308 | 84.157 | 79.7957 | 80.7493 | 0.009219 | 0.99759 | 0.023757 | 23.8141 | 0.031451 | 0.15688 | 0.006119 | 0.43457 | 0.43457 | 0.014692
2019-03-29 13:59 | 1553882330.5418 | 343 | 0.20829 | 0.001433 | 3.1e-05 | 100.8078 | 100.7462 | 120.4492 | 21.1339 | 0.002606 | 0.002384 | 251.6026 | 0.26404 | 8.4015 | 0.002334 | 0.00367 | 0.003587 | 5.5e-05 | 0.000992 | 0.29103 | 1.3218 | 2.5473 | 100.2 | 20.91 | 37.308 | 84.3976 | 77.8399 | 79.0507 | 0.009219 | 1.0024 | 0.023801 | 23.7444 | 0.005471 | 0.15707 | 0.006099 | 0.43396 | 0.43396 | 0.014732
2019-03-29 13:59 | 1553882331.5428 | 344 | 0.20833 | 0.001426 | 3.1e-05 | 100.8785 | 100.4895 | 120.3445 | 21.1338 | 0.002525 | 0.002428 | 251.5654 | 0.2643 | 8.404 | 0.002433 | 0.004 | 0.003909 | 5.7e-05 | 0.000995 | 0.29574 | 1.3441 | 2.5474 | 100.2 | 20.91 | 37.308 | 82.4473 | 80.2334 | 80.1108 | 0.009219 | 0.99939 | 0.023755 | 23.7698 | 0.034552 | 0.15723 | 0.005901 | 0.41896 | 0.41896 | 0.015274
2019-03-29 13:59 | 1553882332.5429 | 345 | 0.20829 | 0.001423 | 3.1e-05 | 100.9048 | 100.6919 | 120.4204 | 21.1337 | 0.002604 | 0.002323 | 251.5399 | 0.26441 | 8.4018 | 0.002392 | 0.003411 | 0.003383 | 5.8e-05 | 0.001007 | 0.27743 | 1.3852 | 2.5474 | 100.2 | 20.91 | 37.307 | 84.3411 | 79.2359 | 77.5778 | 0.009219 | 1.0066 | 0.023848 | 23.6918 | 0.018891 | 0.1573 | 0.006145 | 0.43802 | 0.43802 | 0.014616
2019-03-29 13:59 | 1553882333.5433 | 346 | 0.20831 | 0.001423 | 3.2e-05 | 100.9323 | 100.7772 | 120.4018 | 21.1336 | 0.00251 | 0.00236 | 251.4907 | 0.26534 | 8.4015 | 0.002414 | 0.003939 | 0.003825 | 6.2e-05 | 0.001008 | 0.28672 | 1.3332 | 2.5474 | 100.2 | 20.91 | 37.307 | 82.0875 | 79.7809 | 78.4773 | 0.009219 | 1.004 | 0.023816 | 23.7204 | 0.013756 | 0.15785 | 0.006005 | 0.42741 | 0.42741 | 0.015031
2019-03-29 13:59 | 1553882334.5418 | 347 | 0.20831 | 0.001424 | 3.2e-05 | 100.8877 | 100.7363 | 120.452 | 21.1326 | 0.002545 | 0.002414 | 251.4357 | 0.26407 | 8.4011 | 0.002382 | 0.003806 | 0.003683 | 6.3e-05 | 0.001005 | 0.29463 | 1.3833 | 2.5475 | 100.2 | 20.91 | 37.304 | 82.9195 | 79.0105 | 79.772 | 0.009218 | 1.0004 | 0.023777 | 23.769 | 0.013438 | 0.15709 | 0.006024 | 0.42808 | 0.42808 | 0.014936
2019-03-29 13:59 | 1553882335.5419 | 348 | 0.20832 | 0.001417 | 3.1e-05 | 100.9201 | 100.7748 | 120.064 | 21.133 | 0.002573 | 0.002381 | 251.424 | 0.26349 | 8.4035 | 0.002393 | 0.00405 | 0.003943 | 5.9e-05 | 0.001002 | 0.29048 | 1.3488 | 2.5474 | 100.2 | 20.91 | 37.302 | 83.59 | 79.2735 | 78.9789 | 0.009218 | 1.0026 | 0.023766 | 23.704 | 0.01288 | 0.15675 | 0.005953 | 0.42286 | 0.42286 | 0.015087
2019-03-29 13:59 | 1553882336.5429 | 349 | 0.20831 | 0.001417 | 3.2e-05 | 101.0335 | 100.6124 | 120.3126 | 21.1334 | 0.002582 | 0.002375 | 251.3802 | 0.26489 | 8.4021 | 0.002429 | 0.003761 | 0.003723 | 5.9e-05 | 0.001 | 0.30359 | 1.3752 | 2.5475 | 100.2 | 20.91 | 37.3 | 83.8173 | 80.1291 | 78.8365 | 0.009217 | 1.003 | 0.023795 | 23.7237 | 0.03736 | 0.15759 | 0.005997 | 0.42649 | 0.42649 | 0.015038
2019-03-29 13:59 | 1553882337.543 | 350 | 0.2083 | 0.001411 | 3.1e-05 | 100.9973 | 100.9223 | 120.7868 | 21.133 | 0.002564 | 0.00234 | 251.356 | 0.26426 | 8.4021 | 0.002373 | 0.003666 | 0.003611 | 6.1e-05 | 0.000999 | 0.27793 | 1.3215 | 2.5473 | 100.2 | 20.91 | 37.298 | 83.3676 | 78.7956 | 77.9904 | 0.009217 | 1.0054 | 0.023871 | 23.7418 | 0.006646 | 0.15721 | 0.006092 | 0.43464 | 0.43464 | 0.014721
2019-03-29 13:59 | 1553882338.542 | 351 | 0.2083 | 0.001415 | 3.1e-05 | 101.0844 | 100.8677 | 120.5781 | 21.1337 | 0.002583 | 0.002419 | 251.3307 | 0.26458 | 8.4026 | 0.002382 | 0.004066 | 0.004067 | 6.1e-05 | 0.000998 | 0.26952 | 1.3023 | 2.5473 | 100.2 | 20.91 | 37.296 | 83.8239 | 78.999 | 79.8855 | 0.009216 | 1 | 0.023786 | 23.7852 | 0.019197 | 0.1574 | 0.006064 | 0.4311 | 0.4311 | 0.01486
2019-03-29 13:59 | 1553882339.5421 | 352 | 0.20831 | 0.001413 | 3.2e-05 | 101.0829 | 100.8195 | 120.4915 | 21.133 | 0.002545 | 0.002371 | 251.298 | 0.26436 | 8.4005 | 0.002392 | 0.003502 | 0.003394 | 6.2e-05 | 0.001002 | 0.30663 | 1.3503 | 2.5475 | 100.2 | 20.91 | 37.291 | 82.9237 | 79.2346 | 78.7386 | 0.009215 | 1.0033 | 0.023816 | 23.738 | 0.023333 | 0.15727 | 0.006042 | 0.4301 | 0.4301 | 0.014882
2019-03-29 13:59 | 1553882340.5421 | 353 | 0.20831 | 0.001407 | 3.2e-05 | 101.1149 | 100.8151 | 120.4757 | 21.1333 | 0.002532 | 0.002362 | 251.2727 | 0.26435 | 8.402 | 0.002391 | 0.003544 | 0.003524 | 6.1e-05 | 0.001002 | 0.30359 | 1.3516 | 2.5473 | 100.2 | 20.91 | 37.288 | 82.6011 | 79.2238 | 78.527 | 0.009214 | 1.0039 | 0.023822 | 23.7293 | 0.026561 | 0.15726 | 0.00602 | 0.42856 | 0.42856 | 0.014935
2019-03-29 13:59 | 1553882341.5421 | 354 | 0.20832 | 0.001405 | 3.3e-05 | 101.1095 | 100.8704 | 120.6767 | 21.1323 | 0.002527 | 0.002376 | 251.2671 | 0.26472 | 8.4 | 0.002378 | 0.004087 | 0.003966 | 6.2e-05 | 0.001008 | 0.29551 | 1.3166 | 2.5474 | 100.2 | 20.91 | 37.284 | 82.4761 | 78.9127 | 78.855 | 0.009213 | 1.003 | 0.02383 | 23.7602 | 0.02118 | 0.15749 | 0.006004 | 0.42755 | 0.42755 | 0.014992
2019-03-29 13:59 | 1553882342.5412 | 355 | 0.2083 | 0.001404 | 3.2e-05 | 101.1437 | 100.9509 | 120.6203 | 21.1331 | 0.002641 | 0.002355 | 251.2515 | 0.26453 | 8.4012 | 0.002393 | 0.003195 | 0.003136 | 6.1e-05 | 0.001013 | 0.29391 | 1.3214 | 2.5474 | 100.2 | 20.91 | 37.281 | 85.2163 | 79.2768 | 78.342 | 0.009213 | 1.0044 | 0.023842 | 23.7373 | 0.017059 | 0.15737 | 0.00611 | 0.43536 | 0.43536 | 0.014712
2019-03-29 13:59 | 1553882343.5412 | 356 | 0.20828 | 0.001401 | 3.2e-05 | 101.1754 | 100.8588 | 120.5457 | 21.1322 | 0.002583 | 0.002417 | 251.2563 | 0.26459 | 8.4025 | 0.002365 | 0.003878 | 0.003867 | 6.2e-05 | 0.00101 | 0.30431 | 1.3268 | 2.5474 | 100.2 | 20.91 | 37.273 | 83.8411 | 78.5992 | 79.8465 | 0.009211 | 1.0001 | 0.023784 | 23.7807 | 0.028033 | 0.1574 | 0.006213 | 0.44169 | 0.44169 | 0.014504
2019-03-29 13:59 | 1553882344.5422 | 357 | 0.20832 | 0.001399 | 3.3e-05 | 101.1781 | 100.8489 | 120.3713 | 21.1334 | 0.002579 | 0.002381 | 251.2505 | 0.26461 | 8.4029 | 0.002403 | 0.004223 | 0.004144 | 6.2e-05 | 0.001011 | 0.24762 | 1.3322 | 2.5473 | 100.2 | 20.91 | 37.266 | 83.7339 | 79.5146 | 78.9862 | 0.009209 | 1.0026 | 0.023796 | 23.7345 | 0.029149 | 0.15742 | 0.005959 | 0.42373 | 0.42373 | 0.01512
2019-03-29 13:59 | 1553882345.5423 | 358 | 0.20831 | 0.001397 | 3.3e-05 | 101.1978 | 100.9629 | 120.7177 | 21.1323 | 0.002503 | 0.002379 | 251.229 | 0.26526 | 8.4013 | 0.002429 | 0.003994 | 0.003877 | 6.5e-05 | 0.001015 | 0.29286 | 1.341 | 2.5474 | 100.2 | 20.91 | 37.26 | 81.9201 | 80.1406 | 78.9234 | 0.009207 | 1.0028 | 0.023832 | 23.7665 | 0.020791 | 0.15781 | 0.00605 | 0.43084 | 0.43084 | 0.014907
2019-03-29 13:59 | 1553882346.5423 | 359 | 0.2083 | 0.001403 | 3.3e-05 | 101.2338 | 101.0496 | 120.7079 | 21.1322 | 0.002503 | 0.002414 | 251.2076 | 0.26521 | 8.4001 | 0.002312 | 0.003936 | 0.003892 | 6e-05 | 0.001015 | 0.31238 | 1.3921 | 2.5474 | 100.2 | 20.91 | 37.241 | 81.9145 | 77.3284 | 79.7733 | 0.009203 | 1.0003 | 0.023802 | 23.7942 | 0.016288 | 0.15777 | 0.00611 | 0.43459 | 0.43459 | 0.014776
2019-03-29 13:59 | 1553882347.5414 | 360 | 0.20833 | 0.001397 | 3.4e-05 | 101.1823 | 101.0329 | 120.5553 | 21.1327 | 0.002566 | 0.002387 | 251.1904 | 0.2648 | 8.4009 | 0.00242 | 0.003974 | 0.003946 | 6.1e-05 | 0.001022 | 0.29734 | 1.351 | 2.5476 | 100.2 | 20.91 | 37.241 | 83.425 | 79.9141 | 79.1274 | 0.009203 | 1.0022 | 0.023809 | 23.7574 | 0.013219 | 0.15753 | 0.005925 | 0.42147 | 0.42147 | 0.015212
2019-03-29 13:59 | 1553882348.5413 | 361 | 0.20836 | 0.001397 | 3.3e-05 | 101.2059 | 100.9963 | 120.5987 | 21.1328 | 0.002532 | 0.002372 | 251.1625 | 0.26481 | 8.4014 | 0.002368 | 0.003552 | 0.003478 | 6.2e-05 | 0.00102 | 0.2959 | 1.3729 | 2.5475 | 100.2 | 20.91 | 37.217 | 82.6069 | 78.6582 | 78.7682 | 0.009197 | 1.0032 | 0.023826 | 23.7496 | 0.018542 | 0.15754 | 0.005772 | 0.41086 | 0.41086 | 0.015605
2019-03-29 13:59 | 1553882349.5425 | 362 | 0.20834 | 0.001401 | 3.3e-05 | 101.2472 | 100.9643 | 120.7405 | 21.1329 | 0.002584 | 0.002403 | 251.1413 | 0.26456 | 8.4035 | 0.002421 | 0.003504 | 0.003457 | 6.1e-05 | 0.001022 | 0.30718 | 1.3384 | 2.5473 | 100.2 | 20.91 | 37.217 | 83.852 | 79.9415 | 79.5089 | 0.009197 | 1.0011 | 0.023815 | 23.7885 | 0.025034 | 0.15739 | 0.005861 | 0.41705 | 0.41705 | 0.01536
2019-03-29 13:59 | 1553882350.5424 | 363 | 0.20836 | 0.001392 | 3.3e-05 | 101.2427 | 101.2011 | 121.0917 | 21.1319 | 0.002565 | 0.002388 | 251.1065 | 0.26528 | 8.401 | 0.002445 | 0.004107 | 0.004121 | 6.4e-05 | 0.001031 | 0.28655 | 1.3543 | 2.5475 | 100.2 | 20.91 | 37.209 | 83.3963 | 80.5043 | 79.154 | 0.009195 | 1.0021 | 0.023861 | 23.8111 | 0.003679 | 0.15781 | 0.005733 | 0.40873 | 0.40873 | 0.015715
2019-03-29 13:59 | 1553882351.5425 | 364 | 0.20837 | 0.001392 | 3.3e-05 | 101.2694 | 100.9011 | 120.9433 | 21.1326 | 0.002556 | 0.002372 | 251.0597 | 0.26387 | 8.4018 | 0.002545 | 0.004181 | 0.004058 | 6.2e-05 | 0.001032 | 0.31509 | 1.3321 | 2.5474 | 100.2 | 20.91 | 37.194 | 83.1941 | 82.9101 | 78.7625 | 0.009191 | 1.0032 | 0.02386 | 23.7833 | 0.032586 | 0.15698 | 0.005719 | 0.40766 | 0.40766 | 0.015672
2019-03-29 13:59 | 1553882352.5415 | 365 | 0.20835 | 0.001394 | 3.3e-05 | 101.299 | 100.9083 | 120.952 | 21.1324 | 0.002533 | 0.002414 | 251.0192 | 0.26511 | 8.4017 | 0.002418 | 0.004376 | 0.004337 | 6.4e-05 | 0.001034 | 0.29391 | 1.3334 | 2.5475 | 100.2 | 20.91 | 37.187 | 82.6402 | 79.8771 | 79.7718 | 0.009189 | 1.0004 | 0.023827 | 23.8182 | 0.034568 | 0.15772 | 0.005801 | 0.41296 | 0.41296 | 0.015544
2019-03-29 13:59 | 1553882353.5415 | 366 | 0.2083 | 0.001386 | 3.3e-05 | 101.2176 | 100.9895 | 120.7454 | 21.132 | 0.002539 | 0.002326 | 250.9867 | 0.26486 | 8.4023 | 0.002474 | 0.003854 | 0.003777 | 6.4e-05 | 0.001036 | 0.29159 | 1.33 | 2.5474 | 100.2 | 20.91 | 37.174 | 82.7636 | 81.2038 | 77.6564 | 0.009186 | 1.0064 | 0.023878 | 23.7265 | 0.020179 | 0.15757 | 0.006094 | 0.43483 | 0.43483 | 0.014748
2019-03-29 13:59 | 1553882354.5416 | 367 | 0.20837 | 0.001387 | 3.3e-05 | 101.2062 | 101.0311 | 120.9846 | 21.1328 | 0.002534 | 0.002414 | 250.941 | 0.265 | 8.4018 | 0.002436 | 0.003963 | 0.003899 | 6.1e-05 | 0.001045 | 0.30254 | 1.359 | 2.5475 | 100.2 | 20.91 | 37.168 | 82.6569 | 80.2939 | 79.7596 | 0.009185 | 1.0004 | 0.02383 | 23.821 | 0.015486 | 0.15765 | 0.005689 | 0.405 | 0.405 | 0.015843
2019-03-29 13:59 | 1553882355.5416 | 368 | 0.20834 | 0.001383 | 3.4e-05 | 101.2155 | 100.9823 | 120.9248 | 21.1326 | 0.002594 | 0.002347 | 250.9211 | 0.26495 | 8.4021 | 0.002495 | 0.00405 | 0.003988 | 6.2e-05 | 0.00105 | 0.30381 | 1.3815 | 2.5474 | 100.2 | 20.91 | 37.164 | 84.0982 | 81.7123 | 78.1669 | 0.009184 | 1.0049 | 0.023878 | 23.7613 | 0.020628 | 0.15762 | 0.005878 | 0.41931 | 0.41931 | 0.015299
2019-03-29 13:59 | 1553882356.5406 | 369 | 0.20834 | 0.001378 | 3.4e-05 | 101.2781 | 100.9782 | 120.8964 | 21.1327 | 0.00254 | 0.002378 | 250.904 | 0.26525 | 8.4012 | 0.002509 | 0.003751 | 0.003623 | 6.1e-05 | 0.00105 | 0.32947 | 1.3544 | 2.5475 | 100.2 | 20.91 | 37.156 | 82.7962 | 82.0616 | 78.9036 | 0.009182 | 1.0028 | 0.02385 | 23.7835 | 0.026518 | 0.15779 | 0.005878 | 0.41886 | 0.41886 | 0.015333
2019-03-29 13:59 | 1553882357.5406 | 370 | 0.20841 | 0.001376 | 3.4e-05 | 101.2236 | 100.9694 | 120.8304 | 21.1325 | 0.002575 | 0.002326 | 250.8977 | 0.26465 | 8.4006 | 0.002473 | 0.004324 | 0.00423 | 6.1e-05 | 0.001049 | 0.31813 | 1.3995 | 2.5475 | 100.2 | 20.92 | 37.152 | 83.644 | 81.1965 | 77.6569 | 0.009186 | 1.0064 | 0.023886 | 23.7348 | 0.022491 | 0.15744 | 0.005501 | 0.39242 | 0.39242 | 0.016329
2019-03-29 13:59 | 1553882358.5407 | 371 | 0.20832 | 0.001376 | 3.4e-05 | 101.2212 | 101.1653 | 121.0396 | 21.1328 | 0.00255 | 0.002326 | 250.8865 | 0.26441 | 8.403 | 0.002423 | 0.004227 | 0.004196 | 5.8e-05 | 0.001049 | 0.31653 | 1.4171 | 2.5475 | 100.2 | 20.92 | 37.146 | 83.0305 | 79.9983 | 77.644 | 0.009185 | 1.0064 | 0.023907 | 23.7549 | 0.004942 | 0.15729 | 0.006007 | 0.42905 | 0.42905 | 0.014921
2019-03-29 13:59 | 1553882359.5407 | 372 | 0.20832 | 0.001369 | 3.4e-05 | 101.2933 | 100.9496 | 120.7161 | 21.1329 | 0.002546 | 0.002366 | 250.876 | 0.26423 | 8.4021 | 0.00248 | 0.0042 | 0.004104 | 5.6e-05 | 0.001043 | 0.26161 | 1.3449 | 2.5473 | 100.2 | 20.92 | 37.143 | 82.9486 | 81.363 | 78.627 | 0.009184 | 1.0036 | 0.023842 | 23.7564 | 0.030406 | 0.15719 | 0.006052 | 0.43114 | 0.43114 | 0.014839
2019-03-29 13:59 | 1553882360.5407 | 373 | 0.20836 | 0.001372 | 3.5e-05 | 101.1871 | 100.9161 | 120.46 | 21.1324 | 0.002533 | 0.002318 | 250.8766 | 0.26444 | 8.4038 | 0.002385 | 0.004279 | 0.004283 | 5.7e-05 | 0.001037 | 0.28567 | 1.3353 | 2.5475 | 100.2 | 20.92 | 37.141 | 82.6315 | 79.0639 | 77.469 | 0.009184 | 1.0069 | 0.023856 | 23.6921 | 0.023981 | 0.15731 | 0.005754 | 0.41004 | 0.41004 | 0.015615
2019-03-29 13:59 | 1553882361.5407 | 374 | 0.20838 | 0.001373 | 3.5e-05 | 101.238 | 101.0189 | 120.5621 | 21.1327 | 0.002481 | 0.002292 | 250.871 | 0.26454 | 8.4014 | 0.002425 | 0.003557 | 0.003407 | 5.9e-05 | 0.001036 | 0.32493 | 1.3227 | 2.5475 | 100.2 | 20.92 | 37.137 | 81.3784 | 80.0294 | 76.827 | 0.009183 | 1.0088 | 0.023888 | 23.6804 | 0.019378 | 0.15737 | 0.00566 | 0.4038 | 0.4038 | 0.015862
2019-03-29 13:59 | 1553882362.5408 | 375 | 0.20836 | 0.001368 | 3.6e-05 | 101.2251 | 100.9083 | 120.5436 | 21.1324 | 0.002519 | 0.002359 | 250.8697 | 0.26384 | 8.4032 | 0.00252 | 0.003988 | 0.00383 | 5.8e-05 | 0.00103 | 0.30469 | 1.3327 | 2.5474 | 100.2 | 20.92 | 37.135 | 82.2993 | 82.3257 | 78.4418 | 0.009182 | 1.0041 | 0.023831 | 23.7331 | 0.028037 | 0.15696 | 0.005754 | 0.40953 | 0.40953 | 0.015599
2019-03-29 13:59 | 1553882363.5414 | 376 | 0.20833 | 0.00137 | 3.6e-05 | 101.2161 | 101.0114 | 120.9781 | 21.1317 | 0.002569 | 0.002363 | 250.8632 | 0.26508 | 8.405 | 0.002484 | 0.00416 | 0.004158 | 5.7e-05 | 0.001027 | 0.31797 | 1.3455 | 2.5476 | 100.2 | 20.92 | 37.131 | 83.4914 | 81.4564 | 78.5411 | 0.009181 | 1.0039 | 0.023871 | 23.7792 | 0.018112 | 0.15769 | 0.005943 | 0.42375 | 0.42375 | 0.015146
2019-03-29 13:59 | 1553882364.5419 | 377 | 0.20842 | 0.001366 | 3.6e-05 | 101.1104 | 100.9277 | 121.0917 | 21.1318 | 0.00251 | 0.002388 | 250.8627 | 0.26402 | 8.4015 | 0.002489 | 0.003893 | 0.003913 | 5.7e-05 | 0.001029 | 0.29662 | 1.3239 | 2.5474 | 100.2 | 20.92 | 37.131 | 82.0764 | 81.5733 | 79.1383 | 0.009181 | 1.0021 | 0.023862 | 23.8106 | 0.016177 | 0.15707 | 0.005436 | 0.38729 | 0.38729 | 0.016506
2019-03-29 13:59 | 1553882365.5419 | 378 | 0.20835 | 0.001366 | 3.6e-05 | 101.1565 | 100.9881 | 120.8788 | 21.1312 | 0.00246 | 0.002344 | 250.8391 | 0.26425 | 8.4057 | 0.002487 | 0.00383 | 0.003739 | 5.6e-05 | 0.001021 | 0.28208 | 1.3473 | 2.5475 | 100.2 | 20.92 | 37.13 | 80.8686 | 81.5197 | 78.0845 | 0.009181 | 1.0052 | 0.023877 | 23.754 | 0.014902 | 0.1572 | 0.005832 | 0.41584 | 0.41584 | 0.015386
2019-03-29 13:59 | 1553882366.5409 | 379 | 0.20838 | 0.001364 | 3.7e-05 | 101.1179 | 100.7318 | 120.617 | 21.1311 | 0.002593 | 0.002351 | 250.8228 | 0.26467 | 8.4039 | 0.002468 | 0.004592 | 0.004506 | 5.6e-05 | 0.001021 | 0.28479 | 1.3533 | 2.5475 | 100.2 | 20.92 | 37.126 | 84.0782 | 81.0548 | 78.2449 | 0.00918 | 1.0047 | 0.023845 | 23.7337 | 0.034217 | 0.15745 | 0.005693 | 0.40536 | 0.40536 | 0.015809
2019-03-29 13:59 | 1553882367.542 | 380 | 0.20837 | 0.001367 | 3.7e-05 | 101.0808 | 100.9673 | 120.7959 | 21.1318 | 0.002486 | 0.002376 | 250.7931 | 0.26459 | 8.4004 | 0.00236 | 0.00395 | 0.003885 | 5.5e-05 | 0.001022 | 0.32947 | 1.3442 | 2.5475 | 100.2 | 20.92 | 37.124 | 81.5072 | 78.4624 | 78.8608 | 0.00918 | 1.0029 | 0.023842 | 23.7721 | 0.010047 | 0.15741 | 0.005744 | 0.40896 | 0.40896 | 0.015665
2019-03-29 13:59 | 1553882368.542 | 381 | 0.20836 | 0.001362 | 3.7e-05 | 101.1529 | 100.8172 | 120.8215 | 21.1314 | 0.002487 | 0.002344 | 250.753 | 0.26465 | 8.4045 | 0.002448 | 0.004179 | 0.004102 | 5.4e-05 | 0.001023 | 0.28152 | 1.371 | 2.5476 | 100.2 | 20.93 | 37.123 | 81.5253 | 80.5872 | 78.0761 | 0.009185 | 1.0052 | 0.023871 | 23.7481 | 0.029741 | 0.15744 | 0.005816 | 0.41457 | 0.41457 | 0.015457
2019-03-29 13:59 | 1553882369.541 | 382 | 0.2084 | 0.001357 | 3.7e-05 | 101.1128 | 100.7516 | 121.1027 | 21.1322 | 0.002524 | 0.002378 | 250.7483 | 0.26481 | 8.4025 | 0.002345 | 0.003571 | 0.003528 | 5.5e-05 | 0.001024 | 0.3095 | 1.3418 | 2.5474 | 100.2 | 20.93 | 37.12 | 82.4151 | 78.1088 | 78.8976 | 0.009184 | 1.0028 | 0.023871 | 23.8035 | 0.032009 | 0.15754 | 0.005543 | 0.39505 | 0.39505 | 0.01623
2019-03-29 13:59 | 1553882370.5401 | 383 | 0.20839 | 0.001353 | 3.7e-05 | 101.0155 | 100.866 | 121.1569 | 21.1321 | 0.002587 | 0.002346 | 250.721 | 0.26413 | 8.4015 | 0.002408 | 0.003456 | 0.003411 | 5.5e-05 | 0.001024 | 0.33035 | 1.3698 | 2.5474 | 100.2 | 20.92 | 37.117 | 83.9356 | 79.6273 | 78.1274 | 0.009178 | 1.005 | 0.023903 | 23.7828 | 0.013255 | 0.15713 | 0.005597 | 0.39942 | 0.39942 | 0.016011
2019-03-29 14:00 | 1553882371.5407 | 384 | 0.20838 | 0.001353 | 3.7e-05 | 101.0296 | 100.7523 | 121.2001 | 21.1317 | 0.002563 | 0.002277 | 250.7168 | 0.26491 | 8.4017 | 0.002401 | 0.003578 | 0.003519 | 5.8e-05 | 0.001021 | 0.30934 | 1.3569 | 2.5475 | 100.2 | 20.93 | 37.115 | 83.3579 | 79.464 | 76.4666 | 0.009183 | 1.0098 | 0.023963 | 23.7307 | 0.024584 | 0.15759 | 0.005668 | 0.40558 | 0.40558 | 0.015815
2019-03-29 14:00 | 1553882372.5411 | 385 | 0.20841 | 0.001344 | 3.7e-05 | 101.0598 | 100.783 | 121.5961 | 21.1314 | 0.002515 | 0.002279 | 250.6787 | 0.26413 | 8.4041 | 0.002379 | 0.004201 | 0.004121 | 5.4e-05 | 0.00102 | 0.31222 | 1.3655 | 2.5474 | 100.2 | 20.92 | 37.114 | 82.2056 | 78.9337 | 76.5247 | 0.009177 | 1.0096 | 0.024001 | 23.7715 | 0.024534 | 0.15713 | 0.005531 | 0.39628 | 0.39628 | 0.016138
2019-03-29 14:00 | 1553882373.5411 | 386 | 0.20843 | 0.001344 | 3.7e-05 | 101.038 | 100.8213 | 121.6468 | 21.1309 | 0.002538 | 0.002351 | 250.6536 | 0.2638 | 8.4008 | 0.002362 | 0.003607 | 0.003531 | 5.5e-05 | 0.001014 | 0.28998 | 1.347 | 2.5474 | 100.2 | 20.93 | 37.112 | 82.7493 | 78.5206 | 78.2628 | 0.009182 | 1.0046 | 0.023946 | 23.8354 | 0.019208 | 0.15693 | 0.005395 | 0.38567 | 0.38567 | 0.016561
2019-03-29 14:00 | 1553882374.5402 | 387 | 0.20842 | 0.001335 | 3.8e-05 | 100.9296 | 100.566 | 121.4259 | 21.131 | 0.002574 | 0.002345 | 250.6112 | 0.26407 | 8.4024 | 0.002351 | 0.004158 | 0.004143 | 5.3e-05 | 0.001008 | 0.26128 | 1.3422 | 2.5475 | 100.2 | 20.92 | 37.11 | 83.6145 | 78.2448 | 78.1132 | 0.009176 | 1.0051 | 0.02393 | 23.8087 | 0.032287 | 0.1571 | 0.005448 | 0.38911 | 0.38911 | 0.016432
2019-03-29 14:00 | 1553882375.5412 | 388 | 0.2084 | 0.001335 | 3.8e-05 | 100.9374 | 100.6431 | 121.4621 | 21.1306 | 0.002557 | 0.00235 | 250.5843 | 0.26421 | 8.4024 | 0.002363 | 0.003785 | 0.003782 | 5.5e-05 | 0.001008 | 0.30055 | 1.338 | 2.5476 | 100.2 | 20.92 | 37.109 | 83.2154 | 78.5313 | 78.2211 | 0.009176 | 1.0048 | 0.023929 | 23.8159 | 0.026119 | 0.15718 | 0.0056 | 0.40005 | 0.40005 | 0.015991
2019-03-29 14:00 | 1553882376.5412 | 389 | 0.20843 | 0.001333 | 3.8e-05 | 100.9341 | 100.609 | 121.6144 | 21.1306 | 0.002491 | 0.002402 | 250.5565 | 0.26445 | 8.4032 | 0.002372 | 0.003738 | 0.003631 | 5.3e-05 | 0.001011 | 0.29983 | 1.3326 | 2.5474 | 100.2 | 20.92 | 37.108 | 81.6165 | 78.7473 | 79.4709 | 0.009176 | 1.0012 | 0.023902 | 23.8732 | 0.028861 | 0.15732 | 0.005392 | 0.38463 | 0.38463 | 0.016647
2019-03-29 14:00 | 1553882377.5413 | 390 | 0.20847 | 0.001332 | 3.8e-05 | 100.83 | 100.5373 | 121.582 | 21.1309 | 0.002508 | 0.002351 | 250.5079 | 0.26444 | 8.4027 | 0.002391 | 0.003592 | 0.003611 | 5.3e-05 | 0.001009 | 0.30254 | 1.29 | 2.5474 | 100.2 | 20.92 | 37.107 | 82.0303 | 79.2239 | 78.2524 | 0.009175 | 1.0047 | 0.02394 | 23.8287 | 0.025999 | 0.15731 | 0.005157 | 0.36838 | 0.36838 | 0.017381
2019-03-29 14:00 | 1553882378.5403 | 391 | 0.20845 | 0.001323 | 3.8e-05 | 100.8518 | 100.5407 | 121.5383 | 21.1312 | 0.002528 | 0.002393 | 250.4832 | 0.26361 | 8.4024 | 0.002355 | 0.003491 | 0.003479 | 5.4e-05 | 0.001004 | 0.31349 | 1.3056 | 2.5474 | 100.2 | 20.92 | 37.108 | 82.5017 | 78.3433 | 79.2523 | 0.009176 | 1.0018 | 0.023902 | 23.8583 | 0.027634 | 0.15682 | 0.005306 | 0.37852 | 0.37852 | 0.016862
2019-03-29 14:00 | 1553882379.5403 | 392 | 0.20846 | 0.001325 | 3.8e-05 | 100.8009 | 100.494 | 121.3299 | 21.1305 | 0.002541 | 0.002351 | 250.4305 | 0.26435 | 8.4042 | 0.00238 | 0.003338 | 0.003199 | 5.3e-05 | 0.000999 | 0.30702 | 1.3428 | 2.5475 | 100.2 | 20.92 | 37.108 | 82.8097 | 78.9436 | 78.2499 | 0.009176 | 1.0047 | 0.023915 | 23.8039 | 0.02728 | 0.15726 | 0.005237 | 0.37371 | 0.37371 | 0.017127
2019-03-29 14:00 | 1553882380.5414 | 393 | 0.20846 | 0.001319 | 3.8e-05 | 100.7973 | 100.5131 | 120.9874 | 21.1305 | 0.002553 | 0.002351 | 250.4264 | 0.26407 | 8.4034 | 0.002441 | 0.004229 | 0.004125 | 5.7e-05 | 0.001002 | 0.29623 | 1.3221 | 2.5474 | 100.2 | 20.92 | 37.108 | 83.0992 | 80.414 | 78.2614 | 0.009176 | 1.0047 | 0.023881 | 23.7707 | 0.025261 | 0.1571 | 0.005226 | 0.37237 | 0.37237 | 0.017171
2019-03-29 14:00 | 1553882381.5405 | 394 | 0.20843 | 0.001324 | 3.9e-05 | 100.7569 | 100.5035 | 120.9431 | 21.1301 | 0.002507 | 0.002356 | 250.3997 | 0.26394 | 8.4033 | 0.002349 | 0.004459 | 0.004373 | 5.5e-05 | 0.001 | 0.28191 | 1.3633 | 2.5473 | 100.2 | 20.92 | 37.109 | 81.9982 | 78.2057 | 78.3765 | 0.009176 | 1.0043 | 0.023873 | 23.7702 | 0.022524 | 0.15702 | 0.005403 | 0.38491 | 0.38491 | 0.016603
2019-03-29 14:00 | 1553882382.5404 | 395 | 0.20846 | 0.001315 | 3.8e-05 | 100.8042 | 100.4407 | 120.5147 | 21.1304 | 0.002488 | 0.002325 | 250.3843 | 0.26586 | 8.4037 | 0.00236 | 0.003823 | 0.003727 | 5.9e-05 | 0.001003 | 0.29645 | 1.3542 | 2.5475 | 100.2 | 20.92 | 37.112 | 81.5548 | 78.4641 | 77.6222 | 0.009177 | 1.0065 | 0.023856 | 23.7026 | 0.032312 | 0.15816 | 0.005267 | 0.37497 | 0.37497 | 0.017167
2019-03-29 14:00 | 1553882383.5414 | 396 | 0.20846 | 0.001315 | 4e-05 | 100.7057 | 100.4489 | 120.0523 | 21.1302 | 0.002561 | 0.002391 | 250.3366 | 0.26477 | 8.4004 | 0.002337 | 0.003449 | 0.00347 | 5.7e-05 | 0.001013 | 0.30989 | 1.3368 | 2.5474 | 100.2 | 20.92 | 37.115 | 83.3095 | 77.9096 | 79.2225 | 0.009177 | 1.0019 | 0.023756 | 23.711 | 0.02284 | 0.15751 | 0.005228 | 0.37052 | 0.37052 | 0.017302
2019-03-29 14:00 | 1553882384.5405 | 397 | 0.20846 | 0.001309 | 4e-05 | 100.7569 | 100.3182 | 119.9279 | 21.1307 | 0.002535 | 0.002306 | 250.3337 | 0.26576 | 8.4028 | 0.002376 | 0.004029 | 0.003928 | 5.7e-05 | 0.001007 | 0.29839 | 1.3702 | 2.5475 | 100.2 | 20.92 | 37.121 | 82.6883 | 78.8635 | 77.1852 | 0.009179 | 1.0077 | 0.023813 | 23.6301 | 0.039033 | 0.1581 | 0.005259 | 0.37362 | 0.37362 | 0.017223
2019-03-29 14:00 | 1553882385.5415 | 398 | 0.20845 | 0.001312 | 4e-05 | 100.6986 | 100.4892 | 119.8544 | 21.1309 | 0.002478 | 0.00234 | 250.3129 | 0.26497 | 8.4031 | 0.002383 | 0.004076 | 0.00404 | 5.8e-05 | 0.001002 | 0.30879 | 1.3633 | 2.5476 | 100.2 | 20.92 | 37.124 | 81.3041 | 79.0134 | 77.9905 | 0.00918 | 1.0054 | 0.023778 | 23.65 | 0.018617 | 0.15763 | 0.005326 | 0.37783 | 0.37783 | 0.01698
2019-03-29 14:00 | 1553882386.5396 | 399 | 0.20848 | 0.001307 | 4e-05 | 100.6824 | 100.2308 | 120.0406 | 21.1305 | 0.002513 | 0.002333 | 250.3 | 0.26471 | 8.4034 | 0.002444 | 0.003777 | 0.003734 | 5.6e-05 | 0.000996 | 0.29214 | 1.3664 | 2.5474 | 100.2 | 20.92 | 37.128 | 82.1585 | 80.4973 | 77.8164 | 0.009181 | 1.0059 | 0.023803 | 23.6625 | 0.040207 | 0.15747 | 0.005157 | 0.36609 | 0.36609 | 0.017507
2019-03-29 14:00 | 1553882387.5396 | 400 | 0.20845 | 0.00131 | 4.1e-05 | 100.633 | 100.3527 | 120.0563 | 21.1306 | 0.002565 | 0.002336 | 250.3027 | 0.26482 | 8.4027 | 0.002361 | 0.004617 | 0.004555 | 5.9e-05 | 0.000991 | 0.30663 | 1.3552 | 2.5475 | 100.2 | 20.92 | 37.136 | 83.4083 | 78.4946 | 77.8874 | 0.009182 | 1.0057 | 0.023802 | 23.6664 | 0.024952 | 0.15754 | 0.005299 | 0.37625 | 0.37625 | 0.017042
2019-03-29 14:00 | 1553882388.5407 | 401 | 0.20842 | 0.001311 | 4.1e-05 | 100.6228 | 100.1438 | 120.0138 | 21.1297 | 0.002548 | 0.002355 | 250.2802 | 0.26485 | 8.4013 | 0.002367 | 0.003505 | 0.003444 | 6e-05 | 0.000989 | 0.30198 | 1.338 | 2.5474 | 100.2 | 20.92 | 37.14 | 82.9876 | 78.6349 | 78.3512 | 0.009183 | 1.0044 | 0.023782 | 23.6779 | 0.042683 | 0.15756 | 0.005477 | 0.38864 | 0.38864 | 0.0165
2019-03-29 14:00 | 1553882389.5396 | 402 | 0.20846 | 0.001308 | 4.1e-05 | 100.6546 | 100.2946 | 119.9826 | 21.1298 | 0.002521 | 0.002347 | 250.2708 | 0.26512 | 8.4013 | 0.002351 | 0.003581 | 0.003525 | 5.8e-05 | 0.000988 | 0.31885 | 1.3158 | 2.5474 | 100.2 | 20.92 | 37.148 | 82.3294 | 78.2638 | 78.1648 | 0.009185 | 1.0049 | 0.023785 | 23.6685 | 0.03204 | 0.15772 | 0.005264 | 0.37344 | 0.37344 | 0.01719
2019-03-29 14:00 | 1553882390.5407 | 403 | 0.20846 | 0.001305 | 4.1e-05 | 100.5498 | 100.1073 | 120.2455 | 21.1301 | 0.002585 | 0.002337 | 250.2166 | 0.26508 | 8.402 | 0.002359 | 0.004211 | 0.004099 | 5.9e-05 | 0.000992 | 0.32084 | 1.3462 | 2.5474 | 100.2 | 20.93 | 37.152 | 83.8788 | 78.4419 | 77.9095 | 0.009192 | 1.0057 | 0.02382 | 23.6858 | 0.039451 | 0.15769 | 0.005278 | 0.37496 | 0.37496 | 0.017117
2019-03-29 14:00 | 1553882391.5407 | 404 | 0.20844 | 0.001301 | 4.2e-05 | 100.5866 | 100.0691 | 120.3971 | 21.1296 | 0.002565 | 0.002379 | 250.1884 | 0.26483 | 8.4026 | 0.002272 | 0.004352 | 0.004302 | 5.9e-05 | 0.000991 | 0.32117 | 1.3352 | 2.5473 | 100.2 | 20.92 | 37.155 | 83.3972 | 76.3517 | 78.9365 | 0.009187 | 1.0027 | 0.0238 | 23.7354 | 0.046141 | 0.15755 | 0.005371 | 0.38128 | 0.38128 | 0.016818
2019-03-29 14:00 | 1553882392.5398 | 405 | 0.20848 | 0.001306 | 4.2e-05 | 100.5417 | 100.1264 | 120.8147 | 21.1295 | 0.002528 | 0.002437 | 250.1552 | 0.26453 | 8.4024 | 0.002264 | 0.003895 | 0.003803 | 5.9e-05 | 0.000995 | 0.33124 | 1.3256 | 2.5474 | 100.2 | 20.92 | 37.162 | 82.4992 | 76.169 | 80.3203 | 0.009189 | 0.9988 | 0.023795 | 23.8232 | 0.037024 | 0.15737 | 0.005158 | 0.36598 | 0.36598 | 0.017501
2019-03-29 14:00 | 1553882393.5398 | 406 | 0.20847 | 0.001304 | 4.2e-05 | 100.5046 | 100.1728 | 121.252 | 21.1298 | 0.002556 | 0.002297 | 250.1498 | 0.26477 | 8.4019 | 0.002314 | 0.004019 | 0.003989 | 5.8e-05 | 0.000995 | 0.33555 | 1.3332 | 2.5474 | 100.2 | 20.92 | 37.166 | 83.1733 | 77.3542 | 76.9614 | 0.00919 | 1.0084 | 0.023952 | 23.7526 | 0.029576 | 0.15751 | 0.005211 | 0.3722 | 0.3722 | 0.017224
2019-03-29 14:00 | 1553882394.5408 | 407 | 0.20848 | 0.001301 | 4.3e-05 | 100.5318 | 100.1322 | 121.3893 | 21.1292 | 0.002467 | 0.002355 | 250.1285 | 0.26498 | 8.4046 | 0.002312 | 0.003558 | 0.003423 | 5.8e-05 | 0.000994 | 0.32444 | 1.3411 | 2.5475 | 100.2 | 20.92 | 37.166 | 81.0453 | 77.3123 | 78.357 | 0.00919 | 1.0044 | 0.023918 | 23.8134 | 0.035627 | 0.15764 | 0.005147 | 0.36702 | 0.36702 | 0.017481
2019-03-29 14:00 | 1553882395.5408 | 408 | 0.20846 | 0.001296 | 4.3e-05 | 100.4396 | 100.0353 | 121.4238 | 21.1303 | 0.002511 | 0.002337 | 250.1287 | 0.26447 | 8.3981 | 0.002304 | 0.003398 | 0.003325 | 5.8e-05 | 0.000995 | 0.30574 | 1.3649 | 2.5474 | 100.2 | 20.92 | 37.173 | 82.0968 | 77.1334 | 77.9152 | 0.009192 | 1.0056 | 0.023936 | 23.8018 | 0.036083 | 0.15733 | 0.005276 | 0.37654 | 0.37654 | 0.017006
2019-03-29 14:00 | 1553882396.5409 | 409 | 0.20847 | 0.001295 | 4.4e-05 | 100.4214 | 100.1056 | 121.1374 | 21.1291 | 0.002546 | 0.002326 | 250.1001 | 0.26413 | 8.4017 | 0.002328 | 0.003313 | 0.003289 | 5.9e-05 | 0.000997 | 0.3183 | 1.3809 | 2.5475 | 100.2 | 20.92 | 37.173 | 82.9394 | 77.7098 | 77.6591 | 0.009192 | 1.0064 | 0.023917 | 23.765 | 0.028173 | 0.15713 | 0.005188 | 0.36992 | 0.36992 | 0.017288
2019-03-29 14:00 | 1553882397.5409 | 410 | 0.20849 | 0.001296 | 4.4e-05 | 100.3965 | 100.0063 | 120.9952 | 21.1296 | 0.002537 | 0.00232 | 250.0572 | 0.26443 | 8.4008 | 0.002249 | 0.004245 | 0.004148 | 5.9e-05 | 0.000995 | 0.3095 | 1.4169 | 2.5473 | 100.2 | 20.92 | 37.178 | 82.7146 | 75.8099 | 77.4999 | 0.009193 | 1.0068 | 0.023908 | 23.7457 | 0.034838 | 0.15731 | 0.005092 | 0.3629 | 0.3629 | 0.017643
2019-03-29 14:00 | 1553882398.5399 | 411 | 0.20847 | 0.001289 | 4.4e-05 | 100.4181 | 100.0138 | 120.913 | 21.131 | 0.002507 | 0.002322 | 250.0323 | 0.26314 | 8.4031 | 0.00238 | 0.004654 | 0.004576 | 5.8e-05 | 0.000992 | 0.29374 | 1.3662 | 2.5474 | 100.2 | 20.92 | 37.178 | 82.0053 | 78.9457 | 77.5519 | 0.009193 | 1.0067 | 0.023898 | 23.7394 | 0.036086 | 0.15654 | 0.005238 | 0.37317 | 0.37317 | 0.017073
2019-03-29 14:00 | 1553882399.54 | 412 | 0.20846 | 0.001293 | 4.4e-05 | 100.3846 | 100.0445 | 120.9011 | 21.13 | 0.002496 | 0.002311 | 250.0156 | 0.26453 | 8.4042 | 0.00243 | 0.003905 | 0.003807 | 5.9e-05 | 0.000986 | 0.30845 | 1.3864 | 2.5473 | 100.2 | 20.92 | 37.179 | 81.7424 | 80.146 | 77.2991 | 0.009193 | 1.0074 | 0.023905 | 23.7297 | 0.030353 | 0.15737 | 0.005253 | 0.37434 | 0.37434 | 0.01711
2019-03-29 14:00 | 1553882400.5415 | 413 | 0.20847 | 0.001281 | 4.5e-05 | 100.3887 | 100.1615 | 120.8846 | 21.1292 | 0.002511 | 0.00237 | 250.0185 | 0.26374 | 8.4017 | 0.002353 | 0.004111 | 0.004039 | 5.5e-05 | 0.000982 | 0.29839 | 1.3481 | 2.5473 | 100.2 | 20.92 | 37.181 | 82.0878 | 78.3032 | 78.7172 | 0.009194 | 1.0033 | 0.023856 | 23.776 | 0.020267 | 0.1569 | 0.005257 | 0.37386 | 0.37386 | 0.017081
2019-03-29 14:00 | 1553882401.5391 | 414 | 0.20846 | 0.001283 | 4.5e-05 | 100.3615 | 99.965 | 120.599 | 21.1295 | 0.002587 | 0.002327 | 250.0006 | 0.26382 | 8.4041 | 0.002443 | 0.003814 | 0.003769 | 5.7e-05 | 0.000978 | 0.31885 | 1.3484 | 2.5473 | 100.2 | 20.92 | 37.182 | 83.9299 | 80.4522 | 77.6727 | 0.009194 | 1.0063 | 0.023863 | 23.7126 | 0.035411 | 0.15695 | 0.005268 | 0.37467 | 0.37467 | 0.017049
2019-03-29 14:00 | 1553882402.5402 | 415 | 0.20852 | 0.00128 | 4.6e-05 | 100.3567 | 100.0899 | 120.2323 | 21.1292 | 0.002475 | 0.002331 | 249.9717 | 0.26481 | 8.401 | 0.002307 | 0.004071 | 0.003987 | 5.5e-05 | 0.000978 | 0.30541 | 1.3022 | 2.5473 | 100.2 | 20.92 | 37.183 | 81.2424 | 77.2063 | 77.7801 | 0.009194 | 1.006 | 0.023823 | 23.6802 | 0.023815 | 0.15754 | 0.004945 | 0.35096 | 0.35096 | 0.018269
2019-03-29 14:00 | 1553882403.5391 | 416 | 0.20849 | 0.001282 | 4.6e-05 | 100.2804 | 99.8936 | 120.2779 | 21.1296 | 0.002514 | 0.002339 | 249.9404 | 0.26489 | 8.4014 | 0.002342 | 0.00367 | 0.003595 | 5.7e-05 | 0.000975 | 0.32084 | 1.3103 | 2.5472 | 100.2 | 20.92 | 37.184 | 82.1771 | 78.0424 | 77.9705 | 0.009194 | 1.0055 | 0.023821 | 23.6911 | 0.034563 | 0.15758 | 0.005078 | 0.36036 | 0.36036 | 0.017797
2019-03-29 14:00 | 1553882404.5403 | 417 | 0.20851 | 0.001278 | 4.6e-05 | 100.31 | 99.9216 | 120.6269 | 21.13 | 0.00252 | 0.002348 | 249.9186 | 0.26396 | 8.4026 | 0.002376 | 0.003905 | 0.003975 | 5.2e-05 | 0.000973 | 0.28871 | 1.3378 | 2.5475 | 100.2 | 20.92 | 37.185 | 82.3116 | 78.8583 | 78.172 | 0.009195 | 1.0049 | 0.023849 | 23.7322 | 0.034701 | 0.15703 | 0.00497 | 0.35311 | 0.35311 | 0.018099
2019-03-29 14:00 | 1553882405.5402 | 418 | 0.2085 | 0.001274 | 4.6e-05 | 100.3238 | 99.9121 | 120.8921 | 21.1287 | 0.002484 | 0.002254 | 249.8887 | 0.26459 | 8.4021 | 0.002333 | 0.003423 | 0.00328 | 5.1e-05 | 0.000967 | 0.29856 | 1.3248 | 2.5475 | 100.2 | 20.92 | 37.187 | 81.4409 | 77.831 | 75.9118 | 0.009195 | 1.0114 | 0.023952 | 23.6818 | 0.036784 | 0.15741 | 0.005045 | 0.36002 | 0.36002 | 0.017795
2019-03-29 14:00 | 1553882406.5392 | 419 | 0.20849 | 0.001271 | 4.7e-05 | 100.3441 | 100.0383 | 121.1471 | 21.1297 | 0.002482 | 0.002325 | 249.863 | 0.26442 | 8.4026 | 0.002347 | 0.003698 | 0.003618 | 5.4e-05 | 0.00096 | 0.32388 | 1.3114 | 2.5473 | 100.2 | 20.92 | 37.188 | 81.4029 | 78.1501 | 77.6285 | 0.009195 | 1.0065 | 0.023919 | 23.7649 | 0.027301 | 0.1573 | 0.005095 | 0.36308 | 0.36308 | 0.017633
2019-03-29 14:00 | 1553882407.5393 | 420 | 0.20847 | 0.00127 | 4.7e-05 | 100.2265 | 99.8776 | 120.9605 | 21.1305 | 0.002619 | 0.002374 | 249.8419 | 0.26543 | 8.401 | 0.002276 | 0.004354 | 0.004294 | 5.2e-05 | 0.000961 | 0.28998 | 1.335 | 2.5474 | 100.2 | 20.92 | 37.185 | 84.6846 | 76.4421 | 78.8001 | 0.009195 | 1.0031 | 0.02386 | 23.7863 | 0.031193 | 0.15791 | 0.005241 | 0.3726 | 0.3726 | 0.017249
2019-03-29 14:00 | 1553882408.5402 | 421 | 0.2085 | 0.001268 | 4.7e-05 | 100.2526 | 99.8349 | 120.8971 | 21.1304 | 0.002477 | 0.002348 | 249.8452 | 0.26474 | 8.4033 | 0.002363 | 0.004017 | 0.003998 | 5.2e-05 | 0.000959 | 0.28528 | 1.3382 | 2.5473 | 100.2 | 20.92 | 37.184 | 81.2921 | 78.5325 | 78.1755 | 0.009194 | 1.0049 | 0.023875 | 23.7589 | 0.037338 | 0.15749 | 0.005076 | 0.36104 | 0.36104 | 0.017754
2019-03-29 14:00 | 1553882409.5403 | 422 | 0.20851 | 0.001264 | 4.8e-05 | 100.3283 | 99.8493 | 120.8386 | 21.1295 | 0.002483 | 0.002366 | 249.8307 | 0.26459 | 8.4 | 0.002337 | 0.003694 | 0.003638 | 5.6e-05 | 0.000965 | 0.27865 | 1.3729 | 2.5474 | 100.2 | 20.92 | 37.183 | 81.416 | 77.9106 | 78.6074 | 0.009194 | 1.0037 | 0.023855 | 23.7678 | 0.042808 | 0.1574 | 0.004984 | 0.3541 | 0.3541 | 0.018092
2019-03-29 14:00 | 1553882410.5404 | 423 | 0.20851 | 0.001262 | 4.8e-05 | 100.2223 | 99.7325 | 120.6715 | 21.1296 | 0.002534 | 0.002324 | 249.8067 | 0.26499 | 8.4011 | 0.002351 | 0.00348 | 0.003409 | 5e-05 | 0.000966 | 0.31133 | 1.3783 | 2.5474 | 100.2 | 20.92 | 37.181 | 82.6507 | 78.2497 | 77.594 | 0.009194 | 1.0066 | 0.023873 | 23.7171 | 0.043818 | 0.15764 | 0.00502 | 0.35687 | 0.35687 | 0.017979
2019-03-29 14:00 | 1553882411.5394 | 424 | 0.20855 | 0.001262 | 4.9e-05 | 100.2077 | 99.9295 | 120.7823 | 21.1291 | 0.002487 | 0.002331 | 249.789 | 0.26465 | 8.4 | 0.002349 | 0.004247 | 0.00407 | 5e-05 | 0.000961 | 0.28816 | 1.3517 | 2.5474 | 100.2 | 20.92 | 37.18 | 81.5115 | 78.2149 | 77.7612 | 0.009193 | 1.0061 | 0.023878 | 23.7336 | 0.024865 | 0.15744 | 0.004752 | 0.3378 | 0.3378 | 0.018969
2019-03-29 14:00 | 1553882412.5394 | 425 | 0.20855 | 0.001257 | 4.8e-05 | 100.2382 | 99.8513 | 120.7717 | 21.1301 | 0.002532 | 0.00236 | 249.7695 | 0.26461 | 8.4013 | 0.002347 | 0.003822 | 0.003862 | 4.9e-05 | 0.000955 | 0.29927 | 1.3331 | 2.5473 | 100.2 | 20.92 | 37.179 | 82.6059 | 78.1661 | 78.4614 | 0.009193 | 1.0041 | 0.023853 | 23.7563 | 0.034588 | 0.15741 | 0.004786 | 0.33986 | 0.33986 | 0.018851
2019-03-29 14:00 | 1553882413.5394 | 426 | 0.20855 | 0.001259 | 4.9e-05 | 100.172 | 99.7537 | 120.8891 | 21.1296 | 0.002472 | 0.002349 | 249.7602 | 0.26513 | 8.4019 | 0.002279 | 0.004343 | 0.004283 | 5.2e-05 | 0.000954 | 0.286 | 1.3372 | 2.5474 | 100.2 | 20.92 | 37.181 | 81.1524 | 76.5191 | 78.2139 | 0.009194 | 1.0048 | 0.023873 | 23.7594 | 0.037431 | 0.15773 | 0.00475 | 0.33755 | 0.33755 | 0.019018
2019-03-29 14:00 | 1553882414.5395 | 427 | 0.20855 | 0.001252 | 4.9e-05 | 100.1173 | 99.781 | 120.7896 | 21.1292 | 0.002526 | 0.002316 | 249.7389 | 0.26538 | 8.4021 | 0.002291 | 0.004114 | 0.004072 | 5e-05 | 0.000948 | 0.30574 | 1.3296 | 2.5475 | 100.2 | 20.92 | 37.181 | 82.4498 | 76.8229 | 77.4151 | 0.009194 | 1.0071 | 0.02389 | 23.7226 | 0.030091 | 0.15787 | 0.004779 | 0.33981 | 0.33981 | 0.018909
2019-03-29 14:00 | 1553882415.5395 | 428 | 0.20854 | 0.00125 | 4.9e-05 | 100.1247 | 99.6537 | 120.7476 | 21.129 | 0.002474 | 0.002329 | 249.7076 | 0.26487 | 8.4002 | 0.002333 | 0.004294 | 0.004237 | 4.9e-05 | 0.000948 | 0.29407 | 1.3179 | 2.5474 | 100.2 | 20.92 | 37.183 | 81.2064 | 77.8118 | 77.7268 | 0.009194 | 1.0062 | 0.023876 | 23.7291 | 0.042178 | 0.15757 | 0.00484 | 0.34401 | 0.34401 | 0.018643
2019-03-29 14:00 | 1553882416.5391 | 429 | 0.20858 | 0.001249 | 5e-05 | 100.0679 | 99.6868 | 120.4598 | 21.1291 | 0.002497 | 0.002313 | 249.6828 | 0.26428 | 8.401 | 0.002389 | 0.003737 | 0.003657 | 5.1e-05 | 0.000947 | 0.31437 | 1.3076 | 2.5475 | 100.2 | 20.91 | 37.184 | 81.7656 | 79.1621 | 77.3473 | 0.009189 | 1.0073 | 0.02386 | 23.6879 | 0.034125 | 0.15722 | 0.004612 | 0.32745 | 0.32745 | 0.019541
2019-03-29 14:00 | 1553882417.5396 | 430 | 0.2086 | 0.001244 | 5e-05 | 100.1068 | 99.6622 | 120.4624 | 21.1298 | 0.002495 | 0.00235 | 249.6687 | 0.2659 | 8.4025 | 0.002358 | 0.003822 | 0.003757 | 4.8e-05 | 0.000942 | 0.30486 | 1.3238 | 2.5474 | 100.2 | 20.92 | 37.186 | 81.7071 | 78.4308 | 78.2256 | 0.009195 | 1.0048 | 0.023831 | 23.7179 | 0.039808 | 0.15818 | 0.004519 | 0.3203 | 0.3203 | 0.0201
2019-03-29 14:00 | 1553882418.5397 | 431 | 0.20858 | 0.001246 | 5e-05 | 100.1247 | 99.6745 | 120.4912 | 21.1297 | 0.002488 | 0.002309 | 249.6429 | 0.26362 | 8.4031 | 0.002432 | 0.0037 | 0.003564 | 4.9e-05 | 0.000939 | 0.31636 | 1.3063 | 2.5473 | 100.2 | 20.91 | 37.19 | 81.5486 | 80.1911 | 77.2527 | 0.00919 | 1.0075 | 0.023866 | 23.6878 | 0.04031 | 0.15683 | 0.004637 | 0.32923 | 0.32923 | 0.019387
2019-03-29 14:00 | 1553882419.5386 | 432 | 0.20855 | 0.001245 | 5.1e-05 | 100.014 | 99.666 | 120.7795 | 21.1292 | 0.002439 | 0.002365 | 249.6328 | 0.26491 | 8.4029 | 0.002376 | 0.003905 | 0.003866 | 4.5e-05 | 0.000934 | 0.28191 | 1.3605 | 2.5474 | 100.2 | 20.91 | 37.193 | 80.3692 | 78.8637 | 78.5932 | 0.009191 | 1.0037 | 0.023849 | 23.7615 | 0.031176 | 0.1576 | 0.00476 | 0.33776 | 0.33776 | 0.01899
2019-03-29 14:00 | 1553882420.5387 | 433 | 0.20857 | 0.001245 | 5.2e-05 | 100.0475 | 99.625 | 121.0212 | 21.13 | 0.002454 | 0.002298 | 249.6162 | 0.26419 | 8.404 | 0.00241 | 0.003817 | 0.003751 | 4.9e-05 | 0.000941 | 0.33572 | 1.3285 | 2.5474 | 100.2 | 20.91 | 37.196 | 80.738 | 79.6611 | 76.9779 | 0.009192 | 1.0083 | 0.023928 | 23.7306 | 0.03785 | 0.15716 | 0.004658 | 0.33157 | 0.33157 | 0.019292
2019-03-29 14:00 | 1553882421.5389 | 434 | 0.20858 | 0.001235 | 5.2e-05 | 100.0068 | 99.6022 | 121.1121 | 21.1297 | 0.002512 | 0.002314 | 249.5862 | 0.26476 | 8.4045 | 0.002344 | 0.00365 | 0.003478 | 4.6e-05 | 0.000939 | 0.3194 | 1.35 | 2.5474 | 100.2 | 20.91 | 37.198 | 82.1142 | 78.0907 | 77.3649 | 0.009192 | 1.0072 | 0.023924 | 23.7526 | 0.036263 | 0.1575 | 0.004634 | 0.32974 | 0.32974 | 0.019441
2019-03-29 14:00 | 1553882422.5398 | 435 | 0.20857 | 0.001238 | 5.1e-05 | 99.9826 | 99.553 | 121.24 | 21.1287 | 0.002522 | 0.00232 | 249.5703 | 0.26462 | 8.4026 | 0.002409 | 0.003914 | 0.003858 | 4.5e-05 | 0.000938 | 0.31636 | 1.3523 | 2.5473 | 100.2 | 20.91 | 37.201 | 82.3534 | 79.6435 | 77.5098 | 0.009193 | 1.0068 | 0.023932 | 23.77 | 0.038509 | 0.15742 | 0.004679 | 0.33313 | 0.33313 | 0.019233
2019-03-29 14:00 | 1553882423.5398 | 436 | 0.20856 | 0.001231 | 5.2e-05 | 99.9852 | 99.5725 | 120.9938 | 21.1301 | 0.002588 | 0.00238 | 249.5433 | 0.2654 | 8.4033 | 0.002424 | 0.004091 | 0.004038 | 4.9e-05 | 0.000939 | 0.29966 | 1.3529 | 2.5474 | 100.2 | 20.91 | 37.204 | 83.9389 | 80.0106 | 78.9518 | 0.009194 | 1.0027 | 0.023858 | 23.7947 | 0.037002 | 0.15789 | 0.004739 | 0.33633 | 0.33633 | 0.019106
2019-03-29 14:00 | 1553882424.5388 | 437 | 0.20856 | 0.001231 | 5.2e-05 | 100.0203 | 99.553 | 120.8072 | 21.1298 | 0.002503 | 0.002373 | 249.5351 | 0.26467 | 8.4042 | 0.002431 | 0.003768 | 0.003724 | 4.8e-05 | 0.00094 | 0.32189 | 1.3458 | 2.5475 | 100.2 | 20.91 | 37.205 | 81.9083 | 80.1656 | 78.7786 | 0.009194 | 1.0032 | 0.023846 | 23.7705 | 0.041882 | 0.15745 | 0.004749 | 0.33688 | 0.33688 | 0.019023
2019-03-29 14:00 | 1553882425.5388 | 438 | 0.20859 | 0.001232 | 5.2e-05 | 99.9604 | 99.5687 | 120.5842 | 21.1304 | 0.002457 | 0.002322 | 249.5069 | 0.26458 | 8.4002 | 0.002383 | 0.004236 | 0.004087 | 4.9e-05 | 0.000941 | 0.32803 | 1.3901 | 2.5475 | 100.2 | 20.91 | 37.206 | 80.811 | 79.0153 | 77.5614 | 0.009194 | 1.0067 | 0.023865 | 23.7074 | 0.035116 | 0.1574 | 0.004547 | 0.32268 | 0.32268 | 0.019853
2019-03-29 14:00 | 1553882426.5388 | 439 | 0.20856 | 0.001225 | 5.3e-05 | 99.9203 | 99.4991 | 120.7307 | 21.1293 | 0.002498 | 0.002332 | 249.4835 | 0.26604 | 8.4024 | 0.00236 | 0.004185 | 0.004117 | 4.9e-05 | 0.000946 | 0.31581 | 1.3607 | 2.5476 | 100.2 | 20.91 | 37.207 | 81.7799 | 78.4676 | 77.8066 | 0.009194 | 1.006 | 0.023871 | 23.7301 | 0.037783 | 0.15827 | 0.004773 | 0.33894 | 0.33894 | 0.019005
2019-03-29 14:00 | 1553882427.5391 | 440 | 0.20856 | 0.001226 | 5.3e-05 | 99.92 | 99.4977 | 120.9487 | 21.1286 | 0.002485 | 0.002344 | 249.4716 | 0.26424 | 8.4031 | 0.002413 | 0.003622 | 0.003581 | 4.8e-05 | 0.00094 | 0.29463 | 1.3705 | 2.5474 | 100.2 | 20.91 | 37.209 | 81.4797 | 79.7503 | 78.0889 | 0.009195 | 1.0051 | 0.023883 | 23.7611 | 0.037879 | 0.1572 | 0.00475 | 0.33747 | 0.33747 | 0.018959
2019-03-29 14:00 | 1553882428.54 | 441 | 0.20855 | 0.001223 | 5.3e-05 | 99.9326 | 99.4039 | 121.0731 | 21.1283 | 0.002413 | 0.002309 | 249.44 | 0.26492 | 8.403 | 0.002417 | 0.004323 | 0.004267 | 4.9e-05 | 0.000944 | 0.29518 | 1.3473 | 2.5474 | 100.2 | 20.91 | 37.213 | 79.7358 | 79.8451 | 77.2318 | 0.009196 | 1.0076 | 0.023925 | 23.7443 | 0.047446 | 0.1576 | 0.004817 | 0.34282 | 0.34282 | 0.01871
2019-03-29 14:00 | 1553882429.5389 | 442 | 0.20856 | 0.001218 | 5.3e-05 | 99.8146 | 99.3626 | 121.1269 | 21.1292 | 0.00249 | 0.002291 | 249.4049 | 0.26389 | 8.4023 | 0.002431 | 0.003759 | 0.003676 | 4.8e-05 | 0.000942 | 0.32101 | 1.3638 | 2.5475 | 100.2 | 20.91 | 37.217 | 81.5926 | 80.182 | 76.8181 | 0.009197 | 1.0088 | 0.023944 | 23.7355 | 0.0406 | 0.15699 | 0.004746 | 0.338 | 0.338 | 0.018904
2019-03-29 14:00 | 1553882430.5389 | 443 | 0.2086 | 0.001213 | 5.3e-05 | 99.8434 | 99.4851 | 121.0668 | 21.1296 | 0.00247 | 0.002331 | 249.3803 | 0.26478 | 8.4019 | 0.002363 | 0.003864 | 0.003858 | 5.7e-05 | 0.000936 | 0.31581 | 1.3689 | 2.5473 | 100.2 | 20.91 | 37.221 | 81.1159 | 78.5412 | 77.7843 | 0.009198 | 1.006 | 0.023905 | 23.7623 | 0.032152 | 0.15752 | 0.00452 | 0.32127 | 0.32127 | 0.019955
2019-03-29 14:01 | 1553882431.54 | 444 | 0.20855 | 0.001213 | 5.3e-05 | 99.7563 | 99.5698 | 120.7731 | 21.1292 | 0.002537 | 0.002297 | 249.373 | 0.26508 | 8.403 | 0.002339 | 0.003892 | 0.003788 | 5.5e-05 | 0.000932 | 0.30309 | 1.3753 | 2.5474 | 100.2 | 20.91 | 37.227 | 82.7305 | 77.9586 | 76.9615 | 0.009199 | 1.0084 | 0.023904 | 23.7057 | 0.016739 | 0.1577 | 0.004808 | 0.34193 | 0.34193 | 0.018771
2019-03-29 14:01 | 1553882432.539 | 445 | 0.20854 | 0.001211 | 5.3e-05 | 99.8101 | 99.394 | 120.521 | 21.1292 | 0.002494 | 0.002293 | 249.3509 | 0.26462 | 8.4044 | 0.002376 | 0.004568 | 0.004453 | 5.5e-05 | 0.000925 | 0.29645 | 1.3659 | 2.5474 | 100.2 | 20.91 | 37.233 | 81.6808 | 78.8488 | 76.8548 | 0.009201 | 1.0087 | 0.023883 | 23.6773 | 0.037372 | 0.15742 | 0.004866 | 0.34573 | 0.34573 | 0.018532
2019-03-29 14:01 | 1553882433.5381 | 446 | 0.20855 | 0.00121 | 5.3e-05 | 99.7928 | 99.4595 | 120.3112 | 21.1292 | 0.002531 | 0.002318 | 249.3356 | 0.26435 | 8.3999 | 0.002314 | 0.003737 | 0.003732 | 5.1e-05 | 0.000929 | 0.3126 | 1.3462 | 2.5473 | 100.2 | 20.91 | 37.247 | 82.5839 | 77.3716 | 77.455 | 0.009204 | 1.007 | 0.023842 | 23.677 | 0.029921 | 0.15726 | 0.004827 | 0.34237 | 0.34237 | 0.018694
2019-03-29 14:01 | 1553882434.5381 | 447 | 0.20855 | 0.001212 | 5.3e-05 | 99.7904 | 99.4469 | 120.2577 | 21.1282 | 0.002545 | 0.00231 | 249.314 | 0.26401 | 8.4038 | 0.002397 | 0.003965 | 0.003805 | 5.1e-05 | 0.000932 | 0.30469 | 1.3438 | 2.5474 | 100.2 | 20.91 | 37.253 | 82.9059 | 79.3671 | 77.2599 | 0.009206 | 1.0075 | 0.023843 | 23.6651 | 0.030842 | 0.15706 | 0.004862 | 0.34486 | 0.34486 | 0.018536
2019-03-29 14:01 | 1553882435.5382 | 448 | 0.20855 | 0.001207 | 5.4e-05 | 99.7194 | 99.3459 | 120.3394 | 21.1288 | 0.002555 | 0.002296 | 249.301 | 0.26449 | 8.4013 | 0.002394 | 0.003689 | 0.003734 | 5.1e-05 | 0.000931 | 0.30414 | 1.3802 | 2.5474 | 100.2 | 20.91 | 37.27 | 83.1663 | 79.2959 | 76.9393 | 0.00921 | 1.0084 | 0.023862 | 23.6623 | 0.033572 | 0.15735 | 0.004847 | 0.34403 | 0.34403 | 0.018615
2019-03-29 14:01 | 1553882436.5381 | 449 | 0.20855 | 0.0012 | 5.4e-05 | 99.7182 | 99.3346 | 120.4229 | 21.1291 | 0.002436 | 0.00233 | 249.2869 | 0.26486 | 8.4018 | 0.002384 | 0.003283 | 0.003167 | 5.4e-05 | 0.000934 | 0.29695 | 1.3779 | 2.5475 | 100.2 | 20.91 | 37.279 | 80.2864 | 79.0421 | 77.7439 | 0.009212 | 1.0061 | 0.023843 | 23.6977 | 0.034478 | 0.15756 | 0.004844 | 0.34356 | 0.34356 | 0.018666
2019-03-29 14:01 | 1553882437.5382 | 450 | 0.20858 | 0.001196 | 5.4e-05 | 99.7314 | 99.2759 | 120.5682 | 21.129 | 0.002513 | 0.002392 | 249.2655 | 0.26601 | 8.4034 | 0.002333 | 0.003751 | 0.003567 | 5.3e-05 | 0.000943 | 0.31277 | 1.3632 | 2.5474 | 100.2 | 20.9 | 37.288 | 82.1556 | 77.8157 | 79.2408 | 0.009209 | 1.0019 | 0.023807 | 23.7625 | 0.040947 | 0.15825 | 0.004666 | 0.33025 | 0.33025 | 0.019503
2019-03-29 14:01 | 1553882438.5392 | 451 | 0.20858 | 0.0012 | 5.5e-05 | 99.7733 | 99.396 | 120.7231 | 21.128 | 0.002485 | 0.002307 | 249.2444 | 0.26463 | 8.4044 | 0.002349 | 0.003905 | 0.003887 | 4.9e-05 | 0.000944 | 0.30934 | 1.3448 | 2.5474 | 100.2 | 20.9 | 37.308 | 81.4674 | 78.2114 | 77.191 | 0.009214 | 1.0077 | 0.023892 | 23.7085 | 0.033887 | 0.15743 | 0.004636 | 0.32931 | 0.32931 | 0.019457
2019-03-29 14:01 | 1553882439.5392 | 452 | 0.2086 | 0.001194 | 5.5e-05 | 99.7392 | 99.3657 | 120.952 | 21.1295 | 0.002462 | 0.002306 | 249.2364 | 0.26514 | 8.4029 | 0.002379 | 0.004043 | 0.003968 | 5e-05 | 0.000943 | 0.3063 | 1.3308 | 2.5474 | 100.2 | 20.9 | 37.317 | 80.9121 | 78.9332 | 77.1698 | 0.009216 | 1.0078 | 0.023915 | 23.7303 | 0.033562 | 0.15773 | 0.004575 | 0.3252 | 0.3252 | 0.019741
2019-03-29 14:01 | 1553882440.5383 | 453 | 0.20861 | 0.001193 | 5.5e-05 | 99.741 | 99.4052 | 121.0478 | 21.1288 | 0.002431 | 0.00231 | 249.2115 | 0.26535 | 8.4034 | 0.002364 | 0.003591 | 0.003531 | 4.9e-05 | 0.000949 | 0.31421 | 1.3231 | 2.5473 | 100.2 | 20.9 | 37.326 | 80.1668 | 78.5609 | 77.2661 | 0.009218 | 1.0075 | 0.023921 | 23.7429 | 0.03016 | 0.15786 | 0.004461 | 0.31707 | 0.31707 | 0.020263
2019-03-29 14:01 | 1553882441.5384 | 454 | 0.20855 | 0.001196 | 5.6e-05 | 99.7446 | 99.3749 | 121.3473 | 21.1283 | 0.002474 | 0.002278 | 249.1863 | 0.26524 | 8.404 | 0.002397 | 0.003717 | 0.003758 | 5e-05 | 0.00095 | 0.2636 | 1.365 | 2.5473 | 100.2 | 20.9 | 37.341 | 81.2025 | 79.3561 | 76.4875 | 0.009222 | 1.0097 | 0.023977 | 23.7459 | 0.033215 | 0.15779 | 0.004833 | 0.34455 | 0.34455 | 0.018639
2019-03-29 14:01 | 1553882442.5399 | 455 | 0.20854 | 0.0012 | 5.6e-05 | 99.7015 | 99.2827 | 121.7343 | 21.1282 | 0.002445 | 0.002308 | 249.1684 | 0.26498 | 8.4009 | 0.00237 | 0.003968 | 0.003826 | 5.1e-05 | 0.000954 | 0.32029 | 1.3502 | 2.5474 | 100.2 | 20.9 | 37.348 | 80.5132 | 78.7028 | 77.2225 | 0.009224 | 1.0076 | 0.02399 | 23.8087 | 0.037647 | 0.15764 | 0.004916 | 0.35067 | 0.35067 | 0.018296
2019-03-29 14:01 | 1553882443.5398 | 456 | 0.20861 | 0.00119 | 5.7e-05 | 99.6778 | 99.3469 | 121.9212 | 21.1293 | 0.002467 | 0.002302 | 249.1473 | 0.26458 | 8.4017 | 0.002405 | 0.004017 | 0.003893 | 5.5e-05 | 0.000954 | 0.30198 | 1.362 | 2.5474 | 100.2 | 20.9 | 37.359 | 81.05 | 79.552 | 77.0808 | 0.009226 | 1.008 | 0.024014 | 23.8221 | 0.029747 | 0.1574 | 0.004504 | 0.32132 | 0.32132 | 0.019937
2019-03-29 14:01 | 1553882444.5395 | 457 | 0.2086 | 0.001187 | 5.7e-05 | 99.6398 | 99.0401 | 122.0803 | 21.1288 | 0.002494 | 0.0023 | 249.1211 | 0.26486 | 8.4021 | 0.0023 | 0.004325 | 0.004296 | 5.2e-05 | 0.000959 | 0.31996 | 1.3773 | 2.5474 | 100.2 | 20.91 | 37.369 | 81.6866 | 77.0173 | 77.0235 | 0.009234 | 1.0082 | 0.024031 | 23.8357 | 0.054001 | 0.15756 | 0.004586 | 0.32743 | 0.32743 | 0.019585
2019-03-29 14:01 | 1553882445.5385 | 458 | 0.20862 | 0.001192 | 5.7e-05 | 99.7194 | 99.3704 | 122.1939 | 21.1293 | 0.002507 | 0.002344 | 249.0922 | 0.26447 | 8.4016 | 0.002273 | 0.0039 | 0.003899 | 5.1e-05 | 0.000954 | 0.29734 | 1.323 | 2.5473 | 100.2 | 20.91 | 37.369 | 82.0092 | 76.3708 | 78.0764 | 0.009234 | 1.0052 | 0.024006 | 23.8826 | 0.03136 | 0.15733 | 0.004431 | 0.31592 | 0.31592 | 0.020269
2019-03-29 14:01 | 1553882446.5395 | 459 | 0.20858 | 0.001189 | 5.8e-05 | 99.7135 | 99.3489 | 122.1815 | 21.1291 | 0.002478 | 0.002321 | 249.0758 | 0.26455 | 8.4045 | 0.002364 | 0.004412 | 0.004283 | 5.3e-05 | 0.000956 | 0.34147 | 1.3035 | 2.5474 | 100.2 | 20.9 | 37.377 | 81.3062 | 78.5611 | 77.5397 | 0.009231 | 1.0067 | 0.024023 | 23.8632 | 0.032758 | 0.15738 | 0.004669 | 0.33324 | 0.33324 | 0.019222
2019-03-29 14:01 | 1553882447.5395 | 460 | 0.20858 | 0.001188 | 5.8e-05 | 99.6575 | 99.2343 | 122.0338 | 21.1293 | 0.002477 | 0.002372 | 249.0329 | 0.26519 | 8.4008 | 0.002295 | 0.003701 | 0.003663 | 5.2e-05 | 0.000958 | 0.29966 | 1.3092 | 2.5474 | 100.2 | 20.9 | 37.381 | 81.2893 | 76.9085 | 78.7593 | 0.009232 | 1.0032 | 0.023967 | 23.8902 | 0.038066 | 0.15776 | 0.004659 | 0.33175 | 0.33175 | 0.019355
2019-03-29 14:01 | 1553882448.5375 | 461 | 0.20859 | 0.001186 | 5.9e-05 | 99.6799 | 99.1605 | 121.801 | 21.1279 | 0.002485 | 0.00228 | 249.0254 | 0.26549 | 8.403 | 0.002321 | 0.004122 | 0.003995 | 5.2e-05 | 0.000963 | 0.32925 | 1.3138 | 2.5473 | 100.2 | 20.91 | 37.388 | 81.4704 | 77.527 | 76.5526 | 0.009239 | 1.0096 | 0.02402 | 23.7924 | 0.046727 | 0.15794 | 0.004596 | 0.32787 | 0.32787 | 0.019606
2019-03-29 14:01 | 1553882449.5376 | 462 | 0.20857 | 0.001188 | 5.9e-05 | 99.6868 | 99.307 | 121.4036 | 21.1293 | 0.002437 | 0.00223 | 248.9821 | 0.26521 | 8.401 | 0.002308 | 0.004011 | 0.00399 | 5.2e-05 | 0.000968 | 0.34435 | 1.3279 | 2.5474 | 100.2 | 20.9 | 37.391 | 80.3177 | 77.2096 | 75.3505 | 0.009234 | 1.013 | 0.024022 | 23.7127 | 0.034148 | 0.15777 | 0.004719 | 0.33679 | 0.33679 | 0.019066
2019-03-29 14:01 | 1553882450.5386 | 463 | 0.20861 | 0.001181 | 5.9e-05 | 99.6644 | 99.3858 | 121.3658 | 21.1293 | 0.002473 | 0.002306 | 248.9558 | 0.26493 | 8.4036 | 0.002311 | 0.003524 | 0.003405 | 4.7e-05 | 0.000968 | 0.3235 | 1.3523 | 2.5474 | 100.2 | 20.9 | 37.394 | 81.1903 | 77.2828 | 77.1689 | 0.009235 | 1.0078 | 0.023956 | 23.7708 | 0.025036 | 0.15761 | 0.004528 | 0.32214 | 0.32214 | 0.019913
2019-03-29 14:01 | 1553882451.5376 | 464 | 0.2086 | 0.001184 | 6e-05 | 99.6823 | 99.322 | 121.4355 | 21.1286 | 0.002535 | 0.002289 | 248.9267 | 0.26446 | 8.4017 | 0.002253 | 0.004022 | 0.00394 | 5.1e-05 | 0.000968 | 0.31957 | 1.3175 | 2.5473 | 100.2 | 20.91 | 37.401 | 82.6768 | 75.9071 | 76.7581 | 0.009242 | 1.009 | 0.023977 | 23.7637 | 0.032393 | 0.15733 | 0.00459 | 0.3268 | 0.3268 | 0.019594
2019-03-29 14:01 | 1553882452.5386 | 465 | 0.2086 | 0.001184 | 6e-05 | 99.6874 | 99.0592 | 121.3069 | 21.1277 | 0.002541 | 0.002266 | 248.9021 | 0.26475 | 8.4012 | 0.002313 | 0.003885 | 0.003823 | 4.9e-05 | 0.000964 | 0.29286 | 1.3579 | 2.5473 | 100.2 | 20.91 | 37.404 | 82.8282 | 77.3342 | 76.2047 | 0.009243 | 1.0106 | 0.023983 | 23.7323 | 0.056547 | 0.1575 | 0.004554 | 0.32432 | 0.32432 | 0.019765
2019-03-29 14:01 | 1553882453.5379 | 466 | 0.20861 | 0.001181 | 6e-05 | 99.6736 | 99.1445 | 121.6374 | 21.1279 | 0.002468 | 0.002274 | 248.8795 | 0.26489 | 8.4018 | 0.002279 | 0.004188 | 0.004122 | 4.8e-05 | 0.000962 | 0.32101 | 1.3631 | 2.5474 | 100.2 | 20.91 | 37.411 | 81.0626 | 76.5151 | 76.3959 | 0.009245 | 1.01 | 0.024009 | 23.7711 | 0.04761 | 0.15759 | 0.00453 | 0.32291 | 0.32291 | 0.019862
2019-03-29 14:01 | 1553882454.5384 | 467 | 0.2086 | 0.001177 | 6e-05 | 99.5982 | 99.267 | 121.8017 | 21.1289 | 0.002508 | 0.002312 | 248.8557 | 0.26455 | 8.4022 | 0.002294 | 0.003603 | 0.003555 | 5e-05 | 0.000957 | 0.30757 | 1.3794 | 2.5474 | 100.2 | 20.91 | 37.413 | 82.0173 | 76.8778 | 77.3137 | 0.009245 | 1.0074 | 0.023994 | 23.8184 | 0.029792 | 0.15738 | 0.004595 | 0.32738 | 0.32738 | 0.019566
2019-03-29 14:01 | 1553882455.5378 | 468 | 0.20861 | 0.00118 | 6.1e-05 | 99.629 | 99.252 | 122.0073 | 21.1278 | 0.002504 | 0.002332 | 248.8372 | 0.26446 | 8.4005 | 0.002215 | 0.003678 | 0.003572 | 4.8e-05 | 0.000949 | 0.30757 | 1.3672 | 2.5474 | 100.2 | 20.91 | 37.416 | 81.932 | 74.974 | 77.7965 | 0.009246 | 1.006 | 0.023998 | 23.8549 | 0.033914 | 0.15733 | 0.004533 | 0.32293 | 0.32293 | 0.019829
2019-03-29 14:01 | 1553882456.5378 | 469 | 0.2086 | 0.001177 | 6e-05 | 99.6171 | 99.2483 | 122.2136 | 21.1276 | 0.002488 | 0.002244 | 248.8275 | 0.26533 | 8.4024 | 0.002309 | 0.003448 | 0.003412 | 4.7e-05 | 0.00095 | 0.29374 | 1.3555 | 2.5475 | 100.2 | 20.91 | 37.421 | 81.5427 | 77.2531 | 75.6895 | 0.009247 | 1.0121 | 0.02409 | 23.8033 | 0.033177 | 0.15785 | 0.004571 | 0.32696 | 0.32696 | 0.019648
2019-03-29 14:01 | 1553882457.5384 | 470 | 0.20861 | 0.001177 | 6.1e-05 | 99.6871 | 99.2015 | 122.1787 | 21.1286 | 0.002504 | 0.002347 | 248.8346 | 0.26553 | 8.4009 | 0.002234 | 0.003586 | 0.003584 | 4.9e-05 | 0.000946 | 0.32068 | 1.3667 | 2.5475 | 100.2 | 20.91 | 37.425 | 81.9346 | 75.4402 | 78.1474 | 0.009248 | 1.005 | 0.024002 | 23.8836 | 0.043679 | 0.15796 | 0.004525 | 0.32244 | 0.32244 | 0.019939
2019-03-29 14:01 | 1553882458.5389 | 471 | 0.2086 | 0.001172 | 6.1e-05 | 99.6317 | 99.2503 | 122.273 | 21.128 | 0.002493 | 0.002324 | 248.8163 | 0.26459 | 8.4013 | 0.002238 | 0.004126 | 0.003992 | 4.9e-05 | 0.00094 | 0.29015 | 1.3374 | 2.5474 | 100.2 | 20.91 | 37.425 | 81.6734 | 75.5448 | 77.6074 | 0.009248 | 1.0065 | 0.02403 | 23.8744 | 0.034309 | 0.1574 | 0.004573 | 0.32625 | 0.32625 | 0.019636
2019-03-29 14:01 | 1553882459.5379 | 472 | 0.2086 | 0.001174 | 6.1e-05 | 99.7212 | 99.2353 | 122.4061 | 21.1274 | 0.002452 | 0.002314 | 248.7978 | 0.26481 | 8.4041 | 0.002305 | 0.004217 | 0.004159 | 4.8e-05 | 0.00094 | 0.31404 | 1.3353 | 2.5473 | 100.2 | 20.91 | 37.428 | 80.6896 | 77.1453 | 77.3596 | 0.009249 | 1.0072 | 0.024052 | 23.879 | 0.043694 | 0.15754 | 0.004602 | 0.32866 | 0.32866 | 0.019509
2019-03-29 14:01 | 1553882460.5394 | 473 | 0.20861 | 0.001174 | 6.1e-05 | 99.6611 | 99.2162 | 122.2371 | 21.1281 | 0.002405 | 0.002291 | 248.7698 | 0.26458 | 8.4034 | 0.002294 | 0.003535 | 0.003458 | 5e-05 | 0.00094 | 0.29607 | 1.3345 | 2.5474 | 100.2 | 20.91 | 37.428 | 79.5408 | 76.8848 | 76.817 | 0.009249 | 1.0088 | 0.024054 | 23.844 | 0.040019 | 0.1574 | 0.004512 | 0.32218 | 0.32218 | 0.019884
2019-03-29 14:01 | 1553882461.5389 | 474 | 0.20861 | 0.001173 | 6.1e-05 | 99.6763 | 99.1732 | 122.0679 | 21.1287 | 0.002411 | 0.002305 | 248.7493 | 0.26485 | 8.4035 | 0.002334 | 0.003922 | 0.003874 | 4.5e-05 | 0.000936 | 0.31526 | 1.3591 | 2.5473 | 100.2 | 20.91 | 37.431 | 79.6841 | 77.8341 | 77.1489 | 0.00925 | 1.0078 | 0.024026 | 23.8388 | 0.045266 | 0.15756 | 0.004545 | 0.32415 | 0.32415 | 0.019783
2019-03-29 14:01 | 1553882462.538 | 475 | 0.2086 | 0.001171 | 6.1e-05 | 99.6503 | 99.1756 | 121.9695 | 21.129 | 0.002506 | 0.002268 | 248.7388 | 0.26459 | 8.4013 | 0.002242 | 0.004043 | 0.003934 | 4.8e-05 | 0.000936 | 0.29463 | 1.3505 | 2.5475 | 100.2 | 20.91 | 37.431 | 81.9708 | 75.6299 | 76.2521 | 0.00925 | 1.0104 | 0.024047 | 23.7986 | 0.042714 | 0.1574 | 0.004556 | 0.32526 | 0.32526 | 0.019696
2019-03-29 14:01 | 1553882463.5382 | 476 | 0.20859 | 0.001172 | 6.2e-05 | 99.6817 | 99.0691 | 121.5259 | 21.1275 | 0.002474 | 0.002285 | 248.7187 | 0.26504 | 8.4016 | 0.00227 | 0.004202 | 0.004121 | 4.7e-05 | 0.00094 | 0.32156 | 1.3899 | 2.5475 | 100.2 | 20.91 | 37.432 | 81.216 | 76.2941 | 76.673 | 0.00925 | 1.0092 | 0.023989 | 23.7696 | 0.055143 | 0.15767 | 0.004639 | 0.33039 | 0.33039 | 0.019423
2019-03-29 14:01 | 1553882464.5381 | 477 | 0.20861 | 0.001166 | 6.2e-05 | 99.6347 | 99.1872 | 121.1546 | 21.1274 | 0.002486 | 0.002316 | 248.6938 | 0.26511 | 8.4027 | 0.002296 | 0.003589 | 0.003575 | 4.5e-05 | 0.000934 | 0.30143 | 1.3607 | 2.5475 | 100.2 | 20.91 | 37.432 | 81.4966 | 76.9209 | 77.4127 | 0.00925 | 1.0071 | 0.023927 | 23.7584 | 0.04027 | 0.15771 | 0.00452 | 0.32096 | 0.32096 | 0.019999
2019-03-29 14:01 | 1553882465.5381 | 478 | 0.2086 | 0.001167 | 6.3e-05 | 99.6689 | 99.1394 | 120.945 | 21.1277 | 0.002449 | 0.00232 | 248.6669 | 0.2651 | 8.4022 | 0.002239 | 0.003809 | 0.003679 | 4.6e-05 | 0.000936 | 0.29286 | 1.3561 | 2.5474 | 100.2 | 20.91 | 37.432 | 80.5985 | 75.5719 | 77.4949 | 0.00925 | 1.0068 | 0.023903 | 23.7406 | 0.047643 | 0.15771 | 0.004609 | 0.327 | 0.327 | 0.019628
2019-03-29 14:01 | 1553882466.5371 | 479 | 0.2086 | 0.001165 | 6.3e-05 | 99.6605 | 99.0496 | 120.6663 | 21.1276 | 0.002432 | 0.002277 | 248.636 | 0.26461 | 8.4019 | 0.00222 | 0.00353 | 0.003461 | 4.9e-05 | 0.00093 | 0.31868 | 1.3205 | 2.5474 | 100.2 | 20.9 | 37.431 | 80.186 | 75.1024 | 76.4658 | 0.009244 | 1.0098 | 0.023911 | 23.6784 | 0.054992 | 0.15742 | 0.00459 | 0.32573 | 0.32573 | 0.019669
2019-03-29 14:01 | 1553882467.5384 | 480 | 0.20864 | 0.001172 | 6.3e-05 | 99.7215 | 99.1237 | 120.7173 | 21.1271 | 0.002526 | 0.002329 | 248.6253 | 0.26525 | 8.4025 | 0.002256 | 0.003461 | 0.003369 | 4.9e-05 | 0.000932 | 0.31404 | 1.3648 | 2.5474 | 100.2 | 20.9 | 37.429 | 82.4555 | 75.9766 | 77.7219 | 0.009244 | 1.0062 | 0.023873 | 23.7259 | 0.053786 | 0.1578 | 0.00436 | 0.30874 | 0.30874 | 0.020802
2019-03-29 14:01 | 1553882468.5382 | 481 | 0.20861 | 0.00117 | 6.3e-05 | 99.6835 | 99.2305 | 120.6293 | 21.1283 | 0.0025 | 0.002311 | 248.5912 | 0.26491 | 8.4011 | 0.002267 | 0.004027 | 0.003916 | 4.4e-05 | 0.000928 | 0.30917 | 1.3536 | 2.5475 | 100.2 | 20.9 | 37.428 | 81.8405 | 76.2372 | 77.2825 | 0.009243 | 1.0075 | 0.023879 | 23.7024 | 0.040741 | 0.1576 | 0.004505 | 0.3192 | 0.3192 | 0.020095
2019-03-29 14:01 | 1553882469.5383 | 482 | 0.2086 | 0.001168 | 6.3e-05 | 99.7374 | 99.2824 | 120.4276 | 21.1278 | 0.00244 | 0.002307 | 248.5752 | 0.26468 | 8.4027 | 0.002262 | 0.003578 | 0.003621 | 4.5e-05 | 0.000931 | 0.2943 | 1.356 | 2.5475 | 100.2 | 20.9 | 37.425 | 80.395 | 76.112 | 77.1862 | 0.009243 | 1.0077 | 0.023862 | 23.6793 | 0.0409 | 0.15746 | 0.004588 | 0.32491 | 0.32491 | 0.019724
2019-03-29 14:01 | 1553882470.5373 | 483 | 0.20862 | 0.001169 | 6.3e-05 | 99.6464 | 99.1001 | 120.3884 | 21.1283 | 0.002501 | 0.002295 | 248.5658 | 0.26554 | 8.404 | 0.002277 | 0.004595 | 0.004486 | 4.4e-05 | 0.000926 | 0.28943 | 1.3476 | 2.5474 | 100.2 | 20.9 | 37.422 | 81.8679 | 76.4782 | 76.9054 | 0.009242 | 1.0085 | 0.023868 | 23.666 | 0.049169 | 0.15797 | 0.00448 | 0.31723 | 0.31723 | 0.020267
2019-03-29 14:01 | 1553882471.5373 | 484 | 0.20863 | 0.001164 | 6.4e-05 | 99.7395 | 99.366 | 120.3478 | 21.1273 | 0.002498 | 0.0023 | 248.5565 | 0.26471 | 8.4015 | 0.002249 | 0.003857 | 0.00384 | 4.5e-05 | 0.000926 | 0.32189 | 1.321 | 2.5473 | 100.2 | 20.9 | 37.419 | 81.7856 | 75.7944 | 77.0274 | 0.009241 | 1.0082 | 0.02386 | 23.6661 | 0.033557 | 0.15747 | 0.004426 | 0.31325 | 0.31325 | 0.02046
2019-03-29 14:01 | 1553882472.5383 | 485 | 0.20863 | 0.00117 | 6.4e-05 | 99.7078 | 99.28 | 120.4145 | 21.1272 | 0.002542 | 0.002333 | 248.5438 | 0.2643 | 8.4042 | 0.002277 | 0.003604 | 0.003501 | 4.5e-05 | 0.000927 | 0.29302 | 1.3324 | 2.5475 | 100.2 | 20.9 | 37.416 | 82.8339 | 76.4834 | 77.8178 | 0.00924 | 1.0059 | 0.02384 | 23.6994 | 0.038457 | 0.15723 | 0.004411 | 0.31193 | 0.31193 | 0.020515
2019-03-29 14:01 | 1553882473.5384 | 486 | 0.20861 | 0.001164 | 6.4e-05 | 99.7081 | 99.1674 | 120.7224 | 21.128 | 0.002527 | 0.002277 | 248.5175 | 0.26469 | 8.4014 | 0.002298 | 0.003608 | 0.003557 | 4.8e-05 | 0.00093 | 0.32029 | 1.2987 | 2.5473 | 100.2 | 20.9 | 37.407 | 82.482 | 76.968 | 76.4779 | 0.009238 | 1.0098 | 0.023916 | 23.6843 | 0.048636 | 0.15746 | 0.004507 | 0.31976 | 0.31976 | 0.020042
2019-03-29 14:01 | 1553882474.5374 | 487 | 0.20859 | 0.001162 | 6.5e-05 | 99.7048 | 99.0547 | 120.7518 | 21.1285 | 0.002446 | 0.00228 | 248.511 | 0.26433 | 8.4031 | 0.002239 | 0.00376 | 0.003682 | 4.5e-05 | 0.00093 | 0.30237 | 1.366 | 2.5474 | 100.2 | 20.9 | 37.407 | 80.5236 | 75.5585 | 76.5538 | 0.009238 | 1.0096 | 0.023916 | 23.6898 | 0.058505 | 0.15725 | 0.004656 | 0.33044 | 0.33044 | 0.019368
2019-03-29 14:01 | 1553882475.5384 | 488 | 0.20865 | 0.001156 | 6.5e-05 | 99.6698 | 99.2291 | 120.9323 | 21.1284 | 0.002437 | 0.002321 | 248.4951 | 0.26463 | 8.4021 | 0.002259 | 0.003589 | 0.003458 | 4.4e-05 | 0.000933 | 0.32493 | 1.3938 | 2.5474 | 100.2 | 20.9 | 37.405 | 80.3223 | 76.0519 | 77.529 | 0.009238 | 1.0067 | 0.023901 | 23.7405 | 0.039629 | 0.15743 | 0.004288 | 0.3038 | 0.3038 | 0.021091
2019-03-29 14:01 | 1553882476.5384 | 489 | 0.20862 | 0.001156 | 6.5e-05 | 99.7658 | 99.3148 | 120.9626 | 21.128 | 0.002472 | 0.002312 | 248.4805 | 0.26504 | 8.4027 | 0.002262 | 0.004086 | 0.004019 | 4.6e-05 | 0.000941 | 0.31725 | 1.3556 | 2.5473 | 100.2 | 20.9 | 37.401 | 81.1494 | 76.1067 | 77.3229 | 0.009237 | 1.0073 | 0.023911 | 23.7365 | 0.040529 | 0.15767 | 0.004472 | 0.31716 | 0.31716 | 0.020233
2019-03-29 14:01 | 1553882477.5375 | 490 | 0.20863 | 0.001153 | 6.5e-05 | 99.7051 | 99.1923 | 121.0097 | 21.1274 | 0.002477 | 0.002237 | 248.4518 | 0.26422 | 8.3995 | 0.002267 | 0.003301 | 0.003276 | 4.6e-05 | 0.000937 | 0.31133 | 1.3235 | 2.5473 | 100.2 | 20.9 | 37.397 | 81.2867 | 76.227 | 75.5185 | 0.009236 | 1.0126 | 0.023977 | 23.6799 | 0.046121 | 0.15719 | 0.004449 | 0.3164 | 0.3164 | 0.02022
2019-03-29 14:01 | 1553882478.5375 | 491 | 0.20862 | 0.001154 | 6.5e-05 | 99.6853 | 99.0421 | 121.2736 | 21.128 | 0.002455 | 0.002298 | 248.4327 | 0.26459 | 8.4017 | 0.002213 | 0.004511 | 0.004446 | 4.6e-05 | 0.000935 | 0.29822 | 1.3458 | 2.5474 | 100.2 | 20.89 | 37.394 | 80.761 | 74.9422 | 76.9836 | 0.009229 | 1.0083 | 0.023953 | 23.7555 | 0.0579 | 0.1574 | 0.004498 | 0.31955 | 0.31955 | 0.020048
2019-03-29 14:01 | 1553882479.5376 | 492 | 0.20863 | 0.001152 | 6.4e-05 | 99.7437 | 99.2786 | 121.5789 | 21.1286 | 0.002449 | 0.002318 | 248.4019 | 0.26403 | 8.4023 | 0.002284 | 0.003163 | 0.003045 | 4.8e-05 | 0.000939 | 0.29214 | 1.3234 | 2.5474 | 100.2 | 20.89 | 37.391 | 80.608 | 76.6546 | 77.4537 | 0.009228 | 1.007 | 0.023967 | 23.8013 | 0.041801 | 0.15707 | 0.004436 | 0.31532 | 0.31532 | 0.020274
2019-03-29 14:01 | 1553882480.5375 | 493 | 0.20864 | 0.001151 | 6.5e-05 | 99.6814 | 99.1251 | 121.959 | 21.1278 | 0.002481 | 0.002306 | 248.3778 | 0.26344 | 8.4016 | 0.002251 | 0.004408 | 0.004332 | 4.8e-05 | 0.000944 | 0.30055 | 1.3317 | 2.5472 | 100.2 | 20.89 | 37.389 | 81.3737 | 75.8446 | 77.1651 | 0.009228 | 1.0078 | 0.024014 | 23.8287 | 0.050064 | 0.15672 | 0.004377 | 0.3117 | 0.3117 | 0.020464
2019-03-29 14:01 | 1553882481.5366 | 494 | 0.20862 | 0.001151 | 6.5e-05 | 99.6853 | 99.1196 | 121.8414 | 21.1281 | 0.002453 | 0.00237 | 248.3597 | 0.26454 | 8.4033 | 0.002266 | 0.003746 | 0.003699 | 4.9e-05 | 0.000941 | 0.31078 | 1.3382 | 2.5473 | 100.2 | 20.89 | 37.388 | 80.6948 | 76.2027 | 78.7051 | 0.009228 | 1.0034 | 0.02395 | 23.8695 | 0.050906 | 0.15738 | 0.004493 | 0.31922 | 0.31922 | 0.020065
2019-03-29 14:01 | 1553882482.5366 | 495 | 0.20863 | 0.001152 | 6.4e-05 | 99.723 | 99.0626 | 121.8482 | 21.1274 | 0.002513 | 0.002244 | 248.3428 | 0.26419 | 8.4028 | 0.0023 | 0.003947 | 0.003864 | 5.1e-05 | 0.000939 | 0.31222 | 1.3464 | 2.5475 | 100.2 | 20.89 | 37.384 | 82.1353 | 77.0237 | 75.6852 | 0.009227 | 1.0121 | 0.024054 | 23.7675 | 0.059434 | 0.15717 | 0.004441 | 0.31687 | 0.31687 | 0.020187
2019-03-29 14:01 | 1553882483.5378 | 496 | 0.20862 | 0.001152 | 6.5e-05 | 99.6626 | 99.0855 | 121.544 | 21.1283 | 0.00242 | 0.002367 | 248.3293 | 0.26487 | 8.4001 | 0.002282 | 0.003084 | 0.002974 | 5e-05 | 0.00094 | 0.30829 | 1.3799 | 2.5474 | 100.2 | 20.89 | 37.382 | 79.9144 | 76.5856 | 78.625 | 0.009226 | 1.0036 | 0.023924 | 23.8376 | 0.051945 | 0.15757 | 0.004482 | 0.31804 | 0.31804 | 0.020164
2019-03-29 14:01 | 1553882484.5367 | 497 | 0.20863 | 0.001151 | 6.5e-05 | 99.741 | 99.2199 | 121.3834 | 21.1279 | 0.002509 | 0.002321 | 248.3092 | 0.26469 | 8.4013 | 0.002205 | 0.003533 | 0.00349 | 5.2e-05 | 0.000938 | 0.30309 | 1.3774 | 2.5475 | 100.2 | 20.89 | 37.38 | 82.0413 | 74.7442 | 77.5426 | 0.009226 | 1.0067 | 0.023945 | 23.7852 | 0.04685 | 0.15746 | 0.004426 | 0.31431 | 0.31431 | 0.02039
2019-03-29 14:01 | 1553882485.5367 | 498 | 0.20861 | 0.001151 | 6.5e-05 | 99.6623 | 99.2155 | 121.2548 | 21.1266 | 0.00253 | 0.00228 | 248.3059 | 0.26419 | 8.4006 | 0.002246 | 0.003602 | 0.003452 | 5.1e-05 | 0.000945 | 0.32084 | 1.3728 | 2.5475 | 100.2 | 20.89 | 37.379 | 82.5526 | 75.7337 | 76.545 | 0.009226 | 1.0096 | 0.023966 | 23.7388 | 0.040188 | 0.15716 | 0.004554 | 0.32377 | 0.32377 | 0.019757
2019-03-29 14:01 | 1553882486.5367 | 499 | 0.20862 | 0.001145 | 6.5e-05 | 99.7625 | 99.2517 | 121.2064 | 21.1271 | 0.002498 | 0.002285 | 248.2816 | 0.2638 | 8.4034 | 0.002352 | 0.004457 | 0.004441 | 4.5e-05 | 0.000947 | 0.3115 | 1.3683 | 2.5474 | 100.2 | 20.88 | 37.38 | 81.791 | 78.2667 | 76.6712 | 0.00922 | 1.0092 | 0.023957 | 23.7383 | 0.045922 | 0.15693 | 0.004473 | 0.31782 | 0.31782 | 0.020097
2019-03-29 14:01 | 1553882487.5372 | 500 | 0.20862 | 0.001151 | 6.6e-05 | 99.729 | 99.2721 | 121.2194 | 21.1274 | 0.002493 | 0.00228 | 248.2502 | 0.26437 | 8.4021 | 0.002274 | 0.004417 | 0.004301 | 4.8e-05 | 0.000949 | 0.29374 | 1.3438 | 2.5475 | 100.2 | 20.88 | 37.377 | 81.6596 | 76.394 | 76.5528 | 0.009219 | 1.0096 | 0.023962 | 23.7356 | 0.04107 | 0.15728 | 0.004458 | 0.3168 | 0.3168 | 0.020205
2019-03-29 14:01 | 1553882488.5368 | 501 | 0.20863 | 0.001149 | 6.6e-05 | 99.7314 | 99.2626 | 120.9384 | 21.1285 | 0.002513 | 0.002314 | 248.2206 | 0.26432 | 8.4036 | 0.002323 | 0.003629 | 0.003533 | 4.6e-05 | 0.000951 | 0.31758 | 1.2997 | 2.5474 | 100.2 | 20.88 | 37.375 | 82.1402 | 77.5859 | 77.3567 | 0.009219 | 1.0072 | 0.023907 | 23.7353 | 0.042146 | 0.15724 | 0.004398 | 0.31174 | 0.31174 | 0.020529
2019-03-29 14:01 | 1553882489.5369 | 502 | 0.20864 | 0.001152 | 6.6e-05 | 99.8293 | 99.282 | 120.6654 | 21.1279 | 0.00245 | 0.002303 | 248.1989 | 0.26387 | 8.4028 | 0.002256 | 0.003826 | 0.003818 | 5e-05 | 0.000944 | 0.31133 | 1.3556 | 2.5474 | 100.2 | 20.88 | 37.374 | 80.6384 | 75.9798 | 77.0888 | 0.009219 | 1.008 | 0.023889 | 23.6994 | 0.049168 | 0.15698 | 0.004341 | 0.30736 | 0.30736 | 0.020787
2019-03-29 14:01 | 1553882490.5369 | 503 | 0.20869 | 0.001151 | 6.6e-05 | 99.7835 | 99.2264 | 120.3124 | 21.1273 | 0.002461 | 0.002294 | 248.1853 | 0.26528 | 8.3997 | 0.002211 | 0.003735 | 0.003635 | 4.9e-05 | 0.000939 | 0.3131 | 1.3143 | 2.5474 | 100.2 | 20.88 | 37.37 | 80.8983 | 74.8808 | 76.8756 | 0.009218 | 1.0086 | 0.023862 | 23.6575 | 0.050077 | 0.15782 | 0.004088 | 0.28891 | 0.28891 | 0.022232
2019-03-29 14:02 | 1553882491.5369 | 504 | 0.20861 | 0.001146 | 6.7e-05 | 99.8505 | 99.3012 | 120.7926 | 21.1274 | 0.002467 | 0.002285 | 248.1592 | 0.26427 | 8.4003 | 0.002216 | 0.003971 | 0.003905 | 4.6e-05 | 0.000937 | 0.32068 | 1.3679 | 2.5474 | 100.2 | 20.88 | 37.369 | 81.031 | 75.0173 | 76.6589 | 0.009217 | 1.0093 | 0.023917 | 23.6973 | 0.04935 | 0.15721 | 0.004525 | 0.32091 | 0.32091 | 0.019939
2019-03-29 14:02 | 1553882492.537 | 505 | 0.20862 | 0.001146 | 6.7e-05 | 99.7715 | 99.2715 | 121.014 | 21.127 | 0.002487 | 0.002307 | 248.1185 | 0.26476 | 8.4013 | 0.002302 | 0.003077 | 0.003011 | 4.6e-05 | 0.000936 | 0.31509 | 1.3878 | 2.5476 | 100.2 | 20.88 | 37.364 | 81.5258 | 77.0723 | 77.1987 | 0.009216 | 1.0077 | 0.02392 | 23.7373 | 0.044943 | 0.1575 | 0.004479 | 0.31766 | 0.31766 | 0.02018
2019-03-29 14:02 | 1553882493.537 | 506 | 0.20864 | 0.001146 | 6.8e-05 | 99.7952 | 99.2483 | 121.1515 | 21.1274 | 0.002431 | 0.002264 | 248.1017 | 0.26427 | 8.4031 | 0.002319 | 0.004503 | 0.004452 | 4.7e-05 | 0.000931 | 0.28152 | 1.368 | 2.5475 | 100.2 | 20.88 | 37.363 | 80.1632 | 77.4766 | 76.1561 | 0.009216 | 1.0107 | 0.023969 | 23.7155 | 0.049154 | 0.15721 | 0.004383 | 0.3114 | 0.3114 | 0.020548
2019-03-29 14:02 | 1553882494.5371 | 507 | 0.20861 | 0.001147 | 6.8e-05 | 99.8535 | 99.479 | 121.4498 | 21.1275 | 0.00241 | 0.00228 | 248.0598 | 0.26499 | 8.3989 | 0.002254 | 0.003942 | 0.003821 | 4.7e-05 | 0.000929 | 0.33411 | 1.3527 | 2.5474 | 100.2 | 20.87 | 37.364 | 79.6579 | 75.9318 | 76.5396 | 0.00921 | 1.0096 | 0.023986 | 23.7577 | 0.033616 | 0.15764 | 0.004566 | 0.3247 | 0.3247 | 0.01976
2019-03-29 14:02 | 1553882495.536 | 508 | 0.20864 | 0.001148 | 6.8e-05 | 99.8931 | 99.2329 | 121.402 | 21.1268 | 0.002463 | 0.002329 | 248.0374 | 0.26426 | 8.4034 | 0.002207 | 0.004071 | 0.00407 | 4.7e-05 | 0.000933 | 0.29551 | 1.3643 | 2.5474 | 100.2 | 20.87 | 37.36 | 80.954 | 74.7952 | 77.7201 | 0.009209 | 1.0062 | 0.023941 | 23.793 | 0.059307 | 0.15721 | 0.004335 | 0.30751 | 0.30751 | 0.020807
2019-03-29 14:02 | 1553882496.5361 | 509 | 0.20861 | 0.001144 | 6.9e-05 | 99.9104 | 99.1728 | 121.2595 | 21.1274 | 0.002445 | 0.002298 | 248.0226 | 0.26458 | 8.4007 | 0.002305 | 0.003393 | 0.003291 | 4.6e-05 | 0.00093 | 0.3402 | 1.3486 | 2.5474 | 100.2 | 20.87 | 37.36 | 80.5179 | 77.1421 | 76.9893 | 0.009209 | 1.0083 | 0.023951 | 23.7543 | 0.066277 | 0.1574 | 0.004521 | 0.32101 | 0.32101 | 0.019956
2019-03-29 14:02 | 1553882497.5361 | 510 | 0.20863 | 0.001143 | 6.8e-05 | 99.8781 | 99.2547 | 121.4062 | 21.1271 | 0.002474 | 0.002352 | 247.9987 | 0.26433 | 8.4018 | 0.002258 | 0.003962 | 0.003953 | 4.8e-05 | 0.00093 | 0.32189 | 1.3739 | 2.5474 | 100.2 | 20.87 | 37.357 | 81.2108 | 76.0225 | 78.2838 | 0.009209 | 1.0046 | 0.023922 | 23.8126 | 0.055998 | 0.15725 | 0.004417 | 0.31313 | 0.31313 | 0.020439
2019-03-29 14:02 | 1553882498.5372 | 511 | 0.20863 | 0.001147 | 6.8e-05 | 99.9107 | 99.3462 | 121.3557 | 21.1264 | 0.002469 | 0.002339 | 247.9825 | 0.26416 | 8.403 | 0.002368 | 0.004025 | 0.003843 | 4.9e-05 | 0.000928 | 0.30182 | 1.3959 | 2.5476 | 100.2 | 20.87 | 37.356 | 81.0791 | 78.6691 | 77.9629 | 0.009208 | 1.0055 | 0.023928 | 23.7967 | 0.050681 | 0.15715 | 0.004452 | 0.31574 | 0.31574 | 0.020257
2019-03-29 14:02 | 1553882499.5372 | 512 | 0.20863 | 0.001143 | 6.9e-05 | 99.9122 | 99.3312 | 121.6564 | 21.1269 | 0.002481 | 0.00226 | 247.9755 | 0.26337 | 8.404 | 0.002361 | 0.003985 | 0.003914 | 5.1e-05 | 0.000926 | 0.31797 | 1.3446 | 2.5474 | 100.2 | 20.87 | 37.354 | 81.3791 | 78.4951 | 76.0542 | 0.009208 | 1.011 | 0.024023 | 23.7614 | 0.052167 | 0.15668 | 0.004434 | 0.31562 | 0.31562 | 0.020204
2019-03-29 14:02 | 1553882500.5372 | 513 | 0.20863 | 0.001143 | 6.9e-05 | 99.9745 | 99.3093 | 121.7355 | 21.1264 | 0.002483 | 0.002298 | 247.9513 | 0.2655 | 8.4038 | 0.002317 | 0.003744 | 0.003823 | 5.2e-05 | 0.000931 | 0.31078 | 1.3285 | 2.5474 | 100.2 | 20.86 | 37.352 | 81.4151 | 77.4398 | 76.9713 | 0.009202 | 1.0084 | 0.023999 | 23.8003 | 0.059706 | 0.15795 | 0.004451 | 0.31657 | 0.31657 | 0.020307
2019-03-29 14:02 | 1553882501.5378 | 514 | 0.20864 | 0.001145 | 6.9e-05 | 100.029 | 99.55 | 121.5048 | 21.1273 | 0.002488 | 0.002329 | 247.9331 | 0.26432 | 8.403 | 0.002244 | 0.004084 | 0.003963 | 5.1e-05 | 0.000932 | 0.29031 | 1.3358 | 2.5474 | 100.2 | 20.86 | 37.352 | 81.5355 | 75.6719 | 77.7175 | 0.009202 | 1.0062 | 0.023951 | 23.803 | 0.042934 | 0.15724 | 0.004395 | 0.31187 | 0.31187 | 0.020521
2019-03-29 14:02 | 1553882502.5363 | 515 | 0.20861 | 0.001143 | 6.9e-05 | 99.973 | 99.4459 | 121.3114 | 21.1272 | 0.002495 | 0.002332 | 247.9071 | 0.26497 | 8.4035 | 0.002319 | 0.004211 | 0.004138 | 4.9e-05 | 0.000938 | 0.28816 | 1.3687 | 2.5475 | 100.2 | 20.86 | 37.351 | 81.6998 | 77.4864 | 77.8038 | 0.009202 | 1.006 | 0.023929 | 23.787 | 0.047285 | 0.15763 | 0.004553 | 0.32295 | 0.32295 | 0.019866
2019-03-29 14:02 | 1553882503.5373 | 516 | 0.20863 | 0.001139 | 6.9e-05 | 100.0023 | 99.552 | 121.4238 | 21.1261 | 0.002431 | 0.002328 | 247.8978 | 0.26467 | 8.4019 | 0.002341 | 0.003864 | 0.003776 | 4.9e-05 | 0.000946 | 0.29231 | 1.3212 | 2.5474 | 100.2 | 20.86 | 37.351 | 80.1773 | 78.0109 | 77.7 | 0.009202 | 1.0063 | 0.023943 | 23.7945 | 0.040368 | 0.15745 | 0.004448 | 0.31562 | 0.31562 | 0.020304
2019-03-29 14:02 | 1553882504.5374 | 517 | 0.20863 | 0.001137 | 6.9e-05 | 100.0053 | 99.1937 | 121.2945 | 21.1264 | 0.002445 | 0.002302 | 247.9041 | 0.26537 | 8.401 | 0.002309 | 0.003992 | 0.003926 | 5.2e-05 | 0.000948 | 0.2943 | 1.3459 | 2.5475 | 100.2 | 20.86 | 37.351 | 80.5193 | 77.2425 | 77.0647 | 0.009202 | 1.0081 | 0.023952 | 23.7603 | 0.072891 | 0.15787 | 0.004438 | 0.31498 | 0.31498 | 0.020399
2019-03-29 14:02 | 1553882505.5364 | 518 | 0.20864 | 0.001138 | 7e-05 | 99.9562 | 99.2367 | 121.1161 | 21.1257 | 0.002396 | 0.002308 | 247.8867 | 0.26409 | 8.4015 | 0.002346 | 0.004136 | 0.004003 | 5.1e-05 | 0.000947 | 0.27041 | 1.3305 | 2.5474 | 100.2 | 20.85 | 37.347 | 79.3428 | 78.122 | 77.2248 | 0.009195 | 1.0076 | 0.023929 | 23.7482 | 0.064622 | 0.1571 | 0.00435 | 0.30835 | 0.30835 | 0.020737
2019-03-29 14:02 | 1553882506.5374 | 519 | 0.20865 | 0.00114 | 7.1e-05 | 100.0227 | 99.222 | 121.1839 | 21.1257 | 0.002468 | 0.00231 | 247.8819 | 0.26438 | 8.4033 | 0.002305 | 0.004007 | 0.004017 | 5.2e-05 | 0.000946 | 0.29662 | 1.3752 | 2.5473 | 100.2 | 20.85 | 37.346 | 81.0542 | 77.1561 | 77.2575 | 0.009195 | 1.0075 | 0.023935 | 23.756 | 0.07189 | 0.15728 | 0.004295 | 0.30437 | 0.30437 | 0.021031
2019-03-29 14:02 | 1553882507.5374 | 520 | 0.20866 | 0.001141 | 7.1e-05 | 100.0239 | 99.2448 | 120.8633 | 21.1259 | 0.002548 | 0.002286 | 247.8892 | 0.26441 | 8.4014 | 0.002358 | 0.003521 | 0.003449 | 5.2e-05 | 0.000947 | 0.29983 | 1.3526 | 2.5474 | 100.2 | 20.85 | 37.346 | 82.995 | 78.4076 | 76.6933 | 0.009195 | 1.0092 | 0.023922 | 23.7054 | 0.069936 | 0.1573 | 0.004236 | 0.30003 | 0.30003 | 0.021338
2019-03-29 14:02 | 1553882508.5365 | 521 | 0.20864 | 0.001138 | 7.2e-05 | 100.081 | 99.4257 | 120.9466 | 21.1262 | 0.002449 | 0.002292 | 247.8717 | 0.2653 | 8.4034 | 0.002409 | 0.003616 | 0.003477 | 5.3e-05 | 0.000941 | 0.29734 | 1.3591 | 2.5475 | 100.2 | 20.85 | 37.345 | 80.6071 | 79.6365 | 76.837 | 0.009194 | 1.0087 | 0.023926 | 23.7185 | 0.058759 | 0.15782 | 0.004391 | 0.31113 | 0.31113 | 0.020646
2019-03-29 14:02 | 1553882509.5375 | 522 | 0.20864 | 0.00114 | 7.2e-05 | 100.1173 | 99.3834 | 120.6778 | 21.1268 | 0.002405 | 0.002294 | 247.8516 | 0.26426 | 8.404 | 0.00236 | 0.004136 | 0.004113 | 5.3e-05 | 0.000946 | 0.30309 | 1.3338 | 2.5473 | 100.2 | 20.85 | 37.344 | 79.5443 | 78.4731 | 76.8734 | 0.009194 | 1.0086 | 0.023898 | 23.6933 | 0.065804 | 0.15721 | 0.004361 | 0.30863 | 0.30863 | 0.020732
2019-03-29 14:02 | 1553882510.5366 | 523 | 0.20864 | 0.001139 | 7.2e-05 | 100.0622 | 99.5097 | 120.706 | 21.1269 | 0.002492 | 0.002314 | 247.8398 | 0.26433 | 8.4039 | 0.002356 | 0.003836 | 0.003783 | 5.2e-05 | 0.000949 | 0.28423 | 1.3756 | 2.5474 | 100.2 | 20.85 | 37.342 | 81.6317 | 78.371 | 77.3694 | 0.009194 | 1.0072 | 0.023884 | 23.7129 | 0.049524 | 0.15725 | 0.004363 | 0.30859 | 0.30859 | 0.02074
2019-03-29 14:02 | 1553882511.5366 | 524 | 0.20867 | 0.001138 | 7.1e-05 | 100.0463 | 99.579 | 120.7438 | 21.1265 | 0.002415 | 0.002315 | 247.8315 | 0.26413 | 8.4025 | 0.002338 | 0.003843 | 0.003712 | 5.3e-05 | 0.000956 | 0.29407 | 1.3773 | 2.5475 | 100.2 | 20.85 | 37.341 | 79.7887 | 77.9387 | 77.3806 | 0.009193 | 1.0072 | 0.023887 | 23.717 | 0.04188 | 0.15713 | 0.004199 | 0.29689 | 0.29689 | 0.021541
2019-03-29 14:02 | 1553882512.5366 | 525 | 0.20869 | 0.001137 | 7.1e-05 | 100.1098 | 99.551 | 120.833 | 21.1257 | 0.002496 | 0.002328 | 247.8259 | 0.26417 | 8.4028 | 0.002361 | 0.004299 | 0.004174 | 5.1e-05 | 0.000953 | 0.2943 | 1.371 | 2.5475 | 100.2 | 20.85 | 37.34 | 81.7438 | 78.4877 | 77.7003 | 0.009193 | 1.0063 | 0.023885 | 23.7365 | 0.050066 | 0.15716 | 0.00406 | 0.28693 | 0.28693 | 0.022292
2019-03-29 14:02 | 1553882513.5356 | 526 | 0.20867 | 0.001138 | 7e-05 | 100.1454 | 99.5684 | 121.0978 | 21.1263 | 0.002449 | 0.002279 | 247.7865 | 0.26495 | 8.4038 | 0.002329 | 0.004082 | 0.003986 | 5.2e-05 | 0.000948 | 0.27865 | 1.374 | 2.5474 | 100.2 | 20.85 | 37.338 | 80.601 | 77.7124 | 76.5292 | 0.009193 | 1.0096 | 0.023951 | 23.7228 | 0.051685 | 0.15762 | 0.004173 | 0.29591 | 0.29591 | 0.021679
2019-03-29 14:02 | 1553882514.5357 | 527 | 0.20867 | 0.001139 | 7e-05 | 100.1454 | 99.6261 | 121.2172 | 21.1269 | 0.002448 | 0.002242 | 247.7467 | 0.26495 | 8.3997 | 0.002358 | 0.003258 | 0.003254 | 5.1e-05 | 0.000947 | 0.30198 | 1.3457 | 2.5474 | 100.2 | 20.85 | 37.337 | 80.5704 | 78.4225 | 75.6327 | 0.009192 | 1.0122 | 0.023994 | 23.7041 | 0.046505 | 0.15762 | 0.004207 | 0.29889 | 0.29889 | 0.021463
2019-03-29 14:02 | 1553882515.5371 | 528 | 0.20866 | 0.001135 | 6.9e-05 | 100.1616 | 99.8233 | 121.3405 | 21.1266 | 0.002446 | 0.00229 | 247.7381 | 0.26452 | 8.4038 | 0.00238 | 0.003679 | 0.003592 | 5e-05 | 0.000948 | 0.2912 | 1.3225 | 2.5474 | 100.2 | 20.85 | 37.335 | 80.5381 | 78.9405 | 76.7843 | 0.009192 | 1.0089 | 0.023966 | 23.7553 | 0.030255 | 0.15736 | 0.004253 | 0.30186 | 0.30186 | 0.021217
2019-03-29 14:02 | 1553882516.5367 | 529 | 0.20864 | 0.001136 | 7e-05 | 100.1978 | 99.6602 | 121.4069 | 21.1252 | 0.002549 | 0.002316 | 247.6921 | 0.26473 | 8.4015 | 0.00227 | 0.004016 | 0.003959 | 4.8e-05 | 0.000953 | 0.30574 | 1.3261 | 2.5474 | 100.2 | 20.85 | 37.334 | 82.9995 | 76.2911 | 77.4091 | 0.009192 | 1.0071 | 0.023952 | 23.783 | 0.048118 | 0.15749 | 0.004399 | 0.31214 | 0.31214 | 0.020535
2019-03-29 14:02 | 1553882517.5368 | 530 | 0.20867 | 0.001133 | 6.9e-05 | 100.1259 | 99.8312 | 121.4388 | 21.1257 | 0.002499 | 0.002292 | 247.6356 | 0.2642 | 8.4021 | 0.002381 | 0.004116 | 0.003998 | 4.9e-05 | 0.000951 | 0.29679 | 1.3549 | 2.5476 | 100.2 | 20.85 | 37.331 | 81.7983 | 78.9656 | 76.8279 | 0.009191 | 1.0088 | 0.023975 | 23.7664 | 0.02637 | 0.15718 | 0.004175 | 0.2964 | 0.2964 | 0.021583
2019-03-29 14:02 | 1553882518.5362 | 531 | 0.20867 | 0.001132 | 6.9e-05 | 100.1747 | 99.6568 | 121.5261 | 21.1241 | 0.002472 | 0.002301 | 247.6219 | 0.26479 | 8.4015 | 0.002327 | 0.00426 | 0.004287 | 5.1e-05 | 0.000948 | 0.28479 | 1.3798 | 2.5476 | 100.2 | 20.85 | 37.325 | 81.1588 | 77.6619 | 77.0385 | 0.009189 | 1.0082 | 0.023976 | 23.7821 | 0.046367 | 0.15752 | 0.004224 | 0.29995 | 0.29995 | 0.021374
2019-03-29 14:02 | 1553882519.5373 | 532 | 0.20867 | 0.001133 | 7e-05 | 100.2582 | 99.693 | 121.6421 | 21.1248 | 0.002459 | 0.002287 | 247.6041 | 0.26448 | 8.4014 | 0.002338 | 0.004023 | 0.003889 | 5e-05 | 0.000949 | 0.29175 | 1.366 | 2.5474 | 100.2 | 20.85 | 37.32 | 80.8336 | 77.9411 | 76.7218 | 0.009188 | 1.0091 | 0.023998 | 23.7827 | 0.050574 | 0.15734 | 0.004182 | 0.29717 | 0.29717 | 0.021549
2019-03-29 14:02 | 1553882520.5368 | 533 | 0.20867 | 0.001134 | 7e-05 | 100.2406 | 99.7704 | 121.6975 | 21.125 | 0.002457 | 0.002243 | 247.5929 | 0.26442 | 8.4032 | 0.0023 | 0.004531 | 0.004566 | 5e-05 | 0.000944 | 0.2907 | 1.3566 | 2.5475 | 100.2 | 20.85 | 37.315 | 80.7946 | 77.0189 | 75.6533 | 0.009187 | 1.0122 | 0.024041 | 23.7517 | 0.042051 | 0.1573 | 0.004181 | 0.29759 | 0.29759 | 0.021514
2019-03-29 14:02 | 1553882521.5369 | 534 | 0.20867 | 0.001128 | 7e-05 | 100.2478 | 99.5445 | 121.9118 | 21.1242 | 0.002459 | 0.002302 | 247.5582 | 0.26453 | 8.4017 | 0.002308 | 0.003661 | 0.003521 | 5.3e-05 | 0.000945 | 0.29479 | 1.3836 | 2.5473 | 100.2 | 20.85 | 37.309 | 80.8508 | 77.2139 | 77.0767 | 0.009186 | 1.008 | 0.024013 | 23.8211 | 0.06297 | 0.15737 | 0.004193 | 0.2981 | 0.2981 | 0.021486
2019-03-29 14:02 | 1553882522.5369 | 535 | 0.20867 | 0.001128 | 7e-05 | 100.3166 | 99.8213 | 122.0742 | 21.1251 | 0.002509 | 0.002329 | 247.5398 | 0.26487 | 8.4003 | 0.002251 | 0.003963 | 0.003921 | 4.8e-05 | 0.000941 | 0.27256 | 1.3438 | 2.5473 | 100.2 | 20.85 | 37.296 | 82.0465 | 75.8562 | 77.716 | 0.009182 | 1.0062 | 0.024007 | 23.8587 | 0.044274 | 0.15757 | 0.004189 | 0.29771 | 0.29771 | 0.021541
2019-03-29 14:02 | 1553882523.5359 | 536 | 0.20867 | 0.001128 | 7e-05 | 100.2588 | 99.7141 | 121.6813 | 21.1257 | 0.002446 | 0.002312 | 247.5159 | 0.26476 | 8.4038 | 0.002347 | 0.003318 | 0.003219 | 4.8e-05 | 0.000934 | 0.3126 | 1.3476 | 2.5475 | 100.2 | 20.85 | 37.29 | 80.5295 | 78.1587 | 77.3043 | 0.009181 | 1.0074 | 0.023982 | 23.8063 | 0.048729 | 0.1575 | 0.00422 | 0.29964 | 0.29964 | 0.021393
2019-03-29 14:02 | 1553882524.5369 | 537 | 0.20867 | 0.001125 | 7e-05 | 100.3112 | 99.7104 | 121.5839 | 21.1234 | 0.002424 | 0.002301 | 247.499 | 0.26472 | 8.3998 | 0.002309 | 0.002971 | 0.0029 | 4.7e-05 | 0.000931 | 0.30237 | 1.3313 | 2.5473 | 100.2 | 20.85 | 37.273 | 80.0134 | 77.2482 | 77.0423 | 0.009177 | 1.0081 | 0.023982 | 23.7878 | 0.053738 | 0.15748 | 0.004229 | 0.30029 | 0.30029 | 0.021345
2019-03-29 14:02 | 1553882525.537 | 538 | 0.20865 | 0.001133 | 7e-05 | 100.395 | 99.7684 | 121.6052 | 21.124 | 0.002488 | 0.002289 | 247.4568 | 0.26344 | 8.4035 | 0.002318 | 0.003936 | 0.003791 | 4.3e-05 | 0.00093 | 0.27041 | 1.3684 | 2.5475 | 100.2 | 20.85 | 37.265 | 81.5416 | 77.4468 | 76.7522 | 0.009175 | 1.009 | 0.023994 | 23.7801 | 0.056006 | 0.15672 | 0.004305 | 0.30593 | 0.30593 | 0.020849
2019-03-29 14:02 | 1553882526.536 | 539 | 0.20863 | 0.001128 | 7.1e-05 | 100.3576 | 99.81 | 121.6991 | 21.1239 | 0.002461 | 0.002296 | 247.427 | 0.26408 | 8.4015 | 0.002315 | 0.003828 | 0.003727 | 4.5e-05 | 0.000927 | 0.27743 | 1.3695 | 2.5474 | 100.2 | 20.84 | 37.237 | 80.8817 | 77.3766 | 76.9229 | 0.009162 | 1.0085 | 0.023997 | 23.7951 | 0.048939 | 0.1571 | 0.004476 | 0.31826 | 0.31826 | 0.020091
2019-03-29 14:02 | 1553882527.5351 | 540 | 0.20864 | 0.001131 | 7e-05 | 100.4307 | 99.8172 | 121.9547 | 21.1242 | 0.002405 | 0.002303 | 247.4082 | 0.26396 | 8.401 | 0.00224 | 0.003941 | 0.003925 | 4.4e-05 | 0.000929 | 0.29839 | 1.3632 | 2.5474 | 100.2 | 20.84 | 37.237 | 79.5574 | 75.5901 | 77.1063 | 0.009162 | 1.008 | 0.024016 | 23.8263 | 0.054804 | 0.15703 | 0.004385 | 0.31194 | 0.31194 | 0.020489
2019-03-29 14:02 | 1553882528.5351 | 541 | 0.2086 | 0.00113 | 7e-05 | 100.3953 | 99.854 | 122.2211 | 21.1238 | 0.002404 | 0.002272 | 247.3717 | 0.26412 | 8.4012 | 0.002368 | 0.00359 | 0.003563 | 3.9e-05 | 0.000927 | 0.28102 | 1.3223 | 2.5474 | 100.2 | 20.84 | 37.221 | 79.5328 | 78.6463 | 76.3412 | 0.009158 | 1.0102 | 0.024069 | 23.8262 | 0.048356 | 0.15712 | 0.004611 | 0.32901 | 0.32901 | 0.019437
2019-03-29 14:02 | 1553882529.5351 | 542 | 0.20863 | 0.001131 | 7e-05 | 100.4414 | 99.7684 | 122.262 | 21.1239 | 0.002449 | 0.002313 | 247.3568 | 0.26353 | 8.4002 | 0.002318 | 0.003481 | 0.003365 | 4.4e-05 | 0.000926 | 0.27793 | 1.3542 | 2.5474 | 100.2 | 20.84 | 37.221 | 80.5955 | 77.4599 | 77.3273 | 0.009158 | 1.0073 | 0.024039 | 23.8638 | 0.060138 | 0.15677 | 0.004477 | 0.31893 | 0.31893 | 0.020006
2019-03-29 14:02 | 1553882530.5361 | 543 | 0.20864 | 0.001127 | 7e-05 | 100.5067 | 99.9028 | 122.3439 | 21.1244 | 0.002499 | 0.00229 | 247.3232 | 0.26333 | 8.4012 | 0.002419 | 0.003755 | 0.003655 | 4.4e-05 | 0.000921 | 0.27383 | 1.3479 | 2.5474 | 100.2 | 20.84 | 37.205 | 81.8009 | 79.8724 | 76.787 | 0.009154 | 1.0089 | 0.024065 | 23.8534 | 0.0539 | 0.15666 | 0.004381 | 0.3124 | 0.3124 | 0.020409
2019-03-29 14:02 | 1553882531.5368 | 544 | 0.20867 | 0.001135 | 7e-05 | 100.5246 | 100.0547 | 121.7374 | 21.1235 | 0.002453 | 0.0023 | 247.2984 | 0.26446 | 8.4014 | 0.002382 | 0.00379 | 0.003712 | 4.1e-05 | 0.00092 | 0.2677 | 1.355 | 2.5473 | 100.2 | 20.84 | 37.199 | 80.7091 | 79.0038 | 77.0261 | 0.009153 | 1.0082 | 0.023997 | 23.8023 | 0.041911 | 0.15733 | 0.004197 | 0.29821 | 0.29821 | 0.021472
2019-03-29 14:02 | 1553882532.5353 | 545 | 0.20866 | 0.00113 | 7e-05 | 100.5459 | 100.0564 | 121.5642 | 21.124 | 0.002464 | 0.00224 | 247.2761 | 0.26432 | 8.4016 | 0.002397 | 0.004243 | 0.004266 | 4.3e-05 | 0.00092 | 0.28191 | 1.3162 | 2.5475 | 100.2 | 20.84 | 37.188 | 80.9708 | 79.3536 | 75.5735 | 0.00915 | 1.0124 | 0.02403 | 23.736 | 0.043649 | 0.15724 | 0.004303 | 0.30627 | 0.30627 | 0.020896
2019-03-29 14:02 | 1553882533.5362 | 546 | 0.20866 | 0.00113 | 7e-05 | 100.6013 | 99.9435 | 121.5402 | 21.1241 | 0.002427 | 0.002313 | 247.2769 | 0.26467 | 8.4008 | 0.002376 | 0.004245 | 0.004148 | 4.3e-05 | 0.000922 | 0.29231 | 1.3296 | 2.5474 | 100.2 | 20.84 | 37.183 | 80.0749 | 78.8525 | 77.34 | 0.009149 | 1.0073 | 0.023967 | 23.7937 | 0.058679 | 0.15745 | 0.004276 | 0.30352 | 0.30352 | 0.021113
2019-03-29 14:02 | 1553882534.5364 | 547 | 0.20866 | 0.001128 | 7e-05 | 100.5977 | 100.1315 | 121.5562 | 21.1236 | 0.002526 | 0.002319 | 247.262 | 0.2639 | 8.4026 | 0.002378 | 0.003568 | 0.003572 | 3.7e-05 | 0.000924 | 0.29247 | 1.4006 | 2.5475 | 100.2 | 20.83 | 37.173 | 82.4513 | 78.891 | 77.4683 | 0.009141 | 1.0069 | 0.023964 | 23.7996 | 0.041545 | 0.15699 | 0.004304 | 0.3055 | 0.3055 | 0.020915
2019-03-29 14:02 | 1553882535.5363 | 548 | 0.20863 | 0.001126 | 7e-05 | 100.6318 | 100.0844 | 121.6839 | 21.1242 | 0.002451 | 0.002238 | 247.2353 | 0.26444 | 8.4016 | 0.002452 | 0.004264 | 0.004161 | 4e-05 | 0.000927 | 0.27151 | 1.374 | 2.5474 | 100.2 | 20.83 | 37.167 | 80.6396 | 80.6785 | 75.5318 | 0.009139 | 1.0125 | 0.024043 | 23.7462 | 0.048787 | 0.15732 | 0.00444 | 0.31633 | 0.31633 | 0.020241
2019-03-29 14:02 | 1553882536.5354 | 549 | 0.20864 | 0.001124 | 7e-05 | 100.677 | 100.2844 | 121.794 | 21.1241 | 0.002472 | 0.002287 | 247.215 | 0.26447 | 8.4015 | 0.002421 | 0.003891 | 0.003792 | 4.1e-05 | 0.00093 | 0.27975 | 1.3679 | 2.5473 | 100.2 | 20.83 | 37.167 | 81.1456 | 79.9266 | 76.7117 | 0.009139 | 1.0091 | 0.024014 | 23.7971 | 0.034948 | 0.15733 | 0.004387 | 0.31209 | 0.31209 | 0.020518
2019-03-29 14:02 | 1553882537.5364 | 550 | 0.20865 | 0.001122 | 7.1e-05 | 100.7467 | 100.2063 | 121.8578 | 21.1237 | 0.002468 | 0.002309 | 247.1982 | 0.26442 | 8.4009 | 0.002427 | 0.004092 | 0.004058 | 4e-05 | 0.000932 | 0.27759 | 1.3707 | 2.5474 | 100.2 | 20.83 | 37.165 | 81.0546 | 80.0791 | 77.2329 | 0.009139 | 1.0076 | 0.024002 | 23.8211 | 0.048116 | 0.1573 | 0.004357 | 0.30971 | 0.30971 | 0.020671
2019-03-29 14:02 | 1553882538.5354 | 551 | 0.20863 | 0.001126 | 7.2e-05 | 100.827 | 100.4639 | 121.7163 | 21.1242 | 0.002444 | 0.002316 | 247.2012 | 0.26426 | 8.4037 | 0.002314 | 0.004371 | 0.004332 | 4.2e-05 | 0.000928 | 0.3194 | 1.3419 | 2.5474 | 100.2 | 20.83 | 37.165 | 80.4718 | 77.355 | 77.4165 | 0.009139 | 1.0071 | 0.023982 | 23.8135 | 0.032264 | 0.15721 | 0.004434 | 0.31499 | 0.31499 | 0.020313
2019-03-29 14:02 | 1553882539.5365 | 552 | 0.20867 | 0.001124 | 7.1e-05 | 100.7665 | 100.3837 | 121.6041 | 21.1246 | 0.002481 | 0.002299 | 247.1547 | 0.26364 | 8.4035 | 0.002398 | 0.003904 | 0.003805 | 3.8e-05 | 0.000926 | 0.32012 | 1.3591 | 2.5475 | 100.2 | 20.83 | 37.164 | 81.3644 | 79.3783 | 76.996 | 0.009139 | 1.0083 | 0.023985 | 23.7883 | 0.034042 | 0.15684 | 0.004238 | 0.30094 | 0.30094 | 0.021212
2019-03-29 14:02 | 1553882540.5365 | 553 | 0.20863 | 0.00113 | 7.1e-05 | 100.8701 | 100.3544 | 121.7052 | 21.1241 | 0.002477 | 0.002274 | 247.1108 | 0.26384 | 8.4026 | 0.002392 | 0.003882 | 0.003812 | 4.1e-05 | 0.000919 | 0.27975 | 1.336 | 2.5474 | 100.2 | 20.83 | 37.163 | 81.2875 | 79.233 | 76.404 | 0.009138 | 1.01 | 0.024016 | 23.778 | 0.045846 | 0.15696 | 0.004486 | 0.31917 | 0.31917 | 0.020015
2019-03-29 14:02 | 1553882541.5355 | 554 | 0.20861 | 0.001129 | 7.1e-05 | 100.8362 | 100.3342 | 121.8655 | 21.125 | 0.002443 | 0.002316 | 247.094 | 0.26439 | 8.403 | 0.002318 | 0.004323 | 0.004263 | 3.9e-05 | 0.00092 | 0.27328 | 1.3582 | 2.5474 | 100.2 | 20.83 | 37.164 | 80.4586 | 77.4656 | 77.4096 | 0.009139 | 1.0071 | 0.023997 | 23.8279 | 0.044642 | 0.15729 | 0.004568 | 0.32485 | 0.32485 | 0.019706
2019-03-29 14:02 | 1553882542.5356 | 555 | 0.20871 | 0.001133 | 7.2e-05 | 100.8365 | 100.5612 | 121.6951 | 21.1237 | 0.002454 | 0.002292 | 247.0806 | 0.26384 | 8.4002 | 0.00235 | 0.004096 | 0.004026 | 4e-05 | 0.000924 | 0.31293 | 1.33 | 2.5473 | 100.2 | 20.83 | 37.165 | 80.7225 | 78.2128 | 76.833 | 0.009139 | 1.0088 | 0.024 | 23.7916 | 0.024458 | 0.15696 | 0.003982 | 0.28262 | 0.28262 | 0.022603
2019-03-29 14:02 | 1553882543.5346 | 556 | 0.20863 | 0.001129 | 7.2e-05 | 100.897 | 100.4803 | 121.4801 | 21.124 | 0.002427 | 0.00231 | 247.0398 | 0.26388 | 8.4014 | 0.002336 | 0.004078 | 0.003979 | 4.5e-05 | 0.000926 | 0.28008 | 1.3077 | 2.5472 | 100.2 | 20.83 | 37.165 | 80.0856 | 77.8793 | 77.2519 | 0.009139 | 1.0075 | 0.023964 | 23.7848 | 0.037017 | 0.15699 | 0.004436 | 0.31491 | 0.31491 | 0.02029
2019-03-29 14:02 | 1553882544.5346 | 557 | 0.20862 | 0.001129 | 7.2e-05 | 100.9219 | 100.4909 | 121.436 | 21.1238 | 0.002442 | 0.00231 | 247.0095 | 0.2649 | 8.401 | 0.002322 | 0.003865 | 0.003771 | 4.4e-05 | 0.000922 | 0.29767 | 1.3785 | 2.5473 | 100.2 | 20.83 | 37.165 | 80.4248 | 77.545 | 77.2675 | 0.009139 | 1.0075 | 0.023959 | 23.781 | 0.038278 | 0.15759 | 0.0045 | 0.31938 | 0.31938 | 0.020082
2019-03-29 14:02 | 1553882545.5357 | 558 | 0.20863 | 0.001128 | 7.2e-05 | 100.9278 | 100.5141 | 121.3231 | 21.1245 | 0.0024 | 0.002284 | 246.9756 | 0.26478 | 8.4015 | 0.002389 | 0.00353 | 0.003437 | 4e-05 | 0.000919 | 0.29479 | 1.3489 | 2.5475 | 100.2 | 20.83 | 37.167 | 79.4313 | 79.1646 | 76.6319 | 0.009139 | 1.0093 | 0.02397 | 23.7484 | 0.036743 | 0.15752 | 0.004487 | 0.3186 | 0.3186 | 0.020123
2019-03-29 14:02 | 1553882546.5347 | 559 | 0.20863 | 0.001129 | 7.2e-05 | 101.0293 | 100.4847 | 121.2909 | 21.1243 | 0.002457 | 0.002276 | 246.9647 | 0.26416 | 8.4049 | 0.002386 | 0.004134 | 0.00405 | 4.3e-05 | 0.000916 | 0.28351 | 1.3366 | 2.5474 | 100.2 | 20.83 | 37.168 | 80.8058 | 79.0966 | 76.4486 | 0.00914 | 1.0099 | 0.023973 | 23.739 | 0.048343 | 0.15715 | 0.004434 | 0.31483 | 0.31483 | 0.020316
2019-03-29 14:02 | 1553882547.5357 | 560 | 0.20863 | 0.001121 | 7.1e-05 | 101.0637 | 100.6697 | 121.0067 | 21.1243 | 0.002429 | 0.002294 | 246.9517 | 0.26388 | 8.4012 | 0.002346 | 0.004229 | 0.004217 | 4.7e-05 | 0.000914 | 0.30182 | 1.3081 | 2.5474 | 100.2 | 20.83 | 37.17 | 80.1227 | 78.1268 | 76.8811 | 0.00914 | 1.0086 | 0.02393 | 23.7259 | 0.03494 | 0.15698 | 0.004499 | 0.31898 | 0.31898 | 0.02003
2019-03-29 14:02 | 1553882548.5357 | 561 | 0.20863 | 0.001131 | 7.1e-05 | 101 | 100.5264 | 120.8375 | 21.1239 | 0.002475 | 0.002302 | 246.9366 | 0.26383 | 8.4029 | 0.002402 | 0.003379 | 0.003333 | 4.4e-05 | 0.000913 | 0.29087 | 1.3616 | 2.5474 | 100.2 | 20.83 | 37.17 | 81.2194 | 79.4766 | 77.0796 | 0.00914 | 1.008 | 0.023907 | 23.716 | 0.042041 | 0.15696 | 0.004479 | 0.31724 | 0.31724 | 0.020136
2019-03-29 14:02 | 1553882549.5358 | 562 | 0.20863 | 0.001131 | 7.1e-05 | 101.0673 | 100.6086 | 120.7748 | 21.1233 | 0.002449 | 0.002259 | 246.9214 | 0.26476 | 8.402 | 0.002415 | 0.003762 | 0.003657 | 4.3e-05 | 0.000916 | 0.28263 | 1.3847 | 2.5475 | 100.2 | 20.83 | 37.169 | 80.606 | 79.7817 | 76.033 | 0.00914 | 1.0111 | 0.023936 | 23.6744 | 0.040688 | 0.1575 | 0.004481 | 0.31775 | 0.31775 | 0.020175
2019-03-29 14:02 | 1553882550.5348 | 563 | 0.20862 | 0.001135 | 7.2e-05 | 101.0679 | 100.6974 | 120.8809 | 21.1244 | 0.002436 | 0.002323 | 246.9292 | 0.26504 | 8.403 | 0.002312 | 0.003872 | 0.003776 | 4.4e-05 | 0.000916 | 0.29623 | 1.3922 | 2.5476 | 100.2 | 20.83 | 37.167 | 80.2952 | 77.3163 | 77.565 | 0.009139 | 1.0066 | 0.023894 | 23.7367 | 0.032855 | 0.15767 | 0.004502 | 0.31872 | 0.31872 | 0.020135
2019-03-29 14:02 | 1553882551.5348 | 564 | 0.20862 | 0.001132 | 7.2e-05 | 101.196 | 100.7557 | 121.3726 | 21.1241 | 0.002462 | 0.002283 | 246.8836 | 0.26516 | 8.4017 | 0.002224 | 0.003324 | 0.003292 | 4.7e-05 | 0.000915 | 0.30773 | 1.3312 | 2.5473 | 100.2 | 20.83 | 37.167 | 80.9257 | 75.2004 | 76.6034 | 0.009139 | 1.0094 | 0.023976 | 23.7523 | 0.039004 | 0.15774 | 0.004495 | 0.31929 | 0.31929 | 0.020107
2019-03-29 14:03 | 1553882552.5348 | 565 | 0.20861 | 0.001133 | 7.2e-05 | 101.2164 | 100.8124 | 121.7785 | 21.1236 | 0.002453 | 0.002305 | 246.8548 | 0.26492 | 8.4025 | 0.002263 | 0.004326 | 0.004269 | 4.2e-05 | 0.000913 | 0.29214 | 1.3924 | 2.5475 | 100.2 | 20.83 | 37.166 | 80.6998 | 76.1369 | 77.151 | 0.009139 | 1.0078 | 0.023997 | 23.8106 | 0.035775 | 0.1576 | 0.004551 | 0.32356 | 0.32356 | 0.019824
2019-03-29 14:03 | 1553882553.5359 | 566 | 0.20863 | 0.001129 | 7.2e-05 | 101.2131 | 100.6725 | 121.939 | 21.1236 | 0.002409 | 0.002256 | 246.8452 | 0.26478 | 8.4025 | 0.002347 | 0.003723 | 0.00371 | 4.6e-05 | 0.000915 | 0.27649 | 1.3607 | 2.5475 | 100.2 | 20.83 | 37.163 | 79.6496 | 78.1589 | 75.9701 | 0.009138 | 1.0112 | 0.024054 | 23.7861 | 0.047907 | 0.15752 | 0.004489 | 0.31986 | 0.31986 | 0.020043
2019-03-29 14:03 | 1553882554.5359 | 567 | 0.20864 | 0.001131 | 7.2e-05 | 101.2439 | 100.8752 | 122.028 | 21.125 | 0.002428 | 0.002275 | 246.8173 | 0.26578 | 8.3999 | 0.002288 | 0.004547 | 0.004431 | 4.9e-05 | 0.000908 | 0.30342 | 1.4026 | 2.5474 | 100.2 | 20.83 | 37.163 | 80.1015 | 76.7462 | 76.4257 | 0.009138 | 1.0099 | 0.024047 | 23.8103 | 0.032638 | 0.15812 | 0.004404 | 0.31367 | 0.31367 | 0.020516
2019-03-29 14:03 | 1553882555.5353 | 568 | 0.20862 | 0.001129 | 7.2e-05 | 101.3227 | 100.8274 | 121.9705 | 21.1246 | 0.002454 | 0.002311 | 246.8076 | 0.26361 | 8.4035 | 0.002343 | 0.004362 | 0.004376 | 4.5e-05 | 0.000908 | 0.30807 | 1.3782 | 2.5475 | 100.2 | 20.83 | 37.161 | 80.7209 | 78.0531 | 77.2929 | 0.009138 | 1.0074 | 0.024011 | 23.8342 | 0.043828 | 0.15682 | 0.0045 | 0.3201 | 0.3201 | 0.01994
2019-03-29 14:03 | 1553882556.534 | 569 | 0.20862 | 0.001135 | 7.1e-05 | 101.3011 | 100.9114 | 121.7508 | 21.1236 | 0.002459 | 0.002278 | 246.7889 | 0.26526 | 8.4029 | 0.002288 | 0.003127 | 0.003023 | 4.5e-05 | 0.000903 | 0.27007 | 1.3521 | 2.5474 | 100.2 | 20.83 | 37.16 | 80.8416 | 76.7362 | 76.483 | 0.009138 | 1.0098 | 0.024017 | 23.7852 | 0.03448 | 0.1578 | 0.004528 | 0.32228 | 0.32228 | 0.019929
2019-03-29 14:03 | 1553882557.534 | 570 | 0.20863 | 0.001131 | 7.1e-05 | 101.4427 | 101.0728 | 121.8247 | 21.1244 | 0.002416 | 0.002284 | 246.7673 | 0.26524 | 8.4004 | 0.002281 | 0.003971 | 0.003904 | 4.7e-05 | 0.000903 | 0.2995 | 1.3581 | 2.5474 | 100.2 | 20.83 | 37.159 | 79.8053 | 76.5727 | 76.6443 | 0.009137 | 1.0093 | 0.024019 | 23.7979 | 0.032676 | 0.15779 | 0.004433 | 0.31543 | 0.31543 | 0.02036
2019-03-29 14:03 | 1553882558.5345 | 571 | 0.20863 | 0.001133 | 7.1e-05 | 101.4538 | 101.0342 | 122.0791 | 21.1241 | 0.002489 | 0.002342 | 246.7446 | 0.26446 | 8.4021 | 0.002311 | 0.003953 | 0.003927 | 4.7e-05 | 0.000903 | 0.2995 | 1.3333 | 2.5473 | 100.2 | 20.83 | 37.157 | 81.5737 | 77.2903 | 78.0401 | 0.009137 | 1.0053 | 0.023996 | 23.8702 | 0.037066 | 0.15733 | 0.004478 | 0.31837 | 0.31837 | 0.020112
2019-03-29 14:03 | 1553882559.5341 | 572 | 0.20863 | 0.001135 | 7.1e-05 | 101.5454 | 100.8953 | 122.3085 | 21.1234 | 0.002449 | 0.002316 | 246.7239 | 0.26534 | 8.4011 | 0.002272 | 0.004 | 0.003944 | 4.5e-05 | 0.000908 | 0.29717 | 1.3325 | 2.5474 | 100.2 | 20.83 | 37.156 | 80.613 | 76.3557 | 77.4032 | 0.009137 | 1.0071 | 0.024041 | 23.8709 | 0.057443 | 0.15785 | 0.004479 | 0.31906 | 0.31906 | 0.020136
2019-03-29 14:03 | 1553882560.5341 | 573 | 0.20861 | 0.001136 | 7.1e-05 | 101.4271 | 101.0294 | 122.0721 | 21.1232 | 0.002423 | 0.002274 | 246.6941 | 0.26438 | 8.4013 | 0.0023 | 0.004437 | 0.004327 | 4.7e-05 | 0.000911 | 0.28816 | 1.3492 | 2.5473 | 100.2 | 20.82 | 37.153 | 79.9744 | 77.0189 | 76.4031 | 0.00913 | 1.01 | 0.024052 | 23.8138 | 0.03514 | 0.15728 | 0.004577 | 0.32628 | 0.32628 | 0.019619
2019-03-29 14:03 | 1553882561.5342 | 574 | 0.20863 | 0.001133 | 7.1e-05 | 101.4906 | 101.0724 | 121.8773 | 21.1234 | 0.002433 | 0.002322 | 246.6809 | 0.26447 | 8.4014 | 0.002377 | 0.003785 | 0.003796 | 4.4e-05 | 0.000911 | 0.31349 | 1.3391 | 2.5475 | 100.2 | 20.83 | 37.153 | 80.2279 | 78.8766 | 77.5609 | 0.009136 | 1.0067 | 0.023993 | 23.8342 | 0.036928 | 0.15733 | 0.004481 | 0.31854 | 0.31854 | 0.020103
2019-03-29 14:03 | 1553882562.6051 | 575 | 0.20863 | 0.001138 | 7.2e-05 | 101.5439 | 101.183 | 121.7911 | 21.1234 | 0.002511 | 0.002279 | 246.6615 | 0.26538 | 8.4027 | 0.002308 | 0.004033 | 0.00389 | 4.7e-05 | 0.000911 | 0.32211 | 1.3334 | 2.5473 | 100.2 | 20.83 | 37.15 | 82.1002 | 77.2108 | 76.5124 | 0.009135 | 1.0097 | 0.02402 | 23.7901 | 0.031842 | 0.15787 | 0.004465 | 0.31771 | 0.31771 | 0.020224
2019-03-29 14:03 | 1553882563.5484 | 576 | 0.20861 | 0.001135 | 7.1e-05 | 101.566 | 101.2243 | 122.0514 | 21.1236 | 0.002436 | 0.002299 | 246.6356 | 0.26484 | 8.4016 | 0.002282 | 0.003767 | 0.003778 | 4.6e-05 | 0.000917 | 0.32068 | 1.3649 | 2.5474 | 100.2 | 20.83 | 37.147 | 80.2874 | 76.5959 | 76.988 | 0.009134 | 1.0083 | 0.02403 | 23.8317 | 0.030143 | 0.15755 | 0.004595 | 0.32722 | 0.32722 | 0.019597
2019-03-29 14:03 | 1553882564.5475 | 577 | 0.20858 | 0.001134 | 7.2e-05 | 101.6208 | 101.0786 | 122.1937 | 21.123 | 0.002503 | 0.002292 | 246.6108 | 0.26518 | 8.4033 | 0.002337 | 0.003774 | 0.003748 | 4.7e-05 | 0.000917 | 0.26936 | 1.3858 | 2.5474 | 100.2 | 20.83 | 37.144 | 81.9048 | 77.9113 | 76.8378 | 0.009134 | 1.0087 | 0.024049 | 23.8405 | 0.047852 | 0.15776 | 0.004726 | 0.33692 | 0.33692 | 0.019057
2019-03-29 14:03 | 1553882565.5479 | 578 | 0.20864 | 0.001134 | 7.1e-05 | 101.5753 | 100.9503 | 122.5446 | 21.1226 | 0.002419 | 0.002316 | 246.5985 | 0.26483 | 8.4015 | 0.002309 | 0.004151 | 0.004054 | 4.9e-05 | 0.000918 | 0.29231 | 1.3368 | 2.5475 | 100.2 | 20.83 | 37.142 | 79.8769 | 77.2342 | 77.4079 | 0.009133 | 1.0071 | 0.024064 | 23.8941 | 0.05521 | 0.15755 | 0.004369 | 0.31142 | 0.31142 | 0.02059
2019-03-29 14:03 | 1553882566.5486 | 579 | 0.20861 | 0.001135 | 7.2e-05 | 101.6953 | 101.0612 | 122.8704 | 21.123 | 0.002465 | 0.002301 | 246.5888 | 0.26499 | 8.4007 | 0.002346 | 0.004118 | 0.004081 | 4.7e-05 | 0.000918 | 0.29574 | 1.3557 | 2.5474 | 100.2 | 20.83 | 37.137 | 80.9785 | 78.1169 | 77.0538 | 0.009132 | 1.0081 | 0.024108 | 23.9138 | 0.055951 | 0.15764 | 0.004603 | 0.32886 | 0.32886 | 0.019509
2019-03-29 14:03 | 1553882567.5492 | 580 | 0.20859 | 0.001138 | 7.1e-05 | 101.6693 | 101.1571 | 123.1692 | 21.1241 | 0.002499 | 0.002291 | 246.5585 | 0.2651 | 8.4025 | 0.0023 | 0.003947 | 0.003865 | 4.7e-05 | 0.000919 | 0.27367 | 1.3542 | 2.5475 | 100.2 | 20.83 | 37.134 | 81.8145 | 77.0315 | 76.8131 | 0.009131 | 1.0088 | 0.024145 | 23.9346 | 0.045176 | 0.15771 | 0.004701 | 0.33649 | 0.33649 | 0.019075
2019-03-29 14:03 | 1553882568.5485 | 581 | 0.2086 | 0.001139 | 7.2e-05 | 101.6343 | 101.1642 | 123.0321 | 21.1234 | 0.002459 | 0.002294 | 246.529 | 0.26417 | 8.4012 | 0.002277 | 0.003125 | 0.003041 | 4.7e-05 | 0.000924 | 0.29535 | 1.309 | 2.5473 | 100.2 | 20.83 | 37.131 | 80.8424 | 76.4725 | 76.8812 | 0.00913 | 1.0086 | 0.02413 | 23.9236 | 0.041461 | 0.15716 | 0.004609 | 0.3296 | 0.3296 | 0.019406
2019-03-29 14:03 | 1553882569.5476 | 582 | 0.20861 | 0.001142 | 7.2e-05 | 101.6205 | 101.1465 | 122.6455 | 21.1237 | 0.002493 | 0.002278 | 246.523 | 0.26419 | 8.4016 | 0.002296 | 0.003829 | 0.00376 | 4.8e-05 | 0.00093 | 0.29856 | 1.3194 | 2.5474 | 100.2 | 20.83 | 37.128 | 81.6591 | 76.9259 | 76.4821 | 0.00913 | 1.0098 | 0.024105 | 23.8724 | 0.041818 | 0.15717 | 0.004575 | 0.32684 | 0.32684 | 0.019572
2019-03-29 14:03 | 1553882570.5486 | 583 | 0.20866 | 0.00114 | 7.1e-05 | 101.6552 | 101.2871 | 122.485 | 21.1233 | 0.002452 | 0.002279 | 246.4981 | 0.26394 | 8.4038 | 0.002238 | 0.004267 | 0.004099 | 4.7e-05 | 0.000933 | 0.32925 | 1.3484 | 2.5474 | 100.2 | 20.83 | 37.123 | 80.6862 | 75.5347 | 76.513 | 0.009128 | 1.0097 | 0.024089 | 23.8578 | 0.032448 | 0.15702 | 0.004288 | 0.30584 | 0.30584 | 0.020896
2019-03-29 14:03 | 1553882571.5487 | 584 | 0.2086 | 0.001142 | 7.1e-05 | 101.6726 | 101.2315 | 122.1763 | 21.1234 | 0.002447 | 0.00229 | 246.5014 | 0.26411 | 8.4041 | 0.002305 | 0.004256 | 0.004172 | 4.5e-05 | 0.000933 | 0.30071 | 1.355 | 2.5474 | 100.2 | 20.83 | 37.118 | 80.5485 | 77.1306 | 76.7774 | 0.009127 | 1.0089 | 0.024049 | 23.8367 | 0.038889 | 0.15712 | 0.004661 | 0.33233 | 0.33233 | 0.019242
2019-03-29 14:03 | 1553882572.5482 | 585 | 0.20862 | 0.001141 | 7e-05 | 101.654 | 101.1366 | 121.824 | 21.1229 | 0.002485 | 0.002295 | 246.4804 | 0.26355 | 8.4013 | 0.00231 | 0.003305 | 0.00322 | 4.5e-05 | 0.000933 | 0.27671 | 1.3262 | 2.5473 | 100.2 | 20.83 | 37.115 | 81.4666 | 77.2541 | 76.9103 | 0.009127 | 1.0085 | 0.02401 | 23.8068 | 0.045644 | 0.15679 | 0.004489 | 0.3194 | 0.3194 | 0.019979
2019-03-29 14:03 | 1553882573.548 | 586 | 0.20859 | 0.00114 | 7e-05 | 101.6534 | 101.3332 | 121.4571 | 21.1228 | 0.002508 | 0.00232 | 246.443 | 0.26433 | 8.4015 | 0.002319 | 0.003955 | 0.003818 | 4.7e-05 | 0.00093 | 0.30917 | 1.3332 | 2.5474 | 100.2 | 20.83 | 37.108 | 82.0238 | 77.4831 | 77.4954 | 0.009125 | 1.0068 | 0.023954 | 23.7908 | 0.028222 | 0.15725 | 0.004671 | 0.33176 | 0.33176 | 0.019291
2019-03-29 14:03 | 1553882574.5478 | 587 | 0.20863 | 0.00114 | 7e-05 | 101.7085 | 101.0458 | 121.2616 | 21.1225 | 0.002378 | 0.002277 | 246.4278 | 0.26355 | 8.4039 | 0.002338 | 0.003468 | 0.003428 | 4.3e-05 | 0.000927 | 0.29999 | 1.3217 | 2.5473 | 100.2 | 20.83 | 37.106 | 78.8923 | 77.9396 | 76.4748 | 0.009124 | 1.0098 | 0.023969 | 23.737 | 0.058469 | 0.15678 | 0.004434 | 0.31491 | 0.31491 | 0.020263
2019-03-29 14:03 | 1553882575.5478 | 588 | 0.20864 | 0.001136 | 7e-05 | 101.6941 | 100.9332 | 120.7999 | 21.123 | 0.00248 | 0.002303 | 246.415 | 0.2647 | 8.4021 | 0.002263 | 0.004111 | 0.003953 | 4.6e-05 | 0.000924 | 0.29391 | 1.3394 | 2.5473 | 100.2 | 20.84 | 37.102 | 81.3479 | 76.1317 | 77.0965 | 0.009129 | 1.008 | 0.023902 | 23.7129 | 0.06718 | 0.15747 | 0.00438 | 0.31012 | 0.31012 | 0.020666
2019-03-29 14:03 | 1553882576.5468 | 589 | 0.20859 | 0.001138 | 7.1e-05 | 101.5858 | 101.139 | 120.7311 | 21.1226 | 0.002475 | 0.002338 | 246.3928 | 0.26369 | 8.404 | 0.002383 | 0.004309 | 0.004374 | 4.7e-05 | 0.000922 | 0.29407 | 1.3523 | 2.5474 | 100.2 | 20.84 | 37.102 | 81.2175 | 79.0179 | 77.9287 | 0.009129 | 1.0056 | 0.023867 | 23.7343 | 0.039425 | 0.15687 | 0.004702 | 0.33272 | 0.33272 | 0.019189
2019-03-29 14:03 | 1553882577.5469 | 590 | 0.20861 | 0.001145 | 7e-05 | 101.6091 | 100.8905 | 120.4729 | 21.1227 | 0.002447 | 0.002294 | 246.3682 | 0.2639 | 8.4026 | 0.002309 | 0.003622 | 0.003582 | 4.4e-05 | 0.000918 | 0.30165 | 1.3408 | 2.5475 | 100.2 | 20.84 | 37.1 | 80.5467 | 77.2382 | 76.8843 | 0.009128 | 1.0086 | 0.023877 | 23.6736 | 0.063481 | 0.15699 | 0.004553 | 0.32222 | 0.32222 | 0.01983
2019-03-29 14:03 | 1553882578.5469 | 591 | 0.20864 | 0.001145 | 7e-05 | 101.5639 | 100.9933 | 120.2495 | 21.1225 | 0.00245 | 0.002311 | 246.3409 | 0.26506 | 8.4008 | 0.00231 | 0.004098 | 0.004003 | 4.7e-05 | 0.000917 | 0.29231 | 1.3229 | 2.5474 | 100.2 | 20.84 | 37.1 | 80.63 | 77.2582 | 77.294 | 0.009128 | 1.0074 | 0.023841 | 23.6655 | 0.050399 | 0.15769 | 0.004394 | 0.31034 | 0.31034 | 0.02068
2019-03-29 14:03 | 1553882579.5469 | 592 | 0.20859 | 0.001142 | 7e-05 | 101.6447 | 101.095 | 120.4959 | 21.1222 | 0.002443 | 0.00231 | 246.3071 | 0.26402 | 8.4004 | 0.002276 | 0.003295 | 0.003224 | 4.6e-05 | 0.00092 | 0.28799 | 1.3532 | 2.5475 | 100.2 | 20.84 | 37.098 | 80.4684 | 76.4559 | 77.268 | 0.009128 | 1.0075 | 0.023866 | 23.6888 | 0.04851 | 0.15707 | 0.004702 | 0.33276 | 0.33276 | 0.019211
2019-03-29 14:03 | 1553882580.548 | 593 | 0.20858 | 0.001147 | 7e-05 | 101.569 | 101.0905 | 120.3743 | 21.1217 | 0.00246 | 0.002265 | 246.2869 | 0.26388 | 8.4022 | 0.002234 | 0.004699 | 0.004602 | 4.5e-05 | 0.00092 | 0.31526 | 1.3592 | 2.5474 | 100.2 | 20.84 | 37.094 | 80.8609 | 75.4327 | 76.1811 | 0.009127 | 1.0106 | 0.023891 | 23.6401 | 0.042236 | 0.15699 | 0.004741 | 0.33596 | 0.33596 | 0.019018
2019-03-29 14:03 | 1553882581.5471 | 594 | 0.20862 | 0.001142 | 7e-05 | 101.6007 | 101.1776 | 120.1953 | 21.1216 | 0.002497 | 0.002331 | 246.2512 | 0.26409 | 8.4034 | 0.002316 | 0.00411 | 0.004118 | 4.7e-05 | 0.000921 | 0.30182 | 1.3817 | 2.5474 | 100.2 | 20.84 | 37.094 | 81.7467 | 77.393 | 77.7688 | 0.009127 | 1.0061 | 0.02382 | 23.6761 | 0.037333 | 0.15711 | 0.00452 | 0.31912 | 0.31912 | 0.020037
2019-03-29 14:03 | 1553882582.5481 | 595 | 0.20859 | 0.001142 | 7e-05 | 101.5708 | 100.8953 | 120.1354 | 21.1215 | 0.002456 | 0.002285 | 246.2071 | 0.26453 | 8.4005 | 0.002271 | 0.003942 | 0.003933 | 4.3e-05 | 0.00092 | 0.2891 | 1.3687 | 2.5473 | 100.2 | 20.84 | 37.09 | 80.7641 | 76.3351 | 76.6672 | 0.009126 | 1.0092 | 0.023851 | 23.6331 | 0.059684 | 0.15737 | 0.004703 | 0.33263 | 0.33263 | 0.019255
2019-03-29 14:03 | 1553882583.5481 | 596 | 0.20858 | 0.001146 | 7e-05 | 101.4921 | 100.9711 | 120.3755 | 21.1204 | 0.0024 | 0.002319 | 246.183 | 0.26421 | 8.4005 | 0.002264 | 0.003672 | 0.003555 | 4.2e-05 | 0.000931 | 0.31116 | 1.377 | 2.5474 | 100.2 | 20.84 | 37.09 | 79.4165 | 76.158 | 77.4691 | 0.009126 | 1.0069 | 0.023848 | 23.6838 | 0.046034 | 0.15718 | 0.00474 | 0.33527 | 0.33527 | 0.01908
2019-03-29 14:03 | 1553882584.5472 | 597 | 0.20857 | 0.001148 | 7e-05 | 101.437 | 100.955 | 120.2222 | 21.1203 | 0.002452 | 0.002263 | 246.1808 | 0.26416 | 8.4027 | 0.002332 | 0.004644 | 0.004549 | 4.1e-05 | 0.000937 | 0.26233 | 1.3609 | 2.5474 | 100.2 | 20.85 | 37.085 | 80.6708 | 77.7829 | 76.1266 | 0.00913 | 1.0108 | 0.023878 | 23.6233 | 0.0426 | 0.15715 | 0.004801 | 0.34004 | 0.34004 | 0.01881
2019-03-29 14:03 | 1553882585.5471 | 598 | 0.20856 | 0.001146 | 7e-05 | 101.4331 | 101.0107 | 120.4725 | 21.1214 | 0.002475 | 0.002338 | 246.1471 | 0.26472 | 8.3997 | 0.002318 | 0.004155 | 0.004112 | 4.7e-05 | 0.000938 | 0.28998 | 1.3901 | 2.5473 | 100.2 | 20.85 | 37.084 | 81.2286 | 77.4537 | 77.9354 | 0.00913 | 1.0056 | 0.023841 | 23.7091 | 0.03733 | 0.15749 | 0.004884 | 0.34547 | 0.34547 | 0.018553
2019-03-29 14:03 | 1553882586.5471 | 599 | 0.20858 | 0.001139 | 6.9e-05 | 101.4089 | 100.868 | 120.4225 | 21.1207 | 0.002471 | 0.002298 | 246.1016 | 0.26466 | 8.3996 | 0.002274 | 0.003964 | 0.00391 | 4.5e-05 | 0.000942 | 0.32068 | 1.3373 | 2.5475 | 100.2 | 20.85 | 37.078 | 81.1219 | 76.3874 | 76.9703 | 0.009129 | 1.0084 | 0.023869 | 23.6715 | 0.047833 | 0.15745 | 0.004746 | 0.336 | 0.336 | 0.019072
2019-03-29 14:03 | 1553882587.5482 | 600 | 0.20856 | 0.001149 | 6.9e-05 | 101.3577 | 100.8059 | 120.38 | 21.1211 | 0.002441 | 0.002352 | 246.0673 | 0.26486 | 8.4018 | 0.00229 | 0.003904 | 0.003812 | 4.3e-05 | 0.000938 | 0.30807 | 1.3118 | 2.5474 | 100.2 | 20.85 | 37.075 | 80.4121 | 76.7727 | 78.2655 | 0.009128 | 1.0046 | 0.023821 | 23.7111 | 0.048827 | 0.15756 | 0.004886 | 0.34532 | 0.34532 | 0.01857
2019-03-29 14:03 | 1553882588.5473 | 601 | 0.20857 | 0.001141 | 6.9e-05 | 101.2799 | 100.9424 | 120.4443 | 21.1205 | 0.00247 | 0.002316 | 246.0422 | 0.26446 | 8.4011 | 0.00225 | 0.003668 | 0.003603 | 4.8e-05 | 0.000935 | 0.28119 | 1.3666 | 2.5474 | 100.2 | 20.85 | 37.072 | 81.1008 | 75.8157 | 77.4108 | 0.009127 | 1.0071 | 0.023856 | 23.6886 | 0.029851 | 0.15733 | 0.004834 | 0.34209 | 0.34209 | 0.018718
2019-03-29 14:03 | 1553882589.5464 | 602 | 0.20856 | 0.001145 | 7e-05 | 101.3394 | 100.9025 | 120.4079 | 21.1209 | 0.002494 | 0.002269 | 246.0317 | 0.26478 | 8.4012 | 0.002261 | 0.003049 | 0.003074 | 4.5e-05 | 0.000933 | 0.30055 | 1.3235 | 2.5473 | 100.2 | 20.85 | 37.065 | 81.6759 | 76.079 | 76.2808 | 0.009125 | 1.0103 | 0.023891 | 23.6468 | 0.038649 | 0.15752 | 0.004865 | 0.34478 | 0.34478 | 0.018595
2019-03-29 14:03 | 1553882590.5473 | 603 | 0.20857 | 0.001146 | 7e-05 | 101.2694 | 100.6585 | 120.5912 | 21.1213 | 0.002502 | 0.002268 | 246.0154 | 0.26362 | 8.4008 | 0.00227 | 0.003964 | 0.00385 | 4.6e-05 | 0.000936 | 0.31509 | 1.3674 | 2.5476 | 100.2 | 20.85 | 37.06 | 81.8758 | 76.3054 | 76.2461 | 0.009124 | 1.0104 | 0.023911 | 23.6636 | 0.054122 | 0.15683 | 0.004833 | 0.34278 | 0.34278 | 0.018621
2019-03-29 14:03 | 1553882591.5464 | 604 | 0.20858 | 0.001147 | 7.1e-05 | 101.2852 | 100.6755 | 120.8581 | 21.1203 | 0.00249 | 0.002306 | 246.0025 | 0.26426 | 8.4021 | 0.002259 | 0.003856 | 0.003732 | 4.8e-05 | 0.000931 | 0.28672 | 1.319 | 2.5474 | 100.2 | 20.85 | 37.053 | 81.5731 | 76.0447 | 77.1588 | 0.009123 | 1.0078 | 0.023906 | 23.7207 | 0.054007 | 0.15721 | 0.00478 | 0.33888 | 0.33888 | 0.018881
2019-03-29 14:03 | 1553882592.5474 | 605 | 0.2086 | 0.001148 | 7.1e-05 | 101.2023 | 100.7506 | 121.1191 | 21.1205 | 0.002469 | 0.002307 | 245.9985 | 0.26335 | 8.4018 | 0.002208 | 0.003675 | 0.003641 | 4.2e-05 | 0.000931 | 0.27793 | 1.3095 | 2.5473 | 100.2 | 20.85 | 37.05 | 81.0813 | 74.8065 | 77.1893 | 0.009122 | 1.0077 | 0.023931 | 23.7473 | 0.040014 | 0.15667 | 0.004655 | 0.33028 | 0.33028 | 0.019306
2019-03-29 14:03 | 1553882593.5474 | 606 | 0.20857 | 0.001148 | 7.1e-05 | 101.1592 | 100.5612 | 120.9567 | 21.1203 | 0.002459 | 0.002274 | 245.9993 | 0.26455 | 8.4005 | 0.002243 | 0.003401 | 0.003357 | 4.7e-05 | 0.000925 | 0.31725 | 1.3058 | 2.5472 | 100.2 | 20.85 | 37.045 | 80.8408 | 75.6527 | 76.4049 | 0.009121 | 1.01 | 0.023942 | 23.7048 | 0.053035 | 0.15738 | 0.004811 | 0.34159 | 0.34159 | 0.018752
2019-03-29 14:03 | 1553882594.5479 | 607 | 0.20857 | 0.001148 | 7.1e-05 | 101.1643 | 100.525 | 121.2733 | 21.1199 | 0.002473 | 0.002312 | 245.957 | 0.26459 | 8.4019 | 0.00228 | 0.003618 | 0.003588 | 4.4e-05 | 0.000929 | 0.28766 | 1.3649 | 2.5474 | 100.2 | 20.85 | 37.038 | 81.1774 | 76.5378 | 77.3064 | 0.009119 | 1.0074 | 0.023942 | 23.7664 | 0.056704 | 0.15741 | 0.004802 | 0.34095 | 0.34095 | 0.01879
2019-03-29 14:03 | 1553882595.5465 | 608 | 0.20858 | 0.001152 | 7.1e-05 | 101.0757 | 100.3858 | 120.9884 | 21.1215 | 0.002462 | 0.002312 | 245.9007 | 0.26366 | 8.4029 | 0.002351 | 0.003866 | 0.003724 | 4.5e-05 | 0.000932 | 0.2959 | 1.3541 | 2.5474 | 100.2 | 20.85 | 37.031 | 80.9239 | 78.2489 | 77.3111 | 0.009117 | 1.0074 | 0.023914 | 23.7386 | 0.061265 | 0.15685 | 0.004777 | 0.33877 | 0.33877 | 0.018844
2019-03-29 14:03 | 1553882596.5476 | 609 | 0.20854 | 0.001152 | 7.3e-05 | 101.114 | 100.4636 | 120.8701 | 21.1208 | 0.002406 | 0.002309 | 245.8723 | 0.2643 | 8.4029 | 0.002358 | 0.00398 | 0.003951 | 4.9e-05 | 0.000931 | 0.30381 | 1.3515 | 2.5475 | 100.2 | 20.85 | 37.031 | 79.5781 | 78.4142 | 77.2378 | 0.009117 | 1.0076 | 0.023904 | 23.7245 | 0.057724 | 0.15723 | 0.004979 | 0.35306 | 0.35306 | 0.018126
2019-03-29 14:03 | 1553882597.5466 | 610 | 0.20862 | 0.001159 | 7.5e-05 | 101.0147 | 100.3588 | 120.8109 | 21.12 | 0.002422 | 0.002274 | 245.8464 | 0.26455 | 8.4008 | 0.002269 | 0.003528 | 0.003506 | 4.8e-05 | 0.000935 | 0.31349 | 1.3704 | 2.5474 | 100.2 | 20.85 | 37.024 | 79.9477 | 76.2784 | 76.3827 | 0.009115 | 1.01 | 0.023928 | 23.6898 | 0.058262 | 0.15738 | 0.004513 | 0.31977 | 0.31977 | 0.020031
2019-03-29 14:03 | 1553882598.5477 | 611 | 0.20851 | 0.001159 | 7.8e-05 | 101.0338 | 100.2728 | 120.6041 | 21.1201 | 0.002475 | 0.002275 | 245.799 | 0.26383 | 8.402 | 0.002344 | 0.003934 | 0.003811 | 4.6e-05 | 0.000937 | 0.30071 | 1.3962 | 2.5473 | 100.2 | 20.85 | 37.02 | 81.2196 | 78.0862 | 76.4116 | 0.009114 | 1.01 | 0.023906 | 23.6705 | 0.067627 | 0.15696 | 0.005137 | 0.36407 | 0.36407 | 0.017546
2019-03-29 14:03 | 1553882599.5476 | 612 | 0.20842 | 0.001165 | 8.1e-05 | 100.909 | 100.2612 | 120.3063 | 21.1197 | 0.002523 | 0.002322 | 245.7454 | 0.26422 | 8.4039 | 0.00228 | 0.003472 | 0.003374 | 4.7e-05 | 0.000937 | 0.30287 | 1.3558 | 2.5473 | 100.2 | 20.85 | 37.013 | 82.3742 | 76.5268 | 77.5354 | 0.009113 | 1.0067 | 0.023839 | 23.6792 | 0.057604 | 0.15719 | 0.005694 | 0.4027 | 0.4027 | 0.015887
2019-03-29 14:03 | 1553882600.5471 | 613 | 0.20825 | 0.001176 | 8.5e-05 | 100.9219 | 100.3373 | 119.9202 | 21.1206 | 0.002446 | 0.00236 | 245.6726 | 0.2641 | 8.4022 | 0.002291 | 0.002939 | 0.00289 | 4.8e-05 | 0.000936 | 0.30901 | 1.3463 | 2.5474 | 100.2 | 20.85 | 37.01 | 80.5191 | 76.7934 | 78.4609 | 0.009112 | 1.0041 | 0.023769 | 23.6723 | 0.051958 | 0.15711 | 0.006688 | 0.47219 | 0.47219 | 0.013542
2019-03-29 14:03 | 1553882601.5467 | 614 | 0.20802 | 0.001203 | 8.8e-05 | 100.9222 | 100.4219 | 119.7455 | 21.1196 | 0.002539 | 0.002297 | 245.6004 | 0.26454 | 8.4035 | 0.00231 | 0.003713 | 0.003545 | 4.6e-05 | 0.000937 | 0.31365 | 1.3227 | 2.5475 | 100.2 | 20.85 | 37.004 | 82.7727 | 77.2585 | 76.9577 | 0.00911 | 1.0084 | 0.023803 | 23.6045 | 0.044443 | 0.15738 | 0.008056 | 0.5703 | 0.5703 | 0.011231
2019-03-29 14:03 | 1553882602.5478 | 615 | 0.2079 | 0.001242 | 9.2e-05 | 100.8898 | 100.4721 | 119.1134 | 21.1204 | 0.00244 | 0.002334 | 245.5331 | 0.26483 | 8.4012 | 0.002297 | 0.003171 | 0.003124 | 4.8e-05 | 0.000939 | 0.30254 | 1.3096 | 2.5474 | 100.2 | 20.85 | 37 | 80.3792 | 76.9429 | 77.8355 | 0.009109 | 1.0059 | 0.02371 | 23.5716 | 0.03711 | 0.15755 | 0.008692 | 0.61314 | 0.61314 | 0.010458
2019-03-29 14:03 | 1553882603.5468 | 616 | 0.20782 | 0.001292 | 9.7e-05 | 100.8165 | 100.4179 | 118.3668 | 21.1199 | 0.002436 | 0.00232 | 245.4598 | 0.26446 | 8.4009 | 0.002224 | 0.003819 | 0.003725 | 4.6e-05 | 0.00094 | 0.28296 | 1.3395 | 2.5474 | 100.2 | 20.85 | 36.998 | 80.2953 | 75.2025 | 77.4981 | 0.009109 | 1.0068 | 0.023647 | 23.4863 | 0.035438 | 0.15733 | 0.009104 | 0.64042 | 0.64042 | 0.009998
2019-03-29 14:03 | 1553882604.5458 | 617 | 0.20755 | 0.001352 | 0.000103 | 100.8042 | 100.5025 | 117.0406 | 21.1198 | 0.002475 | 0.002376 | 245.381 | 0.2643 | 8.4036 | 0.002337 | 0.003402 | 0.003374 | 4.7e-05 | 0.000942 | 0.27848 | 1.3392 | 2.5474 | 100.2 | 20.85 | 36.994 | 81.2186 | 77.9034 | 78.8553 | 0.009108 | 1.003 | 0.023469 | 23.3995 | 0.026813 | 0.15723 | 0.010651 | 0.74407 | 0.74407 | 0.0086
2019-03-29 14:03 | 1553882605.5464 | 618 | 0.20727 | 0.001426 | 0.000111 | 100.7417 | 100.583 | 115.5644 | 21.1192 | 0.002501 | 0.002415 | 245.3076 | 0.26387 | 8.4017 | 0.002227 | 0.003974 | 0.003902 | 4e-05 | 0.00094 | 0.29822 | 1.3219 | 1.8597 | 100.2 | 20.85 | 36.995 | 81.8387 | 75.272 | 79.7816 | 0.009108 | 1.0003 | 0.02329 | 23.282 | 0.014095 | 0.15697 | 0.012227 | 0.84791 | 0.84791 | 0.007535
2019-03-29 14:03 | 1553882606.5459 | 619 | 0.20689 | 0.001512 | 0.000123 | 100.7845 | 100.4103 | 114.2582 | 21.1194 | 0.002597 | 0.00249 | 245.2085 | 0.26374 | 8.4017 | 0.002304 | 0.004599 | 0.004545 | 4.3e-05 | 0.00094 | 0.26825 | 1.3572 | 2.5474 | 100.2 | 20.86 | 36.999 | 84.1496 | 77.1048 | 81.5855 | 0.009115 | 0.99524 | 0.023099 | 23.2092 | 0.033265 | 0.1569 | 0.01439 | 0.98997 | 0.98997 | 0.00645
2019-03-29 14:03 | 1553882607.5459 | 620 | 0.20638 | 0.001618 | 0.000136 | 100.6824 | 100.4407 | 112.7587 | 21.1192 | 0.002625 | 0.002488 | 245.1076 | 0.26449 | 8.4038 | 0.002317 | 0.004094 | 0.004098 | 4.2e-05 | 0.000929 | 0.25658 | 1.3301 | 2.5474 | 100.2 | 20.86 | 36.999 | 84.835 | 77.436 | 81.5245 | 0.009115 | 0.99541 | 0.022949 | 23.0544 | 0.021496 | 0.15735 | 0.017324 | 1.1844 | 1.1844 | 0.005407
2019-03-29 14:03 | 1553882608.5463 | 621 | 0.20539 | 0.001754 | 0.000151 | 100.6779 | 99.7755 | 111.5156 | 21.1192 | 0.002635 | 0.002583 | 244.9945 | 0.26467 | 8.4026 | 0.002328 | 0.004038 | 0.003906 | 4.4e-05 | 0.000928 | 0.27383 | 1.3286 | 2.5474 | 100.2 | 20.85 | 37 | 85.0786 | 77.687 | 83.8262 | 0.009109 | 0.98899 | 0.022748 | 23.0012 | 0.080529 | 0.15745 | 0.023045 | 1.5626 | 1.5626 | 0.004101
2019-03-29 14:03 | 1553882609.5459 | 622 | 0.20445 | 0.00215 | 0.000163 | 100.6399 | 99.2971 | 110.3725 | 21.119 | 0.002747 | 0.002629 | 244.8547 | 0.26402 | 8.4017 | 0.002398 | 0.00323 | 0.003247 | 4.1e-05 | 0.000926 | 0.25962 | 1.3839 | 2.5474 | 100.2 | 20.85 | 37.006 | 87.7597 | 79.3839 | 84.937 | 0.009111 | 0.98592 | 0.022596 | 22.9186 | 0.12015 | 0.15707 | 0.028195 | 1.8992 | 1.8992 | 0.003366
2019-03-29 14:03 | 1553882610.5459 | 623 | 0.20323 | 0.002736 | 0.000173 | 100.6141 | 99.0377 | 109.2217 | 21.1187 | 0.002747 | 0.002673 | 244.7287 | 0.26407 | 8.4007 | 0.00255 | 0.003856 | 0.003722 | 4.3e-05 | 0.000923 | 0.2844 | 1.3252 | 2.5474 | 100.2 | 20.86 | 37.01 | 87.7736 | 83.0175 | 85.9872 | 0.009118 | 0.98304 | 0.022445 | 22.8322 | 0.14126 | 0.1571 | 0.034749 | 2.3249 | 2.3249 | 0.00275
2019-03-29 14:04 | 1553882611.546 | 624 | 0.20171 | 0.003441 | 0.000181 | 100.554 | 98.7742 | 107.7868 | 21.1187 | 0.00282 | 0.002787 | 244.5869 | 0.26415 | 8.4014 | 0.002643 | 0.003809 | 0.003741 | 4.2e-05 | 0.00092 | 0.29048 | 1.317 | 2.5473 | 100.2 | 20.86 | 37.013 | 89.5319 | 85.2631 | 88.7334 | 0.009118 | 0.97558 | 0.022212 | 22.7683 | 0.15973 | 0.15714 | 0.04294 | 2.8425 | 2.8425 | 0.00225
2019-03-29 14:04 | 1553882612.546 | 625 | 0.20071 | 0.004255 | 0.000188 | 100.5106 | 98.6404 | 106.6881 | 21.1185 | 0.002863 | 0.00279 | 244.4531 | 0.26445 | 8.4026 | 0.002733 | 0.003693 | 0.003635 | 4.4e-05 | 0.00092 | 0.28119 | 1.3223 | 2.5474 | 100.2 | 20.85 | 37.017 | 90.5584 | 87.4336 | 88.7941 | 0.009114 | 0.97541 | 0.022097 | 22.6539 | 0.168 | 0.15732 | 0.047891 | 3.153 | 3.153 | 0.002031
2019-03-29 14:04 | 1553882613.546 | 626 | 0.2001 | 0.005103 | 0.000197 | 100.513 | 98.5926 | 105.8088 | 21.119 | 0.002858 | 0.002818 | 244.2893 | 0.26521 | 8.4025 | 0.002782 | 0.003949 | 0.003886 | 4.3e-05 | 0.000925 | 0.27721 | 1.338 | 2.5473 | 100.2 | 20.85 | 37.026 | 90.444 | 88.6104 | 89.4806 | 0.009116 | 0.97357 | 0.021985 | 22.5817 | 0.17254 | 0.15777 | 0.050499 | 3.3071 | 3.3071 | 0.001942
2019-03-29 14:04 | 1553882614.5471 | 627 | 0.19955 | 0.005871 | 0.000205 | 100.4878 | 99.0001 | 105.2316 | 21.1185 | 0.00292 | 0.002888 | 244.1184 | 0.26407 | 8.4015 | 0.002947 | 0.004168 | 0.004124 | 4.6e-05 | 0.000923 | 0.28854 | 1.3511 | 2.5473 | 100.2 | 20.85 | 37.032 | 91.9408 | 92.5858 | 91.1501 | 0.009117 | 0.96911 | 0.021874 | 22.5718 | 0.13341 | 0.1571 | 0.05282 | 3.4409 | 3.4409 | 0.001858
2019-03-29 14:04 | 1553882615.5471 | 628 | 0.19929 | 0.006429 | 0.000212 | 100.5208 | 98.7288 | 105.1403 | 21.1178 | 0.003036 | 0.002927 | 243.9692 | 0.26398 | 8.4027 | 0.002865 | 0.0042 | 0.004171 | 4.1e-05 | 0.000927 | 0.29247 | 1.3172 | 2.5473 | 100.2 | 20.85 | 37.036 | 94.7173 | 90.6071 | 92.1004 | 0.009118 | 0.96659 | 0.021837 | 22.5914 | 0.16089 | 0.15704 | 0.053697 | 3.4914 | 3.4914 | 0.001831
2019-03-29 14:04 | 1553882616.5467 | 629 | 0.19889 | 0.006934 | 0.000219 | 100.4693 | 98.7868 | 104.8394 | 21.1178 | 0.003009 | 0.002938 | 243.8053 | 0.26467 | 8.401 | 0.00298 | 0.004531 | 0.004395 | 4.5e-05 | 0.000923 | 0.29607 | 1.3375 | 2.5473 | 100.2 | 20.85 | 37.047 | 94.0801 | 93.3675 | 92.3529 | 0.009121 | 0.96592 | 0.021798 | 22.5669 | 0.15105 | 0.15745 | 0.055452 | 3.5984 | 3.5984 | 0.001781
2019-03-29 14:04 | 1553882617.5473 | 630 | 0.19864 | 0.007257 | 0.000224 | 100.445 | 98.8445 | 104.4244 | 21.118 | 0.003087 | 0.002971 | 243.6328 | 0.26453 | 8.4002 | 0.003041 | 0.004976 | 0.004837 | 4.4e-05 | 0.000921 | 0.29087 | 1.3446 | 2.5474 | 100.2 | 20.85 | 37.051 | 95.968 | 94.8553 | 93.1712 | 0.009122 | 0.96376 | 0.02173 | 22.5474 | 0.14367 | 0.15737 | 0.05656 | 3.6585 | 3.6585 | 0.001751
2019-03-29 14:04 | 1553882618.5472 | 631 | 0.19842 | 0.007584 | 0.000228 | 100.3669 | 98.6138 | 104.1479 | 21.1183 | 0.003124 | 0.003016 | 243.4679 | 0.26353 | 8.403 | 0.003119 | 0.00442 | 0.004401 | 4.2e-05 | 0.000918 | 0.2589 | 1.3348 | 2.5475 | 100.2 | 20.85 | 37.056 | 96.8596 | 96.721 | 94.2545 | 0.009123 | 0.96092 | 0.021669 | 22.5508 | 0.15761 | 0.15677 | 0.057507 | 3.7089 | 3.7089 | 0.00172
2019-03-29 14:04 | 1553882619.5463 | 632 | 0.19807 | 0.007818 | 0.00023 | 100.2834 | 98.479 | 103.5857 | 21.1169 | 0.003206 | 0.003072 | 243.2926 | 0.2642 | 8.401 | 0.003086 | 0.004981 | 0.004862 | 4.3e-05 | 0.000921 | 0.27544 | 1.3111 | 2.5474 | 100.2 | 20.85 | 37.064 | 98.8321 | 95.9404 | 95.5919 | 0.009125 | 0.95743 | 0.021572 | 22.5307 | 0.1624 | 0.15718 | 0.059273 | 3.8052 | 3.8052 | 0.001681
2019-03-29 14:04 | 1553882620.5453 | 633 | 0.19757 | 0.008059 | 0.000232 | 100.2388 | 98.3514 | 102.9337 | 21.1168 | 0.003256 | 0.003153 | 243.1136 | 0.26468 | 8.4023 | 0.003087 | 0.004858 | 0.004835 | 4.3e-05 | 0.000925 | 0.28943 | 1.3056 | 2.5473 | 100.2 | 20.85 | 37.069 | 100.0492 | 95.9683 | 97.5509 | 0.009126 | 0.95237 | 0.021447 | 22.5193 | 0.17003 | 0.15746 | 0.061966 | 3.9544 | 3.9544 | 0.001621
2019-03-29 14:04 | 1553882621.5456 | 634 | 0.1972 | 0.008288 | 0.000231 | 100.1496 | 98.193 | 102.5356 | 21.1171 | 0.003315 | 0.003168 | 242.941 | 0.26441 | 8.4024 | 0.00314 | 0.004372 | 0.004277 | 4.4e-05 | 0.000923 | 0.31725 | 1.3378 | 2.5473 | 100.2 | 20.85 | 37.076 | 101.4742 | 97.2355 | 97.9182 | 0.009128 | 0.95143 | 0.021395 | 22.4868 | 0.17648 | 0.15729 | 0.063917 | 4.0686 | 4.0686 | 0.001573
2019-03-29 14:04 | 1553882622.5453 | 635 | 0.19687 | 0.008524 | 0.000231 | 100.125 | 98.0838 | 102.2645 | 21.1176 | 0.00334 | 0.003241 | 242.7481 | 0.26424 | 8.4021 | 0.003202 | 0.004736 | 0.004665 | 4.6e-05 | 0.000927 | 0.31045 | 1.3045 | 2.5474 | 100.2 | 20.85 | 37.08 | 102.0816 | 98.7332 | 99.6766 | 0.009129 | 0.94694 | 0.021316 | 22.5102 | 0.18424 | 0.15719 | 0.065552 | 4.1569 | 4.1569 | 0.001539
2019-03-29 14:04 | 1553882623.5456 | 636 | 0.19653 | 0.008758 | 0.000231 | 100.1038 | 98.1807 | 102.1427 | 21.117 | 0.003396 | 0.003305 | 242.5515 | 0.26403 | 8.3988 | 0.00326 | 0.004752 | 0.004604 | 4.5e-05 | 0.000928 | 0.30613 | 1.367 | 2.5474 | 100.2 | 20.85 | 37.084 | 103.4354 | 100.1476 | 101.2166 | 0.00913 | 0.94305 | 0.021259 | 22.5432 | 0.1735 | 0.15707 | 0.067275 | 4.2544 | 4.2544 | 0.001503
2019-03-29 14:04 | 1553882624.5459 | 637 | 0.19653 | 0.008988 | 0.00023 | 100.0263 | 97.9811 | 102.0908 | 21.1165 | 0.003484 | 0.003296 | 242.3656 | 0.26459 | 8.4019 | 0.003346 | 0.005616 | 0.00549 | 4.7e-05 | 0.000928 | 0.2912 | 1.3792 | 2.5475 | 100.2 | 20.85 | 37.09 | 105.5635 | 102.2107 | 101.0188 | 0.009132 | 0.94355 | 0.02126 | 22.5315 | 0.18478 | 0.1574 | 0.067031 | 4.2391 | 4.2391 | 0.001511
2019-03-29 14:04 | 1553882625.5459 | 638 | 0.1962 | 0.009218 | 0.000228 | 100.005 | 97.7384 | 102.0863 | 21.1179 | 0.00349 | 0.003376 | 242.1768 | 0.26466 | 8.4024 | 0.003389 | 0.004969 | 0.004962 | 4.4e-05 | 0.000928 | 0.2912 | 1.3389 | 2.5473 | 100.2 | 20.85 | 37.093 | 105.7129 | 103.265 | 102.9562 | 0.009132 | 0.93868 | 0.021204 | 22.5893 | 0.20506 | 0.15745 | 0.068747 | 4.336 | 4.336 | 0.001478
2019-03-29 14:04 | 1553882626.5455 | 639 | 0.196 | 0.009441 | 0.000226 | 99.9392 | 97.6128 | 102.1082 | 21.1166 | 0.003539 | 0.003338 | 241.9914 | 0.26453 | 8.3999 | 0.003435 | 0.004659 | 0.004623 | 4.3e-05 | 0.000937 | 0.32189 | 1.3633 | 2.5475 | 100.2 | 20.85 | 37.101 | 106.9032 | 104.3659 | 102.0359 | 0.009134 | 0.94099 | 0.021233 | 22.5641 | 0.21067 | 0.15737 | 0.069631 | 4.3974 | 4.3974 | 0.001457
2019-03-29 14:04 | 1553882627.546 | 640 | 0.19578 | 0.009657 | 0.000224 | 99.9763 | 97.6254 | 102.3218 | 21.1169 | 0.00356 | 0.003431 | 241.8086 | 0.26496 | 8.4022 | 0.003455 | 0.004585 | 0.004479 | 4.6e-05 | 0.00094 | 0.31222 | 1.3307 | 2.5475 | 100.2 | 20.85 | 37.103 | 107.3957 | 104.8693 | 104.267 | 0.009135 | 0.93542 | 0.021192 | 22.6547 | 0.21283 | 0.15762 | 0.070694 | 4.4558 | 4.4558 | 0.00144
2019-03-29 14:04 | 1553882628.547 | 641 | 0.19558 | 0.009869 | 0.000222 | 99.8464 | 97.3193 | 102.8097 | 21.1175 | 0.003616 | 0.003423 | 241.6172 | 0.26454 | 8.4027 | 0.003553 | 0.005479 | 0.005407 | 4.6e-05 | 0.000944 | 0.29336 | 1.3509 | 2.5474 | 100.2 | 20.85 | 37.105 | 108.7727 | 107.2364 | 104.0733 | 0.009135 | 0.93591 | 0.021248 | 22.7028 | 0.2293 | 0.15737 | 0.07165 | 4.5278 | 4.5278 | 0.001415
2019-03-29 14:04 | 1553882629.5467 | 642 | 0.19553 | 0.01007 | 0.000221 | 99.8446 | 97.578 | 103.4562 | 21.1166 | 0.003709 | 0.003509 | 241.4523 | 0.2643 | 8.4029 | 0.003589 | 0.004997 | 0.004947 | 4.8e-05 | 0.000945 | 0.28943 | 1.4014 | 2.5474 | 100.2 | 20.85 | 37.109 | 111.0303 | 108.1004 | 106.1554 | 0.009136 | 0.93077 | 0.021256 | 22.8369 | 0.20539 | 0.15723 | 0.071681 | 4.5315 | 4.5315 | 0.001412
2019-03-29 14:04 | 1553882630.5456 | 643 | 0.19548 | 0.010256 | 0.00022 | 99.7916 | 97.1944 | 104.2319 | 21.1171 | 0.003741 | 0.003523 | 241.3126 | 0.26473 | 8.4032 | 0.0036 | 0.00509 | 0.005034 | 4.8e-05 | 0.000944 | 0.27489 | 1.3661 | 2.5473 | 100.2 | 20.85 | 37.111 | 111.7964 | 108.3734 | 106.5053 | 0.009137 | 0.92991 | 0.021326 | 22.9329 | 0.23588 | 0.15749 | 0.071796 | 4.5537 | 4.5537 | 0.001408
2019-03-29 14:04 | 1553882631.5466 | 644 | 0.19544 | 0.010429 | 0.00022 | 99.7658 | 97.4995 | 104.781 | 21.1162 | 0.003694 | 0.003521 | 241.1543 | 0.26422 | 8.3989 | 0.003643 | 0.004878 | 0.004745 | 4.9e-05 | 0.000942 | 0.28479 | 1.3223 | 2.5474 | 100.2 | 20.85 | 37.112 | 110.6639 | 109.4201 | 106.4486 | 0.009137 | 0.93005 | 0.021383 | 22.9915 | 0.20553 | 0.15718 | 0.071839 | 4.5687 | 4.5687 | 0.0014
2019-03-29 14:04 | 1553882632.5468 | 645 | 0.19536 | 0.010583 | 0.000221 | 99.7446 | 96.9876 | 105.1161 | 21.1161 | 0.003702 | 0.003507 | 240.9891 | 0.26428 | 8.4043 | 0.003632 | 0.004613 | 0.004529 | 4.9e-05 | 0.000942 | 0.30093 | 1.3161 | 2.5474 | 100.2 | 20.85 | 37.116 | 110.8499 | 109.1456 | 106.1234 | 0.009138 | 0.93085 | 0.021427 | 23.0184 | 0.25072 | 0.15722 | 0.072118 | 4.5956 | 4.5956 | 0.001392
2019-03-29 14:04 | 1553882633.5457 | 646 | 0.19584 | 0.01071 | 0.000223 | 99.6075 | 96.9135 | 105.1967 | 21.1166 | 0.003778 | 0.003515 | 240.8404 | 0.26403 | 8.4005 | 0.003659 | 0.004721 | 0.004687 | 5.1e-05 | 0.00094 | 0.30558 | 1.3194 | 2.5473 | 100.2 | 20.85 | 37.118 | 112.7057 | 109.8093 | 106.3245 | 0.009139 | 0.93035 | 0.021429 | 23.0333 | 0.24525 | 0.15707 | 0.069086 | 4.4037 | 4.4037 | 0.001452
2019-03-29 14:04 | 1553882634.5458 | 647 | 0.19613 | 0.010802 | 0.000225 | 99.6446 | 96.8971 | 105.7527 | 21.1167 | 0.003812 | 0.003604 | 240.6983 | 0.26409 | 8.4003 | 0.00366 | 0.004841 | 0.004794 | 5.1e-05 | 0.000938 | 0.29551 | 1.3156 | 2.5474 | 100.2 | 20.85 | 37.122 | 113.5355 | 109.8347 | 108.4685 | 0.009139 | 0.92513 | 0.021425 | 23.1592 | 0.25009 | 0.15711 | 0.06731 | 4.2901 | 4.2901 | 0.00149
2019-03-29 14:04 | 1553882635.5454 | 648 | 0.19636 | 0.010863 | 0.000228 | 99.5767 | 96.7722 | 106.3071 | 21.1162 | 0.003847 | 0.003565 | 240.5591 | 0.2638 | 8.4034 | 0.003668 | 0.005184 | 0.005017 | 4.9e-05 | 0.000938 | 0.31614 | 1.3347 | 2.5474 | 100.2 | 20.85 | 37.124 | 114.3845 | 110.0177 | 107.5164 | 0.00914 | 0.92744 | 0.021508 | 23.1909 | 0.25553 | 0.15693 | 0.06583 | 4.2122 | 4.2122 | 0.001516
2019-03-29 14:04 | 1553882636.5461 | 649 | 0.19652 | 0.010877 | 0.000232 | 99.4958 | 96.5876 | 106.6245 | 21.1164 | 0.003839 | 0.003511 | 240.4221 | 0.26426 | 8.4013 | 0.00368 | 0.005367 | 0.005273 | 5.2e-05 | 0.000938 | 0.28263 | 1.3434 | 2.5474 | 100.2 | 20.85 | 37.126 | 114.1941 | 110.3264 | 106.2159 | 0.00914 | 0.93062 | 0.021577 | 23.1858 | 0.26535 | 0.15721 | 0.064899 | 4.166 | 4.166 | 0.001536
2019-03-29 14:04 | 1553882637.5454 | 650 | 0.19688 | 0.01085 | 0.000235 | 99.4641 | 96.548 | 107.0395 | 21.1168 | 0.003844 | 0.00351 | 240.2929 | 0.26392 | 8.4018 | 0.003708 | 0.004148 | 0.004127 | 5.3e-05 | 0.00094 | 0.3131 | 1.3369 | 2.5475 | 100.2 | 20.85 | 37.129 | 114.322 | 110.995 | 106.1925 | 0.009141 | 0.93068 | 0.02162 | 23.2301 | 0.26616 | 0.15701 | 0.062764 | 4.0372 | 4.0372 | 0.001583
2019-03-29 14:04 | 1553882638.5448 | 651 | 0.19722 | 0.010785 | 0.000236 | 99.4114 | 96.5497 | 107.4894 | 21.1168 | 0.003872 | 0.003609 | 240.1661 | 0.26374 | 8.401 | 0.003641 | 0.004295 | 0.004226 | 5.3e-05 | 0.00094 | 0.28871 | 1.3003 | 2.5475 | 100.2 | 20.85 | 37.13 | 115.0013 | 109.3624 | 108.5922 | 0.009141 | 0.92483 | 0.021597 | 23.3524 | 0.26126 | 0.1569 | 0.060815 | 3.908 | 3.908 | 0.001634
2019-03-29 14:04 | 1553882639.546 | 652 | 0.1976 | 0.010691 | 0.000235 | 99.3836 | 96.3671 | 107.9915 | 21.1167 | 0.003881 | 0.003545 | 240.0423 | 0.26354 | 8.4046 | 0.003691 | 0.005186 | 0.00508 | 5.2e-05 | 0.000945 | 0.29806 | 1.3341 | 2.5473 | 100.2 | 20.85 | 37.136 | 115.2098 | 110.5973 | 107.0513 | 0.009143 | 0.92858 | 0.021691 | 23.3596 | 0.27569 | 0.15678 | 0.058663 | 3.7867 | 3.7867 | 0.001685
2019-03-29 14:04 | 1553882640.5456 | 653 | 0.19827 | 0.010557 | 0.000233 | 99.2531 | 96.3395 | 108.4591 | 21.1157 | 0.003839 | 0.003572 | 239.8867 | 0.26372 | 8.4005 | 0.003639 | 0.005087 | 0.004965 | 5.2e-05 | 0.000938 | 0.25752 | 1.3666 | 2.5473 | 100.2 | 20.85 | 37.137 | 114.1832 | 109.3289 | 107.704 | 0.009143 | 0.92698 | 0.021719 | 23.4302 | 0.2665 | 0.15688 | 0.054852 | 3.5459 | 3.5459 | 0.001801
2019-03-29 14:04 | 1553882641.5451 | 654 | 0.19848 | 0.010387 | 0.000228 | 99.2788 | 96.5023 | 108.9602 | 21.1172 | 0.003776 | 0.003544 | 239.7529 | 0.26396 | 8.4022 | 0.003645 | 0.005032 | 0.004971 | 5.1e-05 | 0.000936 | 0.3095 | 1.3355 | 2.5474 | 100.2 | 20.85 | 37.14 | 112.6676 | 109.4604 | 107.0094 | 0.009144 | 0.92868 | 0.021789 | 23.4628 | 0.25372 | 0.15703 | 0.053814 | 3.4905 | 3.4905 | 0.001831
2019-03-29 14:04 | 1553882642.5464 | 655 | 0.19906 | 0.010174 | 0.000222 | 99.2226 | 96.6337 | 109.6944 | 21.1166 | 0.003887 | 0.00355 | 239.6204 | 0.2632 | 8.4029 | 0.003631 | 0.005749 | 0.0057 | 5e-05 | 0.00093 | 0.3011 | 1.351 | 2.5474 | 100.2 | 20.85 | 37.144 | 115.3641 | 109.1226 | 107.152 | 0.009145 | 0.92833 | 0.021859 | 23.5462 | 0.23648 | 0.15658 | 0.050602 | 3.2932 | 3.2932 | 0.001935
2019-03-29 14:04 | 1553882643.5454 | 656 | 0.1995 | 0.009928 | 0.000214 | 99.1301 | 96.7104 | 110.5512 | 21.117 | 0.003808 | 0.003556 | 239.5178 | 0.26405 | 8.4033 | 0.003562 | 0.004889 | 0.004831 | 4.9e-05 | 0.000934 | 0.31133 | 1.3192 | 2.5474 | 100.2 | 20.85 | 37.145 | 113.4249 | 107.4587 | 107.3004 | 0.009145 | 0.92797 | 0.02194 | 23.6426 | 0.22103 | 0.15709 | 0.048252 | 3.1526 | 3.1526 | 0.002028
2019-03-29 14:04 | 1553882644.5457 | 657 | 0.19996 | 0.009531 | 0.000208 | 99.1241 | 96.9094 | 110.7425 | 21.1172 | 0.003856 | 0.00354 | 239.4298 | 0.26432 | 8.3992 | 0.003565 | 0.005093 | 0.004964 | 5.1e-05 | 0.00093 | 0.33107 | 1.3341 | 2.5473 | 100.2 | 20.84 | 37.148 | 114.6148 | 107.5227 | 106.9149 | 0.00914 | 0.92891 | 0.02197 | 23.651 | 0.20211 | 0.15724 | 0.046026 | 3.0118 | 3.0118 | 0.002125
2019-03-29 14:04 | 1553882645.5463 | 658 | 0.20046 | 0.009214 | 0.0002 | 99.0549 | 96.9449 | 110.9791 | 21.1165 | 0.003925 | 0.00352 | 239.3401 | 0.26315 | 8.4029 | 0.003594 | 0.00484 | 0.004704 | 5.2e-05 | 0.000931 | 0.3126 | 1.3701 | 2.5472 | 100.2 | 20.84 | 37.15 | 116.2827 | 108.2296 | 106.43 | 0.009141 | 0.9301 | 0.022007 | 23.6612 | 0.19259 | 0.15655 | 0.043395 | 2.845 | 2.845 | 0.00224
2019-03-29 14:04 | 1553882646.5456 | 659 | 0.20087 | 0.008788 | 0.000191 | 99.1073 | 97.1971 | 111.4049 | 21.1162 | 0.003878 | 0.003569 | 239.2439 | 0.26376 | 8.4018 | 0.003583 | 0.005293 | 0.005222 | 4.9e-05 | 0.000933 | 0.30093 | 1.3297 | 2.5474 | 100.2 | 20.84 | 37.152 | 115.1318 | 107.974 | 107.6115 | 0.009141 | 0.92721 | 0.022015 | 23.7434 | 0.17408 | 0.15691 | 0.041471 | 2.7204 | 2.7204 | 0.002348
2019-03-29 14:04 | 1553882647.5461 | 660 | 0.20119 | 0.008543 | 0.000184 | 99.0741 | 97.1176 | 111.525 | 21.117 | 0.003845 | 0.003533 | 239.138 | 0.26355 | 8.4033 | 0.003514 | 0.005247 | 0.005178 | 4.8e-05 | 0.000935 | 0.29518 | 1.2973 | 2.5473 | 100.2 | 20.84 | 37.153 | 114.345 | 106.2989 | 106.7594 | 0.009141 | 0.92929 | 0.022052 | 23.7296 | 0.1784 | 0.15679 | 0.039908 | 2.6227 | 2.6227 | 0.002433
2019-03-29 14:04 | 1553882648.5466 | 661 | 0.20149 | 0.00829 | 0.000177 | 99.0142 | 97.2118 | 111.3394 | 21.117 | 0.003841 | 0.003516 | 239.0598 | 0.26433 | 8.4013 | 0.003429 | 0.004579 | 0.004551 | 4.8e-05 | 0.000934 | 0.3027 | 1.2939 | 2.5474 | 100.2 | 20.84 | 37.154 | 114.2289 | 104.2279 | 106.3344 | 0.009142 | 0.93033 | 0.022046 | 23.6966 | 0.16432 | 0.15725 | 0.038424 | 2.5248 | 2.5248 | 0.002535
2019-03-29 14:04 | 1553882649.5462 | 662 | 0.20191 | 0.007925 | 0.00017 | 98.9567 | 97.1446 | 111.1528 | 21.116 | 0.003775 | 0.003512 | 238.9914 | 0.26403 | 8.4008 | 0.003509 | 0.005196 | 0.005109 | 5e-05 | 0.000925 | 0.26178 | 1.3175 | 2.5474 | 100.2 | 20.84 | 37.154 | 112.6316 | 106.1721 | 106.2461 | 0.009142 | 0.93055 | 0.02203 | 23.6739 | 0.16532 | 0.15707 | 0.036343 | 2.3867 | 2.3867 | 0.002678
2019-03-29 14:04 | 1553882650.5447 | 663 | 0.20216 | 0.007639 | 0.000165 | 98.9723 | 97.3971 | 110.9525 | 21.1161 | 0.00387 | 0.003478 | 238.9141 | 0.26278 | 8.4028 | 0.003437 | 0.004996 | 0.004937 | 5.1e-05 | 0.000932 | 0.31868 | 1.324 | 2.5475 | 100.2 | 20.84 | 37.156 | 114.9467 | 104.4102 | 105.416 | 0.009142 | 0.93259 | 0.022034 | 23.6267 | 0.1435 | 0.15633 | 0.035201 | 2.3124 | 2.3124 | 0.002751
2019-03-29 14:04 | 1553882651.5443 | 664 | 0.20243 | 0.007396 | 0.00016 | 98.9026 | 97.4213 | 110.9831 | 21.1169 | 0.003846 | 0.003458 | 238.8298 | 0.26316 | 8.401 | 0.003439 | 0.004527 | 0.004406 | 5.1e-05 | 0.00094 | 0.31133 | 1.336 | 2.5474 | 100.2 | 20.84 | 37.156 | 114.3521 | 104.4786 | 104.9387 | 0.009142 | 0.93376 | 0.022051 | 23.6151 | 0.13497 | 0.15655 | 0.033858 | 2.2261 | 2.2261 | 0.002862
2019-03-29 14:04 | 1553882652.5454 | 665 | 0.20273 | 0.007154 | 0.000157 | 98.9298 | 97.5302 | 110.9981 | 21.1154 | 0.003814 | 0.003482 | 238.7697 | 0.26446 | 8.401 | 0.003404 | 0.005012 | 0.004925 | 5.3e-05 | 0.000945 | 0.28943 | 1.377 | 2.5477 | 100.2 | 20.84 | 37.156 | 113.5705 | 103.6099 | 105.513 | 0.009142 | 0.93235 | 0.022036 | 23.6346 | 0.12745 | 0.15733 | 0.032364 | 2.1266 | 2.1266 | 0.003011
2019-03-29 14:04 | 1553882653.5445 | 666 | 0.20301 | 0.006916 | 0.000153 | 98.8577 | 97.2661 | 111.3044 | 21.1157 | 0.003826 | 0.003475 | 238.695 | 0.26365 | 8.4006 | 0.003443 | 0.005214 | 0.005176 | 5.2e-05 | 0.000953 | 0.29927 | 1.3192 | 2.5475 | 100.2 | 20.84 | 37.157 | 113.8774 | 104.5639 | 105.3528 | 0.009142 | 0.93274 | 0.022071 | 23.6622 | 0.14518 | 0.15684 | 0.031012 | 2.0411 | 2.0411 | 0.003128
2019-03-29 14:04 | 1553882654.5448 | 667 | 0.20315 | 0.006694 | 0.00015 | 98.884 | 97.4169 | 111.7706 | 21.1162 | 0.00385 | 0.003511 | 238.6111 | 0.26355 | 8.4016 | 0.003424 | 0.004863 | 0.004764 | 5e-05 | 0.000955 | 0.29966 | 1.3489 | 2.5472 | 100.2 | 20.84 | 37.158 | 114.455 | 104.0972 | 106.2085 | 0.009143 | 0.93064 | 0.022092 | 23.7385 | 0.1337 | 0.15678 | 0.030416 | 2.004 | 2.004 | 0.003184
2019-03-29 14:04 | 1553882655.5445 | 668 | 0.20333 | 0.006495 | 0.000147 | 98.9035 | 97.3623 | 112.0891 | 21.1158 | 0.003829 | 0.003405 | 238.5056 | 0.26366 | 8.4047 | 0.003404 | 0.004964 | 0.004909 | 5.3e-05 | 0.000944 | 0.28888 | 1.3896 | 2.5474 | 100.2 | 20.85 | 37.16 | 113.9494 | 103.611 | 103.6544 | 0.009149 | 0.93695 | 0.022198 | 23.6921 | 0.14048 | 0.15685 | 0.029592 | 1.9592 | 1.9592 | 0.003259
2019-03-29 14:04 | 1553882656.5446 | 669 | 0.20346 | 0.006303 | 0.000143 | 98.7774 | 97.2766 | 112.0464 | 21.1162 | 0.003741 | 0.00343 | 238.4207 | 0.26483 | 8.4016 | 0.003284 | 0.004694 | 0.004541 | 5.2e-05 | 0.000946 | 0.31061 | 1.3506 | 2.5475 | 100.2 | 20.85 | 37.16 | 111.8012 | 100.7182 | 104.2593 | 0.009149 | 0.93544 | 0.022176 | 23.7066 | 0.13695 | 0.15755 | 0.029089 | 1.9241 | 1.9241 | 0.003333
2019-03-29 14:04 | 1553882657.5459 | 670 | 0.20352 | 0.006133 | 0.000138 | 98.9136 | 97.5169 | 112.5512 | 21.1166 | 0.003802 | 0.00335 | 238.3429 | 0.26376 | 8.4032 | 0.003318 | 0.004115 | 0.004113 | 5.2e-05 | 0.000954 | 0.32278 | 1.3343 | 2.5474 | 100.2 | 20.85 | 37.162 | 113.2942 | 101.5394 | 102.3193 | 0.009149 | 0.94028 | 0.022283 | 23.6988 | 0.1272 | 0.15691 | 0.02892 | 1.9224 | 1.9224 | 0.003322
2019-03-29 14:04 | 1553882658.5456 | 671 | 0.20358 | 0.005983 | 0.000133 | 98.9136 | 97.7377 | 113.0413 | 21.1158 | 0.003741 | 0.003427 | 238.2647 | 0.26469 | 8.4005 | 0.003301 | 0.004725 | 0.004646 | 6.2e-05 | 0.000959 | 0.26537 | 1.3004 | 2.5474 | 100.2 | 20.85 | 37.163 | 111.8 | 101.1171 | 104.1818 | 0.00915 | 0.93564 | 0.022277 | 23.8092 | 0.10697 | 0.15746 | 0.028747 | 1.9106 | 1.9106 | 0.003354
2019-03-29 14:04 | 1553882659.5461 | 672 | 0.20371 | 0.005839 | 0.00013 | 98.8975 | 97.8073 | 113.5164 | 21.1159 | 0.003805 | 0.003361 | 238.1892 | 0.26396 | 8.4 | 0.003292 | 0.005037 | 0.004973 | 5.5e-05 | 0.000965 | 0.30182 | 1.3259 | 2.5475 | 100.2 | 20.85 | 37.167 | 113.3584 | 100.9029 | 102.5763 | 0.009151 | 0.93963 | 0.022371 | 23.8083 | 0.099142 | 0.15703 | 0.028131 | 1.8778 | 1.8778 | 0.003404
2019-03-29 14:04 | 1553882660.5446 | 673 | 0.20384 | 0.005709 | 0.000127 | 98.9235 | 97.9196 | 114.0922 | 21.1157 | 0.003794 | 0.003354 | 238.1189 | 0.26467 | 8.4023 | 0.00321 | 0.004346 | 0.004263 | 5.9e-05 | 0.000969 | 0.27793 | 1.3199 | 2.5472 | 100.2 | 20.85 | 37.167 | 113.0878 | 98.9249 | 102.4206 | 0.009151 | 0.94002 | 0.022432 | 23.8637 | 0.091235 | 0.15745 | 0.02751 | 1.8415 | 1.8415 | 0.00348
2019-03-29 14:04 | 1553882661.5456 | 674 | 0.20402 | 0.005591 | 0.000126 | 98.8966 | 97.822 | 114.2504 | 21.1154 | 0.003719 | 0.003412 | 238.0454 | 0.26426 | 8.4005 | 0.00321 | 0.004888 | 0.004805 | 5.9e-05 | 0.000975 | 0.28567 | 1.3388 | 2.5473 | 100.2 | 20.85 | 37.17 | 111.2746 | 98.9176 | 103.8268 | 0.009151 | 0.93652 | 0.022406 | 23.9249 | 0.097719 | 0.15721 | 0.026625 | 1.7801 | 1.7801 | 0.003594
2019-03-29 14:04 | 1553882662.5447 | 675 | 0.20404 | 0.005473 | 0.000123 | 98.9582 | 97.9073 | 114.4072 | 21.116 | 0.00373 | 0.003355 | 237.972 | 0.26495 | 8.4029 | 0.0032 | 0.004848 | 0.004796 | 5.7e-05 | 0.000978 | 0.29518 | 1.3468 | 2.5475 | 100.2 | 20.85 | 37.17 | 111.5325 | 98.6953 | 102.4357 | 0.009151 | 0.93999 | 0.022463 | 23.8971 | 0.095496 | 0.15762 | 0.02661 | 1.7838 | 1.7838 | 0.003596
2019-03-29 14:04 | 1553882663.5451 | 676 | 0.20406 | 0.005363 | 0.000124 | 98.9478 | 97.7875 | 114.3074 | 21.1153 | 0.003704 | 0.003306 | 237.8911 | 0.2644 | 8.4038 | 0.003151 | 0.004938 | 0.0049 | 6e-05 | 0.000979 | 0.30397 | 1.3385 | 2.5475 | 100.2 | 20.84 | 37.172 | 110.9084 | 97.5009 | 101.2408 | 0.009146 | 0.94299 | 0.022489 | 23.8486 | 0.1055 | 0.15729 | 0.026628 | 1.787 | 1.787 | 0.003582
2019-03-29 14:04 | 1553882664.5447 | 677 | 0.20406 | 0.005262 | 0.000124 | 98.9352 | 98.1111 | 114.072 | 21.1157 | 0.003757 | 0.003426 | 237.8101 | 0.26473 | 8.403 | 0.003056 | 0.004532 | 0.004499 | 5.8e-05 | 0.000979 | 0.31852 | 1.3355 | 2.5473 | 100.2 | 20.85 | 37.174 | 112.189 | 95.2038 | 104.1439 | 0.009152 | 0.93573 | 0.022379 | 23.9163 | 0.074818 | 0.15749 | 0.026788 | 1.789 | 1.789 | 0.003583
2019-03-29 14:04 | 1553882665.5457 | 678 | 0.20405 | 0.00517 | 0.000124 | 98.9361 | 97.9049 | 113.9863 | 21.1164 | 0.003762 | 0.0034 | 237.7263 | 0.26359 | 8.4037 | 0.003081 | 0.004202 | 0.004157 | 5.9e-05 | 0.000979 | 0.28799 | 1.3253 | 2.5474 | 100.2 | 20.85 | 37.176 | 112.3216 | 95.8216 | 103.5307 | 0.009153 | 0.93725 | 0.022389 | 23.8878 | 0.093712 | 0.15681 | 0.026941 | 1.8 | 1.8 | 0.003546
2019-03-29 14:04 | 1553882666.5452 | 679 | 0.20403 | 0.005093 | 0.000126 | 98.9828 | 98.0046 | 113.7357 | 21.1157 | 0.003677 | 0.003327 | 237.6252 | 0.26416 | 8.4039 | 0.003033 | 0.004844 | 0.004735 | 5.8e-05 | 0.000975 | 0.29822 | 1.3508 | 2.5475 | 100.2 | 20.84 | 37.178 | 110.2533 | 94.6409 | 101.7492 | 0.009148 | 0.94171 | 0.022417 | 23.8051 | 0.088833 | 0.15715 | 0.027131 | 1.8149 | 1.8149 | 0.003524
2019-03-29 14:04 | 1553882667.544 | 680 | 0.20397 | 0.005035 | 0.000131 | 98.9792 | 97.9913 | 113.2341 | 21.1163 | 0.003742 | 0.003376 | 237.54 | 0.26394 | 8.4025 | 0.003056 | 0.004313 | 0.004288 | 6e-05 | 0.000973 | 0.28672 | 1.348 | 2.5473 | 100.2 | 20.85 | 37.18 | 111.8311 | 95.1998 | 102.9498 | 0.009154 | 0.9387 | 0.022332 | 23.7905 | 0.089723 | 0.15702 | 0.027572 | 1.8371 | 1.8371 | 0.003479
2019-03-29 14:04 | 1553882668.5449 | 681 | 0.20365 | 0.005007 | 0.000136 | 98.9442 | 97.893 | 112.4123 | 21.1159 | 0.00379 | 0.003396 | 237.4379 | 0.26446 | 8.4008 | 0.002989 | 0.004484 | 0.004351 | 6e-05 | 0.000977 | 0.30541 | 1.323 | 2.5475 | 100.2 | 20.85 | 37.181 | 112.987 | 93.5846 | 103.4367 | 0.009154 | 0.93749 | 0.022237 | 23.7194 | 0.095534 | 0.15733 | 0.029544 | 1.96 | 1.96 | 0.003267
2019-03-29 14:04 | 1553882669.5443 | 682 | 0.20338 | 0.005 | 0.000142 | 98.9576 | 97.6776 | 111.8065 | 21.1154 | 0.00373 | 0.003433 | 237.3189 | 0.26334 | 8.4025 | 0.002946 | 0.004761 | 0.004719 | 6e-05 | 0.000978 | 0.28263 | 1.3331 | 2.5474 | 100.2 | 20.85 | 37.183 | 111.5288 | 92.5649 | 104.3331 | 0.009155 | 0.93526 | 0.02215 | 23.6835 | 0.11645 | 0.15666 | 0.031136 | 2.0573 | 2.0573 | 0.003099
2019-03-29 14:04 | 1553882670.5453 | 683 | 0.20316 | 0.005021 | 0.000148 | 99.0259 | 97.4408 | 111.2103 | 21.1143 | 0.003757 | 0.003436 | 237.2002 | 0.26433 | 8.4025 | 0.002975 | 0.004203 | 0.00422 | 6e-05 | 0.000982 | 0.29822 | 1.3314 | 2.5474 | 100.2 | 20.85 | 37.191 | 112.1831 | 93.2585 | 104.398 | 0.009156 | 0.9351 | 0.022089 | 23.6223 | 0.14433 | 0.15725 | 0.032419 | 2.1359 | 2.1359 | 0.002997
2019-03-29 14:04 | 1553882671.545 | 684 | 0.20285 | 0.005071 | 0.000153 | 98.9391 | 97.0541 | 110.8983 | 21.1145 | 0.003742 | 0.003389 | 237.1066 | 0.26405 | 8.401 | 0.002984 | 0.005132 | 0.005018 | 5.8e-05 | 0.000974 | 0.30934 | 1.3486 | 2.5473 | 100.2 | 20.85 | 37.191 | 111.8262 | 93.4682 | 103.2505 | 0.009156 | 0.93795 | 0.022092 | 23.5533 | 0.17205 | 0.15709 | 0.034247 | 2.2563 | 2.2563 | 0.002834
2019-03-29 14:05 | 1553882672.545 | 685 | 0.20243 | 0.005143 | 0.000158 | 98.8508 | 96.6821 | 110.411 | 21.1144 | 0.00375 | 0.003471 | 237.0032 | 0.26394 | 8.4015 | 0.002965 | 0.004658 | 0.004544 | 5.7e-05 | 0.000963 | 0.32045 | 1.3395 | 2.5474 | 100.2 | 20.85 | 37.192 | 112.014 | 93.023 | 105.2499 | 0.009157 | 0.933 | 0.021985 | 23.5638 | 0.19842 | 0.15702 | 0.03663 | 2.4013 | 2.4013 | 0.002661
2019-03-29 14:05 | 1553882673.544 | 686 | 0.20205 | 0.005252 | 0.000161 | 98.9508 | 96.634 | 110.0866 | 21.1151 | 0.003856 | 0.003406 | 236.8951 | 0.2634 | 8.4041 | 0.003034 | 0.004729 | 0.004697 | 5.5e-05 | 0.000957 | 0.29894 | 1.3526 | 2.5475 | 100.2 | 20.85 | 37.197 | 114.6042 | 94.6678 | 103.6748 | 0.009158 | 0.9369 | 0.021998 | 23.4802 | 0.21191 | 0.1567 | 0.038796 | 2.5447 | 2.5447 | 0.002506
2019-03-29 14:05 | 1553882674.5441 | 687 | 0.20175 | 0.005387 | 0.000161 | 98.9696 | 96.5695 | 109.9953 | 21.114 | 0.003773 | 0.003409 | 236.7743 | 0.26507 | 8.3999 | 0.003026 | 0.004367 | 0.004261 | 5.6e-05 | 0.000949 | 0.32139 | 1.363 | 2.5473 | 100.2 | 20.85 | 37.2 | 112.5878 | 94.4844 | 103.7501 | 0.009159 | 0.93671 | 0.021987 | 23.4728 | 0.21959 | 0.15769 | 0.04037 | 2.6465 | 2.6465 | 0.002425
2019-03-29 14:05 | 1553882675.5444 | 688 | 0.2017 | 0.005544 | 0.000161 | 99.078 | 96.7292 | 109.6712 | 21.114 | 0.003798 | 0.003409 | 236.6695 | 0.26465 | 8.3999 | 0.003097 | 0.004625 | 0.004609 | 6e-05 | 0.000942 | 0.29502 | 1.387 | 2.5475 | 100.2 | 20.85 | 37.208 | 113.1907 | 96.1966 | 103.7329 | 0.009161 | 0.93675 | 0.021955 | 23.4376 | 0.2146 | 0.15744 | 0.040491 | 2.6506 | 2.6506 | 0.002418
2019-03-29 14:05 | 1553882676.5451 | 689 | 0.2017 | 0.005709 | 0.000158 | 99.0813 | 96.9388 | 109.7306 | 21.1131 | 0.003763 | 0.003444 | 236.6025 | 0.26384 | 8.404 | 0.003072 | 0.004792 | 0.004723 | 5.7e-05 | 0.000937 | 0.29695 | 1.3428 | 2.5472 | 100.2 | 20.85 | 37.212 | 112.3476 | 95.5874 | 104.5786 | 0.009162 | 0.93465 | 0.021937 | 23.4702 | 0.19554 | 0.15696 | 0.040286 | 2.6351 | 2.6351 | 0.002424
2019-03-29 14:05 | 1553882677.5442 | 690 | 0.20177 | 0.005865 | 0.000155 | 99.1633 | 97.2548 | 109.905 | 21.1142 | 0.00382 | 0.003443 | 236.5199 | 0.26476 | 8.4007 | 0.003086 | 0.00479 | 0.004695 | 5.3e-05 | 0.000933 | 0.33052 | 1.3341 | 2.5473 | 100.2 | 20.85 | 37.217 | 113.7225 | 95.9409 | 104.5695 | 0.009163 | 0.93468 | 0.021954 | 23.4886 | 0.17382 | 0.1575 | 0.039735 | 2.6013 | 2.6013 | 0.002464
2019-03-29 14:05 | 1553882678.5451 | 691 | 0.20193 | 0.006003 | 0.000151 | 99.1669 | 97.6343 | 110.297 | 21.1133 | 0.003831 | 0.003397 | 236.4237 | 0.26396 | 8.4014 | 0.003174 | 0.005003 | 0.004933 | 5.5e-05 | 0.000934 | 0.29463 | 1.3514 | 2.5475 | 100.2 | 20.85 | 37.228 | 113.9864 | 98.0562 | 103.4444 | 0.009166 | 0.93747 | 0.022026 | 23.4954 | 0.13931 | 0.15703 | 0.038567 | 2.5334 | 2.5334 | 0.002523
2019-03-29 14:05 | 1553882679.5448 | 692 | 0.20212 | 0.006107 | 0.000147 | 99.1953 | 97.9507 | 110.2876 | 21.1142 | 0.003779 | 0.003379 | 236.3313 | 0.26335 | 8.4013 | 0.003221 | 0.005024 | 0.004927 | 5.2e-05 | 0.000939 | 0.29623 | 1.3831 | 2.5474 | 100.2 | 20.85 | 37.236 | 112.7326 | 99.1963 | 103.0233 | 0.009168 | 0.93852 | 0.022038 | 23.4812 | 0.11294 | 0.15667 | 0.037325 | 2.4533 | 2.4533 | 0.002599
2019-03-29 14:05 | 1553882680.5442 | 693 | 0.2023 | 0.006169 | 0.000143 | 99.2136 | 98.1329 | 110.3195 | 21.1142 | 0.003806 | 0.003432 | 236.2588 | 0.26433 | 8.4003 | 0.00319 | 0.004317 | 0.004272 | 5.1e-05 | 0.000937 | 0.29159 | 1.375 | 2.5474 | 100.2 | 20.85 | 37.251 | 113.3866 | 98.4435 | 104.31 | 0.009171 | 0.93532 | 0.022003 | 23.5248 | 0.097959 | 0.15725 | 0.036164 | 2.3736 | 2.3736 | 0.002696
2019-03-29 14:05 | 1553882681.5447 | 694 | 0.2025 | 0.00619 | 0.000138 | 99.2232 | 98.0179 | 110.2251 | 21.1133 | 0.003808 | 0.003382 | 236.1713 | 0.26471 | 8.4005 | 0.003175 | 0.004055 | 0.004008 | 5.1e-05 | 0.000945 | 0.32029 | 1.379 | 2.5474 | 100.2 | 20.85 | 37.259 | 113.4254 | 98.0859 | 103.0875 | 0.009173 | 0.93836 | 0.022029 | 23.4766 | 0.10931 | 0.15748 | 0.034997 | 2.3 | 2.3 | 0.002787
2019-03-29 14:05 | 1553882682.5443 | 695 | 0.20275 | 0.006167 | 0.000133 | 99.2645 | 98.1643 | 110.3167 | 21.1124 | 0.00375 | 0.003386 | 236.0843 | 0.26363 | 8.4031 | 0.00321 | 0.004546 | 0.004423 | 5.3e-05 | 0.00094 | 0.30254 | 1.3134 | 2.5474 | 100.2 | 20.85 | 37.267 | 112.0113 | 98.914 | 103.1834 | 0.009175 | 0.93812 | 0.022036 | 23.4893 | 0.099684 | 0.15684 | 0.033529 | 2.2044 | 2.2044 | 0.002896
2019-03-29 14:05 | 1553882683.5443 | 696 | 0.20305 | 0.006124 | 0.000129 | 99.2276 | 98.2828 | 110.5002 | 21.113 | 0.003743 | 0.003368 | 236.0021 | 0.26401 | 8.4031 | 0.003182 | 0.004512 | 0.004463 | 5.1e-05 | 0.00094 | 0.28368 | 1.357 | 2.5475 | 100.2 | 20.85 | 37.287 | 111.8525 | 98.2582 | 102.7572 | 0.00918 | 0.93918 | 0.022067 | 23.4956 | 0.085582 | 0.15706 | 0.031744 | 2.0902 | 2.0902 | 0.003058
2019-03-29 14:05 | 1553882684.5434 | 697 | 0.20335 | 0.006055 | 0.000126 | 99.2372 | 98.236 | 110.4049 | 21.1125 | 0.003772 | 0.003402 | 235.9309 | 0.26423 | 8.4013 | 0.003229 | 0.004361 | 0.004361 | 5.2e-05 | 0.00095 | 0.29695 | 1.3801 | 2.5473 | 100.2 | 20.85 | 37.297 | 112.5493 | 99.39 | 103.571 | 0.009183 | 0.93715 | 0.022033 | 23.5108 | 0.090701 | 0.15719 | 0.030072 | 1.9773 | 1.9773 | 0.003236
2019-03-29 14:05 | 1553882685.5448 | 698 | 0.20352 | 0.005964 | 0.000121 | 99.2273 | 98.1101 | 110.4789 | 21.1118 | 0.003769 | 0.003408 | 235.8619 | 0.26359 | 8.4024 | 0.003157 | 0.004382 | 0.004345 | 4.9e-05 | 0.000951 | 0.28102 | 1.3898 | 2.5475 | 100.2 | 20.85 | 37.308 | 112.4743 | 97.6479 | 103.7164 | 0.009185 | 0.93679 | 0.022036 | 23.5232 | 0.10129 | 0.15681 | 0.029189 | 1.9197 | 1.9197 | 0.003324
2019-03-29 14:05 | 1553882686.5438 | 699 | 0.20379 | 0.005855 | 0.000118 | 99.1839 | 98.1691 | 110.7028 | 21.1132 | 0.003722 | 0.003382 | 235.7984 | 0.26357 | 8.4024 | 0.003149 | 0.004887 | 0.004853 | 5e-05 | 0.000955 | 0.29717 | 1.3769 | 2.5474 | 100.2 | 20.84 | 37.329 | 111.3262 | 97.4572 | 103.0756 | 0.009185 | 0.93839 | 0.022077 | 23.527 | 0.09199 | 0.1568 | 0.027662 | 1.8228 | 1.8228 | 0.003501
2019-03-29 14:05 | 1553882687.5434 | 700 | 0.20386 | 0.005737 | 0.000114 | 99.2845 | 98.2114 | 110.8626 | 21.1131 | 0.003788 | 0.003336 | 235.7271 | 0.26395 | 8.4039 | 0.003069 | 0.004879 | 0.004833 | 5.2e-05 | 0.000955 | 0.27577 | 1.3671 | 2.5474 | 100.2 | 20.84 | 37.34 | 112.9519 | 95.5107 | 101.9625 | 0.009188 | 0.94117 | 0.022126 | 23.5092 | 0.097201 | 0.15702 | 0.027432 | 1.8118 | 1.8118 | 0.003527
2019-03-29 14:05 | 1553882688.5435 | 701 | 0.20398 | 0.005602 | 0.000111 | 99.219 | 98.1531 | 111.0877 | 21.1128 | 0.003686 | 0.003367 | 235.6574 | 0.26376 | 8.4014 | 0.003146 | 0.004155 | 0.004025 | 5.1e-05 | 0.000954 | 0.27367 | 1.3848 | 2.5474 | 100.2 | 20.84 | 37.361 | 110.4562 | 97.3794 | 102.7217 | 0.009193 | 0.93927 | 0.022126 | 23.5568 | 0.09661 | 0.15691 | 0.026869 | 1.7748 | 1.7748 | 0.003598
2019-03-29 14:05 | 1553882689.5445 | 702 | 0.20408 | 0.005467 | 0.000109 | 99.3285 | 98.4012 | 111.379 | 21.1124 | 0.003692 | 0.003314 | 235.5981 | 0.26409 | 8.4029 | 0.003098 | 0.004791 | 0.00476 | 5.1e-05 | 0.000952 | 0.31061 | 1.3281 | 2.5474 | 100.2 | 20.84 | 37.372 | 110.6159 | 96.2092 | 101.4458 | 0.009195 | 0.94247 | 0.022193 | 23.5476 | 0.083899 | 0.1571 | 0.0264 | 1.7492 | 1.7492 | 0.003656
2019-03-29 14:05 | 1553882690.5446 | 703 | 0.20417 | 0.005342 | 0.000106 | 99.3441 | 98.3565 | 111.6349 | 21.1128 | 0.003726 | 0.003288 | 235.5396 | 0.26411 | 8.4004 | 0.003072 | 0.004284 | 0.00423 | 5.3e-05 | 0.000958 | 0.30182 | 1.3638 | 2.5474 | 100.2 | 20.84 | 37.383 | 111.445 | 95.5901 | 100.8114 | 0.009198 | 0.94407 | 0.022237 | 23.5547 | 0.089365 | 0.15712 | 0.026032 | 1.7284 | 1.7284 | 0.0037
2019-03-29 14:05 | 1553882691.5435 | 704 | 0.20422 | 0.005221 | 0.000103 | 99.3827 | 98.4094 | 112.1816 | 21.1112 | 0.003656 | 0.003354 | 235.4836 | 0.26291 | 8.4021 | 0.003047 | 0.00457 | 0.00444 | 5.4e-05 | 0.000964 | 0.31741 | 1.3436 | 2.5473 | 100.2 | 20.84 | 37.402 | 109.7299 | 94.9818 | 102.4163 | 0.009203 | 0.94003 | 0.022244 | 23.6629 | 0.088032 | 0.1564 | 0.025918 | 1.7215 | 1.7215 | 0.003698
2019-03-29 14:05 | 1553882692.5436 | 705 | 0.20427 | 0.005105 | 0.000102 | 99.4474 | 98.5176 | 112.3477 | 21.1122 | 0.003702 | 0.003317 | 235.4143 | 0.26539 | 8.4016 | 0.003063 | 0.004148 | 0.004053 | 5.3e-05 | 0.000972 | 0.32388 | 1.3713 | 2.5474 | 100.2 | 20.84 | 37.411 | 110.8446 | 95.3766 | 101.5124 | 0.009205 | 0.9423 | 0.022287 | 23.6519 | 0.084021 | 0.15788 | 0.025719 | 1.7116 | 1.7116 | 0.003754
2019-03-29 14:05 | 1553882693.5447 | 706 | 0.20445 | 0.004999 | 0.0001 | 99.4581 | 98.4998 | 112.4273 | 21.1116 | 0.003685 | 0.003351 | 235.3485 | 0.26392 | 8.401 | 0.003002 | 0.003906 | 0.003829 | 5.7e-05 | 0.000972 | 0.2995 | 1.3718 | 2.5474 | 100.2 | 20.83 | 37.428 | 110.4306 | 93.9091 | 102.3309 | 0.009203 | 0.94025 | 0.022271 | 23.6861 | 0.086602 | 0.15701 | 0.024765 | 1.647 | 1.647 | 0.00388
2019-03-29 14:05 | 1553882694.5447 | 707 | 0.20453 | 0.004902 | 9.9e-05 | 99.5153 | 98.5933 | 112.6879 | 21.1118 | 0.003645 | 0.003303 | 235.291 | 0.26334 | 8.4018 | 0.003041 | 0.004458 | 0.004363 | 5.8e-05 | 0.000973 | 0.29463 | 1.3121 | 2.5474 | 100.2 | 20.83 | 37.435 | 109.4719 | 94.8492 | 101.1651 | 0.009205 | 0.94318 | 0.022331 | 23.6767 | 0.083254 | 0.15666 | 0.02442 | 1.6285 | 1.6285 | 0.003915
2019-03-29 14:05 | 1553882695.5447 | 708 | 0.20456 | 0.004808 | 9.7e-05 | 99.4749 | 98.5019 | 113.0531 | 21.1109 | 0.003679 | 0.003332 | 235.2343 | 0.26381 | 8.4032 | 0.003057 | 0.004358 | 0.004292 | 5.7e-05 | 0.000976 | 0.31365 | 1.342 | 2.5474 | 100.2 | 20.83 | 37.443 | 110.2894 | 95.2305 | 101.8701 | 0.009207 | 0.9414 | 0.022346 | 23.7374 | 0.087923 | 0.15694 | 0.024383 | 1.6272 | 1.6272 | 0.003925
2019-03-29 14:05 | 1553882696.5437 | 709 | 0.20459 | 0.004724 | 9.6e-05 | 99.4914 | 98.5169 | 113.2371 | 21.1115 | 0.003633 | 0.003273 | 235.1523 | 0.26442 | 8.4024 | 0.003041 | 0.004605 | 0.004537 | 5.6e-05 | 0.000979 | 0.28871 | 1.3376 | 2.5474 | 100.2 | 20.83 | 37.455 | 109.1742 | 94.8436 | 100.4426 | 0.00921 | 0.945 | 0.022407 | 23.7114 | 0.08804 | 0.1573 | 0.024293 | 1.6257 | 1.6257 | 0.003938
2019-03-29 14:05 | 1553882697.5438 | 710 | 0.20459 | 0.004658 | 9.5e-05 | 99.4839 | 98.5466 | 113.3582 | 21.1116 | 0.003663 | 0.003285 | 235.0711 | 0.26413 | 8.4027 | 0.003037 | 0.004099 | 0.004 | 5.9e-05 | 0.000984 | 0.27057 | 1.3543 | 2.5474 | 100.2 | 20.83 | 37.461 | 109.9052 | 94.7568 | 100.7385 | 0.009212 | 0.94425 | 0.02241 | 23.7335 | 0.084672 | 0.15713 | 0.024359 | 1.6304 | 1.6304 | 0.003923
2019-03-29 14:05 | 1553882698.5438 | 711 | 0.20459 | 0.004592 | 9.4e-05 | 99.5333 | 98.5998 | 113.6774 | 21.1114 | 0.003706 | 0.003289 | 235.0057 | 0.26428 | 8.4037 | 0.002965 | 0.004135 | 0.004 | 5.6e-05 | 0.000992 | 0.33212 | 1.3359 | 2.5475 | 100.2 | 20.83 | 37.473 | 110.9518 | 93.0194 | 100.8335 | 0.009215 | 0.94401 | 0.022439 | 23.7699 | 0.08428 | 0.15722 | 0.024471 | 1.64 | 1.64 | 0.003902
2019-03-29 14:05 | 1553882699.5438 | 712 | 0.2046 | 0.004543 | 9.3e-05 | 99.5144 | 98.608 | 113.8511 | 21.1123 | 0.003672 | 0.00327 | 234.9399 | 0.26409 | 8.4025 | 0.002909 | 0.004882 | 0.004845 | 5.7e-05 | 0.000994 | 0.33107 | 1.3194 | 2.5473 | 100.2 | 20.83 | 37.477 | 110.1213 | 91.6491 | 100.3713 | 0.009216 | 0.94518 | 0.02247 | 23.7733 | 0.081842 | 0.1571 | 0.024471 | 1.6424 | 1.6424 | 0.003893
2019-03-29 14:05 | 1553882700.5441 | 713 | 0.20458 | 0.004497 | 9.2e-05 | 99.5153 | 98.492 | 113.9157 | 21.1109 | 0.003665 | 0.003281 | 234.8881 | 0.26458 | 8.4016 | 0.00295 | 0.004098 | 0.003961 | 5.9e-05 | 0.000994 | 0.2975 | 1.3228 | 2.5474 | 100.2 | 20.83 | 37.482 | 109.9495 | 92.6411 | 100.6406 | 0.009217 | 0.9445 | 0.022468 | 23.7887 | 0.092455 | 0.1574 | 0.024623 | 1.6525 | 1.6525 | 0.003877
2019-03-29 14:05 | 1553882701.543 | 714 | 0.2046 | 0.004461 | 9.1e-05 | 99.5455 | 98.55 | 113.7662 | 21.111 | 0.003681 | 0.003275 | 234.8123 | 0.26221 | 8.4021 | 0.002984 | 0.004718 | 0.004631 | 5.9e-05 | 0.000995 | 0.30431 | 1.3199 | 2.5472 | 100.2 | 20.83 | 37.485 | 110.3408 | 93.4563 | 100.5026 | 0.009217 | 0.94485 | 0.022458 | 23.7686 | 0.089904 | 0.15599 | 0.024558 | 1.6473 | 1.6473 | 0.003854
2019-03-29 14:05 | 1553882702.544 | 715 | 0.20461 | 0.004438 | 9e-05 | 99.5027 | 98.5623 | 113.7387 | 21.1113 | 0.003632 | 0.003252 | 234.7269 | 0.26446 | 8.4009 | 0.002924 | 0.004499 | 0.004419 | 5.7e-05 | 0.000999 | 0.32421 | 1.3803 | 2.5475 | 100.2 | 20.83 | 37.489 | 109.1426 | 92.0244 | 99.9317 | 0.009218 | 0.94629 | 0.022472 | 23.7476 | 0.084942 | 0.15733 | 0.024544 | 1.6476 | 1.6476 | 0.003887
2019-03-29 14:05 | 1553882703.544 | 716 | 0.20465 | 0.004423 | 8.9e-05 | 99.5132 | 98.5104 | 113.8223 | 21.1124 | 0.003669 | 0.003338 | 234.6594 | 0.26336 | 8.4033 | 0.002891 | 0.004079 | 0.004076 | 5.9e-05 | 0.00099 | 0.29966 | 1.4092 | 2.5475 | 100.2 | 20.83 | 37.495 | 110.0417 | 91.2365 | 102.0097 | 0.00922 | 0.94105 | 0.022418 | 23.8224 | 0.090593 | 0.15667 | 0.024296 | 1.627 | 1.627 | 0.003919
2019-03-29 14:05 | 1553882704.5431 | 717 | 0.20465 | 0.004398 | 8.9e-05 | 99.5246 | 98.5981 | 113.719 | 21.1115 | 0.003645 | 0.003295 | 234.5999 | 0.26465 | 8.4038 | 0.002901 | 0.005092 | 0.004913 | 5.6e-05 | 0.000988 | 0.30973 | 1.3619 | 2.5476 | 100.2 | 20.83 | 37.499 | 109.4621 | 91.4671 | 100.9825 | 0.009221 | 0.94364 | 0.022439 | 23.779 | 0.083655 | 0.15744 | 0.024363 | 1.633 | 1.633 | 0.003924
2019-03-29 14:05 | 1553882705.544 | 718 | 0.20462 | 0.004377 | 8.8e-05 | 99.4554 | 98.5356 | 113.4193 | 21.1105 | 0.003613 | 0.003282 | 234.5347 | 0.26432 | 8.4044 | 0.002847 | 0.004612 | 0.004696 | 5.9e-05 | 0.000983 | 0.34219 | 1.3675 | 2.5475 | 100.2 | 20.83 | 37.505 | 108.698 | 90.1741 | 100.6528 | 0.009222 | 0.94447 | 0.022419 | 23.7371 | 0.083106 | 0.15724 | 0.024559 | 1.6447 | 1.6447 | 0.003891
2019-03-29 14:05 | 1553882706.544 | 719 | 0.2046 | 0.004354 | 8.8e-05 | 99.5476 | 98.3565 | 113.1627 | 21.1105 | 0.00367 | 0.003282 | 234.4531 | 0.26395 | 8.4013 | 0.002806 | 0.003995 | 0.003837 | 5.7e-05 | 0.000984 | 0.29717 | 1.3625 | 2.5475 | 100.2 | 20.83 | 37.508 | 110.069 | 89.1865 | 100.6731 | 0.009223 | 0.94442 | 0.022393 | 23.7109 | 0.10767 | 0.15702 | 0.024705 | 1.6526 | 1.6526 | 0.003867
2019-03-29 14:05 | 1553882707.5441 | 720 | 0.20459 | 0.004336 | 8.8e-05 | 99.4845 | 98.3671 | 112.8573 | 21.1106 | 0.003627 | 0.003252 | 234.3918 | 0.26363 | 8.4034 | 0.002923 | 0.004664 | 0.0046 | 5.7e-05 | 0.00098 | 0.28655 | 1.3418 | 2.5472 | 100.2 | 20.83 | 37.511 | 109.0267 | 92.0056 | 99.9491 | 0.009224 | 0.94625 | 0.022384 | 23.656 | 0.10103 | 0.15684 | 0.024759 | 1.6556 | 1.6556 | 0.003855
2019-03-29 14:05 | 1553882708.5431 | 721 | 0.20457 | 0.004332 | 8.8e-05 | 99.568 | 98.4858 | 112.5024 | 21.1116 | 0.003648 | 0.003274 | 234.2982 | 0.26401 | 8.4015 | 0.002892 | 0.004539 | 0.004479 | 5.6e-05 | 0.00098 | 0.30879 | 1.3143 | 2.5473 | 100.2 | 20.83 | 37.518 | 109.5433 | 91.2412 | 100.464 | 0.009226 | 0.94495 | 0.022334 | 23.635 | 0.097746 | 0.15706 | 0.02492 | 1.6627 | 1.6627 | 0.003845
2019-03-29 14:05 | 1553882709.5436 | 722 | 0.20457 | 0.004328 | 8.7e-05 | 99.5704 | 98.7005 | 112.2783 | 21.1107 | 0.003706 | 0.003322 | 234.2213 | 0.26409 | 8.4018 | 0.002885 | 0.004182 | 0.004119 | 5.5e-05 | 0.000982 | 0.30093 | 1.3523 | 2.5475 | 100.2 | 20.83 | 37.52 | 110.9433 | 91.0785 | 101.6351 | 0.009226 | 0.94199 | 0.022277 | 23.6485 | 0.078486 | 0.1571 | 0.024906 | 1.6575 | 1.6575 | 0.003858
2019-03-29 14:05 | 1553882710.5432 | 723 | 0.20452 | 0.004322 | 8.8e-05 | 99.5479 | 98.5179 | 111.9938 | 21.1105 | 0.00369 | 0.00325 | 234.1274 | 0.26402 | 8.4017 | 0.002853 | 0.004207 | 0.004079 | 5.3e-05 | 0.000978 | 0.32084 | 1.3556 | 2.5475 | 100.2 | 20.83 | 37.525 | 110.5533 | 90.2998 | 99.8878 | 0.009227 | 0.94641 | 0.022301 | 23.5634 | 0.093032 | 0.15706 | 0.025208 | 1.6793 | 1.6793 | 0.003807
2019-03-29 14:05 | 1553882711.5432 | 724 | 0.20452 | 0.004319 | 8.7e-05 | 99.6123 | 98.5711 | 111.8759 | 21.1102 | 0.003603 | 0.003236 | 234.0482 | 0.26491 | 8.4017 | 0.002817 | 0.004866 | 0.00473 | 5.6e-05 | 0.000975 | 0.29911 | 1.3431 | 2.5474 | 100.2 | 20.82 | 37.528 | 108.4446 | 89.4323 | 99.5471 | 0.009222 | 0.94727 | 0.022299 | 23.5402 | 0.09398 | 0.15759 | 0.025221 | 1.6801 | 1.6801 | 0.003818
2019-03-29 14:05 | 1553882712.5432 | 725 | 0.20441 | 0.004317 | 8.7e-05 | 99.6958 | 98.7698 | 111.6692 | 21.1097 | 0.003582 | 0.003305 | 233.9891 | 0.264 | 8.4032 | 0.002838 | 0.004551 | 0.00448 | 5.6e-05 | 0.000976 | 0.29927 | 1.3875 | 2.5474 | 100.2 | 20.82 | 37.531 | 107.94 | 89.9444 | 101.2218 | 0.009223 | 0.94303 | 0.022228 | 23.5712 | 0.08347 | 0.15705 | 0.025884 | 1.7189 | 1.7189 | 0.003719
2019-03-29 14:05 | 1553882713.5422 | 726 | 0.2044 | 0.004322 | 8.8e-05 | 99.6084 | 98.5493 | 111.6987 | 21.1111 | 0.003673 | 0.003305 | 233.9076 | 0.26461 | 8.3997 | 0.002851 | 0.004364 | 0.004316 | 5.2e-05 | 0.000968 | 0.29767 | 1.3395 | 2.5474 | 100.2 | 20.82 | 37.538 | 110.1486 | 90.2685 | 101.2256 | 0.009225 | 0.94303 | 0.022231 | 23.5745 | 0.095612 | 0.15742 | 0.025933 | 1.7223 | 1.7223 | 0.00372
2019-03-29 14:05 | 1553882714.5432 | 727 | 0.2044 | 0.004338 | 8.8e-05 | 99.6641 | 98.6739 | 111.5166 | 21.1103 | 0.003582 | 0.003228 | 233.8292 | 0.26309 | 8.4022 | 0.002826 | 0.004473 | 0.004369 | 5.6e-05 | 0.000968 | 0.26753 | 1.3639 | 2.5474 | 100.2 | 20.82 | 37.541 | 107.9399 | 89.6494 | 99.3644 | 0.009226 | 0.94774 | 0.022269 | 23.4966 | 0.089311 | 0.15651 | 0.025948 | 1.7262 | 1.7262 | 0.00369
2019-03-29 14:05 | 1553882715.5427 | 728 | 0.20439 | 0.004354 | 8.8e-05 | 99.712 | 98.7995 | 111.552 | 21.1105 | 0.003625 | 0.003234 | 233.7385 | 0.26435 | 8.404 | 0.002833 | 0.004425 | 0.004311 | 5.5e-05 | 0.000969 | 0.29391 | 1.3107 | 2.5475 | 100.2 | 20.82 | 37.545 | 108.9863 | 89.8247 | 99.4985 | 0.009227 | 0.94739 | 0.022268 | 23.5046 | 0.082231 | 0.15726 | 0.025963 | 1.7271 | 1.7271 | 0.003706
2019-03-29 14:05 | 1553882716.5425 | 729 | 0.2044 | 0.004365 | 8.8e-05 | 99.7913 | 98.7892 | 111.3948 | 21.1096 | 0.003637 | 0.003302 | 233.6574 | 0.2645 | 8.4037 | 0.002852 | 0.003993 | 0.003975 | 5.4e-05 | 0.000969 | 0.30431 | 1.3068 | 2.5474 | 100.2 | 20.82 | 37.551 | 109.2797 | 90.2896 | 101.144 | 0.009228 | 0.94323 | 0.022203 | 23.5398 | 0.090271 | 0.15735 | 0.025901 | 1.718 | 1.718 | 0.003728
2019-03-29 14:05 | 1553882717.5424 | 730 | 0.20435 | 0.004385 | 8.9e-05 | 99.8954 | 98.7991 | 111.5877 | 21.11 | 0.003629 | 0.003298 | 233.5756 | 0.26448 | 8.4028 | 0.002857 | 0.004495 | 0.004376 | 5.5e-05 | 0.000973 | 0.28512 | 1.3274 | 2.5475 | 100.2 | 20.82 | 37.554 | 109.0757 | 90.3938 | 101.0541 | 0.009229 | 0.94346 | 0.022225 | 23.5573 | 0.098707 | 0.15734 | 0.026184 | 1.7384 | 1.7384 | 0.003684
2019-03-29 14:05 | 1553882718.5438 | 731 | 0.20438 | 0.004405 | 8.9e-05 | 99.9281 | 99.0694 | 111.5271 | 21.1099 | 0.003652 | 0.003235 | 233.523 | 0.26458 | 8.4027 | 0.002801 | 0.004043 | 0.004036 | 5.4e-05 | 0.000973 | 0.29319 | 1.3201 | 2.5474 | 100.2 | 20.82 | 37.559 | 109.6267 | 89.0474 | 99.5168 | 0.00923 | 0.94735 | 0.022265 | 23.5025 | 0.07719 | 0.1574 | 0.02595 | 1.726 | 1.726 | 0.003712
2019-03-29 14:05 | 1553882719.5429 | 732 | 0.20437 | 0.004424 | 9e-05 | 100.0152 | 99.0565 | 111.3028 | 21.1107 | 0.00358 | 0.003233 | 233.4464 | 0.26506 | 8.401 | 0.002747 | 0.004458 | 0.004339 | 5.3e-05 | 0.000976 | 0.30453 | 1.3369 | 2.5474 | 100.2 | 20.82 | 37.561 | 107.8859 | 87.751 | 99.4676 | 0.00923 | 0.94747 | 0.022244 | 23.4773 | 0.086154 | 0.15769 | 0.026026 | 1.7294 | 1.7294 | 0.003711
2019-03-29 14:05 | 1553882720.5439 | 733 | 0.20436 | 0.004431 | 8.9e-05 | 100.0275 | 99.0844 | 111.4147 | 21.1117 | 0.003659 | 0.003307 | 233.3631 | 0.26443 | 8.4042 | 0.002824 | 0.004768 | 0.004594 | 5.4e-05 | 0.000972 | 0.30917 | 1.3433 | 2.5474 | 100.2 | 20.82 | 37.563 | 109.806 | 89.611 | 101.2685 | 0.009231 | 0.94292 | 0.022202 | 23.5458 | 0.084725 | 0.15731 | 0.026077 | 1.7295 | 1.7295 | 0.003702
2019-03-29 14:05 | 1553882721.5426 | 734 | 0.20435 | 0.004437 | 8.9e-05 | 100.102 | 99.0797 | 111.2248 | 21.1097 | 0.003573 | 0.003265 | 233.2852 | 0.26387 | 8.4027 | 0.002736 | 0.004572 | 0.004542 | 5.5e-05 | 0.000981 | 0.28241 | 1.3798 | 2.5474 | 100.2 | 20.82 | 37.568 | 107.7171 | 87.4946 | 100.2597 | 0.009232 | 0.94546 | 0.022213 | 23.494 | 0.091819 | 0.15698 | 0.026123 | 1.7334 | 1.7334 | 0.003686
2019-03-29 14:05 | 1553882722.5426 | 735 | 0.20438 | 0.004439 | 8.9e-05 | 100.1277 | 99.1725 | 111.5929 | 21.1107 | 0.00362 | 0.003303 | 233.2202 | 0.26508 | 8.4008 | 0.002805 | 0.005037 | 0.004936 | 5.6e-05 | 0.000979 | 0.3282 | 1.3482 | 2.5474 | 100.2 | 20.82 | 37.569 | 108.8646 | 89.1429 | 101.1758 | 0.009232 | 0.94315 | 0.022222 | 23.5617 | 0.085742 | 0.15769 | 0.025932 | 1.7214 | 1.7214 | 0.003728
2019-03-29 14:05 | 1553882723.543 | 736 | 0.20437 | 0.004444 | 8.9e-05 | 100.1158 | 99.2223 | 111.4147 | 21.1095 | 0.003575 | 0.003209 | 233.1506 | 0.26442 | 8.4025 | 0.002782 | 0.00493 | 0.004898 | 5.4e-05 | 0.000986 | 0.2891 | 1.3484 | 2.5473 | 100.2 | 20.82 | 37.57 | 107.7672 | 88.6051 | 98.8891 | 0.009233 | 0.94895 | 0.022273 | 23.4709 | 0.08018 | 0.1573 | 0.026003 | 1.73 | 1.73 | 0.003701
2019-03-29 14:05 | 1553882724.544 | 737 | 0.20437 | 0.004449 | 9.1e-05 | 100.2487 | 99.2551 | 111.7323 | 21.1104 | 0.003579 | 0.003231 | 233.0629 | 0.26419 | 8.4023 | 0.002819 | 0.003778 | 0.003588 | 5.5e-05 | 0.000983 | 0.30414 | 1.3403 | 2.5474 | 100.2 | 20.82 | 37.57 | 107.8627 | 89.4799 | 99.431 | 0.009233 | 0.94757 | 0.022288 | 23.5214 | 0.089093 | 0.15717 | 0.025949 | 1.7276 | 1.7276 | 0.003703
2019-03-29 14:05 | 1553882725.5437 | 738 | 0.20433 | 0.004451 | 9.1e-05 | 100.2678 | 99.1616 | 111.8637 | 21.1094 | 0.003618 | 0.003213 | 232.9697 | 0.26408 | 8.402 | 0.002807 | 0.004304 | 0.004314 | 5.2e-05 | 0.000984 | 0.30453 | 1.3799 | 2.5475 | 100.2 | 20.82 | 37.57 | 108.8127 | 89.1946 | 98.9837 | 0.009233 | 0.94871 | 0.022315 | 23.5211 | 0.099233 | 0.1571 | 0.026205 | 1.7467 | 1.7467 | 0.003661
2019-03-29 14:05 | 1553882726.5427 | 739 | 0.20434 | 0.00445 | 9.1e-05 | 100.3079 | 99.2332 | 112.1642 | 21.1103 | 0.003589 | 0.003226 | 232.9088 | 0.26441 | 8.4003 | 0.00281 | 0.004391 | 0.004378 | 5.3e-05 | 0.000988 | 0.32012 | 1.3455 | 2.5474 | 100.2 | 20.82 | 37.569 | 108.0989 | 89.2737 | 99.3085 | 0.009232 | 0.94788 | 0.022335 | 23.563 | 0.096348 | 0.1573 | 0.026143 | 1.7442 | 1.7442 | 0.003671
2019-03-29 14:05 | 1553882727.5431 | 740 | 0.20431 | 0.004453 | 9e-05 | 100.3241 | 99.1691 | 112.2527 | 21.1102 | 0.003592 | 0.003236 | 232.8236 | 0.26395 | 8.4041 | 0.002791 | 0.004107 | 0.004006 | 5.4e-05 | 0.000988 | 0.31332 | 1.3315 | 2.5473 | 100.2 | 20.82 | 37.569 | 108.1683 | 88.808 | 99.5618 | 0.009232 | 0.94723 | 0.022336 | 23.5803 | 0.10357 | 0.15702 | 0.026333 | 1.7569 | 1.7569 | 0.003638
2019-03-29 14:05 | 1553882728.5432 | 741 | 0.20433 | 0.004462 | 9.2e-05 | 100.3388 | 99.2196 | 112.2684 | 21.1095 | 0.003591 | 0.003249 | 232.7618 | 0.26469 | 8.4048 | 0.002846 | 0.004243 | 0.004173 | 5.6e-05 | 0.000982 | 0.28313 | 1.3264 | 2.5475 | 100.2 | 20.82 | 37.567 | 108.1501 | 90.1286 | 99.8694 | 0.009232 | 0.94645 | 0.022328 | 23.5917 | 0.10033 | 0.15746 | 0.026215 | 1.7484 | 1.7484 | 0.003666
2019-03-29 14:05 | 1553882729.5429 | 742 | 0.20434 | 0.00447 | 9.2e-05 | 100.3953 | 99.3861 | 112.3461 | 21.1106 | 0.003614 | 0.003258 | 232.6886 | 0.26469 | 8.4038 | 0.002817 | 0.003929 | 0.003887 | 5.5e-05 | 0.000982 | 0.28528 | 1.3477 | 2.5473 | 100.2 | 20.82 | 37.566 | 108.7138 | 89.4365 | 100.076 | 0.009232 | 0.94593 | 0.02233 | 23.6063 | 0.090367 | 0.15746 | 0.026131 | 1.7429 | 1.7429 | 0.003677
2019-03-29 14:05 | 1553882730.5419 | 743 | 0.20436 | 0.004474 | 9.2e-05 | 100.3878 | 99.4527 | 112.4041 | 21.1097 | 0.003548 | 0.003236 | 232.6185 | 0.26489 | 8.4047 | 0.002865 | 0.004405 | 0.004288 | 5.4e-05 | 0.000981 | 0.28152 | 1.3639 | 2.5474 | 100.2 | 20.82 | 37.566 | 107.1153 | 90.585 | 99.5567 | 0.009232 | 0.94725 | 0.022351 | 23.596 | 0.083713 | 0.15758 | 0.025985 | 1.7348 | 1.7348 | 0.003697
2019-03-29 14:06 | 1553882731.5419 | 744 | 0.20433 | 0.004468 | 9.2e-05 | 100.3896 | 99.3326 | 112.6092 | 21.1096 | 0.003597 | 0.003277 | 232.5486 | 0.26487 | 8.4035 | 0.002783 | 0.00458 | 0.004505 | 5.2e-05 | 0.000979 | 0.30254 | 1.3815 | 2.5474 | 100.2 | 20.82 | 37.564 | 108.2952 | 88.6247 | 100.539 | 0.009231 | 0.94476 | 0.022342 | 23.6486 | 0.094684 | 0.15757 | 0.026162 | 1.7459 | 1.7459 | 0.003673
2019-03-29 14:06 | 1553882732.5433 | 745 | 0.20434 | 0.004467 | 9.2e-05 | 100.3711 | 99.3087 | 112.3245 | 21.1102 | 0.003641 | 0.003219 | 232.4873 | 0.26491 | 8.4019 | 0.00279 | 0.004172 | 0.00418 | 5.2e-05 | 0.000975 | 0.32892 | 1.3648 | 2.5475 | 100.2 | 20.82 | 37.562 | 109.3741 | 88.7798 | 99.1439 | 0.009231 | 0.9483 | 0.022356 | 23.5746 | 0.095182 | 0.15759 | 0.026106 | 1.7433 | 1.7433 | 0.003679
2019-03-29 14:06 | 1553882733.5429 | 746 | 0.20436 | 0.004471 | 9.2e-05 | 100.4384 | 99.3756 | 112.2196 | 21.1097 | 0.003604 | 0.003204 | 232.418 | 0.2644 | 8.4012 | 0.002927 | 0.00397 | 0.003892 | 5.4e-05 | 0.000971 | 0.28982 | 1.3222 | 2.5475 | 100.2 | 20.82 | 37.56 | 108.4624 | 92.101 | 98.7646 | 0.00923 | 0.94927 | 0.022357 | 23.5516 | 0.095159 | 0.15729 | 0.025993 | 1.7358 | 1.7358 | 0.003688
2019-03-29 14:06 | 1553882734.5419 | 747 | 0.20428 | 0.004472 | 9.2e-05 | 100.3929 | 99.3827 | 112.0748 | 21.1101 | 0.00361 | 0.003261 | 232.3361 | 0.26366 | 8.4039 | 0.002867 | 0.004502 | 0.004323 | 5.5e-05 | 0.00097 | 0.30735 | 1.3168 | 2.5473 | 100.2 | 20.82 | 37.561 | 108.6085 | 90.6411 | 100.1518 | 0.00923 | 0.94574 | 0.022301 | 23.5802 | 0.090461 | 0.15685 | 0.026449 | 1.7618 | 1.7618 | 0.003624
2019-03-29 14:06 | 1553882735.5431 | 748 | 0.20422 | 0.004476 | 9.3e-05 | 100.3917 | 99.4063 | 111.7825 | 21.1097 | 0.003626 | 0.003252 | 232.2797 | 0.26444 | 8.4041 | 0.002837 | 0.004358 | 0.004305 | 5.5e-05 | 0.000969 | 0.30254 | 1.3383 | 2.5474 | 100.2 | 20.82 | 37.562 | 109.0023 | 89.9154 | 99.9434 | 0.009231 | 0.94627 | 0.022278 | 23.5429 | 0.088235 | 0.15732 | 0.026821 | 1.7847 | 1.7847 | 0.003588
2019-03-29 14:06 | 1553882736.542 | 749 | 0.20425 | 0.004482 | 9.3e-05 | 100.4345 | 99.3609 | 111.5114 | 21.1097 | 0.003637 | 0.003264 | 232.2179 | 0.26455 | 8.4027 | 0.002987 | 0.004649 | 0.004627 | 5.6e-05 | 0.000971 | 0.31957 | 1.3685 | 2.5475 | 100.2 | 20.82 | 37.565 | 109.2766 | 93.544 | 100.2251 | 0.009231 | 0.94555 | 0.022242 | 23.5232 | 0.096134 | 0.15738 | 0.02665 | 1.7706 | 1.7706 | 0.003618
2019-03-29 14:06 | 1553882737.5432 | 750 | 0.20424 | 0.004492 | 9.2e-05 | 100.4016 | 99.4428 | 111.5518 | 21.1096 | 0.00363 | 0.003223 | 232.1651 | 0.26387 | 8.4019 | 0.002815 | 0.004631 | 0.00457 | 5.7e-05 | 0.000971 | 0.30541 | 1.3521 | 2.5474 | 100.2 | 20.82 | 37.565 | 109.0904 | 89.3966 | 99.2303 | 0.009231 | 0.94808 | 0.022276 | 23.4961 | 0.08583 | 0.15697 | 0.026689 | 1.7758 | 1.7758 | 0.003598
2019-03-29 14:06 | 1553882738.5431 | 751 | 0.20424 | 0.004504 | 9.2e-05 | 100.3926 | 99.4271 | 111.7041 | 21.1097 | 0.003617 | 0.003291 | 232.0889 | 0.26417 | 8.402 | 0.002847 | 0.003781 | 0.003723 | 5.7e-05 | 0.000968 | 0.30055 | 1.3335 | 2.5473 | 100.2 | 20.82 | 37.566 | 108.7871 | 90.1529 | 100.8843 | 0.009232 | 0.94388 | 0.022242 | 23.5643 | 0.086442 | 0.15716 | 0.026661 | 1.7713 | 1.7713 | 0.003611
2019-03-29 14:06 | 1553882739.5421 | 752 | 0.20433 | 0.004513 | 9.2e-05 | 100.3546 | 99.4046 | 111.9567 | 21.1096 | 0.003622 | 0.003253 | 232.0006 | 0.26453 | 8.4005 | 0.002772 | 0.004734 | 0.004682 | 5.4e-05 | 0.000966 | 0.32333 | 1.289 | 2.5474 | 100.2 | 20.82 | 37.568 | 108.9099 | 88.3528 | 99.9544 | 0.009232 | 0.94624 | 0.022295 | 23.5616 | 0.085081 | 0.15737 | 0.026104 | 1.7384 | 1.7384 | 0.003684
2019-03-29 14:06 | 1553882740.5431 | 753 | 0.20428 | 0.004513 | 9.2e-05 | 100.3382 | 99.2438 | 112.0553 | 21.1102 | 0.003551 | 0.003236 | 231.9459 | 0.26415 | 8.4027 | 0.002886 | 0.004637 | 0.004501 | 5.7e-05 | 0.000966 | 0.26056 | 1.3095 | 2.5474 | 100.2 | 20.82 | 37.57 | 107.1915 | 91.1054 | 99.5478 | 0.009233 | 0.94727 | 0.022317 | 23.5591 | 0.09809 | 0.15714 | 0.026413 | 1.7607 | 1.7607 | 0.003632
2019-03-29 14:06 | 1553882741.5435 | 754 | 0.20429 | 0.004518 | 9.2e-05 | 100.2681 | 99.0237 | 112.3489 | 21.1094 | 0.003581 | 0.003186 | 231.8776 | 0.2642 | 8.4008 | 0.002884 | 0.004052 | 0.003965 | 5.9e-05 | 0.000969 | 0.30702 | 1.3098 | 2.5473 | 100.2 | 20.82 | 37.573 | 107.9067 | 91.0507 | 98.3347 | 0.009233 | 0.95036 | 0.022382 | 23.5515 | 0.1117 | 0.15718 | 0.026347 | 1.7614 | 1.7614 | 0.003632
2019-03-29 14:06 | 1553882742.5436 | 755 | 0.20435 | 0.004514 | 9.2e-05 | 100.2292 | 99.0803 | 112.6289 | 21.109 | 0.003578 | 0.003172 | 231.7751 | 0.26392 | 8.4024 | 0.002856 | 0.004678 | 0.004597 | 5.5e-05 | 0.000966 | 0.29983 | 1.3131 | 2.5473 | 100.2 | 20.82 | 37.573 | 107.8337 | 90.389 | 97.9955 | 0.009233 | 0.95123 | 0.022421 | 23.5701 | 0.10312 | 0.15701 | 0.026027 | 1.743 | 1.743 | 0.003666
2019-03-29 14:06 | 1553882743.5413 | 756 | 0.20428 | 0.004516 | 9.3e-05 | 100.1062 | 98.9841 | 112.9024 | 21.1089 | 0.003618 | 0.003271 | 231.695 | 0.26432 | 8.4029 | 0.002899 | 0.003955 | 0.003953 | 5.5e-05 | 0.000964 | 0.29839 | 1.2998 | 2.5474 | 100.2 | 20.82 | 37.574 | 108.806 | 91.416 | 100.3947 | 0.009234 | 0.94512 | 0.022376 | 23.6748 | 0.10082 | 0.15724 | 0.026396 | 1.7642 | 1.7642 | 0.003628
2019-03-29 14:06 | 1553882744.5413 | 757 | 0.20434 | 0.004513 | 9.3e-05 | 100.0828 | 98.7199 | 113.1744 | 21.1087 | 0.003666 | 0.003299 | 231.6361 | 0.26394 | 8.4013 | 0.002833 | 0.004307 | 0.004214 | 5.9e-05 | 0.000961 | 0.30342 | 1.3245 | 2.5474 | 100.2 | 20.82 | 37.576 | 109.9688 | 89.8237 | 101.0655 | 0.009234 | 0.94343 | 0.022382 | 23.7246 | 0.12264 | 0.15702 | 0.026046 | 1.7413 | 1.7413 | 0.00367
2019-03-29 14:06 | 1553882745.5413 | 758 | 0.2043 | 0.004511 | 9.3e-05 | 100.1173 | 98.9435 | 113.3885 | 21.1098 | 0.003573 | 0.003239 | 231.5856 | 0.26277 | 8.4016 | 0.002836 | 0.004757 | 0.004603 | 5.8e-05 | 0.000958 | 0.27455 | 1.32 | 2.5473 | 100.2 | 20.82 | 37.577 | 107.7164 | 89.9099 | 99.632 | 0.009234 | 0.94706 | 0.022447 | 23.7015 | 0.10548 | 0.15632 | 0.026321 | 1.7647 | 1.7647 | 0.003605
2019-03-29 14:06 | 1553882746.5427 | 759 | 0.20434 | 0.004509 | 9.5e-05 | 100.0556 | 98.8449 | 113.7171 | 21.1098 | 0.003568 | 0.003274 | 231.5177 | 0.26375 | 8.4013 | 0.002785 | 0.004069 | 0.003991 | 5.8e-05 | 0.000962 | 0.31116 | 1.3133 | 2.5474 | 100.2 | 20.82 | 37.575 | 107.6021 | 88.6787 | 100.4731 | 0.009234 | 0.94492 | 0.022454 | 23.7626 | 0.1089 | 0.1569 | 0.026073 | 1.7485 | 1.7485 | 0.003652
2019-03-29 14:06 | 1553882747.5415 | 760 | 0.20433 | 0.004503 | 9.5e-05 | 99.9263 | 98.7609 | 113.6659 | 21.1082 | 0.003647 | 0.003234 | 231.4425 | 0.26391 | 8.4006 | 0.002857 | 0.004498 | 0.00445 | 5.6e-05 | 0.00096 | 0.28799 | 1.2994 | 2.5473 | 100.2 | 20.82 | 37.572 | 109.5203 | 90.4007 | 99.5011 | 0.009233 | 0.94739 | 0.022478 | 23.7263 | 0.10493 | 0.157 | 0.026152 | 1.7557 | 1.7557 | 0.003639
2019-03-29 14:06 | 1553882748.5414 | 761 | 0.20432 | 0.004496 | 9.5e-05 | 99.8631 | 98.7213 | 113.4399 | 21.1081 | 0.003603 | 0.003284 | 231.3692 | 0.26344 | 8.4041 | 0.002791 | 0.004557 | 0.00437 | 5.5e-05 | 0.00096 | 0.3178 | 1.3825 | 2.5473 | 100.2 | 20.82 | 37.569 | 108.4418 | 88.8076 | 100.7022 | 0.009232 | 0.94434 | 0.02242 | 23.7409 | 0.10286 | 0.15672 | 0.026209 | 1.7549 | 1.7549 | 0.003635
2019-03-29 14:06 | 1553882749.5418 | 762 | 0.20437 | 0.004494 | 9.6e-05 | 99.7937 | 98.6155 | 113.2667 | 21.1089 | 0.003621 | 0.003221 | 231.3166 | 0.2643 | 8.4014 | 0.002752 | 0.004187 | 0.004111 | 5.5e-05 | 0.000959 | 0.33268 | 1.3659 | 2.5476 | 100.2 | 20.82 | 37.562 | 108.8698 | 87.8811 | 99.1873 | 0.009231 | 0.94819 | 0.022448 | 23.6746 | 0.10623 | 0.15723 | 0.025934 | 1.7387 | 1.7387 | 0.00368
2019-03-29 14:06 | 1553882750.5419 | 763 | 0.20435 | 0.004487 | 9.6e-05 | 99.7335 | 98.3312 | 113.197 | 21.1078 | 0.00359 | 0.003234 | 231.2377 | 0.26366 | 8.4016 | 0.002754 | 0.004472 | 0.004414 | 5.5e-05 | 0.000961 | 0.30718 | 1.3526 | 2.5474 | 100.2 | 20.82 | 37.558 | 108.1209 | 87.9301 | 99.5063 | 0.00923 | 0.94737 | 0.022431 | 23.6775 | 0.12666 | 0.15685 | 0.026022 | 1.7433 | 1.7433 | 0.003662
2019-03-29 14:06 | 1553882751.5416 | 764 | 0.20433 | 0.004482 | 9.5e-05 | 99.6859 | 98.3868 | 112.9083 | 21.108 | 0.003581 | 0.003255 | 231.1807 | 0.26325 | 8.4027 | 0.002761 | 0.00362 | 0.003629 | 5.6e-05 | 0.000964 | 0.31708 | 1.3695 | 2.5474 | 100.2 | 20.82 | 37.541 | 107.8983 | 88.0942 | 100.0172 | 0.009226 | 0.94608 | 0.022387 | 23.6635 | 0.11733 | 0.15661 | 0.026169 | 1.7498 | 1.7498 | 0.003643
2019-03-29 14:06 | 1553882752.5415 | 765 | 0.20436 | 0.004473 | 9.5e-05 | 99.4707 | 98.1647 | 112.7104 | 21.1083 | 0.003646 | 0.003293 | 231.1262 | 0.26315 | 8.4032 | 0.00282 | 0.004107 | 0.004027 | 5.4e-05 | 0.000961 | 0.30901 | 1.3624 | 2.5474 | 100.2 | 20.82 | 37.541 | 109.4738 | 89.5194 | 100.9325 | 0.009226 | 0.94376 | 0.02234 | 23.6717 | 0.11822 | 0.15655 | 0.025996 | 1.7346 | 1.7346 | 0.003673
2019-03-29 14:06 | 1553882753.5416 | 766 | 0.20439 | 0.004473 | 9.5e-05 | 99.3971 | 98.1063 | 112.5228 | 21.1077 | 0.003608 | 0.003284 | 231.0465 | 0.26355 | 8.4007 | 0.00281 | 0.004448 | 0.004334 | 5.5e-05 | 0.00096 | 0.29479 | 1.3406 | 2.5475 | 100.2 | 20.82 | 37.535 | 108.5534 | 89.2611 | 100.7169 | 0.009224 | 0.94431 | 0.022328 | 23.6452 | 0.11691 | 0.15678 | 0.025822 | 1.722 | 1.722 | 0.003706
2019-03-29 14:06 | 1553882754.543 | 767 | 0.20437 | 0.004469 | 9.5e-05 | 99.2211 | 98.2213 | 112.3818 | 21.1088 | 0.003584 | 0.003264 | 230.9732 | 0.26266 | 8.4026 | 0.002762 | 0.004368 | 0.004275 | 5.7e-05 | 0.000962 | 0.29231 | 1.3336 | 2.5473 | 100.2 | 20.82 | 37.523 | 107.991 | 88.1084 | 100.2336 | 0.009221 | 0.94553 | 0.022329 | 23.6151 | 0.090581 | 0.15626 | 0.025962 | 1.7314 | 1.7314 | 0.003673
2019-03-29 14:06 | 1553882755.5416 | 768 | 0.20437 | 0.004463 | 9.5e-05 | 99.0938 | 97.7828 | 112.4642 | 21.1083 | 0.003583 | 0.003269 | 230.9009 | 0.26361 | 8.4011 | 0.002704 | 0.004311 | 0.004161 | 5.5e-05 | 0.000965 | 0.29734 | 1.3378 | 2.5473 | 100.2 | 20.83 | 37.51 | 107.9663 | 86.7109 | 100.3447 | 0.009224 | 0.94525 | 0.022334 | 23.6272 | 0.11913 | 0.15682 | 0.025945 | 1.7307 | 1.7307 | 0.003688
2019-03-29 14:06 | 1553882756.5416 | 769 | 0.20438 | 0.004463 | 9.4e-05 | 98.9092 | 97.5073 | 112.4013 | 21.1077 | 0.003565 | 0.003232 | 230.8214 | 0.26307 | 8.4021 | 0.002745 | 0.003696 | 0.003642 | 5.4e-05 | 0.000958 | 0.30071 | 1.3861 | 2.5474 | 100.2 | 20.83 | 37.51 | 107.5246 | 87.7083 | 99.4465 | 0.009224 | 0.94753 | 0.022354 | 23.5922 | 0.12768 | 0.1565 | 0.025892 | 1.7287 | 1.7287 | 0.003685
2019-03-29 14:06 | 1553882757.5417 | 770 | 0.2044 | 0.004464 | 9.5e-05 | 98.6643 | 97.3872 | 112.4207 | 21.1081 | 0.003622 | 0.003275 | 230.7504 | 0.26412 | 8.4028 | 0.00277 | 0.00462 | 0.004472 | 5.6e-05 | 0.000957 | 0.29717 | 1.3928 | 2.5475 | 100.2 | 20.83 | 37.496 | 108.8989 | 88.2982 | 100.4952 | 0.00922 | 0.94487 | 0.022325 | 23.6274 | 0.11653 | 0.15712 | 0.025758 | 1.7175 | 1.7175 | 0.003723
2019-03-29 14:06 | 1553882758.5419 | 771 | 0.20439 | 0.004457 | 9.5e-05 | 98.5299 | 97.1258 | 112.5714 | 21.1064 | 0.003601 | 0.003216 | 230.6772 | 0.26362 | 8.4004 | 0.002751 | 0.004329 | 0.004319 | 5.5e-05 | 0.000954 | 0.29822 | 1.3617 | 2.5474 | 100.2 | 20.83 | 37.49 | 108.3811 | 87.8428 | 99.0734 | 0.009219 | 0.94848 | 0.022382 | 23.5983 | 0.12838 | 0.15683 | 0.02584 | 1.7274 | 1.7274 | 0.003695
2019-03-29 14:06 | 1553882759.5418 | 772 | 0.2044 | 0.004451 | 9.5e-05 | 98.2818 | 96.7483 | 112.7606 | 21.1065 | 0.003521 | 0.003236 | 230.6099 | 0.26322 | 8.4048 | 0.002768 | 0.004654 | 0.004597 | 5.1e-05 | 0.000953 | 0.29789 | 1.3389 | 2.5474 | 100.2 | 20.83 | 37.479 | 106.4563 | 88.2625 | 99.5355 | 0.009216 | 0.9473 | 0.022387 | 23.6327 | 0.14066 | 0.15659 | 0.025783 | 1.724 | 1.724 | 0.003697
2019-03-29 14:06 | 1553882760.5422 | 773 | 0.20438 | 0.004443 | 9.5e-05 | 98.0998 | 96.6149 | 112.8632 | 21.1062 | 0.003617 | 0.003274 | 230.546 | 0.26448 | 8.4042 | 0.002651 | 0.004738 | 0.004668 | 5.6e-05 | 0.000948 | 0.29247 | 1.3261 | 2.5474 | 100.2 | 20.83 | 37.474 | 108.7773 | 85.4342 | 100.4751 | 0.009215 | 0.94492 | 0.022369 | 23.6733 | 0.13642 | 0.15734 | 0.025913 | 1.7313 | 1.7313 | 0.003699
2019-03-29 14:06 | 1553882761.5421 | 774 | 0.20439 | 0.004442 | 9.5e-05 | 97.9474 | 96.5906 | 113.1247 | 21.1063 | 0.003593 | 0.003306 | 230.4693 | 0.26404 | 8.4005 | 0.002627 | 0.00434 | 0.004217 | 5.4e-05 | 0.000946 | 0.27721 | 1.3454 | 2.5473 | 100.2 | 20.83 | 37.47 | 108.2006 | 84.8662 | 101.2477 | 0.009214 | 0.94297 | 0.022372 | 23.7252 | 0.12477 | 0.15707 | 0.025843 | 1.7268 | 1.7268 | 0.003702
2019-03-29 14:06 | 1553882762.5419 | 775 | 0.20439 | 0.00444 | 9.5e-05 | 97.6562 | 96.2569 | 113.1526 | 21.1062 | 0.003654 | 0.003246 | 230.4041 | 0.26337 | 8.4015 | 0.002689 | 0.004108 | 0.004015 | 5.6e-05 | 0.000941 | 0.27007 | 1.3368 | 2.5473 | 100.2 | 20.83 | 37.462 | 109.6698 | 86.3585 | 99.7871 | 0.009212 | 0.94666 | 0.022419 | 23.6818 | 0.12909 | 0.15668 | 0.025869 | 1.7322 | 1.7322 | 0.003681
2019-03-29 14:06 | 1553882763.542 | 776 | 0.20439 | 0.004444 | 9.4e-05 | 97.3637 | 95.9156 | 113.0676 | 21.1063 | 0.003569 | 0.003267 | 230.345 | 0.26287 | 8.4045 | 0.002704 | 0.003505 | 0.003479 | 5.5e-05 | 0.000938 | 0.29607 | 1.3306 | 2.5474 | 100.2 | 20.83 | 37.458 | 107.6233 | 86.7169 | 100.2949 | 0.009211 | 0.94537 | 0.022395 | 23.689 | 0.13403 | 0.15638 | 0.025881 | 1.7312 | 1.7312 | 0.003677
2019-03-29 14:06 | 1553882764.5409 | 777 | 0.2044 | 0.004451 | 9.4e-05 | 97.1356 | 95.7265 | 113.0867 | 21.106 | 0.00361 | 0.003283 | 230.2926 | 0.2645 | 8.4031 | 0.002645 | 0.004716 | 0.004584 | 5.4e-05 | 0.000935 | 0.28263 | 1.307 | 2.5474 | 100.2 | 20.83 | 37.455 | 108.5974 | 85.2967 | 100.6736 | 0.00921 | 0.94442 | 0.022385 | 23.703 | 0.1307 | 0.15735 | 0.025789 | 1.7243 | 1.7243 | 0.003714
2019-03-29 14:06 | 1553882765.5409 | 778 | 0.2044 | 0.004452 | 9.4e-05 | 96.9303 | 95.363 | 113.0536 | 21.1061 | 0.003617 | 0.003282 | 230.2014 | 0.26343 | 8.4023 | 0.002715 | 0.004729 | 0.00464 | 5.6e-05 | 0.000929 | 0.29856 | 1.3359 | 2.5474 | 100.19 | 20.83 | 37.449 | 108.7709 | 86.9751 | 100.6523 | 0.00921 | 0.94447 | 0.022383 | 23.6988 | 0.1458 | 0.15671 | 0.025805 | 1.7252 | 1.7252 | 0.003697
2019-03-29 14:06 | 1553882766.5421 | 779 | 0.20437 | 0.004448 | 9.4e-05 | 96.6184 | 95.134 | 113.1512 | 21.107 | 0.003573 | 0.003282 | 230.1124 | 0.26435 | 8.4005 | 0.002703 | 0.004046 | 0.004012 | 5.4e-05 | 0.000931 | 0.30502 | 1.3795 | 2.5474 | 100.19 | 20.83 | 37.447 | 107.7208 | 86.6839 | 100.669 | 0.009209 | 0.94443 | 0.022392 | 23.7096 | 0.13848 | 0.15726 | 0.025978 | 1.7375 | 1.7375 | 0.003684
2019-03-29 14:06 | 1553882767.5424 | 780 | 0.20439 | 0.004438 | 9.4e-05 | 96.3469 | 94.6818 | 113.2136 | 21.1054 | 0.003587 | 0.003283 | 230.0335 | 0.26425 | 8.403 | 0.002697 | 0.003892 | 0.003797 | 5.6e-05 | 0.000936 | 0.31924 | 1.402 | 2.5476 | 100.2 | 20.83 | 37.441 | 108.054 | 86.5512 | 100.6713 | 0.009207 | 0.94442 | 0.022398 | 23.7162 | 0.15593 | 0.1572 | 0.02584 | 1.7287 | 1.7287 | 0.003701
2019-03-29 14:06 | 1553882768.542 | 781 | 0.20439 | 0.00443 | 9.4e-05 | 96.044 | 94.3197 | 113.3416 | 21.1055 | 0.003522 | 0.003228 | 229.9487 | 0.26372 | 8.3994 | 0.002675 | 0.003816 | 0.003818 | 5.7e-05 | 0.000943 | 0.32908 | 1.3807 | 2.5476 | 100.2 | 20.83 | 37.439 | 106.4798 | 86.0167 | 99.3409 | 0.009206 | 0.9478 | 0.022451 | 23.6873 | 0.16204 | 0.15689 | 0.025893 | 1.7363 | 1.7363 | 0.003678
2019-03-29 14:06 | 1553882769.541 | 782 | 0.20436 | 0.004421 | 9.4e-05 | 95.7153 | 93.9907 | 113.3101 | 21.1054 | 0.003636 | 0.003235 | 229.8565 | 0.26404 | 8.4026 | 0.002678 | 0.004116 | 0.004026 | 5.5e-05 | 0.00095 | 0.27489 | 1.3807 | 2.5474 | 100.2 | 20.83 | 37.436 | 109.2321 | 86.0946 | 99.5258 | 0.009205 | 0.94733 | 0.022442 | 23.6899 | 0.16263 | 0.15708 | 0.026039 | 1.7454 | 1.7454 | 0.003663
2019-03-29 14:06 | 1553882770.5414 | 783 | 0.20433 | 0.00442 | 9.4e-05 | 95.4552 | 93.689 | 113.1599 | 21.1055 | 0.003592 | 0.003327 | 229.7712 | 0.26381 | 8.4036 | 0.002689 | 0.004334 | 0.004212 | 5.4e-05 | 0.000951 | 0.26897 | 1.3885 | 2.5475 | 100.2 | 20.83 | 37.434 | 108.1762 | 86.3569 | 101.7508 | 0.009205 | 0.9417 | 0.022361 | 23.7448 | 0.16705 | 0.15694 | 0.026226 | 1.7516 | 1.7516 | 0.003647
2019-03-29 14:06 | 1553882771.5426 | 784 | 0.20434 | 0.004419 | 9.5e-05 | 95.1301 | 93.3573 | 112.885 | 21.1056 | 0.003626 | 0.003259 | 229.6867 | 0.2649 | 8.4013 | 0.002731 | 0.004125 | 0.004054 | 5.4e-05 | 0.000945 | 0.29966 | 1.4067 | 2.5474 | 100.2 | 20.83 | 37.431 | 109.0054 | 87.3674 | 100.1043 | 0.009204 | 0.94586 | 0.022383 | 23.6638 | 0.16826 | 0.15759 | 0.026195 | 1.7512 | 1.7512 | 0.003663
2019-03-29 14:06 | 1553882772.5412 | 785 | 0.20431 | 0.004431 | 9.5e-05 | 94.7535 | 92.7795 | 112.786 | 21.1065 | 0.003638 | 0.003281 | 229.5973 | 0.26499 | 8.4034 | 0.002723 | 0.003849 | 0.003759 | 5.5e-05 | 0.000945 | 0.29645 | 1.3903 | 2.5474 | 100.2 | 20.84 | 37.428 | 109.2922 | 87.1739 | 100.6245 | 0.009209 | 0.94454 | 0.022357 | 23.6699 | 0.18832 | 0.15764 | 0.026363 | 1.7604 | 1.7604 | 0.003645
2019-03-29 14:06 | 1553882773.5412 | 786 | 0.20433 | 0.004439 | 9.5e-05 | 94.4207 | 92.4914 | 112.7017 | 21.1063 | 0.0036 | 0.00326 | 229.525 | 0.26379 | 8.4016 | 0.002746 | 0.003901 | 0.003812 | 5.8e-05 | 0.000941 | 0.31885 | 1.3654 | 2.5474 | 100.2 | 20.84 | 37.426 | 108.3638 | 87.7178 | 100.126 | 0.009209 | 0.9458 | 0.022364 | 23.6453 | 0.18465 | 0.15693 | 0.026251 | 1.7535 | 1.7535 | 0.003643
2019-03-29 14:06 | 1553882774.5416 | 787 | 0.20432 | 0.004443 | 9.6e-05 | 94.0321 | 91.9891 | 112.7503 | 21.1061 | 0.003588 | 0.003229 | 229.4654 | 0.26452 | 8.4029 | 0.00269 | 0.004608 | 0.004485 | 5.8e-05 | 0.000939 | 0.30685 | 1.3526 | 2.5473 | 100.2 | 20.83 | 37.424 | 108.086 | 86.3685 | 99.3782 | 0.009202 | 0.9477 | 0.022391 | 23.6267 | 0.19648 | 0.15736 | 0.026263 | 1.7564 | 1.7564 | 0.003647
2019-03-29 14:06 | 1553882775.5416 | 788 | 0.20436 | 0.004442 | 9.5e-05 | 93.6227 | 91.6324 | 112.4966 | 21.1058 | 0.003612 | 0.003245 | 229.4071 | 0.26397 | 8.4027 | 0.002743 | 0.004665 | 0.004567 | 5.5e-05 | 0.000943 | 0.29319 | 1.3538 | 2.5474 | 100.19 | 20.84 | 37.421 | 108.6603 | 87.6494 | 99.7655 | 0.009208 | 0.94672 | 0.022354 | 23.6123 | 0.1922 | 0.15704 | 0.026022 | 1.7374 | 1.7374 | 0.003679
2019-03-29 14:06 | 1553882776.5413 | 789 | 0.20439 | 0.004445 | 9.5e-05 | 93.3176 | 91.2242 | 112.2769 | 21.1054 | 0.003567 | 0.003261 | 229.3457 | 0.26395 | 8.4019 | 0.002772 | 0.004166 | 0.004115 | 5.4e-05 | 0.000949 | 0.30486 | 1.3388 | 2.5475 | 100.19 | 20.84 | 37.42 | 107.5655 | 88.3519 | 100.1597 | 0.009208 | 0.94572 | 0.022321 | 23.6017 | 0.20294 | 0.15702 | 0.025885 | 1.7257 | 1.7257 | 0.003703
2019-03-29 14:06 | 1553882777.5413 | 790 | 0.20439 | 0.004447 | 9.4e-05 | 92.9845 | 90.9614 | 112.0194 | 21.1064 | 0.003665 | 0.003299 | 229.2792 | 0.26419 | 8.4008 | 0.002734 | 0.004505 | 0.004459 | 5.7e-05 | 0.000946 | 0.28512 | 1.3349 | 2.5473 | 100.19 | 20.84 | 37.416 | 109.9434 | 87.4447 | 101.06 | 0.009207 | 0.94344 | 0.022268 | 23.603 | 0.19675 | 0.15716 | 0.025834 | 1.7183 | 1.7183 | 0.003723
2019-03-29 14:06 | 1553882778.5403 | 791 | 0.20437 | 0.004433 | 9.3e-05 | 92.5564 | 90.5331 | 111.6008 | 21.105 | 0.003629 | 0.003288 | 229.2375 | 0.26436 | 8.4032 | 0.002748 | 0.004575 | 0.004463 | 5.4e-05 | 0.000946 | 0.27383 | 1.3363 | 2.5474 | 100.19 | 20.83 | 37.413 | 109.0811 | 87.7642 | 100.7957 | 0.009201 | 0.94411 | 0.022234 | 23.5506 | 0.1977 | 0.15727 | 0.026006 | 1.7272 | 1.7272 | 0.003706
2019-03-29 14:06 | 1553882779.5418 | 792 | 0.2044 | 0.004424 | 9.2e-05 | 92.2616 | 90.0546 | 111.449 | 21.1052 | 0.003596 | 0.003277 | 229.1865 | 0.264 | 8.4026 | 0.002656 | 0.004712 | 0.004637 | 5.8e-05 | 0.000944 | 0.30989 | 1.353 | 2.5475 | 100.19 | 20.84 | 37.411 | 108.2673 | 85.5567 | 100.5441 | 0.009206 | 0.94474 | 0.022227 | 23.5266 | 0.21656 | 0.15705 | 0.025818 | 1.7141 | 1.7141 | 0.003729
2019-03-29 14:06 | 1553882780.5404 | 793 | 0.20437 | 0.004423 | 9.2e-05 | 91.8853 | 89.6349 | 111.7032 | 21.1049 | 0.003585 | 0.003275 | 229.1185 | 0.26418 | 8.4011 | 0.002634 | 0.003999 | 0.00393 | 5.7e-05 | 0.000942 | 0.28279 | 1.3432 | 2.5473 | 100.19 | 20.84 | 37.409 | 107.9889 | 85.0317 | 100.4983 | 0.009205 | 0.94486 | 0.022253 | 23.552 | 0.2218 | 0.15716 | 0.025994 | 1.7279 | 1.7279 | 0.003702
2019-03-29 14:06 | 1553882781.5414 | 794 | 0.20438 | 0.00442 | 9.2e-05 | 91.5842 | 89.5857 | 111.7889 | 21.1051 | 0.003601 | 0.003316 | 229.0607 | 0.26354 | 8.4042 | 0.00273 | 0.003889 | 0.003772 | 5.3e-05 | 0.000934 | 0.32892 | 1.3602 | 2.5475 | 100.19 | 20.84 | 37.406 | 108.3886 | 87.3317 | 101.4694 | 0.009205 | 0.94241 | 0.022233 | 23.5916 | 0.19734 | 0.15678 | 0.02596 | 1.7241 | 1.7241 | 0.003701
2019-03-29 14:06 | 1553882782.5404 | 795 | 0.2044 | 0.004413 | 9.2e-05 | 91.2208 | 89.1359 | 111.7081 | 21.1049 | 0.003588 | 0.003312 | 229.0173 | 0.26356 | 8.4016 | 0.002689 | 0.003767 | 0.003733 | 5.4e-05 | 0.000929 | 0.29264 | 1.3479 | 2.5475 | 100.19 | 20.84 | 37.404 | 108.0849 | 86.3486 | 101.3923 | 0.009204 | 0.94261 | 0.022227 | 23.5807 | 0.20681 | 0.15679 | 0.025824 | 1.7146 | 1.7146 | 0.003722
2019-03-29 14:06 | 1553882783.5417 | 796 | 0.20447 | 0.004412 | 9.2e-05 | 90.782 | 88.5106 | 112.0459 | 21.1053 | 0.003617 | 0.003326 | 228.9807 | 0.26348 | 8.4017 | 0.002685 | 0.003897 | 0.003808 | 5.6e-05 | 0.000929 | 0.28672 | 1.369 | 2.5473 | 100.19 | 20.84 | 37.401 | 108.7866 | 86.2607 | 101.7169 | 0.009203 | 0.94179 | 0.022251 | 23.6266 | 0.22664 | 0.15675 | 0.025413 | 1.6891 | 1.6891 | 0.003777
2019-03-29 14:06 | 1553882784.5406 | 797 | 0.20447 | 0.004412 | 9.2e-05 | 90.5746 | 88.2567 | 112.1041 | 21.1037 | 0.003588 | 0.003286 | 228.8925 | 0.26447 | 8.4025 | 0.002787 | 0.004398 | 0.004396 | 5.7e-05 | 0.000932 | 0.27975 | 1.3616 | 2.5473 | 100.19 | 20.84 | 37.398 | 108.0713 | 88.7063 | 100.7475 | 0.009203 | 0.94423 | 0.022286 | 23.6021 | 0.23187 | 0.15733 | 0.025439 | 1.6935 | 1.6935 | 0.003781
2019-03-29 14:06 | 1553882785.5409 | 798 | 0.2045 | 0.004409 | 9.2e-05 | 90.3174 | 87.9243 | 112.3426 | 21.1044 | 0.003601 | 0.003314 | 228.8322 | 0.26347 | 8.4019 | 0.002721 | 0.003959 | 0.003761 | 6.1e-05 | 0.000936 | 0.27903 | 1.3906 | 2.5474 | 100.19 | 20.84 | 37.396 | 108.3957 | 87.1162 | 101.4359 | 0.009202 | 0.94249 | 0.022289 | 23.6489 | 0.2402 | 0.15674 | 0.025284 | 1.6834 | 1.6834 | 0.003789
2019-03-29 14:06 | 1553882786.5417 | 799 | 0.20451 | 0.004401 | 9.1e-05 | 90.2621 | 87.8298 | 112.3395 | 21.1054 | 0.003649 | 0.003256 | 228.7879 | 0.26355 | 8.4023 | 0.002688 | 0.004332 | 0.004236 | 5.6e-05 | 0.000937 | 0.30502 | 1.3751 | 2.5475 | 100.19 | 20.84 | 37.393 | 109.5553 | 86.3227 | 100.0304 | 0.009201 | 0.94605 | 0.022331 | 23.6042 | 0.24433 | 0.15678 | 0.025189 | 1.6802 | 1.6802 | 0.003798
2019-03-29 14:06 | 1553882787.5418 | 800 | 0.20451 | 0.004386 | 9.1e-05 | 90.4198 | 88.1066 | 112.2139 | 21.105 | 0.003579 | 0.003285 | 228.739 | 0.26326 | 8.4016 | 0.002694 | 0.004676 | 0.004707 | 5.6e-05 | 0.000937 | 0.29839 | 1.4053 | 2.5475 | 100.19 | 20.84 | 37.391 | 107.8652 | 86.4782 | 100.7362 | 0.009201 | 0.94426 | 0.022297 | 23.6133 | 0.23181 | 0.15661 | 0.025225 | 1.6801 | 1.6801 | 0.003794
2019-03-29 14:06 | 1553882788.5407 | 801 | 0.20453 | 0.004372 | 9.1e-05 | 90.7093 | 88.3765 | 112.2431 | 21.1048 | 0.003597 | 0.003295 | 228.6574 | 0.26346 | 8.4032 | 0.002734 | 0.004026 | 0.00398 | 5.4e-05 | 0.000945 | 0.30071 | 1.3655 | 2.5475 | 100.19 | 20.84 | 37.385 | 108.2803 | 87.4338 | 100.9651 | 0.009199 | 0.94368 | 0.022293 | 23.6236 | 0.23303 | 0.15673 | 0.025149 | 1.6748 | 1.6748 | 0.003809
2019-03-29 14:06 | 1553882789.5417 | 802 | 0.20455 | 0.004355 | 9e-05 | 91.1226 | 88.9338 | 112.3487 | 21.1047 | 0.003589 | 0.003289 | 228.587 | 0.26343 | 8.4024 | 0.002695 | 0.003935 | 0.003755 | 5.4e-05 | 0.00094 | 0.31188 | 1.333 | 2.5473 | 100.19 | 20.84 | 37.383 | 108.0973 | 86.5084 | 100.8339 | 0.009199 | 0.94401 | 0.022308 | 23.6306 | 0.21747 | 0.15671 | 0.02504 | 1.6686 | 1.6686 | 0.003822
2019-03-29 14:06 | 1553882790.5406 | 803 | 0.2045 | 0.004343 | 9e-05 | 91.4004 | 89.2605 | 112.1651 | 21.105 | 0.003631 | 0.003295 | 228.5059 | 0.2638 | 8.4022 | 0.002779 | 0.004184 | 0.00413 | 5.5e-05 | 0.000942 | 0.2995 | 1.3461 | 2.5474 | 100.19 | 20.84 | 37.381 | 109.1151 | 88.5203 | 100.9709 | 0.009199 | 0.94367 | 0.022285 | 23.6156 | 0.21191 | 0.15693 | 0.025307 | 1.6847 | 1.6847 | 0.003791
2019-03-29 14:06 | 1553882791.5407 | 804 | 0.20451 | 0.004333 | 9.1e-05 | 91.8452 | 89.8075 | 112.1292 | 21.1041 | 0.003627 | 0.003279 | 228.4223 | 0.26392 | 8.402 | 0.002711 | 0.004001 | 0.003928 | 5.4e-05 | 0.00094 | 0.30038 | 1.3176 | 2.5474 | 100.2 | 20.84 | 37.378 | 109.0173 | 86.8856 | 100.59 | 0.009197 | 0.94463 | 0.022293 | 23.5998 | 0.20068 | 0.15701 | 0.02529 | 1.6841 | 1.6841 | 0.003794
2019-03-29 14:07 | 1553882792.5412 | 805 | 0.20451 | 0.004324 | 9e-05 | 92.1984 | 90.3181 | 112.4067 | 21.104 | 0.00369 | 0.003296 | 228.3508 | 0.26272 | 8.4026 | 0.002744 | 0.004588 | 0.004477 | 5.5e-05 | 0.000944 | 0.29839 | 1.3619 | 2.5474 | 100.2 | 20.84 | 37.37 | 110.5581 | 87.6759 | 100.9982 | 0.009195 | 0.9436 | 0.022308 | 23.6419 | 0.1843 | 0.15629 | 0.025289 | 1.6853 | 1.6853 | 0.003774
2019-03-29 14:07 | 1553882793.5398 | 806 | 0.20451 | 0.004321 | 8.9e-05 | 92.5995 | 90.6341 | 112.2076 | 21.1041 | 0.003614 | 0.003268 | 228.278 | 0.26413 | 8.4022 | 0.002739 | 0.004051 | 0.003976 | 5.6e-05 | 0.000939 | 0.31581 | 1.3492 | 2.5476 | 100.2 | 20.84 | 37.36 | 108.7012 | 87.5696 | 100.3249 | 0.009192 | 0.9453 | 0.022309 | 23.5997 | 0.19189 | 0.15713 | 0.025297 | 1.6859 | 1.6859 | 0.003793
2019-03-29 14:07 | 1553882794.5399 | 807 | 0.20458 | 0.004326 | 8.9e-05 | 92.9947 | 91.055 | 112.0884 | 21.1025 | 0.003623 | 0.00329 | 228.2186 | 0.26412 | 8.4028 | 0.002759 | 0.004542 | 0.00441 | 5.3e-05 | 0.000938 | 0.32604 | 1.3481 | 2.5474 | 100.2 | 20.84 | 37.36 | 108.9243 | 88.0383 | 100.86 | 0.009192 | 0.94395 | 0.022281 | 23.604 | 0.18854 | 0.15712 | 0.024899 | 1.6574 | 1.6574 | 0.003859
2019-03-29 14:07 | 1553882795.5409 | 808 | 0.20462 | 0.004317 | 8.8e-05 | 93.2721 | 91.6307 | 112.3151 | 21.1035 | 0.003619 | 0.003312 | 228.1432 | 0.26424 | 8.4024 | 0.002796 | 0.003936 | 0.003912 | 5.4e-05 | 0.000938 | 0.29175 | 1.3373 | 2.5475 | 100.2 | 20.84 | 37.356 | 108.8291 | 88.9264 | 101.3846 | 0.009191 | 0.94262 | 0.022288 | 23.6444 | 0.15881 | 0.1572 | 0.024659 | 1.6419 | 1.6419 | 0.003897
2019-03-29 14:07 | 1553882796.5413 | 809 | 0.20467 | 0.00431 | 8.8e-05 | 93.6544 | 91.8556 | 112.585 | 21.1034 | 0.003578 | 0.003275 | 228.066 | 0.26388 | 8.404 | 0.002684 | 0.00415 | 0.004153 | 5.2e-05 | 0.000941 | 0.28031 | 1.3263 | 2.5474 | 100.2 | 20.85 | 37.353 | 107.8243 | 86.2276 | 100.4952 | 0.009196 | 0.94487 | 0.022341 | 23.6447 | 0.17347 | 0.15699 | 0.024393 | 1.6281 | 1.6281 | 0.003924
2019-03-29 14:07 | 1553882797.5415 | 810 | 0.20471 | 0.004298 | 8.7e-05 | 93.9717 | 92.1263 | 112.7376 | 21.1024 | 0.003613 | 0.003283 | 228.0007 | 0.26422 | 8.4027 | 0.002745 | 0.004328 | 0.004246 | 5.4e-05 | 0.000939 | 0.29407 | 1.3681 | 2.5475 | 100.2 | 20.85 | 37.345 | 108.6774 | 87.6962 | 100.6848 | 0.009194 | 0.94439 | 0.022351 | 23.6667 | 0.1774 | 0.15719 | 0.024163 | 1.6134 | 1.6134 | 0.003965
2019-03-29 14:07 | 1553882798.541 | 811 | 0.20468 | 0.004284 | 8.7e-05 | 94.2875 | 92.4877 | 112.8566 | 21.1026 | 0.0036 | 0.00326 | 227.9251 | 0.26419 | 8.4033 | 0.002769 | 0.004099 | 0.004046 | 5.3e-05 | 0.000943 | 0.2943 | 1.3621 | 2.5476 | 100.2 | 20.85 | 37.343 | 108.3742 | 88.275 | 100.1185 | 0.009194 | 0.94582 | 0.022379 | 23.6613 | 0.17239 | 0.15716 | 0.024311 | 1.6254 | 1.6254 | 0.003935
2019-03-29 14:07 | 1553882799.541 | 812 | 0.20469 | 0.00427 | 8.7e-05 | 94.5727 | 93.1433 | 112.8815 | 21.1033 | 0.003548 | 0.003253 | 227.8779 | 0.26399 | 8.4026 | 0.002747 | 0.004306 | 0.004165 | 5.3e-05 | 0.000942 | 0.29662 | 1.3183 | 2.5474 | 100.2 | 20.85 | 37.338 | 107.1083 | 87.7562 | 99.9567 | 0.009193 | 0.94623 | 0.022387 | 23.6588 | 0.13623 | 0.15705 | 0.024289 | 1.6245 | 1.6245 | 0.003935
2019-03-29 14:07 | 1553882800.5401 | 813 | 0.20474 | 0.004255 | 8.7e-05 | 94.9451 | 93.4102 | 112.7475 | 21.1035 | 0.00363 | 0.003292 | 227.8169 | 0.26401 | 8.4029 | 0.002704 | 0.004481 | 0.004499 | 5.5e-05 | 0.00094 | 0.30414 | 1.3415 | 2.5474 | 100.2 | 20.85 | 37.336 | 109.0922 | 86.7079 | 100.8949 | 0.009192 | 0.94386 | 0.022345 | 23.6744 | 0.14578 | 0.15706 | 0.024002 | 1.6023 | 1.6023 | 0.00399
2019-03-29 14:07 | 1553882801.5412 | 814 | 0.20476 | 0.004236 | 8.7e-05 | 95.3124 | 93.8132 | 112.6186 | 21.1031 | 0.003661 | 0.0033 | 227.7691 | 0.26409 | 8.4019 | 0.002726 | 0.004625 | 0.004492 | 5.5e-05 | 0.000934 | 0.31598 | 1.3675 | 2.5474 | 100.2 | 20.85 | 37.335 | 109.8418 | 87.2454 | 101.0891 | 0.009192 | 0.94337 | 0.022327 | 23.667 | 0.1418 | 0.1571 | 0.023908 | 1.5947 | 1.5947 | 0.00401
2019-03-29 14:07 | 1553882802.5412 | 815 | 0.20477 | 0.004221 | 8.8e-05 | 95.5456 | 94.2108 | 112.463 | 21.1025 | 0.003621 | 0.003295 | 227.7117 | 0.26385 | 8.4004 | 0.002666 | 0.003983 | 0.003913 | 5.2e-05 | 0.000935 | 0.28694 | 1.375 | 2.5474 | 100.2 | 20.85 | 37.333 | 108.869 | 85.8031 | 100.9706 | 0.009191 | 0.94367 | 0.022315 | 23.6469 | 0.12583 | 0.15697 | 0.023885 | 1.5923 | 1.5923 | 0.004012
2019-03-29 14:07 | 1553882803.5402 | 816 | 0.20479 | 0.0042 | 8.7e-05 | 95.9542 | 94.5183 | 112.6073 | 21.103 | 0.003596 | 0.003294 | 227.6367 | 0.26513 | 8.403 | 0.002776 | 0.003799 | 0.003657 | 5.3e-05 | 0.000938 | 0.30558 | 1.346 | 2.5474 | 100.2 | 20.85 | 37.332 | 108.2738 | 88.4393 | 100.9578 | 0.009191 | 0.9437 | 0.02233 | 23.6617 | 0.13485 | 0.15773 | 0.023806 | 1.5881 | 1.5881 | 0.004042
2019-03-29 14:07 | 1553882804.5402 | 817 | 0.20481 | 0.004189 | 8.8e-05 | 96.1873 | 95.0023 | 112.6989 | 21.1024 | 0.003613 | 0.003296 | 227.552 | 0.26359 | 8.4039 | 0.002807 | 0.004606 | 0.004553 | 5.2e-05 | 0.000939 | 0.30646 | 1.3462 | 2.5474 | 100.2 | 20.85 | 37.331 | 108.6767 | 89.1996 | 100.9904 | 0.009191 | 0.94362 | 0.022338 | 23.6723 | 0.11088 | 0.15681 | 0.023669 | 1.5795 | 1.5795 | 0.004041
2019-03-29 14:07 | 1553882805.5403 | 818 | 0.20486 | 0.004175 | 8.7e-05 | 96.4846 | 95.2839 | 112.5261 | 21.1025 | 0.003593 | 0.003291 | 227.496 | 0.26338 | 8.4039 | 0.002839 | 0.004878 | 0.00474 | 5.3e-05 | 0.000946 | 0.32947 | 1.3781 | 2.5475 | 100.2 | 20.85 | 37.328 | 108.1913 | 89.9701 | 100.879 | 0.00919 | 0.9439 | 0.022324 | 23.6507 | 0.11201 | 0.15668 | 0.023379 | 1.5592 | 1.5592 | 0.00409
2019-03-29 14:07 | 1553882806.5403 | 819 | 0.20487 | 0.004158 | 8.8e-05 | 96.8084 | 95.5361 | 112.4778 | 21.1029 | 0.003584 | 0.003238 | 227.4385 | 0.2642 | 8.4029 | 0.002761 | 0.004534 | 0.004512 | 5e-05 | 0.000939 | 0.30901 | 1.3374 | 2.5475 | 100.2 | 20.85 | 37.326 | 107.9712 | 88.0924 | 99.6026 | 0.00919 | 0.94713 | 0.022357 | 23.6052 | 0.11834 | 0.15718 | 0.023333 | 1.5584 | 1.5584 | 0.004105
2019-03-29 14:07 | 1553882807.5393 | 820 | 0.20486 | 0.004133 | 8.7e-05 | 97.0985 | 95.9958 | 112.547 | 21.1028 | 0.003603 | 0.003262 | 227.3784 | 0.264 | 8.4012 | 0.002695 | 0.00453 | 0.004433 | 5.4e-05 | 0.000938 | 0.30469 | 1.3704 | 2.5475 | 100.2 | 20.85 | 37.324 | 108.4414 | 86.494 | 100.1823 | 0.009189 | 0.94566 | 0.022347 | 23.6308 | 0.10216 | 0.15706 | 0.023441 | 1.5649 | 1.5649 | 0.004085
2019-03-29 14:07 | 1553882808.5394 | 821 | 0.20486 | 0.004113 | 8.7e-05 | 97.3697 | 96.3613 | 112.9275 | 21.1032 | 0.003609 | 0.003251 | 227.327 | 0.26361 | 8.4013 | 0.002762 | 0.004548 | 0.004504 | 4.9e-05 | 0.000938 | 0.30287 | 1.3683 | 2.5475 | 100.2 | 20.85 | 37.322 | 108.5915 | 88.1201 | 99.8964 | 0.009189 | 0.94639 | 0.022393 | 23.6617 | 0.093115 | 0.15682 | 0.02348 | 1.5708 | 1.5708 | 0.004063
2019-03-29 14:07 | 1553882809.5404 | 822 | 0.20487 | 0.004096 | 8.7e-05 | 97.6382 | 96.536 | 113.0331 | 21.1018 | 0.003557 | 0.003324 | 227.2372 | 0.26386 | 8.4033 | 0.002851 | 0.003673 | 0.00373 | 5.3e-05 | 0.000937 | 0.28313 | 1.3462 | 2.5473 | 100.2 | 20.85 | 37.32 | 107.3163 | 90.2437 | 101.6758 | 0.009188 | 0.94189 | 0.02235 | 23.7291 | 0.10154 | 0.15697 | 0.02341 | 1.5631 | 1.5631 | 0.004087
2019-03-29 14:07 | 1553882810.5394 | 823 | 0.20489 | 0.004079 | 8.7e-05 | 97.8789 | 96.8633 | 113.3465 | 21.1021 | 0.003592 | 0.003204 | 227.1858 | 0.26361 | 8.4025 | 0.002787 | 0.004475 | 0.004344 | 5.1e-05 | 0.000934 | 0.28926 | 1.3347 | 2.5474 | 100.2 | 20.84 | 37.315 | 108.1705 | 88.7096 | 98.7596 | 0.009181 | 0.94928 | 0.022469 | 23.6694 | 0.093287 | 0.15682 | 0.02334 | 1.5667 | 1.5667 | 0.004074
2019-03-29 14:07 | 1553882811.5393 | 824 | 0.20497 | 0.004067 | 8.8e-05 | 98.027 | 97.3067 | 113.5922 | 21.1034 | 0.003569 | 0.003294 | 227.1373 | 0.26359 | 8.4039 | 0.002788 | 0.004084 | 0.003944 | 5e-05 | 0.000934 | 0.30486 | 1.357 | 2.5475 | 100.2 | 20.84 | 37.312 | 107.6211 | 88.7286 | 100.9354 | 0.009181 | 0.94376 | 0.022428 | 23.7642 | 0.065972 | 0.15681 | 0.022836 | 1.5301 | 1.5301 | 0.004171
2019-03-29 14:07 | 1553882812.5408 | 825 | 0.20501 | 0.004045 | 8.7e-05 | 98.2916 | 97.4343 | 113.6563 | 21.1029 | 0.003516 | 0.003261 | 227.0633 | 0.26448 | 8.4008 | 0.002766 | 0.003909 | 0.003868 | 5.2e-05 | 0.000934 | 0.30934 | 1.3685 | 2.5475 | 100.2 | 20.84 | 37.305 | 106.3143 | 88.1963 | 100.1549 | 0.009179 | 0.94573 | 0.022457 | 23.7461 | 0.078359 | 0.15734 | 0.022661 | 1.5203 | 1.5203 | 0.004212
2019-03-29 14:07 | 1553882813.5395 | 826 | 0.20504 | 0.00402 | 8.7e-05 | 98.5506 | 97.605 | 113.5993 | 21.104 | 0.003578 | 0.003239 | 226.9991 | 0.26345 | 8.4023 | 0.002773 | 0.003852 | 0.003709 | 5e-05 | 0.000939 | 0.28982 | 1.3416 | 2.5474 | 100.2 | 20.84 | 37.3 | 107.8327 | 88.3787 | 99.6194 | 0.009178 | 0.94709 | 0.022468 | 23.7231 | 0.086238 | 0.15673 | 0.022497 | 1.51 | 1.51 | 0.004224
2019-03-29 14:07 | 1553882814.5405 | 827 | 0.2051 | 0.004002 | 8.7e-05 | 98.6969 | 97.6855 | 113.5162 | 21.1022 | 0.003568 | 0.003287 | 226.9519 | 0.26392 | 8.401 | 0.00276 | 0.004509 | 0.004381 | 4.5e-05 | 0.000936 | 0.30702 | 1.3658 | 2.5474 | 100.19 | 20.84 | 37.297 | 107.5806 | 88.069 | 100.7734 | 0.009178 | 0.94417 | 0.022425 | 23.7511 | 0.092135 | 0.15701 | 0.022179 | 1.4858 | 1.4858 | 0.004301
2019-03-29 14:07 | 1553882815.5399 | 828 | 0.20517 | 0.00398 | 8.7e-05 | 98.9127 | 98.1022 | 113.2984 | 21.1021 | 0.003568 | 0.003249 | 226.9074 | 0.26301 | 8.4016 | 0.002804 | 0.004414 | 0.004371 | 5e-05 | 0.000929 | 0.26681 | 1.3085 | 2.5473 | 100.19 | 20.84 | 37.292 | 107.5805 | 89.1119 | 99.86 | 0.009177 | 0.94648 | 0.022431 | 23.6993 | 0.073597 | 0.15647 | 0.021792 | 1.4603 | 1.4603 | 0.004361
2019-03-29 14:07 | 1553882816.5409 | 829 | 0.20519 | 0.00396 | 8.6e-05 | 99.1857 | 98.3544 | 113.0869 | 21.1027 | 0.003565 | 0.003332 | 226.8514 | 0.26393 | 8.4008 | 0.002785 | 0.005051 | 0.004986 | 5.2e-05 | 0.000925 | 0.29463 | 1.3155 | 2.5474 | 100.19 | 20.84 | 37.29 | 107.5242 | 88.6701 | 101.8746 | 0.009176 | 0.94139 | 0.02235 | 23.741 | 0.075283 | 0.15701 | 0.021665 | 1.4465 | 1.4465 | 0.004418
2019-03-29 14:07 | 1553882817.5406 | 830 | 0.20522 | 0.003937 | 8.6e-05 | 99.3689 | 98.6295 | 113.2707 | 21.1031 | 0.003595 | 0.003263 | 226.7807 | 0.264 | 8.4037 | 0.002839 | 0.004547 | 0.004412 | 4.9e-05 | 0.000919 | 0.29231 | 1.3631 | 2.5475 | 100.19 | 20.84 | 37.285 | 108.2313 | 89.9637 | 100.1911 | 0.009175 | 0.94564 | 0.022418 | 23.7069 | 0.066807 | 0.15705 | 0.02153 | 1.4419 | 1.4419 | 0.004433
2019-03-29 14:07 | 1553882818.5396 | 831 | 0.20519 | 0.003909 | 8.5e-05 | 99.5578 | 98.8059 | 113.0134 | 21.103 | 0.003579 | 0.00326 | 226.7275 | 0.26424 | 8.4035 | 0.002756 | 0.004485 | 0.0044 | 4.4e-05 | 0.000916 | 0.28495 | 1.3569 | 2.5475 | 100.19 | 20.84 | 37.281 | 107.8443 | 87.977 | 100.1134 | 0.009174 | 0.94583 | 0.022395 | 23.6775 | 0.067805 | 0.1572 | 0.021725 | 1.4535 | 1.4535 | 0.004402
2019-03-29 14:07 | 1553882819.54 | 832 | 0.20522 | 0.003887 | 8.6e-05 | 99.7329 | 98.8937 | 112.9411 | 21.1029 | 0.003544 | 0.003249 | 226.6997 | 0.26329 | 8.4018 | 0.002804 | 0.004778 | 0.004654 | 4.5e-05 | 0.000911 | 0.29645 | 1.3614 | 2.5475 | 100.19 | 20.84 | 37.277 | 107.0017 | 89.1329 | 99.8647 | 0.009173 | 0.94647 | 0.022395 | 23.6621 | 0.075587 | 0.15663 | 0.021591 | 1.4446 | 1.4446 | 0.004413
2019-03-29 14:07 | 1553882820.5397 | 833 | 0.2052 | 0.003858 | 8.5e-05 | 99.926 | 99.2653 | 112.8162 | 21.1031 | 0.003546 | 0.003273 | 226.6494 | 0.26318 | 8.4028 | 0.002774 | 0.004002 | 0.004004 | 4.7e-05 | 0.000915 | 0.30541 | 1.3845 | 2.5475 | 100.19 | 20.84 | 37.272 | 107.0429 | 88.4005 | 100.4322 | 0.009172 | 0.94503 | 0.022366 | 23.667 | 0.059333 | 0.15656 | 0.021775 | 1.455 | 1.455 | 0.00438
2019-03-29 14:07 | 1553882821.5397 | 834 | 0.20519 | 0.003843 | 8.6e-05 | 100.0439 | 99.4889 | 112.8421 | 21.1027 | 0.003539 | 0.003279 | 226.581 | 0.26327 | 8.4043 | 0.002778 | 0.004302 | 0.004204 | 4.7e-05 | 0.000911 | 0.32211 | 1.3631 | 2.5475 | 100.19 | 20.84 | 37.27 | 106.8707 | 88.4899 | 100.577 | 0.009171 | 0.94466 | 0.022364 | 23.6743 | 0.049763 | 0.15662 | 0.02181 | 1.4572 | 1.4572 | 0.004374
2019-03-29 14:07 | 1553882822.5399 | 835 | 0.20523 | 0.003832 | 8.5e-05 | 100.1654 | 99.6418 | 113.2416 | 21.1032 | 0.003526 | 0.003229 | 226.5275 | 0.26372 | 8.4028 | 0.002768 | 0.004532 | 0.00441 | 4.7e-05 | 0.000915 | 0.29878 | 1.3194 | 2.5475 | 100.19 | 20.84 | 37.264 | 106.5666 | 88.2583 | 99.3746 | 0.00917 | 0.94771 | 0.02244 | 23.678 | 0.046886 | 0.15689 | 0.021605 | 1.4485 | 1.4485 | 0.004408
2019-03-29 14:07 | 1553882823.5402 | 836 | 0.20535 | 0.003819 | 8.5e-05 | 100.401 | 99.8093 | 113.4714 | 21.1017 | 0.003542 | 0.003258 | 226.4771 | 0.26325 | 8.4026 | 0.002724 | 0.004535 | 0.004404 | 4.7e-05 | 0.000917 | 0.26449 | 1.354 | 2.5475 | 100.19 | 20.84 | 37.261 | 106.9595 | 87.2005 | 100.0674 | 0.009169 | 0.94595 | 0.022442 | 23.724 | 0.052868 | 0.15661 | 0.020911 | 1.402 | 1.402 | 0.004546
2019-03-29 14:07 | 1553882824.5388 | 837 | 0.20536 | 0.003807 | 8.5e-05 | 100.5405 | 100.1028 | 113.6347 | 21.1028 | 0.003528 | 0.003201 | 226.4234 | 0.26415 | 8.4048 | 0.002741 | 0.004419 | 0.004302 | 4.5e-05 | 0.000924 | 0.29446 | 1.3523 | 2.5474 | 100.19 | 20.84 | 37.258 | 106.6039 | 87.602 | 98.6967 | 0.009168 | 0.94944 | 0.022499 | 23.6974 | 0.039022 | 0.15714 | 0.020863 | 1.4024 | 1.4024 | 0.004561
2019-03-29 14:07 | 1553882825.5389 | 838 | 0.20537 | 0.003788 | 8.4e-05 | 100.6303 | 100.0465 | 114.1215 | 21.1025 | 0.003507 | 0.003246 | 226.3611 | 0.26374 | 8.4016 | 0.002756 | 0.004121 | 0.004106 | 4.8e-05 | 0.000923 | 0.29103 | 1.3231 | 2.5475 | 100.19 | 20.84 | 37.253 | 106.0951 | 87.9649 | 99.7804 | 0.009167 | 0.94668 | 0.022515 | 23.7827 | 0.05204 | 0.1569 | 0.020791 | 1.3985 | 1.3985 | 0.004566
2019-03-29 14:07 | 1553882826.5389 | 839 | 0.20541 | 0.003762 | 8.4e-05 | 100.7117 | 100.2373 | 114.5497 | 21.1025 | 0.003574 | 0.003236 | 226.2949 | 0.26377 | 8.4033 | 0.00286 | 0.004671 | 0.004489 | 4.6e-05 | 0.000928 | 0.29878 | 1.3423 | 2.5475 | 100.19 | 20.84 | 37.249 | 107.7304 | 90.4756 | 99.5417 | 0.009166 | 0.94729 | 0.022564 | 23.8197 | 0.042234 | 0.15692 | 0.020581 | 1.3874 | 1.3874 | 0.004603
2019-03-29 14:07 | 1553882827.5389 | 840 | 0.20543 | 0.003741 | 8.3e-05 | 100.8147 | 100.3356 | 114.5309 | 21.1021 | 0.003574 | 0.003208 | 226.2441 | 0.26446 | 8.402 | 0.002781 | 0.004263 | 0.004195 | 4.9e-05 | 0.000931 | 0.28672 | 1.3282 | 2.5474 | 100.19 | 20.84 | 37.246 | 107.7237 | 88.561 | 98.87 | 0.009165 | 0.949 | 0.022583 | 23.7962 | 0.042608 | 0.15733 | 0.020532 | 1.3852 | 1.3852 | 0.004622
2019-03-29 14:07 | 1553882828.5391 | 841 | 0.20545 | 0.003716 | 8.3e-05 | 100.909 | 100.4479 | 114.5786 | 21.1024 | 0.003495 | 0.00323 | 226.1908 | 0.26413 | 8.4011 | 0.002794 | 0.004699 | 0.004607 | 5e-05 | 0.000934 | 0.30735 | 1.2832 | 2.5472 | 100.19 | 20.84 | 37.24 | 105.8252 | 88.8918 | 99.3864 | 0.009164 | 0.94768 | 0.022572 | 23.8177 | 0.040965 | 0.15713 | 0.020417 | 1.3769 | 1.3769 | 0.004645
2019-03-29 14:07 | 1553882829.539 | 842 | 0.20553 | 0.003694 | 8.2e-05 | 100.9635 | 100.7226 | 114.7389 | 21.1023 | 0.003493 | 0.003227 | 226.1373 | 0.26334 | 8.4033 | 0.002735 | 0.00425 | 0.00414 | 4.7e-05 | 0.000935 | 0.27688 | 1.3362 | 2.5475 | 100.19 | 20.84 | 37.238 | 105.7762 | 87.4641 | 99.3165 | 0.009163 | 0.94786 | 0.022589 | 23.8321 | 0.021362 | 0.15666 | 0.020004 | 1.3501 | 1.3501 | 0.004723
2019-03-29 14:07 | 1553882830.539 | 843 | 0.20555 | 0.003662 | 8.2e-05 | 101.0164 | 100.5987 | 114.7851 | 21.1028 | 0.003476 | 0.003223 | 226.0824 | 0.26445 | 8.4015 | 0.002808 | 0.004581 | 0.004614 | 4.6e-05 | 0.000931 | 0.31133 | 1.3253 | 2.5475 | 100.19 | 20.84 | 37.233 | 105.3455 | 89.2093 | 99.2362 | 0.009162 | 0.94806 | 0.022596 | 23.8344 | 0.037063 | 0.15732 | 0.019926 | 1.3453 | 1.3453 | 0.00476
2019-03-29 14:07 | 1553882831.5401 | 844 | 0.20563 | 0.003635 | 8.1e-05 | 101.1203 | 100.653 | 114.4152 | 21.1015 | 0.003499 | 0.003168 | 226.0328 | 0.26411 | 8.4025 | 0.002744 | 0.003928 | 0.00379 | 4.7e-05 | 0.00094 | 0.29502 | 1.3534 | 2.5474 | 100.19 | 20.84 | 37.23 | 105.9149 | 87.6679 | 97.9029 | 0.009161 | 0.95147 | 0.022601 | 23.7533 | 0.041432 | 0.15712 | 0.019481 | 1.3154 | 1.3154 | 0.004861
2019-03-29 14:07 | 1553882832.5401 | 845 | 0.20564 | 0.003605 | 8.1e-05 | 101.1122 | 100.8666 | 114.0957 | 21.1018 | 0.003519 | 0.003212 | 225.9792 | 0.26338 | 8.4047 | 0.002817 | 0.004719 | 0.004662 | 4.9e-05 | 0.000942 | 0.30237 | 1.3465 | 2.5473 | 100.19 | 20.84 | 37.226 | 106.3896 | 89.4326 | 98.9498 | 0.00916 | 0.94879 | 0.022537 | 23.7535 | 0.021751 | 0.15669 | 0.019396 | 1.306 | 1.306 | 0.004883
2019-03-29 14:07 | 1553882833.5401 | 846 | 0.20567 | 0.003578 | 8.1e-05 | 101.1362 | 100.7021 | 114.0845 | 21.1022 | 0.00347 | 0.003182 | 225.9302 | 0.2634 | 8.404 | 0.002821 | 0.004733 | 0.004671 | 5e-05 | 0.000944 | 0.31741 | 1.3623 | 2.5473 | 100.19 | 20.84 | 37.22 | 105.2106 | 89.5291 | 98.2288 | 0.009159 | 0.95063 | 0.022558 | 23.7293 | 0.038469 | 0.1567 | 0.019253 | 1.2976 | 1.2976 | 0.004915
2019-03-29 14:07 | 1553882834.5402 | 847 | 0.20569 | 0.00355 | 8.1e-05 | 101.0958 | 100.8503 | 114.0004 | 21.1019 | 0.00356 | 0.003168 | 225.8914 | 0.26384 | 8.4029 | 0.002791 | 0.004434 | 0.004295 | 4.7e-05 | 0.000942 | 0.29502 | 1.3646 | 2.5475 | 100.19 | 20.84 | 37.217 | 107.381 | 88.8093 | 97.9009 | 0.009158 | 0.95147 | 0.02256 | 23.7101 | 0.021748 | 0.15696 | 0.019216 | 1.2952 | 1.2952 | 0.004932
2019-03-29 14:07 | 1553882835.5402 | 848 | 0.20572 | 0.003523 | 8.1e-05 | 101.2074 | 100.7312 | 113.8636 | 21.1013 | 0.003518 | 0.003173 | 225.844 | 0.26425 | 8.4002 | 0.002718 | 0.004348 | 0.004297 | 4.9e-05 | 0.000941 | 0.28512 | 1.3514 | 2.5475 | 100.19 | 20.84 | 37.215 | 106.3607 | 87.0376 | 98.0148 | 0.009158 | 0.95118 | 0.022543 | 23.6995 | 0.042191 | 0.1572 | 0.019038 | 1.2822 | 1.2822 | 0.00499
2019-03-29 14:07 | 1553882836.5392 | 849 | 0.20573 | 0.003501 | 8.1e-05 | 101.2556 | 100.8503 | 113.5143 | 21.1014 | 0.003517 | 0.003103 | 225.7975 | 0.26421 | 8.4016 | 0.002716 | 0.004364 | 0.004356 | 4.7e-05 | 0.000942 | 0.30021 | 1.3724 | 2.5474 | 100.19 | 20.84 | 37.206 | 106.342 | 87.0058 | 96.3396 | 0.009155 | 0.95549 | 0.022559 | 23.6097 | 0.035879 | 0.15718 | 0.019034 | 1.2829 | 1.2829 | 0.004987
2019-03-29 14:07 | 1553882837.5382 | 850 | 0.20576 | 0.003479 | 8.2e-05 | 101.1191 | 100.7414 | 113.3204 | 21.1009 | 0.00352 | 0.003171 | 225.7588 | 0.26404 | 8.4038 | 0.002717 | 0.004442 | 0.004348 | 4.9e-05 | 0.000939 | 0.27671 | 1.4191 | 2.5475 | 100.19 | 20.84 | 37.204 | 106.421 | 87.0223 | 97.9806 | 0.009155 | 0.95127 | 0.02249 | 23.6418 | 0.033474 | 0.15707 | 0.018842 | 1.2659 | 1.2659 | 0.00505
2019-03-29 14:07 | 1553882838.5393 | 851 | 0.20577 | 0.003458 | 8.2e-05 | 101.1547 | 100.8097 | 113.2517 | 21.1016 | 0.00355 | 0.003192 | 225.7267 | 0.26357 | 8.4029 | 0.002707 | 0.004525 | 0.004432 | 5e-05 | 0.000937 | 0.29535 | 1.3471 | 2.5475 | 100.19 | 20.84 | 37.199 | 107.146 | 86.7935 | 98.4721 | 0.009154 | 0.95001 | 0.022468 | 23.6503 | 0.030566 | 0.1568 | 0.018829 | 1.2638 | 1.2638 | 0.005049
2019-03-29 14:07 | 1553882839.5397 | 852 | 0.20581 | 0.003444 | 8.2e-05 | 101.1553 | 100.767 | 113.2406 | 21.1005 | 0.00352 | 0.00316 | 225.6698 | 0.26395 | 8.4029 | 0.002773 | 0.004227 | 0.00417 | 4.9e-05 | 0.000936 | 0.27903 | 1.3421 | 2.5475 | 100.19 | 20.84 | 37.197 | 106.43 | 88.3798 | 97.6924 | 0.009153 | 0.95201 | 0.022491 | 23.6243 | 0.034405 | 0.15702 | 0.018616 | 1.2508 | 1.2508 | 0.00511
2019-03-29 14:07 | 1553882840.5384 | 853 | 0.20583 | 0.003418 | 8.2e-05 | 101.1398 | 100.8243 | 113.3397 | 21.1016 | 0.00344 | 0.003194 | 225.6342 | 0.26324 | 8.401 | 0.002669 | 0.004876 | 0.004746 | 5e-05 | 0.000935 | 0.30254 | 1.4034 | 2.5475 | 100.19 | 20.84 | 37.194 | 104.4707 | 85.8787 | 98.53 | 0.009152 | 0.94986 | 0.022475 | 23.6613 | 0.027941 | 0.1566 | 0.018496 | 1.2418 | 1.2418 | 0.005132
2019-03-29 14:07 | 1553882841.5393 | 854 | 0.20585 | 0.0034 | 8.1e-05 | 101.1089 | 100.7144 | 113.8443 | 21.1006 | 0.003436 | 0.003189 | 225.5786 | 0.26348 | 8.4026 | 0.002687 | 0.004544 | 0.00445 | 4.6e-05 | 0.000935 | 0.27544 | 1.3546 | 2.5474 | 100.19 | 20.84 | 37.194 | 104.3776 | 86.3 | 98.4123 | 0.009152 | 0.95017 | 0.022529 | 23.7102 | 0.034968 | 0.15675 | 0.01843 | 1.2404 | 1.2404 | 0.005143
2019-03-29 14:07 | 1553882842.5388 | 855 | 0.20591 | 0.003376 | 8.1e-05 | 101.1365 | 100.7243 | 113.9415 | 21.0997 | 0.003485 | 0.003145 | 225.5357 | 0.26387 | 8.4014 | 0.002684 | 0.004316 | 0.004217 | 5.1e-05 | 0.000934 | 0.30342 | 1.3368 | 2.5473 | 100.19 | 20.84 | 37.194 | 105.5755 | 86.2375 | 97.3352 | 0.009152 | 0.95293 | 0.022571 | 23.6859 | 0.036525 | 0.15697 | 0.018112 | 1.2213 | 1.2213 | 0.005231
2019-03-29 14:07 | 1553882843.5385 | 856 | 0.20596 | 0.003355 | 8e-05 | 101.0985 | 100.6223 | 114.0279 | 21.1009 | 0.003444 | 0.003136 | 225.4925 | 0.26437 | 8.4019 | 0.002631 | 0.003943 | 0.003956 | 4.8e-05 | 0.000938 | 0.27743 | 1.365 | 2.5474 | 100.19 | 20.84 | 37.196 | 104.5666 | 84.9586 | 97.1271 | 0.009153 | 0.95346 | 0.022586 | 23.6882 | 0.042229 | 0.15727 | 0.017822 | 1.2025 | 1.2025 | 0.005323
2019-03-29 14:07 | 1553882844.5389 | 857 | 0.20599 | 0.003329 | 8e-05 | 101.1206 | 100.6492 | 113.9425 | 21.1007 | 0.003436 | 0.003142 | 225.4372 | 0.26408 | 8.4019 | 0.002668 | 0.00404 | 0.003935 | 5e-05 | 0.000941 | 0.26537 | 1.3646 | 2.5475 | 100.19 | 20.84 | 37.196 | 104.3896 | 85.8521 | 97.275 | 0.009153 | 0.95308 | 0.022573 | 23.6841 | 0.041792 | 0.1571 | 0.017702 | 1.1937 | 1.1937 | 0.005356
2019-03-29 14:07 | 1553882845.5386 | 858 | 0.20602 | 0.003311 | 7.9e-05 | 101.0371 | 100.6806 | 113.8225 | 21.1004 | 0.003484 | 0.003166 | 225.3915 | 0.26346 | 8.4028 | 0.002674 | 0.005005 | 0.004822 | 4.9e-05 | 0.000939 | 0.30165 | 1.3379 | 2.5474 | 100.19 | 20.84 | 37.195 | 105.5359 | 85.9843 | 97.8422 | 0.009153 | 0.95163 | 0.022544 | 23.6897 | 0.031612 | 0.15673 | 0.017519 | 1.1799 | 1.1799 | 0.005407
2019-03-29 14:07 | 1553882846.5396 | 859 | 0.20605 | 0.003285 | 7.9e-05 | 101.0784 | 100.6779 | 114.0915 | 21.0999 | 0.003465 | 0.003082 | 225.3542 | 0.26396 | 8.4005 | 0.002634 | 0.004647 | 0.004628 | 5.2e-05 | 0.000943 | 0.31094 | 1.3066 | 2.5474 | 100.19 | 20.84 | 37.195 | 105.0737 | 85.0161 | 95.8304 | 0.009153 | 0.95681 | 0.022632 | 23.6533 | 0.03551 | 0.15703 | 0.017341 | 1.1724 | 1.1724 | 0.005451
2019-03-29 14:07 | 1553882847.5396 | 860 | 0.20613 | 0.003254 | 7.9e-05 | 100.9964 | 100.4851 | 114.6081 | 21.0998 | 0.003466 | 0.003174 | 225.3094 | 0.26312 | 8.4027 | 0.002754 | 0.004564 | 0.004479 | 5.5e-05 | 0.000942 | 0.3095 | 1.3174 | 2.5474 | 100.19 | 20.84 | 37.195 | 105.1004 | 87.9161 | 98.0438 | 0.009153 | 0.95111 | 0.022615 | 23.7778 | 0.045398 | 0.15653 | 0.016911 | 1.1425 | 1.1425 | 0.005576
2019-03-29 14:07 | 1553882848.5386 | 861 | 0.20616 | 0.003229 | 7.8e-05 | 100.944 | 100.5182 | 114.9414 | 21.1005 | 0.003461 | 0.003133 | 225.2772 | 0.26385 | 8.4026 | 0.002662 | 0.004048 | 0.003971 | 5.2e-05 | 0.000949 | 0.30309 | 1.3414 | 2.5473 | 100.19 | 20.84 | 37.195 | 104.9808 | 85.702 | 97.042 | 0.009153 | 0.95368 | 0.022679 | 23.7802 | 0.037811 | 0.15697 | 0.016764 | 1.1358 | 1.1358 | 0.005625
2019-03-29 14:07 | 1553882849.5396 | 862 | 0.20615 | 0.003198 | 7.7e-05 | 100.9105 | 100.5728 | 115.0783 | 21.0994 | 0.003442 | 0.003097 | 225.2405 | 0.26327 | 8.4004 | 0.002631 | 0.00432 | 0.004248 | 4.8e-05 | 0.000951 | 0.30093 | 1.3724 | 2.5474 | 100.19 | 20.84 | 37.194 | 104.5211 | 84.9485 | 96.1809 | 0.009152 | 0.95591 | 0.022719 | 23.7667 | 0.029982 | 0.15662 | 0.016897 | 1.1469 | 1.1469 | 0.005558
2019-03-29 14:07 | 1553882850.5391 | 863 | 0.20619 | 0.003167 | 7.7e-05 | 100.8976 | 100.6568 | 115.1564 | 21.0997 | 0.003449 | 0.003137 | 225.1976 | 0.26391 | 8.4022 | 0.002658 | 0.00456 | 0.004512 | 5.1e-05 | 0.000956 | 0.29894 | 1.3377 | 2.5473 | 100.19 | 20.84 | 37.194 | 104.693 | 85.6144 | 97.1487 | 0.009152 | 0.95341 | 0.022697 | 23.8059 | 0.021377 | 0.157 | 0.016659 | 1.1296 | 1.1296 | 0.005657
2019-03-29 14:08 | 1553882851.5387 | 864 | 0.20619 | 0.003138 | 7.7e-05 | 100.8494 | 100.6861 | 115.1257 | 21.1001 | 0.003407 | 0.003105 | 225.1678 | 0.26299 | 8.4048 | 0.002679 | 0.004646 | 0.004524 | 5e-05 | 0.000956 | 0.27455 | 1.3747 | 2.5476 | 100.19 | 20.84 | 37.191 | 103.6892 | 86.0995 | 96.364 | 0.009152 | 0.95543 | 0.022718 | 23.7775 | 0.014496 | 0.15645 | 0.016741 | 1.1362 | 1.1362 | 0.005604
2019-03-29 14:08 | 1553882852.54 | 865 | 0.20619 | 0.003114 | 7.7e-05 | 100.8228 | 100.396 | 115.0717 | 21.1001 | 0.003451 | 0.003145 | 225.1322 | 0.26417 | 8.4005 | 0.002614 | 0.004226 | 0.004105 | 5e-05 | 0.000955 | 0.2844 | 1.3405 | 2.5473 | 100.19 | 20.84 | 37.19 | 104.7467 | 84.5428 | 97.3419 | 0.009152 | 0.95291 | 0.022682 | 23.8033 | 0.037941 | 0.15716 | 0.016724 | 1.1333 | 1.1333 | 0.005644
2019-03-29 14:08 | 1553882853.5382 | 866 | 0.20619 | 0.00309 | 7.6e-05 | 100.7147 | 100.3899 | 114.982 | 21.1001 | 0.00347 | 0.003109 | 225.0897 | 0.26365 | 8.4011 | 0.002611 | 0.003914 | 0.003897 | 4.9e-05 | 0.000953 | 0.3079 | 1.3271 | 2.5473 | 100.19 | 20.84 | 37.184 | 105.2037 | 84.4826 | 96.4722 | 0.00915 | 0.95515 | 0.0227 | 23.7661 | 0.028897 | 0.15684 | 0.01679 | 1.1387 | 1.1387 | 0.005606
2019-03-29 14:08 | 1553882854.5387 | 867 | 0.20621 | 0.003068 | 7.7e-05 | 100.7931 | 100.4172 | 114.9809 | 21.1001 | 0.003408 | 0.003126 | 225.069 | 0.26473 | 8.4 | 0.002662 | 0.004259 | 0.004188 | 5.1e-05 | 0.000956 | 0.2625 | 1.3346 | 2.5474 | 100.19 | 20.84 | 37.184 | 103.701 | 85.703 | 96.8796 | 0.00915 | 0.9541 | 0.022688 | 23.7791 | 0.033427 | 0.15749 | 0.016667 | 1.1297 | 1.1297 | 0.005674
2019-03-29 14:08 | 1553882855.5389 | 868 | 0.20626 | 0.003055 | 7.7e-05 | 100.6465 | 100.394 | 114.9766 | 21.1002 | 0.003447 | 0.003115 | 225.0434 | 0.26402 | 8.4018 | 0.00263 | 0.004009 | 0.003866 | 5e-05 | 0.000954 | 0.28766 | 1.3576 | 2.5476 | 100.19 | 20.84 | 37.182 | 104.6564 | 84.9344 | 96.6187 | 0.00915 | 0.95477 | 0.022695 | 23.7703 | 0.022469 | 0.15706 | 0.016408 | 1.1125 | 1.1125 | 0.005746
2019-03-29 14:08 | 1553882856.5378 | 869 | 0.2063 | 0.003045 | 7.7e-05 | 100.6713 | 100.4926 | 114.9811 | 21.1005 | 0.003462 | 0.003122 | 225.0018 | 0.26297 | 8.4029 | 0.002657 | 0.004227 | 0.004227 | 5.4e-05 | 0.000953 | 0.30397 | 1.3437 | 2.5475 | 100.19 | 20.84 | 37.179 | 105.0074 | 85.5918 | 96.7889 | 0.009149 | 0.95433 | 0.02269 | 23.7762 | 0.015893 | 0.15644 | 0.016148 | 1.0946 | 1.0946 | 0.005817
2019-03-29 14:08 | 1553882857.5392 | 870 | 0.20633 | 0.003026 | 7.8e-05 | 100.6333 | 100.3697 | 115.0255 | 21.0993 | 0.003426 | 0.00306 | 224.9563 | 0.26385 | 8.4047 | 0.002603 | 0.003937 | 0.003867 | 5e-05 | 0.000952 | 0.31492 | 1.3261 | 2.5474 | 100.19 | 20.84 | 37.178 | 104.1325 | 84.2913 | 95.2981 | 0.009149 | 0.9582 | 0.022741 | 23.7328 | 0.023458 | 0.15696 | 0.016017 | 1.0881 | 1.0881 | 0.005871
2019-03-29 14:08 | 1553882858.5389 | 871 | 0.20639 | 0.003009 | 7.7e-05 | 100.6396 | 100.2243 | 114.9968 | 21.0985 | 0.003436 | 0.003081 | 224.9103 | 0.26435 | 8.4016 | 0.00256 | 0.00418 | 0.004082 | 5.2e-05 | 0.000951 | 0.28528 | 1.3208 | 2.5474 | 100.19 | 20.84 | 37.176 | 104.3748 | 83.2475 | 95.8025 | 0.009148 | 0.95689 | 0.022722 | 23.7461 | 0.036982 | 0.15726 | 0.015676 | 1.0639 | 1.0639 | 0.006016
2019-03-29 14:08 | 1553882859.5379 | 872 | 0.20639 | 0.002987 | 7.8e-05 | 100.5689 | 100.1848 | 114.9191 | 21.0984 | 0.003386 | 0.003045 | 224.8641 | 0.26416 | 8.4028 | 0.002625 | 0.003729 | 0.003742 | 5.1e-05 | 0.000939 | 0.27903 | 1.368 | 2.5475 | 100.19 | 20.84 | 37.176 | 103.1784 | 84.8054 | 94.9365 | 0.009148 | 0.95914 | 0.022741 | 23.7102 | 0.034235 | 0.15715 | 0.015732 | 1.0687 | 1.0687 | 0.005985
2019-03-29 14:08 | 1553882860.5393 | 873 | 0.2064 | 0.002965 | 7.8e-05 | 100.5498 | 100.0452 | 114.7492 | 21.0982 | 0.003386 | 0.003041 | 224.8187 | 0.26328 | 8.4006 | 0.002595 | 0.003911 | 0.003754 | 4.8e-05 | 0.000936 | 0.27815 | 1.3446 | 2.5475 | 100.19 | 20.84 | 37.176 | 103.1585 | 84.0765 | 94.8274 | 0.009148 | 0.95942 | 0.022728 | 23.6891 | 0.045002 | 0.15663 | 0.015694 | 1.0654 | 1.0654 | 0.005983
2019-03-29 14:08 | 1553882861.538 | 874 | 0.20644 | 0.002939 | 7.8e-05 | 100.448 | 100.0233 | 114.8832 | 21.0988 | 0.0034 | 0.003114 | 224.7812 | 0.26462 | 8.4016 | 0.002563 | 0.00432 | 0.00421 | 5.1e-05 | 0.000935 | 0.28208 | 1.3433 | 2.5474 | 100.19 | 20.84 | 37.175 | 103.511 | 83.311 | 96.5938 | 0.009148 | 0.95484 | 0.022687 | 23.7598 | 0.037897 | 0.15742 | 0.015438 | 1.0462 | 1.0462 | 0.006124
2019-03-29 14:08 | 1553882862.5391 | 875 | 0.20649 | 0.002924 | 7.8e-05 | 100.4735 | 100.1568 | 115.2013 | 21.0991 | 0.003398 | 0.003055 | 224.739 | 0.26354 | 8.4018 | 0.002641 | 0.004424 | 0.004344 | 4.8e-05 | 0.000936 | 0.29518 | 1.3191 | 2.5473 | 100.19 | 20.84 | 37.175 | 103.4533 | 85.1886 | 95.179 | 0.009148 | 0.95851 | 0.022762 | 23.7471 | 0.028238 | 0.15678 | 0.015176 | 1.0317 | 1.0317 | 0.006185
2019-03-29 14:08 | 1553882863.5576 | 876 | 0.20652 | 0.002902 | 7.8e-05 | 100.3696 | 99.9213 | 115.1133 | 21.0988 | 0.003359 | 0.003096 | 224.7123 | 0.26359 | 8.4014 | 0.002666 | 0.004134 | 0.004008 | 5.3e-05 | 0.000935 | 0.2891 | 1.3569 | 2.5474 | 100.19 | 20.84 | 37.175 | 102.5126 | 85.7884 | 96.1608 | 0.009148 | 0.95596 | 0.022723 | 23.7696 | 0.040041 | 0.15681 | 0.014997 | 1.0177 | 1.0177 | 0.006271
2019-03-29 14:08 | 1553882864.5491 | 877 | 0.20657 | 0.002878 | 7.9e-05 | 100.4079 | 100.0499 | 115.563 | 21.0988 | 0.003373 | 0.003106 | 224.656 | 0.26365 | 8.4016 | 0.002582 | 0.004596 | 0.004528 | 5.2e-05 | 0.000943 | 0.30381 | 1.3628 | 2.5474 | 100.19 | 20.84 | 37.172 | 102.8463 | 83.7713 | 96.3981 | 0.009147 | 0.95534 | 0.02276 | 23.8237 | 0.031945 | 0.15684 | 0.014751 | 1.0026 | 1.0026 | 0.006367
2019-03-29 14:08 | 1553882865.5498 | 878 | 0.20661 | 0.002851 | 7.9e-05 | 100.2918 | 100.0684 | 115.8904 | 21.099 | 0.003382 | 0.003028 | 224.605 | 0.26388 | 8.402 | 0.00262 | 0.003693 | 0.003581 | 4.8e-05 | 0.000941 | 0.32101 | 1.3339 | 2.5473 | 100.19 | 20.84 | 37.171 | 103.0681 | 84.6961 | 94.5229 | 0.009147 | 0.96022 | 0.02285 | 23.7968 | 0.019945 | 0.15699 | 0.014544 | 0.99238 | 0.99238 | 0.006438
2019-03-29 14:08 | 1553882866.5489 | 879 | 0.20668 | 0.002827 | 7.8e-05 | 100.3055 | 99.9104 | 115.9949 | 21.098 | 0.003342 | 0.003028 | 224.5544 | 0.2633 | 8.403 | 0.002672 | 0.004494 | 0.004395 | 5.3e-05 | 0.000944 | 0.2907 | 1.3092 | 2.5473 | 100.19 | 20.84 | 37.171 | 102.0993 | 85.95 | 94.504 | 0.009147 | 0.96027 | 0.022861 | 23.8069 | 0.035309 | 0.15664 | 0.014137 | 0.96506 | 0.96506 | 0.006606
2019-03-29 14:08 | 1553882867.5484 | 880 | 0.20674 | 0.002799 | 7.9e-05 | 100.2427 | 99.9639 | 115.8766 | 21.0979 | 0.003382 | 0.003032 | 224.5096 | 0.26348 | 8.4023 | 0.002618 | 0.005004 | 0.004923 | 5e-05 | 0.000946 | 0.30198 | 1.3293 | 2.5473 | 100.19 | 20.84 | 37.171 | 103.0655 | 84.6341 | 94.6119 | 0.009147 | 0.95998 | 0.022846 | 23.7983 | 0.024906 | 0.15674 | 0.013858 | 0.94524 | 0.94524 | 0.006749
2019-03-29 14:08 | 1553882868.548 | 881 | 0.20675 | 0.002772 | 8e-05 | 100.2262 | 99.7247 | 115.8296 | 21.0978 | 0.003316 | 0.003055 | 224.4503 | 0.26377 | 8.4035 | 0.00259 | 0.003905 | 0.003875 | 5.1e-05 | 0.000948 | 0.29286 | 1.333 | 2.5474 | 100.19 | 20.84 | 37.171 | 101.4717 | 83.9701 | 95.1661 | 0.009147 | 0.95854 | 0.022824 | 23.8114 | 0.044869 | 0.15692 | 0.013834 | 0.94267 | 0.94267 | 0.006775
2019-03-29 14:08 | 1553882869.549 | 882 | 0.20676 | 0.002742 | 8.1e-05 | 100.1164 | 99.7517 | 115.7193 | 21.0974 | 0.003319 | 0.00302 | 224.4298 | 0.2637 | 8.4033 | 0.002638 | 0.004407 | 0.004377 | 5.1e-05 | 0.00095 | 0.29717 | 1.3385 | 2.5474 | 100.19 | 20.84 | 37.172 | 101.5466 | 85.1166 | 94.3131 | 0.009147 | 0.96076 | 0.02284 | 23.7724 | 0.032642 | 0.15687 | 0.013763 | 0.93842 | 0.93842 | 0.006804
2019-03-29 14:08 | 1553882870.5484 | 883 | 0.20677 | 0.002713 | 8.1e-05 | 100.1196 | 99.768 | 115.8566 | 21.0974 | 0.003304 | 0.003057 | 224.4036 | 0.26357 | 8.403 | 0.002631 | 0.004144 | 0.004112 | 5e-05 | 0.000947 | 0.29999 | 1.344 | 2.5474 | 100.19 | 20.84 | 37.172 | 101.1926 | 84.948 | 95.2251 | 0.009147 | 0.95839 | 0.022825 | 23.816 | 0.031467 | 0.1568 | 0.013778 | 0.93877 | 0.93877 | 0.006798
2019-03-29 14:08 | 1553882871.5494 | 884 | 0.20675 | 0.002684 | 8.1e-05 | 100.0789 | 99.7547 | 115.7775 | 21.0976 | 0.003347 | 0.00304 | 224.3445 | 0.26461 | 8.4017 | 0.002621 | 0.004148 | 0.004012 | 5.2e-05 | 0.000944 | 0.3115 | 1.3315 | 2.5473 | 100.19 | 20.84 | 37.173 | 102.2199 | 84.7204 | 94.8063 | 0.009147 | 0.95948 | 0.02283 | 23.7944 | 0.029023 | 0.15741 | 0.013918 | 0.94855 | 0.94855 | 0.006754
2019-03-29 14:08 | 1553882872.5491 | 885 | 0.20677 | 0.002663 | 8.3e-05 | 100.0469 | 99.725 | 115.9179 | 21.0968 | 0.003292 | 0.00303 | 224.2944 | 0.26405 | 8.4026 | 0.002611 | 0.004111 | 0.003997 | 5.3e-05 | 0.00095 | 0.30215 | 1.3861 | 2.5475 | 100.19 | 20.84 | 37.173 | 100.9031 | 84.4756 | 94.5622 | 0.009147 | 0.96011 | 0.022852 | 23.8009 | 0.028823 | 0.15709 | 0.013819 | 0.94259 | 0.94259 | 0.006783
2019-03-29 14:08 | 1553882873.5491 | 886 | 0.20683 | 0.002641 | 8.3e-05 | 100.0275 | 99.6264 | 116.2289 | 21.099 | 0.003374 | 0.003016 | 224.2503 | 0.26443 | 8.4042 | 0.002555 | 0.00446 | 0.004408 | 5.2e-05 | 0.000948 | 0.31133 | 1.3262 | 2.5474 | 100.19 | 20.84 | 37.175 | 102.874 | 83.1193 | 94.22 | 0.009148 | 0.96101 | 0.022893 | 23.8217 | 0.035934 | 0.15731 | 0.013463 | 0.91987 | 0.91987 | 0.00696
2019-03-29 14:08 | 1553882874.5491 | 887 | 0.20688 | 0.002613 | 8.4e-05 | 99.9526 | 99.6346 | 116.3157 | 21.0975 | 0.003347 | 0.003022 | 224.2374 | 0.26438 | 8.403 | 0.002592 | 0.003714 | 0.003733 | 5e-05 | 0.000952 | 0.29927 | 1.2916 | 2.5472 | 100.19 | 20.84 | 37.175 | 102.2339 | 84.022 | 94.3783 | 0.009148 | 0.96059 | 0.022896 | 23.8357 | 0.028505 | 0.15728 | 0.013235 | 0.90433 | 0.90433 | 0.007078
2019-03-29 14:08 | 1553882875.5486 | 888 | 0.20689 | 0.002592 | 8.5e-05 | 99.9254 | 99.4909 | 116.607 | 21.0984 | 0.003301 | 0.002997 | 224.2045 | 0.2628 | 8.4045 | 0.002639 | 0.0037 | 0.003592 | 5.3e-05 | 0.000951 | 0.27041 | 1.3748 | 2.5475 | 100.19 | 20.84 | 37.175 | 101.1008 | 85.1481 | 93.7592 | 0.009148 | 0.96221 | 0.022944 | 23.8455 | 0.038975 | 0.15634 | 0.013177 | 0.90217 | 0.90217 | 0.007053
2019-03-29 14:08 | 1553882876.5492 | 889 | 0.2069 | 0.002572 | 8.6e-05 | 99.9173 | 99.4384 | 116.6885 | 21.0984 | 0.003213 | 0.003 | 224.1924 | 0.26388 | 8.4046 | 0.002599 | 0.004002 | 0.003932 | 5.3e-05 | 0.000942 | 0.29999 | 1.3388 | 2.5474 | 100.19 | 20.84 | 37.177 | 98.98 | 84.1837 | 93.8384 | 0.009148 | 0.96201 | 0.02295 | 23.8564 | 0.042978 | 0.15698 | 0.013138 | 0.89967 | 0.89967 | 0.007102
2019-03-29 14:08 | 1553882877.5482 | 890 | 0.20691 | 0.002554 | 8.7e-05 | 99.9032 | 99.3776 | 116.7904 | 21.0964 | 0.003291 | 0.003001 | 224.1569 | 0.2639 | 8.401 | 0.002596 | 0.003548 | 0.003493 | 5.1e-05 | 0.000942 | 0.30342 | 1.355 | 2.5474 | 100.19 | 20.84 | 37.177 | 100.8659 | 84.1084 | 93.8685 | 0.009148 | 0.96193 | 0.022959 | 23.8677 | 0.047185 | 0.15699 | 0.013126 | 0.89912 | 0.89912 | 0.007107
2019-03-29 14:08 | 1553882878.5483 | 891 | 0.20692 | 0.002534 | 8.8e-05 | 99.867 | 99.4141 | 117.0239 | 21.0986 | 0.003284 | 0.003041 | 224.1349 | 0.26338 | 8.4024 | 0.002588 | 0.00362 | 0.003523 | 4.9e-05 | 0.000948 | 0.3178 | 1.3494 | 2.5473 | 100.19 | 20.84 | 37.176 | 100.6985 | 83.9232 | 94.8285 | 0.009148 | 0.95942 | 0.022952 | 23.9228 | 0.040655 | 0.15668 | 0.013079 | 0.8956 | 0.8956 | 0.00712
2019-03-29 14:08 | 1553882879.5497 | 892 | 0.20696 | 0.002517 | 8.8e-05 | 99.9032 | 99.3677 | 117.023 | 21.0975 | 0.003291 | 0.003064 | 224.091 | 0.26313 | 8.4011 | 0.002564 | 0.004059 | 0.003936 | 5.4e-05 | 0.00095 | 0.28191 | 1.3415 | 2.5473 | 100.19 | 20.84 | 37.175 | 100.8765 | 83.3309 | 95.3794 | 0.009148 | 0.95798 | 0.022935 | 23.9406 | 0.048075 | 0.15654 | 0.01287 | 0.88054 | 0.88054 | 0.007235
2019-03-29 14:08 | 1553882880.5483 | 893 | 0.20697 | 0.002499 | 8.9e-05 | 99.8101 | 99.2623 | 117.1166 | 21.0985 | 0.003267 | 0.003005 | 224.0446 | 0.26404 | 8.4034 | 0.002543 | 0.004313 | 0.004238 | 5e-05 | 0.000957 | 0.29734 | 1.372 | 2.5475 | 100.19 | 20.84 | 37.175 | 100.2841 | 82.8374 | 93.9519 | 0.009148 | 0.96171 | 0.022988 | 23.9038 | 0.049234 | 0.15708 | 0.01278 | 0.87633 | 0.87633 | 0.007295
2019-03-29 14:08 | 1553882881.5494 | 894 | 0.20701 | 0.002485 | 8.9e-05 | 99.7766 | 99.4005 | 117.2744 | 21.0984 | 0.003248 | 0.002995 | 224.0331 | 0.26386 | 8.4005 | 0.002496 | 0.004215 | 0.004125 | 5.2e-05 | 0.000958 | 0.28998 | 1.3379 | 2.5473 | 100.19 | 20.84 | 37.174 | 99.8303 | 81.7132 | 93.7289 | 0.009148 | 0.96229 | 0.023011 | 23.9126 | 0.033783 | 0.15697 | 0.0126 | 0.8648 | 0.8648 | 0.007387
2019-03-29 14:08 | 1553882882.5493 | 895 | 0.20702 | 0.002467 | 9e-05 | 99.7392 | 99.3155 | 117.2734 | 21.0984 | 0.003197 | 0.002967 | 224.0109 | 0.26391 | 8.4024 | 0.002631 | 0.004832 | 0.004775 | 5.5e-05 | 0.000958 | 0.28584 | 1.3233 | 2.5474 | 100.19 | 20.84 | 37.173 | 98.5866 | 84.9531 | 93.0426 | 0.009147 | 0.9641 | 0.023032 | 23.8901 | 0.038079 | 0.157 | 0.012533 | 0.86086 | 0.86086 | 0.007423
2019-03-29 14:08 | 1553882883.5474 | 896 | 0.20705 | 0.002454 | 9.1e-05 | 99.7542 | 99.4875 | 117.161 | 21.0983 | 0.003235 | 0.00295 | 223.9905 | 0.26338 | 8.4012 | 0.002667 | 0.00369 | 0.003622 | 5.4e-05 | 0.000959 | 0.28838 | 1.3549 | 2.5474 | 100.19 | 20.84 | 37.169 | 99.5086 | 85.8254 | 92.6295 | 0.009146 | 0.96519 | 0.023034 | 23.8652 | 0.023943 | 0.15669 | 0.012402 | 0.85188 | 0.85188 | 0.007486
2019-03-29 14:08 | 1553882884.5474 | 897 | 0.20708 | 0.002431 | 9.2e-05 | 99.72 | 99.4698 | 117.0335 | 21.0977 | 0.003234 | 0.003015 | 223.9417 | 0.26477 | 8.404 | 0.002563 | 0.004148 | 0.004015 | 5.4e-05 | 0.000956 | 0.30182 | 1.3416 | 2.5475 | 100.19 | 20.84 | 37.169 | 99.4807 | 83.3138 | 94.1983 | 0.009146 | 0.96106 | 0.022973 | 23.9033 | 0.022478 | 0.15751 | 0.012251 | 0.83913 | 0.83913 | 0.00764
2019-03-29 14:08 | 1553882885.5475 | 898 | 0.20708 | 0.002418 | 9.3e-05 | 99.7126 | 99.3309 | 116.8159 | 21.0981 | 0.003238 | 0.002951 | 223.9104 | 0.26517 | 8.4025 | 0.002597 | 0.004034 | 0.004068 | 5.3e-05 | 0.000949 | 0.29015 | 1.329 | 2.5473 | 100.19 | 20.84 | 37.167 | 99.5846 | 84.1397 | 92.6537 | 0.009146 | 0.96512 | 0.023 | 23.8308 | 0.034306 | 0.15775 | 0.012218 | 0.83785 | 0.83785 | 0.007663
2019-03-29 14:08 | 1553882886.5484 | 899 | 0.20716 | 0.002394 | 9.4e-05 | 99.7042 | 99.2776 | 116.7502 | 21.0977 | 0.003233 | 0.002951 | 223.8675 | 0.2637 | 8.404 | 0.002597 | 0.004318 | 0.004208 | 5.1e-05 | 0.000947 | 0.30021 | 1.3481 | 2.5474 | 100.19 | 20.85 | 37.164 | 99.4674 | 84.1284 | 92.6516 | 0.009151 | 0.96513 | 0.022993 | 23.824 | 0.03835 | 0.15687 | 0.011775 | 0.80698 | 0.80698 | 0.007912
2019-03-29 14:08 | 1553882887.5486 | 900 | 0.20719 | 0.002372 | 9.5e-05 | 99.7488 | 99.2097 | 116.6162 | 21.0984 | 0.003229 | 0.00299 | 223.8663 | 0.26395 | 8.4035 | 0.002598 | 0.003696 | 0.003679 | 5.5e-05 | 0.00095 | 0.29999 | 1.3795 | 2.5475 | 100.19 | 20.85 | 37.164 | 99.3808 | 84.1539 | 93.5996 | 0.009151 | 0.96263 | 0.02295 | 23.8412 | 0.048471 | 0.15702 | 0.011649 | 0.79677 | 0.79677 | 0.008021
2019-03-29 14:08 | 1553882888.5485 | 901 | 0.20721 | 0.002352 | 9.6e-05 | 99.6853 | 99.2121 | 116.3463 | 21.0978 | 0.003208 | 0.002913 | 223.8404 | 0.26408 | 8.4032 | 0.002685 | 0.004374 | 0.004251 | 5e-05 | 0.000952 | 0.29983 | 1.3224 | 2.5475 | 100.19 | 20.84 | 37.159 | 98.8519 | 86.2441 | 91.75 | 0.009144 | 0.96751 | 0.022982 | 23.7535 | 0.042564 | 0.1571 | 0.011539 | 0.79022 | 0.79022 | 0.008091
2019-03-29 14:08 | 1553882889.5481 | 902 | 0.20724 | 0.002328 | 9.7e-05 | 99.6308 | 99.2749 | 116.5584 | 21.0977 | 0.003158 | 0.002903 | 223.8113 | 0.26413 | 8.4007 | 0.002651 | 0.004506 | 0.004451 | 5.3e-05 | 0.000951 | 0.30685 | 1.3447 | 2.5474 | 100.19 | 20.84 | 37.157 | 97.6491 | 85.4284 | 91.4997 | 0.009143 | 0.96818 | 0.023011 | 23.767 | 0.032014 | 0.15713 | 0.011382 | 0.78029 | 0.78029 | 0.008196
2019-03-29 14:08 | 1553882890.5487 | 903 | 0.20723 | 0.002311 | 9.9e-05 | 99.6512 | 99.2575 | 116.4134 | 21.0975 | 0.003201 | 0.002905 | 223.7773 | 0.26358 | 8.4029 | 0.002629 | 0.00426 | 0.004102 | 5.1e-05 | 0.000954 | 0.29856 | 1.3448 | 2.5475 | 100.19 | 20.84 | 37.154 | 98.6884 | 84.9029 | 91.556 | 0.009143 | 0.96803 | 0.022995 | 23.754 | 0.03541 | 0.1568 | 0.011454 | 0.78466 | 0.78466 | 0.008133
2019-03-29 14:08 | 1553882891.5487 | 904 | 0.20726 | 0.002283 | 0.0001 | 99.58 | 99.0752 | 116.2841 | 21.0978 | 0.00316 | 0.002929 | 223.7432 | 0.2631 | 8.4046 | 0.002628 | 0.004013 | 0.00398 | 5.1e-05 | 0.000955 | 0.29679 | 1.3441 | 2.5474 | 100.19 | 20.84 | 37.152 | 97.7041 | 84.8862 | 92.1388 | 0.009142 | 0.96648 | 0.022963 | 23.7598 | 0.045452 | 0.15652 | 0.011328 | 0.77482 | 0.77482 | 0.008222
2019-03-29 14:08 | 1553882892.548 | 905 | 0.20726 | 0.002263 | 0.000101 | 99.6317 | 99.0923 | 115.82 | 21.0968 | 0.003202 | 0.002893 | 223.734 | 0.26409 | 8.4019 | 0.002601 | 0.00427 | 0.004225 | 5.3e-05 | 0.000948 | 0.30469 | 1.3301 | 2.5474 | 100.19 | 20.84 | 37.149 | 98.7198 | 84.2242 | 91.2564 | 0.009141 | 0.96882 | 0.022945 | 23.6837 | 0.048561 | 0.1571 | 0.01135 | 0.77565 | 0.77565 | 0.008244
2019-03-29 14:08 | 1553882893.5487 | 906 | 0.20727 | 0.002242 | 0.000103 | 99.6069 | 99.1571 | 115.8127 | 21.0968 | 0.003187 | 0.002966 | 223.6991 | 0.26398 | 8.4005 | 0.002572 | 0.004412 | 0.004339 | 5.2e-05 | 0.000953 | 0.32261 | 1.3598 | 2.5475 | 100.19 | 20.84 | 37.146 | 98.3601 | 83.544 | 93.0122 | 0.009141 | 0.96418 | 0.022889 | 23.7399 | 0.040479 | 0.15704 | 0.011295 | 0.7699 | 0.7699 | 0.008302
2019-03-29 14:08 | 1553882894.5487 | 907 | 0.20728 | 0.002222 | 0.000104 | 99.6569 | 99.2892 | 115.5081 | 21.0973 | 0.003207 | 0.002842 | 223.6711 | 0.26361 | 8.4036 | 0.002615 | 0.004838 | 0.004764 | 5.3e-05 | 0.000945 | 0.29214 | 1.3812 | 2.5474 | 100.19 | 20.84 | 37.144 | 98.834 | 84.5633 | 90.0252 | 0.00914 | 0.97211 | 0.022953 | 23.6117 | 0.033061 | 0.15682 | 0.011256 | 0.76929 | 0.76929 | 0.008297
2019-03-29 14:08 | 1553882895.5488 | 908 | 0.20733 | 0.002206 | 0.000105 | 99.5602 | 99.1459 | 115.274 | 21.0972 | 0.003146 | 0.002928 | 223.6259 | 0.2634 | 8.4006 | 0.002623 | 0.003417 | 0.003289 | 5.7e-05 | 0.00095 | 0.29983 | 1.3294 | 2.5474 | 100.19 | 20.84 | 37.141 | 97.3736 | 84.759 | 92.1037 | 0.009139 | 0.96658 | 0.022865 | 23.6552 | 0.037301 | 0.1567 | 0.01096 | 0.74601 | 0.74601 | 0.008549
2019-03-29 14:08 | 1553882896.5489 | 909 | 0.20734 | 0.002184 | 0.000107 | 99.6305 | 99.2742 | 115.2912 | 21.0966 | 0.003256 | 0.00291 | 223.6197 | 0.26395 | 8.4011 | 0.00264 | 0.004605 | 0.00453 | 5.6e-05 | 0.000942 | 0.29679 | 1.3731 | 2.5474 | 100.19 | 20.85 | 37.135 | 100.0312 | 85.1736 | 91.6676 | 0.009144 | 0.96773 | 0.02288 | 23.6428 | 0.032049 | 0.15702 | 0.010948 | 0.74557 | 0.74557 | 0.008572
2019-03-29 14:08 | 1553882897.5479 | 910 | 0.20737 | 0.002171 | 0.000109 | 99.6467 | 99.2957 | 115.2262 | 21.0971 | 0.003167 | 0.002942 | 223.5944 | 0.26413 | 8.4033 | 0.0026 | 0.00412 | 0.004134 | 5.3e-05 | 0.000939 | 0.2656 | 1.3525 | 2.5473 | 100.19 | 20.85 | 37.134 | 97.8617 | 84.2073 | 92.4493 | 0.009143 | 0.96566 | 0.022849 | 23.6615 | 0.031563 | 0.15713 | 0.010773 | 0.73247 | 0.73247 | 0.008731
2019-03-29 14:08 | 1553882898.5489 | 911 | 0.20738 | 0.00215 | 0.000111 | 99.5895 | 99.182 | 115.4797 | 21.0959 | 0.003208 | 0.00289 | 223.5642 | 0.26446 | 8.4017 | 0.002546 | 0.004617 | 0.004542 | 5.2e-05 | 0.000945 | 0.27057 | 1.3547 | 2.5473 | 100.19 | 20.85 | 37.135 | 98.8609 | 82.9051 | 91.1964 | 0.009144 | 0.96898 | 0.022913 | 23.6469 | 0.036673 | 0.15733 | 0.010756 | 0.73325 | 0.73325 | 0.008733
2019-03-29 14:08 | 1553882899.547 | 912 | 0.20741 | 0.002138 | 0.000113 | 99.6243 | 99.1332 | 115.5275 | 21.096 | 0.00315 | 0.002877 | 223.5472 | 0.26465 | 8.4022 | 0.002627 | 0.004468 | 0.004425 | 5.2e-05 | 0.000945 | 0.28998 | 1.3756 | 2.5476 | 100.19 | 20.85 | 37.135 | 97.4492 | 84.8524 | 90.8625 | 0.009144 | 0.96987 | 0.022929 | 23.6409 | 0.044193 | 0.15744 | 0.010619 | 0.72417 | 0.72417 | 0.008848
2019-03-29 14:08 | 1553882900.5469 | 913 | 0.20741 | 0.002119 | 0.000115 | 99.6228 | 99.0725 | 116.1129 | 21.0968 | 0.003148 | 0.002939 | 223.5253 | 0.26345 | 8.4013 | 0.002619 | 0.003515 | 0.003392 | 5.3e-05 | 0.000942 | 0.3079 | 1.3324 | 2.5474 | 100.19 | 20.85 | 37.136 | 97.4004 | 84.6564 | 92.366 | 0.009144 | 0.96588 | 0.022939 | 23.7497 | 0.049542 | 0.15673 | 0.010593 | 0.72266 | 0.72266 | 0.008827
2019-03-29 14:08 | 1553882901.5487 | 914 | 0.20744 | 0.002104 | 0.000116 | 99.6859 | 99.2438 | 116.5153 | 21.0961 | 0.003108 | 0.002904 | 223.5192 | 0.26249 | 8.4029 | 0.00269 | 0.004116 | 0.004069 | 5.3e-05 | 0.000942 | 0.2943 | 1.345 | 2.5474 | 100.19 | 20.85 | 37.141 | 96.4369 | 86.3588 | 91.5122 | 0.009145 | 0.96814 | 0.023006 | 23.763 | 0.039756 | 0.15616 | 0.010476 | 0.71657 | 0.71657 | 0.008869
2019-03-29 14:08 | 1553882902.547 | 915 | 0.20745 | 0.00209 | 0.000118 | 99.6449 | 99.25 | 116.9091 | 21.0969 | 0.003117 | 0.002878 | 223.4916 | 0.26362 | 8.4007 | 0.002601 | 0.003864 | 0.00374 | 5.3e-05 | 0.000942 | 0.28152 | 1.4014 | 2.5476 | 100.19 | 20.85 | 37.144 | 96.6618 | 84.2205 | 90.9057 | 0.009146 | 0.96976 | 0.023064 | 23.7833 | 0.035522 | 0.15683 | 0.010428 | 0.71494 | 0.71494 | 0.008928
2019-03-29 14:08 | 1553882903.5471 | 916 | 0.20744 | 0.002076 | 0.00012 | 99.5931 | 99.2913 | 117.0183 | 21.0961 | 0.003173 | 0.00288 | 223.4656 | 0.26327 | 8.4018 | 0.002598 | 0.003919 | 0.003885 | 5.4e-05 | 0.000942 | 0.28838 | 1.3725 | 2.5474 | 100.19 | 20.85 | 37.146 | 98.0128 | 84.1603 | 90.9524 | 0.009146 | 0.96963 | 0.023073 | 23.7959 | 0.027152 | 0.15662 | 0.01048 | 0.71876 | 0.71876 | 0.008869
2019-03-29 14:08 | 1553882904.5475 | 917 | 0.20746 | 0.00206 | 0.000121 | 99.6036 | 99.1913 | 117.5793 | 21.0957 | 0.00315 | 0.002911 | 223.4425 | 0.2634 | 8.4025 | 0.002643 | 0.003299 | 0.003161 | 5.3e-05 | 0.000941 | 0.3298 | 1.3434 | 2.5473 | 100.19 | 20.85 | 37.154 | 97.4661 | 85.2454 | 91.7034 | 0.009148 | 0.96764 | 0.023105 | 23.8775 | 0.037105 | 0.1567 | 0.010371 | 0.71207 | 0.71207 | 0.008957
2019-03-29 14:08 | 1553882905.5482 | 918 | 0.2075 | 0.002046 | 0.000123 | 99.6204 | 99.3653 | 118.0194 | 21.0966 | 0.003089 | 0.002892 | 223.4308 | 0.26443 | 8.4029 | 0.002671 | 0.003911 | 0.003817 | 5.4e-05 | 0.000946 | 0.29175 | 1.3529 | 2.5475 | 100.19 | 20.85 | 37.171 | 95.9785 | 85.9142 | 91.2454 | 0.009152 | 0.96885 | 0.023162 | 23.9071 | 0.022929 | 0.15731 | 0.010156 | 0.69887 | 0.69887 | 0.009161
2019-03-29 14:08 | 1553882906.5482 | 919 | 0.20755 | 0.002031 | 0.000125 | 99.647 | 99.2459 | 118.3151 | 21.0964 | 0.003107 | 0.002902 | 223.391 | 0.26325 | 8.403 | 0.00261 | 0.003689 | 0.003665 | 5.2e-05 | 0.000942 | 0.28152 | 1.3361 | 2.5473 | 100.19 | 20.85 | 37.171 | 96.4208 | 84.4547 | 91.4657 | 0.009152 | 0.96827 | 0.023184 | 23.9443 | 0.03608 | 0.15661 | 0.009896 | 0.68131 | 0.68131 | 0.009355
2019-03-29 14:08 | 1553882907.5472 | 920 | 0.20759 | 0.002013 | 0.000127 | 99.7281 | 99.2213 | 118.4602 | 21.0974 | 0.003075 | 0.00292 | 223.3384 | 0.26348 | 8.4028 | 0.002597 | 0.004399 | 0.00431 | 5.5e-05 | 0.000941 | 0.28783 | 1.4132 | 2.5475 | 100.19 | 20.85 | 37.178 | 95.6518 | 84.1282 | 91.9171 | 0.009154 | 0.96707 | 0.023184 | 23.9738 | 0.045573 | 0.15674 | 0.009666 | 0.66517 | 0.66517 | 0.009591
2019-03-29 14:08 | 1553882908.5482 | 921 | 0.20762 | 0.001999 | 0.000129 | 99.6889 | 99.2571 | 118.5724 | 21.0962 | 0.003104 | 0.002854 | 223.3127 | 0.26375 | 8.4035 | 0.002622 | 0.004326 | 0.004292 | 5.4e-05 | 0.000938 | 0.2907 | 1.3821 | 2.5475 | 100.19 | 20.85 | 37.186 | 96.3501 | 84.733 | 90.3211 | 0.009156 | 0.97132 | 0.023246 | 23.9326 | 0.038825 | 0.1569 | 0.009505 | 0.6556 | 0.6556 | 0.009741
2019-03-29 14:08 | 1553882909.5486 | 922 | 0.20762 | 0.001974 | 0.000132 | 99.647 | 99.1793 | 118.5165 | 21.0976 | 0.003109 | 0.002824 | 223.28 | 0.26383 | 8.403 | 0.002542 | 0.004391 | 0.004268 | 5.4e-05 | 0.000932 | 0.28926 | 1.3327 | 2.5473 | 100.19 | 20.85 | 37.2 | 96.4729 | 82.8107 | 89.5896 | 0.00916 | 0.97328 | 0.023264 | 23.9029 | 0.042079 | 0.15695 | 0.009492 | 0.65507 | 0.65507 | 0.009751
2019-03-29 14:08 | 1553882910.5487 | 923 | 0.20764 | 0.001955 | 0.000136 | 99.6668 | 99.4455 | 118.2271 | 21.0963 | 0.003074 | 0.002866 | 223.2675 | 0.26335 | 8.4032 | 0.0026 | 0.004192 | 0.004097 | 5.3e-05 | 0.000939 | 0.29734 | 1.3524 | 2.5475 | 100.19 | 20.85 | 37.206 | 95.6335 | 84.1942 | 90.6095 | 0.009161 | 0.97055 | 0.023203 | 23.9072 | 0.019877 | 0.15667 | 0.009399 | 0.64665 | 0.64665 | 0.009861
2019-03-29 14:08 | 1553882911.5483 | 924 | 0.20766 | 0.001943 | 0.00014 | 99.6832 | 99.3766 | 117.9353 | 21.0965 | 0.003161 | 0.002856 | 223.2355 | 0.26419 | 8.402 | 0.002616 | 0.003678 | 0.0036 | 5.6e-05 | 0.000943 | 0.26537 | 1.3334 | 2.5473 | 100.19 | 20.84 | 37.22 | 97.7287 | 84.5921 | 90.3761 | 0.009159 | 0.97117 | 0.023182 | 23.8701 | 0.027557 | 0.15716 | 0.009301 | 0.63901 | 0.63901 | 0.01001
2019-03-29 14:09 | 1553882912.5483 | 925 | 0.20764 | 0.001917 | 0.000143 | 99.6608 | 99.2469 | 118.1879 | 21.0957 | 0.003066 | 0.002853 | 223.1911 | 0.26416 | 8.4035 | 0.002554 | 0.004474 | 0.004374 | 5.2e-05 | 0.000948 | 0.31542 | 1.32 | 2.5473 | 100.19 | 20.84 | 37.225 | 95.4324 | 83.0981 | 90.2956 | 0.00916 | 0.97138 | 0.023209 | 23.893 | 0.037223 | 0.15715 | 0.009428 | 0.64838 | 0.64838 | 0.009865
2019-03-29 14:09 | 1553882913.5488 | 926 | 0.20769 | 0.001902 | 0.000146 | 99.7015 | 99.3216 | 118.267 | 21.0963 | 0.0031 | 0.002854 | 223.1586 | 0.2643 | 8.4026 | 0.002597 | 0.003632 | 0.003574 | 5.3e-05 | 0.000944 | 0.29966 | 1.3107 | 2.5474 | 100.19 | 20.84 | 37.231 | 96.2522 | 84.1379 | 90.3188 | 0.009162 | 0.97132 | 0.023216 | 23.9017 | 0.034142 | 0.15723 | 0.009182 | 0.63132 | 0.63132 | 0.010137
2019-03-29 14:09 | 1553882914.5485 | 927 | 0.20773 | 0.001888 | 0.000149 | 99.7814 | 99.4445 | 118.1501 | 21.0964 | 0.003111 | 0.002841 | 223.1489 | 0.26325 | 8.4035 | 0.002578 | 0.003795 | 0.003724 | 5.6e-05 | 0.000943 | 0.31238 | 1.328 | 2.5475 | 100.19 | 20.84 | 37.244 | 96.5273 | 83.6677 | 90.0056 | 0.009165 | 0.97216 | 0.023215 | 23.8796 | 0.030252 | 0.15661 | 0.008948 | 0.61484 | 0.61484 | 0.010367
2019-03-29 14:09 | 1553882915.5475 | 928 | 0.20773 | 0.001869 | 0.000152 | 99.7365 | 99.5145 | 117.9579 | 21.0961 | 0.003062 | 0.002822 | 223.1552 | 0.26343 | 8.4033 | 0.00263 | 0.004545 | 0.004438 | 5.6e-05 | 0.00094 | 0.30055 | 1.3246 | 2.5474 | 100.19 | 20.84 | 37.251 | 95.3226 | 84.9384 | 89.5593 | 0.009167 | 0.97336 | 0.02321 | 23.8455 | 0.019935 | 0.15671 | 0.008967 | 0.61582 | 0.61582 | 0.010357
2019-03-29 14:09 | 1553882916.5478 | 929 | 0.20772 | 0.001853 | 0.000154 | 99.7572 | 99.2926 | 117.7067 | 21.0948 | 0.003117 | 0.002792 | 223.1399 | 0.26342 | 8.4039 | 0.002598 | 0.004887 | 0.004781 | 5.3e-05 | 0.000937 | 0.30342 | 1.3368 | 2.5474 | 100.19 | 20.84 | 37.258 | 96.6612 | 84.1636 | 88.8361 | 0.009168 | 0.9753 | 0.023209 | 23.7964 | 0.041749 | 0.15671 | 0.009051 | 0.62152 | 0.62152 | 0.010262
2019-03-29 14:09 | 1553882917.5478 | 930 | 0.20773 | 0.001842 | 0.000155 | 99.7964 | 99.5551 | 117.5964 | 21.095 | 0.003094 | 0.002762 | 223.1249 | 0.26331 | 8.4028 | 0.002628 | 0.00436 | 0.004271 | 5.5e-05 | 0.00094 | 0.30198 | 1.3334 | 2.5474 | 100.19 | 20.84 | 37.271 | 96.1148 | 84.8834 | 88.1136 | 0.009171 | 0.97725 | 0.023221 | 23.7614 | 0.021653 | 0.15664 | 0.00899 | 0.61756 | 0.61756 | 0.010323
2019-03-29 14:09 | 1553882918.5465 | 931 | 0.20773 | 0.001827 | 0.000158 | 99.6874 | 99.4459 | 117.5821 | 21.0963 | 0.003074 | 0.002846 | 223.0923 | 0.26361 | 8.4033 | 0.002582 | 0.004157 | 0.004039 | 5.4e-05 | 0.000936 | 0.29407 | 1.3559 | 2.5475 | 100.19 | 20.84 | 37.277 | 95.6335 | 83.7844 | 90.1336 | 0.009173 | 0.97182 | 0.023155 | 23.8263 | 0.0217 | 0.15682 | 0.009019 | 0.61765 | 0.61765 | 0.010334
2019-03-29 14:09 | 1553882919.5479 | 932 | 0.20774 | 0.001814 | 0.000159 | 99.7305 | 99.4793 | 117.7053 | 21.095 | 0.003108 | 0.002841 | 223.0921 | 0.26334 | 8.404 | 0.002642 | 0.004572 | 0.00445 | 5.6e-05 | 0.000935 | 0.30807 | 1.336 | 2.5474 | 100.19 | 20.84 | 37.288 | 96.446 | 85.2095 | 90.0018 | 0.009176 | 0.97217 | 0.023171 | 23.8345 | 0.022559 | 0.15666 | 0.008948 | 0.61307 | 0.61307 | 0.0104
2019-03-29 14:09 | 1553882920.547 | 933 | 0.20777 | 0.001804 | 0.00016 | 99.7593 | 99.465 | 117.9435 | 21.0956 | 0.00303 | 0.00278 | 223.0833 | 0.26424 | 8.4015 | 0.002609 | 0.003752 | 0.003702 | 4.8e-05 | 0.000934 | 0.31382 | 1.3351 | 2.5474 | 100.19 | 20.85 | 37.294 | 94.5578 | 84.433 | 88.5251 | 0.009183 | 0.97614 | 0.023242 | 23.81 | 0.026424 | 0.1572 | 0.00876 | 0.60175 | 0.60175 | 0.010632
2019-03-29 14:09 | 1553882921.5466 | 934 | 0.2078 | 0.001793 | 0.000163 | 99.8952 | 99.4988 | 118.2639 | 21.0958 | 0.003011 | 0.002848 | 223.07 | 0.26445 | 8.4027 | 0.002625 | 0.004004 | 0.003843 | 5.2e-05 | 0.000933 | 0.29159 | 1.3309 | 2.5475 | 100.19 | 20.85 | 37.299 | 94.1074 | 84.8146 | 90.1675 | 0.009184 | 0.97173 | 0.023221 | 23.8964 | 0.035563 | 0.15732 | 0.008647 | 0.59316 | 0.59316 | 0.010795
2019-03-29 14:09 | 1553882922.5467 | 935 | 0.20786 | 0.001781 | 0.000165 | 99.8155 | 99.4261 | 118.3247 | 21.0958 | 0.003077 | 0.002847 | 223.05 | 0.26379 | 8.4043 | 0.002624 | 0.003963 | 0.003941 | 5.1e-05 | 0.000923 | 0.27472 | 1.3085 | 2.5473 | 100.19 | 20.85 | 37.312 | 95.6864 | 84.7733 | 90.1394 | 0.009187 | 0.9718 | 0.023228 | 23.9017 | 0.034968 | 0.15693 | 0.008283 | 0.5679 | 0.5679 | 0.011247
2019-03-29 14:09 | 1553882923.5473 | 936 | 0.20785 | 0.001767 | 0.000167 | 99.7847 | 99.3844 | 118.4449 | 21.0953 | 0.003014 | 0.002804 | 223.0259 | 0.26357 | 8.4034 | 0.002544 | 0.004045 | 0.003995 | 5e-05 | 0.000924 | 0.32372 | 1.3265 | 2.5474 | 100.19 | 20.85 | 37.319 | 94.1655 | 82.8589 | 89.1168 | 0.009189 | 0.97454 | 0.023272 | 23.8801 | 0.035951 | 0.1568 | 0.008356 | 0.57396 | 0.57396 | 0.011119
2019-03-29 14:09 | 1553882924.5478 | 937 | 0.20789 | 0.001752 | 0.000169 | 99.8386 | 99.4715 | 118.3369 | 21.0962 | 0.002957 | 0.002821 | 223.0172 | 0.26419 | 8.4017 | 0.002555 | 0.004068 | 0.004022 | 5.3e-05 | 0.000927 | 0.29574 | 1.3256 | 2.5474 | 100.19 | 20.85 | 37.332 | 92.8122 | 83.1316 | 89.5326 | 0.009192 | 0.97343 | 0.023248 | 23.8829 | 0.032951 | 0.15716 | 0.008118 | 0.55664 | 0.55664 | 0.011491
2019-03-29 14:09 | 1553882925.5478 | 938 | 0.20789 | 0.001735 | 0.000172 | 99.8547 | 99.6168 | 118.2916 | 21.0957 | 0.003039 | 0.002769 | 222.98 | 0.26404 | 8.402 | 0.002529 | 0.004071 | 0.003998 | 5.6e-05 | 0.000927 | 0.2891 | 1.3391 | 2.5473 | 100.19 | 20.85 | 37.338 | 94.785 | 82.4993 | 88.2809 | 0.009194 | 0.9768 | 0.023284 | 23.8371 | 0.021335 | 0.15707 | 0.008157 | 0.56005 | 0.56005 | 0.011415
2019-03-29 14:09 | 1553882926.5468 | 939 | 0.20791 | 0.001725 | 0.000175 | 99.8811 | 99.4609 | 118.1928 | 21.0954 | 0.003009 | 0.002783 | 222.9504 | 0.2635 | 8.4029 | 0.002578 | 0.00413 | 0.004007 | 5.2e-05 | 0.000923 | 0.31205 | 1.3081 | 2.5473 | 100.19 | 20.85 | 37.343 | 94.0492 | 83.6675 | 88.6085 | 0.009195 | 0.97591 | 0.023264 | 23.8379 | 0.037709 | 0.15676 | 0.008056 | 0.5523 | 0.5523 | 0.011552
2019-03-29 14:09 | 1553882927.5472 | 940 | 0.20795 | 0.001703 | 0.000179 | 99.8655 | 99.6844 | 118.317 | 21.0955 | 0.003045 | 0.002821 | 222.9273 | 0.26325 | 8.4035 | 0.002531 | 0.004159 | 0.004117 | 5.4e-05 | 0.000922 | 0.29214 | 1.292 | 2.5473 | 100.19 | 20.85 | 37.355 | 94.9163 | 82.5356 | 89.5132 | 0.009198 | 0.97348 | 0.023247 | 23.8803 | 0.016234 | 0.15661 | 0.007803 | 0.5341 | 0.5341 | 0.011934
2019-03-29 14:09 | 1553882928.548 | 941 | 0.208 | 0.001684 | 0.000181 | 99.8547 | 99.5278 | 118.032 | 21.0961 | 0.00299 | 0.002785 | 222.9228 | 0.26428 | 8.404 | 0.002522 | 0.004199 | 0.004089 | 5.2e-05 | 0.000923 | 0.29479 | 1.3516 | 2.5475 | 100.19 | 20.85 | 37.36 | 93.589 | 82.3357 | 88.6654 | 0.009199 | 0.97576 | 0.023246 | 23.8236 | 0.029336 | 0.15722 | 0.007529 | 0.51484 | 0.51484 | 0.012429
2019-03-29 14:09 | 1553882929.5469 | 942 | 0.20795 | 0.001667 | 0.000183 | 99.9412 | 99.5848 | 118.3085 | 21.0953 | 0.002981 | 0.002771 | 222.8806 | 0.26468 | 8.4019 | 0.002619 | 0.003787 | 0.003746 | 5.5e-05 | 0.000923 | 0.29911 | 1.3656 | 2.5474 | 100.19 | 20.85 | 37.37 | 93.3714 | 84.6706 | 88.3156 | 0.009201 | 0.97671 | 0.023285 | 23.84 | 0.031961 | 0.15746 | 0.007846 | 0.53769 | 0.53769 | 0.011919
2019-03-29 14:09 | 1553882930.546 | 943 | 0.20802 | 0.001655 | 0.000185 | 99.9263 | 99.6906 | 118.3808 | 21.0954 | 0.003001 | 0.002789 | 222.8651 | 0.26479 | 8.4018 | 0.002593 | 0.004175 | 0.004065 | 5.4e-05 | 0.000922 | 0.30359 | 1.3524 | 2.5474 | 100.19 | 20.85 | 37.377 | 93.8502 | 84.0279 | 88.7437 | 0.009203 | 0.97555 | 0.023278 | 23.8614 | 0.021124 | 0.15752 | 0.007414 | 0.50722 | 0.50722 | 0.01264
2019-03-29 14:09 | 1553882931.547 | 944 | 0.20804 | 0.001635 | 0.000187 | 99.935 | 99.465 | 118.2548 | 21.0941 | 0.003053 | 0.002784 | 222.829 | 0.26357 | 8.4037 | 0.00255 | 0.003192 | 0.003128 | 5.4e-05 | 0.000922 | 0.29479 | 1.3537 | 2.5476 | 100.19 | 20.85 | 37.382 | 95.1036 | 82.9944 | 88.6245 | 0.009204 | 0.97587 | 0.023269 | 23.8447 | 0.042165 | 0.1568 | 0.007338 | 0.50168 | 0.50168 | 0.012721
2019-03-29 14:09 | 1553882932.5464 | 945 | 0.20806 | 0.001612 | 0.000187 | 99.8691 | 99.4377 | 118.1928 | 21.094 | 0.002997 | 0.00279 | 222.8279 | 0.26416 | 8.403 | 0.002557 | 0.003909 | 0.003889 | 5.3e-05 | 0.000927 | 0.28191 | 1.4142 | 2.5475 | 100.19 | 20.84 | 37.393 | 93.7668 | 83.1768 | 88.7664 | 0.009201 | 0.97549 | 0.023259 | 23.8431 | 0.038724 | 0.15715 | 0.007237 | 0.49436 | 0.49436 | 0.012938
2019-03-29 14:09 | 1553882933.547 | 946 | 0.20802 | 0.001606 | 0.00019 | 99.9077 | 99.5998 | 118.4909 | 21.0952 | 0.002959 | 0.002772 | 222.8363 | 0.26367 | 8.4016 | 0.002616 | 0.003904 | 0.003811 | 5.5e-05 | 0.000927 | 0.32875 | 1.3641 | 2.5474 | 100.19 | 20.85 | 37.398 | 92.8478 | 84.594 | 88.3496 | 0.009208 | 0.97661 | 0.023301 | 23.8594 | 0.027612 | 0.15686 | 0.007466 | 0.51109 | 0.51109 | 0.012491
2019-03-29 14:09 | 1553882934.547 | 947 | 0.20806 | 0.00159 | 0.000192 | 99.9092 | 99.7626 | 118.7052 | 21.0948 | 0.003025 | 0.002818 | 222.802 | 0.26507 | 8.4034 | 0.002545 | 0.003843 | 0.003747 | 5.6e-05 | 0.000924 | 0.29999 | 1.3604 | 2.5475 | 100.19 | 20.84 | 37.403 | 94.4469 | 82.8891 | 89.4595 | 0.009204 | 0.97362 | 0.023287 | 23.9176 | 0.013138 | 0.15769 | 0.00724 | 0.49495 | 0.49495 | 0.012967
2019-03-29 14:09 | 1553882935.5464 | 948 | 0.20809 | 0.001576 | 0.000193 | 99.9242 | 99.6083 | 119.012 | 21.0953 | 0.002969 | 0.002782 | 222.7956 | 0.2639 | 8.4015 | 0.002569 | 0.004102 | 0.004146 | 5.4e-05 | 0.000928 | 0.31526 | 1.3233 | 2.5474 | 100.19 | 20.84 | 37.418 | 93.0928 | 83.4629 | 88.5756 | 0.009208 | 0.976 | 0.023345 | 23.9193 | 0.028319 | 0.15699 | 0.007106 | 0.48674 | 0.48674 | 0.013127
2019-03-29 14:09 | 1553882936.5465 | 949 | 0.20808 | 0.001559 | 0.000195 | 99.8913 | 99.6663 | 119.1423 | 21.0947 | 0.002956 | 0.002775 | 222.7797 | 0.26411 | 8.4017 | 0.002568 | 0.003597 | 0.003481 | 5.6e-05 | 0.00093 | 0.32261 | 1.3094 | 2.5474 | 100.19 | 20.84 | 37.418 | 92.7863 | 83.4357 | 88.418 | 0.009208 | 0.97643 | 0.023363 | 23.9272 | 0.020163 | 0.15712 | 0.007194 | 0.49309 | 0.49309 | 0.012968
2019-03-29 14:09 | 1553882937.5462 | 950 | 0.20812 | 0.001546 | 0.000196 | 99.9089 | 99.6008 | 119.0479 | 21.0947 | 0.002996 | 0.002747 | 222.7636 | 0.26478 | 8.4051 | 0.002659 | 0.004205 | 0.004183 | 5.2e-05 | 0.000936 | 0.27887 | 1.3134 | 2.5473 | 100.18 | 20.84 | 37.433 | 93.7511 | 85.6206 | 87.7441 | 0.009212 | 0.97825 | 0.023376 | 23.8954 | 0.027627 | 0.15752 | 0.00697 | 0.47759 | 0.47759 | 0.013424
2019-03-29 14:09 | 1553882938.5472 | 951 | 0.20812 | 0.001532 | 0.000198 | 99.9098 | 99.6831 | 119.0869 | 21.0951 | 0.002947 | 0.002727 | 222.7487 | 0.26366 | 8.402 | 0.002602 | 0.003369 | 0.003255 | 5.4e-05 | 0.000945 | 0.30198 | 1.3036 | 2.5475 | 100.18 | 20.84 | 37.433 | 92.5585 | 84.264 | 87.2571 | 0.009212 | 0.97957 | 0.023395 | 23.8832 | 0.020324 | 0.15685 | 0.006963 | 0.47737 | 0.47737 | 0.013373
2019-03-29 14:09 | 1553882939.5476 | 952 | 0.20813 | 0.001516 | 0.000201 | 99.9433 | 99.5302 | 118.8144 | 21.0951 | 0.002963 | 0.002758 | 222.744 | 0.26437 | 8.4024 | 0.002579 | 0.004078 | 0.004022 | 5.3e-05 | 0.000947 | 0.29983 | 1.3173 | 2.5473 | 100.19 | 20.84 | 37.437 | 92.9526 | 83.6897 | 88.0054 | 0.009212 | 0.97754 | 0.023344 | 23.8806 | 0.037055 | 0.15727 | 0.006928 | 0.47374 | 0.47374 | 0.013512
2019-03-29 14:09 | 1553882940.5477 | 953 | 0.20814 | 0.001505 | 0.000203 | 99.917 | 99.5803 | 118.5426 | 21.0938 | 0.003013 | 0.002721 | 222.7215 | 0.26433 | 8.4022 | 0.002581 | 0.004791 | 0.004745 | 5.5e-05 | 0.000948 | 0.28816 | 1.3439 | 2.5474 | 100.19 | 20.84 | 37.446 | 94.1429 | 83.7469 | 87.1115 | 0.009215 | 0.97997 | 0.023347 | 23.8237 | 0.030189 | 0.15725 | 0.006902 | 0.47183 | 0.47183 | 0.013564
2019-03-29 14:09 | 1553882941.5473 | 954 | 0.20816 | 0.001489 | 0.000205 | 99.9335 | 99.7093 | 118.447 | 21.0947 | 0.002895 | 0.00273 | 222.7074 | 0.26406 | 8.4014 | 0.002609 | 0.00437 | 0.004294 | 5.3e-05 | 0.000959 | 0.30071 | 1.3594 | 2.5473 | 100.19 | 20.84 | 37.452 | 91.307 | 84.4197 | 87.3301 | 0.009216 | 0.97938 | 0.02333 | 23.8214 | 0.020083 | 0.15709 | 0.006795 | 0.4639 | 0.4639 | 0.013782
2019-03-29 14:09 | 1553882942.5467 | 955 | 0.20818 | 0.001477 | 0.000207 | 99.9592 | 99.5465 | 118.4423 | 21.0941 | 0.002915 | 0.00269 | 222.6978 | 0.26361 | 8.4039 | 0.002532 | 0.004244 | 0.004166 | 5.4e-05 | 0.000959 | 0.30773 | 1.3433 | 2.5474 | 100.19 | 20.84 | 37.459 | 91.791 | 82.5669 | 86.3788 | 0.009218 | 0.98197 | 0.02336 | 23.7894 | 0.037002 | 0.15682 | 0.006667 | 0.45537 | 0.45537 | 0.014016
2019-03-29 14:09 | 1553882943.5467 | 956 | 0.20817 | 0.001461 | 0.000212 | 99.9128 | 99.6715 | 118.1517 | 21.095 | 0.002957 | 0.002749 | 222.6867 | 0.26469 | 8.4009 | 0.002522 | 0.004348 | 0.004294 | 5.3e-05 | 0.000956 | 0.28512 | 1.3432 | 2.5474 | 100.19 | 20.84 | 37.463 | 92.7898 | 82.3332 | 87.7983 | 0.009219 | 0.97811 | 0.023286 | 23.8071 | 0.021633 | 0.15746 | 0.006747 | 0.45931 | 0.45931 | 0.013953
2019-03-29 14:09 | 1553882944.5464 | 957 | 0.20819 | 0.001456 | 0.000216 | 99.8957 | 99.5708 | 118.2351 | 21.0945 | 0.002969 | 0.002767 | 222.6649 | 0.26395 | 8.4024 | 0.002615 | 0.003735 | 0.003641 | 5.4e-05 | 0.000958 | 0.28174 | 1.3723 | 2.5474 | 100.19 | 20.84 | 37.466 | 93.0914 | 84.5693 | 88.2279 | 0.009219 | 0.97694 | 0.02328 | 23.8297 | 0.029145 | 0.15702 | 0.006644 | 0.45177 | 0.45177 | 0.014146
2019-03-29 14:09 | 1553882945.5475 | 958 | 0.20821 | 0.001444 | 0.000221 | 99.9921 | 99.61 | 118.1046 | 21.0941 | 0.00295 | 0.002776 | 222.6233 | 0.26401 | 8.4028 | 0.002663 | 0.00342 | 0.003289 | 5.3e-05 | 0.000955 | 0.29717 | 1.366 | 2.5474 | 100.19 | 20.84 | 37.472 | 92.631 | 85.7198 | 88.4478 | 0.009221 | 0.97635 | 0.02326 | 23.8237 | 0.034246 | 0.15706 | 0.006516 | 0.44213 | 0.44213 | 0.014458
2019-03-29 14:09 | 1553882946.5454 | 959 | 0.20821 | 0.001429 | 0.000225 | 99.8787 | 99.4834 | 118.2372 | 21.0948 | 0.002926 | 0.00273 | 222.6099 | 0.26404 | 8.3999 | 0.002574 | 0.003991 | 0.003933 | 5.6e-05 | 0.000953 | 0.29048 | 1.3837 | 2.5474 | 100.19 | 20.84 | 37.475 | 92.0529 | 83.5716 | 87.327 | 0.009222 | 0.97938 | 0.023309 | 23.8002 | 0.035469 | 0.15707 | 0.006507 | 0.44215 | 0.44215 | 0.014459
2019-03-29 14:09 | 1553882947.5456 | 960 | 0.20826 | 0.001415 | 0.00023 | 99.9873 | 99.5895 | 118.5937 | 21.0947 | 0.00292 | 0.002719 | 222.6153 | 0.26528 | 8.4052 | 0.002617 | 0.004191 | 0.004186 | 5.2e-05 | 0.000947 | 0.29839 | 1.3545 | 2.5474 | 100.19 | 20.84 | 37.48 | 91.909 | 84.6241 | 87.0758 | 0.009223 | 0.98007 | 0.023353 | 23.8277 | 0.035657 | 0.15781 | 0.006234 | 0.42361 | 0.42361 | 0.015162
2019-03-29 14:09 | 1553882948.5455 | 961 | 0.20831 | 0.001395 | 0.000233 | 99.9631 | 99.6943 | 119.0754 | 21.0945 | 0.002956 | 0.002696 | 222.6189 | 0.2633 | 8.4014 | 0.00263 | 0.003419 | 0.003351 | 5.4e-05 | 0.000941 | 0.2912 | 1.3658 | 2.5474 | 100.19 | 20.84 | 37.483 | 92.7674 | 84.9149 | 86.5204 | 0.009224 | 0.98158 | 0.023418 | 23.8576 | 0.024082 | 0.15664 | 0.00594 | 0.40388 | 0.40388 | 0.015785
2019-03-29 14:09 | 1553882949.5466 | 962 | 0.20833 | 0.001382 | 0.000236 | 99.9415 | 99.5397 | 119.3059 | 21.0944 | 0.002942 | 0.00268 | 222.5928 | 0.26508 | 8.4063 | 0.002658 | 0.004313 | 0.004256 | 5.2e-05 | 0.000939 | 0.28766 | 1.3421 | 2.5475 | 100.19 | 20.85 | 37.485 | 92.4317 | 85.5991 | 86.127 | 0.00923 | 0.98265 | 0.023454 | 23.8676 | 0.036035 | 0.1577 | 0.005838 | 0.39716 | 0.39716 | 0.016161
2019-03-29 14:09 | 1553882950.5455 | 963 | 0.2084 | 0.001355 | 0.00024 | 99.9649 | 99.5022 | 119.3559 | 21.0939 | 0.002985 | 0.002729 | 222.5655 | 0.26373 | 8.4016 | 0.002631 | 0.003995 | 0.003901 | 5.2e-05 | 0.000942 | 0.30879 | 1.3954 | 2.5474 | 100.19 | 20.85 | 37.489 | 93.4687 | 84.9519 | 87.3084 | 0.009231 | 0.97943 | 0.02342 | 23.9119 | 0.041499 | 0.15689 | 0.00544 | 0.36841 | 0.36841 | 0.017333
2019-03-29 14:09 | 1553882951.5456 | 964 | 0.2084 | 0.001335 | 0.000244 | 99.9667 | 99.5831 | 119.4392 | 21.093 | 0.00293 | 0.002673 | 222.5385 | 0.26358 | 8.4031 | 0.002588 | 0.004217 | 0.004154 | 5.2e-05 | 0.000951 | 0.30414 | 1.3371 | 2.5475 | 100.19 | 20.84 | 37.49 | 92.145 | 83.9186 | 85.9678 | 0.009225 | 0.98309 | 0.023472 | 23.8757 | 0.034391 | 0.1568 | 0.005477 | 0.37163 | 0.37163 | 0.017173
2019-03-29 14:09 | 1553882952.5463 | 965 | 0.20842 | 0.001305 | 0.000249 | 99.9592 | 99.5971 | 119.6413 | 21.0942 | 0.002923 | 0.002718 | 222.5184 | 0.26395 | 8.4061 | 0.002609 | 0.0044 | 0.004321 | 5e-05 | 0.000953 | 0.31094 | 1.3397 | 2.5475 | 100.19 | 20.84 | 37.491 | 91.9826 | 84.4167 | 87.0502 | 0.009226 | 0.98014 | 0.023456 | 23.9319 | 0.032465 | 0.15702 | 0.005382 | 0.36437 | 0.36437 | 0.017539
2019-03-29 14:09 | 1553882953.5457 | 966 | 0.20843 | 0.001287 | 0.000255 | 99.9909 | 99.6001 | 119.512 | 21.0945 | 0.002923 | 0.002672 | 222.4873 | 0.26487 | 8.4013 | 0.00259 | 0.003654 | 0.003637 | 5.1e-05 | 0.000951 | 0.31188 | 1.3691 | 2.5475 | 100.19 | 20.84 | 37.491 | 91.9729 | 83.9631 | 85.9389 | 0.009226 | 0.98317 | 0.02348 | 23.882 | 0.035028 | 0.15757 | 0.005329 | 0.36061 | 0.36061 | 0.017784
2019-03-29 14:09 | 1553882954.5456 | 967 | 0.20847 | 0.001263 | 0.00026 | 100.0224 | 99.4192 | 119.0369 | 21.0953 | 0.002872 | 0.002703 | 222.4553 | 0.26397 | 8.4027 | 0.002574 | 0.004404 | 0.004373 | 5.1e-05 | 0.000947 | 0.29142 | 1.3993 | 2.5474 | 100.18 | 20.84 | 37.49 | 90.7572 | 83.5791 | 86.692 | 0.009226 | 0.98111 | 0.023409 | 23.8595 | 0.054098 | 0.15704 | 0.005157 | 0.34716 | 0.34716 | 0.018411
2019-03-29 14:09 | 1553882955.5467 | 968 | 0.20847 | 0.001238 | 0.000264 | 99.92 | 99.4633 | 119.1028 | 21.0946 | 0.002891 | 0.002672 | 222.4433 | 0.26439 | 8.4019 | 0.002585 | 0.004238 | 0.004122 | 5e-05 | 0.000954 | 0.29087 | 1.3556 | 2.5475 | 100.18 | 20.84 | 37.489 | 91.2091 | 83.8343 | 85.9265 | 0.009226 | 0.9832 | 0.02344 | 23.8407 | 0.040978 | 0.15728 | 0.005142 | 0.34637 | 0.34637 | 0.018481
2019-03-29 14:09 | 1553882956.5468 | 969 | 0.20847 | 0.001217 | 0.000267 | 99.9849 | 99.5872 | 119.0073 | 21.0946 | 0.002924 | 0.002699 | 222.4494 | 0.26377 | 8.4011 | 0.0026 | 0.004374 | 0.00428 | 5.1e-05 | 0.000957 | 0.29717 | 1.3877 | 2.5474 | 100.18 | 20.84 | 37.487 | 91.9924 | 84.2134 | 86.5883 | 0.009226 | 0.9814 | 0.023409 | 23.8531 | 0.035657 | 0.15692 | 0.005207 | 0.35028 | 0.35028 | 0.018232
2019-03-29 14:09 | 1553882957.5468 | 970 | 0.20846 | 0.001198 | 0.000269 | 99.9694 | 99.3605 | 118.996 | 21.0945 | 0.002918 | 0.002648 | 222.4473 | 0.26354 | 8.4049 | 0.002598 | 0.004254 | 0.004184 | 5e-05 | 0.000956 | 0.2959 | 1.3427 | 2.5473 | 100.18 | 20.84 | 37.485 | 91.8492 | 84.1462 | 85.3697 | 0.009225 | 0.98473 | 0.023448 | 23.8115 | 0.054641 | 0.15678 | 0.005254 | 0.35409 | 0.35409 | 0.018021
2019-03-29 14:09 | 1553882958.5469 | 971 | 0.20846 | 0.00118 | 0.000271 | 99.9254 | 99.3455 | 118.6261 | 21.0953 | 0.002899 | 0.002689 | 222.4121 | 0.26323 | 8.4034 | 0.002616 | 0.00437 | 0.004301 | 5.2e-05 | 0.000957 | 0.30807 | 1.3564 | 2.5474 | 100.18 | 20.84 | 37.479 | 91.4069 | 84.5841 | 86.347 | 0.009224 | 0.98205 | 0.02338 | 23.8068 | 0.052055 | 0.1566 | 0.005257 | 0.35308 | 0.35308 | 0.018051
2019-03-29 14:09 | 1553882959.5477 | 972 | 0.20849 | 0.001165 | 0.000274 | 100.0239 | 99.525 | 118.4851 | 21.0939 | 0.002854 | 0.002659 | 222.395 | 0.26342 | 8.4005 | 0.002576 | 0.003627 | 0.003523 | 5.2e-05 | 0.000957 | 0.31454 | 1.3837 | 2.5474 | 100.18 | 20.84 | 37.476 | 90.3056 | 83.6176 | 85.6323 | 0.009223 | 0.98401 | 0.023389 | 23.769 | 0.044718 | 0.15671 | 0.005092 | 0.3415 | 0.3415 | 0.018677
2019-03-29 14:09 | 1553882960.5469 | 973 | 0.20846 | 0.001151 | 0.000276 | 99.9655 | 99.6025 | 118.4886 | 21.0953 | 0.002949 | 0.002646 | 222.3992 | 0.26401 | 8.402 | 0.002542 | 0.004825 | 0.004739 | 5.3e-05 | 0.000955 | 0.29551 | 1.32 | 2.5474 | 100.18 | 20.84 | 37.468 | 92.6165 | 82.8022 | 85.3012 | 0.009221 | 0.98492 | 0.0234 | 23.7584 | 0.032537 | 0.15706 | 0.00531 | 0.35684 | 0.35684 | 0.017914
2019-03-29 14:09 | 1553882961.5459 | 974 | 0.20849 | 0.001139 | 0.000278 | 100.0478 | 99.4711 | 118.5144 | 21.0946 | 0.002882 | 0.002663 | 222.4184 | 0.26337 | 8.4028 | 0.002598 | 0.003917 | 0.003861 | 5e-05 | 0.000955 | 0.31277 | 1.3104 | 2.5474 | 100.18 | 20.84 | 37.464 | 90.9997 | 84.1676 | 85.7088 | 0.00922 | 0.9838 | 0.023389 | 23.7745 | 0.051707 | 0.15668 | 0.005168 | 0.34661 | 0.34661 | 0.018398
2019-03-29 14:09 | 1553882962.5453 | 975 | 0.2085 | 0.001134 | 0.000278 | 100.0625 | 99.5575 | 118.1618 | 21.0947 | 0.002903 | 0.002633 | 222.3965 | 0.26331 | 8.4031 | 0.002604 | 0.003863 | 0.003804 | 5.5e-05 | 0.000947 | 0.30055 | 1.3418 | 2.5474 | 100.18 | 20.84 | 37.458 | 91.4992 | 84.3025 | 84.9954 | 0.009218 | 0.98576 | 0.023378 | 23.7155 | 0.045257 | 0.15665 | 0.00506 | 0.33887 | 0.33887 | 0.018814
2019-03-29 14:09 | 1553882963.5461 | 976 | 0.20848 | 0.00112 | 0.000279 | 100.0044 | 99.5575 | 118.186 | 21.0946 | 0.002848 | 0.002674 | 222.3874 | 0.26346 | 8.4013 | 0.002547 | 0.003903 | 0.003866 | 5.1e-05 | 0.000946 | 0.26648 | 1.3508 | 2.5473 | 100.18 | 20.84 | 37.448 | 90.1647 | 82.9316 | 85.9779 | 0.009216 | 0.98306 | 0.023348 | 23.7504 | 0.040065 | 0.15673 | 0.005223 | 0.3498 | 0.3498 | 0.018236
2019-03-29 14:09 | 1553882964.5459 | 977 | 0.20847 | 0.001115 | 0.000278 | 100.0403 | 99.7223 | 118.421 | 21.0942 | 0.002921 | 0.002614 | 222.3845 | 0.26326 | 8.4008 | 0.002549 | 0.004231 | 0.004096 | 5.3e-05 | 0.000949 | 0.28224 | 1.3543 | 2.5475 | 100.18 | 20.84 | 37.442 | 91.9357 | 82.982 | 84.5479 | 0.009214 | 0.98699 | 0.023418 | 23.7266 | 0.028479 | 0.15661 | 0.005307 | 0.35684 | 0.35684 | 0.017863
2019-03-29 14:09 | 1553882965.5461 | 978 | 0.20846 | 0.001112 | 0.000277 | 100.0355 | 99.6756 | 118.4616 | 21.0945 | 0.002893 | 0.002656 | 222.3822 | 0.26321 | 8.4015 | 0.002527 | 0.003502 | 0.00344 | 5.2e-05 | 0.000947 | 0.28423 | 1.3683 | 2.5475 | 100.18 | 20.84 | 37.431 | 91.2511 | 82.4457 | 85.5588 | 0.009212 | 0.98421 | 0.023389 | 23.7642 | 0.032246 | 0.15659 | 0.005345 | 0.35911 | 0.35911 | 0.017747
2019-03-29 14:09 | 1553882966.545 | 979 | 0.2085 | 0.001109 | 0.000276 | 100.0373 | 99.5708 | 118.9397 | 21.0947 | 0.002842 | 0.002599 | 222.3789 | 0.2641 | 8.4004 | 0.002575 | 0.003698 | 0.003616 | 5.4e-05 | 0.000942 | 0.26128 | 1.3441 | 2.5474 | 100.18 | 20.84 | 37.425 | 90.0188 | 83.6073 | 84.1832 | 0.00921 | 0.988 | 0.023481 | 23.7664 | 0.041813 | 0.15711 | 0.005099 | 0.34321 | 0.34321 | 0.018632
2019-03-29 14:09 | 1553882967.5455 | 980 | 0.20853 | 0.001102 | 0.000274 | 100.0484 | 99.65 | 119.0653 | 21.0949 | 0.002813 | 0.002656 | 222.3568 | 0.26455 | 8.404 | 0.002537 | 0.004538 | 0.004417 | 5.4e-05 | 0.000939 | 0.25713 | 1.3558 | 2.5475 | 100.18 | 20.84 | 37.42 | 89.336 | 82.6816 | 85.5543 | 0.009209 | 0.98422 | 0.023449 | 23.8246 | 0.035694 | 0.15738 | 0.004944 | 0.33193 | 0.33193 | 0.019298
2019-03-29 14:09 | 1553882968.5461 | 981 | 0.20852 | 0.001098 | 0.000272 | 100.075 | 99.5227 | 119.3312 | 21.0953 | 0.002855 | 0.002606 | 222.3476 | 0.264 | 8.4023 | 0.002567 | 0.004125 | 0.004121 | 5.1e-05 | 0.000941 | 0.27743 | 1.3903 | 2.5475 | 100.18 | 20.84 | 37.409 | 90.353 | 83.4122 | 84.3553 | 0.009206 | 0.98752 | 0.023514 | 23.8112 | 0.049509 | 0.15706 | 0.005016 | 0.33809 | 0.33809 | 0.018906
2019-03-29 14:09 | 1553882969.5462 | 982 | 0.20854 | 0.001092 | 0.00027 | 100.0149 | 99.8213 | 119.3857 | 21.0947 | 0.002854 | 0.002651 | 222.3343 | 0.26323 | 8.4041 | 0.002616 | 0.004338 | 0.004223 | 5.5e-05 | 0.000938 | 0.28031 | 1.3891 | 2.5474 | 100.18 | 20.84 | 37.404 | 90.3221 | 84.5929 | 85.4349 | 0.009205 | 0.98455 | 0.023484 | 23.8526 | 0.017331 | 0.1566 | 0.004925 | 0.33138 | 0.33138 | 0.019233
2019-03-29 14:09 | 1553882970.5456 | 983 | 0.20855 | 0.001092 | 0.000269 | 100.0053 | 99.6322 | 119.608 | 21.0948 | 0.002869 | 0.002686 | 222.3219 | 0.26361 | 8.4034 | 0.002567 | 0.00396 | 0.003912 | 5.5e-05 | 0.000938 | 0.30397 | 1.392 | 2.5476 | 100.18 | 20.84 | 37.399 | 90.6715 | 83.4188 | 86.2647 | 0.009204 | 0.98228 | 0.023479 | 23.9024 | 0.033433 | 0.15682 | 0.004836 | 0.32508 | 0.32508 | 0.019634
2019-03-29 14:10 | 1553882971.5463 | 984 | 0.20853 | 0.001093 | 0.000269 | 100.0505 | 99.6895 | 119.5669 | 21.0954 | 0.002836 | 0.002622 | 222.3002 | 0.2637 | 8.4014 | 0.002588 | 0.003946 | 0.003884 | 5.7e-05 | 0.000941 | 0.2891 | 1.3575 | 2.5474 | 100.18 | 20.84 | 37.395 | 89.8858 | 83.9093 | 84.7346 | 0.009203 | 0.98648 | 0.023525 | 23.8474 | 0.032328 | 0.15687 | 0.004975 | 0.33552 | 0.33552 | 0.019029
2019-03-29 14:10 | 1553882972.5454 | 985 | 0.20854 | 0.001091 | 0.00027 | 100.0538 | 99.6892 | 119.726 | 21.0953 | 0.002842 | 0.002669 | 222.2806 | 0.26357 | 8.4008 | 0.002557 | 0.004118 | 0.003968 | 5.5e-05 | 0.000939 | 0.28584 | 1.3722 | 2.5473 | 100.18 | 20.84 | 37.391 | 90.0404 | 83.179 | 85.8702 | 0.009202 | 0.98336 | 0.023503 | 23.9011 | 0.032653 | 0.1568 | 0.004908 | 0.33044 | 0.33044 | 0.019312
2019-03-29 14:10 | 1553882973.5457 | 986 | 0.20857 | 0.001087 | 0.000272 | 100.1083 | 99.6878 | 119.8497 | 21.0944 | 0.002806 | 0.00263 | 222.2575 | 0.26366 | 8.4014 | 0.00248 | 0.004522 | 0.004469 | 5.5e-05 | 0.000936 | 0.29999 | 1.3645 | 2.5472 | 100.18 | 20.84 | 37.39 | 89.1615 | 81.325 | 84.9141 | 0.009202 | 0.98598 | 0.023547 | 23.8816 | 0.037645 | 0.15685 | 0.004761 | 0.32056 | 0.32056 | 0.019915
2019-03-29 14:10 | 1553882974.5464 | 987 | 0.20854 | 0.001085 | 0.000273 | 100.0361 | 99.8011 | 119.854 | 21.0952 | 0.002806 | 0.002616 | 222.2762 | 0.264 | 8.401 | 0.00254 | 0.004175 | 0.004149 | 5.4e-05 | 0.00093 | 0.31636 | 1.3771 | 2.5474 | 100.18 | 20.83 | 37.392 | 89.1685 | 82.7553 | 84.5815 | 0.009196 | 0.9869 | 0.023558 | 23.8709 | 0.021036 | 0.15706 | 0.004929 | 0.33251 | 0.33251 | 0.019224
2019-03-29 14:10 | 1553882975.5454 | 988 | 0.20854 | 0.001085 | 0.000275 | 100.0526 | 99.7226 | 119.8852 | 21.0952 | 0.002821 | 0.00261 | 222.2565 | 0.26429 | 8.4029 | 0.002493 | 0.004088 | 0.00404 | 5.3e-05 | 0.000926 | 0.27472 | 1.3917 | 2.5475 | 100.18 | 20.83 | 37.398 | 89.5292 | 81.6291 | 84.4466 | 0.009198 | 0.98727 | 0.023566 | 23.8695 | 0.029546 | 0.15723 | 0.00489 | 0.32976 | 0.32976 | 0.019406
2019-03-29 14:10 | 1553882976.5464 | 989 | 0.20856 | 0.001084 | 0.000275 | 100.0352 | 99.7445 | 119.8598 | 21.0952 | 0.002864 | 0.002647 | 222.2157 | 0.26422 | 8.4018 | 0.002454 | 0.003767 | 0.003655 | 5.4e-05 | 0.000921 | 0.29789 | 1.347 | 2.5473 | 100.18 | 20.83 | 37.403 | 90.5657 | 80.6855 | 85.3332 | 0.009199 | 0.98483 | 0.023534 | 23.8965 | 0.026035 | 0.15718 | 0.004818 | 0.32425 | 0.32425 | 0.01973
2019-03-29 14:10 | 1553882977.5448 | 990 | 0.20853 | 0.001076 | 0.000275 | 100.134 | 99.6619 | 119.8657 | 21.095 | 0.002812 | 0.002599 | 222.1835 | 0.26433 | 8.4008 | 0.002502 | 0.004206 | 0.004184 | 5.6e-05 | 0.000922 | 0.28854 | 1.3096 | 2.5474 | 100.18 | 20.83 | 37.414 | 89.2971 | 81.8464 | 84.1701 | 0.009202 | 0.98804 | 0.023573 | 23.8583 | 0.042272 | 0.15725 | 0.004998 | 0.33746 | 0.33746 | 0.018966
2019-03-29 14:10 | 1553882978.5454 | 991 | 0.20858 | 0.001077 | 0.000274 | 100.0251 | 99.8714 | 119.7535 | 21.0954 | 0.002815 | 0.002633 | 222.1681 | 0.26365 | 8.4021 | 0.002515 | 0.003596 | 0.00349 | 5.1e-05 | 0.000921 | 0.27831 | 1.3444 | 2.5475 | 100.18 | 20.83 | 37.419 | 89.3693 | 82.1602 | 84.9922 | 0.009203 | 0.98577 | 0.023535 | 23.8746 | 0.013747 | 0.15684 | 0.004686 | 0.31496 | 0.31496 | 0.020268
2019-03-29 14:10 | 1553882979.5445 | 992 | 0.20857 | 0.001078 | 0.000271 | 100.0685 | 99.7226 | 119.6094 | 21.0955 | 0.002853 | 0.00261 | 222.1745 | 0.26469 | 8.4027 | 0.002527 | 0.004078 | 0.003969 | 5.1e-05 | 0.000921 | 0.30055 | 1.3609 | 2.5474 | 100.18 | 20.83 | 37.423 | 90.2954 | 82.4514 | 84.4545 | 0.009204 | 0.98725 | 0.023538 | 23.8423 | 0.030964 | 0.15746 | 0.004753 | 0.31991 | 0.31991 | 0.020033
2019-03-29 14:10 | 1553882980.5459 | 993 | 0.20862 | 0.001072 | 0.000267 | 100.0774 | 99.8244 | 119.6354 | 21.0947 | 0.002831 | 0.002654 | 222.1414 | 0.26359 | 8.4022 | 0.00253 | 0.003665 | 0.003636 | 5.4e-05 | 0.000925 | 0.29767 | 1.3874 | 2.5474 | 100.18 | 20.83 | 37.431 | 89.7734 | 82.5209 | 85.5114 | 0.009206 | 0.98434 | 0.023506 | 23.8801 | 0.022649 | 0.15681 | 0.004486 | 0.3009 | 0.3009 | 0.021211
2019-03-29 14:10 | 1553882981.5456 | 994 | 0.20862 | 0.001068 | 0.000262 | 100.0493 | 99.6476 | 119.807 | 21.0937 | 0.002836 | 0.002643 | 222.1094 | 0.26333 | 8.4039 | 0.002529 | 0.003358 | 0.003226 | 5.4e-05 | 0.000928 | 0.3131 | 1.4162 | 2.5476 | 100.18 | 20.83 | 37.434 | 89.8717 | 82.5033 | 85.2472 | 0.009207 | 0.98507 | 0.023532 | 23.8884 | 0.035988 | 0.15666 | 0.004493 | 0.30202 | 0.30202 | 0.021111
2019-03-29 14:10 | 1553882982.5456 | 995 | 0.20862 | 0.001065 | 0.000258 | 100.0873 | 99.6052 | 119.5922 | 21.095 | 0.00282 | 0.002633 | 222.1018 | 0.26487 | 8.4007 | 0.002435 | 0.003887 | 0.003808 | 5.3e-05 | 0.000924 | 0.30486 | 1.3808 | 2.5474 | 100.18 | 20.83 | 37.44 | 89.4907 | 80.2346 | 84.9895 | 0.009208 | 0.98578 | 0.023519 | 23.8584 | 0.043186 | 0.15757 | 0.004495 | 0.30223 | 0.30223 | 0.021219
2019-03-29 14:10 | 1553882983.5447 | 996 | 0.20863 | 0.001056 | 0.000255 | 100.0783 | 99.8155 | 119.4127 | 21.0953 | 0.002816 | 0.002604 | 222.0914 | 0.26427 | 8.4003 | 0.002585 | 0.004048 | 0.003975 | 5.3e-05 | 0.000923 | 0.28385 | 1.3424 | 2.5473 | 100.18 | 20.83 | 37.442 | 89.4137 | 83.846 | 84.3069 | 0.009209 | 0.98766 | 0.023524 | 23.8178 | 0.023524 | 0.15721 | 0.004424 | 0.2974 | 0.2974 | 0.021515
2019-03-29 14:10 | 1553882984.5447 | 997 | 0.20863 | 0.001054 | 0.000254 | 100.1496 | 99.8131 | 119.5498 | 21.094 | 0.002849 | 0.002644 | 222.0593 | 0.26373 | 8.4032 | 0.002528 | 0.004066 | 0.004007 | 5.6e-05 | 0.000927 | 0.29374 | 1.3244 | 2.5474 | 100.18 | 20.83 | 37.444 | 90.1979 | 82.4608 | 85.2632 | 0.009209 | 0.98502 | 0.023506 | 23.8633 | 0.030103 | 0.15689 | 0.004475 | 0.30086 | 0.30086 | 0.021224
2019-03-29 14:10 | 1553882985.5457 | 998 | 0.20863 | 0.001046 | 0.000254 | 100.0912 | 99.9291 | 119.5643 | 21.0936 | 0.002724 | 0.002609 | 222.0433 | 0.2643 | 8.4011 | 0.002573 | 0.00373 | 0.003604 | 5.7e-05 | 0.000922 | 0.29264 | 1.3319 | 2.5474 | 100.18 | 20.83 | 37.443 | 87.1964 | 83.5635 | 84.4318 | 0.009209 | 0.98731 | 0.023535 | 23.837 | 0.014496 | 0.15723 | 0.004423 | 0.29756 | 0.29756 | 0.021506
2019-03-29 14:10 | 1553882986.5447 | 999 | 0.2086 | 0.001047 | 0.000254 | 100.1274 | 99.68 | 119.2122 | 21.0942 | 0.002768 | 0.002677 | 222.0268 | 0.2645 | 8.4026 | 0.002459 | 0.004004 | 0.00396 | 5.5e-05 | 0.000926 | 0.27793 | 1.3374 | 2.5473 | 100.18 | 20.82 | 37.441 | 88.245 | 80.8054 | 86.0541 | 0.009203 | 0.98286 | 0.023447 | 23.8559 | 0.04006 | 0.15735 | 0.00461 | 0.30965 | 0.30965 | 0.020682
2019-03-29 14:10 | 1553882987.5448 | 1000 | 0.20863 | 0.001043 | 0.000254 | 100.1849 | 99.6479 | 118.9392 | 21.0943 | 0.002829 | 0.002667 | 222.0324 | 0.26383 | 8.4042 | 0.002589 | 0.003541 | 0.003441 | 5.4e-05 | 0.000936 | 0.28871 | 1.3476 | 2.5473 | 100.18 | 20.83 | 37.438 | 89.7072 | 83.9371 | 85.8097 | 0.009208 | 0.98352 | 0.023428 | 23.8204 | 0.048072 | 0.15696 | 0.004482 | 0.3004 | 0.3004 | 0.021265
2019-03-29 14:10 | 1553882988.5447 | 1001 | 0.20862 | 0.001042 | 0.000252 | 100.0924 | 99.7469 | 118.8055 | 21.0935 | 0.002785 | 0.002589 | 222.0263 | 0.26392 | 8.4008 | 0.002572 | 0.003783 | 0.003679 | 5.4e-05 | 0.000935 | 0.26665 | 1.3664 | 2.5473 | 100.18 | 20.82 | 37.432 | 88.6581 | 83.5259 | 83.9394 | 0.009201 | 0.98867 | 0.023476 | 23.7449 | 0.030931 | 0.15701 | 0.004532 | 0.30461 | 0.30461 | 0.020978
2019-03-29 14:10 | 1553882989.5462 | 1002 | 0.20861 | 0.001043 | 0.000251 | 100.0951 | 99.8407 | 118.9118 | 21.0938 | 0.002814 | 0.002581 | 222.0037 | 0.26373 | 8.4044 | 0.002638 | 0.00415 | 0.004058 | 5.8e-05 | 0.000933 | 0.29048 | 1.3363 | 2.5473 | 100.18 | 20.83 | 37.43 | 89.3646 | 85.11 | 83.7375 | 0.009206 | 0.98923 | 0.023493 | 23.7488 | 0.022759 | 0.15689 | 0.004582 | 0.30839 | 0.30839 | 0.020706
2019-03-29 14:10 | 1553882990.5449 | 1003 | 0.2086 | 0.001045 | 0.00025 | 100.1292 | 99.5759 | 118.8418 | 21.0936 | 0.002808 | 0.002648 | 221.9918 | 0.26449 | 8.4005 | 0.002484 | 0.003701 | 0.003695 | 5.5e-05 | 0.000941 | 0.29048 | 1.3293 | 2.5473 | 100.18 | 20.82 | 37.427 | 89.1966 | 81.4232 | 85.3511 | 0.009199 | 0.98478 | 0.023433 | 23.7954 | 0.049566 | 0.15735 | 0.004664 | 0.31343 | 0.31343 | 0.020432
2019-03-29 14:10 | 1553882991.5459 | 1004 | 0.20859 | 0.001047 | 0.000249 | 100.1418 | 99.8158 | 119.0345 | 21.0937 | 0.002864 | 0.002548 | 221.9792 | 0.26402 | 8.4025 | 0.002572 | 0.004477 | 0.004314 | 5.5e-05 | 0.000946 | 0.29446 | 1.3103 | 2.5475 | 100.18 | 20.83 | 37.427 | 90.5517 | 83.5347 | 82.9426 | 0.009205 | 0.99144 | 0.023531 | 23.7346 | 0.029163 | 0.15706 | 0.004693 | 0.31686 | 0.31686 | 0.020175
2019-03-29 14:10 | 1553882992.5444 | 1005 | 0.2086 | 0.001044 | 0.000248 | 100.1388 | 99.8046 | 119.1465 | 21.0938 | 0.002808 | 0.002591 | 221.9697 | 0.26402 | 8.4018 | 0.002566 | 0.003737 | 0.003697 | 5.5e-05 | 0.000944 | 0.30541 | 1.3304 | 2.5474 | 100.18 | 20.83 | 37.426 | 89.2187 | 83.3966 | 83.9756 | 0.009205 | 0.98857 | 0.023508 | 23.7802 | 0.029905 | 0.15707 | 0.00467 | 0.31496 | 0.31496 | 0.020297
2019-03-29 14:10 | 1553882993.5453 | 1006 | 0.20861 | 0.001047 | 0.000248 | 100.1966 | 99.7885 | 119.2467 | 21.0939 | 0.002798 | 0.002545 | 221.9561 | 0.26522 | 8.4019 | 0.00255 | 0.0038 | 0.003787 | 5.4e-05 | 0.000942 | 0.26985 | 1.3107 | 2.5473 | 100.18 | 20.83 | 37.423 | 88.9679 | 82.9959 | 82.8913 | 0.009204 | 0.99159 | 0.023554 | 23.754 | 0.036502 | 0.15778 | 0.004595 | 0.31035 | 0.31035 | 0.020692
2019-03-29 14:10 | 1553882994.544 | 1007 | 0.20858 | 0.001049 | 0.000248 | 100.1978 | 99.9598 | 119.0153 | 21.0945 | 0.002863 | 0.002558 | 221.9495 | 0.26409 | 8.4044 | 0.002593 | 0.004458 | 0.0043 | 5.3e-05 | 0.000947 | 0.30359 | 1.3722 | 2.5475 | 100.18 | 20.83 | 37.421 | 90.5399 | 84.0395 | 83.197 | 0.009204 | 0.99074 | 0.023521 | 23.7412 | 0.021265 | 0.15711 | 0.004747 | 0.32062 | 0.32062 | 0.019943
2019-03-29 14:10 | 1553882995.545 | 1008 | 0.20861 | 0.001046 | 0.000248 | 100.2328 | 99.9885 | 118.566 | 21.0944 | 0.002803 | 0.002639 | 221.9387 | 0.26414 | 8.4021 | 0.002622 | 0.004388 | 0.004325 | 5.3e-05 | 0.000949 | 0.29894 | 1.3463 | 2.5475 | 100.18 | 20.83 | 37.416 | 89.0839 | 84.7313 | 85.1404 | 0.009202 | 0.98536 | 0.023413 | 23.7608 | 0.021826 | 0.15714 | 0.004563 | 0.30617 | 0.30617 | 0.020888
2019-03-29 14:10 | 1553882996.5442 | 1009 | 0.20861 | 0.00105 | 0.000248 | 100.2382 | 99.8332 | 118.609 | 21.0938 | 0.002791 | 0.002611 | 221.9391 | 0.26441 | 8.4041 | 0.002569 | 0.004524 | 0.004449 | 5.4e-05 | 0.000943 | 0.30989 | 1.337 | 2.5474 | 100.18 | 20.83 | 37.413 | 88.8007 | 83.449 | 84.4733 | 0.009202 | 0.9872 | 0.023439 | 23.743 | 0.036209 | 0.1573 | 0.004601 | 0.30923 | 0.30923 | 0.020703
2019-03-29 14:10 | 1553882997.5451 | 1010 | 0.20861 | 0.001053 | 0.000248 | 100.1445 | 99.923 | 118.9385 | 21.0946 | 0.002737 | 0.002591 | 221.9199 | 0.26409 | 8.4042 | 0.002611 | 0.003977 | 0.003914 | 5.2e-05 | 0.000949 | 0.30071 | 1.3669 | 2.5475 | 100.18 | 20.83 | 37.408 | 87.4972 | 84.4741 | 83.9826 | 0.0092 | 0.98855 | 0.023488 | 23.7596 | 0.019806 | 0.15711 | 0.004577 | 0.30823 | 0.30823 | 0.020746
2019-03-29 14:10 | 1553882998.5447 | 1011 | 0.20861 | 0.001052 | 0.000246 | 100.1927 | 99.8994 | 119.3082 | 21.0939 | 0.00276 | 0.002602 | 221.9137 | 0.26416 | 8.4035 | 0.00258 | 0.003969 | 0.003854 | 5.2e-05 | 0.000948 | 0.29607 | 1.3662 | 2.5476 | 100.18 | 20.83 | 37.4 | 88.0533 | 83.7306 | 84.2496 | 0.009198 | 0.98782 | 0.023515 | 23.8054 | 0.026218 | 0.15715 | 0.004601 | 0.31032 | 0.31032 | 0.020611
2019-03-29 14:10 | 1553882999.5452 | 1012 | 0.20862 | 0.001053 | 0.000245 | 100.1651 | 99.8865 | 119.5681 | 21.0943 | 0.002805 | 0.002603 | 221.9006 | 0.26405 | 8.4019 | 0.002638 | 0.004244 | 0.004145 | 5.1e-05 | 0.000944 | 0.3095 | 1.3381 | 2.5474 | 100.18 | 20.83 | 37.396 | 89.1464 | 85.13 | 84.2827 | 0.009197 | 0.98773 | 0.02354 | 23.8324 | 0.02492 | 0.15708 | 0.004498 | 0.30347 | 0.30347 | 0.021067
2019-03-29 14:10 | 1553883000.5445 | 1013 | 0.20859 | 0.00105 | 0.000243 | 100.1816 | 99.9626 | 119.6514 | 21.094 | 0.002787 | 0.002576 | 221.893 | 0.26479 | 8.4046 | 0.002618 | 0.003872 | 0.00387 | 5.3e-05 | 0.000939 | 0.27831 | 1.3325 | 2.5473 | 100.18 | 20.83 | 37.392 | 88.708 | 84.6374 | 83.6154 | 0.009196 | 0.98957 | 0.02357 | 23.8185 | 0.019578 | 0.15752 | 0.004699 | 0.3182 | 0.3182 | 0.020148
2019-03-29 14:10 | 1553883001.5442 | 1014 | 0.20861 | 0.001051 | 0.000242 | 100.2106 | 99.881 | 119.2775 | 21.0942 | 0.002799 | 0.002578 | 221.9114 | 0.26407 | 8.404 | 0.002638 | 0.004555 | 0.004402 | 5.3e-05 | 0.000941 | 0.28622 | 1.3226 | 2.5473 | 100.18 | 20.83 | 37.381 | 89.0026 | 85.126 | 83.6701 | 0.009194 | 0.98942 | 0.023531 | 23.7831 | 0.029471 | 0.15709 | 0.00458 | 0.30933 | 0.30933 | 0.02067
2019-03-29 14:10 | 1553883002.5443 | 1015 | 0.20867 | 0.001053 | 0.000241 | 100.2307 | 100.0039 | 119.4739 | 21.0946 | 0.002796 | 0.002601 | 221.9383 | 0.26523 | 8.4008 | 0.002572 | 0.00416 | 0.004143 | 5.6e-05 | 0.000938 | 0.26808 | 1.3131 | 2.5473 | 100.18 | 20.83 | 37.374 | 88.9108 | 83.5238 | 84.2315 | 0.009192 | 0.98787 | 0.023532 | 23.8214 | 0.020265 | 0.15779 | 0.004189 | 0.28176 | 0.28176 | 0.022792
2019-03-29 14:10 | 1553883003.5453 | 1016 | 0.20859 | 0.001049 | 0.000241 | 100.2005 | 99.8984 | 119.3467 | 21.0935 | 0.002825 | 0.002636 | 221.9446 | 0.26461 | 8.4024 | 0.00264 | 0.0044 | 0.004266 | 5.7e-05 | 0.000943 | 0.29607 | 1.3609 | 2.5475 | 100.18 | 20.83 | 37.363 | 89.6208 | 85.1605 | 85.0693 | 0.009189 | 0.98556 | 0.023492 | 23.8366 | 0.027004 | 0.15741 | 0.004737 | 0.31997 | 0.31997 | 0.020023
2019-03-29 14:10 | 1553883004.5443 | 1017 | 0.2086 | 0.001046 | 0.00024 | 100.2166 | 99.9308 | 119.4195 | 21.0948 | 0.00279 | 0.002614 | 221.922 | 0.264 | 8.4008 | 0.002584 | 0.003738 | 0.003668 | 5.4e-05 | 0.000948 | 0.33483 | 1.3694 | 2.5474 | 100.18 | 20.83 | 37.358 | 88.779 | 83.8094 | 84.5415 | 0.009188 | 0.98701 | 0.023517 | 23.8263 | 0.025544 | 0.15705 | 0.004633 | 0.313 | 0.313 | 0.020421
2019-03-29 14:10 | 1553883005.5454 | 1018 | 0.20862 | 0.001047 | 0.000238 | 100.2391 | 100.0001 | 119.7958 | 21.095 | 0.002691 | 0.0026 | 221.91 | 0.265 | 8.4009 | 0.002592 | 0.004349 | 0.004247 | 5.5e-05 | 0.000947 | 0.28224 | 1.3845 | 2.5473 | 100.18 | 20.83 | 37.354 | 86.385 | 84.0071 | 84.2031 | 0.009187 | 0.98795 | 0.023565 | 23.8525 | 0.021349 | 0.15765 | 0.004559 | 0.30849 | 0.30849 | 0.020799
2019-03-29 14:10 | 1553883006.5454 | 1019 | 0.20867 | 0.001049 | 0.000237 | 100.219 | 99.9175 | 119.7488 | 21.0946 | 0.002751 | 0.002612 | 221.9012 | 0.26412 | 8.4018 | 0.00259 | 0.00429 | 0.00423 | 5.4e-05 | 0.000946 | 0.29247 | 1.344 | 2.5474 | 100.18 | 20.83 | 37.347 | 87.84 | 83.974 | 84.4889 | 0.009185 | 0.98716 | 0.023551 | 23.8573 | 0.02695 | 0.15712 | 0.004218 | 0.28428 | 0.28428 | 0.022495
2019-03-29 14:10 | 1553883007.5458 | 1020 | 0.20868 | 0.001047 | 0.000235 | 100.26 | 99.9322 | 119.608 | 21.0943 | 0.002776 | 0.002597 | 221.888 | 0.26355 | 8.4009 | 0.002557 | 0.004199 | 0.004143 | 5.6e-05 | 0.00095 | 0.29336 | 1.3366 | 2.5475 | 100.18 | 20.83 | 37.344 | 88.4285 | 83.1593 | 84.1406 | 0.009185 | 0.98812 | 0.023549 | 23.8317 | 0.029295 | 0.15678 | 0.004205 | 0.28345 | 0.28345 | 0.022512
2019-03-29 14:10 | 1553883008.5434 | 1021 | 0.20862 | 0.001045 | 0.000233 | 100.2487 | 99.8875 | 119.4819 | 21.0938 | 0.002745 | 0.00265 | 221.8728 | 0.2637 | 8.4037 | 0.002564 | 0.004446 | 0.004311 | 5.2e-05 | 0.000949 | 0.29822 | 1.326 | 2.5472 | 100.18 | 20.83 | 37.341 | 87.7011 | 83.3483 | 85.416 | 0.009184 | 0.9846 | 0.023494 | 23.8616 | 0.032283 | 0.15688 | 0.004555 | 0.30762 | 0.30762 | 0.020756
2019-03-29 14:10 | 1553883009.5434 | 1022 | 0.20859 | 0.001045 | 0.000232 | 100.2681 | 100.022 | 119.2251 | 21.0945 | 0.002709 | 0.002572 | 221.856 | 0.26535 | 8.4021 | 0.002611 | 0.004207 | 0.004223 | 5e-05 | 0.000949 | 0.29623 | 1.3307 | 2.5474 | 100.18 | 20.83 | 37.341 | 86.8342 | 84.463 | 83.5307 | 0.009184 | 0.98981 | 0.023531 | 23.7732 | 0.021986 | 0.15786 | 0.004733 | 0.32073 | 0.32073 | 0.020032
2019-03-29 14:10 | 1553883010.5444 | 1023 | 0.20866 | 0.001039 | 0.00023 | 100.2813 | 99.9585 | 119.1552 | 21.0936 | 0.002694 | 0.002548 | 221.8459 | 0.26443 | 8.4008 | 0.002581 | 0.003477 | 0.003285 | 5.6e-05 | 0.000953 | 0.30541 | 1.33 | 2.5473 | 100.18 | 20.82 | 37.34 | 86.4695 | 83.7443 | 82.9518 | 0.009178 | 0.99142 | 0.023543 | 23.7469 | 0.028839 | 0.15731 | 0.004324 | 0.29208 | 0.29208 | 0.02192
2019-03-29 14:10 | 1553883011.5448 | 1024 | 0.20862 | 0.001043 | 0.000229 | 100.2609 | 99.9714 | 119.5143 | 21.0938 | 0.002782 | 0.002594 | 221.828 | 0.26427 | 8.4012 | 0.002568 | 0.003821 | 0.003806 | 5.4e-05 | 0.000952 | 0.29607 | 1.3487 | 2.5475 | 100.18 | 20.82 | 37.338 | 88.585 | 83.4435 | 84.0583 | 0.009178 | 0.98835 | 0.023542 | 23.8196 | 0.025863 | 0.15721 | 0.004549 | 0.30804 | 0.30804 | 0.020772
2019-03-29 14:10 | 1553883012.5435 | 1025 | 0.20864 | 0.001041 | 0.000228 | 100.2666 | 99.8837 | 119.5359 | 21.0927 | 0.002763 | 0.002597 | 221.7895 | 0.26399 | 8.4011 | 0.002519 | 0.003727 | 0.003616 | 5.4e-05 | 0.000954 | 0.29015 | 1.3803 | 2.5475 | 100.18 | 20.82 | 37.338 | 88.1155 | 82.2521 | 84.1407 | 0.009178 | 0.98812 | 0.023541 | 23.8245 | 0.034222 | 0.15705 | 0.004453 | 0.30135 | 0.30135 | 0.02121
2019-03-29 14:10 | 1553883013.5445 | 1026 | 0.20867 | 0.00104 | 0.000227 | 100.2753 | 100.0878 | 119.8436 | 21.0953 | 0.002759 | 0.002583 | 221.7795 | 0.2638 | 8.4022 | 0.002461 | 0.004274 | 0.004245 | 5.1e-05 | 0.000952 | 0.30845 | 1.3325 | 2.5474 | 100.18 | 20.82 | 37.341 | 88.032 | 80.8579 | 83.8077 | 0.009178 | 0.98904 | 0.023583 | 23.844 | 0.016738 | 0.15693 | 0.004269 | 0.28887 | 0.28887 | 0.022111
2019-03-29 14:10 | 1553883014.544 | 1027 | 0.20863 | 0.001041 | 0.000227 | 100.2882 | 99.9346 | 120.0779 | 21.0945 | 0.002715 | 0.002579 | 221.7636 | 0.26502 | 8.402 | 0.002518 | 0.003735 | 0.003599 | 5.2e-05 | 0.000948 | 0.29645 | 1.3486 | 2.5474 | 100.18 | 20.82 | 37.342 | 86.959 | 82.2336 | 83.6978 | 0.009178 | 0.98935 | 0.023609 | 23.8637 | 0.031591 | 0.15766 | 0.004499 | 0.30553 | 0.30553 | 0.021002
2019-03-29 14:10 | 1553883015.5437 | 1028 | 0.20866 | 0.001036 | 0.000226 | 100.252 | 100.0592 | 120.061 | 21.0945 | 0.002722 | 0.002622 | 221.7515 | 0.26436 | 8.3999 | 0.002538 | 0.004199 | 0.004121 | 5e-05 | 0.000943 | 0.28766 | 1.3261 | 2.5475 | 100.18 | 20.82 | 37.345 | 87.1434 | 82.7031 | 84.7283 | 0.009179 | 0.9865 | 0.023574 | 23.8964 | 0.017218 | 0.15727 | 0.004307 | 0.29155 | 0.29155 | 0.021954
2019-03-29 14:10 | 1553883016.5437 | 1029 | 0.20864 | 0.001037 | 0.000226 | 100.3115 | 99.9592 | 120.2131 | 21.0941 | 0.002761 | 0.002553 | 221.7416 | 0.26406 | 8.4023 | 0.002484 | 0.004311 | 0.004278 | 5.2e-05 | 0.000945 | 0.31006 | 1.3285 | 2.5474 | 100.18 | 20.82 | 37.347 | 88.0886 | 81.4261 | 83.0826 | 0.00918 | 0.99105 | 0.023643 | 23.8565 | 0.031474 | 0.15709 | 0.004454 | 0.30278 | 0.30278 | 0.021116
2019-03-29 14:10 | 1553883017.544 | 1030 | 0.20864 | 0.001037 | 0.000225 | 100.2963 | 99.9404 | 120.1359 | 21.0934 | 0.002743 | 0.002547 | 221.7385 | 0.26383 | 8.403 | 0.002578 | 0.003894 | 0.003797 | 5.8e-05 | 0.000945 | 0.29302 | 1.3222 | 2.5474 | 100.18 | 20.82 | 37.349 | 87.6436 | 83.6647 | 82.9172 | 0.00918 | 0.99151 | 0.023641 | 23.8433 | 0.031793 | 0.15695 | 0.004438 | 0.30169 | 0.30169 | 0.021173
2019-03-29 14:10 | 1553883018.5437 | 1031 | 0.20864 | 0.001032 | 0.000225 | 100.2567 | 99.98 | 119.9162 | 21.0934 | 0.002781 | 0.002557 | 221.7181 | 0.26365 | 8.4021 | 0.002443 | 0.00434 | 0.004261 | 5.3e-05 | 0.000944 | 0.3235 | 1.3326 | 2.5474 | 100.18 | 20.82 | 37.35 | 88.5699 | 80.4405 | 83.1741 | 0.00918 | 0.9908 | 0.023611 | 23.8301 | 0.024726 | 0.15684 | 0.004451 | 0.30222 | 0.30222 | 0.021122
2019-03-29 14:10 | 1553883019.5447 | 1032 | 0.20868 | 0.001034 | 0.000223 | 100.3082 | 99.9022 | 120.0159 | 21.0939 | 0.002733 | 0.002595 | 221.7081 | 0.26481 | 8.4001 | 0.002479 | 0.00362 | 0.003512 | 5.4e-05 | 0.000945 | 0.29878 | 1.3369 | 2.5475 | 100.18 | 20.82 | 37.353 | 87.3982 | 81.3043 | 84.0724 | 0.009181 | 0.98831 | 0.023591 | 23.87 | 0.036282 | 0.15754 | 0.004204 | 0.28463 | 0.28463 | 0.022527
2019-03-29 14:10 | 1553883020.5438 | 1033 | 0.20871 | 0.001034 | 0.000221 | 100.243 | 99.9773 | 119.784 | 21.093 | 0.002743 | 0.002553 | 221.7144 | 0.26361 | 8.4036 | 0.00256 | 0.003639 | 0.003622 | 5.4e-05 | 0.000941 | 0.3027 | 1.3204 | 2.5473 | 100.18 | 20.82 | 37.353 | 87.6548 | 83.2491 | 83.0614 | 0.009181 | 0.99111 | 0.023602 | 23.8132 | 0.023742 | 0.15682 | 0.004029 | 0.27247 | 0.27247 | 0.023425
2019-03-29 14:10 | 1553883021.544 | 1034 | 0.20866 | 0.001036 | 0.00022 | 100.2837 | 99.9902 | 119.6359 | 21.0941 | 0.002779 | 0.002579 | 221.7217 | 0.26381 | 8.403 | 0.002545 | 0.003873 | 0.003734 | 5.3e-05 | 0.000938 | 0.30702 | 1.3634 | 2.5476 | 100.18 | 20.82 | 37.354 | 88.5118 | 82.8891 | 83.6996 | 0.009181 | 0.98934 | 0.023566 | 23.8198 | 0.026213 | 0.15694 | 0.004312 | 0.29215 | 0.29215 | 0.021864
2019-03-29 14:10 | 1553883022.5449 | 1035 | 0.20869 | 0.001036 | 0.000219 | 100.316 | 99.9158 | 119.6542 | 21.0932 | 0.002749 | 0.002522 | 221.7132 | 0.26377 | 8.4014 | 0.002516 | 0.003901 | 0.003888 | 5.2e-05 | 0.00094 | 0.32261 | 1.3422 | 2.5474 | 100.18 | 20.82 | 37.354 | 87.7871 | 82.1841 | 82.3336 | 0.009181 | 0.99314 | 0.023613 | 23.7759 | 0.035754 | 0.15692 | 0.004181 | 0.28353 | 0.28353 | 0.022525
2019-03-29 14:10 | 1553883023.5439 | 1036 | 0.20869 | 0.001036 | 0.000217 | 100.2394 | 99.9875 | 119.5016 | 21.0939 | 0.002767 | 0.002593 | 221.6732 | 0.26314 | 8.4013 | 0.002505 | 0.004074 | 0.004018 | 5.5e-05 | 0.000946 | 0.28783 | 1.2849 | 2.5473 | 100.18 | 20.82 | 37.353 | 88.2259 | 81.9078 | 84.0411 | 0.009181 | 0.98839 | 0.023541 | 23.8178 | 0.022505 | 0.15654 | 0.004172 | 0.28211 | 0.28211 | 0.022585
2019-03-29 14:10 | 1553883024.5439 | 1037 | 0.2087 | 0.001034 | 0.000215 | 100.3274 | 100.1534 | 119.5101 | 21.0929 | 0.002747 | 0.002551 | 221.673 | 0.26424 | 8.4008 | 0.002542 | 0.003958 | 0.003853 | 5.2e-05 | 0.000945 | 0.27455 | 1.3439 | 2.5475 | 100.18 | 20.82 | 37.352 | 87.7415 | 82.7968 | 83.0353 | 0.009181 | 0.99118 | 0.023575 | 23.7851 | 0.015529 | 0.1572 | 0.004082 | 0.27625 | 0.27625 | 0.02316
2019-03-29 14:10 | 1553883025.5433 | 1038 | 0.20871 | 0.001031 | 0.000213 | 100.3489 | 100.0957 | 119.3958 | 21.0932 | 0.002748 | 0.00253 | 221.6887 | 0.26439 | 8.4017 | 0.002446 | 0.003579 | 0.003552 | 5.1e-05 | 0.000947 | 0.29407 | 1.3564 | 2.5475 | 100.18 | 20.82 | 37.351 | 87.7636 | 80.5119 | 82.5185 | 0.009181 | 0.99262 | 0.023581 | 23.7564 | 0.022603 | 0.15728 | 0.004072 | 0.27577 | 0.27577 | 0.023213
2019-03-29 14:10 | 1553883026.5434 | 1039 | 0.20869 | 0.001035 | 0.000211 | 100.278 | 100.0895 | 119.5688 | 21.0936 | 0.00273 | 0.002542 | 221.6733 | 0.26435 | 8.4038 | 0.002519 | 0.004674 | 0.004569 | 5.3e-05 | 0.000943 | 0.30182 | 1.3961 | 2.5475 | 100.18 | 20.82 | 37.347 | 87.3228 | 82.2467 | 82.8177 | 0.00918 | 0.99179 | 0.023588 | 23.7836 | 0.016826 | 0.15726 | 0.004149 | 0.28143 | 0.28143 | 0.022744
2019-03-29 14:10 | 1553883027.5431 | 1040 | 0.20873 | 0.001033 | 0.000209 | 100.3217 | 100.0749 | 119.8619 | 21.0942 | 0.002734 | 0.00256 | 221.6463 | 0.2637 | 8.4038 | 0.002568 | 0.004083 | 0.004099 | 5.4e-05 | 0.000946 | 0.29463 | 1.3651 | 2.5476 | 100.18 | 20.82 | 37.346 | 87.4374 | 83.4431 | 83.2332 | 0.009179 | 0.99063 | 0.023604 | 23.8267 | 0.022035 | 0.15687 | 0.003939 | 0.26682 | 0.26682 | 0.023929
2019-03-29 14:10 | 1553883028.5437 | 1041 | 0.20876 | 0.00103 | 0.000208 | 100.2906 | 100.0421 | 120.0401 | 21.0931 | 0.002707 | 0.002572 | 221.6377 | 0.26456 | 8.4018 | 0.002466 | 0.003609 | 0.003511 | 5.3e-05 | 0.000945 | 0.31222 | 1.3514 | 2.5474 | 100.18 | 20.82 | 37.345 | 86.7844 | 80.9829 | 83.5209 | 0.009179 | 0.98984 | 0.023612 | 23.854 | 0.022188 | 0.15739 | 0.003774 | 0.25527 | 0.25527 | 0.025094
2019-03-29 14:10 | 1553883029.5441 | 1042 | 0.20869 | 0.001032 | 0.000208 | 100.3504 | 100.0646 | 120.1704 | 21.0936 | 0.002741 | 0.002569 | 221.6457 | 0.2649 | 8.4033 | 0.00243 | 0.003706 | 0.003626 | 5.3e-05 | 0.00094 | 0.30862 | 1.326 | 2.5474 | 100.18 | 20.82 | 37.342 | 87.6024 | 80.1287 | 83.4534 | 0.009178 | 0.99002 | 0.023627 | 23.8647 | 0.025512 | 0.15759 | 0.004163 | 0.28307 | 0.28307 | 0.022658
2019-03-29 14:10 | 1553883030.543 | 1043 | 0.20873 | 0.00103 | 0.000207 | 100.3205 | 100.223 | 120.5128 | 21.0941 | 0.002677 | 0.002557 | 221.6249 | 0.265 | 8.3999 | 0.002437 | 0.003554 | 0.003541 | 5.3e-05 | 0.00094 | 0.31061 | 1.3453 | 2.5474 | 100.18 | 20.82 | 37.341 | 86.0602 | 80.2789 | 83.1583 | 0.009178 | 0.99084 | 0.02367 | 23.8888 | 0.008699 | 0.15765 | 0.003943 | 0.26801 | 0.26801 | 0.023941
2019-03-29 14:10 | 1553883031.5435 | 1044 | 0.20868 | 0.001029 | 0.000205 | 100.3618 | 99.9646 | 120.6004 | 21.0939 | 0.002703 | 0.00254 | 221.5937 | 0.26532 | 8.4011 | 0.002475 | 0.003976 | 0.003833 | 5.6e-05 | 0.000939 | 0.3095 | 1.317 | 2.5473 | 100.18 | 20.82 | 37.338 | 86.6853 | 81.1872 | 82.7689 | 0.009178 | 0.99193 | 0.023692 | 23.8844 | 0.035469 | 0.15784 | 0.004244 | 0.28972 | 0.28972 | 0.022173
2019-03-29 14:11 | 1553883032.5432 | 1045 | 0.20869 | 0.001026 | 0.000203 | 100.31 | 100.0653 | 120.6271 | 21.0933 | 0.002767 | 0.00252 | 221.5823 | 0.26385 | 8.4015 | 0.002477 | 0.003794 | 0.003706 | 5.2e-05 | 0.000943 | 0.28119 | 1.3469 | 2.5474 | 100.18 | 20.82 | 37.336 | 88.2176 | 81.2489 | 82.2832 | 0.009177 | 0.99328 | 0.02371 | 23.8707 | 0.021848 | 0.15697 | 0.004209 | 0.28759 | 0.28759 | 0.022214
2019-03-29 14:11 | 1553883033.5432 | 1046 | 0.2087 | 0.001031 | 0.0002 | 100.3723 | 100.0595 | 120.7307 | 21.094 | 0.002719 | 0.002547 | 221.5907 | 0.26384 | 8.4021 | 0.002515 | 0.003514 | 0.003506 | 5.5e-05 | 0.000944 | 0.3194 | 1.3516 | 2.5474 | 100.18 | 20.82 | 37.334 | 87.0633 | 82.163 | 82.922 | 0.009177 | 0.9915 | 0.023699 | 23.9024 | 0.027917 | 0.15696 | 0.004128 | 0.28191 | 0.28191 | 0.022661
2019-03-29 14:11 | 1553883034.5437 | 1047 | 0.20869 | 0.001031 | 0.000197 | 100.3154 | 100.1998 | 120.7675 | 21.0941 | 0.002707 | 0.002554 | 221.5829 | 0.26471 | 8.4023 | 0.00243 | 0.003789 | 0.003687 | 5.4e-05 | 0.000946 | 0.34523 | 1.3513 | 2.5473 | 100.18 | 20.82 | 37.331 | 86.7885 | 80.1121 | 83.1098 | 0.009176 | 0.99098 | 0.023697 | 23.9124 | 0.010317 | 0.15747 | 0.004224 | 0.28883 | 0.28883 | 0.02219
2019-03-29 14:11 | 1553883035.5443 | 1048 | 0.20864 | 0.001034 | 0.000194 | 100.4651 | 100.0588 | 120.6856 | 21.0943 | 0.002707 | 0.002528 | 221.5647 | 0.26388 | 8.4042 | 0.002521 | 0.00385 | 0.003728 | 5.2e-05 | 0.000942 | 0.28655 | 1.3491 | 2.5474 | 100.18 | 20.82 | 37.33 | 86.7849 | 82.3113 | 82.4666 | 0.009176 | 0.99277 | 0.02371 | 23.8827 | 0.036243 | 0.15698 | 0.004479 | 0.30737 | 0.30737 | 0.020787
2019-03-29 14:11 | 1553883036.5436 | 1049 | 0.20867 | 0.001035 | 0.000191 | 100.2795 | 100.0353 | 120.2265 | 21.0934 | 0.002746 | 0.002555 | 221.556 | 0.26452 | 8.4009 | 0.002484 | 0.003795 | 0.003719 | 5.3e-05 | 0.000944 | 0.29518 | 1.3539 | 2.5474 | 100.18 | 20.82 | 37.327 | 87.7233 | 81.4182 | 83.1264 | 0.009175 | 0.99093 | 0.023643 | 23.8593 | 0.02181 | 0.15736 | 0.004313 | 0.29487 | 0.29487 | 0.02172
2019-03-29 14:11 | 1553883037.5434 | 1050 | 0.20869 | 0.00103 | 0.000189 | 100.3421 | 100.1977 | 120.0319 | 21.0933 | 0.002735 | 0.00254 | 221.5445 | 0.26417 | 8.4027 | 0.002492 | 0.003465 | 0.003428 | 5.2e-05 | 0.00094 | 0.27721 | 1.3701 | 2.5475 | 100.18 | 20.82 | 37.324 | 87.4546 | 81.606 | 82.7534 | 0.009174 | 0.99197 | 0.023636 | 23.8275 | 0.012875 | 0.15716 | 0.004183 | 0.28571 | 0.28571 | 0.022387
2019-03-29 14:11 | 1553883038.5437 | 1051 | 0.20869 | 0.001033 | 0.000187 | 100.3902 | 100.3141 | 119.903 | 21.0938 | 0.002687 | 0.002526 | 221.5339 | 0.26411 | 8.4044 | 0.002498 | 0.004236 | 0.004122 | 5.2e-05 | 0.000941 | 0.32836 | 1.3734 | 2.5476 | 100.18 | 20.82 | 37.324 | 86.2972 | 81.7517 | 82.4342 | 0.009174 | 0.99286 | 0.023634 | 23.804 | 0.006787 | 0.15712 | 0.004186 | 0.28603 | 0.28603 | 0.022357
2019-03-29 14:11 | 1553883039.5431 | 1052 | 0.20868 | 0.001031 | 0.000184 | 100.4196 | 100.2274 | 119.9298 | 21.0933 | 0.002748 | 0.002563 | 221.5267 | 0.26446 | 8.4023 | 0.002445 | 0.00403 | 0.003914 | 5.3e-05 | 0.000939 | 0.31526 | 1.3645 | 2.5474 | 100.18 | 20.82 | 37.321 | 87.7677 | 80.4754 | 83.3174 | 0.009173 | 0.9904 | 0.023607 | 23.8362 | 0.017133 | 0.15733 | 0.004259 | 0.29102 | 0.29102 | 0.022003
2019-03-29 14:11 | 1553883040.5428 | 1053 | 0.20867 | 0.001033 | 0.000183 | 100.3798 | 99.894 | 119.7953 | 21.094 | 0.002762 | 0.002533 | 221.5069 | 0.26408 | 8.4015 | 0.002555 | 0.004176 | 0.004134 | 5.2e-05 | 0.000943 | 0.30574 | 1.3587 | 2.5475 | 100.18 | 20.82 | 37.32 | 88.1073 | 83.1262 | 82.5997 | 0.009173 | 0.9924 | 0.023618 | 23.7989 | 0.043392 | 0.1571 | 0.004307 | 0.29459 | 0.29459 | 0.021705
2019-03-29 14:11 | 1553883041.5435 | 1054 | 0.2087 | 0.001036 | 0.000181 | 100.4244 | 100.3073 | 119.6387 | 21.094 | 0.002723 | 0.002614 | 221.4918 | 0.26392 | 8.4007 | 0.002474 | 0.004273 | 0.004174 | 5.3e-05 | 0.000946 | 0.28838 | 1.3312 | 2.5475 | 100.18 | 20.82 | 37.321 | 87.1694 | 81.1749 | 84.5337 | 0.009173 | 0.98703 | 0.023539 | 23.8479 | 0.010435 | 0.15701 | 0.004147 | 0.28246 | 0.28246 | 0.022624
2019-03-29 14:11 | 1553883042.5425 | 1055 | 0.20869 | 0.001037 | 0.00018 | 100.4196 | 100.1254 | 119.5192 | 21.0938 | 0.002702 | 0.002517 | 221.4777 | 0.26439 | 8.4016 | 0.002522 | 0.004056 | 0.004036 | 5.1e-05 | 0.000947 | 0.2891 | 1.3575 | 2.5474 | 100.18 | 20.82 | 37.323 | 86.6686 | 82.325 | 82.1969 | 0.009174 | 0.99352 | 0.023604 | 23.758 | 0.026244 | 0.15729 | 0.004212 | 0.28793 | 0.28793 | 0.022233
2019-03-29 14:11 | 1553883043.5428 | 1056 | 0.20874 | 0.001035 | 0.000178 | 100.4208 | 100.2435 | 119.1918 | 21.0935 | 0.002674 | 0.002551 | 221.4632 | 0.26507 | 8.4003 | 0.002528 | 0.003913 | 0.00382 | 5.1e-05 | 0.000948 | 0.29645 | 1.3274 | 2.5474 | 100.18 | 20.82 | 37.324 | 85.9918 | 82.4822 | 83.0337 | 0.009174 | 0.99119 | 0.023544 | 23.7533 | 0.015808 | 0.15769 | 0.003895 | 0.26488 | 0.26488 | 0.02423
2019-03-29 14:11 | 1553883044.5425 | 1057 | 0.20866 | 0.001039 | 0.000177 | 100.428 | 100.1189 | 118.8897 | 21.0935 | 0.002715 | 0.002563 | 221.453 | 0.26433 | 8.4032 | 0.002504 | 0.004353 | 0.004277 | 5e-05 | 0.00095 | 0.3063 | 1.3272 | 2.5473 | 100.18 | 20.82 | 37.325 | 86.9675 | 81.9033 | 83.3215 | 0.009174 | 0.99039 | 0.023505 | 23.7328 | 0.02757 | 0.15725 | 0.004404 | 0.30045 | 0.30045 | 0.021301
2019-03-29 14:11 | 1553883045.5428 | 1058 | 0.2087 | 0.001037 | 0.000173 | 100.3977 | 100.1383 | 119.3078 | 21.0941 | 0.00268 | 0.002546 | 221.4333 | 0.2646 | 8.4014 | 0.002494 | 0.004417 | 0.004348 | 5.2e-05 | 0.000955 | 0.32333 | 1.3004 | 2.5473 | 100.18 | 20.82 | 37.327 | 86.1298 | 81.6624 | 82.8939 | 0.009175 | 0.99158 | 0.02356 | 23.7602 | 0.023139 | 0.15741 | 0.004138 | 0.28253 | 0.28253 | 0.022676
2019-03-29 14:11 | 1553883046.5426 | 1059 | 0.20871 | 0.001039 | 0.000171 | 100.3507 | 100.137 | 119.676 | 21.0938 | 0.002706 | 0.002518 | 221.4262 | 0.26463 | 8.4029 | 0.002489 | 0.003185 | 0.003099 | 4.9e-05 | 0.000954 | 0.30287 | 1.3388 | 2.5475 | 100.18 | 20.82 | 37.329 | 86.7483 | 81.5342 | 82.2313 | 0.009175 | 0.99343 | 0.023618 | 23.7747 | 0.019073 | 0.15743 | 0.004086 | 0.2797 | 0.2797 | 0.022908
2019-03-29 14:11 | 1553883047.5427 | 1060 | 0.20872 | 0.001039 | 0.000169 | 100.4369 | 100.2564 | 119.7554 | 21.0948 | 0.002696 | 0.002555 | 221.4186 | 0.26483 | 8.4026 | 0.002497 | 0.004111 | 0.004068 | 4.9e-05 | 0.000952 | 0.32278 | 1.3936 | 2.5476 | 100.18 | 20.82 | 37.334 | 86.5026 | 81.7301 | 83.1184 | 0.009177 | 0.99095 | 0.023597 | 23.8122 | 0.01609 | 0.15755 | 0.004065 | 0.27807 | 0.27807 | 0.02306
2019-03-29 14:11 | 1553883048.5427 | 1061 | 0.20872 | 0.00104 | 0.000168 | 100.4339 | 100.2441 | 120.222 | 21.0939 | 0.002695 | 0.002557 | 221.408 | 0.26331 | 8.4039 | 0.002543 | 0.003155 | 0.003056 | 5e-05 | 0.00095 | 0.30254 | 1.3548 | 2.5474 | 100.18 | 20.82 | 37.336 | 86.4892 | 82.8242 | 83.1666 | 0.009177 | 0.99082 | 0.023641 | 23.8602 | 0.01692 | 0.15665 | 0.004072 | 0.27913 | 0.27913 | 0.02284
2019-03-29 14:11 | 1553883049.5438 | 1062 | 0.20869 | 0.00104 | 0.000167 | 100.4124 | 100.1663 | 120.4741 | 21.094 | 0.002768 | 0.002496 | 221.3923 | 0.26485 | 8.4056 | 0.002542 | 0.004282 | 0.004197 | 5e-05 | 0.000945 | 0.29142 | 1.3518 | 2.5474 | 100.18 | 20.81 | 37.338 | 88.2549 | 82.8104 | 81.6981 | 0.009172 | 0.99492 | 0.023715 | 23.8359 | 0.021947 | 0.15756 | 0.004223 | 0.29077 | 0.29077 | 0.022054
2019-03-29 14:11 | 1553883050.5432 | 1063 | 0.20869 | 0.001037 | 0.000165 | 100.4516 | 100.1028 | 120.6288 | 21.0936 | 0.002686 | 0.00255 | 221.3914 | 0.26394 | 8.4027 | 0.002553 | 0.00394 | 0.003868 | 4.7e-05 | 0.00095 | 0.29607 | 1.342 | 2.5473 | 100.18 | 20.82 | 37.345 | 86.2751 | 83.0773 | 83.0103 | 0.009179 | 0.99125 | 0.023686 | 23.8953 | 0.031109 | 0.15702 | 0.004216 | 0.29006 | 0.29006 | 0.022032
2019-03-29 14:11 | 1553883051.5432 | 1064 | 0.20867 | 0.001041 | 0.000163 | 100.475 | 100.3076 | 120.7241 | 21.0928 | 0.002733 | 0.002534 | 221.3932 | 0.26343 | 8.4016 | 0.0025 | 0.004343 | 0.004145 | 5.1e-05 | 0.000957 | 0.32244 | 1.3878 | 2.5475 | 100.18 | 20.82 | 37.348 | 87.3978 | 81.7883 | 82.615 | 0.00918 | 0.99236 | 0.023709 | 23.8914 | 0.014909 | 0.15671 | 0.004337 | 0.29905 | 0.29905 | 0.021329
2019-03-29 14:11 | 1553883052.5429 | 1065 | 0.20869 | 0.001047 | 0.00016 | 100.4313 | 100.2322 | 120.6919 | 21.0929 | 0.00267 | 0.002464 | 221.382 | 0.26439 | 8.4025 | 0.002565 | 0.003564 | 0.003543 | 5.2e-05 | 0.000957 | 0.29319 | 1.3727 | 2.5475 | 100.18 | 20.81 | 37.351 | 85.8818 | 83.3505 | 80.9288 | 0.009175 | 0.99708 | 0.023762 | 23.8316 | 0.017746 | 0.15729 | 0.004215 | 0.29118 | 0.29118 | 0.021985
2019-03-29 14:11 | 1553883053.5432 | 1066 | 0.20868 | 0.001045 | 0.000157 | 100.4941 | 100.2728 | 120.9098 | 21.0927 | 0.002679 | 0.002512 | 221.3651 | 0.26439 | 8.4013 | 0.002425 | 0.003621 | 0.003564 | 5e-05 | 0.000957 | 0.28241 | 1.3969 | 2.5475 | 100.18 | 20.81 | 37.353 | 86.1065 | 80 | 82.0898 | 0.009176 | 0.99382 | 0.023745 | 23.8922 | 0.019719 | 0.15728 | 0.004284 | 0.29611 | 0.29611 | 0.021618
2019-03-29 14:11 | 1553883054.5433 | 1067 | 0.2087 | 0.001041 | 0.000154 | 100.4941 | 100.2551 | 120.937 | 21.0938 | 0.002739 | 0.002525 | 221.3538 | 0.26394 | 8.4008 | 0.002525 | 0.00379 | 0.003693 | 5.1e-05 | 0.000954 | 0.32316 | 1.3723 | 2.5475 | 100.18 | 20.81 | 37.358 | 87.5448 | 82.4005 | 82.3998 | 0.009177 | 0.99296 | 0.023737 | 23.9053 | 0.021302 | 0.15702 | 0.004178 | 0.28861 | 0.28861 | 0.022142
2019-03-29 14:11 | 1553883055.542 | 1068 | 0.2087 | 0.001037 | 0.000153 | 100.5073 | 100.2735 | 120.5614 | 21.0932 | 0.002737 | 0.002526 | 221.3513 | 0.26405 | 8.4024 | 0.002495 | 0.00414 | 0.004117 | 4.9e-05 | 0.000953 | 0.30414 | 1.3316 | 2.5474 | 100.18 | 20.82 | 37.365 | 87.5043 | 81.6692 | 82.4256 | 0.009184 | 0.99288 | 0.023699 | 23.869 | 0.020831 | 0.15709 | 0.004206 | 0.29023 | 0.29023 | 0.022029
2019-03-29 14:11 | 1553883056.5422 | 1069 | 0.20869 | 0.001041 | 0.000152 | 100.4307 | 100.2527 | 120.2929 | 21.094 | 0.002684 | 0.002533 | 221.3376 | 0.26443 | 8.4009 | 0.002421 | 0.00372 | 0.003632 | 4.7e-05 | 0.000958 | 0.31725 | 1.3285 | 2.5473 | 100.18 | 20.82 | 37.365 | 86.2185 | 79.91 | 82.6029 | 0.009184 | 0.99239 | 0.023667 | 23.8483 | 0.015866 | 0.15731 | 0.004216 | 0.29061 | 0.29061 | 0.022032
2019-03-29 14:11 | 1553883057.5434 | 1070 | 0.20869 | 0.001045 | 0.000151 | 100.4453 | 100.1762 | 120.0809 | 21.0936 | 0.002715 | 0.002481 | 221.3463 | 0.26461 | 8.3981 | 0.002403 | 0.00437 | 0.004233 | 5.2e-05 | 0.000952 | 0.31509 | 1.323 | 2.5474 | 100.18 | 20.82 | 37.372 | 86.9589 | 79.477 | 81.3499 | 0.009186 | 0.9959 | 0.023688 | 23.7853 | 0.023995 | 0.15742 | 0.004255 | 0.29368 | 0.29368 | 0.021816
2019-03-29 14:11 | 1553883058.5426 | 1071 | 0.20865 | 0.001038 | 0.000151 | 100.4381 | 100.1646 | 119.6542 | 21.0929 | 0.002753 | 0.002528 | 221.3541 | 0.26526 | 8.4039 | 0.00245 | 0.00372 | 0.003728 | 4.8e-05 | 0.000947 | 0.27184 | 1.3092 | 2.5474 | 100.18 | 20.82 | 37.376 | 87.8812 | 80.5967 | 82.4653 | 0.009187 | 0.99277 | 0.023608 | 23.7803 | 0.024392 | 0.15781 | 0.004507 | 0.31057 | 0.31057 | 0.02068
2019-03-29 14:11 | 1553883059.5421 | 1072 | 0.20867 | 0.001046 | 0.000149 | 100.4627 | 100.0755 | 119.2737 | 21.0939 | 0.002639 | 0.002523 | 221.3386 | 0.26351 | 8.403 | 0.002537 | 0.003736 | 0.003683 | 5.1e-05 | 0.000948 | 0.30038 | 1.3588 | 2.5474 | 100.18 | 20.82 | 37.381 | 85.1306 | 82.6915 | 82.3477 | 0.009188 | 0.9931 | 0.023575 | 23.7386 | 0.034535 | 0.15676 | 0.004367 | 0.30032 | 0.30032 | 0.021245
2019-03-29 14:11 | 1553883060.5434 | 1073 | 0.20869 | 0.00104 | 0.000148 | 100.4339 | 100.2138 | 118.9916 | 21.0931 | 0.002676 | 0.002511 | 221.319 | 0.26495 | 8.4024 | 0.002472 | 0.00401 | 0.003941 | 4.8e-05 | 0.00095 | 0.29142 | 1.3376 | 2.5474 | 100.18 | 20.82 | 37.392 | 86.0419 | 81.1177 | 82.0582 | 0.009191 | 0.99391 | 0.023557 | 23.7008 | 0.019631 | 0.15762 | 0.004258 | 0.29244 | 0.29244 | 0.021937
2019-03-29 14:11 | 1553883061.5432 | 1074 | 0.20866 | 0.001042 | 0.000146 | 100.4797 | 100.252 | 118.9153 | 21.0935 | 0.002638 | 0.002487 | 221.2971 | 0.26565 | 8.3994 | 0.002478 | 0.004055 | 0.003993 | 5.1e-05 | 0.000955 | 0.2792 | 1.319 | 2.5473 | 100.18 | 20.82 | 37.398 | 85.1191 | 81.2702 | 81.4769 | 0.009192 | 0.99554 | 0.023568 | 23.6738 | 0.020298 | 0.15803 | 0.004443 | 0.30576 | 0.30576 | 0.021036
2019-03-29 14:11 | 1553883062.5438 | 1075 | 0.20871 | 0.001044 | 0.000145 | 100.4708 | 100.1025 | 118.9477 | 21.0936 | 0.002696 | 0.002522 | 221.2857 | 0.2644 | 8.4027 | 0.002444 | 0.00389 | 0.00381 | 4.6e-05 | 0.000954 | 0.29983 | 1.3041 | 2.5473 | 100.18 | 20.82 | 37.404 | 86.521 | 80.4494 | 82.3263 | 0.009194 | 0.99316 | 0.023543 | 23.7054 | 0.032845 | 0.15729 | 0.004139 | 0.28406 | 0.28406 | 0.022536
2019-03-29 14:11 | 1553883063.5431 | 1076 | 0.20869 | 0.001049 | 0.000144 | 100.5297 | 100.2516 | 118.9369 | 21.0947 | 0.002689 | 0.002496 | 221.2836 | 0.26476 | 8.4039 | 0.002536 | 0.004363 | 0.004292 | 4.9e-05 | 0.000955 | 0.30735 | 1.3361 | 2.5475 | 100.18 | 20.82 | 37.41 | 86.3419 | 82.6652 | 81.7096 | 0.009195 | 0.99489 | 0.023563 | 23.6837 | 0.024777 | 0.15751 | 0.004231 | 0.29086 | 0.29086 | 0.02204
2019-03-29 14:11 | 1553883064.5432 | 1077 | 0.2087 | 0.001049 | 0.000143 | 100.5016 | 100.2499 | 118.8559 | 21.0948 | 0.002686 | 0.002511 | 221.2786 | 0.26433 | 8.404 | 0.002492 | 0.004223 | 0.00418 | 4.6e-05 | 0.000953 | 0.30414 | 1.321 | 2.5473 | 100.18 | 20.82 | 37.429 | 86.2748 | 81.62 | 82.0748 | 0.0092 | 0.99386 | 0.023543 | 23.6879 | 0.022425 | 0.15725 | 0.004213 | 0.28938 | 0.28938 | 0.022117
2019-03-29 14:11 | 1553883065.5433 | 1078 | 0.20868 | 0.001053 | 0.000142 | 100.4968 | 100.3267 | 118.7433 | 21.0946 | 0.002673 | 0.002475 | 221.2575 | 0.26497 | 8.4017 | 0.002541 | 0.003471 | 0.003386 | 4.9e-05 | 0.000946 | 0.29031 | 1.3392 | 2.5474 | 100.18 | 20.82 | 37.429 | 85.9616 | 82.7898 | 81.1966 | 0.0092 | 0.99633 | 0.023561 | 23.6473 | 0.01515 | 0.15763 | 0.004301 | 0.29588 | 0.29588 | 0.021682
2019-03-29 14:11 | 1553883066.5422 | 1079 | 0.2087 | 0.001046 | 0.000142 | 100.5043 | 100.2929 | 118.8407 | 21.0938 | 0.0027 | 0.00248 | 221.2512 | 0.26429 | 8.4023 | 0.002522 | 0.003999 | 0.003932 | 4.8e-05 | 0.000943 | 0.28335 | 1.3396 | 2.5473 | 100.18 | 20.82 | 37.446 | 86.6119 | 82.3265 | 81.3205 | 0.009204 | 0.99598 | 0.023566 | 23.6612 | 0.018829 | 0.15723 | 0.004192 | 0.28827 | 0.28827 | 0.022198
2019-03-29 14:11 | 1553883067.5435 | 1080 | 0.20867 | 0.001049 | 0.000142 | 100.527 | 100.3728 | 118.9892 | 21.0939 | 0.002746 | 0.002486 | 221.2413 | 0.26452 | 8.4034 | 0.002493 | 0.003923 | 0.003821 | 4.9e-05 | 0.000947 | 0.30989 | 1.3336 | 2.5475 | 100.18 | 20.82 | 37.446 | 87.712 | 81.6355 | 81.4649 | 0.009204 | 0.99557 | 0.023576 | 23.6808 | 0.013734 | 0.15737 | 0.004395 | 0.3028 | 0.3028 | 0.021152
2019-03-29 14:11 | 1553883068.5424 | 1081 | 0.20873 | 0.001047 | 0.000142 | 100.5384 | 100.2271 | 118.9869 | 21.0941 | 0.002694 | 0.002495 | 221.2012 | 0.26424 | 8.4038 | 0.002552 | 0.003922 | 0.003826 | 4.7e-05 | 0.000944 | 0.2975 | 1.337 | 2.5473 | 100.18 | 20.82 | 37.452 | 86.4763 | 83.0382 | 81.6749 | 0.009206 | 0.99499 | 0.023569 | 23.6876 | 0.027742 | 0.15719 | 0.004029 | 0.27672 | 0.27672 | 0.02312
2019-03-29 14:11 | 1553883069.5423 | 1082 | 0.20875 | 0.00105 | 0.000142 | 100.522 | 100.4615 | 119.2244 | 21.0937 | 0.002707 | 0.002482 | 221.1767 | 0.26383 | 8.403 | 0.002516 | 0.004593 | 0.00451 | 5.1e-05 | 0.00094 | 0.28567 | 1.3347 | 2.5473 | 100.18 | 20.82 | 37.461 | 86.7888 | 82.1804 | 81.3742 | 0.009208 | 0.99583 | 0.023602 | 23.7012 | 0.005377 | 0.15696 | 0.003898 | 0.26793 | 0.26793 | 0.023842
2019-03-29 14:11 | 1553883070.5434 | 1083 | 0.20873 | 0.001048 | 0.000142 | 100.4546 | 100.2858 | 119.48 | 21.0932 | 0.002754 | 0.002505 | 221.1857 | 0.26374 | 8.4025 | 0.002444 | 0.003908 | 0.003826 | 5e-05 | 0.000946 | 0.26344 | 1.3322 | 2.5475 | 100.18 | 20.82 | 37.466 | 87.9114 | 80.4449 | 81.9072 | 0.009209 | 0.99433 | 0.02361 | 23.7444 | 0.015045 | 0.1569 | 0.004016 | 0.27637 | 0.27637 | 0.023106
2019-03-29 14:11 | 1553883071.5415 | 1084 | 0.20872 | 0.001047 | 0.00014 | 100.6076 | 100.4342 | 119.7742 | 21.0938 | 0.002695 | 0.00252 | 221.1853 | 0.26426 | 8.4027 | 0.002406 | 0.003759 | 0.003765 | 4.9e-05 | 0.000954 | 0.29911 | 1.3485 | 2.5474 | 100.18 | 20.82 | 37.469 | 86.4791 | 79.5373 | 82.2837 | 0.00921 | 0.99328 | 0.023626 | 23.7862 | 0.015423 | 0.15721 | 0.004099 | 0.28252 | 0.28252 | 0.022647
2019-03-29 14:11 | 1553883072.5425 | 1085 | 0.20873 | 0.001041 | 0.000139 | 100.6213 | 100.4598 | 119.8601 | 21.094 | 0.002652 | 0.002533 | 221.169 | 0.2635 | 8.4038 | 0.002493 | 0.004279 | 0.004136 | 5e-05 | 0.000955 | 0.28512 | 1.3755 | 2.5476 | 100.18 | 20.82 | 37.476 | 85.4575 | 81.6397 | 82.6005 | 0.009211 | 0.9924 | 0.023624 | 23.8053 | 0.014368 | 0.15676 | 0.004048 | 0.27888 | 0.27888 | 0.022878
2019-03-29 14:11 | 1553883073.5415 | 1086 | 0.20872 | 0.001042 | 0.000138 | 100.533 | 100.4394 | 119.891 | 21.0942 | 0.002666 | 0.002473 | 221.1655 | 0.26384 | 8.4013 | 0.002496 | 0.004054 | 0.003928 | 5.1e-05 | 0.000955 | 0.3011 | 1.4164 | 2.5475 | 100.18 | 20.81 | 37.478 | 85.7981 | 81.6966 | 81.1476 | 0.009206 | 0.99647 | 0.023676 | 23.7597 | 0.008338 | 0.15696 | 0.004116 | 0.28442 | 0.28442 | 0.02246
2019-03-29 14:11 | 1553883074.5415 | 1087 | 0.20872 | 0.001042 | 0.000137 | 100.5456 | 100.3465 | 119.9317 | 21.0945 | 0.002629 | 0.002556 | 221.1447 | 0.26383 | 8.4044 | 0.00252 | 0.004004 | 0.004006 | 4.8e-05 | 0.000954 | 0.30486 | 1.4463 | 2.5476 | 100.18 | 20.81 | 37.482 | 84.9062 | 82.2695 | 83.1352 | 0.009207 | 0.99091 | 0.023614 | 23.8303 | 0.017727 | 0.15696 | 0.00406 | 0.27978 | 0.27978 | 0.022832
2019-03-29 14:11 | 1553883075.542 | 1088 | 0.20873 | 0.001045 | 0.000135 | 100.5291 | 100.325 | 120.015 | 21.0947 | 0.002701 | 0.002512 | 221.1409 | 0.26371 | 8.402 | 0.002491 | 0.004104 | 0.00401 | 4.8e-05 | 0.000952 | 0.30055 | 1.4194 | 2.5476 | 100.18 | 20.82 | 37.484 | 86.6299 | 81.5865 | 82.0911 | 0.009213 | 0.99382 | 0.023657 | 23.8036 | 0.018179 | 0.15688 | 0.004018 | 0.27743 | 0.27743 | 0.023015
2019-03-29 14:11 | 1553883076.5422 | 1089 | 0.20875 | 0.001046 | 0.000135 | 100.5504 | 100.4049 | 119.7817 | 21.0942 | 0.002702 | 0.00248 | 221.1356 | 0.264 | 8.4024 | 0.00258 | 0.004015 | 0.003966 | 4.9e-05 | 0.000948 | 0.29048 | 1.4007 | 2.5476 | 100.18 | 20.81 | 37.488 | 86.6482 | 83.7202 | 81.3099 | 0.009209 | 0.99601 | 0.02366 | 23.7543 | 0.012953 | 0.15705 | 0.003895 | 0.26876 | 0.26876 | 0.023784
2019-03-29 14:11 | 1553883077.5426 | 1090 | 0.20877 | 0.001044 | 0.000133 | 100.5755 | 100.4578 | 119.5986 | 21.0942 | 0.002683 | 0.00252 | 221.1295 | 0.26472 | 8.4028 | 0.002519 | 0.004062 | 0.003995 | 5e-05 | 0.000951 | 0.29407 | 1.3568 | 2.5476 | 100.18 | 20.81 | 37.489 | 86.2124 | 82.2511 | 82.291 | 0.009209 | 0.99326 | 0.023609 | 23.769 | 0.010478 | 0.15748 | 0.003808 | 0.26211 | 0.26211 | 0.024453
2019-03-29 14:11 | 1553883078.5416 | 1091 | 0.20873 | 0.001048 | 0.000133 | 100.5893 | 100.3383 | 119.8342 | 21.0943 | 0.00268 | 0.002551 | 221.1369 | 0.26482 | 8.402 | 0.002475 | 0.003704 | 0.003628 | 5e-05 | 0.00095 | 0.29662 | 1.327 | 2.5474 | 100.18 | 20.81 | 37.489 | 86.137 | 81.204 | 83.0367 | 0.009209 | 0.99118 | 0.023607 | 23.8173 | 0.022344 | 0.15754 | 0.004024 | 0.27736 | 0.27736 | 0.023118
2019-03-29 14:11 | 1553883079.5431 | 1092 | 0.20873 | 0.001043 | 0.000133 | 100.607 | 100.3452 | 119.9561 | 21.0952 | 0.002627 | 0.002487 | 221.1533 | 0.26562 | 8.4014 | 0.002506 | 0.004232 | 0.004207 | 4.9e-05 | 0.00095 | 0.26737 | 1.3412 | 2.5473 | 100.18 | 20.81 | 37.491 | 84.8562 | 81.9557 | 81.4793 | 0.009209 | 0.99553 | 0.023671 | 23.7773 | 0.023306 | 0.15801 | 0.00401 | 0.27713 | 0.27713 | 0.023207
2019-03-29 14:11 | 1553883080.5431 | 1093 | 0.20877 | 0.001042 | 0.000133 | 100.607 | 100.4554 | 120.0415 | 21.0952 | 0.002645 | 0.002476 | 221.1395 | 0.26392 | 8.4011 | 0.002496 | 0.004138 | 0.004074 | 5.2e-05 | 0.00095 | 0.27704 | 1.3298 | 2.5473 | 100.18 | 20.81 | 37.492 | 85.2839 | 81.693 | 81.2306 | 0.00921 | 0.99623 | 0.023688 | 23.7774 | 0.013485 | 0.15701 | 0.003768 | 0.26013 | 0.26013 | 0.024565
2019-03-29 14:11 | 1553883081.5417 | 1094 | 0.20873 | 0.001041 | 0.000134 | 100.61 | 100.3271 | 119.8371 | 21.0949 | 0.002694 | 0.002519 | 221.1146 | 0.26418 | 8.4008 | 0.002458 | 0.003663 | 0.00359 | 4.8e-05 | 0.000942 | 0.26825 | 1.332 | 2.5473 | 100.18 | 20.81 | 37.492 | 86.4614 | 80.789 | 82.2691 | 0.00921 | 0.99332 | 0.023633 | 23.7919 | 0.025185 | 0.15716 | 0.004057 | 0.27999 | 0.27999 | 0.022845
2019-03-29 14:11 | 1553883082.5421 | 1095 | 0.2087 | 0.001043 | 0.000134 | 100.6246 | 100.3588 | 119.7591 | 21.0943 | 0.0027 | 0.002485 | 221.1103 | 0.26419 | 8.4024 | 0.002491 | 0.003697 | 0.003641 | 4.7e-05 | 0.000939 | 0.29983 | 1.3213 | 2.5475 | 100.18 | 20.81 | 37.492 | 86.5989 | 81.5907 | 81.4433 | 0.00921 | 0.99564 | 0.023653 | 23.7566 | 0.023659 | 0.15717 | 0.004231 | 0.29253 | 0.29253 | 0.021867
2019-03-29 14:11 | 1553883083.5422 | 1096 | 0.20868 | 0.001043 | 0.000132 | 100.6198 | 100.4288 | 119.9528 | 21.0936 | 0.002684 | 0.002565 | 221.0795 | 0.26469 | 8.4021 | 0.002516 | 0.004978 | 0.004751 | 5e-05 | 0.000935 | 0.28263 | 1.2994 | 2.5474 | 100.18 | 20.81 | 37.491 | 86.2229 | 82.1929 | 83.3692 | 0.009209 | 0.99026 | 0.023608 | 23.8402 | 0.017 | 0.15746 | 0.004346 | 0.30027 | 0.30027 | 0.021344
2019-03-29 14:11 | 1553883084.5429 | 1097 | 0.20869 | 0.001048 | 0.000131 | 100.5944 | 100.4134 | 120.0873 | 21.096 | 0.002667 | 0.002458 | 221.089 | 0.26459 | 8.4032 | 0.002427 | 0.003625 | 0.003599 | 4.8e-05 | 0.000935 | 0.2912 | 1.3308 | 2.5473 | 100.18 | 20.82 | 37.491 | 85.8278 | 80.0554 | 80.7922 | 0.009215 | 0.99747 | 0.023707 | 23.7672 | 0.016106 | 0.1574 | 0.004245 | 0.29436 | 0.29436 | 0.021764
2019-03-29 14:11 | 1553883085.5432 | 1098 | 0.20869 | 0.001044 | 0.00013 | 100.5794 | 100.4264 | 120.2035 | 21.0936 | 0.002633 | 0.002517 | 221.0746 | 0.26493 | 8.4035 | 0.002446 | 0.004031 | 0.004011 | 5e-05 | 0.000935 | 0.27704 | 1.362 | 2.5476 | 100.18 | 20.81 | 37.49 | 84.9867 | 80.5092 | 82.2171 | 0.009209 | 0.99347 | 0.023671 | 23.8265 | 0.01362 | 0.1576 | 0.00428 | 0.29654 | 0.29654 | 0.021631
2019-03-29 14:11 | 1553883086.542 | 1099 | 0.20873 | 0.001047 | 0.000129 | 100.5716 | 100.3871 | 120.614 | 21.0942 | 0.00269 | 0.00251 | 221.0493 | 0.2651 | 8.4008 | 0.002512 | 0.003623 | 0.003528 | 5.2e-05 | 0.000932 | 0.28296 | 1.3642 | 2.5474 | 100.18 | 20.81 | 37.489 | 86.357 | 82.0944 | 82.05 | 0.009209 | 0.99393 | 0.023717 | 23.8616 | 0.016424 | 0.15771 | 0.004055 | 0.28112 | 0.28112 | 0.022833
2019-03-29 14:11 | 1553883087.542 | 1100 | 0.20874 | 0.001053 | 0.000128 | 100.6665 | 100.3619 | 120.89 | 21.0953 | 0.00259 | 0.00251 | 221.022 | 0.26495 | 8.4018 | 0.002544 | 0.003877 | 0.003812 | 5e-05 | 0.000928 | 0.26521 | 1.3785 | 2.5475 | 100.18 | 20.82 | 37.487 | 83.9597 | 82.8498 | 82.0308 | 0.009214 | 0.99399 | 0.023745 | 23.8882 | 0.02711 | 0.15762 | 0.003995 | 0.27725 | 0.27725 | 0.023138
2019-03-29 14:11 | 1553883088.541 | 1101 | 0.20872 | 0.001053 | 0.000128 | 100.6402 | 100.3807 | 120.9447 | 21.094 | 0.002692 | 0.00247 | 221.0045 | 0.26411 | 8.4023 | 0.002527 | 0.004312 | 0.004168 | 5.2e-05 | 0.000929 | 0.29679 | 1.366 | 2.5475 | 100.18 | 20.82 | 37.486 | 86.4261 | 82.4553 | 81.0688 | 0.009214 | 0.99669 | 0.023782 | 23.8612 | 0.023096 | 0.15712 | 0.004067 | 0.28278 | 0.28278 | 0.022614
2019-03-29 14:11 | 1553883089.5424 | 1102 | 0.20872 | 0.001045 | 0.000128 | 100.668 | 100.4373 | 120.8065 | 21.0944 | 0.002613 | 0.002507 | 220.9876 | 0.26466 | 8.401 | 0.002514 | 0.004117 | 0.004035 | 5.2e-05 | 0.000935 | 0.30127 | 1.3217 | 2.5475 | 100.18 | 20.82 | 37.483 | 84.516 | 82.1258 | 81.9667 | 0.009213 | 0.99417 | 0.023739 | 23.8778 | 0.020523 | 0.15745 | 0.004126 | 0.28647 | 0.28647 | 0.022369
2019-03-29 14:11 | 1553883090.5411 | 1103 | 0.2087 | 0.001045 | 0.000129 | 100.6713 | 100.5888 | 120.7358 | 21.0943 | 0.002698 | 0.002503 | 220.9813 | 0.26595 | 8.402 | 0.002555 | 0.003993 | 0.003922 | 5.7e-05 | 0.000935 | 0.2959 | 1.3422 | 2.5475 | 100.18 | 20.82 | 37.481 | 86.5599 | 83.1241 | 81.8752 | 0.009213 | 0.99442 | 0.023735 | 23.8678 | 0.007331 | 0.15822 | 0.004208 | 0.29227 | 0.29227 | 0.022033
2019-03-29 14:12 | 1553883091.5413 | 1104 | 0.20872 | 0.001045 | 0.000129 | 100.6189 | 100.4684 | 120.514 | 21.0945 | 0.002633 | 0.002507 | 220.9959 | 0.26423 | 8.4023 | 0.002537 | 0.003939 | 0.003922 | 5e-05 | 0.000929 | 0.30309 | 1.3873 | 2.5475 | 100.18 | 20.82 | 37.479 | 84.9983 | 82.6976 | 81.9613 | 0.009212 | 0.99418 | 0.02371 | 23.8487 | 0.013395 | 0.15719 | 0.004117 | 0.28544 | 0.28544 | 0.022413
2019-03-29 14:12 | 1553883092.542 | 1105 | 0.20871 | 0.001042 | 0.000128 | 100.7204 | 100.5745 | 120.4462 | 21.0931 | 0.002676 | 0.002479 | 220.9938 | 0.26411 | 8.4034 | 0.002539 | 0.004947 | 0.004765 | 5.3e-05 | 0.000924 | 0.27865 | 1.4169 | 2.5476 | 100.18 | 20.82 | 37.477 | 86.0381 | 82.7244 | 81.2973 | 0.009212 | 0.99604 | 0.023725 | 23.8197 | 0.012966 | 0.15712 | 0.004179 | 0.29011 | 0.29011 | 0.022042
2019-03-29 14:12 | 1553883093.5425 | 1106 | 0.2087 | 0.001039 | 0.000128 | 100.8033 | 100.4888 | 120.5581 | 21.0945 | 0.002653 | 0.002507 | 221.0124 | 0.26495 | 8.4023 | 0.002573 | 0.004434 | 0.004375 | 5.3e-05 | 0.000919 | 0.33572 | 1.4176 | 2.5475 | 100.18 | 20.82 | 37.476 | 85.4689 | 83.5644 | 81.9748 | 0.009211 | 0.99414 | 0.023714 | 23.8535 | 0.027948 | 0.15762 | 0.004201 | 0.29153 | 0.29153 | 0.022005
2019-03-29 14:12 | 1553883094.5422 | 1107 | 0.20871 | 0.00104 | 0.000127 | 100.7461 | 100.5868 | 120.829 | 21.0946 | 0.002676 | 0.002449 | 220.9929 | 0.26444 | 8.4006 | 0.002532 | 0.00362 | 0.003509 | 5.4e-05 | 0.000913 | 0.2907 | 1.3521 | 2.5474 | 100.18 | 20.82 | 37.474 | 86.0372 | 82.566 | 80.5851 | 0.009211 | 0.99805 | 0.023787 | 23.8335 | 0.014159 | 0.15731 | 0.004144 | 0.28844 | 0.28844 | 0.022198
2019-03-29 14:12 | 1553883095.5422 | 1108 | 0.20871 | 0.001039 | 0.000126 | 100.7632 | 100.6004 | 120.9276 | 21.0937 | 0.002675 | 0.002534 | 220.9783 | 0.26467 | 8.4008 | 0.002536 | 0.004081 | 0.004024 | 5.5e-05 | 0.000909 | 0.31454 | 1.3875 | 2.5476 | 100.17 | 20.82 | 37.474 | 86.0105 | 82.6634 | 82.6079 | 0.009212 | 0.99238 | 0.023729 | 23.9113 | 0.01446 | 0.15745 | 0.004176 | 0.29002 | 0.29002 | 0.022097
2019-03-29 14:12 | 1553883096.5412 | 1109 | 0.20872 | 0.001042 | 0.000126 | 100.747 | 100.4489 | 121.009 | 21.0935 | 0.002669 | 0.002413 | 220.9824 | 0.2644 | 8.4022 | 0.002554 | 0.003963 | 0.003904 | 5.2e-05 | 0.000909 | 0.32084 | 1.3369 | 2.5474 | 100.17 | 20.82 | 37.476 | 85.8598 | 83.103 | 79.699 | 0.009212 | 1.0006 | 0.023835 | 23.8214 | 0.026509 | 0.15729 | 0.004104 | 0.28626 | 0.28626 | 0.022364
2019-03-29 14:12 | 1553883097.5422 | 1110 | 0.20873 | 0.001042 | 0.000126 | 100.7758 | 100.7008 | 120.9917 | 21.0935 | 0.002609 | 0.002516 | 220.9889 | 0.265 | 8.4032 | 0.002451 | 0.004443 | 0.004311 | 5.4e-05 | 0.000902 | 0.28152 | 1.3397 | 2.5473 | 100.17 | 20.82 | 37.48 | 84.4168 | 80.6122 | 82.1858 | 0.009213 | 0.99355 | 0.023749 | 23.9035 | 0.006659 | 0.15765 | 0.004024 | 0.27951 | 0.27951 | 0.022956
2019-03-29 14:12 | 1553883098.5422 | 1111 | 0.20873 | 0.00104 | 0.000125 | 100.791 | 100.4226 | 120.8858 | 21.0934 | 0.002683 | 0.002486 | 220.9963 | 0.26443 | 8.4026 | 0.002551 | 0.00407 | 0.004025 | 5.2e-05 | 0.000902 | 0.26344 | 1.3455 | 2.5474 | 100.17 | 20.82 | 37.483 | 86.1882 | 83.0273 | 81.4626 | 0.009214 | 0.99558 | 0.023763 | 23.8687 | 0.032755 | 0.15731 | 0.004068 | 0.28283 | 0.28283 | 0.022636
2019-03-29 14:12 | 1553883099.5414 | 1112 | 0.20873 | 0.001045 | 0.000124 | 100.8566 | 100.6766 | 120.7992 | 21.0939 | 0.00268 | 0.002476 | 220.9827 | 0.26446 | 8.4016 | 0.002484 | 0.00376 | 0.003689 | 5.5e-05 | 0.000904 | 0.31061 | 1.328 | 2.5474 | 100.17 | 20.82 | 37.486 | 86.1327 | 81.4148 | 81.2175 | 0.009215 | 0.99627 | 0.023763 | 23.8519 | 0.015982 | 0.15733 | 0.00406 | 0.28231 | 0.28231 | 0.022681
2019-03-29 14:12 | 1553883100.5424 | 1113 | 0.20873 | 0.001048 | 0.000124 | 100.8425 | 100.6527 | 120.694 | 21.0947 | 0.002643 | 0.00246 | 220.9595 | 0.2651 | 8.4004 | 0.002531 | 0.004444 | 0.004339 | 5.3e-05 | 0.000896 | 0.30917 | 1.3518 | 2.5474 | 100.17 | 20.82 | 37.495 | 85.2475 | 82.5418 | 80.8496 | 0.009217 | 0.9973 | 0.023765 | 23.8291 | 0.016857 | 0.15771 | 0.004038 | 0.2808 | 0.2808 | 0.022858
2019-03-29 14:12 | 1553883101.5426 | 1114 | 0.20874 | 0.001045 | 0.000123 | 100.8455 | 100.6759 | 120.6058 | 21.0961 | 0.002656 | 0.00247 | 220.9377 | 0.26468 | 8.4002 | 0.002551 | 0.004103 | 0.004063 | 5.4e-05 | 0.000902 | 0.30326 | 1.3684 | 2.5475 | 100.17 | 20.83 | 37.501 | 85.5462 | 83.0189 | 81.0848 | 0.009224 | 0.99664 | 0.023748 | 23.8283 | 0.015061 | 0.15746 | 0.003989 | 0.27716 | 0.27716 | 0.023122
2019-03-29 14:12 | 1553883102.5408 | 1115 | 0.20872 | 0.001048 | 0.000121 | 100.8476 | 100.5305 | 120.606 | 21.095 | 0.002625 | 0.002466 | 220.9302 | 0.26453 | 8.403 | 0.00252 | 0.004398 | 0.004311 | 5.3e-05 | 0.0009 | 0.30309 | 1.3592 | 2.5475 | 100.17 | 20.82 | 37.516 | 84.8021 | 82.2918 | 80.9864 | 0.009222 | 0.99692 | 0.023752 | 23.825 | 0.028173 | 0.15737 | 0.004106 | 0.28564 | 0.28564 | 0.022423
2019-03-29 14:12 | 1553883103.5405 | 1116 | 0.20875 | 0.00105 | 0.000119 | 100.9326 | 100.6298 | 120.3143 | 21.0938 | 0.002653 | 0.00247 | 220.919 | 0.26391 | 8.4023 | 0.002527 | 0.003396 | 0.003323 | 5.2e-05 | 0.000902 | 0.30574 | 1.3335 | 2.5475 | 100.17 | 20.82 | 37.524 | 85.4847 | 82.4575 | 81.0833 | 0.009224 | 0.99665 | 0.02372 | 23.7995 | 0.026878 | 0.157 | 0.003944 | 0.2738 | 0.2738 | 0.023337
2019-03-29 14:12 | 1553883104.5415 | 1117 | 0.20875 | 0.001049 | 0.000119 | 100.9718 | 100.6858 | 119.9441 | 21.0943 | 0.00263 | 0.00242 | 220.9142 | 0.26425 | 8.4011 | 0.002472 | 0.003237 | 0.003089 | 5.3e-05 | 0.000905 | 0.31006 | 1.3954 | 2.5475 | 100.17 | 20.82 | 37.532 | 84.9255 | 81.1275 | 79.8674 | 0.009226 | 1.0001 | 0.023724 | 23.722 | 0.025378 | 0.1572 | 0.003904 | 0.27104 | 0.27104 | 0.023606
2019-03-29 14:12 | 1553883105.5408 | 1118 | 0.20877 | 0.00105 | 0.000117 | 100.9132 | 100.6189 | 119.6882 | 21.0948 | 0.002667 | 0.002449 | 220.9013 | 0.2645 | 8.4018 | 0.002438 | 0.00467 | 0.004567 | 5.6e-05 | 0.000904 | 0.3131 | 1.3553 | 2.5474 | 100.17 | 20.82 | 37.548 | 85.8134 | 80.3023 | 80.5846 | 0.00923 | 0.99805 | 0.023675 | 23.7207 | 0.026124 | 0.15735 | 0.003839 | 0.26594 | 0.26594 | 0.024082
2019-03-29 14:12 | 1553883106.542 | 1119 | 0.20874 | 0.001053 | 0.000116 | 100.9506 | 100.7714 | 119.246 | 21.095 | 0.002651 | 0.002495 | 220.8816 | 0.26504 | 8.4013 | 0.00248 | 0.003067 | 0.003005 | 5.4e-05 | 0.000905 | 0.31166 | 1.3669 | 2.5475 | 100.17 | 20.82 | 37.556 | 85.4227 | 81.328 | 81.6803 | 0.009232 | 0.99497 | 0.023594 | 23.7135 | 0.015889 | 0.15767 | 0.003963 | 0.27397 | 0.27397 | 0.023424
2019-03-29 14:12 | 1553883107.5416 | 1120 | 0.20875 | 0.00105 | 0.000115 | 101.0048 | 100.7912 | 118.9489 | 21.0943 | 0.002628 | 0.002505 | 220.868 | 0.26405 | 8.4002 | 0.002442 | 0.003684 | 0.003596 | 5.3e-05 | 0.000908 | 0.30735 | 1.3159 | 2.5474 | 100.17 | 20.82 | 37.578 | 84.8757 | 80.4163 | 81.919 | 0.009237 | 0.9943 | 0.023557 | 23.6919 | 0.018931 | 0.15709 | 0.003924 | 0.27082 | 0.27082 | 0.023607
2019-03-29 14:12 | 1553883108.5406 | 1121 | 0.20875 | 0.001043 | 0.000113 | 100.9243 | 100.882 | 119.0294 | 21.0948 | 0.002663 | 0.002465 | 220.8543 | 0.26482 | 8.4008 | 0.002437 | 0.00412 | 0.003997 | 5.6e-05 | 0.000908 | 0.29927 | 1.352 | 2.5474 | 100.17 | 20.82 | 37.578 | 85.7285 | 80.2782 | 80.9533 | 0.009237 | 0.99701 | 0.023597 | 23.6677 | 0.003745 | 0.15754 | 0.003967 | 0.2744 | 0.2744 | 0.023366
2019-03-29 14:12 | 1553883109.5417 | 1122 | 0.20876 | 0.001048 | 0.000112 | 100.9751 | 100.683 | 118.728 | 21.095 | 0.002699 | 0.002531 | 220.8492 | 0.26446 | 8.404 | 0.002452 | 0.003846 | 0.003795 | 5.8e-05 | 0.000911 | 0.29878 | 1.3473 | 2.5474 | 100.17 | 20.83 | 37.597 | 86.5918 | 80.6424 | 82.5428 | 0.009248 | 0.99256 | 0.023514 | 23.6907 | 0.025913 | 0.15733 | 0.003896 | 0.26848 | 0.26848 | 0.02385
2019-03-29 14:12 | 1553883110.5409 | 1123 | 0.20876 | 0.001048 | 0.000111 | 100.9317 | 100.7677 | 118.5491 | 21.0951 | 0.002741 | 0.002507 | 220.8275 | 0.26381 | 8.4004 | 0.002393 | 0.003784 | 0.003806 | 5.4e-05 | 0.000911 | 0.29695 | 1.3404 | 2.5474 | 100.17 | 20.83 | 37.597 | 87.5928 | 79.2288 | 81.9629 | 0.009248 | 0.99418 | 0.023516 | 23.6535 | 0.014551 | 0.15694 | 0.003866 | 0.26644 | 0.26644 | 0.023973
2019-03-29 14:12 | 1553883111.5417 | 1124 | 0.20874 | 0.001047 | 0.000111 | 100.9763 | 100.8612 | 118.74 | 21.0949 | 0.002661 | 0.002451 | 220.8015 | 0.26418 | 8.404 | 0.002484 | 0.003882 | 0.003802 | 5.6e-05 | 0.000909 | 0.31238 | 1.35 | 2.5474 | 100.17 | 20.83 | 37.602 | 85.6723 | 81.412 | 80.6323 | 0.009249 | 0.99792 | 0.023579 | 23.6282 | 0.010206 | 0.15716 | 0.004015 | 0.27776 | 0.27776 | 0.023028
2019-03-29 14:12 | 1553883112.5418 | 1125 | 0.20875 | 0.001045 | 0.00011 | 100.9985 | 100.753 | 118.6918 | 21.0937 | 0.002627 | 0.002452 | 220.8051 | 0.26408 | 8.4029 | 0.002457 | 0.004405 | 0.004307 | 5.7e-05 | 0.000905 | 0.30055 | 1.3767 | 2.5474 | 100.17 | 20.83 | 37.607 | 84.8432 | 80.7628 | 80.652 | 0.00925 | 0.99786 | 0.023574 | 23.6241 | 0.021767 | 0.1571 | 0.003969 | 0.2744 | 0.2744 | 0.023302
2019-03-29 14:12 | 1553883113.5422 | 1126 | 0.20874 | 0.001047 | 0.00011 | 100.9425 | 100.6885 | 118.7064 | 21.0944 | 0.002602 | 0.002447 | 220.7885 | 0.26438 | 8.4012 | 0.002441 | 0.004221 | 0.004163 | 5.8e-05 | 0.000908 | 0.32947 | 1.3656 | 2.5474 | 100.17 | 20.83 | 37.614 | 84.246 | 80.372 | 80.5287 | 0.009252 | 0.99821 | 0.023579 | 23.6214 | 0.022537 | 0.15728 | 0.003984 | 0.27559 | 0.27559 | 0.023227
2019-03-29 14:12 | 1553883114.5408 | 1127 | 0.20876 | 0.00104 | 0.000108 | 101.017 | 100.826 | 118.8604 | 21.0953 | 0.002688 | 0.002492 | 220.7734 | 0.26449 | 8.4008 | 0.002388 | 0.004032 | 0.003953 | 5.8e-05 | 0.000913 | 0.31332 | 1.3447 | 2.5474 | 100.17 | 20.83 | 37.616 | 86.3292 | 79.0981 | 81.6198 | 0.009252 | 0.99514 | 0.023558 | 23.6731 | 0.016929 | 0.15735 | 0.003872 | 0.26755 | 0.26755 | 0.023936
2019-03-29 14:12 | 1553883115.5421 | 1128 | 0.20874 | 0.001044 | 0.000108 | 101.0718 | 100.8059 | 118.9627 | 21.0947 | 0.002588 | 0.002469 | 220.7602 | 0.26548 | 8.3995 | 0.002512 | 0.004203 | 0.004157 | 5.8e-05 | 0.00091 | 0.29302 | 1.3114 | 2.5475 | 100.17 | 20.83 | 37.619 | 83.9077 | 82.0911 | 81.0633 | 0.009253 | 0.9967 | 0.023587 | 23.6647 | 0.023565 | 0.15793 | 0.004035 | 0.27937 | 0.27937 | 0.023008
2019-03-29 14:12 | 1553883116.5419 | 1129 | 0.20875 | 0.001042 | 0.000107 | 101.0012 | 100.7069 | 119.1627 | 21.0953 | 0.002659 | 0.002428 | 220.7493 | 0.2648 | 8.4013 | 0.002478 | 0.003873 | 0.003745 | 5.8e-05 | 0.000908 | 0.28407 | 1.3694 | 2.5473 | 100.17 | 20.83 | 37.621 | 85.6289 | 81.2744 | 80.0602 | 0.009254 | 0.99953 | 0.02364 | 23.6511 | 0.026097 | 0.15753 | 0.003924 | 0.27217 | 0.27217 | 0.023557
2019-03-29 14:12 | 1553883117.5399 | 1130 | 0.20876 | 0.001039 | 0.000106 | 101.079 | 100.6489 | 119.3331 | 21.0955 | 0.00256 | 0.002463 | 220.7425 | 0.26434 | 8.4024 | 0.002492 | 0.004067 | 0.004052 | 5.5e-05 | 0.000906 | 0.30773 | 1.38 | 2.5474 | 100.17 | 20.83 | 37.622 | 83.2515 | 81.6113 | 80.9143 | 0.009254 | 0.99712 | 0.023628 | 23.6966 | 0.038141 | 0.15726 | 0.003879 | 0.26891 | 0.26891 | 0.023801
2019-03-29 14:12 | 1553883118.541 | 1131 | 0.20878 | 0.001038 | 0.000105 | 101.0215 | 100.8967 | 119.2653 | 21.0949 | 0.002691 | 0.002467 | 220.7278 | 0.26437 | 8.4008 | 0.002481 | 0.003745 | 0.003727 | 5.7e-05 | 0.000907 | 0.29463 | 1.3781 | 2.5473 | 100.17 | 20.83 | 37.621 | 86.3959 | 81.3432 | 81.019 | 0.009254 | 0.99683 | 0.023618 | 23.6933 | 0.011062 | 0.15727 | 0.003799 | 0.2632 | 0.2632 | 0.02432
2019-03-29 14:12 | 1553883119.54 | 1132 | 0.20876 | 0.001039 | 0.000105 | 101.1119 | 100.7974 | 119.5413 | 21.0938 | 0.002614 | 0.002467 | 220.7005 | 0.26415 | 8.4014 | 0.002472 | 0.004107 | 0.004034 | 5.8e-05 | 0.000905 | 0.29662 | 1.3759 | 2.5475 | 100.17 | 20.83 | 37.62 | 84.5405 | 81.1263 | 81.0148 | 0.009253 | 0.99684 | 0.023646 | 23.7206 | 0.02787 | 0.15714 | 0.00388 | 0.26931 | 0.26931 | 0.023748
2019-03-29 14:12 | 1553883120.541 | 1133 | 0.20878 | 0.001043 | 0.000105 | 101.1296 | 100.9233 | 119.8141 | 21.0944 | 0.002606 | 0.002489 | 220.7102 | 0.26487 | 8.4019 | 0.002485 | 0.00396 | 0.003888 | 5.7e-05 | 0.000902 | 0.2975 | 1.3746 | 2.5474 | 100.17 | 20.83 | 37.616 | 84.3388 | 81.4396 | 81.5393 | 0.009252 | 0.99537 | 0.023655 | 23.7652 | 0.018265 | 0.15757 | 0.003763 | 0.26106 | 0.26106 | 0.024565
2019-03-29 14:12 | 1553883121.541 | 1134 | 0.2088 | 0.001038 | 0.000105 | 101.1053 | 100.9202 | 119.9037 | 21.095 | 0.002633 | 0.002473 | 220.7234 | 0.265 | 8.4041 | 0.002525 | 0.003489 | 0.003461 | 5.8e-05 | 0.000896 | 0.29214 | 1.4139 | 2.5475 | 100.17 | 20.83 | 37.614 | 84.995 | 82.4033 | 81.1563 | 0.009252 | 0.99644 | 0.023677 | 23.7613 | 0.016392 | 0.15765 | 0.003658 | 0.25383 | 0.25383 | 0.025279
2019-03-29 14:12 | 1553883122.5415 | 1135 | 0.20877 | 0.001034 | 0.000106 | 101.1149 | 100.9704 | 120.0561 | 21.0946 | 0.002593 | 0.00246 | 220.6987 | 0.26416 | 8.401 | 0.002577 | 0.003695 | 0.003636 | 5.9e-05 | 0.000894 | 0.29446 | 1.3384 | 2.5474 | 100.17 | 20.83 | 37.61 | 84.0428 | 83.6549 | 80.844 | 0.009251 | 0.99732 | 0.023702 | 23.7659 | 0.012794 | 0.15715 | 0.003863 | 0.26866 | 0.26866 | 0.023807
2019-03-29 14:12 | 1553883123.5401 | 1136 | 0.20878 | 0.001035 | 0.000106 | 101.1119 | 101.0663 | 120.2176 | 21.0951 | 0.002605 | 0.002484 | 220.6901 | 0.2642 | 8.4021 | 0.002513 | 0.004011 | 0.003846 | 5.8e-05 | 0.000892 | 0.28982 | 1.3229 | 2.5474 | 100.17 | 20.84 | 37.604 | 84.3318 | 82.1078 | 81.4146 | 0.009255 | 0.99572 | 0.023699 | 23.801 | 0.004038 | 0.15718 | 0.003776 | 0.26244 | 0.26244 | 0.024375
2019-03-29 14:12 | 1553883124.5401 | 1137 | 0.20875 | 0.001037 | 0.000106 | 101.1895 | 100.9458 | 120.2093 | 21.0938 | 0.002659 | 0.002436 | 220.6675 | 0.26419 | 8.4031 | 0.002546 | 0.003917 | 0.003874 | 5.5e-05 | 0.000892 | 0.30127 | 1.3268 | 2.5473 | 100.17 | 20.84 | 37.601 | 85.6336 | 82.903 | 80.2637 | 0.009254 | 0.99896 | 0.023737 | 23.7615 | 0.021562 | 0.15717 | 0.003934 | 0.27411 | 0.27411 | 0.023336
2019-03-29 14:12 | 1553883125.5403 | 1138 | 0.20877 | 0.001036 | 0.000105 | 101.1529 | 100.8509 | 120.2528 | 21.0939 | 0.002574 | 0.002465 | 220.6856 | 0.26427 | 8.4026 | 0.002452 | 0.004013 | 0.003898 | 5.6e-05 | 0.000888 | 0.29983 | 1.3165 | 2.5474 | 100.17 | 20.84 | 37.594 | 83.5739 | 80.6467 | 80.9618 | 0.009253 | 0.99699 | 0.023718 | 23.7893 | 0.026744 | 0.15721 | 0.003818 | 0.26563 | 0.26563 | 0.024088
2019-03-29 14:12 | 1553883126.5401 | 1139 | 0.20878 | 0.001035 | 0.000105 | 101.2359 | 101.0683 | 120.7163 | 21.095 | 0.002602 | 0.002442 | 208.7496 | 0.26404 | 8.402 | 0.002478 | 0.003967 | 0.004027 | 5.6e-05 | 0.000884 | 0.28838 | 1.3058 | 2.5473 | 100.17 | 20.84 | 37.591 | 84.256 | 81.2679 | 80.399 | 0.009252 | 0.99857 | 0.023782 | 23.8162 | 0.014812 | 0.15708 | 0.003761 | 0.26234 | 0.26234 | 0.02437
2019-03-29 14:12 | 1553883127.5412 | 1140 | 0.2088 | 0.001033 | 0.000105 | 101.2643 | 101.0158 | 121.0717 | 21.0938 | 0.002649 | 0.002445 | 171.469 | 0.26362 | 8.4039 | 0.002506 | 0.003716 | 0.003597 | 5.8e-05 | 0.000887 | 0.30574 | 1.3323 | 2.5473 | 100.17 | 20.84 | 37.588 | 85.3825 | 81.938 | 80.4721 | 0.009251 | 0.99837 | 0.023815 | 23.8537 | 0.021977 | 0.15683 | 0.003693 | 0.25782 | 0.25782 | 0.024757
2019-03-29 14:12 | 1553883128.5413 | 1141 | 0.2088 | 0.001038 | 0.000104 | 101.2332 | 100.9953 | 121.5501 | 21.0952 | 0.002643 | 0.002464 | 134.1923 | 0.26437 | 8.4002 | 0.002542 | 0.004258 | 0.004182 | 5.7e-05 | 0.000883 | 0.30182 | 1.3435 | 2.5474 | 100.17 | 20.84 | 37.581 | 85.2386 | 82.8161 | 80.9303 | 0.00925 | 0.99708 | 0.023846 | 23.9162 | 0.02104 | 0.15727 | 0.003693 | 0.25823 | 0.25823 | 0.024788
2019-03-29 14:12 | 1553883129.5414 | 1142 | 0.20882 | 0.001034 | 0.000102 | 101.2529 | 101.0332 | 121.5886 | 21.0945 | 0.00265 | 0.002446 | 96.9219 | 0.26359 | 8.402 | 0.002495 | 0.004477 | 0.004385 | 5.9e-05 | 0.000882 | 0.26322 | 1.3723 | 2.5473 | 100.17 | 20.84 | 37.577 | 85.412 | 81.669 | 80.5086 | 0.009249 | 0.99827 | 0.023864 | 23.9057 | 0.019431 | 0.15681 | 0.003559 | 0.24895 | 0.24895 | 0.025636
2019-03-29 14:12 | 1553883130.5413 | 1143 | 0.20884 | 0.001024 | 9.8e-05 | 101.2445 | 101.0728 | 121.8972 | 21.0956 | 0.002593 | 0.002444 | 59.6489 | 0.26441 | 8.4027 | 0.002529 | 0.004829 | 0.004736 | 5.6e-05 | 0.000882 | 0.30326 | 1.356 | 2.5475 | 100.17 | 20.84 | 37.573 | 84.0312 | 82.4991 | 80.4457 | 0.009248 | 0.99844 | 0.023897 | 23.9339 | 0.015186 | 0.15729 | 0.003474 | 0.24341 | 0.24341 | 0.026301
2019-03-29 14:12 | 1553883131.5414 | 1144 | 0.2088 | 0.001029 | 9.2e-05 | 101.2493 | 101.0151 | 121.7839 | 21.0952 | 0.002593 | 0.002368 | 22.3717 | 0.25447 | 8.4015 | 0.002509 | 0.004058 | 0.003982 | 5.6e-05 | 0.000881 | 0.29015 | 1.3813 | 2.5474 | 100.17 | 20.84 | 37.566 | 84.0416 | 82.0151 | 78.6241 | 0.009246 | 1.0036 | 0.023947 | 23.8611 | 0.020715 | 0.15139 | 0.003702 | 0.2607 | 0.2607 | 0.023634
2019-03-29 14:12 | 1553883132.5398 | 1145 | 0.20895 | 0.001026 | 8.5e-05 | 101.2622 | 101.0858 | 121.9378 | 21.0948 | 0.002602 | 0.002434 | -14.9033 | 0.23493 | 8.402 | 0.002533 | 0.004098 | 0.003942 | 5.7e-05 | 0.00088 | 0.30596 | 1.3932 | 2.5475 | 100.17 | 20.84 | 37.558 | 84.2488 | 82.5888 | 80.2095 | 0.009244 | 0.99911 | 0.023909 | 23.9299 | 0.015599 | 0.13976 | 0.002821 | 0.19745 | 0.19745 | 0.028809
2019-03-29 14:12 | 1553883133.5398 | 1146 | 0.20896 | 0.001015 | 7.8e-05 | 101.2625 | 101.0073 | 122.3939 | 21.0951 | 0.00262 | 0.002354 | -52.1884 | 0.23515 | 8.4021 | 0.00249 | 0.003748 | 0.0036 | 5.8e-05 | 0.000882 | 0.26288 | 1.3287 | 2.5475 | 100.17 | 20.84 | 37.558 | 84.6906 | 81.5696 | 78.2958 | 0.009244 | 1.0046 | 0.024018 | 23.9096 | 0.022574 | 0.13989 | 0.002737 | 0.19278 | 0.19278 | 0.029535
2019-03-29 14:12 | 1553883134.5394 | 1147 | 0.20905 | 0.001003 | 7e-05 | 101.2245 | 101.1134 | 122.9469 | 21.095 | 0.002522 | 0.002364 | -89.4432 | 0.23531 | 8.4026 | 0.00254 | 0.004211 | 0.003991 | 5.9e-05 | 0.00088 | 0.26897 | 1.3285 | 2.5474 | 100.17 | 20.84 | 37.554 | 82.3284 | 82.7597 | 78.5413 | 0.009243 | 1.0038 | 0.024064 | 23.972 | 0.009821 | 0.13999 | 0.002213 | 0.15568 | 0.15568 | 0.036597
2019-03-29 14:12 | 1553883135.5398 | 1148 | 0.20917 | 0.000985 | 6.3e-05 | 101.3104 | 101.0861 | 123.5997 | 21.0952 | 0.002587 | 0.002361 | -126.7237 | 0.23509 | 8.4034 | 0.002517 | 0.004095 | 0.003818 | 5.8e-05 | 0.000882 | 0.28854 | 1.3261 | 2.5475 | 100.17 | 20.84 | 37.551 | 83.8988 | 82.2132 | 78.4706 | 0.009242 | 1.0041 | 0.02413 | 24.0331 | 0.019824 | 0.13986 | 0.001576 | 0.1104 | 0.1104 | 0.051558
2019-03-29 14:12 | 1553883136.5396 | 1149 | 0.20919 | 0.00095 | 5.7e-05 | 101.3272 | 101.1356 | 123.6168 | 21.0948 | 0.002575 | 0.002427 | -152.0622 | 0.23533 | 8.4016 | 0.002478 | 0.003742 | 0.003563 | 6e-05 | 0.000883 | 0.29806 | 1.3206 | 2.5472 | 100.17 | 20.84 | 37.546 | 83.6154 | 81.2638 | 80.0448 | 0.009241 | 0.99958 | 0.024078 | 24.0885 | 0.016926 | 0.14 | 0.001488 | 0.10417 | 0.10417 | 0.054698
2019-03-29 14:12 | 1553883137.5406 | 1150 | 0.20928 | 0.000915 | 5e-05 | 101.2703 | 101.1735 | 123.7337 | 21.0945 | 0.002537 | 0.002346 | -152.0512 | 0.23552 | 8.402 | 0.002569 | 0.003347 | 0.003204 | 5.9e-05 | 0.000879 | 0.30525 | 1.317 | 2.5473 | 100.17 | 20.84 | 37.543 | 82.6868 | 83.4684 | 78.0903 | 0.00924 | 1.0051 | 0.024157 | 24.0331 | 0.008555 | 0.14011 | 0.000982 | 0.068182 | 0.068182 | 0.083638
2019-03-29 14:12 | 1553883138.5396 | 1151 | 0.2093 | 0.000875 | 4.5e-05 | 101.3254 | 101.1755 | 123.7922 | 21.0952 | 0.002467 | 0.002322 | -152.0548 | 0.23492 | 8.4038 | 0.002544 | 0.004074 | 0.003904 | 6.1e-05 | 0.000875 | 0.31133 | 1.3305 | 2.5475 | 100.17 | 20.84 | 37.539 | 80.9982 | 82.8575 | 77.5203 | 0.009239 | 1.0068 | 0.024182 | 24.0193 | 0.013237 | 0.13975 | 0.00094 | 0.065588 | 0.065588 | 0.086722
2019-03-29 14:12 | 1553883139.54 | 1152 | 0.2093 | 0.00084 | 3.9e-05 | 101.3283 | 101.1301 | 124.2137 | 21.0957 | 0.002542 | 0.002386 | -152.0553 | 0.23546 | 8.4011 | 0.002514 | 0.00399 | 0.003776 | 6e-05 | 0.000877 | 0.29535 | 1.3474 | 2.5474 | 100.17 | 20.84 | 37.535 | 82.8128 | 82.1452 | 79.0727 | 0.009238 | 1.0023 | 0.02417 | 24.1133 | 0.017515 | 0.14007 | 0.000971 | 0.068091 | 0.068091 | 0.083726
2019-03-29 14:12 | 1553883140.54 | 1153 | 0.20935 | 0.000804 | 3.4e-05 | 101.3032 | 101.2052 | 124.1729 | 21.0946 | 0.002522 | 0.002345 | -152.0561 | 0.23581 | 8.4023 | 0.002524 | 0.00397 | 0.00378 | 5.8e-05 | 0.000885 | 0.3115 | 1.3514 | 2.5474 | 100.17 | 20.84 | 37.533 | 82.328 | 82.367 | 78.0692 | 0.009238 | 1.0052 | 0.0242 | 24.075 | 0.008657 | 0.14028 | 0.000747 | 0.05224 | 0.05224 | 0.10929
2019-03-29 14:12 | 1553883141.5396 | 1154 | 0.20941 | 0.00077 | 2.9e-05 | 101.3245 | 101.2666 | 124.5158 | 21.0958 | 0.00245 | 0.002298 | -152.0645 | 0.23545 | 8.4024 | 0.002545 | 0.004408 | 0.004247 | 6e-05 | 0.00089 | 0.31492 | 1.3378 | 2.5475 | 100.17 | 20.84 | 37.528 | 80.5887 | 82.8736 | 76.9416 | 0.009237 | 1.0084 | 0.024273 | 24.0695 | 0.005106 | 0.14007 | 0.000417 | 0.02864 | 0.02864 | 0.19905
2019-03-29 14:12 | 1553883142.5397 | 1155 | 0.20943 | 0.000734 | 2.5e-05 | 101.2987 | 101.2267 | 124.8163 | 21.0957 | 0.00245 | 0.002293 | -152.0657 | 0.2358 | 8.4013 | 0.002556 | 0.00409 | 0.003905 | 6.3e-05 | 0.000894 | 0.29264 | 1.3466 | 2.5475 | 100.17 | 20.85 | 37.527 | 80.6018 | 83.1512 | 76.8377 | 0.009242 | 1.0087 | 0.024305 | 24.0949 | 0.006361 | 0.14028 | 0.00034 | 0.023299 | 0.023299 | 0.24504
2019-03-29 14:12 | 1553883143.5408 | 1156 | 0.20944 | 0.0007 | 2.1e-05 | 101.3445 | 101.3431 | 124.7327 | 21.0952 | 0.002499 | 0.002328 | -152.0549 | 0.23481 | 8.4021 | 0.00254 | 0.003683 | 0.003508 | 5.8e-05 | 0.000899 | 0.28031 | 1.3519 | 2.5474 | 100.17 | 20.85 | 37.524 | 81.775 | 82.7685 | 77.6558 | 0.009241 | 1.0064 | 0.024269 | 24.115 | 0.000126 | 0.13969 | 0.000358 | 0.024803 | 0.024803 | 0.22923
2019-03-29 14:12 | 1553883144.5398 | 1157 | 0.20943 | 0.000668 | 1.7e-05 | 101.3343 | 101.225 | 124.8076 | 21.0956 | 0.002442 | 0.002258 | -152.0728 | 0.23507 | 8.4029 | 0.002544 | 0.003728 | 0.003573 | 5.9e-05 | 0.000905 | 0.30917 | 1.3059 | 2.5473 | 100.17 | 20.85 | 37.517 | 80.4145 | 82.854 | 75.9896 | 0.00924 | 1.0112 | 0.024334 | 24.0649 | 0.009656 | 0.13984 | 0.000446 | 0.03155 | 0.03155 | 0.1804
2019-03-29 14:12 | 1553883145.5397 | 1158 | 0.20947 | 0.000641 | 1.4e-05 | 101.3478 | 101.2868 | 125.0111 | 21.0949 | 0.002413 | 0.002278 | -152.0927 | 0.23489 | 8.4009 | 0.00253 | 0.003693 | 0.0035 | 5.9e-05 | 0.000906 | 0.30093 | 1.3156 | 2.5475 | 100.17 | 20.85 | 37.514 | 79.7176 | 82.5148 | 76.4777 | 0.009239 | 1.0098 | 0.024337 | 24.1013 | 0.005388 | 0.13973 | 0.000265 | 0.018513 | 0.018513 | 0.30721
2019-03-29 14:12 | 1553883146.5402 | 1159 | 0.20949 | 0.000619 | 1.2e-05 | 101.3298 | 101.2677 | 124.9583 | 21.0954 | 0.002503 | 0.002292 | -152.1074 | 0.23488 | 8.4006 | 0.002612 | 0.003644 | 0.003461 | 5.7e-05 | 0.000906 | 0.2907 | 1.3443 | 2.547 | 100.17 | 20.85 | 37.51 | 81.8749 | 84.4945 | 76.8096 | 0.009238 | 1.0088 | 0.02432 | 24.1077 | 0.005491 | 0.13973 | 0.000161 | 0.011011 | 0.011011 | 0.51649
2019-03-29 14:12 | 1553883147.5412 | 1160 | 0.2095 | 0.000597 | 1.1e-05 | 101.3574 | 101.1687 | 124.8721 | 21.0953 | 0.002405 | 0.002246 | -152.1244 | 0.23463 | 8.4011 | 0.002535 | 0.003976 | 0.003765 | 5.7e-05 | 0.000904 | 0.31332 | 1.3326 | 2.5475 | 100.17 | 20.85 | 37.507 | 79.5068 | 82.6352 | 75.694 | 0.009237 | 1.012 | 0.024351 | 24.0609 | 0.016668 | 0.13958 | 0.000122 | 0.008292 | 0.008292 | 0.68507
2019-03-29 14:12 | 1553883148.5389 | 1161 | 0.20948 | 0.000577 | 9e-06 | 101.3652 | 101.2369 | 124.8102 | 21.0952 | 0.002401 | 0.002269 | -152.1345 | 0.23464 | 8.4029 | 0.002484 | 0.004192 | 0.003984 | 5.8e-05 | 0.000906 | 0.30879 | 1.3371 | 2.5473 | 100.17 | 20.85 | 37.505 | 79.4161 | 81.4099 | 76.2419 | 0.009237 | 1.0105 | 0.024326 | 24.0738 | 0.011322 | 0.13959 | 0.000263 | 0.018642 | 0.018642 | 0.30475
2019-03-29 14:12 | 1553883149.5402 | 1162 | 0.20951 | 0.000562 | 7e-06 | 101.3089 | 101.2253 | 124.6703 | 21.0955 | 0.002418 | 0.002226 | -152.1345 | 0.23542 | 8.4019 | 0.002542 | 0.00398 | 0.003783 | 6.1e-05 | 0.00091 | 0.32853 | 1.355 | 2.5474 | 100.17 | 20.86 | 37.499 | 79.8242 | 82.8078 | 75.2132 | 0.009241 | 1.0134 | 0.024348 | 24.0249 | 0.00738 | 0.14005 | 0.00013 | 0.009039 | 0.009039 | 0.63062
2019-03-29 14:12 | 1553883150.5394 | 1163 | 0.20951 | 0.000546 | 6e-06 | 101.3122 | 101.1431 | 124.8212 | 21.0953 | 0.002416 | 0.00227 | -152.1469 | 0.2348 | 8.4017 | 0.002522 | 0.004387 | 0.004175 | 6e-05 | 0.000919 | 0.27151 | 1.3401 | 2.5473 | 100.17 | 20.86 | 37.496 | 79.7921 | 82.3335 | 76.274 | 0.00924 | 1.0104 | 0.024326 | 24.076 | 0.014941 | 0.13968 | 0.000146 | 0.010294 | 0.010294 | 0.55225
2019-03-29 14:12 | 1553883151.54 | 1164 | 0.20952 | 0.000535 | 6e-06 | 101.3077 | 101.0366 | 125.124 | 21.0947 | 0.002436 | 0.002149 | -152.1554 | 0.23463 | 8.4007 | 0.002524 | 0.003877 | 0.003777 | 5.9e-05 | 0.000925 | 0.28174 | 1.3355 | 2.5474 | 100.17 | 20.86 | 37.493 | 80.2701 | 82.3718 | 73.3759 | 0.009239 | 1.0188 | 0.024457 | 24.005 | 0.023967 | 0.13958 | 0.000106 | 0.007427 | 0.007427 | 0.76493
2019-03-29 14:13 | 1553883152.5389 | 1165 | 0.20951 | 0.000523 | 5e-06 | 101.3307 | 101.2629 | 125.0395 | 21.095 | 0.002316 | 0.002215 | -152.1621 | 0.2355 | 8.4002 | 0.002541 | 0.003674 | 0.003477 | 5.9e-05 | 0.000924 | 0.31277 | 1.3306 | 2.5473 | 100.17 | 20.86 | 37.491 | 77.3905 | 82.7971 | 74.9609 | 0.009239 | 1.0142 | 0.024393 | 24.0517 | 0.005992 | 0.1401 | 0.000186 | 0.013326 | 0.013326 | 0.4279
2019-03-29 14:13 | 1553883153.5401 | 1166 | 0.20951 | 0.000511 | 4e-06 | 101.3014 | 101.1165 | 125.1557 | 21.0954 | 0.002388 | 0.002208 | -152.1677 | 0.23571 | 8.4013 | 0.002544 | 0.004119 | 0.003989 | 5.7e-05 | 0.000922 | 0.30309 | 1.3774 | 2.5475 | 100.17 | 20.86 | 37.487 | 79.1179 | 82.8671 | 74.7735 | 0.009238 | 1.0147 | 0.024411 | 24.0564 | 0.016344 | 0.14022 | 0.000151 | 0.010789 | 0.010789 | 0.52896
2019-03-29 14:13 | 1553883154.5405 | 1167 | 0.20952 | 0.000509 | 3e-06 | 101.3092 | 101.1567 | 125.4075 | 21.0959 | 0.002413 | 0.002255 | -152.1687 | 0.23534 | 8.4026 | 0.002515 | 0.004194 | 0.003954 | 5.8e-05 | 0.000921 | 0.32189 | 1.3362 | 2.5472 | 100.17 | 20.86 | 37.485 | 79.7099 | 82.1711 | 75.9158 | 0.009237 | 1.0114 | 0.024395 | 24.1201 | 0.01347 | 0.14 | 0.000108 | 0.007712 | 0.007712 | 0.73886
2019-03-29 14:13 | 1553883155.5392 | 1168 | 0.20953 | 0.000498 | 3e-06 | 101.3337 | 101.2141 | 125.3355 | 21.0957 | 0.00232 | 0.002221 | -152.153 | 0.23569 | 8.4025 | 0.002534 | 0.00435 | 0.004204 | 5.8e-05 | 0.00092 | 0.29103 | 1.338 | 2.5473 | 100.17 | 20.86 | 37.484 | 77.4871 | 82.6168 | 75.105 | 0.009237 | 1.0138 | 0.024416 | 24.0852 | 0.010568 | 0.14021 | 4.4e-05 | 0.003056 | 0.003056 | 1.8674
2019-03-29 14:13 | 1553883156.5401 | 1169 | 0.20953 | 0.000493 | 2e-06 | 101.267 | 101.084 | 125.5627 | 21.0941 | 0.002328 | 0.00219 | -152.1371 | 0.2348 | 8.4031 | 0.00261 | 0.004261 | 0.004175 | 5.8e-05 | 0.000915 | 0.27743 | 1.3323 | 2.5473 | 100.17 | 20.86 | 37.481 | 77.6689 | 84.4549 | 74.3443 | 0.009236 | 1.016 | 0.024465 | 24.0806 | 0.016173 | 0.13968 | 6.9e-05 | 0.00491 | 0.00491 | 1.1578
2019-03-29 14:13 | 1553883157.5405 | 1170 | 0.20954 | 0.000491 | 2e-06 | 101.2775 | 101.2018 | 125.8364 | 21.0958 | 0.002386 | 0.00219 | -152.1321 | 0.23535 | 8.4029 | 0.002592 | 0.004204 | 0.003955 | 5.7e-05 | 0.000911 | 0.26089 | 1.3534 | 2.5474 | 100.17 | 20.86 | 37.481 | 79.0582 | 84.0251 | 74.356 | 0.009236 | 1.0159 | 0.024492 | 24.1073 | 0.006686 | 0.14001 | 4.9e-05 | 0.003441 | 0.003441 | 1.656
2019-03-29 14:13 | 1553883158.5406 | 1171 | 0.20952 | 0.000485 | 2e-06 | 101.2119 | 101.1711 | 125.8082 | 21.094 | 0.002322 | 0.00223 | -152.1193 | 0.23566 | 8.4012 | 0.002554 | 0.003877 | 0.003735 | 6.1e-05 | 0.000914 | 0.31133 | 1.3809 | 2.5474 | 100.17 | 20.87 | 37.48 | 77.5292 | 83.0902 | 75.3013 | 0.009242 | 1.0132 | 0.024456 | 24.1373 | 0.003609 | 0.14019 | 0.000171 | 0.012474 | 0.012474 | 0.45743
2019-03-29 14:13 | 1553883159.5403 | 1172 | 0.2095 | 0.000477 | 2e-06 | 101.1895 | 101.1028 | 125.7725 | 21.0954 | 0.002331 | 0.002176 | -152.1043 | 0.23436 | 8.4028 | 0.002615 | 0.004248 | 0.003962 | 5.6e-05 | 0.000918 | 0.29446 | 1.3935 | 2.5475 | 100.17 | 20.87 | 37.479 | 77.7286 | 84.5584 | 74.0049 | 0.009242 | 1.017 | 0.024498 | 24.089 | 0.007662 | 0.13942 | 0.000294 | 0.021535 | 0.021535 | 0.2635
2019-03-29 14:13 | 1553883160.5403 | 1173 | 0.20956 | 0.000478 | 1e-06 | 101.2098 | 100.9782 | 125.9643 | 21.0958 | 0.0023 | 0.002197 | -152.0925 | 0.23494 | 8.4014 | 0.00258 | 0.003704 | 0.003516 | 5.8e-05 | 0.000921 | 0.30237 | 1.3451 | 2.5474 | 100.17 | 20.87 | 37.48 | 76.9992 | 83.7349 | 74.5244 | 0.009242 | 1.0154 | 0.024498 | 24.1254 | 0.020488 | 0.13977 | -9.6e-05 | -0.007147 | -0.007147 | -0.79598
2019-03-29 14:13 | 1553883161.5403 | 1174 | 0.20955 | 0.000476 | 1e-06 | 101.2658 | 100.995 | 125.9392 | 21.0951 | 0.002323 | 0.002203 | -152.0875 | 0.23499 | 8.4004 | 0.002613 | 0.003723 | 0.003546 | 5.8e-05 | 0.000917 | 0.31006 | 1.3585 | 2.5473 | 100.17 | 20.87 | 37.479 | 77.538 | 84.5143 | 74.6717 | 0.009242 | 1.015 | 0.02449 | 24.1281 | 0.023953 | 0.1398 | -3.2e-05 | -0.002471 | -0.002471 | -2.3022
2019-03-29 14:13 | 1553883162.5403 | 1175 | 0.20952 | 0.000474 | 1e-06 | 101.2499 | 101.0612 | 126.2004 | 21.0957 | 0.002331 | 0.00217 | -152.0851 | 0.23441 | 8.4028 | 0.002591 | 0.003455 | 0.003339 | 5.9e-05 | 0.000923 | 0.31022 | 1.3783 | 2.5476 | 100.17 | 20.87 | 37.479 | 77.7476 | 83.9875 | 73.8625 | 0.009242 | 1.0174 | 0.024544 | 24.125 | 0.016689 | 0.13945 | 0.00015 | 0.010983 | 0.010983 | 0.51679
2019-03-29 14:13 | 1553883163.5394 | 1176 | 0.20954 | 0.000473 | 1e-06 | 101.2371 | 100.9318 | 125.9774 | 21.0954 | 0.002245 | 0.002157 | -152.0811 | 0.23513 | 8.403 | 0.002546 | 0.003932 | 0.003746 | 5.8e-05 | 0.000918 | 0.29767 | 1.3831 | 2.5475 | 100.17 | 20.87 | 37.479 | 75.6816 | 82.913 | 73.5597 | 0.009242 | 1.0183 | 0.024533 | 24.0931 | 0.027008 | 0.13988 | 3.8e-05 | 0.002772 | 0.002772 | 2.0536
2019-03-29 14:13 | 1553883164.6933 | 1177 | 0.20953 | 0.00047 | 1e-06 | 101.1631 | 101.0482 | 125.7185 | 21.0956 | 0.002338 | 0.002143 | -152.0699 | 0.2349 | 8.4005 | 0.002553 | 0.003996 | 0.003879 | 5.8e-05 | 0.000923 | 0.29286 | 1.3786 | 2.5475 | 100.17 | 20.87 | 37.479 | 77.8986 | 83.0653 | 73.2163 | 0.009242 | 1.0193 | 0.02452 | 24.0564 | 0.010165 | 0.13974 | 0.000102 | 0.007433 | 0.007433 | 0.76514
2019-03-29 14:13 | 1553883165.5841 | 1178 | 0.20955 | 0.000464 | 1e-06 | 101.1778 | 100.9318 | 125.3277 | 21.0957 | 0.002226 | 0.002123 | -152.0728 | 0.23348 | 8.4026 | 0.002512 | 0.004308 | 0.004122 | 6e-05 | 0.000922 | 0.3115 | 1.3687 | 2.5474 | 100.17 | 20.87 | 37.479 | 75.2252 | 82.0812 | 72.7469 | 0.009242 | 1.0207 | 0.024499 | 24.0027 | 0.02177 | 0.1389 | 8e-06 | 0.000558 | 0.000558 | 10.1276
2019-03-29 14:13 | 1553883166.5841 | 1179 | 0.20953 | 0.000466 | 1e-06 | 101.2275 | 100.9895 | 124.9684 | 21.0955 | 0.002313 | 0.002129 | -152.0719 | 0.23481 | 8.4 | 0.002515 | 0.003407 | 0.003301 | 5.9e-05 | 0.000924 | 0.31133 | 1.3419 | 2.5473 | 100.17 | 20.87 | 37.479 | 77.3115 | 82.1627 | 72.8786 | 0.009242 | 1.0203 | 0.024459 | 23.9729 | 0.021051 | 0.13969 | 0.000111 | 0.008121 | 0.008121 | 0.70011
2019-03-29 14:13 | 1553883167.5841 | 1180 | 0.20954 | 0.000469 | 1e-06 | 101.1616 | 100.9578 | 124.5508 | 21.0958 | 0.002298 | 0.002167 | -152.0844 | 0.23514 | 8.4047 | 0.002476 | 0.003904 | 0.003717 | 5.9e-05 | 0.000921 | 0.2855 | 1.3217 | 2.5475 | 100.17 | 20.87 | 37.478 | 76.9479 | 81.2242 | 73.8073 | 0.009241 | 1.0175 | 0.024385 | 23.9649 | 0.018042 | 0.13988 | 8e-05 | 0.005803 | 0.005803 | 0.98107
2019-03-29 14:13 | 1553883168.5852 | 1181 | 0.20953 | 0.000464 | 1e-06 | 101.1945 | 101.1533 | 124.5724 | 21.0945 | 0.002304 | 0.002157 | -152.0844 | 0.23526 | 8.4023 | 0.002482 | 0.004144 | 0.003887 | 5.8e-05 | 0.000922 | 0.28241 | 1.3204 | 2.5475 | 100.17 | 20.87 | 37.478 | 77.0893 | 81.3779 | 73.5643 | 0.009241 | 1.0183 | 0.024396 | 23.9586 | 0.003644 | 0.13995 | 0.000107 | 0.007845 | 0.007845 | 0.72605
2019-03-29 14:13 | 1553883169.5841 | 1182 | 0.20954 | 0.000458 | 1e-06 | 101.1631 | 100.8967 | 124.8127 | 21.0962 | 0.002307 | 0.002138 | -152.1006 | 0.23511 | 8.4016 | 0.002449 | 0.003315 | 0.003118 | 6.3e-05 | 0.00092 | 0.30165 | 1.367 | 2.5475 | 100.17 | 20.87 | 37.476 | 77.1619 | 80.5663 | 73.1054 | 0.009241 | 1.0196 | 0.024436 | 23.9658 | 0.023589 | 0.13987 | 7.6e-05 | 0.005539 | 0.005539 | 1.0277
2019-03-29 14:13 | 1553883170.5852 | 1183 | 0.20955 | 0.000464 | 1e-06 | 101.1221 | 100.9615 | 124.5806 | 21.0947 | 0.00229 | 0.002124 | -152.1054 | 0.23485 | 8.4013 | 0.002485 | 0.003075 | 0.002831 | 5.8e-05 | 0.000919 | 0.28136 | 1.3885 | 2.5477 | 100.17 | 20.87 | 37.475 | 76.7512 | 81.4319 | 72.7733 | 0.009241 | 1.0206 | 0.024425 | 23.932 | 0.014216 | 0.13971 | 3.3e-05 | 0.00238 | 0.00238 | 2.3888
2019-03-29 14:13 | 1553883171.5844 | 1184 | 0.20954 | 0.000458 | 1e-06 | 101.1323 | 101.0206 | 124.5358 | 21.0955 | 0.002302 | 0.002087 | -152.1147 | 0.23443 | 8.4028 | 0.002487 | 0.003964 | 0.003865 | 5.9e-05 | 0.000925 | 0.29031 | 1.4167 | 2.5475 | 100.17 | 20.87 | 37.473 | 77.0459 | 81.4929 | 71.876 | 0.00924 | 1.0232 | 0.024452 | 23.8966 | 0.009886 | 0.13946 | 4.9e-05 | 0.003528 | 0.003528 | 1.6088
2019-03-29 14:13 | 1553883172.5843 | 1185 | 0.20954 | 0.000462 | 0 | 101.0895 | 100.8168 | 124.0133 | 21.0955 | 0.002242 | 0.002108 | -152.135 | 0.23523 | 8.4 | 0.002461 | 0.003677 | 0.003502 | 5.7e-05 | 0.00092 | 0.28528 | 1.3863 | 2.5476 | 100.17 | 20.87 | 37.473 | 75.6004 | 80.8674 | 72.3733 | 0.00924 | 1.0218 | 0.024383 | 23.8636 | 0.024158 | 0.13994 | 8.7e-05 | 0.006366 | 0.006366 | 0.89461
2019-03-29 14:13 | 1553883173.5844 | 1186 | 0.20955 | 0.00046 | 1e-06 | 101.1047 | 101.1462 | 124.0823 | 21.096 | 0.002267 | 0.002084 | -152.1289 | 0.23507 | 8.4025 | 0.002474 | 0.003778 | 0.003637 | 6e-05 | 0.000927 | 0.3115 | 1.42 | 2.5472 | 100.17 | 20.87 | 37.474 | 76.2129 | 81.1713 | 71.801 | 0.00924 | 1.0235 | 0.02441 | 23.8505 | -0.003663 | 0.13985 | 1.3e-05 | 0.000914 | 0.000914 | 6.2243
2019-03-29 14:13 | 1553883174.5852 | 1187 | 0.20957 | 0.000461 | 1e-06 | 101.1024 | 100.8834 | 124.2823 | 21.0948 | 0.002289 | 0.002084 | -152.1312 | 0.23501 | 8.4038 | 0.002516 | 0.002988 | 0.002822 | 6e-05 | 0.00093 | 0.2891 | 1.3744 | 2.5473 | 100.17 | 20.87 | 37.479 | 76.7194 | 82.184 | 71.7999 | 0.009242 | 1.0235 | 0.02443 | 23.8697 | 0.019395 | 0.13981 | -9.9e-05 | -0.007348 | -0.007348 | -0.77439
2019-03-29 14:13 | 1553883175.5843 | 1188 | 0.20957 | 0.000459 | 1e-06 | 101.0287 | 100.8312 | 124.885 | 21.0958 | 0.002223 | 0.002044 | -152.1373 | 0.23494 | 8.4009 | 0.002521 | 0.004016 | 0.003757 | 5.6e-05 | 0.00093 | 0.29856 | 1.3634 | 2.5476 | 100.17 | 20.87 | 37.483 | 75.1565 | 82.2988 | 70.8449 | 0.009243 | 1.0263 | 0.024523 | 23.8943 | 0.017508 | 0.13977 | -9.6e-05 | -0.007144 | -0.007144 | -0.7963
2019-03-29 14:13 | 1553883176.5855 | 1189 | 0.2095 | 0.000461 | 1e-06 | 101.0545 | 100.8622 | 124.9655 | 21.0958 | 0.002275 | 0.002086 | -152.1457 | 0.23509 | 8.4035 | 0.002578 | 0.00351 | 0.003368 | 5.6e-05 | 0.000931 | 0.26001 | 1.3771 | 2.5475 | 100.17 | 20.87 | 37.493 | 76.3967 | 83.6678 | 71.8648 | 0.009245 | 1.0233 | 0.024495 | 23.9374 | 0.017033 | 0.13986 | 0.000331 | 0.024326 | 0.024326 | 0.23399
2019-03-29 14:13 | 1553883177.5854 | 1190 | 0.20959 | 0.000459 | 0 | 100.9991 | 100.7875 | 125.0843 | 21.0963 | 0.002234 | 0.00209 | -152.147 | 0.23556 | 8.4027 | 0.002512 | 0.003939 | 0.003743 | 6e-05 | 0.000936 | 0.26432 | 1.3298 | 2.5474 | 100.17 | 20.87 | 37.499 | 75.4019 | 82.0895 | 71.9517 | 0.009246 | 1.023 | 0.024503 | 23.9518 | 0.018761 | 0.14014 | -0.000201 | -0.014842 | -0.014842 | -0.3843
2019-03-29 14:13 | 1553883178.5855 | 1191 | 0.20957 | 0.000461 | 0 | 101.0958 | 100.8731 | 125.0601 | 21.0956 | 0.002258 | 0.002102 | -152.1392 | 0.23503 | 8.4013 | 0.002501 | 0.003701 | 0.00353 | 5.8e-05 | 0.00094 | 0.3095 | 1.3333 | 2.5475 | 100.17 | 20.88 | 37.505 | 75.9843 | 81.823 | 72.2262 | 0.009254 | 1.0222 | 0.024491 | 23.959 | 0.01972 | 0.13982 | -8.7e-05 | -0.006416 | -0.006416 | -0.88695
2019-03-29 14:13 | 1553883179.5848 | 1192 | 0.20951 | 0.000457 | 1e-06 | 100.9961 | 100.8509 | 125.0953 | 21.096 | 0.002206 | 0.002107 | -152.1346 | 0.2348 | 8.4027 | 0.002493 | 0.003951 | 0.003795 | 5.6e-05 | 0.000936 | 0.29966 | 1.3415 | 2.5475 | 100.17 | 20.88 | 37.518 | 74.7313 | 81.6361 | 72.3505 | 0.009257 | 1.0218 | 0.02449 | 23.9667 | 0.012864 | 0.13968 | 0.000265 | 0.019439 | 0.019439 | 0.29246
2019-03-29 14:13 | 1553883180.5836 | 1193 | 0.20957 | 0.000456 | 1e-06 | 100.965 | 100.8039 | 124.8432 | 21.0947 | 0.002234 | 0.002049 | -152.1474 | 0.23475 | 8.401 | 0.002547 | 0.003689 | 0.003506 | 5.8e-05 | 0.00094 | 0.28888 | 1.346 | 2.5474 | 100.17 | 20.88 | 37.525 | 75.4089 | 82.9185 | 70.9639 | 0.009259 | 1.026 | 0.024515 | 23.8945 | 0.014284 | 0.13965 | -9.3e-05 | -0.006901 | -0.006901 | -0.82362
2019-03-29 14:13 | 1553883181.5841 | 1194 | 0.20956 | 0.000454 | 1e-06 | 100.9281 | 100.5301 | 124.8522 | 21.0947 | 0.002134 | 0.002086 | -152.1514 | 0.23472 | 8.4031 | 0.00254 | 0.003882 | 0.003578 | 5.7e-05 | 0.000936 | 0.30862 | 1.3911 | 2.5475 | 100.17 | 20.88 | 37.532 | 73.0112 | 82.7635 | 71.8627 | 0.00926 | 1.0233 | 0.024484 | 23.9265 | 0.035342 | 0.13964 | -1.6e-05 | -0.001205 | -0.001205 | -4.7169
2019-03-29 14:13 | 1553883182.5846 | 1195 | 0.20956 | 0.000453 | 1e-06 | 100.9542 | 100.7243 | 124.9853 | 21.0957 | 0.002175 | 0.002038 | -152.1272 | 0.23502 | 8.4003 | 0.002544 | 0.004013 | 0.003857 | 5.9e-05 | 0.000933 | 0.31492 | 1.3554 | 2.5473 | 100.17 | 20.88 | 37.545 | 73.992 | 82.8681 | 70.6959 | 0.009264 | 1.0268 | 0.024538 | 23.8987 | 0.020388 | 0.13981 | -5e-05 | -0.003729 | -0.003729 | -1.526
2019-03-29 14:13 | 1553883183.5836 | 1196 | 0.20955 | 0.000458 | 1e-06 | 100.8021 | 100.6339 | 125.173 | 21.0954 | 0.002255 | 0.00208 | -152.1317 | 0.23341 | 8.4031 | 0.002463 | 0.003629 | 0.003572 | 5.7e-05 | 0.000935 | 0.30215 | 1.3677 | 2.5473 | 100.17 | 20.88 | 37.55 | 75.9038 | 80.9111 | 71.7147 | 0.009265 | 1.0237 | 0.02452 | 23.9521 | 0.014941 | 0.13886 | 7e-06 | 0.000463 | 0.000463 | 12.2114
2019-03-29 14:13 | 1553883184.5836 | 1197 | 0.20953 | 0.00045 | 1e-06 | 100.8279 | 100.5748 | 124.9503 | 21.0962 | 0.002261 | 0.002067 | -152.1395 | 0.23373 | 8.4016 | 0.002523 | 0.004258 | 0.004039 | 5.4e-05 | 0.000939 | 0.28335 | 1.374 | 2.5474 | 100.17 | 20.88 | 37.561 | 76.0525 | 82.3538 | 71.4073 | 0.009267 | 1.0246 | 0.024509 | 23.9201 | 0.022474 | 0.13905 | 0.000156 | 0.011431 | 0.011431 | 0.49508
2019-03-29 14:13 | 1553883185.5837 | 1198 | 0.20956 | 0.000458 | 0 | 100.7836 | 100.5837 | 124.3017 | 21.0958 | 0.002144 | 0.002061 | -152.1454 | 0.23432 | 8.4045 | 0.00254 | 0.00338 | 0.003199 | 5.5e-05 | 0.000936 | 0.29103 | 1.4003 | 2.5476 | 100.17 | 20.88 | 37.564 | 73.2585 | 82.7722 | 71.2642 | 0.009268 | 1.0251 | 0.024451 | 23.853 | 0.017754 | 0.1394 | -2.8e-05 | -0.002064 | -0.002064 | -2.7487
2019-03-29 14:13 | 1553883186.5847 | 1199 | 0.20956 | 0.000453 | 1e-06 | 100.7668 | 100.5779 | 124.3841 | 21.0961 | 0.002199 | 0.002053 | -152.148 | 0.23399 | 8.4031 | 0.002494 | 0.003912 | 0.003755 | 5.3e-05 | 0.000941 | 0.30486 | 1.328 | 2.5473 | 100.17 | 20.88 | 37.564 | 74.5598 | 81.6462 | 71.0518 | 0.009268 | 1.0257 | 0.024466 | 23.8535 | 0.016782 | 0.1392 | -4.7e-05 | -0.003517 | -0.003517 | -1.6111
2019-03-29 14:13 | 1553883187.5837 | 1200 | 0.20955 | 0.000451 | 1e-06 | 100.7399 | 100.3878 | 124.6458 | 21.0953 | 0.00222 | 0.002094 | -152.1537 | 0.23379 | 8.4017 | 0.002482 | 0.004486 | 0.004269 | 5.6e-05 | 0.000952 | 0.31404 | 1.3762 | 2.5474 | 100.17 | 20.88 | 37.559 | 75.082 | 81.3623 | 72.049 | 0.009267 | 1.0227 | 0.024457 | 23.9132 | 0.031312 | 0.13908 | 5e-05 | 0.00366 | 0.00366 | 1.5466
2019-03-29 14:13 | 1553883188.5842 | 1201 | 0.20955 | 0.00046 | 1e-06 | 100.7788 | 100.5632 | 124.8057 | 21.0947 | 0.002204 | 0.002088 | -152.1752 | 0.23477 | 8.4006 | 0.002472 | 0.003962 | 0.003777 | 5.6e-05 | 0.000958 | 0.31885 | 1.3672 | 2.5475 | 100.17 | 20.88 | 37.559 | 74.698 | 81.138 | 71.8893 | 0.009267 | 1.0232 | 0.024478 | 23.923 | 0.01915 | 0.13966 | 1e-06 | 3.4e-05 | 3.4e-05 | 166.7499
2019-03-29 14:13 | 1553883189.5838 | 1202 | 0.20954 | 0.000457 | 1e-06 | 100.745 | 100.6779 | 124.5712 | 21.0961 | 0.002189 | 0.002113 | -152.1825 | 0.23593 | 8.4035 | 0.002504 | 0.004421 | 0.004227 | 5.6e-05 | 0.000958 | 0.29695 | 1.355 | 2.5473 | 100.17 | 20.88 | 37.549 | 74.3174 | 81.8957 | 72.4971 | 0.009265 | 1.0214 | 0.024434 | 23.9215 | 0.005954 | 0.14036 | 7.1e-05 | 0.005132 | 0.005132 | 1.1132
2019-03-29 14:13 | 1553883190.5848 | 1203 | 0.20953 | 0.000451 | 1e-06 | 100.7563 | 100.4554 | 125.0517 | 21.0954 | 0.002177 | 0.002048 | -152.1788 | 0.23494 | 8.4028 | 0.00249 | 0.003927 | 0.003726 | 5.9e-05 | 0.000965 | 0.32875 | 1.3448 | 2.5474 | 100.17 | 20.88 | 37.543 | 74.0462 | 81.5684 | 70.9435 | 0.009263 | 1.026 | 0.024536 | 23.9137 | 0.026755 | 0.13976 | 0.000145 | 0.010656 | 0.010656 | 0.5338
2019-03-29 14:13 | 1553883191.5842 | 1204 | 0.20953 | 0.000453 | 1e-06 | 100.6653 | 100.5827 | 124.951 | 21.0961 | 0.002213 | 0.002078 | -152.1724 | 0.23508 | 8.4026 | 0.00245 | 0.00423 | 0.004036 | 5.7e-05 | 0.000963 | 0.30342 | 1.3162 | 2.5474 | 100.17 | 20.88 | 37.535 | 74.8995 | 80.5974 | 71.6591 | 0.009261 | 1.0239 | 0.0245 | 23.9289 | 0.007345 | 0.13985 | 0.000163 | 0.01195 | 0.01195 | 0.47629
2019-03-29 14:13 | 1553883192.5841 | 1205 | 0.20954 | 0.000453 | 0 | 100.6123 | 100.5851 | 125.1606 | 21.0957 | 0.002133 | 0.002059 | -152.1842 | 0.23489 | 8.3998 | 0.002417 | 0.003833 | 0.003647 | 5.6e-05 | 0.000969 | 0.27505 | 1.3116 | 2.5474 | 100.17 | 20.88 | 37.519 | 72.9715 | 79.8126 | 71.2043 | 0.009257 | 1.0252 | 0.024537 | 23.9332 | 0.002424 | 0.13974 | 9.3e-05 | 0.006862 | 0.006862 | 0.82882
2019-03-29 14:13 | 1553883193.584 | 1206 | 0.20956 | 0.000454 | 1e-06 | 100.6165 | 100.4359 | 124.9658 | 21.0956 | 0.002156 | 0.002069 | -152.1799 | 0.23532 | 8.4006 | 0.002384 | 0.004104 | 0.003895 | 5.9e-05 | 0.000968 | 0.31564 | 1.324 | 2.5474 | 100.17 | 20.88 | 37.511 | 73.5335 | 79.0212 | 71.4426 | 0.009255 | 1.0245 | 0.02451 | 23.9228 | 0.016069 | 0.13999 | -5.1e-05 | -0.00382 | -0.00382 | -1.4917
2019-03-29 14:13 | 1553883194.585 | 1207 | 0.20949 | 0.00045 | 0 | 100.5755 | 100.3516 | 125.0343 | 21.096 | 0.002174 | 0.002027 | -152.1876 | 0.23508 | 8.4009 | 0.00241 | 0.003966 | 0.003875 | 5.8e-05 | 0.000971 | 0.30646 | 1.3035 | 2.5473 | 100.17 | 20.88 | 37.497 | 73.9781 | 79.6367 | 70.4421 | 0.009252 | 1.0275 | 0.024552 | 23.8946 | 0.019933 | 0.13985 | 0.000372 | 0.027447 | 0.027447 | 0.20737
2019-03-29 14:13 | 1553883195.5844 | 1208 | 0.2095 | 0.000453 | 1e-06 | 100.5507 | 100.2039 | 125.4624 | 21.0947 | 0.002169 | 0.002055 | -152.1709 | 0.23607 | 8.401 | 0.002466 | 0.003436 | 0.0031 | 5.7e-05 | 0.000971 | 0.2959 | 1.3067 | 2.5474 | 100.17 | 20.89 | 37.488 | 73.8499 | 80.9938 | 71.1033 | 0.009255 | 1.0255 | 0.02457 | 23.9585 | 0.030905 | 0.14044 | 0.000314 | 0.023089 | 0.023089 | 0.24756
2019-03-29 14:13 | 1553883196.5829 | 1209 | 0.20952 | 0.000455 | 1e-06 | 100.5504 | 100.3114 | 125.4885 | 21.0957 | 0.00217 | 0.002098 | -152.1824 | 0.23511 | 8.3998 | 0.00246 | 0.004004 | 0.003923 | 5.7e-05 | 0.000975 | 0.27168 | 1.3416 | 2.5475 | 100.17 | 20.88 | 37.48 | 73.8805 | 80.8294 | 72.1506 | 0.009247 | 1.0224 | 0.024536 | 23.9974 | 0.021287 | 0.13986 | 0.000204 | 0.014989 | 0.014989 | 0.37978
2019-03-29 14:13 | 1553883197.5845 | 1210 | 0.20953 | 0.000458 | 1e-06 | 100.5663 | 100.2011 | 125.4251 | 21.096 | 0.002163 | 0.002006 | -152.1918 | 0.23493 | 8.4029 | 0.002467 | 0.003756 | 0.003532 | 5.4e-05 | 0.00098 | 0.33572 | 1.3496 | 2.5475 | 100.17 | 20.88 | 37.464 | 73.7078 | 80.9982 | 69.9266 | 0.009244 | 1.0291 | 0.024609 | 23.914 | 0.032533 | 0.13976 | 0.000122 | 0.008938 | 0.008938 | 0.6364
2019-03-29 14:13 | 1553883198.5847 | 1211 | 0.20954 | 0.000456 | 1e-06 | 100.4582 | 100.2407 | 125.1322 | 21.0961 | 0.002123 | 0.001989 | -152.1887 | 0.23391 | 8.4033 | 0.002352 | 0.00433 | 0.00416 | 5.8e-05 | 0.000983 | 0.29048 | 1.3285 | 2.5471 | 100.17 | 20.89 | 37.457 | 72.7545 | 78.2484 | 69.5298 | 0.009248 | 1.0302 | 0.024594 | 23.8722 | 0.019384 | 0.13915 | 9.4e-05 | 0.00693 | 0.00693 | 0.8173
2019-03-29 14:13 | 1553883199.5831 | 1212 | 0.20952 | 0.000455 | 1e-06 | 100.5252 | 100.3906 | 125.0646 | 21.0963 | 0.002138 | 0.00198 | -152.1825 | 0.23507 | 8.4026 | 0.002446 | 0.004364 | 0.004193 | 5.6e-05 | 0.000986 | 0.30829 | 1.3762 | 2.5474 | 100.17 | 20.89 | 37.444 | 73.1093 | 80.5041 | 69.314 | 0.009244 | 1.0309 | 0.024595 | 23.8582 | 0.011993 | 0.13984 | 0.000236 | 0.01738 | 0.01738 | 0.32747
2019-03-29 14:13 | 1553883200.5832 | 1213 | 0.20949 | 0.000453 | 0 | 100.4735 | 100.1124 | 124.6062 | 21.0961 | 0.00217 | 0.002057 | -152.1928 | 0.23397 | 8.4009 | 0.002382 | 0.003841 | 0.003708 | 5.7e-05 | 0.00099 | 0.30326 | 1.3522 | 2.5475 | 100.17 | 20.89 | 37.439 | 73.8628 | 78.975 | 71.1523 | 0.009243 | 1.0254 | 0.024485 | 23.8783 | 0.032201 | 0.13919 | 0.000359 | 0.026416 | 0.026416 | 0.21445
2019-03-29 14:13 | 1553883201.5835 | 1214 | 0.20949 | 0.000452 | 1e-06 | 100.4747 | 100.1516 | 124.8885 | 21.0962 | 0.002099 | 0.002033 | -152.1744 | 0.23474 | 8.4025 | 0.00243 | 0.003926 | 0.00372 | 5.8e-05 | 0.00099 | 0.33555 | 1.3491 | 2.5474 | 100.17 | 20.89 | 37.435 | 72.1746 | 80.1274 | 70.5708 | 0.009242 | 1.0271 | 0.024533 | 23.8852 | 0.028801 | 0.13965 | 0.000399 | 0.029381 | 0.029381 | 0.19345
2019-03-29 14:13 | 1553883202.5843 | 1215 | 0.20952 | 0.000458 | 1e-06 | 100.4579 | 100.3028 | 125.0104 | 21.0958 | 0.002157 | 0.002031 | -152.1712 | 0.23464 | 8.404 | 0.002479 | 0.003386 | 0.003257 | 5.8e-05 | 0.000988 | 0.30326 | 1.3314 | 2.5474 | 100.17 | 20.89 | 37.428 | 73.5478 | 81.303 | 70.5346 | 0.00924 | 1.0272 | 0.024546 | 23.8955 | 0.013816 | 0.13959 | 0.000174 | 0.012818 | 0.012818 | 0.44322
2019-03-29 14:13 | 1553883203.5843 | 1216 | 0.20954 | 0.000454 | 0 | 100.3977 | 100.1045 | 125.0505 | 21.0969 | 0.002145 | 0.002009 | -152.1753 | 0.23593 | 8.4043 | 0.002484 | 0.003599 | 0.003349 | 5.6e-05 | 0.000983 | 0.2995 | 1.3293 | 2.5473 | 100.17 | 20.89 | 37.426 | 73.2659 | 81.4209 | 70.0039 | 0.00924 | 1.0288 | 0.024569 | 23.8809 | 0.026158 | 0.14036 | 0.0001 | 0.007384 | 0.007384 | 0.77365
2019-03-29 14:13 | 1553883204.5844 | 1217 | 0.20951 | 0.000457 | 1e-06 | 100.4091 | 100.2643 | 124.7604 | 21.0963 | 0.002091 | 0.002042 | -152.17 | 0.23487 | 8.4046 | 0.002404 | 0.003947 | 0.003714 | 5.6e-05 | 0.000984 | 0.31636 | 1.3294 | 2.5475 | 100.17 | 20.89 | 37.424 | 71.9743 | 79.4929 | 70.7877 | 0.009239 | 1.0265 | 0.024513 | 23.8804 | 0.012911 | 0.13973 | 0.000281 | 0.020639 | 0.020639 | 0.27553
2019-03-29 14:13 | 1553883205.5833 | 1218 | 0.20955 | 0.000456 | 1e-06 | 100.3917 | 100.0677 | 124.5484 | 21.0972 | 0.002095 | 0.002017 | -152.1779 | 0.23525 | 8.4022 | 0.002444 | 0.003604 | 0.003492 | 5.7e-05 | 0.000987 | 0.30237 | 1.3699 | 2.5475 | 100.17 | 20.89 | 37.418 | 72.0734 | 80.4598 | 70.2022 | 0.009238 | 1.0282 | 0.024513 | 23.8398 | 0.028918 | 0.13995 | 1.4e-05 | 0.00101 | 0.00101 | 5.6395
2019-03-29 14:13 | 1553883206.5842 | 1219 | 0.20952 | 0.00045 | 1e-06 | 100.4384 | 100.1827 | 124.8576 | 21.0966 | 0.002095 | 0.001999 | -152.1604 | 0.23539 | 8.4036 | 0.002456 | 0.003633 | 0.003391 | 5.8e-05 | 0.000991 | 0.31365 | 1.4189 | 2.5473 | 100.17 | 20.89 | 37.415 | 72.0635 | 80.7483 | 69.7548 | 0.009237 | 1.0296 | 0.024559 | 23.8538 | 0.022803 | 0.14003 | 0.000235 | 0.017311 | 0.017311 | 0.32924
2019-03-29 14:13 | 1553883207.5843 | 1220 | 0.20948 | 0.000454 | 1e-06 | 100.3699 | 100.1001 | 125.0115 | 21.097 | 0.002162 | 0.002038 | -152.1508 | 0.2347 | 8.403 | 0.002444 | 0.003865 | 0.0037 | 5.8e-05 | 0.000988 | 0.30254 | 1.3885 | 2.5474 | 100.17 | 20.89 | 37.408 | 73.6696 | 80.4464 | 70.7019 | 0.009235 | 1.0267 | 0.024541 | 23.9015 | 0.024074 | 0.13962 | 0.000461 | 0.033976 | 0.033976 | 0.16726
2019-03-29 14:13 | 1553883208.5838 | 1221 | 0.20953 | 0.000454 | 1e-06 | 100.337 | 100.1824 | 125.1446 | 21.0969 | 0.002126 | 0.001953 | -152.1522 | 0.2346 | 8.4028 | 0.002443 | 0.0039 | 0.003631 | 5.8e-05 | 0.000986 | 0.30397 | 1.3712 | 2.5475 | 100.17 | 20.89 | 37.404 | 72.8137 | 80.4391 | 68.6555 | 0.009234 | 1.0329 | 0.024627 | 23.8429 | 0.013792 | 0.13956 | 0.00014 | 0.010297 | 0.010297 | 0.55164
2019-03-29 14:13 | 1553883209.5847 | 1222 | 0.20967 | 0.000453 | 1e-06 | 100.3648 | 100.2551 | 125.1651 | 21.0975 | 0.002077 | 0.002025 | -152.1694 | 0.23522 | 8.4004 | 0.002445 | 0.003758 | 0.003523 | 6e-05 | 0.000989 | 0.27328 | 1.4056 | 2.5474 | 100.17 | 20.89 | 37.401 | 71.6347 | 80.4747 | 70.3918 | 0.009234 | 1.0277 | 0.024567 | 23.9054 | 0.009786 | 0.13994 | -0.000713 | -0.052694 | -0.052694 | -0.10808
2019-03-29 14:13 | 1553883210.5829 | 1223 | 0.2095 | 0.000456 | 0 | 100.2594 | 100.0237 | 125.5519 | 21.0977 | 0.002136 | 0.002034 | -152.1608 | 0.23471 | 8.4006 | 0.002448 | 0.00397 | 0.003853 | 6e-05 | 0.000981 | 0.30071 | 1.3987 | 2.5475 | 100.17 | 20.89 | 37.394 | 73.06 | 80.5632 | 70.5995 | 0.009232 | 1.027 | 0.024597 | 23.9495 | 0.021059 | 0.13963 | 0.000327 | 0.024144 | 0.024144 | 0.23538
2019-03-29 14:13 | 1553883211.5826 | 1224 | 0.2095 | 0.000451 | 1e-06 | 100.3415 | 100.0677 | 125.3571 | 21.097 | 0.002172 | 0.001998 | -152.176 | 0.23463 | 8.4031 | 0.002384 | 0.003574 | 0.003362 | 5.8e-05 | 0.000985 | 0.31078 | 1.37 | 2.5474 | 100.17 | 20.89 | 37.389 | 73.9144 | 79.0156 | 69.742 | 0.009231 | 1.0296 | 0.024609 | 23.901 | 0.024437 | 0.13958 | 0.000336 | 0.024784 | 0.024784 | 0.22922
2019-03-29 14:14 | 1553883212.5825 | 1225 | 0.20949 | 0.000456 | 1e-06 | 100.3418 | 100.0752 | 125.2089 | 21.0975 | 0.002143 | 0.00204 | -152.1942 | 0.23482 | 8.4038 | 0.002404 | 0.004042 | 0.00388 | 6e-05 | 0.000977 | 0.28694 | 1.3284 | 2.5475 | 100.17 | 20.89 | 37.38 | 73.232 | 79.5072 | 70.7561 | 0.009229 | 1.0266 | 0.024558 | 23.9222 | 0.023792 | 0.1397 | 0.000368 | 0.02712 | 0.02712 | 0.20964
2019-03-29 14:14 | 1553883213.5835 | 1226 | 0.20951 | 0.000453 | 0 | 100.249 | 100.0998 | 125.3033 | 21.0973 | 0.002149 | 0.002008 | -152.1918 | 0.23519 | 8.4025 | 0.002452 | 0.00314 | 0.002962 | 6.1e-05 | 0.000978 | 0.30453 | 1.3415 | 2.5475 | 100.17 | 20.89 | 37.376 | 73.3667 | 80.643 | 69.9904 | 0.009228 | 1.0289 | 0.024595 | 23.9046 | 0.013321 | 0.13992 | 0.000289 | 0.021305 | 0.021305 | 0.26728
2019-03-29 14:14 | 1553883214.5827 | 1227 | 0.2095 | 0.000453 | 1e-06 | 100.2508 | 99.9585 | 125.7098 | 21.0979 | 0.002124 | 0.002026 | -152.193 | 0.23487 | 8.4029 | 0.002398 | 0.004183 | 0.003994 | 5.9e-05 | 0.000977 | 0.29767 | 1.3441 | 2.5475 | 100.17 | 20.89 | 37.372 | 72.7577 | 79.3541 | 70.4235 | 0.009227 | 1.0276 | 0.024619 | 23.9584 | 0.026115 | 0.13973 | 0.000315 | 0.023203 | 0.023203 | 0.2451
2019-03-29 14:14 | 1553883215.5826 | 1228 | 0.20952 | 0.000451 | 1e-06 | 100.2394 | 100.0991 | 125.9206 | 21.0973 | 0.002141 | 0.001966 | -152.1844 | 0.23494 | 8.4039 | 0.002518 | 0.003888 | 0.003682 | 6e-05 | 0.000975 | 0.28495 | 1.3582 | 2.5475 | 100.17 | 20.89 | 37.364 | 73.1766 | 82.2385 | 68.9702 | 0.009225 | 1.0319 | 0.024692 | 23.9277 | 0.012527 | 0.13977 | 0.00022 | 0.01624 | 0.01624 | 0.35028
2019-03-29 14:14 | 1553883216.583 | 1229 | 0.20952 | 0.00045 | 0 | 100.2909 | 100.011 | 125.7507 | 21.0979 | 0.002137 | 0.001951 | -152.2134 | 0.23404 | 8.4008 | 0.002466 | 0.003823 | 0.003637 | 5.8e-05 | 0.000983 | 0.30431 | 1.3508 | 2.5474 | 100.17 | 20.89 | 37.361 | 73.0907 | 80.986 | 68.6191 | 0.009224 | 1.033 | 0.024688 | 23.8993 | 0.024992 | 0.13923 | 0.00019 | 0.014054 | 0.014054 | 0.40322
2019-03-29 14:14 | 1553883217.5837 | 1230 | 0.20956 | 0.000456 | 0 | 100.2699 | 99.8223 | 125.6589 | 21.0975 | 0.002142 | 0.002011 | -152.2065 | 0.23548 | 8.4008 | 0.00244 | 0.003721 | 0.003578 | 6.1e-05 | 0.000977 | 0.29623 | 1.3153 | 2.5474 | 100.17 | 20.89 | 37.351 | 73.1901 | 80.3625 | 70.0504 | 0.009221 | 1.0287 | 0.024627 | 23.9406 | 0.040018 | 0.14009 | -5.2e-05 | -0.003871 | -0.003871 | -1.4729
2019-03-29 14:14 | 1553883218.5827 | 1231 | 0.20955 | 0.000454 | 1e-06 | 100.2331 | 99.8728 | 125.6608 | 21.0968 | 0.002133 | 0.00198 | -152.2084 | 0.23489 | 8.4009 | 0.002407 | 0.004049 | 0.003807 | 6.3e-05 | 0.000975 | 0.31564 | 1.3595 | 2.5475 | 100.17 | 20.89 | 37.351 | 72.989 | 79.5736 | 69.3034 | 0.009221 | 1.0309 | 0.024654 | 23.9147 | 0.032208 | 0.13974 | 5.4e-05 | 0.003941 | 0.003941 | 1.4432
2019-03-29 14:14 | 1553883219.5843 | 1232 | 0.20955 | 0.000454 | 1e-06 | 100.2379 | 100.0994 | 125.7634 | 21.097 | 0.002097 | 0.00196 | -152.2075 | 0.23543 | 8.404 | 0.002413 | 0.003991 | 0.003767 | 5.9e-05 | 0.000974 | 0.26913 | 1.3186 | 2.5474 | 100.17 | 20.89 | 37.349 | 72.1191 | 79.7235 | 68.8341 | 0.009221 | 1.0323 | 0.024681 | 23.908 | 0.012363 | 0.14006 | 3.9e-05 | 0.002858 | 0.002858 | 1.9943
2019-03-29 14:14 | 1553883220.5832 | 1233 | 0.20954 | 0.000458 | 0 | 100.2226 | 99.8517 | 125.7453 | 21.097 | 0.00208 | 0.001943 | -152.211 | 0.23416 | 8.4003 | 0.002458 | 0.003743 | 0.003608 | 6.3e-05 | 0.000975 | 0.2724 | 1.3047 | 2.5474 | 100.17 | 20.89 | 37.345 | 71.7185 | 80.7931 | 68.4115 | 0.00922 | 1.0336 | 0.024695 | 23.8915 | 0.033169 | 0.1393 | 0.000113 | 0.008342 | 0.008342 | 0.67963
2019-03-29 14:14 | 1553883221.5838 | 1234 | 0.20955 | 0.000453 | 1e-06 | 100.2349 | 99.965 | 125.8756 | 21.0973 | 0.002094 | 0.001969 | -152.2233 | 0.23494 | 8.4018 | 0.002461 | 0.003728 | 0.003557 | 6.3e-05 | 0.00097 | 0.28191 | 1.3338 | 2.5473 | 100.17 | 20.89 | 37.343 | 72.0406 | 80.8677 | 69.0412 | 0.009219 | 1.0317 | 0.024685 | 23.9259 | 0.02412 | 0.13977 | 2.7e-05 | 0.001924 | 0.001924 | 2.9561
2019-03-29 14:14 | 1553883222.5828 | 1235 | 0.20955 | 0.000453 | 0 | 100.1625 | 100.01 | 126.0291 | 21.098 | 0.002174 | 0.001934 | -152.2059 | 0.2344 | 8.4048 | 0.002433 | 0.004547 | 0.004366 | 5.9e-05 | 0.000971 | 0.32316 | 1.3111 | 2.5472 | 100.17 | 20.89 | 37.341 | 73.9667 | 80.1986 | 68.2033 | 0.009219 | 1.0343 | 0.02473 | 23.9112 | 0.013623 | 0.13945 | 2e-05 | 0.001492 | 0.001492 | 3.8037
2019-03-29 14:14 | 1553883223.5834 | 1236 | 0.20955 | 0.000454 | 0 | 100.1699 | 100.0397 | 126.0828 | 21.0969 | 0.00212 | 0.002001 | -152.2198 | 0.2345 | 8.4032 | 0.002431 | 0.004583 | 0.004341 | 5.8e-05 | 0.00097 | 0.30071 | 1.3829 | 2.5475 | 100.17 | 20.89 | 37.337 | 72.682 | 80.143 | 69.8034 | 0.009218 | 1.0294 | 0.024678 | 23.9723 | 0.011636 | 0.1395 | 1.2e-05 | 0.000878 | 0.000878 | 6.4677
2019-03-29 14:14 | 1553883224.5844 | 1237 | 0.20955 | 0.000457 | 1e-06 | 100.2121 | 99.8766 | 125.8019 | 21.098 | 0.002122 | 0.001924 | -152.215 | 0.23522 | 8.4026 | 0.002422 | 0.004265 | 0.004162 | 6e-05 | 0.000965 | 0.3079 | 1.3813 | 2.5475 | 100.17 | 20.89 | 37.335 | 72.7095 | 79.9193 | 67.9545 | 0.009217 | 1.035 | 0.024717 | 23.8809 | 0.030002 | 0.13993 | 4.9e-05 | 0.003556 | 0.003556 | 1.6015
2019-03-29 14:14 | 1553883225.5829 | 1238 | 0.20956 | 0.000456 | 0 | 100.2106 | 99.9472 | 125.8371 | 21.0974 | 0.002058 | 0.001988 | -152.2352 | 0.23474 | 8.4021 | 0.002443 | 0.003233 | 0.003041 | 6.1e-05 | 0.00096 | 0.3011 | 1.3056 | 2.5474 | 100.17 | 20.9 | 37.332 | 71.1886 | 80.4436 | 69.5068 | 0.009222 | 1.0303 | 0.024664 | 23.9386 | 0.023544 | 0.13965 | -4.9e-05 | -0.003628 | -0.003628 | -1.5664
2019-03-29 14:14 | 1553883226.5821 | 1239 | 0.20955 | 0.000455 | 0 | 100.2211 | 99.7476 | 125.7674 | 21.097 | 0.002112 | 0.002001 | -152.2202 | 0.23461 | 8.4038 | 0.002451 | 0.004528 | 0.004381 | 6.1e-05 | 0.000959 | 0.31222 | 1.364 | 2.5475 | 100.16 | 20.9 | 37.329 | 72.4716 | 80.617 | 69.8028 | 0.009223 | 1.0294 | 0.024647 | 23.9423 | 0.042364 | 0.13957 | 2e-06 | 9.9e-05 | 9.9e-05 | 57.1499
2019-03-29 14:14 | 1553883227.583 | 1240 | 0.20955 | 0.000453 | 1e-06 | 100.1648 | 99.9616 | 125.396 | 21.0976 | 0.002082 | 0.001967 | -152.2205 | 0.23448 | 8.402 | 0.002375 | 0.003429 | 0.003181 | 6e-05 | 0.00096 | 0.28672 | 1.3311 | 2.5474 | 100.16 | 20.9 | 37.326 | 71.7574 | 78.7932 | 68.9938 | 0.009222 | 1.0319 | 0.02464 | 23.8787 | 0.018173 | 0.13949 | 3.8e-05 | 0.002778 | 0.002778 | 2.0435
2019-03-29 14:14 | 1553883228.5833 | 1241 | 0.20954 | 0.000459 | 0 | 100.1711 | 99.9192 | 125.4296 | 21.0972 | 0.00211 | 0.001985 | -152.216 | 0.23559 | 8.4009 | 0.002442 | 0.00353 | 0.003413 | 6e-05 | 0.000964 | 0.31797 | 1.3751 | 2.5475 | 100.16 | 20.9 | 37.322 | 72.4288 | 80.4115 | 69.4209 | 0.009221 | 1.0306 | 0.024627 | 23.8968 | 0.022521 | 0.14016 | 8.6e-05 | 0.006332 | 0.006332 | 0.90085
2019-03-29 14:14 | 1553883229.5831 | 1242 | 0.20954 | 0.00045 | 1e-06 | 100.1559 | 99.909 | 125.6068 | 21.0972 | 0.002073 | 0.001998 | -152.2043 | 0.23395 | 8.4035 | 0.002481 | 0.003069 | 0.002865 | 5.8e-05 | 0.000966 | 0.30541 | 1.368 | 2.5475 | 100.16 | 20.9 | 37.321 | 71.5316 | 81.3433 | 69.729 | 0.009221 | 1.0296 | 0.024634 | 23.9244 | 0.022075 | 0.13918 | 7.4e-05 | 0.005451 | 0.005451 | 1.0391
2019-03-29 14:14 | 1553883230.5821 | 1243 | 0.20956 | 0.000449 | 0 | 100.1373 | 99.8677 | 125.4061 | 21.0968 | 0.002123 | 0.001953 | -152.1967 | 0.23567 | 8.4025 | 0.002466 | 0.003844 | 0.003674 | 6.1e-05 | 0.00097 | 0.29247 | 1.3212 | 2.5474 | 100.16 | 20.9 | 37.318 | 72.7326 | 80.984 | 68.6623 | 0.00922 | 1.0329 | 0.024652 | 23.8681 | 0.024115 | 0.1402 | 0 | -5.3e-05 | -5.3e-05 | -107.8485
2019-03-29 14:14 | 1553883231.5831 | 1244 | 0.20955 | 0.000455 | 1e-06 | 100.2463 | 100.037 | 125.2545 | 21.0972 | 0.002115 | 0.002006 | -152.1821 | 0.23576 | 8.4028 | 0.002483 | 0.003626 | 0.003425 | 6.1e-05 | 0.000968 | 0.31797 | 1.3163 | 2.5474 | 100.16 | 20.9 | 37.316 | 72.5403 | 81.3892 | 69.9239 | 0.009219 | 1.0291 | 0.024592 | 23.8976 | 0.018694 | 0.14025 | 2.1e-05 | 0.001492 | 0.001492 | 3.8253
2019-03-29 14:14 | 1553883232.5837 | 1245 | 0.20953 | 0.000451 | 1e-06 | 100.2035 | 99.9728 | 125.1658 | 21.0973 | 0.00209 | 0.001999 | -152.1931 | 0.23476 | 8.4044 | 0.002441 | 0.003877 | 0.003698 | 5.7e-05 | 0.000972 | 0.3027 | 1.3464 | 2.5475 | 100.16 | 20.9 | 37.316 | 71.9479 | 80.396 | 69.7649 | 0.009219 | 1.0295 | 0.024589 | 23.8836 | 0.020612 | 0.13966 | 0.000167 | 0.012308 | 0.012308 | 0.46181
2019-03-29 14:14 | 1553883233.5822 | 1246 | 0.20956 | 0.000455 | 0 | 100.2041 | 99.9479 | 124.9686 | 21.0977 | 0.002082 | 0.001958 | -152.1957 | 0.23425 | 8.4038 | 0.002591 | 0.003814 | 0.003602 | 6e-05 | 0.000979 | 0.28191 | 1.3296 | 2.5475 | 100.16 | 20.9 | 37.316 | 71.7488 | 83.9849 | 68.7865 | 0.009219 | 1.0325 | 0.024605 | 23.8307 | 0.022895 | 0.13936 | -1.4e-05 | -0.001056 | -0.001056 | -5.3706
2019-03-29 14:14 | 1553883234.5823 | 1247 | 0.20955 | 0.000451 | 0 | 100.2226 | 99.9609 | 125.1247 | 21.0968 | 0.00205 | 0.001995 | -152.2039 | 0.23433 | 8.4015 | 0.002476 | 0.004604 | 0.004499 | 6e-05 | 0.000976 | 0.31813 | 1.3311 | 2.5474 | 100.16 | 20.9 | 37.315 | 70.9785 | 81.2212 | 69.6771 | 0.009219 | 1.0298 | 0.024588 | 23.8766 | 0.023391 | 0.1394 | 1.2e-05 | 0.000869 | 0.000869 | 6.5301
2019-03-29 14:14 | 1553883235.5826 | 1248 | 0.20958 | 0.000456 | 0 | 100.1613 | 99.9912 | 124.9245 | 21.0968 | 0.002054 | 0.002007 | -152.1984 | 0.23491 | 8.4007 | 0.002518 | 0.003608 | 0.003367 | 6.2e-05 | 0.000962 | 0.27793 | 1.3321 | 2.5476 | 100.16 | 20.9 | 37.317 | 71.0821 | 82.2297 | 69.9593 | 0.00922 | 1.029 | 0.024559 | 23.8673 | 0.015195 | 0.13975 | -0.000133 | -0.009851 | -0.009851 | -0.57741
2019-03-29 14:14 | 1553883236.5823 | 1249 | 0.20959 | 0.000454 | 0 | 100.1663 | 99.8796 | 124.7221 | 21.0977 | 0.002044 | 0.001916 | -152.1886 | 0.23384 | 8.4012 | 0.002516 | 0.003664 | 0.003526 | 6.6e-05 | 0.000956 | 0.2959 | 1.3539 | 2.5473 | 100.16 | 20.9 | 37.315 | 70.8506 | 82.1968 | 67.7679 | 0.009219 | 1.0356 | 0.024617 | 23.7717 | 0.025638 | 0.13911 | -0.000204 | -0.015092 | -0.015092 | -0.37516
2019-03-29 14:14 | 1553883237.5824 | 1250 | 0.20958 | 0.000454 | 1e-06 | 100.2388 | 99.9093 | 124.9651 | 21.0968 | 0.002071 | 0.002017 | -152.1781 | 0.23501 | 8.4017 | 0.002481 | 0.00335 | 0.003127 | 6.5e-05 | 0.000957 | 0.323 | 1.3518 | 2.5474 | 100.16 | 20.9 | 37.315 | 71.4961 | 81.3422 | 70.1855 | 0.009219 | 1.0283 | 0.024554 | 23.8791 | 0.029446 | 0.13981 | -0.000144 | -0.010643 | -0.010643 | -0.53465
2019-03-29 14:14 | 1553883238.5823 | 1251 | 0.20956 | 0.000452 | 0 | 100.1107 | 100.0428 | 124.9085 | 21.0985 | 0.002106 | 0.001953 | -152.1832 | 0.2344 | 8.403 | 0.002503 | 0.00333 | 0.003137 | 6.5e-05 | 0.000951 | 0.31222 | 1.3148 | 2.5474 | 100.16 | 20.91 | 37.314 | 72.3265 | 81.8817 | 68.668 | 0.009225 | 1.0328 | 0.024603 | 23.8209 | 0.006068 | 0.13944 | -7e-06 | -0.000512 | -0.000512 | -11.0941
2019-03-29 14:14 | 1553883239.5823 | 1252 | 0.20954 | 0.000448 | 0 | 100.2487 | 99.9694 | 124.6829 | 21.0972 | 0.00204 | 0.002013 | -152.1794 | 0.23469 | 8.3991 | 0.002627 | 0.003654 | 0.003538 | 6.6e-05 | 0.00095 | 0.32947 | 1.3266 | 2.5473 | 100.16 | 20.9 | 37.312 | 70.755 | 84.8497 | 70.1047 | 0.009218 | 1.0285 | 0.02453 | 23.8493 | 0.024951 | 0.13962 | 8.2e-05 | 0.006044 | 0.006044 | 0.94015
2019-03-29 14:14 | 1553883240.5826 | 1253 | 0.20954 | 0.000455 | 0 | 100.2208 | 100.1073 | 124.5226 | 21.0978 | 0.002059 | 0.001961 | -152.1821 | 0.23494 | 8.4045 | 0.002533 | 0.004221 | 0.004037 | 6.2e-05 | 0.000953 | 0.29767 | 1.3726 | 2.5476 | 100.16 | 20.9 | 37.31 | 71.2033 | 82.6011 | 68.8624 | 0.009218 | 1.0323 | 0.024558 | 23.7908 | 0.010139 | 0.13977 | 9.1e-05 | 0.006666 | 0.006666 | 0.85343
2019-03-29 14:14 | 1553883241.5839 | 1254 | 0.20955 | 0.000455 | 1e-06 | 100.2175 | 99.8714 | 124.547 | 21.0982 | 0.002088 | 0.001976 | -152.1723 | 0.23441 | 8.4038 | 0.002569 | 0.004566 | 0.004342 | 6.4e-05 | 0.000952 | 0.31974 | 1.3867 | 2.5475 | 100.16 | 20.9 | 37.308 | 71.8997 | 83.4588 | 69.2238 | 0.009217 | 1.0312 | 0.024548 | 23.8057 | 0.030942 | 0.13945 | 5.5e-05 | 0.004039 | 0.004039 | 1.4053
2019-03-29 14:14 | 1553883242.5814 | 1255 | 0.20954 | 0.000454 | 0 | 100.2196 | 99.8776 | 124.286 | 21.0965 | 0.002068 | 0.002 | -152.1561 | 0.23421 | 8.4033 | 0.002569 | 0.003727 | 0.003592 | 6.2e-05 | 0.00096 | 0.30525 | 1.3826 | 2.5472 | 100.16 | 20.91 | 37.302 | 71.4302 | 83.461 | 69.7836 | 0.009222 | 1.0295 | 0.024502 | 23.8002 | 0.030578 | 0.13933 | 0.000109 | 0.008048 | 0.008048 | 0.70459
2019-03-29 14:14 | 1553883243.5815 | 1256 | 0.20956 | 0.000452 | 1e-06 | 100.2573 | 99.9363 | 124.598 | 21.0993 | 0.002073 | 0.001956 | -152.1472 | 0.23384 | 8.4035 | 0.002506 | 0.004267 | 0.004034 | 6.6e-05 | 0.000966 | 0.28047 | 1.3548 | 2.5474 | 100.16 | 20.91 | 37.298 | 71.5423 | 81.9513 | 68.728 | 0.009221 | 1.0327 | 0.024571 | 23.7933 | 0.028688 | 0.13911 | 0 | -9e-06 | -9e-06 | -615.4987
2019-03-29 14:14 | 1553883244.5826 | 1257 | 0.20954 | 0.000452 | 1e-06 | 100.2172 | 99.9489 | 124.2442 | 21.0987 | 0.002066 | 0.00196 | -152.1456 | 0.23552 | 8.4011 | 0.002492 | 0.004008 | 0.003884 | 6.2e-05 | 0.000965 | 0.31669 | 1.3214 | 2.5475 | 100.16 | 20.91 | 37.293 | 71.3858 | 81.6217 | 68.8212 | 0.009219 | 1.0324 | 0.024532 | 23.7628 | 0.023979 | 0.14011 | 0.000125 | 0.009209 | 0.009209 | 0.61926
2019-03-29 14:14 | 1553883245.5816 | 1258 | 0.20955 | 0.000453 | 1e-06 | 100.2223 | 100.008 | 124.3154 | 21.098 | 0.002065 | 0.001972 | -152.1486 | 0.23453 | 8.4022 | 0.002501 | 0.004088 | 0.003892 | 6.1e-05 | 0.000968 | 0.28528 | 1.3699 | 2.5474 | 100.16 | 20.91 | 37.285 | 71.3452 | 81.8284 | 69.1197 | 0.009217 | 1.0315 | 0.024529 | 23.7799 | 0.019151 | 0.13952 | 6.5e-05 | 0.004732 | 0.004732 | 1.2001
2019-03-29 14:14 | 1553883246.5825 | 1259 | 0.20955 | 0.00045 | 0 | 100.2319 | 99.9261 | 124.7095 | 21.0976 | 0.00203 | 0.001921 | -152.1468 | 0.23515 | 8.4007 | 0.002517 | 0.003679 | 0.00345 | 6.3e-05 | 0.000974 | 0.30055 | 1.3578 | 2.5475 | 100.16 | 20.91 | 37.281 | 70.516 | 82.2038 | 67.891 | 0.009216 | 1.0352 | 0.024612 | 23.7748 | 0.027335 | 0.13989 | 6.5e-05 | 0.004783 | 0.004783 | 1.1903
2019-03-29 14:14 | 1553883247.5816 | 1260 | 0.20955 | 0.000454 | 0 | 100.2262 | 100.0237 | 124.7578 | 21.0982 | 0.002034 | 0.001967 | -152.1571 | 0.23483 | 8.4022 | 0.002479 | 0.00289 | 0.002781 | 6.2e-05 | 0.000969 | 0.286 | 1.3658 | 2.5474 | 100.16 | 20.91 | 37.27 | 70.6165 | 81.3029 | 68.9858 | 0.009214 | 1.0319 | 0.024577 | 23.8175 | 0.018094 | 0.1397 | 3.9e-05 | 0.002864 | 0.002864 | 1.985
2019-03-29 14:14 | 1553883248.5824 | 1261 | 0.20957 | 0.000459 | 1e-06 | 100.2199 | 100.009 | 124.7257 | 21.098 | 0.002103 | 0.002014 | -152.1498 | 0.23497 | 8.4012 | 0.002486 | 0.003921 | 0.003739 | 6.3e-05 | 0.000969 | 0.31238 | 1.3513 | 2.5474 | 100.16 | 20.91 | 37.266 | 72.2761 | 81.4609 | 70.1251 | 0.009213 | 1.0285 | 0.024533 | 23.8541 | 0.018845 | 0.13978 | -6e-05 | -0.004489 | -0.004489 | -1.2673
2019-03-29 14:14 | 1553883249.5827 | 1262 | 0.20953 | 0.000454 | 0 | 100.1759 | 99.9209 | 124.8522 | 21.097 | 0.00205 | 0.001904 | -152.1668 | 0.23474 | 8.4026 | 0.002517 | 0.003495 | 0.003336 | 6e-05 | 0.000971 | 0.2912 | 1.3794 | 2.5474 | 100.16 | 20.91 | 37.263 | 70.9903 | 82.2094 | 67.4778 | 0.009212 | 1.0365 | 0.024641 | 23.774 | 0.022796 | 0.13965 | 0.000148 | 0.010949 | 0.010949 | 0.5191
2019-03-29 14:14 | 1553883250.5817 | 1263 | 0.20957 | 0.000453 | 0 | 100.2499 | 100.0104 | 124.927 | 21.0979 | 0.002041 | 0.001972 | -152.1762 | 0.23388 | 8.4013 | 0.002464 | 0.004001 | 0.003779 | 5.8e-05 | 0.000976 | 0.30215 | 1.3626 | 2.5472 | 100.16 | 20.91 | 37.254 | 70.7844 | 80.9322 | 69.1244 | 0.00921 | 1.0315 | 0.024589 | 23.8385 | 0.021395 | 0.13913 | -9.4e-05 | -0.006998 | -0.006998 | -0.80921
2019-03-29 14:14 | 1553883251.5817 | 1264 | 0.20952 | 0.000452 | 1e-06 | 100.2989 | 99.8182 | 125.0613 | 21.0984 | 0.00205 | 0.001942 | -152.1816 | 0.23474 | 8.4029 | 0.00245 | 0.003574 | 0.00341 | 6.3e-05 | 0.000974 | 0.28926 | 1.3249 | 2.5474 | 100.16 | 20.91 | 37.25 | 70.9916 | 80.5982 | 68.3918 | 0.009209 | 1.0337 | 0.024628 | 23.8258 | 0.042975 | 0.13964 | 0.000195 | 0.014376 | 0.014376 | 0.39534
2019-03-29 14:14 | 1553883252.5831 | 1265 | 0.20957 | 0.000453 | 0 | 100.2529 | 100.0807 | 125.3397 | 21.0977 | 0.00203 | 0.001923 | -152.1943 | 0.23436 | 8.4014 | 0.002446 | 0.003525 | 0.003348 | 6.2e-05 | 0.000974 | 0.30613 | 1.3604 | 2.5473 | 100.16 | 20.91 | 37.246 | 70.5005 | 80.5016 | 67.9285 | 0.009208 | 1.0351 | 0.024672 | 23.8361 | 0.015376 | 0.13942 | -5.5e-05 | -0.004089 | -0.004089 | -1.3878
2019-03-29 14:14 | 1553883253.5832 | 1266 | 0.20958 | 0.00045 | 0 | 100.2268 | 100.0516 | 125.0482 | 21.0981 | 0.002011 | 0.001935 | -152.2087 | 0.23442 | 8.4021 | 0.002434 | 0.003843 | 0.003685 | 6.3e-05 | 0.00098 | 0.2896 | 1.3629 | 2.5474 | 100.16 | 20.91 | 37.239 | 70.0565 | 80.212 | 68.2371 | 0.009206 | 1.0341 | 0.024633 | 23.8191 | 0.015645 | 0.13946 | -0.000115 | -0.008533 | -0.008533 | -0.6652
2019-03-29 14:14 | 1553883254.5824 | 1267 | 0.20954 | 0.000456 | 0 | 100.3549 | 99.9493 | 125.4967 | 21.0976 | 0.00206 | 0.001928 | -152.2131 | 0.23498 | 8.4033 | 0.002473 | 0.003556 | 0.003382 | 6.1e-05 | 0.000974 | 0.27079 | 1.3759 | 2.5474 | 100.16 | 20.91 | 37.235 | 71.2333 | 81.1588 | 68.0479 | 0.009205 | 1.0347 | 0.024684 | 23.8552 | 0.036229 | 0.13979 | 0.000101 | 0.007462 | 0.007462 | 0.76246
2019-03-29 14:14 | 1553883255.5819 | 1268 | 0.20958 | 0.000455 | 1e-06 | 100.2654 | 99.8885 | 125.5596 | 21.0977 | 0.001998 | 0.001962 | -152.2017 | 0.23435 | 8.4032 | 0.002512 | 0.004029 | 0.003933 | 6.3e-05 | 0.000977 | 0.32764 | 1.3877 | 2.5474 | 100.16 | 20.91 | 37.231 | 69.7511 | 82.1017 | 68.8647 | 0.009204 | 1.0323 | 0.02466 | 23.8897 | 0.033687 | 0.13941 | -0.000144 | -0.010739 | -0.010739 | -0.52835
2019-03-29 14:14 | 1553883256.582 | 1269 | 0.20958 | 0.000454 | 1e-06 | 100.2478 | 99.9773 | 125.3376 | 21.0976 | 0.002018 | 0.001968 | -152.1991 | 0.23398 | 8.4047 | 0.002455 | 0.003827 | 0.003667 | 5.9e-05 | 0.000981 | 0.30829 | 1.3667 | 2.5475 | 100.16 | 20.91 | 37.228 | 70.2237 | 80.7111 | 69.0102 | 0.009203 | 1.0318 | 0.024633 | 23.8737 | 0.024169 | 0.1392 | -0.000118 | -0.008771 | -0.008771 | -0.6459
2019-03-29 14:14 | 1553883257.581 | 1270 | 0.20955 | 0.000453 | 0 | 100.2552 | 100.0759 | 125.0139 | 21.0981 | 0.001976 | 0.001905 | -152.2021 | 0.2348 | 8.4032 | 0.002406 | 0.003977 | 0.003755 | 5.9e-05 | 0.000983 | 0.28174 | 1.3762 | 2.5476 | 100.16 | 20.91 | 37.226 | 69.2198 | 79.5391 | 67.4952 | 0.009203 | 1.0364 | 0.024656 | 23.79 | 0.016017 | 0.13968 | 2.8e-05 | 0.002038 | 0.002038 | 2.7889
2019-03-29 14:14 | 1553883258.5824 | 1271 | 0.20957 | 0.000455 | 0 | 100.2484 | 100.0875 | 124.8681 | 21.0973 | 0.001969 | 0.001932 | -152.2166 | 0.23446 | 8.4031 | 0.002425 | 0.004342 | 0.004091 | 5.9e-05 | 0.000988 | 0.31868 | 1.3401 | 2.5474 | 100.16 | 20.91 | 37.219 | 69.0498 | 80.0107 | 68.1543 | 0.009201 | 1.0344 | 0.024618 | 23.7991 | 0.014365 | 0.13948 | -9e-05 | -0.006677 | -0.006677 | -0.85017
2019-03-29 14:14 | 1553883259.581 | 1272 | 0.20956 | 0.000454 | 1e-06 | 100.2041 | 99.9598 | 124.689 | 21.0975 | 0.002009 | 0.001953 | -152.2064 | 0.23437 | 8.4013 | 0.002479 | 0.003665 | 0.003544 | 5.9e-05 | 0.000984 | 0.28639 | 1.337 | 2.5473 | 100.16 | 20.91 | 37.219 | 69.995 | 81.288 | 68.6598 | 0.009201 | 1.0329 | 0.024582 | 23.7996 | 0.021826 | 0.13943 | -4.3e-05 | -0.003194 | -0.003194 | -1.7768
2019-03-29 14:14 | 1553883260.581 | 1273 | 0.20955 | 0.000453 | 1e-06 | 100.2699 | 99.953 | 125.146 | 21.0975 | 0.002031 | 0.001976 | -152.1991 | 0.23401 | 8.3989 | 0.002479 | 0.003337 | 0.003165 | 5.8e-05 | 0.000988 | 0.30127 | 1.3161 | 2.5474 | 100.16 | 20.92 | 37.215 | 70.5259 | 81.2998 | 69.2206 | 0.009206 | 1.0312 | 0.024607 | 23.8628 | 0.028313 | 0.13922 | 3e-05 | 0.002176 | 0.002176 | 2.604
2019-03-29 14:14 | 1553883261.5811 | 1274 | 0.20957 | 0.000452 | 0 | 100.257 | 99.8332 | 124.9059 | 21.0976 | 0.001986 | 0.001923 | -152.2053 | 0.23568 | 8.4017 | 0.002453 | 0.00437 | 0.004215 | 5.9e-05 | 0.000988 | 0.29662 | 1.3264 | 2.5472 | 100.16 | 20.92 | 37.212 | 69.4595 | 80.67 | 67.9426 | 0.009205 | 1.035 | 0.024629 | 23.7953 | 0.037891 | 0.14021 | -0.0001 | -0.007437 | -0.007437 | -0.7673
2019-03-29 14:14 | 1553883262.5822 | 1275 | 0.20954 | 0.000449 | 0 | 100.263 | 100.0489 | 125.0181 | 21.0986 | 0.002001 | 0.00193 | -152.1972 | 0.23445 | 8.4028 | 0.002543 | 0.003773 | 0.003562 | 6.1e-05 | 0.000986 | 0.31133 | 1.3357 | 2.5474 | 100.16 | 20.91 | 37.209 | 69.8183 | 82.835 | 68.1143 | 0.009199 | 1.0345 | 0.024634 | 23.812 | 0.019121 | 0.13947 | 9e-05 | 0.006657 | 0.006657 | 0.85272
2019-03-29 14:14 | 1553883263.5812 | 1276 | 0.20957 | 0.00045 | 0 | 100.2675 | 100.04 | 125.1275 | 21.0981 | 0.002006 | 0.001899 | -152.1906 | 0.23487 | 8.4026 | 0.002403 | 0.004654 | 0.004493 | 5.9e-05 | 0.000986 | 0.30326 | 1.351 | 2.5474 | 100.16 | 20.91 | 37.205 | 69.9404 | 79.4672 | 67.3555 | 0.009198 | 1.0368 | 0.024672 | 23.7959 | 0.020315 | 0.13973 | -8.3e-05 | -0.006177 | -0.006177 | -0.92067
2019-03-29 14:14 | 1553883264.5816 | 1277 | 0.20948 | 0.000449 | 0 | 100.3181 | 100.1192 | 124.6984 | 21.0982 | 0.002007 | 0.001975 | -152.183 | 0.2349 | 8.4034 | 0.002493 | 0.003533 | 0.003348 | 5.6e-05 | 0.000984 | 0.28191 | 1.3471 | 2.5475 | 100.16 | 20.91 | 37.202 | 69.96 | 81.6272 | 69.1964 | 0.009197 | 1.0313 | 0.024564 | 23.8192 | 0.01775 | 0.13974 | 0.000481 | 0.035526 | 0.035526 | 0.16009
2019-03-29 14:14 | 1553883265.5816 | 1278 | 0.20957 | 0.000455 | 0 | 100.313 | 100.0005 | 124.6625 | 21.0978 | 0.00204 | 0.001971 | -152.1759 | 0.23446 | 8.4024 | 0.002452 | 0.004256 | 0.004021 | 5.8e-05 | 0.000983 | 0.29878 | 1.3373 | 2.5473 | 100.16 | 20.92 | 37.197 | 70.7423 | 80.6553 | 69.0986 | 0.009201 | 1.0315 | 0.024564 | 23.8124 | 0.027913 | 0.13948 | -6.4e-05 | -0.004757 | -0.004757 | -1.1933
2019-03-29 14:14 | 1553883266.5822 | 1279 | 0.20958 | 0.000452 | 0 | 100.2481 | 100.023 | 124.7738 | 21.0976 | 0.002064 | 0.001976 | -152.1789 | 0.23402 | 8.4029 | 0.002487 | 0.003249 | 0.003065 | 5.8e-05 | 0.000987 | 0.27848 | 1.3568 | 2.5476 | 100.16 | 20.92 | 37.196 | 71.3222 | 81.4816 | 69.2044 | 0.009201 | 1.0312 | 0.024571 | 23.8267 | 0.020105 | 0.13922 | -0.000117 | -0.008631 | -0.008631 | -0.65652
2019-03-29 14:14 | 1553883267.5822 | 1280 | 0.20973 | 0.00045 | 0 | 100.2361 | 100.096 | 124.8677 | 21.0988 | 0.001982 | 0.0019 | -152.1799 | 0.23476 | 8.4031 | 0.002468 | 0.004061 | 0.003867 | 6e-05 | 0.000988 | 0.29286 | 1.3493 | 2.5474 | 100.16 | 20.92 | 37.194 | 69.3591 | 81.0266 | 67.3852 | 0.009201 | 1.0367 | 0.024646 | 23.7722 | 0.012508 | 0.13966 | -0.001029 | -0.076203 | -0.076203 | -0.07459
2019-03-29 14:14 | 1553883268.5817 | 1281 | 0.20957 | 0.000453 | 1e-06 | 100.2576 | 99.9827 | 125.0989 | 21.0976 | 0.002004 | 0.001937 | -152.1555 | 0.23547 | 8.4038 | 0.002496 | 0.003827 | 0.003686 | 5.3e-05 | 0.000985 | 0.29734 | 1.3284 | 2.5473 | 100.16 | 20.92 | 37.193 | 69.8798 | 81.7009 | 68.2756 | 0.0092 | 1.034 | 0.024636 | 23.8253 | 0.024562 | 0.14008 | -9.7e-05 | -0.007198 | -0.007198 | -0.79203
2019-03-29 14:14 | 1553883269.5824 | 1282 | 0.20952 | 0.00045 | 0 | 100.3169 | 99.9315 | 125.3662 | 21.098 | 0.002022 | 0.001937 | -152.1532 | 0.23513 | 8.4019 | 0.002511 | 0.00391 | 0.003698 | 6.1e-05 | 0.000983 | 0.30558 | 1.3432 | 2.5474 | 100.16 | 20.92 | 37.194 | 70.3245 | 82.0786 | 68.2704 | 0.009201 | 1.034 | 0.024663 | 23.8506 | 0.034429 | 0.13988 | 0.000215 | 0.015938 | 0.015938 | 0.35722
2019-03-29 14:14 | 1553883270.5814 | 1283 | 0.20957 | 0.000452 | 0 | 100.2091 | 100.0404 | 124.8127 | 21.0993 | 0.002071 | 0.001985 | -152.1529 | 0.23503 | 8.4009 | 0.002463 | 0.003717 | 0.003598 | 5.9e-05 | 0.000987 | 0.31006 | 1.325 | 2.5473 | 100.16 | 20.92 | 37.194 | 71.5024 | 80.9187 | 69.419 | 0.009201 | 1.0306 | 0.024567 | 23.8379 | 0.015076 | 0.13982 | -7.6e-05 | -0.005629 | -0.005629 | -1.011
2019-03-29 14:14 | 1553883271.5814 | 1284 | 0.20949 | 0.000455 | 1e-06 | 100.2178 | 100.0083 | 124.9735 | 21.0984 | 0.00205 | 0.001943 | -152.1622 | 0.23446 | 8.4007 | 0.002429 | 0.004181 | 0.003978 | 5.8e-05 | 0.000991 | 0.25857 | 1.3359 | 2.5473 | 100.16 | 20.91 | 37.193 | 70.9801 | 80.107 | 68.4262 | 0.009195 | 1.0336 | 0.024618 | 23.8186 | 0.01872 | 0.13948 | 0.000398 | 0.029405 | 0.029405 | 0.19306
2019-03-29 14:15 | 1553883272.5805 | 1285 | 0.20958 | 0.000454 | 1e-06 | 100.2379 | 99.9561 | 124.882 | 21.0995 | 0.001964 | 0.001939 | -152.1661 | 0.23496 | 8.4017 | 0.002487 | 0.004243 | 0.004088 | 5.6e-05 | 0.000991 | 0.274 | 1.335 | 2.5475 | 100.16 | 20.91 | 37.193 | 68.9272 | 81.4852 | 68.3274 | 0.009195 | 1.0339 | 0.024613 | 23.8065 | 0.025181 | 0.13978 | -0.000156 | -0.01156 | -0.01156 | -0.49213
2019-03-29 14:15 | 1553883273.5805 | 1286 | 0.20958 | 0.000455 | 0 | 100.2549 | 99.9312 | 124.8705 | 21.0996 | 0.00204 | 0.001913 | -152.1813 | 0.23474 | 8.4015 | 0.002513 | 0.003995 | 0.0039 | 5.9e-05 | 0.000989 | 0.31725 | 1.3483 | 2.5475 | 100.16 | 20.91 | 37.196 | 70.7541 | 82.1247 | 67.6933 | 0.009195 | 1.0358 | 0.024635 | 23.7832 | 0.028933 | 0.13965 | -0.000149 | -0.011032 | -0.011032 | -0.51519
2019-03-29 14:15 | 1553883274.5806 | 1287 | 0.20953 | 0.000449 | 1e-06 | 100.2939 | 100.0595 | 125.0566 | 21.0988 | 0.002041 | 0.00197 | -152.1919 | 0.2349 | 8.403 | 0.002539 | 0.00375 | 0.00354 | 5.8e-05 | 0.000989 | 0.28512 | 1.3049 | 2.5474 | 100.16 | 20.91 | 37.196 | 70.7863 | 82.7431 | 69.0683 | 0.009195 | 1.0316 | 0.024603 | 23.8489 | 0.020925 | 0.13974 | 0.000158 | 0.011649 | 0.011649 | 0.48823
2019-03-29 14:15 | 1553883275.5815 | 1288 | 0.20956 | 0.000453 | 1e-06 | 100.2154 | 100.0585 | 124.905 | 21.098 | 0.002034 | 0.001921 | -152.2207 | 0.2342 | 8.4029 | 0.002505 | 0.003892 | 0.003761 | 5.8e-05 | 0.000984 | 0.28766 | 1.3236 | 2.5473 | 100.16 | 20.91 | 37.199 | 70.5998 | 81.9289 | 67.8882 | 0.009196 | 1.0352 | 0.024631 | 23.7933 | 0.01402 | 0.13933 | -3.7e-05 | -0.00274 | -0.00274 | -2.0694
2019-03-29 14:15 | 1553883276.5815 | 1289 | 0.20953 | 0.000456 | 0 | 100.2406 | 99.936 | 124.9033 | 21.0988 | 0.002003 | 0.001955 | -152.2353 | 0.23448 | 8.4012 | 0.002546 | 0.00415 | 0.004032 | 5.9e-05 | 0.000987 | 0.31598 | 1.3277 | 2.5474 | 100.16 | 20.91 | 37.199 | 69.8653 | 82.9109 | 68.7181 | 0.009196 | 1.0327 | 0.024601 | 23.8221 | 0.027223 | 0.13949 | 0.000161 | 0.011854 | 0.011854 | 0.47894
2019-03-29 14:15 | 1553883277.5817 | 1290 | 0.20957 | 0.000454 | 0 | 100.1885 | 99.9875 | 125.0111 | 21.0994 | 0.002046 | 0.001903 | -152.2397 | 0.23464 | 8.4026 | 0.002483 | 0.003799 | 0.003665 | 5.8e-05 | 0.000986 | 0.32388 | 1.3221 | 2.5474 | 100.16 | 20.91 | 37.2 | 70.8897 | 81.4062 | 67.4582 | 0.009196 | 1.0365 | 0.024657 | 23.7884 | 0.017963 | 0.13959 | -6.1e-05 | -0.004537 | -0.004537 | -1.2522
2019-03-29 14:15 | 1553883278.5816 | 1291 | 0.20954 | 0.000452 | 0 | 100.1807 | 99.9793 | 125.1585 | 21.0989 | 0.001984 | 0.001932 | -152.2489 | 0.23588 | 8.4025 | 0.00244 | 0.003848 | 0.003721 | 6e-05 | 0.000986 | 0.30773 | 1.3343 | 2.5476 | 100.16 | 20.91 | 37.198 | 69.4138 | 80.3614 | 68.1458 | 0.009196 | 1.0344 | 0.024647 | 23.8265 | 0.018001 | 0.14033 | 0.000108 | 0.008002 | 0.008002 | 0.71371
2019-03-29 14:15 | 1553883279.5807 | 1292 | 0.20955 | 0.000452 | 0 | 100.2029 | 99.9994 | 125.0829 | 21.0992 | 0.002062 | 0.001941 | -152.2601 | 0.23468 | 8.404 | 0.00255 | 0.00412 | 0.003865 | 5.8e-05 | 0.000986 | 0.29806 | 1.3308 | 2.5474 | 100.16 | 20.91 | 37.196 | 71.2692 | 82.9971 | 68.38 | 0.009195 | 1.0337 | 0.024631 | 23.8274 | 0.018177 | 0.13961 | 7.2e-05 | 0.005323 | 0.005323 | 1.0675
2019-03-29 14:15 | 1553883280.5819 | 1293 | 0.20954 | 0.000452 | 1e-06 | 100.1741 | 99.9185 | 125.5622 | 21.0993 | 0.002038 | 0.001946 | -152.2746 | 0.23502 | 8.4029 | 0.002529 | 0.003508 | 0.003309 | 5.9e-05 | 0.000987 | 0.31421 | 1.3935 | 2.5477 | 100.16 | 20.91 | 37.194 | 70.7111 | 82.5081 | 68.4924 | 0.009195 | 1.0334 | 0.024674 | 23.877 | 0.02285 | 0.13981 | 0.000105 | 0.007784 | 0.007784 | 0.731
2019-03-29 14:15 | 1553883281.5807 | 1294 | 0.20955 | 0.000453 | 1e-06 | 100.1463 | 99.8523 | 125.4693 | 21.0998 | 0.001952 | 0.001865 | -152.2558 | 0.23541 | 8.4001 | 0.002426 | 0.004424 | 0.004222 | 6e-05 | 0.000984 | 0.30829 | 1.3592 | 2.5476 | 100.16 | 20.91 | 37.189 | 68.6311 | 80.0381 | 66.5465 | 0.009194 | 1.0393 | 0.024735 | 23.8001 | 0.026293 | 0.14004 | 3.5e-05 | 0.002544 | 0.002544 | 2.2401
2019-03-29 14:15 | 1553883282.5807 | 1295 | 0.20954 | 0.000453 | 1e-06 | 100.2166 | 99.9697 | 125.6591 | 21.0988 | 0.001952 | 0.001884 | -152.2626 | 0.23499 | 8.4009 | 0.002506 | 0.003119 | 0.003066 | 6e-05 | 0.000987 | 0.31902 | 1.3544 | 2.5474 | 100.16 | 20.91 | 37.187 | 68.6441 | 81.953 | 66.9974 | 0.009193 | 1.0379 | 0.024738 | 23.8339 | 0.022063 | 0.1398 | 0.000122 | 0.008994 | 0.008994 | 0.63259
2019-03-29 14:15 | 1553883283.5809 | 1296 | 0.20952 | 0.000452 | 0 | 100.1897 | 99.8404 | 125.9188 | 21.0984 | 0.002024 | 0.001942 | -152.247 | 0.23501 | 8.4019 | 0.002498 | 0.003538 | 0.003325 | 5.8e-05 | 0.000977 | 0.27599 | 1.3263 | 2.5474 | 100.16 | 20.91 | 37.18 | 70.3742 | 81.7638 | 68.4073 | 0.009191 | 1.0336 | 0.024712 | 23.9079 | 0.031239 | 0.13981 | 0.000257 | 0.019027 | 0.019027 | 0.29906
2019-03-29 14:15 | 1553883284.5812 | 1297 | 0.20953 | 0.000452 | 1e-06 | 100.1193 | 100.0472 | 126.0204 | 21.0992 | 0.002017 | 0.001924 | -152.2372 | 0.23518 | 8.4008 | 0.002497 | 0.004334 | 0.004168 | 5.9e-05 | 0.000971 | 0.28008 | 1.3078 | 2.5472 | 100.16 | 20.91 | 37.177 | 70.1941 | 81.7282 | 67.9699 | 0.009191 | 1.035 | 0.024738 | 23.9022 | 0.006447 | 0.13991 | 0.000173 | 0.01284 | 0.01284 | 0.44346
2019-03-29 14:15 | 1553883285.5823 | 1298 | 0.20953 | 0.000455 | 1e-06 | 100.1777 | 99.8988 | 126.161 | 21.0997 | 0.002036 | 0.001901 | -152.2197 | 0.23432 | 8.4019 | 0.002545 | 0.003694 | 0.003472 | 5.7e-05 | 0.000972 | 0.30326 | 1.328 | 2.5473 | 100.16 | 20.92 | 37.173 | 70.6624 | 82.8906 | 67.4195 | 0.009195 | 1.0366 | 0.024772 | 23.8962 | 0.024943 | 0.1394 | 0.000146 | 0.010858 | 0.010858 | 0.5225
2019-03-29 14:15 | 1553883286.5819 | 1299 | 0.20951 | 0.000451 | 0 | 100.1936 | 99.8175 | 126.0272 | 21.0996 | 0.00201 | 0.001921 | -152.214 | 0.23488 | 8.4021 | 0.002509 | 0.003624 | 0.003534 | 5.9e-05 | 0.000972 | 0.31365 | 1.3164 | 2.5472 | 100.16 | 20.91 | 37.165 | 70.0322 | 82.0186 | 67.8907 | 0.009188 | 1.0352 | 0.024741 | 23.9001 | 0.033635 | 0.13973 | 0.000304 | 0.022604 | 0.022604 | 0.25159
2019-03-29 14:15 | 1553883287.582 | 1300 | 0.20955 | 0.000454 | 0 | 100.2091 | 99.9728 | 126.0103 | 21.0997 | 0.001989 | 0.00189 | -152.2069 | 0.23496 | 8.4019 | 0.002562 | 0.00364 | 0.003385 | 5.6e-05 | 0.000973 | 0.32211 | 1.3307 | 2.5474 | 100.16 | 20.91 | 37.162 | 69.5278 | 83.2931 | 67.1348 | 0.009187 | 1.0375 | 0.024767 | 23.872 | 0.02112 | 0.13978 | 1.6e-05 | 0.001167 | 0.001167 | 4.8769
2019-03-29 14:15 | 1553883288.5813 | 1301 | 0.20959 | 0.000454 | 0 | 100.2184 | 99.9056 | 125.7927 | 21.1 | 0.001931 | 0.001928 | -152.2191 | 0.2347 | 8.401 | 0.002563 | 0.004132 | 0.004017 | 5.8e-05 | 0.000969 | 0.31581 | 1.3341 | 2.5474 | 100.16 | 20.91 | 37.158 | 68.1406 | 83.323 | 68.0486 | 0.009186 | 1.0347 | 0.024713 | 23.8833 | 0.027966 | 0.13962 | -0.000194 | -0.014447 | -0.014447 | -0.39336
2019-03-29 14:15 | 1553883289.581 | 1302 | 0.20955 | 0.000451 | 0 | 100.1981 | 99.7455 | 125.9096 | 21.0996 | 0.001995 | 0.001932 | -152.2143 | 0.23536 | 8.4016 | 0.002527 | 0.003545 | 0.003439 | 5.7e-05 | 0.000971 | 0.30342 | 1.3383 | 2.5475 | 100.16 | 20.91 | 37.151 | 69.6594 | 82.4617 | 68.1628 | 0.009184 | 1.0344 | 0.02472 | 23.8984 | 0.040491 | 0.14002 | 5.1e-05 | 0.003805 | 0.003805 | 1.4976
2019-03-29 14:15 | 1553883290.581 | 1303 | 0.20957 | 0.000452 | 1e-06 | 100.2214 | 99.9834 | 125.8281 | 21.0995 | 0.002014 | 0.001921 | -152.1974 | 0.23448 | 8.4032 | 0.002556 | 0.003313 | 0.003171 | 5.9e-05 | 0.000971 | 0.32228 | 1.3454 | 2.5474 | 100.16 | 20.91 | 37.147 | 70.1342 | 83.1463 | 67.8913 | 0.009183 | 1.0352 | 0.024722 | 23.8812 | 0.021269 | 0.13949 | -4.3e-05 | -0.00321 | -0.00321 | -1.7684
2019-03-29 14:15 | 1553883291.58 | 1304 | 0.20957 | 0.000451 | 1e-06 | 100.1466 | 99.8871 | 126 | 21.1001 | 0.002006 | 0.001904 | -152.215 | 0.23394 | 8.4038 | 0.002591 | 0.00416 | 0.00396 | 5.9e-05 | 0.000968 | 0.31686 | 1.3189 | 2.5475 | 100.16 | 20.91 | 37.139 | 69.925 | 83.9848 | 67.4824 | 0.009181 | 1.0364 | 0.024753 | 23.8832 | 0.023202 | 0.13917 | -6.9e-05 | -0.00518 | -0.00518 | -1.0935
2019-03-29 14:15 | 1553883292.5811 | 1305 | 0.20956 | 0.000451 | 0 | 100.1619 | 99.9237 | 125.6981 | 21.0997 | 0.002005 | 0.001875 | -152.2093 | 0.23459 | 8.403 | 0.002557 | 0.003567 | 0.003406 | 6e-05 | 0.000962 | 0.30381 | 1.3646 | 2.5474 | 100.16 | 20.91 | 37.134 | 69.9177 | 83.1684 | 66.7891 | 0.00918 | 1.0386 | 0.024749 | 23.8303 | 0.021296 | 0.13956 | 1.8e-05 | 0.001323 | 0.001323 | 4.2938
2019-03-29 14:15 | 1553883293.5812 | 1306 | 0.20957 | 0.000452 | 0 | 100.0882 | 100.0131 | 125.242 | 21.1 | 0.002027 | 0.001896 | -152.2218 | 0.23556 | 8.402 | 0.002463 | 0.004483 | 0.004293 | 6e-05 | 0.000952 | 0.31277 | 1.3722 | 2.5473 | 100.16 | 20.91 | 37.13 | 70.4356 | 80.9124 | 67.2996 | 0.009179 | 1.037 | 0.024686 | 23.8048 | 0.006717 | 0.14013 | -9.4e-05 | -0.006979 | -0.006979 | -0.81718
2019-03-29 14:15 | 1553883294.5813 | 1307 | 0.2096 | 0.000454 | 1e-06 | 100.161 | 100.0312 | 125.2362 | 21.0997 | 0.001966 | 0.001906 | -152.2383 | 0.23507 | 8.4043 | 0.002552 | 0.004043 | 0.003886 | 5.9e-05 | 0.000948 | 0.31708 | 1.3616 | 2.5475 | 100.16 | 20.91 | 37.119 | 68.9771 | 83.0474 | 67.5318 | 0.009176 | 1.0363 | 0.024677 | 23.8124 | 0.011597 | 0.13984 | -0.000226 | -0.016817 | -0.016817 | -0.33843
2019-03-29 14:15 | 1553883295.5805 | 1308 | 0.20957 | 0.000451 | 1e-06 | 100.1092 | 99.7752 | 125.5171 | 21.0994 | 0.00208 | 0.001931 | -152.2512 | 0.23549 | 8.4001 | 0.002583 | 0.003461 | 0.003288 | 6.1e-05 | 0.000951 | 0.3413 | 1.3614 | 2.5475 | 100.16 | 20.9 | 37.115 | 71.7116 | 83.7961 | 68.1213 | 0.00917 | 1.0345 | 0.024683 | 23.8597 | 0.029889 | 0.14009 | -5.8e-05 | -0.004373 | -0.004373 | -1.3039
2019-03-29 14:15 | 1553883296.5802 | 1309 | 0.20957 | 0.000453 | 1e-06 | 100.1155 | 99.9482 | 125.8845 | 21.0996 | 0.002012 | 0.001933 | -152.2524 | 0.23452 | 8.4013 | 0.002583 | 0.003838 | 0.00365 | 6.2e-05 | 0.000953 | 0.31454 | 1.3473 | 2.5475 | 100.16 | 20.9 | 37.106 | 70.0781 | 83.8064 | 68.1784 | 0.009167 | 1.0343 | 0.024717 | 23.8966 | 0.014952 | 0.13952 | -3.8e-05 | -0.002848 | -0.002848 | -1.9936
2019-03-29 14:15 | 1553883297.5802 | 1310 | 0.20956 | 0.000447 | 1e-06 | 100.1113 | 99.768 | 125.9899 | 21.1005 | 0.001955 | 0.001906 | -152.261 | 0.23462 | 8.4015 | 0.002601 | 0.003764 | 0.003593 | 6.1e-05 | 0.000959 | 0.29983 | 1.3575 | 2.5474 | 100.16 | 20.9 | 37.102 | 68.7114 | 84.2365 | 67.5233 | 0.009166 | 1.0363 | 0.024751 | 23.8836 | 0.030719 | 0.13958 | 2.3e-05 | 0.001677 | 0.001677 | 3.3875
2019-03-29 14:15 | 1553883298.5817 | 1311 | 0.20959 | 0.000447 | 0 | 100.1696 | 99.7281 | 125.9587 | 21.1001 | 0.002015 | 0.001926 | -152.2773 | 0.2348 | 8.4019 | 0.002529 | 0.003665 | 0.003456 | 6.2e-05 | 0.000965 | 0.30807 | 1.3427 | 2.5473 | 100.16 | 20.9 | 37.097 | 70.1536 | 82.4956 | 68.0198 | 0.009165 | 1.0348 | 0.02473 | 23.8981 | 0.039513 | 0.13968 | -0.000194 | -0.014426 | -0.014426 | -0.39407
2019-03-29 14:15 | 1553883299.5803 | 1312 | 0.20956 | 0.000451 | 0 | 100.0966 | 99.8568 | 125.6868 | 21.0998 | 0.001968 | 0.001925 | -152.2672 | 0.235 | 8.4018 | 0.002594 | 0.003894 | 0.003788 | 6.3e-05 | 0.000962 | 0.31045 | 1.382 | 2.5475 | 100.16 | 20.9 | 37.088 | 69.0178 | 84.0586 | 67.9913 | 0.009163 | 1.0349 | 0.024704 | 23.8713 | 0.021456 | 0.1398 | -2.8e-05 | -0.002081 | -0.002081 | -2.7349
2019-03-29 14:15 | 1553883300.5814 | 1313 | 0.20957 | 0.00045 | 1e-06 | 100.1253 | 99.7595 | 125.4146 | 21.1002 | 0.001996 | 0.001918 | -152.2671 | 0.23392 | 8.4029 | 0.002532 | 0.003913 | 0.003712 | 6.1e-05 | 0.000965 | 0.30646 | 1.4125 | 2.5476 | 100.16 | 20.9 | 37.083 | 69.6843 | 82.5666 | 67.8085 | 0.009162 | 1.0354 | 0.024684 | 23.839 | 0.032741 | 0.13916 | -7.6e-05 | -0.005652 | -0.005652 | -1.002
2019-03-29 14:15 | 1553883301.5814 | 1314 | 0.20957 | 0.000451 | 1e-06 | 100.0245 | 99.8578 | 125.4772 | 21.1002 | 0.001978 | 0.001984 | -152.2602 | 0.23472 | 8.4027 | 0.002551 | 0.003499 | 0.003319 | 6.1e-05 | 0.000974 | 0.30038 | 1.3232 | 2.5474 | 100.16 | 20.9 | 37.078 | 69.2562 | 83.0378 | 69.4113 | 0.009161 | 1.0306 | 0.024632 | 23.901 | 0.014916 | 0.13964 | -8.4e-05 | -0.006236 | -0.006236 | -0.91137
2019-03-29 14:15 | 1553883302.5805 | 1315 | 0.20958 | 0.000451 | 0 | 100.1589 | 99.9871 | 125.8453 | 21.1003 | 0.001963 | 0.001833 | -152.2634 | 0.23582 | 8.4016 | 0.002516 | 0.004105 | 0.003851 | 6.2e-05 | 0.000977 | 0.29535 | 1.3647 | 2.5474 | 100.16 | 20.91 | 37.065 | 68.9061 | 82.1962 | 65.7706 | 0.009163 | 1.0417 | 0.024801 | 23.8085 | 0.015348 | 0.14029 | -0.000108 | -0.008034 | -0.008034 | -0.7107
2019-03-29 14:15 | 1553883303.5814 | 1316 | 0.20955 | 0.000447 | 0 | 100.1406 | 99.9025 | 126.1037 | 21.1003 | 0.001952 | 0.001901 | -152.2562 | 0.23496 | 8.4034 | 0.00254 | 0.003932 | 0.003771 | 6.3e-05 | 0.000983 | 0.32803 | 1.3537 | 2.5474 | 100.16 | 20.91 | 37.058 | 68.6386 | 82.7617 | 67.4093 | 0.009161 | 1.0367 | 0.024766 | 23.8904 | 0.021292 | 0.13978 | 4.2e-05 | 0.003151 | 0.003151 | 1.8053
2019-03-29 14:15 | 1553883304.5808 | 1317 | 0.20957 | 0.000446 | 1e-06 | 100.1307 | 99.9455 | 126.0565 | 21.1003 | 0.002005 | 0.00193 | -152.2351 | 0.2348 | 8.4025 | 0.002504 | 0.004263 | 0.004016 | 6.3e-05 | 0.000987 | 0.29407 | 1.3566 | 2.5474 | 100.16 | 20.91 | 37.042 | 69.9043 | 81.9116 | 68.1058 | 0.009157 | 1.0345 | 0.024736 | 23.9104 | 0.01656 | 0.13968 | -5.5e-05 | -0.00412 | -0.00412 | -1.38
2019-03-29 14:15 | 1553883305.5805 | 1318 | 0.20958 | 0.000448 | 0 | 100.1831 | 99.9745 | 125.677 | 21.1 | 0.002009 | 0.001912 | -152.2195 | 0.23607 | 8.4026 | 0.002492 | 0.003817 | 0.003672 | 6.2e-05 | 0.000985 | 0.31421 | 1.3286 | 2.5474 | 100.16 | 20.91 | 37.033 | 70.0139 | 81.622 | 67.6663 | 0.009155 | 1.0359 | 0.024715 | 23.859 | 0.018643 | 0.14044 | -0.000134 | -0.009989 | -0.009989 | -0.57222
2019-03-29 14:15 | 1553883306.5805 | 1319 | 0.20956 | 0.000447 | 1e-06 | 100.1164 | 99.8267 | 125.3641 | 21.1004 | 0.001975 | 0.001898 | -152.2226 | 0.23443 | 8.4021 | 0.002502 | 0.004279 | 0.004139 | 6.1e-05 | 0.000981 | 0.29264 | 1.3658 | 2.5473 | 100.16 | 20.9 | 37.013 | 69.1862 | 81.8509 | 67.343 | 0.009144 | 1.0369 | 0.024696 | 23.818 | 0.025912 | 0.13946 | 2.8e-05 | 0.002022 | 0.002022 | 2.8071
2019-03-29 14:15 | 1553883307.58 | 1320 | 0.20963 | 0.000449 | 0 | 100.1454 | 99.8616 | 125.2697 | 21.0986 | 0.001962 | 0.001876 | -152.2101 | 0.23502 | 8.4021 | 0.002544 | 0.00405 | 0.003868 | 6.2e-05 | 0.000986 | 0.29214 | 1.4163 | 2.5476 | 100.16 | 20.9 | 37.003 | 68.8778 | 82.8716 | 66.8178 | 0.009142 | 1.0385 | 0.024706 | 23.7906 | 0.025387 | 0.13981 | -0.000399 | -0.029677 | -0.029677 | -0.19174
2019-03-29 14:15 | 1553883308.5797 | 1321 | 0.20953 | 0.000443 | 0 | 100.1185 | 99.7933 | 125.5317 | 21.1 | 0.001925 | 0.00188 | -152.1914 | 0.23524 | 8.4011 | 0.002547 | 0.003713 | 0.003506 | 6.1e-05 | 0.000986 | 0.30486 | 1.3907 | 2.5474 | 100.16 | 20.9 | 36.993 | 67.9892 | 82.9299 | 66.9033 | 0.00914 | 1.0382 | 0.024728 | 23.8185 | 0.029097 | 0.13994 | 0.000158 | 0.011738 | 0.011738 | 0.48525
2019-03-29 14:15 | 1553883309.5806 | 1322 | 0.20958 | 0.000449 | 1e-06 | 100.043 | 99.7762 | 125.5561 | 21.0991 | 0.001942 | 0.001936 | -152.2053 | 0.23486 | 8.4033 | 0.002557 | 0.004078 | 0.003837 | 6.2e-05 | 0.000987 | 0.32781 | 1.3983 | 2.5473 | 100.16 | 20.9 | 36.972 | 68.3847 | 83.1764 | 68.2427 | 0.009134 | 1.0341 | 0.024682 | 23.8677 | 0.023885 | 0.13972 | -0.000153 | -0.011363 | -0.011363 | -0.50043
2019-03-29 14:15 | 1553883310.581 | 1323 | 0.20958 | 0.000452 | 1e-06 | 100.146 | 99.8284 | 125.665 | 21.0996 | 0.001981 | 0.001871 | -152.2114 | 0.23494 | 8.4009 | 0.002481 | 0.004506 | 0.004299 | 6.4e-05 | 0.000983 | 0.31238 | 1.3943 | 2.5475 | 100.16 | 20.9 | 36.96 | 69.3349 | 81.3487 | 66.6814 | 0.009131 | 1.0389 | 0.02475 | 23.8233 | 0.028406 | 0.13977 | -9.4e-05 | -0.007063 | -0.007063 | -0.80543
2019-03-29 14:15 | 1553883311.581 | 1324 | 0.2095 | 0.000446 | 1e-06 | 100.1023 | 99.9476 | 125.6528 | 21.1002 | 0.001957 | 0.001856 | -152.2098 | 0.23538 | 8.4033 | 0.002484 | 0.003318 | 0.003192 | 6.2e-05 | 0.000984 | 0.29286 | 1.3746 | 2.5474 | 100.16 | 20.9 | 36.949 | 68.7469 | 81.4327 | 66.3386 | 0.009129 | 1.0399 | 0.024761 | 23.8102 | 0.013836 | 0.14003 | 0.000368 | 0.027311 | 0.027311 | 0.20867
2019-03-29 14:15 | 1553883312.5798 | 1325 | 0.20952 | 0.000449 | 1e-06 | 100.049 | 99.9049 | 125.4401 | 21.0996 | 0.002001 | 0.001952 | -152.2224 | 0.23533 | 8.4023 | 0.002532 | 0.003922 | 0.003719 | 6.3e-05 | 0.000977 | 0.30397 | 1.3477 | 2.5474 | 100.16 | 20.9 | 36.93 | 69.8029 | 82.5684 | 68.6324 | 0.009124 | 1.033 | 0.024657 | 23.8703 | 0.012893 | 0.14 | 0.00024 | 0.017683 | 0.017683 | 0.32222
2019-03-29 14:15 | 1553883313.5807 | 1326 | 0.20953 | 0.00045 | 1e-06 | 100.0993 | 99.9175 | 125.6833 | 21.0992 | 0.002014 | 0.001878 | -152.2151 | 0.23512 | 8.4012 | 0.002481 | 0.003281 | 0.003057 | 6.4e-05 | 0.000971 | 0.29822 | 1.3117 | 2.5473 | 100.16 | 20.89 | 36.922 | 70.1325 | 81.349 | 66.8535 | 0.009116 | 1.0384 | 0.024745 | 23.8311 | 0.016257 | 0.13987 | 0.000197 | 0.014595 | 0.014595 | 0.39006
2019-03-29 14:15 | 1553883314.5802 | 1327 | 0.20958 | 0.000449 | 1e-06 | 100.073 | 99.8356 | 126.0586 | 21.1004 | 0.001976 | 0.001878 | -152.2294 | 0.23503 | 8.4014 | 0.002469 | 0.00385 | 0.003708 | 6.2e-05 | 0.000978 | 0.30541 | 1.3227 | 2.5474 | 100.16 | 20.89 | 36.908 | 69.2126 | 81.0702 | 66.8498 | 0.009113 | 1.0384 | 0.024782 | 23.8665 | 0.021238 | 0.13982 | -8.7e-05 | -0.006498 | -0.006498 | -0.87577
2019-03-29 14:15 | 1553883315.5802 | 1328 | 0.20954 | 0.000447 | 1e-06 | 100.137 | 99.8063 | 126.2903 | 21.0998 | 0.002025 | 0.001844 | -152.2427 | 0.23514 | 8.4026 | 0.002552 | 0.004083 | 0.003923 | 6.4e-05 | 0.000981 | 0.31902 | 1.3303 | 2.5474 | 100.16 | 20.89 | 36.902 | 70.385 | 83.0663 | 66.0472 | 0.009111 | 1.0408 | 0.024834 | 23.8603 | 0.029592 | 0.13988 | 0.000127 | 0.009475 | 0.009475 | 0.60089
2019-03-29 14:15 | 1553883316.5802 | 1329 | 0.20957 | 0.00045 | 1e-06 | 100.0861 | 99.9319 | 126.7309 | 21.0997 | 0.002 | 0.001894 | -152.2358 | 0.23466 | 8.4013 | 0.002498 | 0.004186 | 0.004018 | 6.1e-05 | 0.000986 | 0.30735 | 1.3601 | 2.5473 | 100.16 | 20.89 | 36.899 | 69.7843 | 81.7519 | 67.2339 | 0.009111 | 1.0372 | 0.024834 | 23.9436 | 0.013797 | 0.1396 | -7.1e-05 | -0.005353 | -0.005353 | -1.0615
2019-03-29 14:15 | 1553883317.5799 | 1330 | 0.20957 | 0.000449 | 1e-06 | 100.0346 | 99.7602 | 127.0198 | 21.0998 | 0.001941 | 0.001903 | -152.2402 | 0.23491 | 8.4026 | 0.002458 | 0.003267 | 0.003027 | 6.3e-05 | 0.000981 | 0.27417 | 1.3634 | 2.5474 | 100.16 | 20.89 | 36.895 | 68.3744 | 80.7952 | 67.4666 | 0.00911 | 1.0365 | 0.024854 | 23.9791 | 0.024573 | 0.13975 | -6.3e-05 | -0.004758 | -0.004758 | -1.1954
2019-03-29 14:15 | 1553883318.5808 | 1331 | 0.20955 | 0.000454 | 1e-06 | 100.0469 | 99.7848 | 127.3602 | 21.0999 | 0.001988 | 0.001956 | -152.2312 | 0.23398 | 8.4037 | 0.002533 | 0.004124 | 0.003927 | 6.3e-05 | 0.000978 | 0.3235 | 1.3188 | 2.5474 | 100.16 | 20.89 | 36.894 | 69.5045 | 82.6 | 68.731 | 0.009109 | 1.0327 | 0.024841 | 24.0557 | 0.023468 | 0.1392 | 8.2e-05 | 0.006049 | 0.006049 | 0.93649
2019-03-29 14:15 | 1553883319.5789 | 1332 | 0.20955 | 0.000449 | 0 | 100.0233 | 99.8472 | 127.6247 | 21.0997 | 0.001962 | 0.001948 | -152.2363 | 0.23526 | 8.4016 | 0.002563 | 0.004268 | 0.004094 | 6e-05 | 0.000975 | 0.30574 | 1.3537 | 2.5475 | 100.16 | 20.89 | 36.892 | 68.8712 | 83.3268 | 68.5382 | 0.009109 | 1.0332 | 0.024874 | 24.0739 | 0.015757 | 0.13996 | 3.2e-05 | 0.00237 | 0.00237 | 2.4036
2019-03-29 14:15 | 1553883320.579 | 1333 | 0.20958 | 0.00045 | 0 | 99.9772 | 99.652 | 127.9047 | 21.0997 | 0.001951 | 0.001908 | -152.2429 | 0.23497 | 8.4033 | 0.002525 | 0.003754 | 0.003604 | 6.2e-05 | 0.000977 | 0.30038 | 1.3425 | 2.5474 | 100.16 | 20.89 | 36.892 | 68.6122 | 82.4 | 67.5686 | 0.009109 | 1.0362 | 0.024937 | 24.0661 | 0.029139 | 0.13978 | -8.8e-05 | -0.006645 | -0.006645 | -0.85612
2019-03-29 14:15 | 1553883321.579 | 1334 | 0.20955 | 0.00045 | 0 | 100.0131 | 99.8588 | 127.7247 | 21.0996 | 0.001999 | 0.001867 | -152.2591 | 0.23495 | 8.4012 | 0.002564 | 0.0039 | 0.00366 | 6.2e-05 | 0.000983 | 0.32892 | 1.3723 | 2.5475 | 100.16 | 20.89 | 36.894 | 69.7716 | 83.3493 | 66.598 | 0.009109 | 1.0391 | 0.024955 | 24.0148 | 0.013807 | 0.13977 | 5.6e-05 | 0.004185 | 0.004185 | 1.3594
2019-03-29 14:15 | 1553883322.58 | 1335 | 0.20954 | 0.000452 | 0 | 99.9598 | 99.8332 | 127.847 | 21.0998 | 0.001964 | 0.00192 | -152.2565 | 0.23526 | 8.4036 | 0.002573 | 0.003913 | 0.003737 | 6.3e-05 | 0.00098 | 0.29695 | 1.3683 | 2.5475 | 100.16 | 20.89 | 36.899 | 68.9165 | 83.5563 | 67.8659 | 0.009111 | 1.0353 | 0.02492 | 24.0711 | 0.011334 | 0.13996 | 0.000121 | 0.009031 | 0.009031 | 0.63072
2019-03-29 14:15 | 1553883323.5791 | 1336 | 0.20955 | 0.000452 | 1e-06 | 99.9547 | 99.6684 | 127.6752 | 21.1001 | 0.001998 | 0.001871 | -152.2586 | 0.23474 | 8.4026 | 0.002579 | 0.003696 | 0.003486 | 6.2e-05 | 0.000992 | 0.29336 | 1.3435 | 2.5475 | 100.16 | 20.89 | 36.901 | 69.7534 | 83.7001 | 66.6834 | 0.009111 | 1.0389 | 0.024947 | 24.0132 | 0.02566 | 0.13964 | 4.9e-05 | 0.003626 | 0.003626 | 1.5675
2019-03-29 14:15 | 1553883324.5801 | 1337 | 0.20953 | 0.000453 | 0 | 99.9398 | 99.5216 | 127.3609 | 21.0999 | 0.001967 | 0.001825 | -152.2632 | 0.23466 | 8.4022 | 0.002632 | 0.00364 | 0.003508 | 6.1e-05 | 0.00099 | 0.31045 | 1.3989 | 2.5474 | 100.16 | 20.89 | 36.903 | 69.003 | 84.9887 | 65.5922 | 0.009112 | 1.0422 | 0.024956 | 23.9451 | 0.0375 | 0.1396 | 0.000163 | 0.012187 | 0.012187 | 0.46622
2019-03-29 14:15 | 1553883325.5806 | 1338 | 0.20955 | 0.000453 | 0 | 100.0328 | 99.7397 | 127.1327 | 21.1001 | 0.002027 | 0.001878 | -152.2586 | 0.23511 | 8.4031 | 0.002559 | 0.003098 | 0.002977 | 6.3e-05 | 0.000993 | 0.30453 | 1.3329 | 2.5474 | 100.16 | 20.89 | 36.91 | 70.4358 | 83.232 | 66.8544 | 0.009113 | 1.0384 | 0.024887 | 23.9682 | 0.026249 | 0.13987 | 6.4e-05 | 0.004794 | 0.004794 | 1.1875
2019-03-29 14:15 | 1553883326.5801 | 1339 | 0.20957 | 0.00045 | 0 | 99.9194 | 99.5612 | 126.6605 | 21.0988 | 0.002005 | 0.001889 | -152.2712 | 0.23465 | 8.4004 | 0.002625 | 0.003555 | 0.003304 | 6.3e-05 | 0.000985 | 0.30541 | 1.346 | 2.5474 | 100.16 | 20.89 | 36.914 | 69.8988 | 84.8169 | 67.1284 | 0.009114 | 1.0375 | 0.024831 | 23.9332 | 0.032121 | 0.13959 | -5.2e-05 | -0.003861 | -0.003861 | -1.4716
2019-03-29 14:15 | 1553883327.5802 | 1340 | 0.20956 | 0.000448 | 1e-06 | 100.0029 | 99.6387 | 126.2985 | 21.0999 | 0.001954 | 0.00186 | -152.2766 | 0.23545 | 8.4034 | 0.002604 | 0.003725 | 0.003539 | 6.3e-05 | 0.000983 | 0.31852 | 1.3282 | 2.5474 | 100.16 | 20.89 | 36.921 | 68.6913 | 84.3131 | 66.4343 | 0.009116 | 1.0396 | 0.024821 | 23.8746 | 0.032636 | 0.14007 | -1.9e-05 | -0.001478 | -0.001478 | -3.8573
2019-03-29 14:15 | 1553883328.5802 | 1341 | 0.20957 | 0.000452 | 0 | 100.0227 | 99.7598 | 125.5498 | 21.1015 | 0.002006 | 0.001896 | -152.2856 | 0.23536 | 8.4021 | 0.002593 | 0.003557 | 0.003423 | 6.3e-05 | 0.000984 | 0.30143 | 1.3705 | 2.5473 | 100.16 | 20.89 | 36.926 | 69.9382 | 84.0494 | 67.2958 | 0.009117 | 1.037 | 0.024716 | 23.8339 | 0.023533 | 0.14002 | -4.8e-05 | -0.00353 | -0.00353 | -1.6144
2019-03-29 14:15 | 1553883329.5793 | 1342 | 0.2096 | 0.000449 | 0 | 99.9499 | 99.6844 | 125.5401 | 21.1 | 0.00195 | 0.00194 | -152.286 | 0.23554 | 8.3999 | 0.002579 | 0.003712 | 0.003512 | 6.5e-05 | 0.000979 | 0.31924 | 1.3405 | 2.5473 | 100.16 | 20.89 | 36.929 | 68.5806 | 83.7053 | 68.3426 | 0.009118 | 1.0338 | 0.024677 | 23.8696 | 0.023792 | 0.14012 | -0.000213 | -0.015792 | -0.015792 | -0.36112
2019-03-29 14:15 | 1553883330.5803 | 1343 | 0.20955 | 0.00045 | 1e-06 | 99.8616 | 99.6472 | 125.2538 | 21.1004 | 0.001998 | 0.001919 | -152.2903 | 0.23453 | 8.4024 | 0.002556 | 0.003783 | 0.003597 | 6.8e-05 | 0.000984 | 0.27936 | 1.3324 | 2.5473 | 100.16 | 20.89 | 36.933 | 69.7554 | 83.1635 | 67.8501 | 0.009119 | 1.0353 | 0.024667 | 23.8252 | 0.019224 | 0.13953 | 4.4e-05 | 0.003173 | 0.003173 | 1.7898
2019-03-29 14:16 | 1553883331.5793 | 1344 | 0.20954 | 0.000449 | 1e-06 | 99.876 | 99.6711 | 125.2892 | 21.1004 | 0.00189 | 0.00191 | -152.2959 | 0.23474 | 8.403 | 0.002611 | 0.004101 | 0.003929 | 6.4e-05 | 0.000985 | 0.3183 | 1.386 | 2.5473 | 100.16 | 20.88 | 36.945 | 67.1488 | 84.4634 | 67.6169 | 0.009116 | 1.036 | 0.024679 | 23.8204 | 0.018367 | 0.13965 | 0.000108 | 0.00793 | 0.00793 | 0.7167
2019-03-29 14:16 | 1553883332.5803 | 1345 | 0.20961 | 0.00045 | 1e-06 | 99.9023 | 99.5906 | 125.0681 | 21.0998 | 0.001994 | 0.001887 | -152.2862 | 0.23573 | 8.402 | 0.002623 | 0.004103 | 0.004018 | 6.3e-05 | 0.000984 | 0.2943 | 1.3466 | 2.5474 | 100.16 | 20.88 | 36.951 | 69.6447 | 84.7601 | 67.0662 | 0.009118 | 1.0377 | 0.024677 | 23.7802 | 0.027957 | 0.14024 | -0.00032 | -0.023768 | -0.023768 | -0.24014
2019-03-29 14:16 | 1553883333.5794 | 1346 | 0.20956 | 0.000449 | 1e-06 | 99.8553 | 99.6626 | 125.0172 | 21.1005 | 0.001952 | 0.001892 | -152.2832 | 0.23489 | 8.4036 | 0.002566 | 0.004701 | 0.004509 | 6.1e-05 | 0.000982 | 0.29391 | 1.3061 | 2.5473 | 100.16 | 20.89 | 36.956 | 68.6412 | 83.4024 | 67.1886 | 0.009125 | 1.0373 | 0.024667 | 23.7796 | 0.017283 | 0.13974 | 1.3e-05 | 0.000921 | 0.000921 | 6.1759
2019-03-29 14:16 | 1553883334.5793 | 1347 | 0.20956 | 0.000447 | 0 | 99.8643 | 99.4459 | 125.1784 | 21.1003 | 0.001956 | 0.001893 | -152.2639 | 0.23472 | 8.4003 | 0.002644 | 0.003676 | 0.003515 | 6.4e-05 | 0.000987 | 0.31581 | 1.3529 | 2.5474 | 100.16 | 20.88 | 36.969 | 68.7328 | 85.2755 | 67.2172 | 0.009122 | 1.0372 | 0.024682 | 23.7959 | 0.037558 | 0.13964 | 2.3e-05 | 0.001723 | 0.001723 | 3.2991
2019-03-29 14:16 | 1553883335.5795 | 1348 | 0.20959 | 0.000449 | 1e-06 | 99.8793 | 99.4285 | 125.3765 | 21.1003 | 0.001947 | 0.001798 | -152.2636 | 0.23583 | 8.4015 | 0.002563 | 0.004297 | 0.004091 | 6.3e-05 | 0.000988 | 0.30718 | 1.3965 | 2.5474 | 100.16 | 20.88 | 36.974 | 68.519 | 83.3273 | 64.9449 | 0.009124 | 1.0442 | 0.024785 | 23.7351 | 0.040465 | 0.1403 | -0.000185 | -0.013825 | -0.013825 | -0.41303
2019-03-29 14:16 | 1553883336.5794 | 1349 | 0.20953 | 0.00045 | 1e-06 | 99.9059 | 99.4718 | 125.3155 | 21.1004 | 0.002023 | 0.001882 | -152.2654 | 0.23448 | 8.4012 | 0.002568 | 0.004012 | 0.003855 | 6e-05 | 0.000997 | 0.31741 | 1.3955 | 2.5473 | 100.16 | 20.89 | 36.981 | 70.3339 | 83.4491 | 66.9561 | 0.009131 | 1.038 | 0.024705 | 23.7998 | 0.038952 | 0.13949 | 0.000166 | 0.012322 | 0.012322 | 0.46074
2019-03-29 14:16 | 1553883337.5785 | 1350 | 0.20954 | 0.000449 | 0 | 99.9727 | 99.7148 | 125.2463 | 21.1001 | 0.001972 | 0.00189 | -152.2667 | 0.2348 | 8.4019 | 0.002613 | 0.004155 | 0.003983 | 6.7e-05 | 0.000999 | 0.31094 | 1.4119 | 2.5473 | 100.16 | 20.89 | 36.991 | 69.1186 | 84.5278 | 67.1547 | 0.009133 | 1.0374 | 0.024691 | 23.8002 | 0.023102 | 0.13968 | 0.000135 | 0.009995 | 0.009995 | 0.56881
2019-03-29 14:16 | 1553883338.5785 | 1351 | 0.20959 | 0.00045 | 0 | 99.9272 | 99.8854 | 125.4394 | 21.0999 | 0.002003 | 0.001846 | -152.2774 | 0.23512 | 8.4016 | 0.002518 | 0.004162 | 0.00404 | 6.4e-05 | 0.001003 | 0.30901 | 1.3981 | 2.5474 | 100.16 | 20.89 | 36.996 | 69.871 | 82.2279 | 66.0957 | 0.009135 | 1.0407 | 0.024749 | 23.7814 | 0.003737 | 0.13987 | -0.00016 | -0.011922 | -0.011922 | -0.4775
2019-03-29 14:16 | 1553883339.5786 | 1352 | 0.2096 | 0.000454 | 1e-06 | 99.8751 | 99.738 | 125.1902 | 21.1006 | 0.001966 | 0.001859 | -152.2686 | 0.23433 | 8.404 | 0.002486 | 0.004519 | 0.004309 | 6.6e-05 | 0.001009 | 0.3194 | 1.3848 | 2.5473 | 100.16 | 20.88 | 37.005 | 68.9728 | 81.4745 | 66.3953 | 0.009131 | 1.0398 | 0.024713 | 23.7683 | 0.012286 | 0.13941 | -0.000264 | -0.01963 | -0.01963 | -0.28904
2019-03-29 14:16 | 1553883340.5789 | 1353 | 0.20954 | 0.000447 | 1e-06 | 99.8907 | 99.6124 | 125.3655 | 21.1005 | 0.001957 | 0.001889 | -152.2565 | 0.23464 | 8.4005 | 0.002589 | 0.00447 | 0.004291 | 6.5e-05 | 0.001012 | 0.34882 | 1.3542 | 2.5473 | 100.16 | 20.88 | 37.008 | 68.769 | 83.9359 | 67.1332 | 0.009132 | 1.0375 | 0.024704 | 23.8108 | 0.02495 | 0.13959 | 0.000113 | 0.008315 | 0.008315 | 0.68321
2019-03-29 14:16 | 1553883341.5786 | 1354 | 0.20945 | 0.000455 | 1e-06 | 99.9565 | 99.6076 | 125.4458 | 21.1003 | 0.001986 | 0.001875 | -152.2431 | 0.23516 | 8.4002 | 0.002566 | 0.003869 | 0.003741 | 6.4e-05 | 0.001006 | 0.31996 | 1.3412 | 2.5474 | 100.16 | 20.88 | 37.011 | 69.444 | 83.399 | 66.7819 | 0.009133 | 1.0386 | 0.024724 | 23.8061 | 0.031274 | 0.1399 | 0.000688 | 0.051084 | 0.051084 | 0.11146
2019-03-29 14:16 | 1553883342.5796 | 1355 | 0.2096 | 0.000449 | 1e-06 | 99.9574 | 99.696 | 125.3202 | 21.1005 | 0.002005 | 0.001844 | -152.2495 | 0.23527 | 8.4021 | 0.002482 | 0.003397 | 0.003157 | 6.6e-05 | 0.000998 | 0.32493 | 1.3491 | 2.5474 | 100.16 | 20.89 | 37.016 | 69.9021 | 81.3784 | 66.0524 | 0.00914 | 1.0408 | 0.024739 | 23.7686 | 0.02342 | 0.13996 | -0.000229 | -0.017037 | -0.017037 | -0.33435
2019-03-29 14:16 | 1553883343.5797 | 1356 | 0.20959 | 0.000445 | 1e-06 | 99.8781 | 99.6465 | 125.2216 | 21.0998 | 0.002008 | 0.001884 | -152.2683 | 0.23444 | 8.4016 | 0.002517 | 0.004027 | 0.003828 | 6.7e-05 | 0.000997 | 0.28224 | 1.3321 | 2.5473 | 100.16 | 20.89 | 37.017 | 69.9778 | 82.2152 | 67.0109 | 0.00914 | 1.0379 | 0.024694 | 23.7928 | 0.02076 | 0.13947 | -0.000176 | -0.013116 | -0.013116 | -0.43276
2019-03-29 14:16 | 1553883344.5787 | 1357 | 0.20941 | 0.000449 | 1e-06 | 99.9263 | 99.5496 | 125.1219 | 21.0993 | 0.001982 | 0.001883 | -152.2621 | 0.23397 | 8.4021 | 0.002493 | 0.00367 | 0.003434 | 6.7e-05 | 0.000996 | 0.28871 | 1.3308 | 2.5472 | 100.16 | 20.88 | 37.017 | 69.3575 | 81.6424 | 66.9666 | 0.009134 | 1.038 | 0.024686 | 23.7818 | 0.033779 | 0.13919 | 0.000883 | 0.065443 | 0.065443 | 0.086565
2019-03-29 14:16 | 1553883345.5787 | 1358 | 0.2094 | 0.000449 | 1e-06 | 99.9703 | 99.566 | 125.0721 | 21.0999 | 0.001973 | 0.001862 | -152.2508 | 0.23513 | 8.4025 | 0.002497 | 0.003941 | 0.003703 | 6.7e-05 | 0.000998 | 0.35458 | 1.3512 | 2.5474 | 100.16 | 20.88 | 37.014 | 69.1388 | 81.735 | 66.471 | 0.009133 | 1.0395 | 0.024699 | 23.7597 | 0.036246 | 0.13988 | 0.000976 | 0.072361 | 0.072361 | 0.078677
2019-03-29 14:16 | 1553883346.5797 | 1359 | 0.20961 | 0.000451 | 1e-06 | 100.0047 | 99.7175 | 125.3711 | 21.1001 | 0.001978 | 0.001883 | -152.253 | 0.23465 | 8.4036 | 0.002583 | 0.004435 | 0.004279 | 6.6e-05 | 0.000996 | 0.28368 | 1.337 | 2.5472 | 100.16 | 20.88 | 37.011 | 69.2552 | 83.8108 | 66.99 | 0.009133 | 1.0379 | 0.02471 | 23.8063 | 0.025722 | 0.13959 | -0.000325 | -0.02416 | -0.02416 | -0.23516
2019-03-29 14:16 | 1553883347.5798 | 1360 | 0.20951 | 0.000449 | 1e-06 | 100.0179 | 99.708 | 125.6202 | 21.0997 | 0.001972 | 0.001858 | -152.255 | 0.23392 | 8.401 | 0.002506 | 0.003381 | 0.003262 | 6.2e-05 | 0.001002 | 0.30901 | 1.341 | 2.5473 | 100.16 | 20.88 | 37.006 | 69.1091 | 81.9493 | 66.3749 | 0.009131 | 1.0398 | 0.024756 | 23.8084 | 0.027757 | 0.13916 | 0.000328 | 0.02434 | 0.02434 | 0.23269
2019-03-29 14:16 | 1553883348.5798 | 1361 | 0.2096 | 0.000449 | 1e-06 | 100.0011 | 99.6373 | 125.7209 | 21.0999 | 0.001984 | 0.001864 | -152.2484 | 0.23429 | 8.4013 | 0.002535 | 0.003782 | 0.003631 | 6.3e-05 | 0.000997 | 0.31006 | 1.3605 | 2.5474 | 100.16 | 20.88 | 37.002 | 69.4172 | 82.657 | 66.525 | 0.00913 | 1.0394 | 0.024761 | 23.8232 | 0.032598 | 0.13938 | -0.0002 | -0.014918 | -0.014918 | -0.38027
2019-03-29 14:16 | 1553883349.5803 | 1362 | 0.20951 | 0.000451 | 1e-06 | 100.0017 | 99.7319 | 125.9392 | 21.1 | 0.001991 | 0.001854 | -152.2532 | 0.23401 | 8.4018 | 0.002493 | 0.003923 | 0.003687 | 6.1e-05 | 0.001004 | 0.31868 | 1.3537 | 2.5474 | 100.16 | 20.88 | 36.989 | 69.5639 | 81.6299 | 66.273 | 0.009127 | 1.0401 | 0.024792 | 23.835 | 0.024169 | 0.13921 | 0.00032 | 0.023756 | 0.023756 | 0.23851
2019-03-29 14:16 | 1553883350.5794 | 1363 | 0.20954 | 0.000451 | 1e-06 | 99.9571 | 99.6667 | 125.9021 | 21.1011 | 0.001974 | 0.001875 | -152.2549 | 0.2344 | 8.403 | 0.002529 | 0.00385 | 0.003668 | 6.4e-05 | 0.001001 | 0.33826 | 1.3683 | 2.5474 | 100.16 | 20.88 | 36.989 | 69.1672 | 82.5105 | 66.7928 | 0.009127 | 1.0385 | 0.024769 | 23.8497 | 0.026027 | 0.13945 | 0.000141 | 0.010387 | 0.010387 | 0.54642
2019-03-29 14:16 | 1553883351.5789 | 1364 | 0.20958 | 0.000452 | 1e-06 | 99.9041 | 99.4223 | 125.7678 | 21.1004 | 0.001975 | 0.001871 | -152.2542 | 0.23505 | 8.3999 | 0.002605 | 0.003864 | 0.003723 | 6.3e-05 | 0.001003 | 0.31133 | 1.3798 | 2.5475 | 100.16 | 20.88 | 36.987 | 69.1852 | 84.3213 | 66.6947 | 0.009127 | 1.0388 | 0.024759 | 23.8336 | 0.043242 | 0.13983 | -0.000105 | -0.007935 | -0.007935 | -0.71726
2019-03-29 14:16 | 1553883352.579 | 1365 | 0.20953 | 0.00045 | 1e-06 | 99.9766 | 99.7783 | 125.7836 | 21.1007 | 0.001935 | 0.001877 | -152.2538 | 0.23411 | 8.4038 | 0.002598 | 0.003813 | 0.003718 | 6.5e-05 | 0.001003 | 0.31365 | 1.3447 | 2.5473 | 100.16 | 20.88 | 36.983 | 68.2387 | 84.1624 | 66.8354 | 0.009126 | 1.0384 | 0.024756 | 23.84 | 0.017758 | 0.13927 | 0.000173 | 0.012853 | 0.012853 | 0.44103
2019-03-29 14:16 | 1553883353.579 | 1366 | 0.20955 | 0.000452 | 1e-06 | 100.0218 | 99.6445 | 125.6528 | 21.1004 | 0.00195 | 0.001907 | -152.2346 | 0.23413 | 8.4047 | 0.002634 | 0.004252 | 0.003949 | 6.5e-05 | 0.001 | 0.32764 | 1.3428 | 2.5474 | 100.16 | 20.88 | 36.981 | 68.5786 | 85.0212 | 67.5638 | 0.009125 | 1.0362 | 0.024716 | 23.8531 | 0.033802 | 0.13928 | 5.5e-05 | 0.004009 | 0.004009 | 1.4141
2019-03-29 14:16 | 1553883354.579 | 1367 | 0.20958 | 0.000448 | 0 | 99.9335 | 99.5162 | 125.3108 | 21.1018 | 0.001952 | 0.001823 | -152.2544 | 0.23423 | 8.4038 | 0.002635 | 0.004008 | 0.003847 | 6.1e-05 | 0.000998 | 0.29286 | 1.3163 | 2.5475 | 100.16 | 20.88 | 36.981 | 68.6275 | 85.043 | 65.5318 | 0.009125 | 1.0424 | 0.024757 | 23.7495 | 0.037428 | 0.13935 | -9.5e-05 | -0.007113 | -0.007113 | -0.79735
2019-03-29 14:16 | 1553883355.5785 | 1368 | 0.20953 | 0.000449 | 0 | 100.0337 | 99.78 | 125.0456 | 21.1011 | 0.001959 | 0.001852 | -152.2701 | 0.23553 | 8.4042 | 0.002616 | 0.003703 | 0.003521 | 5.8e-05 | 0.000995 | 0.29518 | 1.3606 | 2.5475 | 100.16 | 20.88 | 36.987 | 68.8125 | 84.5859 | 66.2303 | 0.009127 | 1.0403 | 0.024705 | 23.7488 | 0.022718 | 0.14012 | 0.000211 | 0.015643 | 0.015643 | 0.36455
2019-03-29 14:16 | 1553883356.5795 | 1369 | 0.20957 | 0.000453 | 0 | 100.0161 | 99.622 | 124.8496 | 21.1025 | 0.00198 | 0.001853 | -152.2656 | 0.23515 | 8.4021 | 0.002564 | 0.004197 | 0.004 | 5.9e-05 | 0.000991 | 0.30541 | 1.3319 | 2.5474 | 100.16 | 20.88 | 36.987 | 69.3099 | 83.3348 | 66.256 | 0.009127 | 1.0402 | 0.024685 | 23.7311 | 0.035315 | 0.13989 | -5.1e-05 | -0.003817 | -0.003817 | -1.4916
2019-03-29 14:16 | 1553883357.5791 | 1370 | 0.20957 | 0.000453 | 1e-06 | 100.0275 | 99.7431 | 124.7825 | 21.1006 | 0.00193 | 0.001864 | -152.268 | 0.23553 | 8.4039 | 0.002582 | 0.004028 | 0.003853 | 6.2e-05 | 0.000988 | 0.30309 | 1.3614 | 2.5471 | 100.16 | 20.88 | 36.996 | 68.1022 | 83.7799 | 66.5114 | 0.009129 | 1.0394 | 0.024669 | 23.7336 | 0.025461 | 0.14012 | -2.6e-05 | -0.001944 | -0.001944 | -2.9339
2019-03-29 14:16 | 1553883358.5781 | 1371 | 0.20954 | 0.00045 | 1e-06 | 99.9852 | 99.7991 | 124.9869 | 21.1004 | 0.001944 | 0.001847 | -152.2681 | 0.23443 | 8.4039 | 0.002605 | 0.003224 | 0.003044 | 6.1e-05 | 0.000983 | 0.32228 | 1.3423 | 2.5472 | 100.16 | 20.88 | 37.002 | 68.4538 | 84.3275 | 66.1094 | 0.00913 | 1.0406 | 0.024704 | 23.739 | 0.016668 | 0.13946 | 0.000148 | 0.010921 | 0.010921 | 0.51975
2019-03-29 14:16 | 1553883359.5782 | 1372 | 0.2096 | 0.000448 | 1e-06 | 100.0301 | 99.7069 | 125.0571 | 21.1012 | 0.001972 | 0.001887 | -152.262 | 0.23522 | 8.4018 | 0.002488 | 0.004446 | 0.004363 | 6.3e-05 | 0.000983 | 0.32278 | 1.3561 | 2.5473 | 100.16 | 20.88 | 37.008 | 69.1228 | 81.5272 | 67.0738 | 0.009132 | 1.0377 | 0.024676 | 23.7794 | 0.028947 | 0.13993 | -0.000213 | -0.015867 | -0.015867 | -0.35893
2019-03-29 14:16 | 1553883360.5788 | 1373 | 0.20955 | 0.000451 | 1e-06 | 99.9625 | 99.7428 | 125.1697 | 21.1014 | 0.00202 | 0.001861 | -152.2699 | 0.23501 | 8.4019 | 0.002501 | 0.003454 | 0.00316 | 6.2e-05 | 0.000995 | 0.33627 | 1.3488 | 2.5473 | 100.16 | 20.88 | 37.019 | 70.2632 | 81.8394 | 66.4562 | 0.009135 | 1.0396 | 0.024709 | 23.7685 | 0.019681 | 0.13981 | 8.1e-05 | 0.005966 | 0.005966 | 0.95383
2019-03-29 14:16 | 1553883361.5792 | 1374 | 0.20956 | 0.000447 | 0 | 99.9553 | 99.767 | 125.2406 | 21.1012 | 0.001944 | 0.001894 | -152.2668 | 0.23455 | 8.4029 | 0.002575 | 0.004204 | 0.004035 | 6.4e-05 | 0.001001 | 0.32029 | 1.3626 | 2.5473 | 100.16 | 20.89 | 37.025 | 68.4437 | 83.6098 | 67.2447 | 0.009142 | 1.0372 | 0.024687 | 23.8028 | 0.016866 | 0.13954 | 2.9e-05 | 0.002134 | 0.002134 | 2.6613
2019-03-29 14:16 | 1553883362.5796 | 1375 | 0.20957 | 0.000454 | 1e-06 | 99.9197 | 99.5926 | 125.5214 | 21.1006 | 0.001975 | 0.001877 | -152.2724 | 0.23495 | 8.4035 | 0.002556 | 0.00404 | 0.003957 | 6.2e-05 | 0.001002 | 0.30215 | 1.3712 | 2.5473 | 100.16 | 20.89 | 37.035 | 69.1802 | 83.1543 | 66.833 | 0.009144 | 1.0384 | 0.02473 | 23.815 | 0.029327 | 0.13977 | -7.4e-05 | -0.005539 | -0.005539 | -1.027
2019-03-29 14:16 | 1553883363.5783 | 1376 | 0.20955 | 0.000449 | 1e-06 | 100.0574 | 99.6953 | 125.6228 | 21.1013 | 0.001946 | 0.001899 | -152.2664 | 0.23541 | 8.4022 | 0.002452 | 0.004512 | 0.004293 | 6.5e-05 | 0.001011 | 0.33428 | 1.3772 | 2.5474 | 100.16 | 20.89 | 37.038 | 68.4866 | 80.6507 | 67.361 | 0.009145 | 1.0368 | 0.024721 | 23.8431 | 0.032423 | 0.14005 | 8.7e-05 | 0.006463 | 0.006463 | 0.88196
2019-03-29 14:16 | 1553883364.5793 | 1377 | 0.20954 | 0.000453 | 0 | 100.0475 | 99.7578 | 125.7009 | 21.1017 | 0.002022 | 0.001879 | -152.2855 | 0.23496 | 8.401 | 0.002509 | 0.003293 | 0.003143 | 6e-05 | 0.001018 | 0.30414 | 1.3503 | 2.5474 | 100.16 | 20.89 | 37.043 | 70.3164 | 82.0301 | 66.8747 | 0.009146 | 1.0383 | 0.024746 | 23.8335 | 0.025938 | 0.13978 | 0.000128 | 0.009494 | 0.009494 | 0.59923
2019-03-29 14:16 | 1553883365.5798 | 1378 | 0.20956 | 0.000451 | 1e-06 | 99.999 | 99.6701 | 125.892 | 21.1006 | 0.001952 | 0.001866 | -152.2823 | 0.23518 | 8.4018 | 0.002456 | 0.00412 | 0.003887 | 6.1e-05 | 0.001022 | 0.32068 | 1.3196 | 2.5474 | 100.16 | 20.89 | 37.048 | 68.6466 | 80.7539 | 66.5677 | 0.009147 | 1.0392 | 0.024776 | 23.8409 | 0.02947 | 0.13991 | 2.5e-05 | 0.00185 | 0.00185 | 3.0776
2019-03-29 14:16 | 1553883366.5785 | 1379 | 0.20952 | 0.000449 | 1e-06 | 100.0185 | 99.7482 | 125.8781 | 21.1015 | 0.001961 | 0.001855 | -152.2717 | 0.23464 | 8.4018 | 0.002465 | 0.003804 | 0.003685 | 6e-05 | 0.001023 | 0.31581 | 1.3142 | 2.5474 | 100.16 | 20.89 | 37.048 | 68.8503 | 80.966 | 66.3158 | 0.009147 | 1.04 | 0.024784 | 23.8307 | 0.024199 | 0.13959 | 0.000239 | 0.017733 | 0.017733 | 0.32038
2019-03-29 14:16 | 1553883367.5784 | 1380 | 0.20956 | 0.00045 | 1e-06 | 99.9814 | 99.7633 | 126.0929 | 21.1016 | 0.001988 | 0.001908 | -152.2734 | 0.23519 | 8.4013 | 0.00252 | 0.003466 | 0.003332 | 5.8e-05 | 0.001025 | 0.28783 | 1.377 | 2.5475 | 100.16 | 20.89 | 37.045 | 69.496 | 82.2771 | 67.5868 | 0.009147 | 1.0361 | 0.024759 | 23.8956 | 0.019533 | 0.13992 | 0 | -5.5e-05 | -5.5e-05 | -102.9723
2019-03-29 14:16 | 1553883368.5785 | 1381 | 0.20955 | 0.00045 | 1e-06 | 100.0313 | 99.654 | 125.8089 | 21.1015 | 0.001985 | 0.001879 | -152.2714 | 0.23451 | 8.404 | 0.002534 | 0.00319 | 0.002969 | 6.1e-05 | 0.001021 | 0.32045 | 1.3588 | 2.5475 | 100.15 | 20.89 | 37.042 | 69.4336 | 82.6203 | 66.8814 | 0.009147 | 1.0383 | 0.024757 | 23.844 | 0.033801 | 0.13951 | 9e-05 | 0.006685 | 0.006685 | 0.8493
2019-03-29 14:16 | 1553883369.578 | 1382 | 0.20957 | 0.000449 | 0 | 99.9706 | 99.7397 | 125.5153 | 21.1015 | 0.00198 | 0.001841 | -152.2791 | 0.23489 | 8.4041 | 0.002538 | 0.00389 | 0.003813 | 5.8e-05 | 0.001027 | 0.29407 | 1.3038 | 2.5475 | 100.15 | 20.89 | 37.037 | 69.3034 | 82.7176 | 65.981 | 0.009146 | 1.041 | 0.02476 | 23.7846 | 0.02068 | 0.13974 | -5.5e-05 | -0.004097 | -0.004097 | -1.3881
2019-03-29 14:16 | 1553883370.5785 | 1383 | 0.20955 | 0.000455 | 1e-06 | 99.9864 | 99.6182 | 125.4254 | 21.1012 | 0.001928 | 0.001852 | -152.2772 | 0.23476 | 8.4032 | 0.002485 | 0.004397 | 0.004179 | 6e-05 | 0.001033 | 0.30143 | 1.3057 | 2.5474 | 100.15 | 20.89 | 37.024 | 68.0598 | 81.4368 | 66.246 | 0.009142 | 1.0402 | 0.024742 | 23.7854 | 0.033002 | 0.13966 | 9.3e-05 | 0.0069 | 0.0069 | 0.82379
2019-03-29 14:16 | 1553883371.5775 | 1384 | 0.20958 | 0.000449 | 0 | 99.9424 | 99.754 | 125.4183 | 21.101 | 0.001875 | 0.001853 | -152.2865 | 0.23478 | 8.4021 | 0.002536 | 0.004069 | 0.003958 | 6e-05 | 0.001039 | 0.29734 | 1.2969 | 2.5473 | 100.15 | 20.89 | 37.018 | 66.7966 | 82.6727 | 66.2466 | 0.009141 | 1.0402 | 0.024741 | 23.7847 | 0.016877 | 0.13967 | -0.000113 | -0.008405 | -0.008405 | -0.67635
2019-03-29 14:16 | 1553883372.5786 | 1385 | 0.20956 | 0.000453 | 1e-06 | 99.9365 | 99.7547 | 125.2301 | 21.1013 | 0.001959 | 0.001882 | -152.295 | 0.23536 | 8.4043 | 0.002543 | 0.004301 | 0.004061 | 6.2e-05 | 0.001037 | 0.30143 | 1.3498 | 2.5475 | 100.15 | 20.89 | 37.004 | 68.8081 | 82.8345 | 66.946 | 0.009138 | 1.0381 | 0.024697 | 23.7913 | 0.01628 | 0.14002 | 1.4e-05 | 0.001015 | 0.001015 | 5.6145
2019-03-29 14:16 | 1553883373.5776 | 1386 | 0.20961 | 0.000447 | 1e-06 | 100.0023 | 99.5315 | 125.2521 | 21.1024 | 0.001934 | 0.00186 | -152.3113 | 0.23535 | 8.4028 | 0.00257 | 0.003715 | 0.003522 | 6.1e-05 | 0.001035 | 0.323 | 1.343 | 2.5474 | 100.15 | 20.89 | 36.997 | 68.2138 | 83.4792 | 66.4345 | 0.009136 | 1.0396 | 0.024718 | 23.7755 | 0.042208 | 0.14001 | -0.000259 | -0.019317 | -0.019317 | -0.295
2019-03-29 14:16 | 1553883374.579 | 1387 | 0.20961 | 0.000449 | 1e-06 | 100.0035 | 99.7186 | 125.3479 | 21.1017 | 0.001978 | 0.001838 | -152.2942 | 0.23576 | 8.4018 | 0.002577 | 0.003581 | 0.0035 | 6.2e-05 | 0.001035 | 0.32516 | 1.3394 | 2.5474 | 100.15 | 20.89 | 36.99 | 69.273 | 83.6555 | 65.8874 | 0.009134 | 1.0413 | 0.024747 | 23.7655 | 0.025523 | 0.14025 | -0.000283 | -0.021076 | -0.021076 | -0.27084
2019-03-29 14:16 | 1553883375.5787 | 1388 | 0.20957 | 0.000452 | 0 | 99.9766 | 99.6742 | 125.5099 | 21.1028 | 0.001936 | 0.001854 | -152.2965 | 0.23572 | 8.4033 | 0.002534 | 0.003917 | 0.003638 | 6.3e-05 | 0.001033 | 0.29535 | 1.3569 | 2.5474 | 100.15 | 20.89 | 36.976 | 68.2455 | 82.6247 | 66.2921 | 0.009131 | 1.0401 | 0.024749 | 23.795 | 0.027094 | 0.14023 | -6.5e-05 | -0.004863 | -0.004863 | -1.1735
2019-03-29 14:16 | 1553883376.5777 | 1389 | 0.20956 | 0.000454 | 0 | 99.9619 | 99.564 | 125.7181 | 21.1018 | 0.001986 | 0.001841 | -152.2976 | 0.23531 | 8.4029 | 0.002483 | 0.003503 | 0.003327 | 6.1e-05 | 0.00102 | 0.31598 | 1.3257 | 2.5474 | 100.15 | 20.89 | 36.968 | 69.466 | 81.3923 | 65.9626 | 0.009129 | 1.0411 | 0.024781 | 23.8032 | 0.03568 | 0.13999 | 3.4e-05 | 0.00251 | 0.00251 | 2.2696
2019-03-29 14:16 | 1553883377.5778 | 1390 | 0.20958 | 0.000454 | 0 | 99.8969 | 99.7745 | 125.6908 | 21.1024 | 0.001912 | 0.001829 | -152.3085 | 0.23536 | 8.4029 | 0.002542 | 0.003888 | 0.003687 | 5.9e-05 | 0.001016 | 0.2975 | 1.3274 | 2.5473 | 100.15 | 20.89 | 36.963 | 67.6772 | 82.8161 | 65.6836 | 0.009127 | 1.0419 | 0.024789 | 23.7908 | 0.010968 | 0.14002 | -0.00012 | -0.008924 | -0.008924 | -0.63861
2019-03-29 14:16 | 1553883378.5778 | 1391 | 0.20959 | 0.000451 | 1e-06 | 99.9715 | 99.6005 | 125.7824 | 21.1019 | 0.001895 | 0.001893 | -152.3018 | 0.23483 | 8.401 | 0.002497 | 0.003399 | 0.003299 | 6.1e-05 | 0.001021 | 0.30718 | 1.3352 | 2.5474 | 100.15 | 20.89 | 36.953 | 67.2594 | 81.744 | 67.2236 | 0.009125 | 1.0372 | 0.024742 | 23.8535 | 0.033257 | 0.1397 | -0.000149 | -0.011154 | -0.011154 | -0.50974
2019-03-29 14:16 | 1553883379.5788 | 1392 | 0.20957 | 0.000455 | 0 | 99.9715 | 99.7486 | 125.7669 | 21.1032 | 0.001954 | 0.001785 | -152.299 | 0.23504 | 8.4005 | 0.002439 | 0.003835 | 0.003639 | 6.2e-05 | 0.001018 | 0.31902 | 1.3013 | 2.5474 | 100.15 | 20.89 | 36.948 | 68.6945 | 80.3502 | 64.6275 | 0.009124 | 1.0452 | 0.024835 | 23.7609 | 0.019965 | 0.13983 | -4.1e-05 | -0.003055 | -0.003055 | -1.8628
2019-03-29 14:16 | 1553883380.5789 | 1393 | 0.20956 | 0.000456 | 1e-06 | 99.8706 | 99.596 | 125.5638 | 21.1016 | 0.001917 | 0.001877 | -152.287 | 0.23562 | 8.4015 | 0.002424 | 0.003598 | 0.003491 | 5.9e-05 | 0.001016 | 0.30879 | 1.3314 | 2.5473 | 100.15 | 20.89 | 36.943 | 67.7905 | 79.9875 | 66.8371 | 0.009122 | 1.0384 | 0.024734 | 23.8192 | 0.024625 | 0.14017 | -7e-06 | -0.00054 | -0.00054 | -10.5638
2019-03-29 14:16 | 1553883381.5779 | 1394 | 0.20956 | 0.000451 | 1e-06 | 99.8999 | 99.7534 | 125.6932 | 21.1031 | 0.002011 | 0.001847 | -152.2765 | 0.2344 | 8.4028 | 0.002517 | 0.003115 | 0.002878 | 6e-05 | 0.001022 | 0.30663 | 1.3387 | 2.5474 | 100.15 | 20.89 | 36.942 | 70.0707 | 82.2122 | 66.1151 | 0.009122 | 1.0406 | 0.024773 | 23.8062 | 0.013133 | 0.13944 | 6e-06 | 0.000396 | 0.000396 | 14.3313
2019-03-29 14:16 | 1553883382.5769 | 1395 | 0.20955 | 0.000455 | 0 | 99.864 | 99.6452 | 125.8519 | 21.1034 | 0.001929 | 0.001808 | -152.2609 | 0.23491 | 8.4024 | 0.002402 | 0.003677 | 0.003558 | 5.9e-05 | 0.00102 | 0.3115 | 1.3008 | 2.5474 | 100.15 | 20.89 | 36.942 | 68.0775 | 79.4528 | 65.1772 | 0.009122 | 1.0435 | 0.024823 | 23.7882 | 0.019623 | 0.13975 | 0.000107 | 0.007954 | 0.007954 | 0.71508
2019-03-29 14:16 | 1553883383.578 | 1396 | 0.20954 | 0.000455 | 0 | 99.9658 | 99.6756 | 126.0622 | 21.1026 | 0.001975 | 0.00187 | -152.2638 | 0.23522 | 8.403 | 0.00246 | 0.003489 | 0.003324 | 6.3e-05 | 0.001023 | 0.28528 | 1.3408 | 2.5474 | 100.15 | 20.89 | 36.943 | 69.1999 | 80.8563 | 66.6671 | 0.009122 | 1.0389 | 0.024789 | 23.8605 | 0.026007 | 0.13994 | 0.000115 | 0.008509 | 0.008509 | 0.66932
2019-03-29 14:16 | 1553883384.577 | 1397 | 0.20952 | 0.000452 | 1e-06 | 99.9502 | 99.5674 | 126.2582 | 21.103 | 0.001976 | 0.001867 | -152.2624 | 0.2336 | 8.4028 | 0.002493 | 0.003718 | 0.003555 | 5.6e-05 | 0.001025 | 0.29479 | 1.399 | 2.5474 | 100.15 | 20.89 | 36.945 | 69.2159 | 81.6467 | 66.6051 | 0.009123 | 1.0391 | 0.024811 | 23.8768 | 0.034328 | 0.13897 | 0.000227 | 0.016864 | 0.016864 | 0.33539
2019-03-29 14:16 | 1553883385.577 | 1398 | 0.20952 | 0.000455 | 0 | 99.967 | 99.6046 | 126.119 | 21.1024 | 0.001918 | 0.001764 | -152.266 | 0.23426 | 8.4021 | 0.002403 | 0.004081 | 0.003906 | 5.9e-05 | 0.001033 | 0.31708 | 1.3642 | 2.5473 | 100.16 | 20.89 | 36.949 | 67.816 | 79.4754 | 64.1265 | 0.009123 | 1.0468 | 0.024888 | 23.7765 | 0.032487 | 0.13936 | 0.000275 | 0.020534 | 0.020534 | 0.27622
2019-03-29 14:16 | 1553883386.5771 | 1399 | 0.20955 | 0.000452 | 1e-06 | 99.935 | 99.6066 | 126.1077 | 21.1033 | 0.001957 | 0.001915 | -152.2726 | 0.23468 | 8.403 | 0.0025 | 0.003521 | 0.003325 | 5.8e-05 | 0.001028 | 0.2959 | 1.373 | 2.5474 | 100.16 | 20.89 | 36.949 | 68.7662 | 81.7985 | 67.757 | 0.009123 | 1.0356 | 0.024754 | 23.903 | 0.029437 | 0.13961 | 6.3e-05 | 0.004638 | 0.004638 | 1.2253
2019-03-29 14:16 | 1553883387.5781 | 1400 | 0.20953 | 0.000452 | 1e-06 | 99.999 | 99.7708 | 125.9274 | 21.1028 | 0.00203 | 0.001843 | -152.2664 | 0.23464 | 8.4018 | 0.00249 | 0.003693 | 0.003484 | 6e-05 | 0.001019 | 0.27616 | 1.3723 | 2.5474 | 100.16 | 20.89 | 36.95 | 70.5102 | 81.5756 | 66.029 | 0.009123 | 1.0409 | 0.024799 | 23.8253 | 0.020439 | 0.13959 | 0.000181 | 0.013456 | 0.013456 | 0.4222
2019-03-29 14:16 | 1553883388.5771 | 1401 | 0.20951 | 0.000454 | 1e-06 | 99.9613 | 99.5964 | 126.1727 | 21.1024 | 0.001882 | 0.001883 | -152.2677 | 0.23554 | 8.4015 | 0.002521 | 0.004418 | 0.004226 | 6e-05 | 0.001012 | 0.28191 | 1.3449 | 2.5474 | 100.16 | 20.89 | 36.955 | 66.965 | 82.3074 | 66.9735 | 0.009124 | 1.038 | 0.024789 | 23.8817 | 0.032714 | 0.14012 | 0.000293 | 0.02176 | 0.02176 | 0.26209
2019-03-29 14:16 | 1553883389.5772 | 1402 | 0.20951 | 0.000454 | 1e-06 | 99.9134 | 99.7042 | 126.6562 | 21.1026 | 0.001999 | 0.00187 | -152.2627 | 0.23484 | 8.4 | 0.002479 | 0.004154 | 0.003969 | 6.2e-05 | 0.001007 | 0.29319 | 1.3313 | 2.5473 | 100.16 | 20.89 | 36.958 | 69.7739 | 81.2971 | 66.6733 | 0.009125 | 1.0389 | 0.024847 | 23.9168 | 0.018747 | 0.13971 | 0.000305 | 0.022751 | 0.022751 | 0.24992
2019-03-29 14:16 | 1553883390.5783 | 1403 | 0.20952 | 0.000456 | 0 | 99.9694 | 99.7407 | 126.8987 | 21.1019 | 0.001928 | 0.001832 | -152.2827 | 0.23545 | 8.4013 | 0.002481 | 0.003823 | 0.003633 | 5.9e-05 | 0.001001 | 0.30326 | 1.3509 | 2.5474 | 100.16 | 20.89 | 36.961 | 68.0567 | 81.3587 | 65.7575 | 0.009126 | 1.0417 | 0.024905 | 23.9074 | 0.020481 | 0.14007 | 0.000252 | 0.018836 | 0.018836 | 0.30265
2019-03-29 14:17 | 1553883391.5772 | 1404 | 0.20951 | 0.000455 | 0 | 99.9754 | 99.6049 | 126.8382 | 21.1028 | 0.001921 | 0.001849 | -152.3015 | 0.23478 | 8.402 | 0.002502 | 0.004049 | 0.003788 | 5.9e-05 | 0.001006 | 0.3095 | 1.3357 | 2.5472 | 100.16 | 20.9 | 36.97 | 67.8969 | 81.8671 | 66.16 | 0.009134 | 1.0405 | 0.024884 | 23.9159 | 0.033206 | 0.13967 | 0.000326 | 0.024341 | 0.024341 | 0.23354
2019-03-29 14:17 | 1553883392.5773 | 1405 | 0.20951 | 0.000453 | 0 | 99.9838 | 99.7346 | 126.9907 | 21.1026 | 0.00201 | 0.001856 | -152.3022 | 0.23433 | 8.4015 | 0.002486 | 0.004245 | 0.004135 | 6e-05 | 0.001006 | 0.31598 | 1.3566 | 2.5473 | 100.16 | 20.9 | 36.974 | 70.0344 | 81.4615 | 66.3246 | 0.009135 | 1.04 | 0.024893 | 23.9361 | 0.022318 | 0.1394 | 0.00034 | 0.025439 | 0.025439 | 0.22302
2019-03-29 14:17 | 1553883393.5773 | 1406 | 0.20952 | 0.000459 | 0 | 100.0164 | 99.7107 | 127.0255 | 21.1023 | 0.001948 | 0.001851 | -152.3071 | 0.23462 | 8.4003 | 0.002471 | 0.003453 | 0.003237 | 6.2e-05 | 0.00101 | 0.31006 | 1.3211 | 2.5474 | 100.16 | 20.9 | 36.981 | 68.5453 | 81.1002 | 66.2159 | 0.009137 | 1.0403 | 0.0249 | 23.9355 | 0.027379 | 0.13958 | 0.000281 | 0.021005 | 0.021005 | 0.27044
2019-03-29 14:17 | 1553883394.5785 | 1407 | 0.20949 | 0.000451 | 1e-06 | 99.979 | 99.694 | 127.214 | 21.1017 | 0.001956 | 0.001888 | -152.301 | 0.23503 | 8.4043 | 0.002543 | 0.003719 | 0.003473 | 6e-05 | 0.001023 | 0.31758 | 1.3589 | 2.5473 | 100.16 | 20.9 | 36.985 | 68.7327 | 82.8474 | 67.1101 | 0.009138 | 1.0376 | 0.024886 | 23.9848 | 0.025532 | 0.13982 | 0.000445 | 0.033196 | 0.033196 | 0.17143
2019-03-29 14:17 | 1553883395.5775 | 1408 | 0.20951 | 0.000454 | 0 | 99.9553 | 99.6769 | 127.6346 | 21.1026 | 0.001951 | 0.001865 | -152.3154 | 0.23483 | 8.4033 | 0.002465 | 0.004289 | 0.004182 | 6.1e-05 | 0.001023 | 0.30685 | 1.376 | 2.5474 | 100.16 | 20.9 | 36.989 | 68.604 | 80.9616 | 66.554 | 0.009139 | 1.0393 | 0.024948 | 24.0048 | 0.024948 | 0.1397 | 0.000334 | 0.02505 | 0.02505 | 0.22697
2019-03-29 14:17 | 1553883396.5774 | 1409 | 0.2095 | 0.000456 | 0 | 99.9664 | 99.7424 | 127.4883 | 21.1018 | 0.001916 | 0.001802 | -152.3163 | 0.23474 | 8.4015 | 0.00247 | 0.003969 | 0.003689 | 6.3e-05 | 0.001033 | 0.31564 | 1.3706 | 2.5474 | 100.16 | 20.9 | 36.997 | 67.7796 | 81.0896 | 65.03 | 0.00914 | 1.044 | 0.024989 | 23.9372 | 0.02006 | 0.13964 | 0.000393 | 0.029491 | 0.029491 | 0.19272
2019-03-29 14:17 | 1553883397.5774 | 1410 | 0.20949 | 0.000451 | 0 | 99.9718 | 99.436 | 127.7686 | 21.1015 | 0.001891 | 0.001834 | -152.3107 | 0.23495 | 8.4027 | 0.002485 | 0.003792 | 0.003683 | 6.5e-05 | 0.001031 | 0.27527 | 1.3671 | 2.5474 | 100.16 | 20.9 | 37.001 | 67.1801 | 81.4527 | 65.803 | 0.009141 | 1.0416 | 0.024988 | 23.9909 | 0.048068 | 0.13977 | 0.000451 | 0.033866 | 0.033866 | 0.16798
2019-03-29 14:17 | 1553883398.5764 | 1411 | 0.20951 | 0.000449 | 1e-06 | 100.017 | 99.7042 | 127.5205 | 21.1022 | 0.001974 | 0.001899 | -152.3247 | 0.23529 | 8.4006 | 0.002467 | 0.004045 | 0.003827 | 6.1e-05 | 0.001028 | 0.27472 | 1.3679 | 2.5474 | 100.16 | 20.9 | 37.007 | 69.1676 | 81.0244 | 67.3626 | 0.009143 | 1.0368 | 0.024907 | 24.0226 | 0.028014 | 0.13997 | 0.000337 | 0.025161 | 0.025161 | 0.22642
2019-03-29 14:17 | 1553883399.5765 | 1412 | 0.20951 | 0.000451 | 1e-06 | 99.9742 | 99.8622 | 127.4083 | 21.1024 | 0.001961 | 0.001918 | -152.3342 | 0.23457 | 8.4017 | 0.00239 | 0.003846 | 0.003716 | 6.2e-05 | 0.001026 | 0.30198 | 1.3159 | 2.5473 | 100.16 | 20.9 | 37.01 | 68.8512 | 79.1525 | 67.8247 | 0.009144 | 1.0354 | 0.024879 | 24.0283 | 0.010021 | 0.13955 | 0.000322 | 0.024036 | 0.024036 | 0.2363
2019-03-29 14:17 | 1553883400.5776 | 1413 | 0.2095 | 0.000451 | 1e-06 | 100.0053 | 99.5223 | 127.701 | 21.1015 | 0.001966 | 0.001838 | -152.3198 | 0.23492 | 8.4 | 0.002442 | 0.003601 | 0.003438 | 6.1e-05 | 0.001025 | 0.3011 | 1.3668 | 2.5474 | 100.16 | 20.9 | 37.011 | 68.9736 | 80.4015 | 65.899 | 0.009144 | 1.0413 | 0.024978 | 23.9879 | 0.043304 | 0.13975 | 0.0004 | 0.029973 | 0.029973 | 0.18977
2019-03-29 14:17 | 1553883401.5765 | 1414 | 0.20952 | 0.000454 | 1e-06 | 99.9386 | 99.7919 | 127.9219 | 21.1022 | 0.00192 | 0.001851 | -152.3143 | 0.23526 | 8.4028 | 0.002462 | 0.003761 | 0.003544 | 6.1e-05 | 0.001024 | 0.29391 | 1.2892 | 2.5472 | 100.16 | 20.9 | 37.012 | 67.8745 | 80.9059 | 66.2119 | 0.009144 | 1.0403 | 0.024988 | 24.0197 | 0.013132 | 0.13996 | 0.000249 | 0.01867 | 0.01867 | 0.30511
2019-03-29 14:17 | 1553883402.5765 | 1415 | 0.20951 | 0.000453 | 1e-06 | 100.0239 | 99.6861 | 127.9782 | 21.1022 | 0.001935 | 0.001849 | -152.3225 | 0.23422 | 8.402 | 0.002462 | 0.003869 | 0.003733 | 6.2e-05 | 0.001025 | 0.29789 | 1.3514 | 2.5474 | 100.16 | 20.9 | 37.012 | 68.2425 | 80.8972 | 66.1744 | 0.009144 | 1.0404 | 0.024995 | 24.0237 | 0.030252 | 0.13934 | 0.000337 | 0.025242 | 0.025242 | 0.22467
2019-03-29 14:17 | 1553883403.5779 | 1416 | 0.20949 | 0.000455 | 0 | 99.9308 | 99.7366 | 128.0932 | 21.1023 | 0.001954 | 0.00186 | -152.3054 | 0.23567 | 8.4039 | 0.002458 | 0.003706 | 0.003528 | 6.4e-05 | 0.001028 | 0.31564 | 1.3475 | 2.5474 | 100.15 | 20.9 | 37.011 | 68.6919 | 80.7976 | 66.4351 | 0.009145 | 1.0396 | 0.024997 | 24.0437 | 0.017393 | 0.1402 | 0.000451 | 0.033833 | 0.033833 | 0.16866
2019-03-29 14:17 | 1553883404.5777 | 1417 | 0.20952 | 0.000448 | 0 | 99.8808 | 99.7612 | 128.0824 | 21.1021 | 0.001903 | 0.001892 | -152.3076 | 0.2344 | 8.402 | 0.002521 | 0.003917 | 0.003771 | 6e-05 | 0.001034 | 0.27417 | 1.3451 | 2.5473 | 100.15 | 20.9 | 37.011 | 67.4709 | 82.3081 | 67.2011 | 0.009145 | 1.0373 | 0.024968 | 24.0698 | 0.010714 | 0.13945 | 0.000298 | 0.022387 | 0.022387 | 0.25352
2019-03-29 14:17 | 1553883405.5771 | 1418 | 0.2095 | 0.000448 | 1e-06 | 99.9484 | 99.651 | 127.6015 | 21.1024 | 0.001962 | 0.001848 | -152.2952 | 0.23522 | 8.4014 | 0.00246 | 0.003959 | 0.003799 | 6.2e-05 | 0.001036 | 0.30917 | 1.3685 | 2.5474 | 100.15 | 20.91 | 37.01 | 68.8869 | 80.8469 | 66.1356 | 0.00915 | 1.0406 | 0.02496 | 23.9869 | 0.026659 | 0.13994 | 0.000402 | 0.030116 | 0.030116 | 0.18911
2019-03-29 14:17 | 1553883406.5777 | 1419 | 0.20953 | 0.000455 | 0 | 99.8841 | 99.7424 | 127.9508 | 21.1019 | 0.001912 | 0.001831 | -152.3052 | 0.23497 | 8.4013 | 0.002427 | 0.004068 | 0.003938 | 6.1e-05 | 0.001038 | 0.31636 | 1.3217 | 2.5474 | 100.15 | 20.91 | 37.014 | 67.6723 | 80.0518 | 65.7301 | 0.009151 | 1.0418 | 0.025009 | 24.0054 | 0.012692 | 0.13978 | 0.000225 | 0.016897 | 0.016897 | 0.3367
2019-03-29 14:17 | 1553883407.5777 | 1420 | 0.20954 | 0.000455 | 0 | 99.9547 | 99.6616 | 128.009 | 21.1025 | 0.001943 | 0.001819 | -152.3081 | 0.23549 | 8.4011 | 0.002468 | 0.003584 | 0.003392 | 6.2e-05 | 0.001046 | 0.3465 | 1.3012 | 2.5472 | 100.15 | 20.91 | 37.016 | 68.4152 | 81.0472 | 65.4347 | 0.009152 | 1.0427 | 0.025025 | 24.0004 | 0.026272 | 0.14009 | 0.00012 | 0.009032 | 0.009032 | 0.63126
2019-03-29 14:17 | 1553883408.5768 | 1421 | 0.20954 | 0.000453 | 0 | 99.9655 | 99.8441 | 127.966 | 21.1023 | 0.001962 | 0.001945 | -152.3256 | 0.23597 | 8.403 | 0.002427 | 0.004378 | 0.004101 | 6.3e-05 | 0.001051 | 0.29518 | 1.3426 | 2.5474 | 100.15 | 20.9 | 37.016 | 68.8682 | 80.0545 | 68.479 | 0.009146 | 1.0334 | 0.024909 | 24.104 | 0.010864 | 0.14038 | 0.000139 | 0.010385 | 0.010385 | 0.55018
2019-03-29 14:17 | 1553883409.5767 | 1422 | 0.20955 | 0.000453 | 0 | 99.9718 | 99.6701 | 127.6285 | 21.102 | 0.001945 | 0.001878 | -152.3204 | 0.23531 | 8.4033 | 0.002458 | 0.003348 | 0.003254 | 6.1e-05 | 0.001055 | 0.29806 | 1.3726 | 2.5473 | 100.15 | 20.91 | 37.021 | 68.4724 | 80.8022 | 66.8639 | 0.009153 | 1.0383 | 0.024936 | 24.0152 | 0.027033 | 0.13999 | 7.3e-05 | 0.005477 | 0.005477 | 1.0403
2019-03-29 14:17 | 1553883410.5768 | 1423 | 0.20953 | 0.000452 | 0 | 99.8883 | 99.4878 | 127.4808 | 21.1024 | 0.001867 | 0.001892 | -152.3213 | 0.23507 | 8.4024 | 0.002406 | 0.003477 | 0.003263 | 6.3e-05 | 0.001053 | 0.29142 | 1.3414 | 2.5473 | 100.15 | 20.91 | 37.024 | 66.5882 | 79.537 | 67.1858 | 0.009154 | 1.0373 | 0.024909 | 24.0127 | 0.035928 | 0.13985 | 0.000194 | 0.014487 | 0.014487 | 0.39289
2019-03-29 14:17 | 1553883411.5772 | 1424 | 0.20954 | 0.000454 | 0 | 99.9754 | 99.7015 | 127.5449 | 21.1027 | 0.001919 | 0.001886 | -152.3124 | 0.23515 | 8.4023 | 0.002403 | 0.003728 | 0.003628 | 6.3e-05 | 0.001053 | 0.29142 | 1.3338 | 2.5473 | 100.15 | 20.91 | 37.029 | 67.8555 | 79.4859 | 67.0426 | 0.009155 | 1.0378 | 0.024921 | 24.0136 | 0.024537 | 0.13989 | 0.000132 | 0.009895 | 0.009895 | 0.57539
2019-03-29 14:17 | 1553883412.5769 | 1425 | 0.20955 | 0.000451 | 0 | 99.8616 | 99.5872 | 127.5346 | 21.1023 | 0.001875 | 0.001848 | -152.2954 | 0.2344 | 8.403 | 0.002376 | 0.003674 | 0.00344 | 6.2e-05 | 0.001057 | 0.2756 | 1.3595 | 2.5473 | 100.15 | 20.91 | 37.033 | 66.8002 | 78.8165 | 66.1307 | 0.009156 | 1.0406 | 0.024953 | 23.9805 | 0.024618 | 0.13944 | 5.9e-05 | 0.004365 | 0.004365 | 1.3003
2019-03-29 14:17 | 1553883413.5773 | 1426 | 0.20952 | 0.00045 | 1e-06 | 99.982 | 99.8261 | 127.2987 | 21.103 | 0.001891 | 0.001812 | -152.2976 | 0.23604 | 8.4018 | 0.002367 | 0.003741 | 0.003558 | 6.2e-05 | 0.001055 | 0.30807 | 1.3435 | 2.5474 | 100.15 | 20.91 | 37.036 | 67.1693 | 78.6136 | 65.2627 | 0.009157 | 1.0432 | 0.024962 | 23.9276 | 0.013958 | 0.14042 | 0.00024 | 0.017983 | 0.017983 | 0.31781
2019-03-29 14:17 | 1553883414.577 | 1427 | 0.20954 | 0.000457 | 1e-06 | 99.9906 | 99.6325 | 127.2762 | 21.1035 | 0.001898 | 0.001877 | -152.3064 | 0.23556 | 8.4013 | 0.00241 | 0.00255 | 0.002368 | 6.2e-05 | 0.001052 | 0.28528 | 1.3457 | 2.5474 | 100.15 | 20.91 | 37.043 | 67.3422 | 79.644 | 66.8459 | 0.009158 | 1.0384 | 0.024902 | 23.9814 | 0.03209 | 0.14013 | 0.000144 | 0.01077 | 0.01077 | 0.52956
2019-03-29 14:17 | 1553883415.576 | 1428 | 0.20954 | 0.000452 | 1e-06 | 99.9754 | 99.8001 | 127.5862 | 21.1028 | 0.001918 | 0.001826 | -152.2981 | 0.23547 | 8.4029 | 0.002418 | 0.002855 | 0.002661 | 6e-05 | 0.001054 | 0.30198 | 1.3371 | 2.5475 | 100.15 | 20.91 | 37.046 | 67.8329 | 79.8452 | 65.5991 | 0.009159 | 1.0422 | 0.024978 | 23.9665 | 0.015693 | 0.14008 | 0.000156 | 0.011649 | 0.011649 | 0.48941
2019-03-29 14:17 | 1553883416.577 | 1429 | 0.20952 | 0.000458 | 0 | 99.9745 | 99.5657 | 127.6419 | 21.1032 | 0.001892 | 0.001837 | -152.2935 | 0.23512 | 8.4022 | 0.002446 | 0.004011 | 0.003829 | 6e-05 | 0.001052 | 0.30182 | 1.3682 | 2.5474 | 100.15 | 20.9 | 37.053 | 67.2067 | 80.517 | 65.8725 | 0.009155 | 1.0414 | 0.024973 | 23.9814 | 0.036651 | 0.13987 | 0.000255 | 0.01917 | 0.01917 | 0.29697
2019-03-29 14:17 | 1553883417.5772 | 1430 | 0.20956 | 0.00045 | 0 | 99.8984 | 99.4913 | 127.4374 | 21.1032 | 0.00193 | 0.001838 | -152.2841 | 0.23586 | 8.4006 | 0.002425 | 0.003965 | 0.003775 | 6.3e-05 | 0.001049 | 0.28567 | 1.3346 | 2.5472 | 100.15 | 20.9 | 37.054 | 68.1031 | 80.0133 | 65.8889 | 0.009155 | 1.0413 | 0.024953 | 23.9628 | 0.036533 | 0.14031 | 1.4e-05 | 0.001003 | 0.001003 | 5.6915
2019-03-29 14:17 | 1553883418.5761 | 1431 | 0.20957 | 0.000455 | 0 | 99.9814 | 99.8076 | 127.5031 | 21.1038 | 0.001919 | 0.001781 | -152.2682 | 0.23504 | 8.4026 | 0.002367 | 0.003249 | 0.003136 | 6.2e-05 | 0.001055 | 0.30702 | 1.3972 | 2.5476 | 100.15 | 20.9 | 37.055 | 67.8503 | 78.6106 | 64.5401 | 0.009156 | 1.0455 | 0.025009 | 23.9212 | 0.015556 | 0.13983 | -5e-05 | -0.003786 | -0.003786 | -1.5032
2019-03-29 14:17 | 1553883419.5771 | 1432 | 0.20957 | 0.000453 | 1e-06 | 99.9296 | 99.5523 | 127.7695 | 21.1031 | 0.001938 | 0.001806 | -152.2699 | 0.23466 | 8.3991 | 0.002463 | 0.004144 | 0.003917 | 6.3e-05 | 0.001048 | 0.29214 | 1.3732 | 2.5474 | 100.15 | 20.9 | 37.058 | 68.2973 | 80.921 | 65.1286 | 0.009156 | 1.0437 | 0.025013 | 23.9671 | 0.033829 | 0.1396 | -1.2e-05 | -0.000899 | -0.000899 | -6.3165
2019-03-29 14:17 | 1553883420.5772 | 1433 | 0.20952 | 0.000453 | 0 | 99.8922 | 99.6482 | 127.6374 | 21.1026 | 0.001925 | 0.001834 | -152.2798 | 0.23463 | 8.4052 | 0.002469 | 0.00357 | 0.00344 | 6.2e-05 | 0.001045 | 0.29894 | 1.3238 | 2.5473 | 100.15 | 20.9 | 37.059 | 67.9871 | 81.0536 | 65.8065 | 0.009157 | 1.0416 | 0.024975 | 23.9787 | 0.021867 | 0.13958 | 0.000288 | 0.021617 | 0.021617 | 0.2628
2019-03-29 14:17 | 1553883421.5761 | 1434 | 0.20958 | 0.000454 | 0 | 99.964 | 99.6537 | 127.3231 | 21.1036 | 0.001908 | 0.001813 | -152.2939 | 0.23416 | 8.402 | 0.002439 | 0.003903 | 0.003721 | 6.5e-05 | 0.001047 | 0.30071 | 1.334 | 2.5473 | 100.15 | 20.9 | 37.061 | 67.595 | 80.349 | 65.2939 | 0.009157 | 1.0431 | 0.024963 | 23.931 | 0.027807 | 0.1393 | -7.2e-05 | -0.005453 | -0.005453 | -1.0398
2019-03-29 14:17 | 1553883422.5772 | 1435 | 0.20952 | 0.000455 | 1e-06 | 99.9179 | 99.5756 | 126.9621 | 21.103 | 0.001991 | 0.001871 | -152.2992 | 0.23489 | 8.4024 | 0.002515 | 0.003547 | 0.003449 | 6.3e-05 | 0.001046 | 0.31758 | 1.2974 | 2.5474 | 100.15 | 20.9 | 37.062 | 69.5748 | 82.1739 | 66.7004 | 0.009157 | 1.0388 | 0.024876 | 23.9467 | 0.030699 | 0.13974 | 0.000277 | 0.020624 | 0.020624 | 0.27576
2019-03-29 14:17 | 1553883423.5763 | 1436 | 0.20957 | 0.000451 | 1e-06 | 99.9035 | 99.5469 | 126.5572 | 21.1026 | 0.00197 | 0.001838 | -152.3001 | 0.23599 | 8.4009 | 0.002503 | 0.004184 | 0.004012 | 6.2e-05 | 0.001048 | 0.32173 | 1.3782 | 2.5473 | 100.15 | 20.9 | 37.062 | 69.0845 | 81.8882 | 65.9063 | 0.009157 | 1.0413 | 0.024866 | 23.8805 | 0.031988 | 0.14039 | -5.4e-05 | -0.004144 | -0.004144 | -1.3788
2019-03-29 14:17 | 1553883424.5762 | 1437 | 0.2096 | 0.000454 | 1e-06 | 99.867 | 99.5063 | 126.2471 | 21.1033 | 0.00194 | 0.001911 | -152.2794 | 0.23456 | 8.4028 | 0.002408 | 0.003449 | 0.003206 | 6.2e-05 | 0.001042 | 0.31741 | 1.3507 | 2.5473 | 100.15 | 20.9 | 37.063 | 68.3496 | 79.5885 | 67.6445 | 0.009158 | 1.0359 | 0.024772 | 23.9123 | 0.032368 | 0.13954 | -0.000233 | -0.017392 | -0.017392 | -0.32654
2019-03-29 14:17 | 1553883425.5764 | 1438 | 0.20961 | 0.000455 | 1e-06 | 99.9218 | 99.5568 | 125.965 | 21.103 | 0.001943 | 0.001861 | -152.2843 | 0.23517 | 8.4016 | 0.002489 | 0.004189 | 0.003954 | 6.5e-05 | 0.00104 | 0.28655 | 1.3238 | 2.5474 | 100.15 | 20.9 | 37.063 | 68.4135 | 81.5376 | 66.4536 | 0.009158 | 1.0396 | 0.024788 | 23.8438 | 0.032734 | 0.1399 | -0.000258 | -0.01928 | -0.01928 | -0.29532
2019-03-29 14:17 | 1553883426.5773 | 1439 | 0.20957 | 0.000454 | 1e-06 | 99.9409 | 99.5063 | 125.508 | 21.1038 | 0.001886 | 0.001856 | -152.2785 | 0.23433 | 8.4017 | 0.002445 | 0.003554 | 0.003505 | 6.1e-05 | 0.001043 | 0.30525 | 1.3337 | 2.5474 | 100.15 | 20.9 | 37.061 | 67.0607 | 80.4808 | 66.332 | 0.009157 | 1.04 | 0.024747 | 23.7962 | 0.038988 | 0.13941 | -3.4e-05 | -0.002568 | -0.002568 | -2.2092
2019-03-29 14:17 | 1553883427.5774 | 1440 | 0.20953 | 0.000455 | 1e-06 | 99.9269 | 99.5831 | 125.6622 | 21.1027 | 0.00193 | 0.001897 | -152.279 | 0.23411 | 8.402 | 0.002527 | 0.002936 | 0.002732 | 6.2e-05 | 0.001041 | 0.29822 | 1.3778 | 2.5475 | 100.15 | 20.9 | 37.061 | 68.1096 | 82.4537 | 67.3148 | 0.009157 | 1.037 | 0.024726 | 23.8453 | 0.030828 | 0.13927 | 0.000174 | 0.012855 | 0.012855 | 0.44093
2019-03-29 14:17 | 1553883428.5774 | 1441 | 0.20957 | 0.000451 | 1e-06 | 99.9371 | 99.6892 | 126.026 | 21.1035 | 0.001949 | 0.001835 | -152.2702 | 0.23448 | 8.4023 | 0.002499 | 0.003171 | 0.003077 | 6.5e-05 | 0.00104 | 0.30901 | 1.3835 | 2.5475 | 100.15 | 20.9 | 37.057 | 68.559 | 81.7744 | 65.83 | 0.009156 | 1.0415 | 0.024816 | 23.8276 | 0.022211 | 0.13949 | -1e-06 | -0.000135 | -0.000135 | -42.0095
2019-03-29 14:17 | 1553883429.5764 | 1442 | 0.20954 | 0.000454 | 1e-06 | 99.9251 | 99.6209 | 126.0863 | 21.1039 | 0.001891 | 0.001841 | -152.2657 | 0.23415 | 8.4022 | 0.002524 | 0.003909 | 0.003673 | 6.1e-05 | 0.001037 | 0.28888 | 1.4078 | 2.5475 | 100.15 | 20.9 | 37.053 | 67.1747 | 82.3908 | 65.9764 | 0.009155 | 1.041 | 0.024817 | 23.8385 | 0.027266 | 0.1393 | 0.000142 | 0.010553 | 0.010553 | 0.53724
2019-03-29 14:17 | 1553883430.5765 | 1443 | 0.20955 | 0.000452 | 1e-06 | 99.9332 | 99.6814 | 125.7434 | 21.1039 | 0.001889 | 0.001862 | -152.2691 | 0.23463 | 8.4 | 0.002487 | 0.003612 | 0.003415 | 6.3e-05 | 0.001032 | 0.30973 | 1.3702 | 2.5474 | 100.15 | 20.9 | 37.05 | 67.117 | 81.5065 | 66.4726 | 0.009155 | 1.0395 | 0.024765 | 23.8235 | 0.022567 | 0.13958 | 8.1e-05 | 0.00597 | 0.00597 | 0.95162
2019-03-29 14:17 | 1553883431.5765 | 1444 | 0.20954 | 0.000447 | 1e-06 | 99.8957 | 99.5148 | 125.4955 | 21.1043 | 0.00195 | 0.001791 | -152.2513 | 0.23491 | 8.4014 | 0.002446 | 0.004123 | 0.003923 | 6.7e-05 | 0.001031 | 0.31669 | 1.3513 | 2.5475 | 100.15 | 20.9 | 37.041 | 68.5861 | 80.5094 | 64.7701 | 0.009152 | 1.0448 | 0.024803 | 23.7402 | 0.034175 | 0.13975 | 0.000145 | 0.010726 | 0.010726 | 0.53028
2019-03-29 14:17 | 1553883432.5755 | 1445 | 0.20951 | 0.000453 | 1e-06 | 99.8823 | 99.5759 | 125.3871 | 21.103 | 0.001892 | 0.001876 | -152.2591 | 0.23445 | 8.4023 | 0.002524 | 0.003912 | 0.003759 | 6.6e-05 | 0.00103 | 0.31022 | 1.3068 | 2.5475 | 100.15 | 20.9 | 37.037 | 67.2013 | 82.3775 | 66.8065 | 0.009151 | 1.0385 | 0.024718 | 23.8014 | 0.027479 | 0.13947 | 0.000318 | 0.023605 | 0.023605 | 0.24049
2019-03-29 14:17 | 1553883433.5766 | 1446 | 0.20955 | 0.000451 | 1e-06 | 99.9113 | 99.7527 | 125.6934 | 21.103 | 0.001887 | 0.001838 | -152.2611 | 0.23399 | 8.4058 | 0.002522 | 0.003936 | 0.003721 | 6.3e-05 | 0.001036 | 0.29159 | 1.3242 | 2.5475 | 100.15 | 20.9 | 37.025 | 67.0787 | 82.3441 | 65.8935 | 0.009148 | 1.0413 | 0.024781 | 23.7984 | 0.014213 | 0.1392 | 6.3e-05 | 0.004626 | 0.004626 | 1.2247
2019-03-29 14:17 | 1553883434.5768 | 1447 | 0.20952 | 0.000445 | 1e-06 | 99.9475 | 99.7155 | 125.8695 | 21.1042 | 0.001909 | 0.001887 | -152.2668 | 0.23446 | 8.4015 | 0.002484 | 0.004135 | 0.003946 | 6.7e-05 | 0.001034 | 0.2959 | 1.383 | 2.5475 | 100.15 | 20.9 | 37.019 | 67.6126 | 81.4224 | 67.0767 | 0.009147 | 1.0377 | 0.024755 | 23.8566 | 0.020791 | 0.13948 | 0.000255 | 0.018885 | 0.018885 | 0.30061
2019-03-29 14:17 | 1553883435.5757 | 1448 | 0.20952 | 0.000449 | 1e-06 | 99.9137 | 99.6462 | 125.9845 | 21.1039 | 0.001882 | 0.001832 | -152.2645 | 0.23468 | 8.4012 | 0.002494 | 0.00394 | 0.00378 | 6.2e-05 | 0.001042 | 0.29518 | 1.3288 | 2.5474 | 100.15 | 20.9 | 37.014 | 66.9578 | 81.6538 | 65.7476 | 0.009146 | 1.0417 | 0.024815 | 23.8208 | 0.02398 | 0.13961 | 0.000249 | 0.01852 | 0.01852 | 0.30682
2019-03-29 14:17 | 1553883436.5756 | 1449 | 0.20952 | 0.00045 | 1e-06 | 99.8835 | 99.6264 | 126.019 | 21.1036 | 0.001948 | 0.001812 | -152.2623 | 0.23448 | 8.401 | 0.002471 | 0.003323 | 0.00323 | 6.1e-05 | 0.001051 | 0.28047 | 1.3339 | 2.5473 | 100.15 | 20.9 | 37 | 68.5434 | 81.1216 | 65.2814 | 0.009142 | 1.0432 | 0.024836 | 23.8077 | 0.02305 | 0.13949 | 0.000272 | 0.020193 | 0.020193 | 0.28116
2019-03-29 14:17 | 1553883437.5756 | 1450 | 0.20953 | 0.000452 | 1e-06 | 99.8775 | 99.5848 | 125.8638 | 21.1033 | 0.001927 | 0.001847 | -152.2772 | 0.23445 | 8.4025 | 0.002523 | 0.003526 | 0.003331 | 6.6e-05 | 0.001057 | 0.3194 | 1.3438 | 2.5473 | 100.15 | 20.9 | 36.994 | 68.0431 | 82.3708 | 66.1101 | 0.009141 | 1.0406 | 0.02479 | 23.8221 | 0.026253 | 0.13947 | 0.000218 | 0.016199 | 0.016199 | 0.35043
2019-03-29 14:17 | 1553883438.5767 | 1451 | 0.20952 | 0.000451 | 1e-06 | 99.9526 | 99.5892 | 125.7272 | 21.1034 | 0.00188 | 0.00183 | -152.2758 | 0.23444 | 8.4028 | 0.002473 | 0.003962 | 0.003797 | 6.3e-05 | 0.001054 | 0.30198 | 1.3393 | 2.5473 | 100.15 | 20.9 | 36.981 | 66.9067 | 81.1527 | 65.7128 | 0.009137 | 1.0419 | 0.024791 | 23.7953 | 0.03258 | 0.13947 | 0.000253 | 0.018791 | 0.018791 | 0.30209
2019-03-29 14:17 | 1553883439.5758 | 1452 | 0.20953 | 0.000447 | 1e-06 | 99.8754 | 99.6247 | 125.6882 | 21.1028 | 0.001897 | 0.001852 | -152.2875 | 0.23496 | 8.4023 | 0.00252 | 0.003992 | 0.003776 | 6.6e-05 | 0.001055 | 0.29679 | 1.3397 | 2.5474 | 100.15 | 20.9 | 36.974 | 67.3298 | 82.2884 | 66.2289 | 0.009136 | 1.0403 | 0.024768 | 23.8097 | 0.02248 | 0.13978 | 0.000205 | 0.015247 | 0.015247 | 0.37313
2019-03-29 14:17 | 1553883440.5758 | 1453 | 0.20955 | 0.000449 | 1e-06 | 99.9254 | 99.4762 | 126.1715 | 21.1036 | 0.001954 | 0.001836 | -152.279 | 0.23392 | 8.4023 | 0.00247 | 0.003927 | 0.003802 | 6e-05 | 0.001056 | 0.29463 | 1.338 | 2.5474 | 100.15 | 20.9 | 36.968 | 68.6893 | 81.082 | 65.8471 | 0.009134 | 1.0414 | 0.02483 | 23.842 | 0.040294 | 0.13916 | 9e-05 | 0.006644 | 0.006644 | 0.85242
2019-03-29 14:17 | 1553883441.5758 | 1454 | 0.20955 | 0.000446 | 1e-06 | 99.8928 | 99.8342 | 126.5077 | 21.1026 | 0.001961 | 0.001829 | -152.2893 | 0.23519 | 8.4031 | 0.002361 | 0.004358 | 0.004162 | 6.2e-05 | 0.00105 | 0.3147 | 1.3434 | 2.5474 | 100.15 | 20.9 | 36.954 | 68.8576 | 78.4715 | 65.6737 | 0.009131 | 1.042 | 0.024869 | 23.8676 | 0.00524 | 0.13991 | 7.7e-05 | 0.005677 | 0.005677 | 1.003
2019-03-29 14:17 | 1553883442.5758 | 1455 | 0.20954 | 0.000451 | 1e-06 | 99.8963 | 99.6421 | 126.6452 | 21.1041 | 0.00195 | 0.001803 | -152.2875 | 0.23453 | 8.4018 | 0.002486 | 0.004006 | 0.003921 | 6.1e-05 | 0.001048 | 0.32173 | 1.3119 | 2.5475 | 100.15 | 20.9 | 36.95 | 68.6009 | 81.4731 | 65.0695 | 0.00913 | 1.0438 | 0.024905 | 23.8593 | 0.022793 | 0.13953 | 0.000162 | 0.012026 | 0.012026 | 0.47219
2019-03-29 14:17 | 1553883443.5769 | 1456 | 0.20951 | 0.000452 | 1e-06 | 99.9089 | 99.7138 | 126.6422 | 21.1044 | 0.001937 | 0.001846 | -152.2884 | 0.23436 | 8.4039 | 0.002517 | 0.004075 | 0.003779 | 6.3e-05 | 0.001055 | 0.32388 | 1.3837 | 2.5475 | 100.15 | 20.9 | 36.945 | 68.2814 | 82.2269 | 66.093 | 0.009129 | 1.0407 | 0.024867 | 23.8951 | 0.017488 | 0.13942 | 0.00033 | 0.024606 | 0.024606 | 0.23061
2019-03-29 14:17 | 1553883444.5749 | 1457 | 0.20954 | 0.000446 | 1e-06 | 99.92 | 99.6363 | 126.5809 | 21.1034 | 0.001928 | 0.001839 | -152.2827 | 0.23528 | 8.4035 | 0.002518 | 0.004279 | 0.004126 | 6.2e-05 | 0.001053 | 0.30989 | 1.39 | 2.5475 | 100.15 | 20.9 | 36.934 | 68.0751 | 82.2485 | 65.9251 | 0.009126 | 1.0412 | 0.024867 | 23.8834 | 0.025431 | 0.13997 | 0.000144 | 0.010738 | 0.010738 | 0.5305
2019-03-29 14:17 | 1553883445.5763 | 1458 | 0.20952 | 0.000451 | 1e-06 | 99.9559 | 99.623 | 126.2255 | 21.1035 | 0.001908 | 0.001829 | -152.2831 | 0.23539 | 8.3996 | 0.002417 | 0.004454 | 0.004211 | 6.4e-05 | 0.001054 | 0.29717 | 1.3689 | 2.5474 | 100.15 | 20.9 | 36.93 | 67.5793 | 79.8188 | 65.6771 | 0.009125 | 1.042 | 0.024842 | 23.8411 | 0.029841 | 0.14003 | 0.000278 | 0.020754 | 0.020754 | 0.27462
2019-03-29 14:17 | 1553883446.5759 | 1459 | 0.20952 | 0.000447 | 1e-06 | 99.8149 | 99.5684 | 126.0859 | 21.1034 | 0.001921 | 0.001887 | -152.2816 | 0.23459 | 8.4015 | 0.002475 | 0.003627 | 0.003514 | 6.2e-05 | 0.00105 | 0.30735 | 1.3422 | 2.5474 | 100.15 | 20.9 | 36.923 | 67.9069 | 81.219 | 67.0706 | 0.009123 | 1.0377 | 0.024777 | 23.8769 | 0.02212 | 0.13956 | 0.000259 | 0.01921 | 0.01921 | 0.29567
2019-03-29 14:17 | 1553883447.575 | 1460 | 0.20954 | 0.000447 | 1e-06 | 99.8574 | 99.3735 | 126.3072 | 21.1039 | 0.00197 | 0.001852 | -152.2685 | 0.23464 | 8.4022 | 0.002494 | 0.004345 | 0.00412 | 6.3e-05 | 0.001044 | 0.32228 | 1.3538 | 2.5473 | 100.15 | 20.9 | 36.92 | 69.0663 | 81.6531 | 66.2392 | 0.009122 | 1.0402 | 0.024829 | 23.8686 | 0.043452 | 0.13959 | 0.000175 | 0.012993 | 0.012993 | 0.43724
2019-03-29 14:17 | 1553883448.575 | 1461 | 0.20954 | 0.000452 | 1e-06 | 99.8311 | 99.5977 | 126.2511 | 21.1044 | 0.001933 | 0.001765 | -152.2782 | 0.23531 | 8.4006 | 0.002529 | 0.003757 | 0.003563 | 5.9e-05 | 0.001043 | 0.31653 | 1.3724 | 2.5475 | 100.15 | 20.9 | 36.917 | 68.1956 | 82.506 | 64.1368 | 0.009122 | 1.0467 | 0.024901 | 23.7893 | 0.020932 | 0.13999 | 0.000129 | 0.009611 | 0.009611 | 0.59283
2019-03-29 14:17 | 1553883449.5756 | 1462 | 0.20953 | 0.000446 | 1e-06 | 99.8622 | 99.4779 | 126.1488 | 21.1043 | 0.001934 | 0.00184 | -152.2768 | 0.23414 | 8.4036 | 0.002522 | 0.003712 | 0.003606 | 6e-05 | 0.001035 | 0.29856 | 1.3898 | 2.5475 | 100.15 | 20.9 | 36.911 | 68.2025 | 82.3391 | 65.947 | 0.00912 | 1.0411 | 0.024824 | 23.8434 | 0.034486 | 0.13929 | 0.0002 | 0.014893 | 0.014893 | 0.38066
2019-03-29 14:17 | 1553883450.5751 | 1463 | 0.20954 | 0.000449 | 1e-06 | 99.8134 | 99.496 | 126.0293 | 21.1044 | 0.001996 | 0.001844 | -152.2654 | 0.2348 | 8.4025 | 0.002482 | 0.00336 | 0.003129 | 5.9e-05 | 0.001042 | 0.31708 | 1.3614 | 2.5475 | 100.15 | 20.9 | 36.908 | 69.703 | 81.3875 | 66.0336 | 0.009119 | 1.0409 | 0.024809 | 23.8351 | 0.028488 | 0.13968 | 0.000155 | 0.01153 | 0.01153 | 0.49305
2019-03-29 14:18 | 1553883451.5751 | 1464 | 0.20955 | 0.000453 | 1e-06 | 99.8538 | 99.6991 | 126.1093 | 21.1048 | 0.001915 | 0.001899 | -152.2595 | 0.23453 | 8.4018 | 0.002529 | 0.003772 | 0.00358 | 6.3e-05 | 0.001046 | 0.30359 | 1.3518 | 2.5474 | 100.15 | 20.91 | 36.904 | 67.7467 | 82.4994 | 67.3762 | 0.009124 | 1.0368 | 0.024768 | 23.8898 | 0.013872 | 0.13953 | 9.5e-05 | 0.007005 | 0.007005 | 0.81066
2019-03-29 14:18 | 1553883452.5752 | 1465 | 0.20954 | 0.00045 | 1e-06 | 99.8155 | 99.609 | 126.3251 | 21.1048 | 0.001936 | 0.001809 | -152.247 | 0.23569 | 8.4014 | 0.002484 | 0.003751 | 0.003619 | 6.1e-05 | 0.00105 | 0.2959 | 1.3246 | 2.5474 | 100.15 | 20.91 | 36.902 | 68.2506 | 81.4202 | 65.1984 | 0.009124 | 1.0434 | 0.024869 | 23.8337 | 0.018526 | 0.14021 | 0.000182 | 0.013545 | 0.013545 | 0.4213
2019-03-29 14:18 | 1553883453.5752 | 1466 | 0.20955 | 0.000447 | 1e-06 | 99.8257 | 99.6841 | 126.4518 | 21.105 | 0.001932 | 0.001844 | -152.2652 | 0.23474 | 8.4025 | 0.002554 | 0.003638 | 0.003362 | 6.4e-05 | 0.001058 | 0.3027 | 1.3059 | 2.5474 | 100.15 | 20.91 | 36.9 | 68.1712 | 83.112 | 66.0414 | 0.009123 | 1.0408 | 0.02485 | 23.8753 | 0.012699 | 0.13964 | 9.8e-05 | 0.007207 | 0.007207 | 0.78857
2019-03-29 14:18 | 1553883454.5754 | 1467 | 0.20956 | 0.000451 | 1e-06 | 99.8814 | 99.4964 | 126.4171 | 21.1049 | 0.001937 | 0.001816 | -152.2827 | 0.235 | 8.4029 | 0.002569 | 0.003841 | 0.003732 | 6.4e-05 | 0.001055 | 0.27815 | 1.3226 | 2.5475 | 100.15 | 20.91 | 36.893 | 68.2725 | 83.4621 | 65.3779 | 0.009121 | 1.0429 | 0.024871 | 23.8487 | 0.034545 | 0.1398 | 6.1e-05 | 0.004499 | 0.004499 | 1.2648
2019-03-29 14:18 | 1553883455.5753 | 1468 | 0.20955 | 0.000449 | 1e-06 | 99.8685 | 99.4411 | 126.4457 | 21.104 | 0.001901 | 0.001824 | -152.2869 | 0.23527 | 8.3993 | 0.00248 | 0.004109 | 0.003892 | 6.3e-05 | 0.001051 | 0.30486 | 1.3507 | 2.5474 | 100.15 | 20.91 | 36.889 | 67.4049 | 81.3179 | 65.5673 | 0.00912 | 1.0423 | 0.024867 | 23.858 | 0.038363 | 0.13996 | 8.1e-05 | 0.00595 | 0.00595 | 0.95742
2019-03-29 14:18 | 1553883456.5763 | 1469 | 0.20956 | 0.000452 | 1e-06 | 99.8371 | 99.5643 | 126.4795 | 21.1039 | 0.001961 | 0.001876 | -152.2834 | 0.23502 | 8.4055 | 0.002544 | 0.004414 | 0.004151 | 5.7e-05 | 0.001054 | 0.29319 | 1.328 | 2.5474 | 100.15 | 20.91 | 36.886 | 68.8589 | 82.876 | 66.8244 | 0.00912 | 1.0384 | 0.024825 | 23.9055 | 0.024473 | 0.13981 | 4.4e-05 | 0.003215 | 0.003215 | 1.7697
2019-03-29 14:18 | 1553883457.5767 | 1470 | 0.20956 | 0.000448 | 1e-06 | 99.8811 | 99.6377 | 126.1044 | 21.1038 | 0.001933 | 0.001804 | -152.3013 | 0.23511 | 8.4044 | 0.002566 | 0.004127 | 0.004023 | 6.4e-05 | 0.001056 | 0.30397 | 1.3516 | 2.5474 | 100.15 | 20.91 | 36.88 | 68.1826 | 83.3931 | 65.081 | 0.009118 | 1.0438 | 0.024852 | 23.8087 | 0.021825 | 0.13987 | 4.5e-05 | 0.003301 | 0.003301 | 1.7248
2019-03-29 14:18 | 1553883458.5764 | 1471 | 0.20957 | 0.000451 | 1e-06 | 99.817 | 99.4844 | 125.7267 | 21.1045 | 0.001882 | 0.001815 | -152.2935 | 0.23511 | 8.4026 | 0.002568 | 0.004431 | 0.004225 | 6.2e-05 | 0.001061 | 0.28871 | 1.3513 | 2.5474 | 100.15 | 20.91 | 36.877 | 66.9639 | 83.4327 | 65.3407 | 0.009117 | 1.043 | 0.024805 | 23.7821 | 0.029853 | 0.13986 | -2.1e-05 | -0.001591 | -0.001591 | -3.5775
2019-03-29 14:18 | 1553883459.5754 | 1472 | 0.20958 | 0.000448 | 1e-06 | 99.8365 | 99.5459 | 125.6249 | 21.1047 | 0.001962 | 0.001831 | -152.2896 | 0.23535 | 8.4017 | 0.002568 | 0.003219 | 0.003074 | 5.8e-05 | 0.001063 | 0.29015 | 1.3438 | 2.5474 | 100.15 | 20.91 | 36.868 | 68.8729 | 83.454 | 65.7345 | 0.009115 | 1.0418 | 0.02478 | 23.7863 | 0.026075 | 0.14001 | -7.3e-05 | -0.005534 | -0.005534 | -1.0297
2019-03-29 14:18 | 1553883460.5754 | 1473 | 0.20957 | 0.000447 | 1e-06 | 99.8685 | 99.5428 | 125.4686 | 21.1039 | 0.001922 | 0.001852 | -152.3 | 0.23502 | 8.4023 | 0.002581 | 0.003875 | 0.003682 | 6.3e-05 | 0.001063 | 0.28385 | 1.3312 | 2.5475 | 100.15 | 20.91 | 36.863 | 67.9283 | 83.7678 | 66.2448 | 0.009114 | 1.0402 | 0.024746 | 23.7894 | 0.02922 | 0.13981 | -9e-06 | -0.000741 | -0.000741 | -7.6823
2019-03-29 14:18 | 1553883461.5749 | 1474 | 0.20959 | 0.000453 | 1e-06 | 99.8488 | 99.494 | 125.419 | 21.1045 | 0.001936 | 0.001803 | -152.2989 | 0.23492 | 8.4021 | 0.002551 | 0.00402 | 0.003825 | 6.5e-05 | 0.001066 | 0.28528 | 1.3258 | 2.5473 | 100.15 | 20.91 | 36.859 | 68.2659 | 83.0256 | 65.0628 | 0.009113 | 1.0439 | 0.024785 | 23.7433 | 0.031837 | 0.13975 | -0.000132 | -0.009872 | -0.009872 | -0.57617
2019-03-29 14:18 | 1553883462.5745 | 1475 | 0.20956 | 0.000446 | 1e-06 | 99.7937 | 99.4407 | 125.438 | 21.1045 | 0.001904 | 0.00181 | -152.3004 | 0.23544 | 8.4029 | 0.002452 | 0.00385 | 0.003664 | 6.4e-05 | 0.001072 | 0.29983 | 1.3246 | 2.5474 | 100.15 | 20.91 | 36.85 | 67.4937 | 80.6563 | 65.2246 | 0.009111 | 1.0434 | 0.024781 | 23.7508 | 0.031689 | 0.14007 | 3.2e-05 | 0.002329 | 0.002329 | 2.4479
2019-03-29 14:18 | 1553883463.5745 | 1476 | 0.20955 | 0.000446 | 1e-06 | 99.8688 | 99.6049 | 125.5026 | 21.1041 | 0.001902 | 0.001795 | -152.2928 | 0.2352 | 8.4026 | 0.002536 | 0.004067 | 0.003927 | 6.1e-05 | 0.001073 | 0.30287 | 1.3168 | 2.5474 | 100.16 | 20.91 | 36.844 | 67.4472 | 82.6736 | 64.865 | 0.009108 | 1.0445 | 0.0248 | 23.7442 | 0.023667 | 0.13992 | 0.000134 | 0.009969 | 0.009969 | 0.57123
2019-03-29 14:18 | 1553883464.5755 | 1477 | 0.2096 | 0.000453 | 0 | 99.8649 | 99.5332 | 125.3495 | 21.1046 | 0.001947 | 0.001823 | -152.2822 | 0.23599 | 8.4003 | 0.00247 | 0.003887 | 0.003665 | 6.3e-05 | 0.001065 | 0.31078 | 1.3318 | 2.5475 | 100.16 | 20.91 | 36.834 | 68.5187 | 81.083 | 65.5359 | 0.009106 | 1.0424 | 0.02476 | 23.7533 | 0.029757 | 0.14039 | -0.000173 | -0.012897 | -0.012897 | -0.44304
2019-03-29 14:18 | 1553883465.6146 | 1478 | 0.20955 | 0.000451 | 1e-06 | 99.8464 | 99.4223 | 125.5246 | 21.1041 | 0.001989 | 0.001814 | -152.2937 | 0.23486 | 8.4047 | 0.002586 | 0.003702 | 0.003544 | 6.2e-05 | 0.001058 | 0.2943 | 1.3586 | 2.5474 | 100.16 | 20.91 | 36.83 | 69.522 | 83.8873 | 65.3278 | 0.009105 | 1.043 | 0.024785 | 23.7626 | 0.038068 | 0.13972 | 9.7e-05 | 0.007166 | 0.007166 | 0.79358
2019-03-29 14:18 | 1553883466.5618 | 1479 | 0.20953 | 0.000451 | 1e-06 | 99.8317 | 99.4834 | 125.5965 | 21.1049 | 0.001907 | 0.001841 | -152.3063 | 0.2349 | 8.402 | 0.002561 | 0.004594 | 0.004352 | 6.3e-05 | 0.001056 | 0.31061 | 1.3267 | 2.5473 | 100.16 | 20.91 | 36.824 | 67.565 | 83.2773 | 65.9663 | 0.009103 | 1.0411 | 0.024769 | 23.7918 | 0.03126 | 0.13974 | 0.000221 | 0.016423 | 0.016423 | 0.34631
2019-03-29 14:18 | 1553883467.5618 | 1480 | 0.20958 | 0.000449 | 1e-06 | 99.797 | 99.4988 | 125.7446 | 21.1051 | 0.001937 | 0.001821 | -152.2946 | 0.23538 | 8.4045 | 0.002568 | 0.003794 | 0.003674 | 6.4e-05 | 0.001058 | 0.29645 | 1.3329 | 2.5474 | 100.16 | 20.91 | 36.816 | 68.2735 | 83.4521 | 65.4982 | 0.009101 | 1.0425 | 0.024801 | 23.7894 | 0.026767 | 0.14003 | -0.000105 | -0.007862 | -0.007862 | -0.72491
2019-03-29 14:18 | 1553883468.5619 | 1481 | 0.20952 | 0.000447 | 1e-06 | 99.8332 | 99.5517 | 126.0978 | 21.1048 | 0.001909 | 0.001825 | -152.3041 | 0.23453 | 8.3996 | 0.002562 | 0.004034 | 0.003852 | 6.3e-05 | 0.00106 | 0.31868 | 1.3496 | 2.5475 | 100.16 | 20.91 | 36.811 | 67.5986 | 83.3019 | 65.5941 | 0.0091 | 1.0422 | 0.024832 | 23.8261 | 0.025259 | 0.13952 | 0.000286 | 0.021259 | 0.021259 | 0.26711
2019-03-29 14:18 | 1553883469.5609 | 1482 | 0.20955 | 0.000449 | 1e-06 | 99.882 | 99.293 | 126.2807 | 21.1051 | 0.001929 | 0.00182 | -152.3025 | 0.23429 | 8.4035 | 0.002543 | 0.004085 | 0.003857 | 6.5e-05 | 0.001059 | 0.29789 | 1.3463 | 2.5474 | 100.16 | 20.91 | 36.802 | 68.0854 | 82.8538 | 65.4567 | 0.009098 | 1.0426 | 0.024855 | 23.8386 | 0.052903 | 0.13938 | 0.000109 | 0.008059 | 0.008059 | 0.70389
2019-03-29 14:18 | 1553883470.5609 | 1483 | 0.20954 | 0.000451 | 1e-06 | 99.7796 | 99.5503 | 126.4506 | 21.1055 | 0.001947 | 0.001821 | -152.2975 | 0.2351 | 8.3999 | 0.002551 | 0.004549 | 0.004497 | 6e-05 | 0.00106 | 0.31581 | 1.3133 | 2.5473 | 100.16 | 20.91 | 36.798 | 68.5258 | 83.0249 | 65.5009 | 0.009097 | 1.0425 | 0.02487 | 23.8562 | 0.02058 | 0.13986 | 0.000139 | 0.010342 | 0.010342 | 0.5504
2019-03-29 14:18 | 1553883471.5619 | 1484 | 0.20959 | 0.000452 | 1e-06 | 99.7557 | 99.5902 | 126.3225 | 21.1043 | 0.001891 | 0.001819 | -152.3086 | 0.23474 | 8.403 | 0.002556 | 0.003854 | 0.003599 | 6.4e-05 | 0.001056 | 0.30359 | 1.3548 | 2.5474 | 100.16 | 20.91 | 36.794 | 67.179 | 83.1641 | 65.4476 | 0.009096 | 1.0427 | 0.02486 | 23.8422 | 0.014846 | 0.13964 | -0.000122 | -0.009118 | -0.009118 | -0.62331
2019-03-29 14:18 | 1553883472.562 | 1485 | 0.20958 | 0.000453 | 1e-06 | 99.7625 | 99.4875 | 125.7932 | 21.1041 | 0.001901 | 0.001803 | -152.3206 | 0.23502 | 8.4022 | 0.002502 | 0.003836 | 0.003718 | 6.1e-05 | 0.001054 | 0.29662 | 1.3696 | 2.5474 | 100.16 | 20.91 | 36.788 | 67.4193 | 81.867 | 65.0524 | 0.009094 | 1.0439 | 0.024822 | 23.7783 | 0.024694 | 0.13981 | -0.000101 | -0.007584 | -0.007584 | -0.75035
2019-03-29 14:18 | 1553883473.561 | 1486 | 0.20957 | 0.00045 | 1e-06 | 99.8021 | 99.6523 | 125.3575 | 21.1036 | 0.001939 | 0.001829 | -152.3137 | 0.23535 | 8.4014 | 0.002473 | 0.004103 | 0.003907 | 6.1e-05 | 0.00106 | 0.30879 | 1.3694 | 2.5474 | 100.16 | 20.91 | 36.786 | 68.3196 | 81.1653 | 65.6864 | 0.009094 | 1.0419 | 0.024756 | 23.7593 | 0.013428 | 0.14001 | -3.9e-05 | -0.00299 | -0.00299 | -1.9061
2019-03-29 14:18 | 1553883474.561 | 1487 | 0.20957 | 0.000453 | 1e-06 | 99.8164 | 99.5274 | 125.5749 | 21.1037 | 0.001912 | 0.001833 | -152.327 | 0.2352 | 8.4043 | 0.002502 | 0.004127 | 0.003891 | 5.9e-05 | 0.001059 | 0.28639 | 1.3927 | 2.5475 | 100.16 | 20.91 | 36.783 | 67.6824 | 81.8486 | 65.7903 | 0.009093 | 1.0416 | 0.024773 | 23.7836 | 0.025934 | 0.13992 | -3.4e-05 | -0.002609 | -0.002609 | -2.1829
2019-03-29 14:18 | 1553883475.5605 | 1488 | 0.20957 | 0.000449 | 1e-06 | 99.7946 | 99.4728 | 125.6796 | 21.1051 | 0.001937 | 0.001826 | -152.3179 | 0.23528 | 8.4025 | 0.002574 | 0.004335 | 0.004183 | 6.5e-05 | 0.001069 | 0.31542 | 1.3447 | 2.5475 | 100.16 | 20.91 | 36.776 | 68.2899 | 83.5949 | 65.6202 | 0.009092 | 1.0421 | 0.02479 | 23.7875 | 0.028885 | 0.13997 | -1.2e-05 | -0.00095 | -0.00095 | -5.9954
2019-03-29 14:18 | 1553883476.5601 | 1489 | 0.20955 | 0.000449 | 2e-06 | 99.8538 | 99.4541 | 125.957 | 21.1042 | 0.001902 | 0.00182 | -152.3146 | 0.23524 | 8.4024 | 0.002628 | 0.003857 | 0.003672 | 6.4e-05 | 0.001067 | 0.29645 | 1.3261 | 2.5474 | 100.16 | 20.91 | 36.776 | 67.4434 | 84.8837 | 65.4747 | 0.009092 | 1.0426 | 0.024823 | 23.8086 | 0.035883 | 0.13995 | 9.3e-05 | 0.006848 | 0.006848 | 0.83172
2019-03-29 14:18 | 1553883477.5601 | 1490 | 0.20958 | 0.000454 | 1e-06 | 99.8077 | 99.4718 | 125.8969 | 21.1032 | 0.001961 | 0.001836 | -152.3159 | 0.2343 | 8.4023 | 0.002593 | 0.003682 | 0.003583 | 6.1e-05 | 0.001068 | 0.31581 | 1.3617 | 2.5474 | 100.16 | 20.92 | 36.773 | 68.8659 | 84.0369 | 65.8412 | 0.009096 | 1.0415 | 0.024803 | 23.8158 | 0.030157 | 0.13939 | -7.8e-05 | -0.005893 | -0.005893 | -0.96269
2019-03-29 14:18 | 1553883478.5612 | 1491 | 0.20956 | 0.000451 | 1e-06 | 99.7736 | 99.6155 | 125.8385 | 21.1041 | 0.001942 | 0.00181 | -152.3084 | 0.23478 | 8.4021 | 0.002486 | 0.004212 | 0.003925 | 6.3e-05 | 0.001069 | 0.2912 | 1.3397 | 2.5474 | 100.16 | 20.92 | 36.77 | 68.4011 | 81.4723 | 65.2366 | 0.009096 | 1.0433 | 0.02482 | 23.7891 | 0.014188 | 0.13967 | 2.3e-05 | 0.001677 | 0.001677 | 3.3905
2019-03-29 14:18 | 1553883479.5602 | 1492 | 0.20954 | 0.000446 | 1e-06 | 99.7769 | 99.6486 | 125.6225 | 21.105 | 0.001956 | 0.00181 | -152.3096 | 0.23542 | 8.4029 | 0.002568 | 0.003926 | 0.003804 | 6.2e-05 | 0.001071 | 0.28313 | 1.3815 | 2.5475 | 100.15 | 20.92 | 36.769 | 68.7398 | 83.4561 | 65.2203 | 0.009096 | 1.0434 | 0.024799 | 23.7681 | 0.011511 | 0.14005 | 0.000171 | 0.012712 | 0.012712 | 0.4484
2019-03-29 14:18 | 1553883480.5612 | 1493 | 0.20953 | 0.000451 | 1e-06 | 99.7886 | 99.5012 | 125.3758 | 21.1047 | 0.001934 | 0.001863 | -152.3153 | 0.23492 | 8.4011 | 0.002531 | 0.003781 | 0.00355 | 6.4e-05 | 0.001063 | 0.27273 | 1.3707 | 2.5474 | 100.15 | 20.92 | 36.765 | 68.2091 | 82.5452 | 66.5115 | 0.009095 | 1.0394 | 0.024727 | 23.79 | 0.025801 | 0.13975 | 0.000235 | 0.017453 | 0.017453 | 0.3259
2019-03-29 14:18 | 1553883481.5618 | 1494 | 0.2095 | 0.000448 | 1e-06 | 99.8221 | 99.5752 | 125.4014 | 21.1044 | 0.001937 | 0.001813 | -152.3115 | 0.23495 | 8.4033 | 0.002548 | 0.003562 | 0.00348 | 6.4e-05 | 0.001064 | 0.27887 | 1.367 | 2.5474 | 100.15 | 20.92 | 36.763 | 68.2821 | 82.9608 | 65.3033 | 0.009095 | 1.0431 | 0.024774 | 23.75 | 0.022151 | 0.13977 | 0.000419 | 0.03111 | 0.03111 | 0.18286
2019-03-29 14:18 | 1553883482.5617 | 1495 | 0.20958 | 0.000449 | 1e-06 | 99.7994 | 99.5203 | 125.7634 | 21.1045 | 0.001898 | 0.001855 | -152.3103 | 0.23528 | 8.4032 | 0.002585 | 0.003996 | 0.003744 | 6.3e-05 | 0.001058 | 0.29679 | 1.3509 | 2.5474 | 100.15 | 20.92 | 36.757 | 67.3463 | 83.8619 | 66.3022 | 0.009093 | 1.04 | 0.024773 | 23.8194 | 0.025049 | 0.13997 | -5.9e-05 | -0.004456 | -0.004456 | -1.2784
2019-03-29 14:18 | 1553883483.5613 | 1496 | 0.20957 | 0.000451 | 1e-06 | 99.7874 | 99.5909 | 126.1356 | 21.1055 | 0.001915 | 0.00181 | -152.3159 | 0.23543 | 8.4036 | 0.002594 | 0.00393 | 0.003766 | 6.3e-05 | 0.001053 | 0.29087 | 1.3335 | 2.5475 | 100.15 | 20.92 | 36.754 | 67.7529 | 84.0759 | 65.2207 | 0.009093 | 1.0434 | 0.024849 | 23.8166 | 0.017629 | 0.14006 | -2e-06 | -0.00023 | -0.00023 | -24.8011
2019-03-29 14:18 | 1553883484.5613 | 1497 | 0.20958 | 0.000454 | 1e-06 | 99.8637 | 99.4745 | 126.2429 | 21.1041 | 0.001902 | 0.001795 | -152.3111 | 0.23516 | 8.402 | 0.002509 | 0.003274 | 0.003141 | 6.8e-05 | 0.00105 | 0.31797 | 1.3562 | 2.5474 | 100.15 | 20.92 | 36.75 | 67.4301 | 82.0346 | 64.8542 | 0.009092 | 1.0445 | 0.024873 | 23.8138 | 0.034927 | 0.1399 | -0.000103 | -0.007787 | -0.007787 | -0.73122
2019-03-29 14:18 | 1553883485.5604 | 1498 | 0.20955 | 0.000454 | 1e-06 | 99.8508 | 99.5186 | 126.2377 | 21.1034 | 0.001897 | 0.001842 | -152.3003 | 0.23524 | 8.4008 | 0.002559 | 0.004038 | 0.003807 | 6.3e-05 | 0.001052 | 0.31006 | 1.3316 | 2.5472 | 100.15 | 20.92 | 36.745 | 67.3102 | 83.221 | 65.9892 | 0.00909 | 1.041 | 0.024831 | 23.8532 | 0.029816 | 0.13995 | 0.000118 | 0.008748 | 0.008748 | 0.65108
2019-03-29 14:18 | 1553883486.5614 | 1499 | 0.20954 | 0.000453 | 1e-06 | 99.7718 | 99.5264 | 126.1232 | 21.1048 | 0.001947 | 0.001817 | -152.2996 | 0.2342 | 8.4035 | 0.00251 | 0.00388 | 0.00381 | 6.7e-05 | 0.001059 | 0.29999 | 1.3217 | 2.5473 | 100.15 | 20.92 | 36.743 | 68.5206 | 82.0432 | 65.3949 | 0.00909 | 1.0428 | 0.024842 | 23.8215 | 0.022029 | 0.13933 | 0.00017 | 0.012621 | 0.012621 | 0.4493
2019-03-29 14:18 | 1553883487.5604 | 1500 | 0.20951 | 0.000452 | 1e-06 | 99.7533 | 99.4585 | 126.1873 | 21.104 | 0.001919 | 0.001838 | -152.2924 | 0.23435 | 8.3999 | 0.00264 | 0.003306 | 0.003067 | 6.6e-05 | 0.00106 | 0.26504 | 1.3604 | 2.5474 | 100.15 | 20.92 | 36.739 | 67.8582 | 85.1754 | 65.8889 | 0.009089 | 1.0413 | 0.02483 | 23.845 | 0.026471 | 0.13941 | 0.000326 | 0.024205 | 0.024205 | 0.23442
2019-03-29 14:18 | 1553883488.5615 | 1501 | 0.20952 | 0.00045 | 2e-06 | 99.7847 | 99.5404 | 126.0455 | 21.104 | 0.001918 | 0.001774 | -152.2941 | 0.23509 | 8.4045 | 0.00252 | 0.003469 | 0.003294 | 6.4e-05 | 0.001067 | 0.29231 | 1.4056 | 2.5476 | 100.15 | 20.92 | 36.737 | 67.8332 | 82.2819 | 64.3536 | 0.009088 | 1.046 | 0.024872 | 23.7775 | 0.021925 | 0.13985 | 0.000305 | 0.022699 | 0.022699 | 0.25076
2019-03-29 14:18 | 1553883489.5605 | 1502 | 0.20952 | 0.000448 | 1e-06 | 99.8479 | 99.4831 | 125.9169 | 21.1036 | 0.001885 | 0.001794 | -152.3179 | 0.23455 | 8.4037 | 0.002549 | 0.003964 | 0.003817 | 6.5e-05 | 0.001069 | 0.30342 | 1.3622 | 2.5474 | 100.15 | 20.92 | 36.736 | 67.0326 | 82.9983 | 64.8373 | 0.009088 | 1.0446 | 0.024842 | 23.7824 | 0.032738 | 0.13953 | 0.000301 | 0.022392 | 0.022392 | 0.25361
2019-03-29 14:18 | 1553883490.5605 | 1503 | 0.20952 | 0.00045 | 1e-06 | 99.8413 | 99.5107 | 125.8887 | 21.1044 | 0.001946 | 0.001811 | -152.3213 | 0.23464 | 8.4001 | 0.002574 | 0.003925 | 0.003686 | 6.5e-05 | 0.001068 | 0.30237 | 1.358 | 2.5475 | 100.15 | 20.92 | 36.733 | 68.499 | 83.6004 | 65.2516 | 0.009087 | 1.0433 | 0.024824 | 23.7943 | 0.029663 | 0.13959 | 0.000293 | 0.021767 | 0.021767 | 0.261
2019-03-29 14:18 | 1553883491.5595 | 1504 | 0.20953 | 0.000452 | 1e-06 | 99.8634 | 99.5169 | 125.7352 | 21.1038 | 0.001895 | 0.001787 | -152.3273 | 0.23537 | 8.3997 | 0.002559 | 0.003449 | 0.003316 | 6.9e-05 | 0.001065 | 0.29806 | 1.3347 | 2.5475 | 100.15 | 20.92 | 36.731 | 67.2768 | 83.2363 | 64.6827 | 0.009087 | 1.045 | 0.02483 | 23.7598 | 0.031096 | 0.14002 | 0.000194 | 0.014383 | 0.014383 | 0.39622
2019-03-29 14:18 | 1553883492.5596 | 1505 | 0.20952 | 0.000449 | 1e-06 | 99.8978 | 99.51 | 125.6859 | 21.1034 | 0.001923 | 0.001857 | -152.332 | 0.23421 | 8.4014 | 0.002494 | 0.00436 | 0.004228 | 6.4e-05 | 0.001068 | 0.29717 | 1.3762 | 2.5475 | 100.15 | 20.92 | 36.729 | 67.942 | 81.6596 | 66.3528 | 0.009086 | 1.0399 | 0.024764 | 23.8138 | 0.034792 | 0.13933 | 0.000285 | 0.021147 | 0.021147 | 0.26815
2019-03-29 14:18 | 1553883493.56 | 1506 | 0.20952 | 0.000453 | 1e-06 | 99.832 | 99.6612 | 125.5502 | 21.1044 | 0.001959 | 0.001821 | -152.3232 | 0.23427 | 8.4007 | 0.002521 | 0.003496 | 0.003282 | 6.3e-05 | 0.001068 | 0.31852 | 1.3409 | 2.5473 | 100.15 | 20.93 | 36.726 | 68.8009 | 82.3197 | 65.4811 | 0.009091 | 1.0426 | 0.024782 | 23.7704 | 0.015314 | 0.13937 | 0.000314 | 0.02335 | 0.02335 | 0.24293
2019-03-29 14:18 | 1553883494.5607 | 1507 | 0.20956 | 0.000452 | 1e-06 | 99.8173 | 99.5547 | 125.4427 | 21.1034 | 0.001879 | 0.001808 | -152.327 | 0.23445 | 8.4045 | 0.00261 | 0.003355 | 0.003214 | 6.3e-05 | 0.001067 | 0.30453 | 1.3191 | 2.5474 | 100.15 | 20.93 | 36.724 | 66.8958 | 84.45 | 65.1751 | 0.009091 | 1.0435 | 0.024783 | 23.7495 | 0.023561 | 0.13947 | 7.3e-05 | 0.005361 | 0.005361 | 1.0588
2019-03-29 14:18 | 1553883495.5607 | 1508 | 0.20952 | 0.000451 | 1e-06 | 99.8931 | 99.5827 | 125.3226 | 21.1036 | 0.001889 | 0.00179 | -152.3264 | 0.23554 | 8.4008 | 0.002484 | 0.004293 | 0.004005 | 6.2e-05 | 0.001069 | 0.28799 | 1.3467 | 2.5474 | 100.15 | 20.93 | 36.719 | 67.1348 | 81.4236 | 64.7507 | 0.00909 | 1.0448 | 0.024786 | 23.7232 | 0.027831 | 0.14012 | 0.00031 | 0.023076 | 0.023076 | 0.24714
2019-03-29 14:18 | 1553883496.5611 | 1509 | 0.20956 | 0.000452 | 1e-06 | 99.8329 | 99.5964 | 125.2866 | 21.1044 | 0.002006 | 0.001839 | -152.3323 | 0.23497 | 8.4034 | 0.002546 | 0.003784 | 0.003631 | 6.1e-05 | 0.001067 | 0.3027 | 1.3743 | 2.5474 | 100.15 | 20.93 | 36.717 | 69.9309 | 82.9241 | 65.9189 | 0.009089 | 1.0412 | 0.02474 | 23.7608 | 0.021216 | 0.13978 | 3.1e-05 | 0.002258 | 0.002258 | 2.5191
2019-03-29 14:18 | 1553883497.5598 | 1510 | 0.20955 | 0.000454 | 1e-06 | 99.8607 | 99.5506 | 125.1815 | 21.1036 | 0.001986 | 0.001852 | -152.3176 | 0.23522 | 8.4018 | 0.002601 | 0.003959 | 0.00378 | 6.1e-05 | 0.001067 | 0.31996 | 1.323 | 2.5474 | 100.15 | 20.93 | 36.713 | 69.4677 | 84.248 | 66.2342 | 0.009088 | 1.0403 | 0.024718 | 23.7618 | 0.027818 | 0.13994 | 8.5e-05 | 0.00622 | 0.00622 | 0.91561
2019-03-29 14:18 | 1553883498.5608 | 1511 | 0.20954 | 0.000446 | 1e-06 | 99.8188 | 99.5213 | 125.1528 | 21.1046 | 0.001966 | 0.001795 | -152.3216 | 0.23615 | 8.4021 | 0.002541 | 0.003804 | 0.003623 | 6.2e-05 | 0.001062 | 0.33627 | 1.3512 | 2.5474 | 100.15 | 20.93 | 36.704 | 68.9882 | 82.7901 | 64.8772 | 0.009086 | 1.0444 | 0.024765 | 23.7116 | 0.026701 | 0.14048 | 0.000169 | 0.012515 | 0.012515 | 0.45687
2019-03-29 14:18 | 1553883499.5608 | 1512 | 0.20955 | 0.000454 | 1e-06 | 99.7853 | 99.4984 | 125.3549 | 21.104 | 0.001903 | 0.001842 | -152.2993 | 0.23537 | 8.4029 | 0.00254 | 0.004229 | 0.004005 | 6.3e-05 | 0.001065 | 0.29878 | 1.3112 | 2.5475 | 100.15 | 20.93 | 36.699 | 67.4564 | 82.7825 | 66.0068 | 0.009085 | 1.041 | 0.024744 | 23.7703 | 0.025752 | 0.14002 | 8e-05 | 0.005897 | 0.005897 | 0.96645
2019-03-29 14:18 | 1553883500.5603 | 1513 | 0.20957 | 0.000451 | 1e-06 | 99.8086 | 99.5523 | 125.3265 | 21.1038 | 0.001954 | 0.001827 | -152.3002 | 0.23509 | 8.4008 | 0.002547 | 0.003587 | 0.003387 | 6.5e-05 | 0.001072 | 0.29927 | 1.303 | 2.5474 | 100.15 | 20.93 | 36.69 | 68.6826 | 82.9422 | 65.646 | 0.009082 | 1.0421 | 0.024754 | 23.755 | 0.022998 | 0.13986 | 1.3e-05 | 0.000888 | 0.000888 | 6.4081
2019-03-29 14:18 | 1553883501.5609 | 1514 | 0.20948 | 0.000447 | 1e-06 | 99.8482 | 99.6465 | 125.3451 | 21.1057 | 0.001911 | 0.001844 | -152.2921 | 0.23498 | 8.4002 | 0.002509 | 0.003157 | 0.003042 | 6.3e-05 | 0.001077 | 0.29734 | 1.3004 | 2.5472 | 100.15 | 20.93 | 36.685 | 67.6499 | 82.0329 | 66.0474 | 0.009081 | 1.0408 | 0.024741 | 23.7708 | 0.018079 | 0.13979 | 0.000525 | 0.038942 | 0.038942 | 0.1461
2019-03-29 14:18 | 1553883502.5599 | 1515 | 0.20953 | 0.000453 | 1e-06 | 99.8167 | 99.5111 | 125.3838 | 21.1038 | 0.001886 | 0.001814 | -152.2808 | 0.23492 | 8.4032 | 0.00261 | 0.003982 | 0.003735 | 6.6e-05 | 0.001072 | 0.31061 | 1.3094 | 2.5473 | 100.15 | 20.94 | 36.681 | 67.0638 | 84.4526 | 65.33 | 0.009086 | 1.043 | 0.024771 | 23.7493 | 0.027433 | 0.13976 | 0.000235 | 0.017442 | 0.017442 | 0.32611
2019-03-29 14:18 | 1553883503.5599 | 1516 | 0.20949 | 0.000447 | 2e-06 | 99.7997 | 99.4936 | 125.5096 | 21.1053 | 0.001919 | 0.001875 | -152.2988 | 0.23597 | 8.4013 | 0.002513 | 0.003793 | 0.003629 | 6.1e-05 | 0.001071 | 0.31564 | 1.3198 | 2.5473 | 100.15 | 20.94 | 36.673 | 67.857 | 82.1256 | 66.7993 | 0.009084 | 1.0385 | 0.02473 | 23.8127 | 0.027469 | 0.14038 | 0.000498 | 0.036907 | 0.036907 | 0.1548
2019-03-29 14:18 | 1553883504.561 | 1517 | 0.2096 | 0.000451 | 1e-06 | 99.8197 | 99.5575 | 125.3334 | 21.1041 | 0.001951 | 0.001848 | -152.2983 | 0.23598 | 8.4001 | 0.00254 | 0.00369 | 0.00356 | 6.3e-05 | 0.001069 | 0.33052 | 1.3219 | 2.5475 | 100.15 | 20.94 | 36.67 | 68.6266 | 82.767 | 66.131 | 0.009083 | 1.0406 | 0.024737 | 23.7726 | 0.023531 | 0.14038 | -0.000203 | -0.01519 | -0.01519 | -0.37615
2019-03-29 14:18 | 1553883505.5601 | 1518 | 0.20953 | 0.000448 | 1e-06 | 99.8332 | 99.5667 | 125.4186 | 21.104 | 0.001894 | 0.001848 | -152.3095 | 0.23497 | 8.4028 | 0.002568 | 0.003476 | 0.003203 | 6.5e-05 | 0.001076 | 0.34075 | 1.3247 | 2.5474 | 100.15 | 20.94 | 36.665 | 67.2349 | 83.4533 | 66.1489 | 0.009082 | 1.0405 | 0.024745 | 23.7813 | 0.02391 | 0.13978 | 0.00025 | 0.018508 | 0.018508 | 0.30739
2019-03-29 14:18 | 1553883506.5601 | 1519 | 0.20948 | 0.000453 | 1e-06 | 99.8039 | 99.4558 | 125.3491 | 21.105 | 0.001929 | 0.001804 | -152.3133 | 0.23448 | 8.4028 | 0.002556 | 0.004564 | 0.004327 | 6.3e-05 | 0.001075 | 0.30613 | 1.3401 | 2.5474 | 100.15 | 20.94 | 36.662 | 68.0923 | 83.1648 | 65.0821 | 0.009081 | 1.0438 | 0.024777 | 23.7373 | 0.031251 | 0.13949 | 0.000545 | 0.040524 | 0.040524 | 0.1401
2019-03-29 14:18 | 1553883507.5601 | 1520 | 0.20953 | 0.00045 | 1e-06 | 99.803 | 99.4186 | 125.2838 | 21.105 | 0.001992 | 0.001823 | -152.3479 | 0.23498 | 8.4008 | 0.002543 | 0.003715 | 0.003536 | 6.2e-05 | 0.001074 | 0.33085 | 1.3287 | 2.5475 | 100.15 | 20.94 | 36.66 | 69.6 | 82.8552 | 65.5332 | 0.009081 | 1.0424 | 0.024754 | 23.747 | 0.034517 | 0.13979 | 0.00021 | 0.015523 | 0.015523 | 0.36651
2019-03-29 14:18 | 1553883508.5596 | 1521 | 0.20955 | 0.000454 | 1e-06 | 99.8751 | 99.4844 | 125.3704 | 21.105 | 0.001973 | 0.001828 | -152.3589 | 0.23443 | 8.403 | 0.002584 | 0.004228 | 0.004095 | 6.1e-05 | 0.001074 | 0.31974 | 1.3295 | 2.5473 | 100.15 | 20.94 | 36.657 | 69.1399 | 83.8256 | 65.6635 | 0.00908 | 1.042 | 0.024758 | 23.7597 | 0.035056 | 0.13946 | 8.2e-05 | 0.00604 | 0.00604 | 0.9397
2019-03-29 14:18 | 1553883509.5592 | 1522 | 0.20954 | 0.000458 | 1e-06 | 99.8221 | 99.5547 | 125.3383 | 21.1039 | 0.001934 | 0.001819 | -152.3667 | 0.23492 | 8.4016 | 0.002548 | 0.003904 | 0.003615 | 6e-05 | 0.001079 | 0.31797 | 1.3563 | 2.5474 | 100.15 | 20.94 | 36.656 | 68.2115 | 82.9657 | 65.4434 | 0.00908 | 1.0427 | 0.024763 | 23.749 | 0.023991 | 0.13975 | 0.000164 | 0.012116 | 0.012116 | 0.46947
2019-03-29 14:18 | 1553883510.5602 | 1523 | 0.20951 | 0.000457 | 1e-06 | 99.8326 | 99.5042 | 125.4324 | 21.1052 | 0.001966 | 0.001805 | -152.3677 | 0.23457 | 8.404 | 0.002534 | 0.003947 | 0.003833 | 6.2e-05 | 0.001075 | 0.28871 | 1.3386 | 2.5474 | 100.15 | 20.94 | 36.654 | 68.9842 | 82.634 | 65.1119 | 0.009079 | 1.0437 | 0.024784 | 23.7463 | 0.029469 | 0.13954 | 0.00036 | 0.026765 | 0.026765 | 0.21219
2019-03-29 14:18 | 1553883511.5592 | 1524 | 0.20953 | 0.000454 | 1e-06 | 99.7997 | 99.408 | 125.9115 | 21.1043 | 0.001878 | 0.001862 | -152.3729 | 0.23604 | 8.402 | 0.002533 | 0.003408 | 0.00324 | 6.3e-05 | 0.001073 | 0.33124 | 1.3447 | 2.5474 | 100.15 | 20.95 | 36.653 | 66.8584 | 82.5931 | 66.4753 | 0.009084 | 1.0395 | 0.024781 | 23.8395 | 0.035174 | 0.14042 | 0.00022 | 0.016319 | 0.016319 | 0.35021
2019-03-29 14:19 | 1553883512.5593 | 1525 | 0.20951 | 0.000453 | 1e-06 | 99.7575 | 99.5759 | 126.134 | 21.1042 | 0.001891 | 0.001851 | -152.3743 | 0.23539 | 8.4024 | 0.002611 | 0.003523 | 0.003288 | 6.2e-05 | 0.00107 | 0.33931 | 1.3465 | 2.5474 | 100.15 | 20.95 | 36.651 | 67.1859 | 84.4888 | 66.2146 | 0.009084 | 1.0403 | 0.024813 | 23.8514 | 0.016294 | 0.14003 | 0.000359 | 0.026735 | 0.026735 | 0.21318
2019-03-29 14:19 | 1553883513.5603 | 1526 | 0.20954 | 0.00045 | 1e-06 | 99.8012 | 99.4363 | 126.0197 | 21.1049 | 0.001895 | 0.001813 | -152.3612 | 0.23604 | 8.4019 | 0.002525 | 0.003707 | 0.003532 | 6e-05 | 0.001069 | 0.30773 | 1.3075 | 2.5474 | 100.15 | 20.95 | 36.651 | 67.2796 | 82.4012 | 65.3076 | 0.009084 | 1.0431 | 0.024835 | 23.8087 | 0.032759 | 0.14042 | 0.000175 | 0.012975 | 0.012975 | 0.44049
2019-03-29 14:19 | 1553883514.5603 | 1527 | 0.20951 | 0.000453 | 1e-06 | 99.7871 | 99.5083 | 126.0917 | 21.105 | 0.001917 | 0.001796 | -152.362 | 0.235 | 8.4028 | 0.002494 | 0.003808 | 0.003646 | 6.1e-05 | 0.001073 | 0.3027 | 1.3401 | 2.5474 | 100.15 | 20.95 | 36.65 | 67.7976 | 81.666 | 64.88 | 0.009084 | 1.0444 | 0.024858 | 23.8004 | 0.025023 | 0.1398 | 0.000355 | 0.026426 | 0.026426 | 0.21532
2019-03-29 14:19 | 1553883515.5603 | 1528 | 0.20954 | 0.000455 | 1e-06 | 99.7889 | 99.5513 | 126.3319 | 21.1056 | 0.001944 | 0.001864 | -152.3528 | 0.23546 | 8.4056 | 0.002522 | 0.003844 | 0.003616 | 6.2e-05 | 0.001071 | 0.31813 | 1.3134 | 2.5474 | 100.15 | 20.95 | 36.65 | 68.4494 | 82.3386 | 66.535 | 0.009084 | 1.0393 | 0.024821 | 23.8813 | 0.021319 | 0.14007 | 0.000159 | 0.011759 | 0.011759 | 0.48483
2019-03-29 14:19 | 1553883516.5606 | 1529 | 0.20954 | 0.000455 | 1e-06 | 99.7979 | 99.5097 | 126.3239 | 21.1055 | 0.001893 | 0.001822 | -152.3552 | 0.23546 | 8.4006 | 0.002535 | 0.003994 | 0.003775 | 6.3e-05 | 0.001069 | 0.25962 | 1.3014 | 2.5472 | 100.15 | 20.95 | 36.649 | 67.2191 | 82.6552 | 65.5214 | 0.009083 | 1.0424 | 0.024857 | 23.8449 | 0.025866 | 0.14007 | 0.000179 | 0.01334 | 0.01334 | 0.42737
2019-03-29 14:19 | 1553883517.5594 | 1530 | 0.20957 | 0.000451 | 1e-06 | 99.8119 | 99.5298 | 126.7454 | 21.1053 | 0.001841 | 0.001861 | -152.3317 | 0.2352 | 8.4031 | 0.002528 | 0.00396 | 0.003776 | 6.1e-05 | 0.001066 | 0.27936 | 1.3573 | 2.5473 | 100.15 | 20.95 | 36.648 | 65.9736 | 82.4741 | 66.4433 | 0.009083 | 1.0396 | 0.024865 | 23.9172 | 0.025317 | 0.13992 | -2.6e-05 | -0.002011 | -0.002011 | -2.8318
2019-03-29 14:19 | 1553883518.5604 | 1531 | 0.20956 | 0.000451 | 1e-06 | 99.8272 | 99.6479 | 126.9001 | 21.1056 | 0.001921 | 0.001788 | -152.305 | 0.23554 | 8.4003 | 0.002528 | 0.004227 | 0.004026 | 6.1e-05 | 0.001068 | 0.33395 | 1.3326 | 2.5475 | 100.15 | 20.95 | 36.648 | 67.9035 | 82.4838 | 64.6907 | 0.009083 | 1.045 | 0.024944 | 23.8699 | 0.016079 | 0.14012 | 3.7e-05 | 0.002717 | 0.002717 | 2.0993
2019-03-29 14:19 | 1553883519.5595 | 1532 | 0.20956 | 0.000455 | 1e-06 | 99.7248 | 99.3083 | 126.9271 | 21.105 | 0.001818 | 0.001841 | -152.3051 | 0.23494 | 8.4034 | 0.002536 | 0.003872 | 0.003773 | 6e-05 | 0.001064 | 0.29927 | 1.3476 | 2.5474 | 100.15 | 20.95 | 36.648 | 65.4124 | 82.6688 | 65.9726 | 0.009083 | 1.0411 | 0.0249 | 23.9177 | 0.037436 | 0.13976 | 7.5e-05 | 0.005569 | 0.005569 | 1.0214
2019-03-29 14:19 | 1553883520.5595 | 1533 | 0.20952 | 0.000453 | 1e-06 | 99.7718 | 99.5111 | 126.8018 | 21.1055 | 0.001917 | 0.001826 | -152.3179 | 0.23559 | 8.3999 | 0.002468 | 0.00386 | 0.003694 | 6e-05 | 0.00107 | 0.29336 | 1.285 | 2.5474 | 100.15 | 20.95 | 36.647 | 67.812 | 81.0439 | 65.6017 | 0.009083 | 1.0422 | 0.024901 | 23.8928 | 0.023409 | 0.14015 | 0.000281 | 0.020969 | 0.020969 | 0.27202
2019-03-29 14:19 | 1553883521.5589 | 1534 | 0.20958 | 0.000454 | 1e-06 | 99.7389 | 99.4636 | 126.2091 | 21.1048 | 0.001933 | 0.001813 | -152.3026 | 0.23533 | 8.4045 | 0.002534 | 0.003795 | 0.003595 | 6e-05 | 0.001073 | 0.30215 | 1.3826 | 2.5474 | 100.16 | 20.95 | 36.647 | 68.1741 | 82.628 | 65.2902 | 0.009082 | 1.0432 | 0.024854 | 23.8259 | 0.024722 | 0.14 | -8.4e-05 | -0.006374 | -0.006374 | -0.89394
2019-03-29 14:19 | 1553883522.5597 | 1535 | 0.20957 | 0.000455 | 1e-06 | 99.7682 | 99.4636 | 126.1924 | 21.1046 | 0.001921 | 0.001809 | -152.3089 | 0.23548 | 8.4029 | 0.00244 | 0.00381 | 0.003595 | 5.8e-05 | 0.001075 | 0.2912 | 1.3786 | 2.5475 | 100.16 | 20.95 | 36.649 | 67.9071 | 80.3655 | 65.1919 | 0.009082 | 1.0435 | 0.024856 | 23.8209 | 0.027352 | 0.14009 | -2.5e-05 | -0.001965 | -0.001965 | -2.9012
2019-03-29 14:19 | 1553883523.5598 | 1536 | 0.20955 | 0.000459 | 1e-06 | 99.7724 | 99.3974 | 126.2345 | 21.1042 | 0.001854 | 0.001773 | -152.316 | 0.23509 | 8.4029 | 0.002547 | 0.004097 | 0.003946 | 6.1e-05 | 0.001072 | 0.31708 | 1.3566 | 2.5474 | 100.16 | 20.95 | 36.649 | 66.2865 | 82.9364 | 64.3335 | 0.009082 | 1.0461 | 0.024892 | 23.7946 | 0.033684 | 0.13985 | 0.000118 | 0.00878 | 0.00878 | 0.64826
2019-03-29 14:19 | 1553883524.5596 | 1537 | 0.20955 | 0.000458 | 1e-06 | 99.7712 | 99.4315 | 126.2537 | 21.1051 | 0.001913 | 0.001838 | -152.3085 | 0.23539 | 8.4027 | 0.002485 | 0.003913 | 0.003666 | 5.7e-05 | 0.001081 | 0.31653 | 1.3365 | 2.5475 | 100.16 | 20.95 | 36.649 | 67.7049 | 81.4389 | 65.9025 | 0.009082 | 1.0413 | 0.024836 | 23.8517 | 0.030506 | 0.14003 | 0.000132 | 0.009803 | 0.009803 | 0.58138
2019-03-29 14:19 | 1553883525.5587 | 1538 | 0.20954 | 0.000459 | 1e-06 | 99.7751 | 99.4237 | 126.0488 | 21.1056 | 0.001969 | 0.001823 | -152.327 | 0.23598 | 8.4015 | 0.002548 | 0.00416 | 0.003999 | 6.1e-05 | 0.001082 | 0.31022 | 1.3229 | 2.5474 | 100.16 | 20.95 | 36.655 | 69.043 | 82.9596 | 65.5362 | 0.009084 | 1.0424 | 0.024829 | 23.8195 | 0.031561 | 0.14038 | 0.000168 | 0.012469 | 0.012469 | 0.45823
2019-03-29 14:19 | 1553883526.5597 | 1539 | 0.20953 | 0.000448 | 1e-06 | 99.6724 | 99.4117 | 126.018 | 21.1061 | 0.001885 | 0.001829 | -152.3173 | 0.23497 | 8.4015 | 0.002528 | 0.004356 | 0.004132 | 5.8e-05 | 0.001087 | 0.29463 | 1.3305 | 2.5475 | 100.16 | 20.95 | 36.658 | 67.0232 | 82.4875 | 65.6935 | 0.009085 | 1.0419 | 0.024821 | 23.8221 | 0.023427 | 0.13978 | 0.000231 | 0.017207 | 0.017207 | 0.33063
2019-03-29 14:19 | 1553883527.5601 | 1540 | 0.20953 | 0.000453 | 1e-06 | 99.7206 | 99.3967 | 125.9734 | 21.1056 | 0.00187 | 0.001836 | -152.3377 | 0.2348 | 8.4025 | 0.002545 | 0.004209 | 0.004026 | 5.6e-05 | 0.001088 | 0.2975 | 1.3512 | 2.5473 | 100.16 | 20.95 | 36.662 | 66.6646 | 82.8988 | 65.8428 | 0.009086 | 1.0415 | 0.024811 | 23.8231 | 0.029102 | 0.13968 | 0.000255 | 0.018933 | 0.018933 | 0.30027
2019-03-29 14:19 | 1553883528.5587 | 1541 | 0.20955 | 0.000453 | 1e-06 | 99.6766 | 99.3247 | 125.7188 | 21.105 | 0.001929 | 0.001854 | -152.3485 | 0.2345 | 8.403 | 0.002546 | 0.004737 | 0.004606 | 6.1e-05 | 0.00109 | 0.29087 | 1.3765 | 2.5474 | 100.16 | 20.95 | 36.675 | 68.0815 | 82.9208 | 66.281 | 0.009089 | 1.0401 | 0.02477 | 23.8144 | 0.031637 | 0.1395 | 0.000129 | 0.009565 | 0.009565 | 0.59359
2019-03-29 14:19 | 1553883529.5598 | 1542 | 0.20954 | 0.000454 | 1e-06 | 99.7102 | 99.3954 | 125.6038 | 21.1053 | 0.00192 | 0.001832 | -152.3349 | 0.23489 | 8.4017 | 0.002505 | 0.00367 | 0.003485 | 5.7e-05 | 0.00109 | 0.32764 | 1.3502 | 2.5475 | 100.16 | 20.95 | 36.681 | 67.879 | 81.9242 | 65.7591 | 0.00909 | 1.0417 | 0.024777 | 23.7852 | 0.028285 | 0.13974 | 0.000184 | 0.013606 | 0.013606 | 0.41801
2019-03-29 14:19 | 1553883530.5589 | 1543 | 0.20956 | 0.000453 | 1e-06 | 99.7084 | 99.2991 | 125.5502 | 21.105 | 0.001878 | 0.001889 | -152.3186 | 0.23492 | 8.4046 | 0.002555 | 0.003897 | 0.003623 | 5.6e-05 | 0.001088 | 0.26936 | 1.3243 | 2.5473 | 100.16 | 20.94 | 36.693 | 66.8716 | 83.1329 | 67.133 | 0.009088 | 1.0375 | 0.024722 | 23.8283 | 0.036789 | 0.13976 | 7.4e-05 | 0.005427 | 0.005427 | 1.0481
2019-03-29 14:19 | 1553883531.5589 | 1544 | 0.20955 | 0.000446 | 1e-06 | 99.7437 | 99.4343 | 125.7444 | 21.1056 | 0.001889 | 0.001831 | -152.3412 | 0.23546 | 8.4052 | 0.002599 | 0.004367 | 0.004204 | 5.5e-05 | 0.001088 | 0.31349 | 1.2798 | 2.5473 | 100.16 | 20.94 | 36.699 | 67.1336 | 84.1859 | 65.7406 | 0.009089 | 1.0418 | 0.024792 | 23.7979 | 0.027791 | 0.14007 | 0.00015 | 0.011137 | 0.011137 | 0.51192
2019-03-29 14:19 | 1553883532.5589 | 1545 | 0.20952 | 0.000449 | 1e-06 | 99.7045 | 99.3053 | 125.6735 | 21.1056 | 0.001874 | 0.001823 | -152.3401 | 0.23559 | 8.4021 | 0.00251 | 0.004034 | 0.003836 | 5.6e-05 | 0.001088 | 0.27992 | 1.3394 | 2.5472 | 100.16 | 20.94 | 36.707 | 66.7683 | 82.0423 | 65.5362 | 0.009091 | 1.0424 | 0.024792 | 23.784 | 0.035886 | 0.14015 | 0.000296 | 0.021945 | 0.021945 | 0.25993
2019-03-29 14:19 | 1553883533.559 | 1546 | 0.20952 | 0.000452 | 1e-06 | 99.77 | 99.5134 | 125.5995 | 21.1058 | 0.001861 | 0.001791 | -152.3371 | 0.23544 | 8.4028 | 0.002558 | 0.004364 | 0.004121 | 5.8e-05 | 0.001086 | 0.28766 | 1.3671 | 2.5474 | 100.16 | 20.94 | 36.724 | 66.4667 | 83.208 | 64.781 | 0.009095 | 1.0447 | 0.024813 | 23.7505 | 0.023033 | 0.14007 | 0.000276 | 0.02052 | 0.02052 | 0.27781
2019-03-29 14:19 | 1553883534.5589 | 1547 | 0.20952 | 0.000453 | 1e-06 | 99.7398 | 99.4107 | 125.6007 | 21.106 | 0.001894 | 0.001851 | -152.3484 | 0.23564 | 8.4016 | 0.002452 | 0.00306 | 0.002887 | 5.7e-05 | 0.00109 | 0.28672 | 1.3251 | 2.5475 | 100.16 | 20.94 | 36.733 | 67.2446 | 80.6515 | 66.225 | 0.009098 | 1.0403 | 0.02476 | 23.8013 | 0.029561 | 0.14018 | 0.000322 | 0.023864 | 0.023864 | 0.23909
2019-03-29 14:19 | 1553883535.559 | 1548 | 0.20954 | 0.000453 | 1e-06 | 99.7625 | 99.3582 | 125.4136 | 21.104 | 0.001919 | 0.001842 | -152.342 | 0.23478 | 8.4043 | 0.002499 | 0.003468 | 0.00335 | 5.5e-05 | 0.001095 | 0.31022 | 1.3605 | 2.5474 | 100.16 | 20.94 | 36.75 | 67.8372 | 81.7849 | 65.9991 | 0.009102 | 1.041 | 0.02475 | 23.7756 | 0.036331 | 0.13967 | 0.00019 | 0.014092 | 0.014092 | 0.40339
2019-03-29 14:19 | 1553883536.5603 | 1549 | 0.20954 | 0.000449 | 1e-06 | 99.7356 | 99.5319 | 125.4742 | 21.1056 | 0.001904 | 0.001827 | -152.3448 | 0.23537 | 8.4014 | 0.002497 | 0.003255 | 0.003056 | 6e-05 | 0.001089 | 0.32853 | 1.3567 | 2.5475 | 100.16 | 20.94 | 36.757 | 67.4954 | 81.7295 | 65.6432 | 0.009104 | 1.0421 | 0.024769 | 23.7689 | 0.01829 | 0.14002 | 0.0002 | 0.014832 | 0.014832 | 0.38423
2019-03-29 14:19 | 1553883537.5586 | 1550 | 0.20955 | 0.000448 | 1e-06 | 99.7661 | 99.4967 | 125.3268 | 21.106 | 0.001915 | 0.001841 | -152.3261 | 0.23429 | 8.4045 | 0.002501 | 0.003554 | 0.003402 | 5.5e-05 | 0.001091 | 0.29502 | 1.3574 | 2.5474 | 100.16 | 20.94 | 36.762 | 67.7517 | 81.8268 | 65.9837 | 0.009105 | 1.041 | 0.024742 | 23.7668 | 0.024188 | 0.13938 | 0.000108 | 0.007951 | 0.007951 | 0.71347
2019-03-29 14:19 | 1553883538.5581 | 1551 | 0.20955 | 0.000449 | 1e-06 | 99.7239 | 99.421 | 125.5106 | 21.1064 | 0.001891 | 0.001826 | -152.3337 | 0.23498 | 8.3997 | 0.002462 | 0.00413 | 0.003927 | 5.7e-05 | 0.001093 | 0.29358 | 1.3414 | 2.5474 | 100.16 | 20.94 | 36.769 | 67.182 | 80.9044 | 65.6072 | 0.009107 | 1.0422 | 0.024774 | 23.771 | 0.027217 | 0.13979 | 0.000132 | 0.009739 | 0.009739 | 0.58417
2019-03-29 14:19 | 1553883539.5591 | 1552 | 0.20953 | 0.000451 | 1e-06 | 99.7766 | 99.4858 | 125.8441 | 21.1058 | 0.001934 | 0.001814 | -152.3337 | 0.235 | 8.4021 | 0.002484 | 0.003867 | 0.003616 | 6e-05 | 0.001082 | 0.29015 | 1.3488 | 2.5474 | 100.16 | 20.94 | 36.771 | 68.1984 | 81.4211 | 65.328 | 0.009107 | 1.043 | 0.024817 | 23.7928 | 0.026109 | 0.1398 | 0.000239 | 0.0178 | 0.0178 | 0.31966
2019-03-29 14:19 | 1553883540.5582 | 1553 | 0.20953 | 0.000451 | 1e-06 | 99.7003 | 99.4165 | 126.0931 | 21.1055 | 0.001872 | 0.001839 | -152.338 | 0.23481 | 8.4032 | 0.002493 | 0.004048 | 0.00399 | 6.3e-05 | 0.001079 | 0.28672 | 1.336 | 2.5474 | 100.16 | 20.94 | 36.772 | 66.7112 | 81.6508 | 65.9231 | 0.009107 | 1.0412 | 0.02482 | 23.8373 | 0.025495 | 0.13969 | 0.000235 | 0.017487 | 0.017487 | 0.32513
2019-03-29 14:19 | 1553883541.5592 | 1554 | 0.20952 | 0.000448 | 1e-06 | 99.706 | 99.4834 | 125.945 | 21.1061 | 0.001923 | 0.001836 | -152.3315 | 0.23461 | 8.4007 | 0.002517 | 0.003758 | 0.003576 | 6.1e-05 | 0.001082 | 0.31758 | 1.3416 | 2.5474 | 100.16 | 20.94 | 36.771 | 67.9561 | 82.223 | 65.8544 | 0.009107 | 1.0414 | 0.024807 | 23.8208 | 0.019989 | 0.13957 | 0.000323 | 0.024017 | 0.024017 | 0.23652
2019-03-29 14:19 | 1553883542.5592 | 1555 | 0.20952 | 0.000445 | 1e-06 | 99.7581 | 99.2906 | 125.5697 | 21.1044 | 0.001926 | 0.00175 | -152.3409 | 0.23377 | 8.4022 | 0.002507 | 0.00435 | 0.00416 | 5.6e-05 | 0.001082 | 0.29264 | 1.3557 | 2.5474 | 100.16 | 20.94 | 36.771 | 68.0146 | 81.9889 | 63.7941 | 0.009107 | 1.0478 | 0.024846 | 23.7129 | 0.042014 | 0.13907 | 0.000323 | 0.024049 | 0.024049 | 0.23535
2019-03-29 14:19 | 1553883543.5592 | 1556 | 0.20952 | 0.000451 | 1e-06 | 99.6775 | 99.3459 | 125.5023 | 21.1056 | 0.001975 | 0.001776 | -152.3411 | 0.23463 | 8.4034 | 0.002553 | 0.003402 | 0.00322 | 6.3e-05 | 0.001081 | 0.29839 | 1.342 | 2.5474 | 100.16 | 20.94 | 36.768 | 69.185 | 83.082 | 64.4013 | 0.009106 | 1.0459 | 0.024817 | 23.7279 | 0.029812 | 0.13958 | 0.000269 | 0.020003 | 0.020003 | 0.28402
2019-03-29 14:19 | 1553883544.5598 | 1557 | 0.20955 | 0.000447 | 1e-06 | 99.691 | 99.2425 | 125.3892 | 21.106 | 0.001896 | 0.001863 | -152.3362 | 0.23429 | 8.4018 | 0.002529 | 0.003771 | 0.003562 | 6.2e-05 | 0.001076 | 0.31581 | 1.338 | 2.5475 | 100.16 | 20.94 | 36.765 | 67.3023 | 82.5114 | 66.5104 | 0.009106 | 1.0394 | 0.024729 | 23.7912 | 0.040336 | 0.13938 | 0.000142 | 0.010528 | 0.010528 | 0.53885
2019-03-29 14:19 | 1553883545.5593 | 1558 | 0.20954 | 0.000449 | 1e-06 | 99.7123 | 99.4769 | 125.4477 | 21.1051 | 0.001948 | 0.001785 | -152.3307 | 0.23485 | 8.4007 | 0.002575 | 0.003393 | 0.003279 | 6.1e-05 | 0.001075 | 0.30613 | 1.3402 | 2.5474 | 100.16 | 20.94 | 36.763 | 68.5497 | 83.6213 | 64.6294 | 0.009105 | 1.0452 | 0.024803 | 23.7308 | 0.021136 | 0.13971 | 0.000209 | 0.015518 | 0.015518 | 0.36644
2019-03-29 14:19 | 1553883546.5583 | 1559 | 0.20953 | 0.00045 | 1e-06 | 99.6949 | 99.5015 | 125.3038 | 21.106 | 0.00188 | 0.001809 | -152.334 | 0.23364 | 8.4025 | 0.002624 | 0.003932 | 0.003713 | 6.1e-05 | 0.001074 | 0.2792 | 1.3374 | 2.5474 | 100.16 | 20.94 | 36.759 | 66.9193 | 84.7988 | 65.208 | 0.009104 | 1.0434 | 0.024768 | 23.7375 | 0.017369 | 0.13899 | 0.000214 | 0.015856 | 0.015856 | 0.35676
2019-03-29 14:19 | 1553883547.5594 | 1560 | 0.20955 | 0.000447 | 1e-06 | 99.6841 | 99.2121 | 125.1942 | 21.1062 | 0.00195 | 0.001844 | -152.3402 | 0.23468 | 8.3997 | 0.002616 | 0.004014 | 0.003842 | 6.1e-05 | 0.001076 | 0.30381 | 1.3421 | 2.5475 | 100.16 | 20.94 | 36.759 | 68.6005 | 84.5992 | 66.0472 | 0.009104 | 1.0408 | 0.024726 | 23.7565 | 0.042456 | 0.13961 | 0.000141 | 0.010418 | 0.010418 | 0.54539
2019-03-29 14:19 | 1553883548.5591 | 1561 | 0.20955 | 0.000447 | 1e-06 | 99.7326 | 99.4448 | 124.8355 | 21.1057 | 0.001919 | 0.001799 | -152.3317 | 0.23474 | 8.4011 | 0.002573 | 0.003757 | 0.003556 | 5.7e-05 | 0.001079 | 0.29048 | 1.3673 | 2.5473 | 100.16 | 20.94 | 36.761 | 67.849 | 83.5714 | 64.9587 | 0.009105 | 1.0442 | 0.024731 | 23.6843 | 0.025846 | 0.13964 | 0.000155 | 0.011483 | 0.011483 | 0.49495
2019-03-29 14:19 | 1553883549.5584 | 1562 | 0.20953 | 0.000449 | 1e-06 | 99.7308 | 99.3431 | 124.8782 | 21.1058 | 0.001878 | 0.001825 | -152.3411 | 0.23495 | 8.4028 | 0.002551 | 0.004505 | 0.004353 | 6e-05 | 0.001075 | 0.29502 | 1.3393 | 2.5474 | 100.16 | 20.93 | 36.764 | 66.8645 | 83.0283 | 65.5934 | 0.0091 | 1.0422 | 0.024712 | 23.7106 | 0.034838 | 0.13977 | 0.000251 | 0.018551 | 0.018551 | 0.30665
2019-03-29 14:19 | 1553883550.5594 | 1563 | 0.20953 | 0.000449 | 1e-06 | 99.7159 | 99.307 | 124.7038 | 21.1052 | 0.00187 | 0.001802 | -152.3308 | 0.23484 | 8.4022 | 0.002574 | 0.00388 | 0.003601 | 5.8e-05 | 0.001075 | 0.28385 | 1.3553 | 2.5474 | 100.16 | 20.93 | 36.767 | 66.6759 | 83.5894 | 65.0402 | 0.0091 | 1.0439 | 0.024715 | 23.6747 | 0.036753 | 0.13971 | 0.000231 | 0.017104 | 0.017104 | 0.33245
2019-03-29 14:19 | 1553883551.5585 | 1564 | 0.2095 | 0.000449 | 1e-06 | 99.7015 | 99.4571 | 124.8709 | 21.1063 | 0.001907 | 0.001835 | -152.3302 | 0.23454 | 8.4024 | 0.002548 | 0.004123 | 0.003998 | 5.8e-05 | 0.00107 | 0.28119 | 1.3508 | 2.5475 | 100.16 | 20.94 | 36.775 | 67.5562 | 82.9603 | 65.824 | 0.009108 | 1.0415 | 0.024703 | 23.718 | 0.021949 | 0.13953 | 0.00045 | 0.033342 | 0.033342 | 0.17032
2019-03-29 14:19 | 1553883552.5587 | 1565 | 0.2095 | 0.000447 | 1e-06 | 99.7371 | 99.3582 | 124.95 | 21.1051 | 0.001934 | 0.00181 | -152.3233 | 0.23414 | 8.4039 | 0.002512 | 0.004058 | 0.00387 | 5.8e-05 | 0.001077 | 0.29358 | 1.3386 | 2.5474 | 100.16 | 20.94 | 36.78 | 68.2158 | 82.1074 | 65.2157 | 0.009109 | 1.0434 | 0.024733 | 23.7042 | 0.034049 | 0.13929 | 0.00041 | 0.030419 | 0.030419 | 0.18637
2019-03-29 14:19 | 1553883553.5576 | 1566 | 0.20954 | 0.00045 | 1e-06 | 99.7242 | 99.3172 | 125.1045 | 21.1055 | 0.001933 | 0.001771 | -152.3207 | 0.2343 | 8.4011 | 0.002579 | 0.003865 | 0.003673 | 6e-05 | 0.001073 | 0.29717 | 1.3481 | 2.5474 | 100.16 | 20.94 | 36.792 | 68.1825 | 83.7201 | 64.2778 | 0.009112 | 1.0463 | 0.024782 | 23.6859 | 0.036583 | 0.13939 | 0.000178 | 0.013159 | 0.013159 | 0.4311
2019-03-29 14:19 | 1553883554.5586 | 1567 | 0.20954 | 0.00045 | 1e-06 | 99.6683 | 99.2704 | 125.1012 | 21.1053 | 0.001922 | 0.001842 | -152.3272 | 0.23457 | 8.4037 | 0.002556 | 0.003414 | 0.003209 | 5.8e-05 | 0.001071 | 0.28982 | 1.4125 | 2.5475 | 100.16 | 20.94 | 36.798 | 67.9258 | 83.1606 | 66.0073 | 0.009114 | 1.0409 | 0.024719 | 23.7463 | 0.035773 | 0.13955 | 0.000181 | 0.013398 | 0.013398 | 0.42392
2019-03-29 14:19 | 1553883555.5576 | 1568 | 0.20955 | 0.000447 | 1e-06 | 99.7338 | 99.2711 | 125.2451 | 21.1059 | 0.001925 | 0.001782 | -152.3408 | 0.23412 | 8.4013 | 0.002479 | 0.003521 | 0.003388 | 6.2e-05 | 0.001072 | 0.32565 | 1.3486 | 2.5474 | 100.16 | 20.94 | 36.804 | 67.9935 | 81.3049 | 64.5436 | 0.009115 | 1.0455 | 0.024786 | 23.7086 | 0.041592 | 0.13928 | 0.000148 | 0.010951 | 0.010951 | 0.51765
2019-03-29 14:19 | 1553883556.5586 | 1569 | 0.20955 | 0.000449 | 1e-06 | 99.7362 | 99.4541 | 125.4345 | 21.1059 | 0.001891 | 0.001774 | -152.3285 | 0.23537 | 8.4008 | 0.002509 | 0.004196 | 0.004045 | 6.1e-05 | 0.001076 | 0.30757 | 1.3131 | 2.5473 | 100.16 | 20.94 | 36.812 | 67.1798 | 82.0293 | 64.3647 | 0.009117 | 1.046 | 0.024812 | 23.7202 | 0.02534 | 0.14002 | 0.000146 | 0.010835 | 0.010835 | 0.52596
2019-03-29 14:19 | 1553883557.5577 | 1570 | 0.20954 | 0.000448 | 1e-06 | 99.7075 | 99.611 | 125.4409 | 21.1061 | 0.001921 | 0.001818 | -152.338 | 0.23435 | 8.4024 | 0.002511 | 0.003828 | 0.003689 | 6e-05 | 0.001079 | 0.30326 | 1.3339 | 2.5474 | 100.16 | 20.94 | 36.816 | 67.8921 | 82.0663 | 65.432 | 0.009118 | 1.0427 | 0.024773 | 23.7583 | 0.008654 | 0.13942 | 0.000168 | 0.012442 | 0.012442 | 0.45605
2019-03-29 14:19 | 1553883558.5588 | 1571 | 0.20957 | 0.000447 | 1e-06 | 99.7515 | 99.3353 | 125.8239 | 21.1063 | 0.001913 | 0.001784 | -152.336 | 0.23587 | 8.4008 | 0.002463 | 0.004747 | 0.004569 | 6e-05 | 0.001081 | 0.3131 | 1.2921 | 2.5474 | 100.16 | 20.94 | 36.82 | 67.7054 | 80.9248 | 64.6073 | 0.009119 | 1.0453 | 0.024841 | 23.7655 | 0.037397 | 0.14032 | 6e-06 | 0.000409 | 0.000409 | 13.9778
2019-03-29 14:19 | 1553883559.5577 | 1572 | 0.20959 | 0.000451 | 1e-06 | 99.7039 | 99.395 | 125.5587 | 21.1057 | 0.001953 | 0.001831 | -152.3315 | 0.23518 | 8.4028 | 0.002491 | 0.004002 | 0.003819 | 5.8e-05 | 0.001079 | 0.30541 | 1.2977 | 2.5473 | 100.16 | 20.93 | 36.825 | 68.6626 | 81.5893 | 65.7431 | 0.009115 | 1.0418 | 0.024773 | 23.7804 | 0.027752 | 0.13991 | -0.000139 | -0.010424 | -0.010424 | -0.54625
2019-03-29 14:19 | 1553883560.5592 | 1573 | 0.20958 | 0.000449 | 1e-06 | 99.6296 | 99.4336 | 125.8659 | 21.1063 | 0.001918 | 0.001807 | -152.3345 | 0.23521 | 8.4026 | 0.00252 | 0.00382 | 0.003731 | 6.1e-05 | 0.001077 | 0.32731 | 1.3148 | 2.5475 | 100.16 | 20.93 | 36.825 | 67.8261 | 82.2817 | 65.1505 | 0.009115 | 1.0436 | 0.024825 | 23.7886 | 0.01762 | 0.13993 | -7.6e-05 | -0.00576 | -0.00576 | -0.98877
2019-03-29 14:19 | 1553883561.5582 | 1574 | 0.20953 | 0.000449 | 1e-06 | 99.6943 | 99.3233 | 126.022 | 21.1061 | 0.001964 | 0.001818 | -152.3267 | 0.23456 | 8.4026 | 0.002565 | 0.003994 | 0.003762 | 5.9e-05 | 0.001071 | 0.29607 | 1.2985 | 2.5474 | 100.16 | 20.94 | 36.823 | 68.9374 | 83.3796 | 65.4195 | 0.00912 | 1.0428 | 0.024831 | 23.8128 | 0.033344 | 0.13954 | 0.000258 | 0.019174 | 0.019174 | 0.29619
2019-03-29 14:19 | 1553883562.5578 | 1575 | 0.20958 | 0.000448 | 1e-06 | 99.7191 | 99.368 | 126.1361 | 21.106 | 0.001883 | 0.001814 | -152.3279 | 0.23414 | 8.4032 | 0.002594 | 0.00409 | 0.003939 | 6.3e-05 | 0.001067 | 0.3131 | 1.3202 | 2.5472 | 100.16 | 20.94 | 36.822 | 66.9825 | 84.0677 | 65.3164 | 0.00912 | 1.0431 | 0.024846 | 23.82 | 0.031548 | 0.13929 | -1.9e-05 | -0.001488 | -0.001488 | -3.8111
2019-03-29 14:19 | 1553883563.5589 | 1576 | 0.20957 | 0.000453 | 1e-06 | 99.6889 | 99.4261 | 126.1471 | 21.1061 | 0.001921 | 0.001818 | -152.339 | 0.23447 | 8.4032 | 0.002486 | 0.003921 | 0.003684 | 6.4e-05 | 0.001064 | 0.29502 | 1.3993 | 2.5474 | 100.16 | 20.93 | 36.822 | 67.8922 | 81.4822 | 65.4157 | 0.009114 | 1.0428 | 0.024843 | 23.8245 | 0.023615 | 0.13949 | 3.3e-05 | 0.002427 | 0.002427 | 2.3389
2019-03-29 14:19 | 1553883564.5589 | 1577 | 0.20957 | 0.000454 | 1e-06 | 99.7527 | 99.437 | 126.283 | 21.1064 | 0.00188 | 0.001757 | -152.3325 | 0.23533 | 8.4008 | 0.002574 | 0.003891 | 0.003809 | 6.5e-05 | 0.001066 | 0.2995 | 1.3947 | 2.5472 | 100.16 | 20.93 | 36.826 | 66.9165 | 83.5891 | 63.9547 | 0.009115 | 1.0473 | 0.024911 | 23.7859 | 0.028351 | 0.14 | -5e-06 | -0.000425 | -0.000425 | -13.4023
2019-03-29 14:19 | 1553883565.5579 | 1578 | 0.20956 | 0.000446 | 1e-06 | 99.7518 | 99.1926 | 126.0338 | 21.1062 | 0.001938 | 0.001824 | -152.3136 | 0.23463 | 8.4022 | 0.002506 | 0.003967 | 0.003777 | 6.6e-05 | 0.001064 | 0.28495 | 1.4028 | 2.5474 | 100.16 | 20.93 | 36.832 | 68.2996 | 81.9462 | 65.5611 | 0.009117 | 1.0423 | 0.024827 | 23.8189 | 0.050278 | 0.13958 | 7.9e-05 | 0.005855 | 0.005855 | 0.97035
2019-03-29 14:19 | 1553883566.5591 | 1579 | 0.20959 | 0.00045 | 1e-06 | 99.7075 | 99.3735 | 126.0704 | 21.107 | 0.001915 | 0.001838 | -152.3175 | 0.2356 | 8.4002 | 0.002506 | 0.00373 | 0.003442 | 6.3e-05 | 0.001061 | 0.33229 | 1.378 | 2.5475 | 100.16 | 20.93 | 36.85 | 67.7596 | 81.9626 | 65.9137 | 0.009121 | 1.0412 | 0.024818 | 23.8348 | 0.030009 | 0.14016 | -8.8e-05 | -0.006588 | -0.006588 | -0.86594
2019-03-29 14:19 | 1553883567.559 | 1580 | 0.20956 | 0.000451 | 1e-06 | 99.6392 | 99.3486 | 126.0413 | 21.107 | 0.001914 | 0.00175 | -152.3173 | 0.23508 | 8.4013 | 0.002501 | 0.003411 | 0.003161 | 6.6e-05 | 0.001063 | 0.28622 | 1.394 | 2.5474 | 100.16 | 20.93 | 36.862 | 67.7294 | 81.834 | 63.789 | 0.009124 | 1.0478 | 0.024893 | 23.7572 | 0.026127 | 0.13985 | 8.2e-05 | 0.006113 | 0.006113 | 0.93117
2019-03-29 14:19 | 1553883568.558 | 1581 |  | 0.000451 |  | 99.6727 | 99.3049 | 125.9331 | 21.1073 | 0.001952 | 0.001826 | -152.3169 | 0.23441 | 8.4021 | 0.002467 | 0.003536 | 0.003449 | 6.1e-05 | 0.00107 | 0.31277 | 1.3542 | 2.5474 | 100.16 | 20.93 | 36.873 | 68.6405 | 81.031 | 65.6088 | 0.009127 | 1.0422 | 0.024815 | 23.8111 | 0.03307 | 0.13945
2019-03-29 14:19 | 1553883569.558 | 1582 |  | 0.000445 |  | 99.7335 | 99.3384 | 125.9636 | 21.1067 | 0.001891 | 0.001838 | -152.3071 | 0.23565 | 8.4003 | 0.002424 | 0.003545 | 0.003321 | 6.5e-05 | 0.001069 | 0.2912 | 1.3204 | 2.5473 | 100.16 | 20.93 | 36.896 | 67.1823 | 79.9809 | 65.9033 | 0.009132 | 1.0413 | 0.024807 | 23.8243 | 0.03551 | 0.14019
2019-03-29 14:19 | 1553883570.558 | 1583 |  |  |  | 99.7033 | 99.5401 | 125.8474 | 21.1072 | 0.001882 | 0.001764 | -152.3131 | 0.23488 | 8.4044 | 0.002501 | 0.003646 | 0.003476 | 6.3e-05 | 0.001074 | 0.2855 | 1.3378 | 2.5474 | 100.16 | 20.94 | 36.907 | 66.9722 | 81.8256 | 64.1312 | 0.009141 | 1.0467 | 0.024861 | 23.751 | 0.014655 | 0.13973
2019-03-29 14:20 | 1553883571.558 | 1584 |  |  |  | 99.7359 | 99.5005 | 125.5106 | 21.1076 | 0.001844 | 0.001811 | -152.3241 | 0.23461 | 8.4045 | 0.002489 | 0.003971 | 0.003712 | 6.4e-05 | 0.001083 | 0.31708 | 1.3448 | 2.5474 | 100.16 | 20.94 | 36.929 | 66.043 | 81.5444 | 65.2588 | 0.009146 | 1.0433 | 0.024786 | 23.7588 | 0.021139 | 0.13957
2019-03-29 14:20 | 1553883572.558 | 1585 |  |  |  | 99.7563 | 99.4172 | 125.4589 | 21.1076 | 0.002025 | 0.001801 | -152.304 | 0.23569 | 8.4035 | 0.002459 | 0.003918 | 0.003752 | 6.5e-05 | 0.001087 | 0.29662 | 1.3993 | 2.5475 | 100.16 | 20.94 | 36.94 | 70.411 | 80.8386 | 65.0043 | 0.009149 | 1.044 | 0.024791 | 23.745 | 0.030454 | 0.14021
2019-03-29 14:20 | 1553883573.558 | 1586 |  |  |  | 99.7392 | 99.3902 |  | 21.1076 | 0.001886 | 0.001806 | -152.2851 | 0.23462 | 8.4042 | 0.002469 | 0.004041 | 0.003894 | 6.5e-05 | 0.001086 | 0.32604 | 1.3491 | 2.5474 | 100.16 | 20.94 | 36.951 | 67.0653 | 81.0727 | 65.133 | 0.009152 | 1.0436 |  |  | 0.03135 | 0.13958
2019-03-29 14:20 | 1553883574.558 | 1587 |  |  |  | 99.7242 | 99.1889 |  | 21.107 | 0.001917 |  | -152.2776 | 0.23499 | 8.4035 | 0.002485 | 0.003554 | 0.003435 | 6.8e-05 | 0.001087 | 0.32045 | 1.3215 | 2.5472 | 100.16 | 20.94 | 36.971 | 67.8082 | 81.4533 |  | 0.009157 |  |  |  | 0.048147 | 0.1398
2019-03-29 14:20 | 1553883575.558 | 1588 |  |  |  | 99.7057 | 99.3097 |  | 21.1068 | 0.00187 |  |  | 0.23503 | 8.4016 | 0.00242 | 0.004361 | 0.004095 | 6.7e-05 | 0.00109 | 0.31542 | 1.3355 | 2.5473 | 100.16 | 20.94 | 36.976 | 66.6659 | 79.8796 |  | 0.009158 |  |  |  | 0.035594 | 0.13982
2019-03-29 14:20 | 1553883576.558 | 1589 |  |  |  | 99.7344 | 99.4721 |  | 21.1075 |  |  |  | 0.23537 | 8.4017 | 0.002527 | 0.003925 | 0.003806 | 6.7e-05 | 0.001095 | 0.30093 | 1.3242 | 2.5474 | 100.16 | 20.94 | 36.978 |  | 82.4662 |  | 0.009158 |  |  |  | 0.023552 | 0.14002
2019-03-29 14:20 | 1553883577.558 | 1590 |  |  |  | 99.6632 | 99.4254 |  | 21.1075 |  |  |  | 0.23502 | 8.4024 | 0.002499 | 0.003388 | 0.003248 | 6.7e-05 | 0.001094 | 0.31166 | 1.3162 | 2.5474 | 100.16 | 20.94 | 36.973 |  | 81.797 |  | 0.009157 |  |  |  | 0.021366 | 0.13981
2019-03-29 14:20 | 1553883578.558 | 1591 |  |  |  | 99.7335 | 99.2657 |  | 21.1065 |  |  |  | 0.23459 | 8.4038 | 0.002522 | 0.00458 | 0.004298 | 6.6e-05 | 0.001097 | 0.30541 | 1.3136 | 2.5473 | 100.16 | 20.94 | 36.967 |  | 82.3419 |  | 0.009156 |  |  |  | 0.042057 | 0.13956
2019-03-29 14:20 | 1553883579.558 | 1592 |  |  |  | 99.7138 | 99.5005 |  | 21.1074 |  |  |  | 0.23523 | 8.4028 | 0.002502 | 0.004296 | 0.004135 | 6.5e-05 | 0.001099 | 0.31437 | 1.3558 | 2.5474 | 100.16 | 20.94 | 36.952 |  | 81.8616 |  | 0.009152 |  |  |  | 0.019153 | 0.13994
2019-03-29 14:20 | 1553883580.558

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-29 13:53 | 0 | 0 | 0 | 0 | 0 | 51.3 | 0.019492
2019-03-29 13:53 | 1 | 0 | 0 | 0 | 0 | 51.4988 | 0.004521
2019-03-29 13:53 | 2 | 0 | 0 | 0 | 0 | 51.2711 | 0.002232
2019-03-29 13:53 | 3 | 0 | 0 | 0 | 0 | 51.8783 | 0.010944
2019-03-29 13:53 | 4 | 0 | 0 | 0 | 0 | 51.382 | 0.025091
2019-03-29 13:53 | 5 | 0 | 0 | 0 | 0 | 51.4416 | 0.094564
2019-03-29 13:53 | 6 | 0 | 0 | 0 | 0 | 51.2418 | 0.39437
2019-03-29 13:53 | 7 | 0 | 0 | 0 | 0 | 51.1971 | 0.7946
2019-03-29 13:53 | 8 | 0 | 0 | 0 | 0 | 51.0719 | 1.5172
2019-03-29 13:53 | 9 | 0 | 0 | 0 | 0 | 50.9907 | 2.1095
2019-03-29 13:53 | 10 | 0 | 0 | 0 | 0 | 50.8893 | 2.3937
2019-03-29 13:53 | 11 | 0 | 0 | 0 | 0 | 50.7883 | 2.9071
2019-03-29 13:53 | 12 | 0 | 0 | 0 | 0 | 50.6804 | 3.2152
2019-03-29 13:53 | 13 | 0.10887 | 0 | 0 | 0 | 50.5684 | 3.3728
2019-03-29 13:53 | 14 | 0.10268 | 0 | 0 | 0 | 50.4631 | 3.3773
2019-03-29 13:53 | 15 | 0.1238 | 0 | 0 | 0 | 50.3564 | 3.3428
2019-03-29 13:53 | 16 | 0.14218 | 0 | 0 | 0 | 50.2165 | 3.3333
2019-03-29 13:53 | 17 | 0.12621 | 0 | 0 | 0 | 50.0806 | 3.3347
2019-03-29 13:53 | 18 | 0.11971 | 0 | 0 | 0 | 49.9625 | 3.32
2019-03-29 13:53 | 19 | 0.12167 | 0 | 0 | 0 | 49.839 | 3.3316
2019-03-29 13:53 | 20 | 0.11566 | 0 | 0 | 0 | 49.7204 | 3.4325
2019-03-29 13:53 | 21 | 0.12448 | 0 | 0 | 0 | 49.6042 | 3.4913
2019-03-29 13:53 | 22 | 0.12907 | 0 | 0 | 0 | 49.4721 | 3.6273
2019-03-29 13:53 | 23 | 0.13134 | 0 | 0 | 0 | 49.3465 | 3.6822
2019-03-29 13:53 | 24 | 0.14566 | 0 | 0 | 0 | 49.2072 | 3.709
2019-03-29 13:54 | 25 | 0.13799 | 0 | 0 | 0 | 49.0606 | 3.8
2019-03-29 13:54 | 26 | 0.12385 | 0 | 0 | 0 | 48.9332 | 3.87
2019-03-29 13:54 | 27 | 0.11828 | 0 | 0 | 0 | 48.8105 | 3.9492
2019-03-29 13:54 | 28 | 0.12546 | 0 | 0 | 0 | 48.6927 | 4.0425
2019-03-29 13:54 | 29 | 0.14784 | 0 | 0 | 0 | 48.5559 | 4.0403
2019-03-29 13:54 | 30 | 0.15195 | 0 | 0 | 0 | 48.4019 | 4.0631
2019-03-29 13:54 | 31 | 0.13083 | 0 | 0 | 0 | 48.2581 | 4.1127
2019-03-29 13:54 | 32 | 0.13823 | 0 | 0 | 0 | 48.1339 | 4.1262
2019-03-29 13:54 | 33 | 0.1437 | 0 | 0 | 0 | 47.982 | 4.0979
2019-03-29 13:54 | 34 | 0.12903 | 0 | 0 | 0 | 47.8512 | 4.0791
2019-03-29 13:54 | 35 | 0.12988 | 0 | 0 | 0 | 47.7207 | 4.0381
2019-03-29 13:54 | 36 | 0.12472 | 0 | 0 | 0 | 47.5921 | 3.9618
2019-03-29 13:54 | 37 | 0.13386 | 0 | 0 | 0 | 47.4679 | 3.8886
2019-03-29 13:54 | 38 | 0.13984 | 0 | 0 | 0 | 47.3251 | 3.8745
2019-03-29 13:54 | 39 | 0.13195 | 0 | 0 | 0 | 47.1917 | 3.7954
2019-03-29 13:54 | 40 | 0.12932 | 0.017418 | 2.3051 | 29.0584 | 47.0605 | 3.7578
2019-03-29 13:54 | 41 | 0.11935 | 0.018349 | 2.4905 | 31.0569 | 46.9345 | 3.7065
2019-03-29 13:54 | 42 | 0.11792 | 0.017935 | 2.5193 | 31.0831 | 46.8195 | 3.6654
2019-03-29 13:54 | 43 | 0.12954 | 0.015846 | 2.2809 | 27.4676 | 46.6954 | 3.5581
2019-03-29 13:54 | 44 | 0.12939 | 0.015296 | 2.2812 | 26.445 | 46.5633 | 3.4218
2019-03-29 13:54 | 45 | 0.11719 | 0.016037 | 2.5029 | 28.2094 | 46.4398 | 3.3058
2019-03-29 13:54 | 46 | 0.11513 | 0.01566 | 2.5335 | 26.978 | 46.3267 | 3.106
2019-03-29 13:54 | 47 | 0.11808 | 0.014608 | 2.4459 | 25.7768 | 46.2083 | 3.0438
2019-03-29 13:54 | 48 | 0.11049 | 0.015052 | 2.5762 | 26.6153 | 46.0932 | 2.9408
2019-03-29 13:54 | 49 | 0.096732 | 0.016443 | 2.9051 | 28.2396 | 45.9889 | 2.7317
2019-03-29 13:54 | 50 | 0.091357 | 0.016601 | 3.0272 | 28.8547 | 45.8978 | 2.6361
2019-03-29 13:54 | 51 | 0.087815 | 0.016629 | 3.0847 | 27.673 | 45.8058 | 2.4301
2019-03-29 13:54 | 52 | 0.074875 | 0.018981 | 3.5205 | 29.6157 | 45.7243 | 2.2175
2019-03-29 13:54 | 53 | 0.064454 | 0.021245 | 3.9499 | 32.6598 | 45.6552 | 2.1051
2019-03-29 13:54 | 54 | 0.070957 | 0.018717 | 3.4465 | 28.7417 | 45.5912 | 2.0394
2019-03-29 13:54 | 55 | 0.072934 | 0.017678 | 3.2071 | 26.722 | 45.5142 | 1.9489
2019-03-29 13:54 | 56 | 0.063902 | 0.019558 | 3.5207 | 29.4716 | 45.4479 | 1.8833
2019-03-29 13:54 | 57 | 0.066243 | 0.018577 | 3.2855 | 27.2215 | 45.3837 | 1.8032
2019-03-29 13:54 | 58 | 0.072808 | 0.016632 | 2.8894 | 23.784 | 45.3144 | 1.7317
2019-03-29 13:54 | 59 | 0.065635 | 0.018154 | 3.0971 | 24.3986 | 45.2414 | 1.6014
2019-03-29 13:54 | 60 | 0.059869 | 0.019556 | 3.2758 | 25.5378 | 45.183 | 1.5289
2019-03-29 13:54 | 61 | 0.059407 | 0.019333 | 3.1831 | 25.8984 | 45.1204 | 1.5386
2019-03-29 13:54 | 62 | 0.055072 | 0.020153 | 3.3221 | 27.0331 | 45.0645 | 1.4888
2019-03-29 13:54 | 63 | 0.07091 | 0.015689 | 2.4924 | 20.3656 | 45.0061 | 1.4441
2019-03-29 13:54 | 64 | 0.05719 | 0.019066 | 2.9769 | 23.5129 | 44.9294 | 1.3447
2019-03-29 13:54 | 65 | 0.050878 | 0.02107 | 3.2386 | 24.9451 | 44.8904 | 1.2692
2019-03-29 13:54 | 66 | 0.052925 | 0.019499 | 3.0081 | 22.7064 | 44.8257 | 1.2017
2019-03-29 13:54 | 67 | 0.040898 | 0.024688 | 3.7488 | 27.9041 | 44.787 | 1.1412
2019-03-29 13:54 | 68 | 0.057768 | 0.017146 | 2.5597 | 19.0014 | 44.7374 | 1.0977
2019-03-29 13:54 | 69 | 0.061591 | 0.015752 | 2.3151 | 17.2654 | 44.6748 | 1.0634
2019-03-29 13:54 | 70 | 0.036982 | 0.026264 | 3.7238 | 28.217 | 44.6208 | 1.0435
2019-03-29 13:54 | 71 | 0.04328 | 0.021413 | 3.0709 | 23.4564 | 44.5935 | 1.0152
2019-03-29 13:54 | 72 | 0.063185 | 0.014651 | 2.0309 | 15.8755 | 44.5312 | 1.0031
2019-03-29 13:54 | 73 | 0.040707 | 0.022218 | 3.0606 | 24.3592 | 44.477 | 0.9916
2019-03-29 13:54 | 74 | 0.03524 | 0.025599 | 3.4213 | 28.1621 | 44.4466 | 0.99243
2019-03-29 13:54 | 75 | 0.032397 | 0.027945 | 3.6406 | 30.663 | 44.4053 | 0.9934
2019-03-29 13:54 | 76 | 0.030473 | 0.030427 | 3.7784 | 32.271 | 44.3811 | 0.98339
2019-03-29 13:54 | 77 | 0.043107 | 0.021499 | 2.6118 | 23.1884 | 44.3413 | 0.99959
2019-03-29 13:54 | 78 | 0.033157 | 0.028007 | 3.3427 | 30.0144 | 44.3001 | 0.9952
2019-03-29 13:54 | 79 | 0.022287 | 0.041737 | 4.8986 | 44.549 | 44.2752 | 0.99284
2019-03-29 13:54 | 80 | 0.033629 | 0.027035 | 3.1956 | 29.7576 | 44.2502 | 1.0007
2019-03-29 13:54 | 81 | 0.036081 | 0.025861 | 2.9483 | 27.3742 | 44.2096 | 0.98768
2019-03-29 13:54 | 82 | 0.036279 | 0.02635 | 2.9033 | 27.7689 | 44.1788 | 1.0074
2019-03-29 13:54 | 83 | 0.038292 | 0.024999 | 2.7234 | 26.1948 | 44.1367 | 1.003
2019-03-29 13:54 | 84 | 0.039207 | 0.024457 | 2.6384 | 25.7074 | 44.1023 | 1.0079
2019-03-29 13:54 | 85 | 0.044929 | 0.021825 | 2.2823 | 22.2291 | 44.057 | 0.99874
2019-03-29 13:55 | 86 | 0.055741 | 0.017582 | 1.826 | 18.1494 | 44.0115 | 1.0117
2019-03-29 13:55 | 87 | 0.049488 | 0.020223 | 2.0409 | 20.2805 | 43.95 | 1.0036
2019-03-29 13:55 | 88 | 0.03203 | 0.030523 | 3.1286 | 31.2306 | 43.9149 | 1.0003
2019-03-29 13:55 | 89 | 0.04876 | 0.01958 | 2.0443 | 20.5845 | 43.878 | 1.0037
2019-03-29 13:55 | 90 | 0.050297 | 0.01895 | 1.9696 | 19.8206 | 43.8204 | 0.99692
2019-03-29 13:55 | 91 | 0.046952 | 0.020308 | 2.0996 | 21.0584 | 43.779 | 0.98873
2019-03-29 13:55 | 92 | 0.045746 | 0.020846 | 2.1437 | 21.7044 | 43.7261 | 0.99288
2019-03-29 13:55 | 93 | 0.048505 | 0.020161 | 2.0175 | 20.6783 | 43.6866 | 1.003
2019-03-29 13:55 | 94 | 0.047852 | 0.020858 | 2.0292 | 20.9197 | 43.6301 | 1.0011
2019-03-29 13:55 | 95 | 0.034733 | 0.028162 | 2.7971 | 28.8303 | 43.5936 | 1.0014
2019-03-29 13:55 | 96 | 0.039526 | 0.024801 | 2.4562 | 25.0838 | 43.5565 | 0.99146
2019-03-29 13:55 | 97 | 0.045888 | 0.021903 | 2.1156 | 21.7509 | 43.5139 | 0.99811
2019-03-29 13:55 | 98 | 0.042559 | 0.024182 | 2.2792 | 23.25 | 43.4672 | 0.9895
2019-03-29 13:55 | 99 | 0.034489 | 0.029988 | 2.8192 | 29.3838 | 43.43 | 1.0134
2019-03-29 13:55 | 100 | 0.036924 | 0.027966 | 2.6244 | 27.103 | 43.3957 | 1.0008
2019-03-29 13:55 | 101 | 0.037725 | 0.028028 | 2.5658 | 26.515 | 43.356 | 1.0003
2019-03-29 13:55 | 102 | 0.045841 | 0.023116 | 2.1107 | 21.8711 | 43.3184 | 1.0026
2019-03-29 13:55 | 103 | 0.056521 | 0.018732 | 1.7042 | 17.5235 | 43.2641 | 0.99044
2019-03-29 13:55 | 104 | 0.045224 | 0.023408 | 2.1225 | 21.7139 | 43.2107 | 0.98198
2019-03-29 13:55 | 105 | 0.040614 | 0.025474 | 2.3574 | 24.1465 | 43.1728 | 0.98068
2019-03-29 13:55 | 106 | 0.029747 | 0.034734 | 3.2026 | 32.862 | 43.1306 | 0.97755
2019-03-29 13:55 | 107 | 0.029642 | 0.034115 | 3.201 | 33.0801 | 43.1104 | 0.98055
2019-03-29 13:55 | 108 | 0.042447 | 0.023785 | 2.2276 | 23.0484 | 43.0683 | 0.97835
2019-03-29 13:55 | 109 | 0.035036 | 0.027475 | 2.6819 | 27.502 | 43.03 | 0.96357
2019-03-29 13:55 | 110 | 0.034143 | 0.028227 | 2.7408 | 28.5333 | 42.997 | 0.9742
2019-03-29 13:55 | 111 | 0.037419 | 0.025067 | 2.4835 | 25.1393 | 42.9607 | 0.94069
2019-03-29 13:55 | 112 | 0.036057 | 0.025415 | 2.5648 | 26.0938 | 42.9235 | 0.94086
2019-03-29 13:55 | 113 | 0.021812 | 0.040943 | 4.2287 | 42.9615 | 42.8911 | 0.93709
2019-03-29 13:55 | 114 | 0.035321 | 0.024605 | 2.5896 | 25.9493 | 42.8732 | 0.91654
2019-03-29 13:55 | 115 | 0.049053 | 0.017251 | 1.8549 | 18.4078 | 42.8201 | 0.90296
2019-03-29 13:55 | 116 | 0.047971 | 0.017192 | 1.8882 | 18.7635 | 42.7789 | 0.9001
2019-03-29 13:55 | 117 | 0.04318 | 0.018599 | 2.0881 | 20.75 | 42.7255 | 0.89599
2019-03-29 13:55 | 118 | 0.038617 | 0.020813 | 2.3221 | 22.7017 | 42.6923 | 0.87666
2019-03-29 13:55 | 119 | 0.047433 | 0.016936 | 1.8806 | 19.0808 | 42.6452 | 0.90506
2019-03-29 13:55 | 120 | 0.046807 | 0.017177 | 1.9015 | 18.9136 | 42.5995 | 0.88528
2019-03-29 13:55 | 121 | 0.042249 | 0.01851 | 2.0917 | 21.059 | 42.5527 | 0.88971
2019-03-29 13:55 | 122 | 0.039258 | 0.019938 | 2.2415 | 22.3212 | 42.5143 | 0.87628
2019-03-29 13:55 | 123 | 0.056086 | 0.014383 | 1.5665 | 15.4606 | 42.47 | 0.86713
2019-03-29 13:55 | 124 | 0.045524 | 0.017753 | 1.9179 | 19.1112 | 42.4087 | 0.87001
2019-03-29 13:55 | 125 | 0.027972 | 0.028827 | 3.0967 | 30.6033 | 42.381 | 0.85603
2019-03-29 13:55 | 126 | 0.024791 | 0.03259 | 3.4825 | 34.2564 | 42.3489 | 0.84924
2019-03-29 13:55 | 127 | 0.034859 | 0.023197 | 2.4687 | 23.9467 | 42.3275 | 0.83476
2019-03-29 13:55 | 128 | 0.049803 | 0.01581 | 1.7206 | 16.9937 | 42.2782 | 0.84634
2019-03-29 13:55 | 129 | 0.041873 | 0.01881 | 2.0344 | 20.3197 | 42.2337 | 0.85085
2019-03-29 13:55 | 130 | 0.030581 | 0.025769 | 2.7675 | 27.4679 | 42.1956 | 0.84001
2019-03-29 13:55 | 131 | 0.030463 | 0.02663 | 2.7678 | 27.4705 | 42.1696 | 0.83683
2019-03-29 13:55 | 132 | 0.040045 | 0.019741 | 2.1024 | 20.6849 | 42.1326 | 0.82832
2019-03-29 13:55 | 133 | 0.027794 | 0.028428 | 3.0071 | 30.0114 | 42.0946 | 0.83414
2019-03-29 13:55 | 134 | 0.019336 | 0.040799 | 4.3103 | 43.1933 | 42.0762 | 0.83519
2019-03-29 13:55 | 135 | 0.022755 | 0.033651 | 3.6456 | 35.9615 | 42.052 | 0.81829
2019-03-29 13:55 | 136 | 0.056021 | 0.013712 | 1.4766 | 14.274 | 42.0239 | 0.79964
2019-03-29 13:55 | 137 | 0.051927 | 0.014343 | 1.5832 | 15.4642 | 41.9493 | 0.80301
2019-03-29 13:55 | 138 | 0.01906 | 0.039051 | 4.2864 | 41.928 | 41.9272 | 0.79917
2019-03-29 13:55 | 139 | 0.029682 | 0.025058 | 2.7397 | 26.4212 | 41.9012 | 0.78423
2019-03-29 13:55 | 140 | 0.036954 | 0.019507 | 2.1851 | 21.3139 | 41.8677 | 0.78765
2019-03-29 13:55 | 141 | 0.032979 | 0.021843 | 2.4424 | 23.6792 | 41.8302 | 0.78092
2019-03-29 13:55 | 142 | 0.026805 | 0.026016 | 2.9871 | 29.2173 | 41.8022 | 0.78317
2019-03-29 13:55 | 143 | 0.037509 | 0.018629 | 2.1276 | 20.3512 | 41.7726 | 0.76335
2019-03-29 13:55 | 144 | 0.043064 | 0.016226 | 1.8441 | 17.9894 | 41.7285 | 0.7747
2019-03-29 13:56 | 145 | 0.03371 | 0.020731 | 2.3509 | 22.5144 | 41.6905 | 0.75897
2019-03-29 13:56 | 146 | 0.016782 | 0.041646 | 4.6909 | 44.0191 | 41.6628 | 0.73875
2019-03-29 13:56 | 147 | 0.019656 | 0.034442 | 3.9907 | 37.6607 | 41.6519 | 0.74025
2019-03-29 13:56 | 148 | 0.033573 | 0.019475 | 2.3164 | 22.1175 | 41.6207 | 0.74255
2019-03-29 13:56 | 149 | 0.025503 | 0.025698 | 3.044 | 28.667 | 41.5897 | 0.7311
2019-03-29 13:56 | 150 | 0.024131 | 0.027106 | 3.1844 | 30.2511 | 41.568 | 0.73001
2019-03-29 13:56 | 151 | 0.04735 | 0.013364 | 1.6201 | 15.5108 | 41.5358 | 0.73443
2019-03-29 13:56 | 152 | 0.045515 | 0.014363 | 1.6734 | 15.6663 | 41.4793 | 0.71306
2019-03-29 13:56 | 153 | 0.029824 | 0.021227 | 2.5494 | 23.8938 | 41.4484 | 0.7126
2019-03-29 13:56 | 154 | 0.035567 | 0.019064 | 2.125 | 19.9722 | 41.4148 | 0.71035
2019-03-29 13:56 | 155 | 0.03739 | 0.018159 | 2.0136 | 18.913 | 41.3777 | 0.70716
2019-03-29 13:56 | 156 | 0.035954 | 0.019502 | 2.081 | 19.2293 | 41.3407 | 0.69137
2019-03-29 13:56 | 157 | 0.039558 | 0.017705 | 1.8866 | 17.4075 | 41.3046 | 0.68861
2019-03-29 13:56 | 158 | 0.048059 | 0.015046 | 1.5464 | 14.8103 | 41.2611 | 0.71177
2019-03-29 13:56 | 159 | 0.027054 | 0.025915 | 2.7364 | 25.2379 | 41.2154 | 0.68279
2019-03-29 13:56 | 160 | 0.018888 | 0.037172 | 3.9081 | 37.2515 | 41.2044 | 0.70359
2019-03-29 13:56 | 161 | 0.016393 | 0.04296 | 4.4969 | 42.193 | 41.1756 | 0.69166
2019-03-29 13:56 | 162 | 0.021732 | 0.031294 | 3.3669 | 31.4416 | 41.1694 | 0.6833
2019-03-29 13:56 | 163 | 0.027056 | 0.025155 | 2.7051 | 25.2011 | 41.1323 | 0.68184
2019-03-29 13:56 | 164 | 0.018271 | 0.035978 | 3.9732 | 37.0075 | 41.1185 | 0.67615
2019-03-29 13:56 | 165 | 0.022088 | 0.029785 | 3.2795 | 30.4804 | 41.0929 | 0.67324
2019-03-29 13:56 | 166 | 0.028555 | 0.02305 | 2.5317 | 23.1482 | 41.0734 | 0.66099
2019-03-29 13:56 | 167 | 0.032881 | 0.02004 | 2.1935 | 19.7464 | 41.0367 | 0.64928
2019-03-29 13:56 | 168 | 0.018188 | 0.035068 | 3.9519 | 35.742 | 41.0116 | 0.65007
2019-03-29 13:56 | 169 | 0.038035 | 0.016747 | 1.8816 | 17.0893 | 40.9927 | 0.65
2019-03-29 13:56 | 170 | 0.04107 | 0.014969 | 1.7354 | 15.7335 | 40.9393 | 0.64618
2019-03-29 13:56 | 171 | 0.026477 | 0.02317 | 2.6714 | 24.0236 | 40.9151 | 0.63607
2019-03-29 13:56 | 172 | 0.025225 | 0.024331 | 2.7911 | 25.4005 | 40.8835 | 0.64073
2019-03-29 13:56 | 173 | 0.022879 | 0.026862 | 3.0626 | 27.9027 | 40.8643 | 0.63838
2019-03-29 13:56 | 174 | 0.026981 | 0.021926 | 2.5842 | 23.4191 | 40.8361 | 0.63188
2019-03-29 13:56 | 175 | 0.030631 | 0.019312 | 2.2609 | 20.3749 | 40.8102 | 0.62409
2019-03-29 13:56 | 176 | 0.035375 | 0.016727 | 1.9542 | 18.0179 | 40.7747 | 0.63738
2019-03-29 13:56 | 177 | 0.03532 | 0.016755 | 1.9474 | 18.0151 | 40.7409 | 0.6363
2019-03-29 13:56 | 178 | 0.028725 | 0.02137 | 2.3806 | 21.7858 | 40.7058 | 0.62579
2019-03-29 13:56 | 179 | 0.022311 | 0.027501 | 3.0411 | 27.6731 | 40.6828 | 0.61741
2019-03-29 13:56 | 180 | 0.04297 | 0.014788 | 1.576 | 14.4919 | 40.655 | 0.62272
2019-03-29 13:56 | 181 | 0.041847 | 0.015196 | 1.6119 | 14.8752 | 40.6018 | 0.62248
2019-03-29 13:56 | 182 | 0.033833 | 0.018819 | 1.9866 | 18.1864 | 40.5733 | 0.6153
2019-03-29 13:56 | 183 | 0.03852 | 0.016479 | 1.7378 | 15.7808 | 40.5316 | 0.60787
2019-03-29 13:56 | 184 | 0.02724 | 0.023348 | 2.4635 | 21.6287 | 40.4995 | 0.58916
2019-03-29 13:56 | 185 | 0.032867 | 0.019363 | 2.0375 | 17.8148 | 40.4732 | 0.58552
2019-03-29 13:56 | 186 | 0.040822 | 0.015079 | 1.6429 | 14.5312 | 40.4331 | 0.59318
2019-03-29 13:56 | 187 | 0.033483 | 0.018359 | 1.9938 | 17.7003 | 40.3953 | 0.59266
2019-03-29 13:56 | 188 | 0.027786 | 0.02217 | 2.4039 | 20.7037 | 40.366 | 0.57527
2019-03-29 13:56 | 189 | 0.029293 | 0.021041 | 2.2692 | 19.6585 | 40.3381 | 0.57585
2019-03-29 13:56 | 190 | 0.024231 | 0.025472 | 2.7322 | 23.0038 | 40.3084 | 0.5574
2019-03-29 13:56 | 191 | 0.035879 | 0.016541 | 1.8335 | 15.7671 | 40.2859 | 0.56571
2019-03-29 13:56 | 192 | 0.035734 | 0.01661 | 1.8352 | 15.8795 | 40.2395 | 0.56744
2019-03-29 13:56 | 193 | 0.019284 | 0.029539 | 3.3848 | 29.7708 | 40.2178 | 0.5741
2019-03-29 13:56 | 194 | 0.039476 | 0.015021 | 1.6469 | 14.1536 | 40.1925 | 0.55873
2019-03-29 13:56 | 195 | 0.048996 | 0.012088 | 1.3216 | 11.7985 | 40.1413 | 0.57808
2019-03-29 13:56 | 196 | 0.031268 | 0.018968 | 2.0705 | 18.5315 | 40.1013 | 0.57944
2019-03-29 13:56 | 197 | 0.027606 | 0.02148 | 2.3355 | 21.1587 | 40.0757 | 0.5841
2019-03-29 13:56 | 198 | 0.036023 | 0.017101 | 1.7866 | 16.2569 | 40.043 | 0.58562
2019-03-29 13:56 | 199 | 0.031318 | 0.01896 | 2.0535 | 18.5335 | 40.0066 | 0.58043
2019-03-29 13:56 | 200 | 0.026171 | 0.022675 | 2.4436 | 21.1353 | 39.9805 | 0.55313
2019-03-29 13:56 | 201 | 0.032846 | 0.018056 | 1.9457 | 16.9616 | 39.9511 | 0.55712
2019-03-29 13:56 | 202 | 0.047556 | 0.012947 | 1.3405 | 11.766 | 39.9127 | 0.55954
2019-03-29 13:56 | 203 | 0.053812 | 0.011438 | 1.1796 | 10.0487 | 39.8587 | 0.54073
2019-03-29 13:56 | 204 | 0.030271 | 0.020384 | 2.0964 | 18.2893 | 39.8124 | 0.55364
2019-03-29 13:56 | 205 | 0.025246 | 0.024392 | 2.4986 | 21.396 | 39.7939 | 0.54017
2019-03-29 13:57 | 206 | 0.034511 | 0.017825 | 1.8248 | 16.3951 | 39.7577 | 0.56581
2019-03-29 13:57 | 207 | 0.046349 | 0.012778 | 1.3569 | 12.1103 | 39.7243 | 0.5613
2019-03-29 13:57 | 208 | 0.039416 | 0.015605 | 1.5903 | 13.6791 | 39.6699 | 0.53917
2019-03-29 13:57 | 209 | 0.022892 | 0.026878 | 2.7266 | 23.2033 | 39.6475 | 0.53117
2019-03-29 13:57 | 210 | 0.040666 | 0.015675 | 1.5333 | 13.5988 | 39.6161 | 0.55301
2019-03-29 13:57 | 211 | 0.047081 | 0.013548 | 1.3214 | 11.8288 | 39.5689 | 0.55691
2019-03-29 13:57 | 212 | 0.023079 | 0.027688 | 2.7038 | 22.9172 | 39.5292 | 0.5289
2019-03-29 13:57 | 213 | 0.021536 | 0.028631 | 2.8928 | 25.7324 | 39.5171 | 0.55417
2019-03-29 13:57 | 214 | 0.051169 | 0.012504 | 1.2197 | 10.5734 | 39.4784 | 0.54103
2019-03-29 13:57 | 215 | 0.05356 | 0.011951 | 1.163 | 9.9608 | 39.4213 | 0.5335
2019-03-29 13:57 | 216 | 0.037665 | 0.017031 | 1.6526 | 14.4477 | 39.377 | 0.54418
2019-03-29 13:57 | 217 | 0.010936 | 0.058627 | 5.6856 | 49.9528 | 39.3482 | 0.5463
2019-03-29 13:57 | 218 | 0.021736 | 0.029528 | 2.8537 | 24.8401 | 39.3454 | 0.53992
2019-03-29 13:57 | 219 | 0.050922 | 0.012133 | 1.211 | 10.1959 | 39.3002 | 0.5192
2019-03-29 13:57 | 220 | 0.038728 | 0.016536 | 1.5906 | 13.5604 | 39.2535 | 0.52517
2019-03-29 13:57 | 221 | 0.022795 | 0.028124 | 2.6928 | 24.4977 | 39.2242 | 0.55842
2019-03-29 13:57 | 222 | 0.032176 | 0.020688 | 1.9091 | 16.5343 | 39.2022 | 0.532
2019-03-29 13:57 | 223 | 0.028166 | 0.024462 | 2.177 | 18.5786 | 39.1629 | 0.52329
2019-03-29 13:57 | 224 | 0.0066473 | 0.10359 | 9.2027 | 78.3354 | 39.1496 | 0.52072
2019-03-29 13:57 | 225 | 0.020392 | 0.033684 | 2.9818 | 24.6456 | 39.1413 | 0.50257
2019-03-29 13:57 | 226 | 0.031321 | 0.021916 | 1.9413 | 16.2187 | 39.1087 | 0.50799
2019-03-29 13:57 | 227 | 0.034734 | 0.01976 | 1.7441 | 15.125 | 39.0806 | 0.52535
2019-03-29 13:57 | 228 | 0.038419 | 0.01788 | 1.58 | 13.4799 | 39.0398 | 0.51789
2019-03-29 13:57 | 229 | 0.026887 | 0.024653 | 2.251 | 18.4837 | 39.007 | 0.49697
2019-03-29 13:57 | 230 | 0.037456 | 0.01775 | 1.6139 | 13.0795 | 38.9811 | 0.4899
2019-03-29 13:57 | 231 | 0.043428 | 0.015295 | 1.3907 | 12.0853 | 38.9332 | 0.52484
2019-03-29 13:57 | 232 | 0.02809 | 0.023615 | 2.1344 | 17.5978 | 38.8992 | 0.49432
2019-03-29 13:57 | 233 | 0.027349 | 0.023424 | 2.1862 | 18.3121 | 38.8735 | 0.50081
2019-03-29 13:57 | 234 | 0.043843 | 0.015097 | 1.3564 | 10.9569 | 38.8405 | 0.48038
2019-03-29 13:57 | 235 | 0.035034 | 0.018933 | 1.6989 | 13.7283 | 38.7914 | 0.48096
2019-03-29 13:57 | 236 | 0.035881 | 0.01786 | 1.6538 | 13.6065 | 38.7686 | 0.48821
2019-03-29 13:57 | 237 | 0.040681 | 0.015722 | 1.4453 | 11.704 | 38.7189 | 0.47613
2019-03-29 13:57 | 238 | 0.02663 | 0.024922 | 2.2026 | 18.0672 | 38.691 | 0.48113
2019-03-29 13:57 | 239 | 0.045822 | 0.014465 | 1.2752 | 10.3675 | 38.6586 | 0.47506
2019-03-29 13:57 | 240 | 0.037414 | 0.017737 | 1.5551 | 13.0885 | 38.6055 | 0.48969
2019-03-29 13:57 | 241 | 0.026194 | 0.026165 | 2.2148 | 18.6937 | 38.5846 | 0.48966
2019-03-29 13:57 | 242 | 0.037097 | 0.018523 | 1.5621 | 12.7015 | 38.5481 | 0.47119
2019-03-29 13:57 | 243 | 0.030779 | 0.022269 | 1.8745 | 15.3174 | 38.5138 | 0.47145
2019-03-29 13:57 | 244 | 0.034132 | 0.020112 | 1.6888 | 13.5936 | 38.4843 | 0.46398
2019-03-29 13:57 | 245 | 0.044354 | 0.015471 | 1.2982 | 10.4777 | 38.4441 | 0.46473
2019-03-29 13:57 | 246 | 0.037243 | 0.019034 | 1.5428 | 12.5914 | 38.4 | 0.46895
2019-03-29 13:57 | 247 | 0.025002 | 0.028317 | 2.2895 | 19.2414 | 38.3711 | 0.48108
2019-03-29 13:57 | 248 | 0.02477 | 0.029592 | 2.3192 | 18.6965 | 38.3471 | 0.46311
2019-03-29 13:57 | 249 | 0.025365 | 0.028818 | 2.2501 | 18.4853 | 38.3208 | 0.46888
2019-03-29 13:57 | 250 | 0.034111 | 0.022114 | 1.6733 | 13.5493 | 38.2947 | 0.46219
2019-03-29 13:57 | 251 | 0.025168 | 0.029096 | 2.2704 | 18.5014 | 38.2567 | 0.46564
2019-03-29 13:57 | 252 | 0.017502 | 0.041921 | 3.2629 | 26.4366 | 38.2438 | 0.46269
2019-03-29 13:57 | 253 | 0.038058 | 0.019311 | 1.4974 | 12.3819 | 38.2154 | 0.47123
2019-03-29 13:57 | 254 | 0.028882 | 0.024598 | 1.9651 | 16.3085 | 38.1741 | 0.47101
2019-03-29 13:57 | 255 | 0.027612 | 0.025746 | 2.0529 | 16.8168 | 38.1561 | 0.46435
2019-03-29 13:57 | 256 | 0.035064 | 0.020274 | 1.6135 | 13.3398 | 38.1167 | 0.46775
2019-03-29 13:57 | 257 | 0.036114 | 0.019681 | 1.5603 | 12.8111 | 38.0873 | 0.46266
2019-03-29 13:57 | 258 | 0.040819 | 0.017963 | 1.3734 | 11.2788 | 38.0441 | 0.46038
2019-03-29 13:57 | 259 | 0.032044 | 0.02288 | 1.7483 | 14.3018 | 38.0088 | 0.45829
2019-03-29 13:57 | 260 | 0.028139 | 0.026044 | 1.9862 | 16.6501 | 37.9788 | 0.46851
2019-03-29 13:57 | 261 | 0.039726 | 0.018453 | 1.4019 | 11.5552 | 37.9489 | 0.45904
2019-03-29 13:57 | 262 | 0.03774 | 0.019481 | 1.479 | 12.5674 | 37.9028 | 0.4743
2019-03-29 13:57 | 263 | 0.017409 | 0.042126 | 3.1775 | 27.1826 | 37.8776 | 0.47321
2019-03-29 13:57 | 264 | 0.02943 | 0.02491 | 1.8826 | 16.4911 | 37.8605 | 0.48533
2019-03-29 13:57 | 265 | 0.037376 | 0.019598 | 1.4811 | 12.7139 | 37.8194 | 0.47519
2019-03-29 13:58 | 266 | 0.02378 | 0.030874 | 2.3333 | 20.4788 | 37.791 | 0.48699
2019-03-29 13:58 | 267 | 0.017786 | 0.041311 | 3.1201 | 26.6664 | 37.7704 | 0.47428
2019-03-29 13:58 | 268 | 0.01738 | 0.043533 | 3.1841 | 27.8253 | 37.7533 | 0.48359
2019-03-29 13:58 | 269 | 0.035806 | 0.021149 | 1.5479 | 13.7548 | 37.7311 | 0.4925
2019-03-29 13:58 | 270 | 0.038734 | 0.01956 | 1.4316 | 12.3544 | 37.6854 | 0.47854
2019-03-29 13:58 | 271 | 0.029874 | 0.024562 | 1.8515 | 16.0059 | 37.6568 | 0.47815
2019-03-29 13:58 | 272 | 0.029873 | 0.024589 | 1.8523 | 16.193 | 37.6234 | 0.48374
2019-03-29 13:58 | 273 | 0.041048 | 0.017328 | 1.3562 | 11.7192 | 37.5941 | 0.48106
2019-03-29 13:58 | 274 | 0.048116 | 0.014805 | 1.1565 | 10.1129 | 37.5429 | 0.48659
2019-03-29 13:58 | 275 | 0.025685 | 0.028705 | 2.1722 | 18.9781 | 37.505 | 0.48746
2019-03-29 13:58 | 276 | 0.019584 | 0.037598 | 2.8562 | 24.0252 | 37.4878 | 0.47051
2019-03-29 13:58 | 277 | 0.027959 | 0.026377 | 1.9973 | 16.8859 | 37.4617 | 0.47211
2019-03-29 13:58 | 278 | 0.039079 | 0.01888 | 1.4287 | 11.8934 | 37.4312 | 0.46478
2019-03-29 13:58 | 279 | 0.034307 | 0.021501 | 1.625 | 13.9463 | 37.3877 | 0.47846
2019-03-29 13:58 | 280 | 0.020487 | 0.035965 | 2.7163 | 22.8889 | 37.365 | 0.46893
2019-03-29 13:58 | 281 | 0.017789 | 0.041503 | 3.1346 | 26.7575 | 37.3439 | 0.47598
2019-03-29 13:58 | 282 | 0.026075 | 0.028319 | 2.125 | 17.9816 | 37.3265 | 0.46887
2019-03-29 13:58 | 283 | 0.027666 | 0.026718 | 2.0035 | 16.6422 | 37.2931 | 0.46043
2019-03-29 13:58 | 284 | 0.026828 | 0.027524 | 2.0653 | 17.1964 | 37.2717 | 0.46134
2019-03-29 13:58 | 285 | 0.034049 | 0.022335 | 1.6209 | 14.101 | 37.2377 | 0.48013
2019-03-29 13:58 | 286 | 0.036071 | 0.020469 | 1.5258 | 12.4402 | 37.2047 | 0.44873
2019-03-29 13:58 | 287 | 0.040652 | 0.018168 | 1.3543 | 11.119 | 37.1653 | 0.45201
2019-03-29 13:58 | 288 | 0.034115 | 0.021671 | 1.6144 | 13.7818 | 37.126 | 0.47016
2019-03-29 13:58 | 289 | 0.03084 | 0.023244 | 1.7827 | 15.0479 | 37.0963 | 0.46407
2019-03-29 13:58 | 290 | 0.037681 | 0.019605 | 1.4537 | 11.8832 | 37.062 | 0.44777
2019-03-29 13:58 | 291 | 0.033483 | 0.022068 | 1.6353 | 13.8122 | 37.0234 | 0.46248
2019-03-29 13:58 | 292 | 0.029714 | 0.024875 | 1.836 | 15.2767 | 36.9949 | 0.45393
2019-03-29 13:58 | 293 | 0.037884 | 0.020107 | 1.4353 | 12.3012 | 36.9611 | 0.46602
2019-03-29 13:58 | 294 | 0.041918 | 0.018098 | 1.2936 | 11.0898 | 36.9204 | 0.46487
2019-03-29 13:58 | 295 | 0.027867 | 0.027176 | 1.9308 | 16.3896 | 36.882 | 0.45674
2019-03-29 13:58 | 296 | 0.011683 | 0.062723 | 4.611 | 38.9587 | 36.8648 | 0.45514
2019-03-29 13:58 | 297 | 0.026309 | 0.027896 | 2.0494 | 17.568 | 36.8507 | 0.46219
2019-03-29 13:58 | 298 | 0.047875 | 0.01533 | 1.124 | 9.5734 | 36.8104 | 0.45832
2019-03-29 13:58 | 299 | 0.040735 | 0.017481 | 1.3221 | 11.4059 | 36.7621 | 0.46462
2019-03-29 13:58 | 300 | 0.03058 | 0.023294 | 1.7712 | 14.8988 | 36.7301 | 0.4556
2019-03-29 13:58 | 301 | 0.03577 | 0.019888 | 1.5022 | 12.6306 | 36.6977 | 0.4518
2019-03-29 13:58 | 302 | 0.024472 | 0.029136 | 2.1977 | 18.3505 | 36.6621 | 0.44907
2019-03-29 13:58 | 303 | 0.020675 | 0.03446 | 2.6133 | 21.647 | 36.6471 | 0.44755
2019-03-29 13:58 | 304 | 0.020732 | 0.035493 | 2.5848 | 21.3667 | 36.6194 | 0.44298
2019-03-29 13:58 | 305 | 0.026208 | 0.028151 | 2.0501 | 17.2151 | 36.6044 | 0.45118
2019-03-29 13:58 | 306 | 0.022698 | 0.032545 | 2.3733 | 19.5688 | 36.5694 | 0.44418
2019-03-29 13:58 | 307 | 0.0082272 | 0.092612 | 6.5402 | 54.6745 | 36.5614 | 0.44982
2019-03-29 13:58 | 308 | 0.017354 | 0.043881 | 3.1051 | 25.7633 | 36.5476 | 0.44711
2019-03-29 13:58 | 309 | 0.017806 | 0.041466 | 3.0198 | 25.3979 | 36.5285 | 0.45224
2019-03-29 13:58 | 310 | 0.0059928 | 0.12699 | 8.9681 | 73.3601 | 36.5145 | 0.43963
2019-03-29 13:58 | 311 | 0.016626 | 0.044357 | 3.2325 | 25.6949 | 36.5107 | 0.42721
2019-03-29 13:58 | 312 | 0.044912 | 0.016423 | 1.2 | 9.8553 | 36.4772 | 0.44262
2019-03-29 13:58 | 313 | 0.039027 | 0.018921 | 1.377 | 10.97 | 36.4293 | 0.42813
2019-03-29 13:58 | 314 | 0.02145 | 0.034437 | 2.5096 | 20.9155 | 36.4023 | 0.44864
2019-03-29 13:58 | 315 | 0.031331 | 0.023538 | 1.713 | 14.9023 | 36.3795 | 0.4669
2019-03-29 13:58 | 316 | 0.054918 | 0.013413 | 0.97283 | 8.3121 | 36.3361 | 0.45649
2019-03-29 13:58 | 317 | 0.05524 | 0.012923 | 0.9695 | 8.0363 | 36.2758 | 0.44392
2019-03-29 13:58 | 318 | 0.031856 | 0.023122 | 1.6702 | 13.5703 | 36.2317 | 0.4323
2019-03-29 13:58 | 319 | 0.028406 | 0.025125 | 1.8735 | 15.9056 | 36.2076 | 0.45181
2019-03-29 13:58 | 320 | 0.033768 | 0.021132 | 1.5779 | 13.1776 | 36.1728 | 0.44498
2019-03-29 13:58 | 321 | 0.015591 | 0.044243 | 3.4044 | 27.7907 | 36.145 | 0.43329
2019-03-29 13:58 | 322 | 0.020228 | 0.035302 | 2.6127 | 21.9113 | 36.1361 | 0.44322
2019-03-29 13:58 | 323 | 0.036425 | 0.019546 | 1.4506 | 11.9863 | 36.1017 | 0.43661
2019-03-29 13:58 | 324 | 0.029823 | 0.023941 | 1.7707 | 14.893 | 36.0688 | 0.44415
2019-03-29 13:58 | 325 | 0.018517 | 0.039711 | 2.8413 | 23.9328 | 36.0435 | 0.44316
2019-03-29 13:59 | 326 | 0.021941 | 0.03246 | 2.3941 | 19.6828 | 36.0282 | 0.43187
2019-03-29 13:59 | 327 | 0.026195 | 0.027213 | 2.003 | 16.6755 | 35.9988 | 0.43682
2019-03-29 13:59 | 328 | 0.034881 | 0.0204 | 1.4995 | 12.7868 | 35.975 | 0.44602
2019-03-29 13:59 | 329 | 0.029799 | 0.023913 | 1.7504 | 14.3816 | 35.9323 | 0.42856
2019-03-29 13:59 | 330 | 0.028738 | 0.024789 | 1.8133 | 15.0596 | 35.9145 | 0.43278
2019-03-29 13:59 | 331 | 0.039278 | 0.017555 | 1.3269 | 11.2354 | 35.8726 | 0.44131
2019-03-29 13:59 | 332 | 0.021347 | 0.032313 | 2.4407 | 20.2226 | 35.8425 | 0.4317
2019-03-29 13:59 | 333 | 0.015527 | 0.044436 | 3.3564 | 27.9024 | 35.8275 | 0.43324
2019-03-29 13:59 | 334 | 0.013918 | 0.047927 | 3.7371 | 31.2688 | 35.81 | 0.43521
2019-03-29 13:59 | 335 | 0.014097 | 0.048972 | 3.6837 | 31.6466 | 35.7992 | 0.44614
2019-03-29 13:59 | 336 | 0.0090647 | 0.076069 | 5.7139 | 46.9222 | 35.7827 | 0.42533
2019-03-29 13:59 | 337 | 0.017498 | 0.039399 | 2.9635 | 24.7254 | 35.7784 | 0.43265
2019-03-29 13:59 | 338 | 0.0059756 | 0.1156 | 8.683 | 72.8172 | 35.7521 | 0.43512
2019-03-29 13:59 | 339 | 0.0025794 | 0.26817 | 20.1432 | 172.4704 | 35.7643 | 0.44487
2019-03-29 13:59 | 340 | 0.01962 | 0.035219 | 2.6418 | 22.0852 | 35.7419 | 0.43332
2019-03-29 13:59 | 341 | 0.024622 | 0.028094 | 2.1043 | 17.7083 | 35.7274 | 0.43601
2019-03-29 13:59 | 342 | 0.031361 | 0.022697 | 1.6464 | 13.857 | 35.6927 | 0.43457
2019-03-29 13:59 | 343 | 0.032097 | 0.022218 | 1.6139 | 13.5201 | 35.6663 | 0.43396
2019-03-29 13:59 | 344 | 0.030321 | 0.023474 | 1.6968 | 13.8177 | 35.6291 | 0.41896
2019-03-29 13:59 | 345 | 0.035884 | 0.019912 | 1.4363 | 12.2065 | 35.6035 | 0.43802
2019-03-29 13:59 | 346 | 0.057661 | 0.012774 | 0.89267 | 7.4125 | 35.5544 | 0.42741
2019-03-29 13:59 | 347 | 0.031204 | 0.023567 | 1.648 | 13.7187 | 35.4993 | 0.42808
2019-03-29 13:59 | 348 | 0.025736 | 0.027668 | 1.9874 | 16.4306 | 35.4876 | 0.42286
2019-03-29 13:59 | 349 | 0.03659 | 0.020113 | 1.3996 | 11.6559 | 35.4439 | 0.42649
2019-03-29 13:59 | 350 | 0.02251 | 0.031773 | 2.2726 | 19.3086 | 35.4196 | 0.43464
2019-03-29 13:59 | 351 | 0.029662 | 0.024026 | 1.7234 | 14.5333 | 35.3944 | 0.4311
2019-03-29 13:59 | 352 | 0.031315 | 0.023522 | 1.6322 | 13.7347 | 35.3617 | 0.4301
2019-03-29 13:59 | 353 | 0.013992 | 0.052656 | 3.6382 | 30.6282 | 35.3363 | 0.42856
2019-03-29 13:59 | 354 | 0.010655 | 0.071331 | 4.7724 | 40.1247 | 35.3308 | 0.42755
2019-03-29 13:59 | 355 | 0.0053729 | 0.13724 | 9.4625 | 81.0294 | 35.3151 | 0.43536
2019-03-29 13:59 | 356 | -0.0025399 | -0.28962 | -19.9254 | -173.9006 | 35.3199 | 0.44169
2019-03-29 13:59 | 357 | 0.014555 | 0.052146 | 3.4739 | 29.113 | 35.3142 | 0.42373
2019-03-29 13:59 | 358 | 0.023109 | 0.032893 | 2.1882 | 18.6438 | 35.2926 | 0.43084
2019-03-29 13:59 | 359 | 0.018362 | 0.041344 | 2.7622 | 23.6678 | 35.2713 | 0.43459
2019-03-29 13:59 | 360 | 0.022777 | 0.03435 | 2.2179 | 18.5038 | 35.2541 | 0.42147
2019-03-29 13:59 | 361 | 0.024296 | 0.031278 | 2.0807 | 16.9108 | 35.2261 | 0.41086
2019-03-29 13:59 | 362 | 0.026452 | 0.028716 | 1.9157 | 15.7664 | 35.205 | 0.41705
2019-03-29 13:59 | 363 | 0.04252 | 0.017899 | 1.1864 | 9.6127 | 35.1701 | 0.40873
2019-03-29 13:59 | 364 | 0.045296 | 0.016801 | 1.1137 | 9 | 35.1233 | 0.40766
2019-03-29 13:59 | 365 | 0.03482 | 0.021826 | 1.4488 | 11.8601 | 35.0828 | 0.41296
2019-03-29 13:59 | 366 | 0.040572 | 0.018772 | 1.2389 | 10.7177 | 35.0504 | 0.43483
2019-03-29 13:59 | 367 | 0.034191 | 0.02223 | 1.4682 | 11.8451 | 35.0047 | 0.405
2019-03-29 13:59 | 368 | 0.017251 | 0.045487 | 2.9075 | 24.3072 | 34.9847 | 0.41931
2019-03-29 13:59 | 369 | 0.011058 | 0.070874 | 4.5139 | 37.8772 | 34.9676 | 0.41886
2019-03-29 13:59 | 370 | 0.0079209 | 0.099096 | 6.3022 | 49.5421 | 34.9613 | 0.39242
2019-03-29 13:59 | 371 | 0.012179 | 0.064506 | 4.1024 | 35.2289 | 34.9501 | 0.42905
2019-03-29 13:59 | 372 | 0.0043811 | 0.17883 | 11.3154 | 98.4088 | 34.9396 | 0.43114
2019-03-29 13:59 | 373 | 0.0019338 | 0.41731 | 25.7061 | 212.0333 | 34.9402 | 0.41004
2019-03-29 13:59 | 374 | 0.0035353 | 0.22857 | 14.0904 | 114.2178 | 34.9346 | 0.4038
2019-03-29 13:59 | 375 | 0.0040283 | 0.20584 | 12.2916 | 101.6631 | 34.9333 | 0.40953
2019-03-29 13:59 | 376 | 0.0020236 | 0.41045 | 24.5456 | 209.4048 | 34.9268 | 0.42375
2019-03-29 13:59 | 377 | 0.012135 | 0.068422 | 4.0798 | 31.916 | 34.9263 | 0.38729
2019-03-29 13:59 | 378 | 0.020704 | 0.040128 | 2.3927 | 20.0855 | 34.9028 | 0.41584
2019-03-29 13:59 | 379 | 0.021518 | 0.039628 | 2.2957 | 18.8378 | 34.8865 | 0.40536
2019-03-29 13:59 | 380 | 0.038986 | 0.02187 | 1.2697 | 10.4899 | 34.8568 | 0.40896
2019-03-29 13:59 | 381 | 0.021378 | 0.039931 | 2.3099 | 19.3926 | 34.8167 | 0.41457
2019-03-29 13:59 | 382 | 0.014969 | 0.057026 | 3.2866 | 26.3905 | 34.812 | 0.39505
2019-03-29 13:59 | 383 | 0.014804 | 0.057742 | 3.3181 | 26.981 | 34.7847 | 0.39942
2019-03-29 13:59 | 384 | 0.020309 | 0.042195 | 2.4247 | 19.97 | 34.7805 | 0.40558
2019-03-29 14:00 | 385 | 0.032993 | 0.026015 | 1.485 | 12.011 | 34.7424 | 0.39628
2019-03-29 14:00 | 386 | 0.033945 | 0.025227 | 1.44 | 11.3615 | 34.7173 | 0.38567
2019-03-29 14:00 | 387 | 0.036042 | 0.024385 | 1.3462 | 10.7958 | 34.6749 | 0.38911
2019-03-29 14:00 | 388 | 0.024375 | 0.036055 | 1.9905 | 16.412 | 34.6479 | 0.40005
2019-03-29 14:00 | 389 | 0.040248 | 0.021812 | 1.2023 | 9.5565 | 34.6201 | 0.38463
2019-03-29 14:00 | 390 | 0.036051 | 0.02439 | 1.3434 | 10.2183 | 34.5716 | 0.36838
2019-03-29 14:00 | 391 | 0.040745 | 0.021546 | 1.1788 | 9.29 | 34.5468 | 0.37852
2019-03-29 14:00 | 392 | 0.028848 | 0.030447 | 1.6683 | 12.9544 | 34.4942 | 0.37371
2019-03-29 14:00 | 393 | 0.012286 | 0.071391 | 3.894 | 30.3089 | 34.49 | 0.37237
2019-03-29 14:00 | 394 | 0.020193 | 0.044564 | 2.3774 | 19.0618 | 34.4634 | 0.38491
2019-03-29 14:00 | 395 | 0.034392 | 0.025476 | 1.3854 | 10.9026 | 34.448 | 0.37497
2019-03-29 14:00 | 396 | 0.026535 | 0.034612 | 1.7881 | 13.9635 | 34.4002 | 0.37052
2019-03-29 14:00 | 397 | 0.0087566 | 0.10514 | 5.4066 | 42.6673 | 34.3973 | 0.37362
2019-03-29 14:00 | 398 | 0.019602 | 0.046897 | 2.4172 | 19.275 | 34.3766 | 0.37783
2019-03-29 14:00 | 399 | 0.0023558 | 0.39062 | 20.057 | 155.3981 | 34.3637 | 0.36609
2019-03-29 14:00 | 400 | 0.0097083 | 0.097154 | 4.878 | 38.7552 | 34.3663 | 0.37625
2019-03-29 14:00 | 401 | 0.014273 | 0.066029 | 3.3178 | 27.2293 | 34.3439 | 0.38864
2019-03-29 14:00 | 402 | 0.032875 | 0.028671 | 1.4373 | 11.3595 | 34.3345 | 0.37344
2019-03-29 14:00 | 403 | 0.044544 | 0.02119 | 1.0599 | 8.4177 | 34.2803 | 0.37496
2019-03-29 14:00 | 404 | 0.030831 | 0.031337 | 1.5254 | 12.3668 | 34.2521 | 0.38128
2019-03-29 14:00 | 405 | 0.018375 | 0.052567 | 2.5686 | 19.917 | 34.2189 | 0.36598
2019-03-29 14:00 | 406 | 0.012891 | 0.075425 | 3.6799 | 28.8728 | 34.2135 | 0.3722
2019-03-29 14:00 | 407 | 0.0094591 | 0.10509 | 4.9964 | 38.8007 | 34.1921 | 0.36702
2019-03-29 14:00 | 408 | 0.011189 | 0.088905 | 4.2108 | 33.6519 | 34.1924 | 0.37654
2019-03-29 14:00 | 409 | 0.039713 | 0.025612 | 1.1846 | 9.3149 | 34.1637 | 0.36992
2019-03-29 14:00 | 410 | 0.035766 | 0.028427 | 1.3158 | 10.1466 | 34.1208 | 0.3629
2019-03-29 14:00 | 411 | 0.020915 | 0.048593 | 2.237 | 17.8423 | 34.0959 | 0.37317
2019-03-29 14:00 | 412 | 0.004503 | 0.22576 | 10.4253 | 83.1304 | 34.0792 | 0.37434
2019-03-29 14:00 | 413 | 0.0049145 | 0.21113 | 9.4443 | 76.073 | 34.0821 | 0.37386
2019-03-29 14:00 | 414 | 0.024938 | 0.041619 | 1.8646 | 15.0242 | 34.0643 | 0.37467
2019-03-29 14:00 | 415 | 0.031836 | 0.03327 | 1.4548 | 11.024 | 34.0353 | 0.35096
2019-03-29 14:00 | 416 | 0.026081 | 0.040606 | 1.7784 | 13.8168 | 34.004 | 0.36036
2019-03-29 14:00 | 417 | 0.025392 | 0.041758 | 1.8231 | 13.9065 | 33.9822 | 0.35311
2019-03-29 14:00 | 418 | 0.028853 | 0.036908 | 1.6063 | 12.4776 | 33.9524 | 0.36002
2019-03-29 14:00 | 419 | 0.025068 | 0.043343 | 1.8419 | 14.4836 | 33.9266 | 0.36308
2019-03-29 14:00 | 420 | 0.0069835 | 0.1552 | 6.5903 | 53.3538 | 33.9056 | 0.3726
2019-03-29 14:00 | 421 | 0.0028212 | 0.38442 | 16.2977 | 127.9724 | 33.9089 | 0.36104
2019-03-29 14:00 | 422 | 0.021314 | 0.051925 | 2.1487 | 16.6139 | 33.8943 | 0.3541
2019-03-29 14:00 | 423 | 0.021438 | 0.051661 | 2.1344 | 16.6461 | 33.8703 | 0.35687
2019-03-29 14:00 | 424 | 0.018931 | 0.059735 | 2.4176 | 17.8436 | 33.8527 | 0.3378
2019-03-29 14:00 | 425 | 0.01357 | 0.081548 | 3.3558 | 25.0452 | 33.8331 | 0.33986
2019-03-29 14:00 | 426 | 0.013592 | 0.08318 | 3.3585 | 24.8336 | 33.8239 | 0.33755
2019-03-29 14:00 | 427 | 0.027869 | 0.040597 | 1.6301 | 12.1931 | 33.8025 | 0.33981
2019-03-29 14:00 | 428 | 0.029751 | 0.038008 | 1.5236 | 11.563 | 33.7712 | 0.34401
2019-03-29 14:00 | 429 | 0.017915 | 0.064363 | 2.5265 | 18.2781 | 33.7465 | 0.32745
2019-03-29 14:00 | 430 | 0.020382 | 0.056505 | 2.2092 | 15.7154 | 33.7324 | 0.3203
2019-03-29 14:00 | 431 | 0.018425 | 0.062598 | 2.4513 | 17.8691 | 33.7065 | 0.32923
2019-03-29 14:00 | 432 | 0.010885 | 0.108 | 4.143 | 31.0301 | 33.6964 | 0.33776
2019-03-29 14:00 | 433 | 0.024989 | 0.048125 | 1.8106 | 13.2685 | 33.6799 | 0.33157
2019-03-29 14:00 | 434 | 0.023175 | 0.051883 | 1.9364 | 14.2281 | 33.6498 | 0.32974
2019-03-29 14:00 | 435 | 0.021831 | 0.054035 | 2.0612 | 15.2592 | 33.634 | 0.33313
2019-03-29 14:00 | 436 | 0.016888 | 0.071003 | 2.6414 | 19.9158 | 33.607 | 0.33633
2019-03-29 14:00 | 437 | 0.017059 | 0.070255 | 2.6135 | 19.748 | 33.5987 | 0.33688
2019-03-29 14:00 | 438 | 0.028384 | 0.042258 | 1.5733 | 11.3685 | 33.5705 | 0.32268
2019-03-29 14:00 | 439 | 0.015602 | 0.078373 | 2.8466 | 21.7237 | 33.5472 | 0.33894
2019-03-29 14:00 | 440 | 0.020521 | 0.059617 | 2.1671 | 16.4447 | 33.5352 | 0.33747
2019-03-29 14:00 | 441 | 0.035855 | 0.034181 | 1.2395 | 9.5614 | 33.5036 | 0.34282
2019-03-29 14:00 | 442 | 0.031562 | 0.038861 | 1.4034 | 10.709 | 33.4685 | 0.338
2019-03-29 14:00 | 443 | 0.01385 | 0.088413 | 3.1798 | 23.196 | 33.444 | 0.32127
2019-03-29 14:00 | 444 | 0.013853 | 0.088395 | 3.1791 | 24.6834 | 33.4366 | 0.34193
2019-03-29 14:01 | 445 | 0.019389 | 0.063098 | 2.2656 | 17.8309 | 33.4145 | 0.34573
2019-03-29 14:01 | 446 | 0.0186 | 0.065661 | 2.3557 | 18.4069 | 33.3993 | 0.34237
2019-03-29 14:01 | 447 | 0.017729 | 0.068891 | 2.4756 | 19.4518 | 33.3776 | 0.34486
2019-03-29 14:01 | 448 | 0.01224 | 0.10175 | 3.5738 | 28.1062 | 33.3647 | 0.34403
2019-03-29 14:01 | 449 | 0.017864 | 0.06966 | 2.4326 | 19.2319 | 33.3505 | 0.34356
2019-03-29 14:01 | 450 | 0.022947 | 0.054147 | 1.8845 | 14.3914 | 33.3292 | 0.33025
2019-03-29 14:01 | 451 | 0.013104 | 0.09692 | 3.323 | 25.1299 | 33.308 | 0.32931
2019-03-29 14:01 | 452 | 0.01531 | 0.083036 | 2.8327 | 21.241 | 33.3001 | 0.3252
2019-03-29 14:01 | 453 | 0.027124 | 0.046881 | 1.598 | 11.6897 | 33.2751 | 0.31707
2019-03-29 14:01 | 454 | 0.021285 | 0.060969 | 2.0462 | 16.1871 | 33.2499 | 0.34455
2019-03-29 14:01 | 455 | 0.018431 | 0.07045 | 2.3723 | 19.0262 | 33.2321 | 0.35067
2019-03-29 14:01 | 456 | 0.023727 | 0.055757 | 1.8292 | 13.542 | 33.211 | 0.32132
2019-03-29 14:01 | 457 | 0.02899 | 0.045667 | 1.4944 | 11.2946 | 33.1847 | 0.32743
2019-03-29 14:01 | 458 | 0.02067 | 0.063983 | 2.1026 | 15.2839 | 33.1559 | 0.31592
2019-03-29 14:01 | 459 | 0.03161 | 0.042603 | 1.3724 | 10.5422 | 33.1394 | 0.33324
2019-03-29 14:01 | 460 | 0.024429 | 0.054998 | 1.7702 | 13.5801 | 33.0965 | 0.33175
2019-03-29 14:01 | 461 | 0.023843 | 0.057449 | 1.8147 | 13.7516 | 33.089 | 0.32787
2019-03-29 14:01 | 462 | 0.037499 | 0.03653 | 1.1559 | 8.9814 | 33.0458 | 0.33679
2019-03-29 14:01 | 463 | 0.026671 | 0.05122 | 1.6111 | 12.0781 | 33.0195 | 0.32214
2019-03-29 14:01 | 464 | 0.027269 | 0.050991 | 1.5812 | 11.9844 | 32.9903 | 0.3268
2019-03-29 14:01 | 465 | 0.023098 | 0.060212 | 1.8671 | 14.041 | 32.9658 | 0.32432
2019-03-29 14:01 | 466 | 0.023516 | 0.059206 | 1.8313 | 13.7314 | 32.9432 | 0.32291
2019-03-29 14:01 | 467 | 0.021999 | 0.06325 | 1.9497 | 14.8813 | 32.9193 | 0.32738
2019-03-29 14:01 | 468 | 0.015 | 0.094326 | 2.8673 | 21.5291 | 32.9009 | 0.32293
2019-03-29 14:01 | 469 | -0.0015317 | -0.91208 | -28.116 | -213.4675 | 32.8912 | 0.32696
2019-03-29 14:01 | 470 | 0.0042342 | 0.33421 | 10.1335 | 76.1526 | 32.8982 | 0.32244
2019-03-29 14:01 | 471 | 0.019675 | 0.072008 | 2.1741 | 16.5818 | 32.8799 | 0.32625
2019-03-29 14:01 | 472 | 0.023839 | 0.059484 | 1.799 | 13.7865 | 32.8615 | 0.32866
2019-03-29 14:01 | 473 | 0.025938 | 0.054675 | 1.6536 | 12.4212 | 32.8335 | 0.32218
2019-03-29 14:01 | 474 | 0.014095 | 0.1005 | 3.0369 | 22.9983 | 32.8129 | 0.32415
2019-03-29 14:01 | 475 | 0.014022 | 0.10111 | 3.05 | 23.1963 | 32.8025 | 0.32526
2019-03-29 14:01 | 476 | 0.023111 | 0.062202 | 1.8477 | 14.296 | 32.7824 | 0.33039
2019-03-29 14:01 | 477 | 0.026009 | 0.055128 | 1.6292 | 12.3407 | 32.7575 | 0.32096
2019-03-29 14:01 | 478 | 0.030792 | 0.047267 | 1.3759 | 10.6197 | 32.7305 | 0.327
2019-03-29 14:01 | 479 | 0.019153 | 0.076019 | 2.209 | 17.0074 | 32.6996 | 0.32573
2019-03-29 14:01 | 480 | 0.022186 | 0.065521 | 1.9154 | 13.916 | 32.689 | 0.30874
2019-03-29 14:01 | 481 | 0.027556 | 0.052765 | 1.5399 | 11.5836 | 32.6549 | 0.3192
2019-03-29 14:01 | 482 | 0.011208 | 0.12963 | 3.7767 | 28.9886 | 32.6389 | 0.32491
2019-03-29 14:01 | 483 | 0.0085323 | 0.17033 | 4.9667 | 37.1802 | 32.6295 | 0.31723
2019-03-29 14:01 | 484 | 0.0098616 | 0.14966 | 4.2774 | 31.7646 | 32.6201 | 0.31325
2019-03-29 14:01 | 485 | 0.021411 | 0.068873 | 1.9786 | 14.5684 | 32.6075 | 0.31193
2019-03-29 14:01 | 486 | 0.016747 | 0.088335 | 2.5247 | 19.0935 | 32.5812 | 0.31976
2019-03-29 14:01 | 487 | 0.0096749 | 0.1553 | 4.3627 | 34.1548 | 32.5747 | 0.33044
2019-03-29 14:01 | 488 | 0.014869 | 0.10099 | 2.8223 | 20.4318 | 32.5587 | 0.3038
2019-03-29 14:01 | 489 | 0.022312 | 0.067326 | 1.8816 | 14.2147 | 32.5442 | 0.31716
2019-03-29 14:01 | 490 | 0.024112 | 0.062471 | 1.7414 | 13.1218 | 32.5155 | 0.3164
2019-03-29 14:01 | 491 | 0.024754 | 0.06079 | 1.696 | 12.9086 | 32.4963 | 0.31955
2019-03-29 14:01 | 492 | 0.028904 | 0.051291 | 1.4508 | 10.9091 | 32.4655 | 0.31532
2019-03-29 14:01 | 493 | 0.020634 | 0.073116 | 2.0346 | 15.1065 | 32.4415 | 0.3117
2019-03-29 14:01 | 494 | 0.017295 | 0.086998 | 2.4208 | 18.4573 | 32.4233 | 0.31922
2019-03-29 14:01 | 495 | 0.014559 | 0.1022 | 2.8908 | 21.7643 | 32.4065 | 0.31687
2019-03-29 14:01 | 496 | 0.017934 | 0.083807 | 2.3341 | 17.7342 | 32.3929 | 0.31804
2019-03-29 14:01 | 497 | 0.010451 | 0.14394 | 4.0052 | 30.0736 | 32.3728 | 0.31431
2019-03-29 14:01 | 498 | 0.011802 | 0.12757 | 3.5498 | 27.4322 | 32.3696 | 0.32377
2019-03-29 14:01 | 499 | 0.029755 | 0.050582 | 1.4002 | 10.6814 | 32.3453 | 0.31782
2019-03-29 14:01 | 500 | 0.031715 | 0.048196 | 1.3208 | 9.9891 | 32.3139 | 0.3168
2019-03-29 14:01 | 501 | 0.026194 | 0.058221 | 1.5928 | 11.9014 | 32.2843 | 0.31174
2019-03-29 14:01 | 502 | 0.015991 | 0.095299 | 2.6139 | 19.2213 | 32.2625 | 0.30736
2019-03-29 14:01 | 503 | 0.01795 | 0.0848 | 2.3239 | 16.0956 | 32.2489 | 0.28891
2019-03-29 14:01 | 504 | 0.036432 | 0.042512 | 1.1427 | 8.8085 | 32.2229 | 0.32091
2019-03-29 14:02 | 505 | 0.027877 | 0.055564 | 1.4935 | 11.395 | 32.1821 | 0.31766
2019-03-29 14:02 | 506 | 0.028974 | 0.054371 | 1.4399 | 10.7476 | 32.1653 | 0.3114
2019-03-29 14:02 | 507 | 0.034902 | 0.045168 | 1.1972 | 9.3034 | 32.1234 | 0.3247
2019-03-29 14:02 | 508 | 0.016245 | 0.096858 | 2.5696 | 18.929 | 32.101 | 0.30751
2019-03-29 14:02 | 509 | 0.019317 | 0.082687 | 2.1543 | 16.6176 | 32.0862 | 0.32101
2019-03-29 14:02 | 510 | 0.021568 | 0.072896 | 1.9255 | 14.5181 | 32.0624 | 0.31313
2019-03-29 14:02 | 511 | 0.0095133 | 0.16531 | 4.3817 | 33.1895 | 32.0461 | 0.31574
2019-03-29 14:02 | 512 | 0.015293 | 0.10476 | 2.7269 | 20.6378 | 32.0392 | 0.31562
2019-03-29 14:02 | 513 | 0.022036 | 0.07263 | 1.8906 | 14.3659 | 32.015 | 0.31657
2019-03-29 14:02 | 514 | 0.023001 | 0.069445 | 1.8109 | 13.5589 | 31.9967 | 0.31187
2019-03-29 14:02 | 515 | 0.019576 | 0.08152 | 2.1221 | 16.4975 | 31.9708 | 0.32295
2019-03-29 14:02 | 516 | -0.0018562 | -0.86024 | -22.3146 | -170.0413 | 31.9614 | 0.31562
2019-03-29 14:02 | 517 | 0.0052824 | 0.30239 | 7.8302 | 59.6274 | 31.9678 | 0.31498
2019-03-29 14:02 | 518 | 0.014104 | 0.11479 | 2.9325 | 21.8625 | 31.9504 | 0.30835
2019-03-29 14:02 | 519 | -0.0043402 | -0.37844 | -9.5485 | -70.1284 | 31.9455 | 0.30437
2019-03-29 14:02 | 520 | 0.0038829 | 0.42277 | 10.6765 | 77.2697 | 31.9528 | 0.30003
2019-03-29 14:02 | 521 | 0.021555 | 0.077243 | 1.9185 | 14.4341 | 31.9353 | 0.31113
2019-03-29 14:02 | 522 | 0.016409 | 0.10135 | 2.5217 | 18.8088 | 31.9152 | 0.30863
2019-03-29 14:02 | 523 | 0.0095565 | 0.17392 | 4.3235 | 32.2914 | 31.9035 | 0.30859
2019-03-29 14:02 | 524 | 0.0038365 | 0.42726 | 10.7615 | 77.3846 | 31.8952 | 0.29689
2019-03-29 14:02 | 525 | 0.022293 | 0.073524 | 1.8502 | 12.8709 | 31.8896 | 0.28693
2019-03-29 14:02 | 526 | 0.04504 | 0.035978 | 0.91913 | 6.57 | 31.8501 | 0.29591
2019-03-29 14:02 | 527 | 0.021071 | 0.07704 | 1.9699 | 14.1848 | 31.8104 | 0.29889
2019-03-29 14:02 | 528 | 0.023843 | 0.067035 | 1.7328 | 12.6605 | 31.8018 | 0.30186
2019-03-29 14:02 | 529 | 0.057973 | 0.027953 | 0.71285 | 5.3843 | 31.7557 | 0.31214
2019-03-29 14:02 | 530 | 0.035635 | 0.044869 | 1.1578 | 8.3176 | 31.6992 | 0.2964
2019-03-29 14:02 | 531 | 0.012733 | 0.12558 | 3.2374 | 23.5567 | 31.6855 | 0.29995
2019-03-29 14:02 | 532 | 0.01284 | 0.12645 | 3.2161 | 23.1437 | 31.6677 | 0.29717
2019-03-29 14:02 | 533 | 0.023704 | 0.068619 | 1.7468 | 12.5545 | 31.6566 | 0.29759
2019-03-29 14:02 | 534 | 0.02806 | 0.057899 | 1.4662 | 10.6239 | 31.6219 | 0.2981
2019-03-29 14:02 | 535 | 0.020423 | 0.07953 | 2.0139 | 14.5774 | 31.6035 | 0.29771
2019-03-29 14:02 | 536 | 0.018779 | 0.086402 | 2.1879 | 15.9564 | 31.5795 | 0.29964
2019-03-29 14:02 | 537 | 0.029955 | 0.054167 | 1.368 | 10.0247 | 31.5626 | 0.30029
2019-03-29 14:02 | 538 | 0.039005 | 0.041618 | 1.0586 | 7.8433 | 31.5205 | 0.30593
2019-03-29 14:02 | 539 | 0.021846 | 0.075379 | 1.8819 | 14.5681 | 31.4907 | 0.31826
2019-03-29 14:02 | 540 | 0.02857 | 0.056873 | 1.444 | 10.9186 | 31.4718 | 0.31194
2019-03-29 14:02 | 541 | 0.025555 | 0.063722 | 1.6165 | 12.8745 | 31.4353 | 0.32901
2019-03-29 14:02 | 542 | 0.023167 | 0.070203 | 1.7824 | 13.7666 | 31.4205 | 0.31893
2019-03-29 14:02 | 543 | 0.03102 | 0.052486 | 1.3279 | 10.0708 | 31.3868 | 0.3124
2019-03-29 14:02 | 544 | 0.024745 | 0.065611 | 1.6717 | 12.0514 | 31.362 | 0.29821
2019-03-29 14:02 | 545 | 0.0092239 | 0.17626 | 4.4712 | 33.2038 | 31.3397 | 0.30627
2019-03-29 14:02 | 546 | 0.0041338 | 0.39225 | 9.9505 | 73.4234 | 31.3405 | 0.30352
2019-03-29 14:02 | 547 | 0.022651 | 0.071579 | 1.8125 | 13.4875 | 31.3256 | 0.3055
2019-03-29 14:02 | 548 | 0.024777 | 0.065652 | 1.6595 | 12.7671 | 31.2989 | 0.31633
2019-03-29 14:02 | 549 | 0.019633 | 0.082751 | 2.088 | 15.8958 | 31.2786 | 0.31209
2019-03-29 14:02 | 550 | 0.0024883 | 0.66192 | 16.4375 | 124.466 | 31.2619 | 0.30971
2019-03-29 14:02 | 551 | 0.020307 | 0.082181 | 2.0196 | 15.5111 | 31.2648 | 0.31499
2019-03-29 14:02 | 552 | 0.051605 | 0.031895 | 0.79345 | 5.8316 | 31.2184 | 0.30094
2019-03-29 14:02 | 553 | 0.030423 | 0.05417 | 1.3548 | 10.4911 | 31.1744 | 0.31917
2019-03-29 14:02 | 554 | 0.010494 | 0.15692 | 3.9211 | 30.9551 | 31.1577 | 0.32485
2019-03-29 14:02 | 555 | 0.027735 | 0.060218 | 1.4891 | 10.1901 | 31.1443 | 0.28262
2019-03-29 14:02 | 556 | 0.037574 | 0.044383 | 1.0936 | 8.381 | 31.1034 | 0.31491
2019-03-29 14:02 | 557 | 0.033134 | 0.05032 | 1.2399 | 9.6391 | 31.0731 | 0.31938
2019-03-29 14:02 | 558 | 0.022496 | 0.074148 | 1.8255 | 14.1622 | 31.0392 | 0.3186
2019-03-29 14:02 | 559 | 0.0098634 | 0.16914 | 4.1677 | 31.9192 | 31.0284 | 0.31483
2019-03-29 14:02 | 560 | 0.014219 | 0.11549 | 2.8654 | 22.4338 | 31.0153 | 0.31898
2019-03-29 14:02 | 561 | 0.017215 | 0.095297 | 2.3855 | 18.4278 | 31.0003 | 0.31724
2019-03-29 14:02 | 562 | -0.00071708 | -2.2906 | -57.3388 | -443.1089 | 30.9851 | 0.31775
2019-03-29 14:02 | 563 | 0.018405 | 0.090342 | 2.2379 | 17.3166 | 30.9928 | 0.31872
2019-03-29 14:02 | 564 | 0.043246 | 0.038581 | 0.95319 | 7.3832 | 30.9472 | 0.31929
2019-03-29 14:02 | 565 | 0.01622 | 0.10295 | 2.5459 | 19.948 | 30.9184 | 0.32356
2019-03-29 14:03 | 566 | 0.018667 | 0.089674 | 2.2096 | 17.1354 | 30.9089 | 0.31986
2019-03-29 14:03 | 567 | 0.019578 | 0.085475 | 2.1099 | 16.0216 | 30.881 | 0.31367
2019-03-29 14:03 | 568 | 0.012456 | 0.13415 | 3.3056 | 25.6989 | 30.8713 | 0.3201
2019-03-29 14:03 | 569 | 0.020829 | 0.079124 | 1.9877 | 15.4723 | 30.8525 | 0.32228
2019-03-29 14:03 | 570 | 0.022517 | 0.073201 | 1.8324 | 14.0086 | 30.831 | 0.31543
2019-03-29 14:03 | 571 | 0.021017 | 0.078349 | 1.9647 | 15.1482 | 30.8083 | 0.31837
2019-03-29 14:03 | 572 | 0.026508 | 0.062236 | 1.5634 | 12.0365 | 30.7876 | 0.31906
2019-03-29 14:03 | 573 | 0.021758 | 0.075859 | 1.9073 | 14.9962 | 30.7577 | 0.32628
2019-03-29 14:03 | 574 | 0.014358 | 0.11467 | 2.8756 | 22.1861 | 30.7446 | 0.31854
2019-03-29 14:03 | 575 | 0.023292 | 0.071766 | 1.7825 | 13.6406 | 30.7251 | 0.31771
2019-03-29 14:03 | 576 | 0.026946 | 0.061196 | 1.5373 | 12.1436 | 30.6992 | 0.32722
2019-03-29 14:03 | 577 | 0.01863 | 0.089829 | 2.2233 | 18.0843 | 30.6744 | 0.33692
2019-03-29 14:03 | 578 | 0.0082208 | 0.20087 | 5.0417 | 37.8823 | 30.6622 | 0.31142
2019-03-29 14:03 | 579 | 0.019883 | 0.084378 | 2.0902 | 16.5403 | 30.6525 | 0.32886
2019-03-29 14:03 | 580 | 0.033546 | 0.049391 | 1.244 | 10.0307 | 30.6221 | 0.33649
2019-03-29 14:03 | 581 | 0.01612 | 0.10417 | 2.5896 | 20.4472 | 30.5927 | 0.3296
2019-03-29 14:03 | 582 | 0.015865 | 0.10573 | 2.6353 | 20.6011 | 30.5866 | 0.32684
2019-03-29 14:03 | 583 | 0.010334 | 0.15996 | 4.036 | 29.5948 | 30.5618 | 0.30584
2019-03-29 14:03 | 584 | 0.0051388 | 0.32114 | 8.1171 | 64.6711 | 30.565 | 0.33233
2019-03-29 14:03 | 585 | 0.033103 | 0.049072 | 1.2569 | 9.6488 | 30.5441 | 0.3194
2019-03-29 14:03 | 586 | 0.027854 | 0.058183 | 1.489 | 11.9104 | 30.5066 | 0.33176
2019-03-29 14:03 | 587 | 0.011306 | 0.14343 | 3.6708 | 27.8537 | 30.4915 | 0.31491
2019-03-29 14:03 | 588 | 0.017117 | 0.094477 | 2.4093 | 18.1181 | 30.4787 | 0.31012
2019-03-29 14:03 | 589 | 0.023983 | 0.06829 | 1.72 | 13.873 | 30.4565 | 0.33272
2019-03-29 14:03 | 590 | 0.02563 | 0.06303 | 1.6201 | 12.5722 | 30.4319 | 0.32222
2019-03-29 14:03 | 591 | 0.03194 | 0.0505 | 1.2981 | 9.7163 | 30.4045 | 0.31034
2019-03-29 14:03 | 592 | 0.026193 | 0.061645 | 1.5804 | 12.7041 | 30.3707 | 0.33276
2019-03-29 14:03 | 593 | 0.026099 | 0.061932 | 1.5947 | 12.8724 | 30.3506 | 0.33596
2019-03-29 14:03 | 594 | 0.042908 | 0.037559 | 0.96289 | 7.4375 | 30.3148 | 0.31912
2019-03-29 14:03 | 595 | 0.036622 | 0.044063 | 1.1296 | 9.0829 | 30.2707 | 0.33263
2019-03-29 14:03 | 596 | 0.008817 | 0.18299 | 4.7078 | 38.0259 | 30.2466 | 0.33527
2019-03-29 14:03 | 597 | 0.015158 | 0.10658 | 2.7467 | 22.4335 | 30.2445 | 0.34004
2019-03-29 14:03 | 598 | 0.043169 | 0.037365 | 0.96127 | 8.0029 | 30.2107 | 0.34547
2019-03-29 14:03 | 599 | 0.041642 | 0.038226 | 0.99158 | 8.0687 | 30.1653 | 0.336
2019-03-29 14:03 | 600 | 0.029997 | 0.052959 | 1.3858 | 11.5119 | 30.1309 | 0.34532
2019-03-29 14:03 | 601 | 0.01654 | 0.096188 | 2.4995 | 20.6827 | 30.1059 | 0.34209
2019-03-29 14:03 | 602 | 0.012505 | 0.12926 | 3.3225 | 27.5717 | 30.0953 | 0.34478
2019-03-29 14:03 | 603 | 0.015803 | 0.10237 | 2.6336 | 21.6906 | 30.079 | 0.34278
2019-03-29 14:03 | 604 | 0.008558 | 0.19169 | 4.8663 | 39.5978 | 30.0662 | 0.33888
2019-03-29 14:03 | 605 | -0.002717 | -0.60442 | -15.3574 | -121.5605 | 30.0621 | 0.33028
2019-03-29 14:03 | 606 | 0.019167 | 0.085717 | 2.1779 | 17.8217 | 30.0629 | 0.34159
2019-03-29 14:03 | 607 | 0.055219 | 0.029754 | 0.75599 | 6.1745 | 30.0206 | 0.34095
2019-03-29 14:03 | 608 | 0.0437 | 0.037553 | 0.95748 | 7.7523 | 29.9643 | 0.33877
2019-03-29 14:03 | 609 | 0.023033 | 0.073224 | 1.8158 | 15.3283 | 29.936 | 0.35306
2019-03-29 14:03 | 610 | 0.03596 | 0.048234 | 1.1713 | 8.8925 | 29.91 | 0.31977
2019-03-29 14:03 | 611 | 0.050686 | 0.035557 | 0.83024 | 7.1828 | 29.8626 | 0.36407
2019-03-29 14:03 | 612 | 0.063734 | 0.029283 | 0.66183 | 6.3184 | 29.8091 | 0.4027
2019-03-29 14:03 | 613 | 0.074512 | 0.026207 | 0.56977 | 6.3371 | 29.7363 | 0.47219
2019-03-29 14:03 | 614 | 0.069203 | 0.029255 | 0.62846 | 8.241 | 29.6641 | 0.5703
2019-03-29 14:03 | 615 | 0.0694 | 0.030379 | 0.64447 | 8.8349 | 29.5968 | 0.61314
2019-03-29 14:03 | 616 | 0.077022 | 0.028784 | 0.60246 | 8.3148 | 29.5235 | 0.64042
2019-03-29 14:03 | 617 | 0.074434 | 0.031388 | 0.64744 | 9.9964 | 29.4446 | 0.74407
2019-03-29 14:03 | 618 | 0.085621 | 0.029182 | 0.58913 | 9.9031 | 29.3713 | 0.84791
2019-03-29 14:03 | 619 | 0.10115 | 0.027147 | 0.5244 | 9.7867 | 29.2722 | 0.98997
2019-03-29 14:03 | 620 | 0.10497 | 0.028739 | 0.53728 | 11.2833 | 29.1712 | 1.1844
2019-03-29 14:03 | 621 | 0.12857 | 0.025822 | 0.47135 | 12.1538 | 29.0581 | 1.5626
2019-03-29 14:03 | 622 | 0.1338 | 0.026605 | 0.55145 | 14.1937 | 28.9184 | 1.8992
2019-03-29 14:03 | 623 | 0.13346 | 0.02812 | 0.69884 | 17.4199 | 28.7923 | 2.3249
2019-03-29 14:03 | 624 | 0.13657 | 0.028453 | 0.85002 | 20.8141 | 28.6505 | 2.8425
2019-03-29 14:04 | 625 | 0.14751 | 0.027219 | 0.96808 | 21.3746 | 28.5168 | 3.153
2019-03-29 14:04 | 626 | 0.17169 | 0.024381 | 0.99243 | 19.2612 | 28.353 | 3.3071
2019-03-29 14:04 | 627 | 0.15945 | 0.027181 | 1.2233 | 21.5798 | 28.182 | 3.4409
2019-03-29 14:04 | 628 | 0.15399 | 0.029057 | 1.3847 | 22.6728 | 28.0328 | 3.4914
2019-03-29 14:04 | 629 | 0.17008 | 0.027128 | 1.3498 | 21.1575 | 27.869 | 3.5984
2019-03-29 14:04 | 630 | 0.16855 | 0.027912 | 1.421 | 21.7057 | 27.6964 | 3.6585
2019-03-29 14:04 | 631 | 0.16914 | 0.028232 | 1.4757 | 21.9281 | 27.5316 | 3.7089
2019-03-29 14:04 | 632 | 0.17859 | 0.026852 | 1.4343 | 21.3071 | 27.3562 | 3.8052
2019-03-29 14:04 | 633 | 0.17436 | 0.027581 | 1.5056 | 22.679 | 27.1772 | 3.9544
2019-03-29 14:04 | 634 | 0.18191 | 0.026259 | 1.4805 | 22.3656 | 27.0047 | 4.0686
2019-03-29 14:04 | 635 | 0.19734 | 0.024116 | 1.3984 | 21.0641 | 26.8117 | 4.1569
2019-03-29 14:04 | 636 | 0.1913 | 0.024812 | 1.4782 | 22.2398 | 26.6152 | 4.2544
2019-03-29 14:04 | 637 | 0.18678 | 0.025303 | 1.5538 | 22.6956 | 26.4293 | 4.2391
2019-03-29 14:04 | 638 | 0.18754 | 0.024915 | 1.5829 | 23.1201 | 26.2404 | 4.336
2019-03-29 14:04 | 639 | 0.1831 | 0.025331 | 1.6629 | 24.0169 | 26.0551 | 4.3974
2019-03-29 14:04 | 640 | 0.18915 | 0.024257 | 1.6433 | 23.557 | 25.8722 | 4.4558
2019-03-29 14:04 | 641 | 0.18092 | 0.025199 | 1.7604 | 25.0262 | 25.6808 | 4.5278
2019-03-29 14:04 | 642 | 0.14854 | 0.030567 | 2.1887 | 30.5078 | 25.516 | 4.5315
2019-03-29 14:04 | 643 | 0.14634 | 0.030987 | 2.27 | 31.1172 | 25.3762 | 4.5537
2019-03-29 14:04 | 644 | 0.16462 | 0.02762 | 2.0575 | 27.7532 | 25.218 | 4.5687
2019-03-29 14:04 | 645 | 0.15811 | 0.028946 | 2.1782 | 29.0651 | 25.0528 | 4.5956
2019-03-29 14:04 | 646 | 0.14431 | 0.032006 | 2.4155 | 30.5159 | 24.904 | 4.4037
2019-03-29 14:04 | 647 | 0.14024 | 0.033224 | 2.5065 | 30.5919 | 24.7619 | 4.2901
2019-03-29 14:04 | 648 | 0.13852 | 0.034216 | 2.5617 | 30.408 | 24.6228 | 4.2122
2019-03-29 14:04 | 649 | 0.13314 | 0.03634 | 2.6774 | 31.2909 | 24.4857 | 4.166
2019-03-29 14:04 | 650 | 0.1276 | 0.038485 | 2.7922 | 31.6401 | 24.3565 | 4.0372
2019-03-29 14:04 | 651 | 0.12244 | 0.040235 | 2.8894 | 31.9187 | 24.2297 | 3.908
2019-03-29 14:04 | 652 | 0.14123 | 0.034886 | 2.494 | 26.8129 | 24.106 | 3.7867
2019-03-29 14:04 | 653 | 0.1475 | 0.033159 | 2.3609 | 24.0395 | 23.9504 | 3.5459
2019-03-29 14:04 | 654 | 0.13383 | 0.035878 | 2.5685 | 26.0811 | 23.8165 | 3.4905
2019-03-29 14:04 | 655 | 0.11861 | 0.039543 | 2.8477 | 27.7649 | 23.6841 | 3.2932
2019-03-29 14:04 | 656 | 0.092652 | 0.048978 | 3.5706 | 34.0257 | 23.5815 | 3.1526
2019-03-29 14:04 | 657 | 0.087137 | 0.050688 | 3.6498 | 34.5636 | 23.4935 | 3.0118
2019-03-29 14:04 | 658 | 0.092336 | 0.046071 | 3.3353 | 30.8118 | 23.4037 | 2.845
2019-03-29 14:04 | 659 | 0.10384 | 0.039137 | 2.8297 | 26.197 | 23.3075 | 2.7204
2019-03-29 14:04 | 660 | 0.093664 | 0.04187 | 3.0549 | 28.0008 | 23.2017 | 2.6227
2019-03-29 14:04 | 661 | 0.070303 | 0.053647 | 3.9484 | 35.914 | 23.1234 | 2.5248
2019-03-29 14:04 | 662 | 0.071406 | 0.050692 | 3.7135 | 33.4249 | 23.055 | 2.3867
2019-03-29 14:04 | 663 | 0.083508 | 0.042079 | 3.0613 | 27.6909 | 22.9778 | 2.3124
2019-03-29 14:04 | 664 | 0.071596 | 0.047629 | 3.4597 | 31.0931 | 22.8935 | 2.2261
2019-03-29 14:04 | 665 | 0.064669 | 0.051707 | 3.7025 | 32.8841 | 22.8333 | 2.1266
2019-03-29 14:04 | 666 | 0.078689 | 0.041477 | 2.9462 | 25.9385 | 22.7586 | 2.0411
2019-03-29 14:04 | 667 | 0.097187 | 0.032955 | 2.3111 | 20.6196 | 22.6748 | 2.004
2019-03-29 14:04 | 668 | 0.097636 | 0.032302 | 2.2428 | 20.066 | 22.5692 | 1.9592
2019-03-29 14:04 | 669 | 0.079576 | 0.038517 | 2.6678 | 24.1795 | 22.4843 | 1.9241
2019-03-29 14:04 | 670 | 0.077638 | 0.038281 | 2.6735 | 24.7615 | 22.4065 | 1.9224
2019-03-29 14:04 | 671 | 0.077304 | 0.037044 | 2.6186 | 24.7156 | 22.3283 | 1.9106
2019-03-29 14:04 | 672 | 0.072413 | 0.038817 | 2.7397 | 25.9317 | 22.2529 | 1.8778
2019-03-29 14:04 | 673 | 0.071471 | 0.038526 | 2.7215 | 25.7656 | 22.1825 | 1.8415
2019-03-29 14:04 | 674 | 0.07308 | 0.037337 | 2.6035 | 24.3588 | 22.1091 | 1.7801
2019-03-29 14:04 | 675 | 0.077119 | 0.034627 | 2.4212 | 23.1303 | 22.0356 | 1.7838
2019-03-29 14:04 | 676 | 0.081344 | 0.033134 | 2.2519 | 21.9691 | 21.9548 | 1.787
2019-03-29 14:04 | 677 | 0.080995 | 0.033114 | 2.2082 | 22.088 | 21.8737 | 1.789
2019-03-29 14:04 | 678 | 0.093972 | 0.028554 | 1.8708 | 19.155 | 21.7899 | 1.8
2019-03-29 14:04 | 679 | 0.09318 | 0.029298 | 1.8609 | 19.4774 | 21.6889 | 1.8149
2019-03-29 14:04 | 680 | 0.090905 | 0.031104 | 1.8786 | 20.2096 | 21.6037 | 1.8371
2019-03-29 14:04 | 681 | 0.11199 | 0.026101 | 1.5101 | 17.5021 | 21.5016 | 1.96
2019-03-29 14:04 | 682 | 0.1224 | 0.024836 | 1.3742 | 16.8071 | 21.3825 | 2.0573
2019-03-29 14:04 | 683 | 0.10533 | 0.029998 | 1.5993 | 20.278 | 21.2638 | 2.1359
2019-03-29 14:04 | 684 | 0.095966 | 0.034042 | 1.773 | 23.5114 | 21.1702 | 2.2563
2019-03-29 14:04 | 685 | 0.10549 | 0.031825 | 1.6279 | 22.7632 | 21.0669 | 2.4013
2019-03-29 14:05 | 686 | 0.1162 | 0.029459 | 1.5101 | 21.8998 | 20.9587 | 2.5447
2019-03-29 14:05 | 687 | 0.11702 | 0.029239 | 1.5374 | 22.6163 | 20.838 | 2.6465
2019-03-29 14:05 | 688 | 0.083275 | 0.041026 | 2.22 | 31.8291 | 20.7331 | 2.6506
2019-03-29 14:05 | 689 | 0.070518 | 0.047506 | 2.6974 | 37.3677 | 20.6661 | 2.6351
2019-03-29 14:05 | 690 | 0.091005 | 0.03614 | 2.1489 | 28.5842 | 20.5835 | 2.6013
2019-03-29 14:05 | 691 | 0.097046 | 0.033124 | 2.0693 | 26.1051 | 20.4874 | 2.5334
2019-03-29 14:05 | 692 | 0.080882 | 0.038712 | 2.5273 | 30.3321 | 20.395 | 2.4533
2019-03-29 14:05 | 693 | 0.078411 | 0.038784 | 2.6292 | 30.2705 | 20.3225 | 2.3736
2019-03-29 14:05 | 694 | 0.088923 | 0.033042 | 2.329 | 25.8654 | 20.235 | 2.3
2019-03-29 14:05 | 695 | 0.085454 | 0.033148 | 2.4153 | 25.7965 | 20.1479 | 2.2044
2019-03-29 14:05 | 696 | 0.076471 | 0.035979 | 2.684 | 27.3332 | 20.0658 | 2.0902
2019-03-29 14:05 | 697 | 0.069639 | 0.03853 | 2.9097 | 28.3933 | 19.9946 | 1.9773
2019-03-29 14:05 | 698 | 0.065389 | 0.039412 | 3.0526 | 29.3586 | 19.9256 | 1.9197
2019-03-29 14:05 | 699 | 0.067114 | 0.037516 | 2.9252 | 27.1594 | 19.8621 | 1.8228
2019-03-29 14:05 | 700 | 0.072051 | 0.033836 | 2.6758 | 25.1466 | 19.7907 | 1.8118
2019-03-29 14:05 | 701 | 0.064392 | 0.036864 | 2.9236 | 27.5622 | 19.721 | 1.7748
2019-03-29 14:05 | 702 | 0.058208 | 0.040167 | 3.1658 | 30.0502 | 19.6618 | 1.7492
2019-03-29 14:05 | 703 | 0.05609 | 0.040617 | 3.2166 | 30.8142 | 19.6033 | 1.7284
2019-03-29 14:05 | 704 | 0.062782 | 0.035272 | 2.8096 | 27.4198 | 19.5473 | 1.7215
2019-03-29 14:05 | 705 | 0.069326 | 0.031693 | 2.4926 | 24.6893 | 19.4779 | 1.7116
2019-03-29 14:05 | 706 | 0.06141 | 0.035052 | 2.7535 | 26.8195 | 19.4122 | 1.647
2019-03-29 14:05 | 707 | 0.054349 | 0.039315 | 3.0591 | 29.9641 | 19.3546 | 1.6285
2019-03-29 14:05 | 708 | 0.069371 | 0.0302 | 2.3523 | 23.4571 | 19.2979 | 1.6272
2019-03-29 14:05 | 709 | 0.084979 | 0.024465 | 1.8918 | 19.1307 | 19.2159 | 1.6257
2019-03-29 14:05 | 710 | 0.073182 | 0.028117 | 2.1664 | 22.2784 | 19.1348 | 1.6304
2019-03-29 14:05 | 711 | 0.06549 | 0.031129 | 2.3897 | 25.0424 | 19.0693 | 1.64
2019-03-29 14:05 | 712 | 0.056833 | 0.035538 | 2.728 | 28.8978 | 19.0035 | 1.6424
2019-03-29 14:05 | 713 | 0.061848 | 0.032303 | 2.4812 | 26.7187 | 18.9517 | 1.6525
2019-03-29 14:05 | 714 | 0.084056 | 0.023499 | 1.8102 | 19.598 | 18.8759 | 1.6473
2019-03-29 14:05 | 715 | 0.077865 | 0.025104 | 1.9453 | 21.1592 | 18.7906 | 1.6476
2019-03-29 14:05 | 716 | 0.06156 | 0.031325 | 2.4463 | 26.4296 | 18.7231 | 1.627
2019-03-29 14:05 | 717 | 0.060364 | 0.031976 | 2.4831 | 27.0531 | 18.6635 | 1.633
2019-03-29 14:05 | 718 | 0.07558 | 0.025229 | 1.9719 | 21.7616 | 18.5983 | 1.6447
2019-03-29 14:05 | 719 | 0.070085 | 0.027176 | 2.1129 | 23.5804 | 18.5167 | 1.6526
2019-03-29 14:05 | 720 | 0.077111 | 0.02469 | 1.9117 | 21.4707 | 18.4555 | 1.6556
2019-03-29 14:05 | 721 | 0.086521 | 0.021955 | 1.6984 | 19.2171 | 18.3619 | 1.6627
2019-03-29 14:05 | 722 | 0.085263 | 0.02197 | 1.7175 | 19.4395 | 18.285 | 1.6575
2019-03-29 14:05 | 723 | 0.089699 | 0.021146 | 1.632 | 18.7219 | 18.191 | 1.6793
2019-03-29 14:05 | 724 | 0.066062 | 0.028383 | 2.2142 | 25.4324 | 18.1118 | 1.6801
2019-03-29 14:05 | 725 | 0.068855 | 0.027145 | 2.1167 | 24.9635 | 18.0527 | 1.7189
2019-03-29 14:05 | 726 | 0.080764 | 0.023412 | 1.8069 | 21.325 | 17.9713 | 1.7223
2019-03-29 14:05 | 727 | 0.085106 | 0.022255 | 1.724 | 20.2832 | 17.8929 | 1.7262
2019-03-29 14:05 | 728 | 0.086851 | 0.021807 | 1.6955 | 19.8862 | 17.8021 | 1.7271
2019-03-29 14:05 | 729 | 0.083068 | 0.022734 | 1.772 | 20.6814 | 17.7211 | 1.718
2019-03-29 14:05 | 730 | 0.065312 | 0.029272 | 2.2663 | 26.6174 | 17.6393 | 1.7384
2019-03-29 14:05 | 731 | 0.061592 | 0.031095 | 2.4185 | 28.0223 | 17.5866 | 1.726
2019-03-29 14:05 | 732 | 0.082408 | 0.02348 | 1.8137 | 20.9856 | 17.5101 | 1.7294
2019-03-29 14:05 | 733 | 0.082243 | 0.023222 | 1.8168 | 21.0286 | 17.4267 | 1.7295
2019-03-29 14:05 | 734 | 0.070605 | 0.027063 | 2.1201 | 24.5507 | 17.3489 | 1.7334
2019-03-29 14:05 | 735 | 0.064751 | 0.029521 | 2.3138 | 26.5855 | 17.2838 | 1.7214
2019-03-29 14:05 | 736 | 0.078568 | 0.024386 | 1.9134 | 22.0195 | 17.2142 | 1.73
2019-03-29 14:05 | 737 | 0.094601 | 0.020722 | 1.592 | 18.262 | 17.1265 | 1.7276
2019-03-29 14:05 | 738 | 0.075474 | 0.026005 | 1.9988 | 23.143 | 17.0334 | 1.7467
2019-03-29 14:05 | 739 | 0.072356 | 0.027149 | 2.0863 | 24.1053 | 16.9724 | 1.7442
2019-03-29 14:05 | 740 | 0.074565 | 0.026057 | 2.0259 | 23.5625 | 16.8872 | 1.7569
2019-03-29 14:05 | 741 | 0.065801 | 0.030173 | 2.2996 | 26.5708 | 16.8255 | 1.7484
2019-03-29 14:05 | 742 | 0.072658 | 0.027328 | 2.0865 | 23.9882 | 16.7522 | 1.7429
2019-03-29 14:05 | 743 | 0.070409 | 0.028227 | 2.1571 | 24.6397 | 16.6821 | 1.7348
2019-03-29 14:05 | 744 | 0.064854 | 0.030632 | 2.3378 | 26.9209 | 16.6123 | 1.7459
2019-03-29 14:06 | 745 | 0.063579 | 0.031266 | 2.3856 | 27.4186 | 16.551 | 1.7433
2019-03-29 14:06 | 746 | 0.078418 | 0.025351 | 1.936 | 22.1353 | 16.4816 | 1.7358
2019-03-29 14:06 | 747 | 0.069752 | 0.028429 | 2.1716 | 25.2577 | 16.3998 | 1.7618
2019-03-29 14:06 | 748 | 0.057732 | 0.034686 | 2.6234 | 30.9138 | 16.3433 | 1.7847
2019-03-29 14:06 | 749 | 0.055764 | 0.035852 | 2.7152 | 31.7511 | 16.2815 | 1.7706
2019-03-29 14:06 | 750 | 0.062728 | 0.031577 | 2.4228 | 28.3097 | 16.2288 | 1.7758
2019-03-29 14:06 | 751 | 0.087024 | 0.022727 | 1.7484 | 20.3537 | 16.1526 | 1.7713
2019-03-29 14:06 | 752 | 0.071367 | 0.027778 | 2.1413 | 24.3579 | 16.0643 | 1.7384
2019-03-29 14:06 | 753 | 0.055821 | 0.03555 | 2.7404 | 31.5422 | 16.0096 | 1.7607
2019-03-29 14:06 | 754 | 0.08841 | 0.022511 | 1.7372 | 19.9233 | 15.9413 | 1.7614
2019-03-29 14:06 | 755 | 0.095955 | 0.020777 | 1.602 | 18.1649 | 15.8388 | 1.743
2019-03-29 14:06 | 756 | 0.068358 | 0.029423 | 2.2452 | 25.8075 | 15.7586 | 1.7642
2019-03-29 14:06 | 757 | 0.051444 | 0.039107 | 2.9822 | 33.8477 | 15.6997 | 1.7413
2019-03-29 14:06 | 758 | 0.057907 | 0.034843 | 2.6559 | 30.4748 | 15.6493 | 1.7647
2019-03-29 14:06 | 759 | 0.073152 | 0.028184 | 2.1021 | 23.9029 | 15.5813 | 1.7485
2019-03-29 14:06 | 760 | 0.076532 | 0.026968 | 2.0087 | 22.941 | 15.5062 | 1.7557
2019-03-29 14:06 | 761 | 0.060661 | 0.033936 | 2.5238 | 28.9304 | 15.4329 | 1.7549
2019-03-29 14:06 | 762 | 0.065865 | 0.031623 | 2.3263 | 26.3986 | 15.3802 | 1.7387
2019-03-29 14:06 | 763 | 0.070309 | 0.029602 | 2.1742 | 24.7949 | 15.3014 | 1.7433
2019-03-29 14:06 | 764 | 0.051815 | 0.039671 | 2.9411 | 33.7698 | 15.2444 | 1.7498
2019-03-29 14:06 | 765 | 0.067421 | 0.030424 | 2.2511 | 25.7276 | 15.1899 | 1.7346
2019-03-29 14:06 | 766 | 0.078677 | 0.026058 | 1.928 | 21.8868 | 15.1102 | 1.722
2019-03-29 14:06 | 767 | 0.071683 | 0.028601 | 2.1143 | 24.1532 | 15.0369 | 1.7314
2019-03-29 14:06 | 768 | 0.076565 | 0.026784 | 1.9773 | 22.6038 | 14.9646 | 1.7307
2019-03-29 14:06 | 769 | 0.075649 | 0.026846 | 2.003 | 22.8517 | 14.885 | 1.7287
2019-03-29 14:06 | 770 | 0.071879 | 0.028518 | 2.1058 | 23.8947 | 14.8141 | 1.7175
2019-03-29 14:06 | 771 | 0.070738 | 0.029052 | 2.1419 | 24.4201 | 14.7408 | 1.7274
2019-03-29 14:06 | 772 | 0.06405 | 0.032093 | 2.3629 | 26.9158 | 14.6735 | 1.724
2019-03-29 14:06 | 773 | 0.070985 | 0.028934 | 2.1265 | 24.389 | 14.6096 | 1.7313
2019-03-29 14:06 | 774 | 0.072482 | 0.02834 | 2.0824 | 23.8237 | 14.5329 | 1.7268
2019-03-29 14:06 | 775 | 0.061752 | 0.033335 | 2.4482 | 28.0505 | 14.4678 | 1.7322
2019-03-29 14:06 | 776 | 0.052029 | 0.039106 | 2.9052 | 33.2731 | 14.4086 | 1.7312
2019-03-29 14:06 | 777 | 0.071438 | 0.028469 | 2.1183 | 24.1373 | 14.3562 | 1.7243
2019-03-29 14:06 | 778 | 0.094162 | 0.021596 | 1.6073 | 18.3214 | 14.265 | 1.7252
2019-03-29 14:06 | 779 | 0.083217 | 0.024447 | 1.8178 | 20.8788 | 14.1761 | 1.7375
2019-03-29 14:06 | 780 | 0.080384 | 0.025315 | 1.8781 | 21.5059 | 14.0972 | 1.7287
2019-03-29 14:06 | 781 | 0.08962 | 0.02276 | 1.6855 | 19.3743 | 14.0124 | 1.7363
2019-03-29 14:06 | 782 | 0.089476 | 0.022787 | 1.6842 | 19.5072 | 13.9201 | 1.7454
2019-03-29 14:06 | 783 | 0.083866 | 0.024224 | 1.7899 | 20.8859 | 13.8348 | 1.7516
2019-03-29 14:06 | 784 | 0.088279 | 0.023281 | 1.7017 | 19.837 | 13.7504 | 1.7512
2019-03-29 14:06 | 785 | 0.082319 | 0.024937 | 1.8278 | 21.3854 | 13.661 | 1.7604
2019-03-29 14:06 | 786 | 0.064622 | 0.031776 | 2.3332 | 27.134 | 13.5887 | 1.7535
2019-03-29 14:06 | 787 | 0.057671 | 0.036024 | 2.62 | 30.4547 | 13.5291 | 1.7564
2019-03-29 14:06 | 788 | 0.059313 | 0.034605 | 2.5427 | 29.2919 | 13.4707 | 1.7374
2019-03-29 14:06 | 789 | 0.066259 | 0.030932 | 2.2743 | 26.0444 | 13.4094 | 1.7257
2019-03-29 14:06 | 790 | 0.053791 | 0.03761 | 2.796 | 31.9437 | 13.3429 | 1.7183
2019-03-29 14:06 | 791 | 0.042848 | 0.046642 | 3.4937 | 40.3092 | 13.3011 | 1.7272
2019-03-29 14:06 | 792 | 0.061092 | 0.032352 | 2.4447 | 28.0579 | 13.2502 | 1.7141
2019-03-29 14:06 | 793 | 0.065542 | 0.03019 | 2.2808 | 26.3626 | 13.1822 | 1.7279
2019-03-29 14:06 | 794 | 0.050316 | 0.039291 | 2.9663 | 34.2652 | 13.1244 | 1.7241
2019-03-29 14:06 | 795 | 0.034504 | 0.05728 | 4.3176 | 49.6921 | 13.081 | 1.7146
2019-03-29 14:06 | 796 | 0.064143 | 0.030846 | 2.3245 | 26.3332 | 13.0443 | 1.6891
2019-03-29 14:06 | 797 | 0.079887 | 0.024806 | 1.8694 | 21.1983 | 12.9562 | 1.6935
2019-03-29 14:06 | 798 | 0.049617 | 0.039944 | 3.0082 | 33.9279 | 12.8958 | 1.6834
2019-03-29 14:06 | 799 | 0.042504 | 0.046209 | 3.5118 | 39.5301 | 12.8516 | 1.6802
2019-03-29 14:06 | 800 | 0.066587 | 0.029451 | 2.2306 | 25.232 | 12.8027 | 1.6801
2019-03-29 14:06 | 801 | 0.077852 | 0.025185 | 1.9014 | 21.512 | 12.721 | 1.6748
2019-03-29 14:06 | 802 | 0.074614 | 0.026007 | 1.9776 | 22.3634 | 12.6507 | 1.6686
2019-03-29 14:06 | 803 | 0.084264 | 0.023005 | 1.7445 | 19.9931 | 12.5695 | 1.6847
2019-03-29 14:06 | 804 | 0.077651 | 0.025251 | 1.8894 | 21.6886 | 12.4859 | 1.6841
2019-03-29 14:06 | 805 | 0.072241 | 0.026861 | 2.028 | 23.3287 | 12.4144 | 1.6853
2019-03-29 14:07 | 806 | 0.064865 | 0.029585 | 2.2571 | 25.9905 | 12.3417 | 1.6859
2019-03-29 14:07 | 807 | 0.066111 | 0.028991 | 2.2144 | 25.0692 | 12.2822 | 1.6574
2019-03-29 14:07 | 808 | 0.078627 | 0.02411 | 1.8586 | 20.882 | 12.2069 | 1.6419
2019-03-29 14:07 | 809 | 0.070597 | 0 | 0 | 0 | 12.1296 | 1.6281
2019-03-29 14:07 | 810 | 0.07182 | 0 | 0 | 0 | 12.0643 | 1.6134
2019-03-29 14:07 | 811 | 0.061086 | 0 | 0 | 0 | 11.9887 | 1.6254
2019-03-29 14:07 | 812 | 0.052839 | 0 | 0 | 0 | 11.9415 | 1.6245
2019-03-29 14:07 | 813 | 0.054723 | 0 | 0 | 0 | 11.8805 | 1.6023
2019-03-29 14:07 | 814 | 0.050014 | 0 | 0 | 0 | 11.8328 | 1.5947
2019-03-29 14:07 | 815 | 0.066224 | 0 | 0 | 0 | 11.7754 | 1.5923
2019-03-29 14:07 | 816 | 0.083708 | 0 | 0 | 0 | 11.7003 | 1.5881
2019-03-29 14:07 | 817 | 0.07101 | 0 | 0 | 0 | 11.6157 | 1.5795
2019-03-29 14:07 | 818 | 0.054158 | 0 | 0 | 0 | 11.5597 | 1.5592
2019-03-29 14:07 | 819 | 0.059683 | 0 | 0 | 0 | 11.5021 | 1.5584
2019-03-29 14:07 | 820 | 0.052734 | 0 | 0 | 0 | 11.442 | 1.5649
2019-03-29 14:07 | 821 | 0.073083 | 0 | 0 | 0 | 11.3907 | 1.5708
2019-03-29 14:07 | 822 | 0.07403 | 0 | 0 | 0 | 11.3008 | 1.5631
2019-03-29 14:07 | 823 | 0.044638 | 0 | 0 | 0 | 11.2495 | 1.5667
2019-03-29 14:07 | 824 | 0.061827 | 0 | 0 | 0 | 11.2009 | 1.5301
2019-03-29 14:07 | 825 | 0.072597 | 0 | 0 | 0 | 11.127 | 1.5203
2019-03-29 14:07 | 826 | 0.055138 | 0 | 0 | 0 | 11.0628 | 1.51
2019-03-29 14:07 | 827 | 0.043481 | 0 | 0 | 0 | 11.0156 | 1.4858
2019-03-29 14:07 | 828 | 0.048809 | 0 | 0 | 0 | 10.9711 | 1.4603
2019-03-29 14:07 | 829 | 0.065752 | 0 | 0 | 0 | 10.915 | 1.4465
2019-03-29 14:07 | 830 | 0.065278 | 0 | 0 | 0 | 10.8444 | 1.4419
2019-03-29 14:07 | 831 | 0.037165 | 0 | 0 | 0 | 10.7912 | 1.4535
2019-03-29 14:07 | 832 | 0.035452 | 0 | 0 | 0 | 10.7634 | 1.4446
2019-03-29 14:07 | 833 | 0.062496 | 0 | 0 | 0 | 10.713 | 1.455
2019-03-29 14:07 | 834 | 0.062663 | 0 | 0 | 0 | 10.6446 | 1.4572
2019-03-29 14:07 | 835 | 0.050442 | 0 | 0 | 0 | 10.5912 | 1.4485
2019-03-29 14:07 | 836 | 0.051085 | 0 | 0 | 0 | 10.5407 | 1.402
2019-03-29 14:07 | 837 | 0.057946 | 0 | 0 | 0 | 10.4871 | 1.4024
2019-03-29 14:07 | 838 | 0.066272 | 0 | 0 | 0 | 10.4247 | 1.3985
2019-03-29 14:07 | 839 | 0.058621 | 0 | 0 | 0 | 10.3585 | 1.3874
2019-03-29 14:07 | 840 | 0.050775 | 0 | 0 | 0 | 10.3077 | 1.3852
2019-03-29 14:07 | 841 | 0.053465 | 0 | 0 | 0 | 10.2544 | 1.3769
2019-03-29 14:07 | 842 | 0.054641 | 0 | 0 | 0 | 10.201 | 1.3501
2019-03-29 14:07 | 843 | 0.052051 | 0 | 0 | 0 | 10.146 | 1.3453
2019-03-29 14:07 | 844 | 0.05151 | 0 | 0 | 0 | 10.0964 | 1.3154
2019-03-29 14:07 | 845 | 0.052502 | 0 | 0 | 0 | 10.0429 | 1.306
2019-03-29 14:07 | 846 | 0.042794 | 0 | 0 | 0 | 9.9938 | 1.2976
2019-03-29 14:07 | 847 | 0.042285 | 0 | 0 | 0 | 9.9551 | 1.2952
2019-03-29 14:07 | 848 | 0.048317 | 0 | 0 | 0 | 9.9077 | 1.2822
2019-03-29 14:07 | 849 | 0.043105 | 0 | 0 | 0 | 9.8612 | 1.2829
2019-03-29 14:07 | 850 | 0.032735 | 0 | 0 | 0 | 9.8224 | 1.2659
2019-03-29 14:07 | 851 | 0.045707 | 0 | 0 | 0 | 9.7903 | 1.2638
2019-03-29 14:07 | 852 | 0.046605 | 0 | 0 | 0 | 9.7334 | 1.2508
2019-03-29 14:07 | 853 | 0.044863 | 0 | 0 | 0 | 9.6979 | 1.2418
2019-03-29 14:07 | 854 | 0.05089 | 0 | 0 | 0 | 9.6423 | 1.2404
2019-03-29 14:07 | 855 | 0.041014 | 0 | 0 | 0 | 9.5994 | 1.2213
2019-03-29 14:07 | 856 | 0.050123 | 0 | 0 | 0 | 9.5561 | 1.2025
2019-03-29 14:07 | 857 | 0.052191 | 0 | 0 | 0 | 9.5008 | 1.1937
2019-03-29 14:07 | 858 | 0.040019 | 0 | 0 | 0 | 9.4552 | 1.1799
2019-03-29 14:07 | 859 | 0.041423 | 0 | 0 | 0 | 9.4179 | 1.1724
2019-03-29 14:07 | 860 | 0.038789 | 0 | 0 | 0 | 9.373 | 1.1425
2019-03-29 14:07 | 861 | 0.032828 | 0 | 0 | 0 | 9.3408 | 1.1358
2019-03-29 14:07 | 862 | 0.041294 | 0 | 0 | 0 | 9.3042 | 1.1469
2019-03-29 14:07 | 863 | 0.036421 | 0 | 0 | 0 | 9.2612 | 1.1296
2019-03-29 14:07 | 864 | 0.031012 | 0 | 0 | 0 | 9.2314 | 1.1362
2019-03-29 14:08 | 865 | 0.041372 | 0 | 0 | 0 | 9.1958 | 1.1333
2019-03-29 14:08 | 866 | 0.031741 | 0 | 0 | 0 | 9.1533 | 1.1387
2019-03-29 14:08 | 867 | 0.019957 | 0 | 0 | 0 | 9.1327 | 1.1297
2019-03-29 14:08 | 868 | 0.033698 | 0 | 0 | 0 | 9.1071 | 1.1125
2019-03-29 14:08 | 869 | 0.044873 | 0 | 0 | 0 | 9.0655 | 1.0946
2019-03-29 14:08 | 870 | 0.046068 | 0 | 0 | 0 | 9.0199 | 1.0881
2019-03-29 14:08 | 871 | 0.046215 | 0 | 0 | 0 | 8.9739 | 1.0639
2019-03-29 14:08 | 872 | 0.046474 | 0 | 0 | 0 | 8.9277 | 1.0687
2019-03-29 14:08 | 873 | 0.040989 | 0 | 0 | 0 | 8.8823 | 1.0654
2019-03-29 14:08 | 874 | 0.0405 | 0 | 0 | 0 | 8.8448 | 1.0462
2019-03-29 14:08 | 875 | 0.032333 | 0 | 0 | 0 | 8.8026 | 1.0317
2019-03-29 14:08 | 876 | 0.040599 | 0 | 0 | 0 | 8.776 | 1.0177
2019-03-29 14:08 | 877 | 0.056163 | 0 | 0 | 0 | 8.7197 | 1.0026
2019-03-29 14:08 | 878 | 0.050881 | 0 | 0 | 0 | 8.6687 | 0.99238
2019-03-29 14:08 | 879 | 0.046467 | 0 | 0 | 0 | 8.618 | 0.96506
2019-03-29 14:08 | 880 | 0.054772 | 0 | 0 | 0 | 8.5733 | 0.94524
2019-03-29 14:08 | 881 | 0.04063 | 0 | 0 | 0 | 8.514 | 0.94267
2019-03-29 14:08 | 882 | 0.017391 | 0 | 0 | 0 | 8.4935 | 0.93842
2019-03-29 14:08 | 883 | 0.04391 | 0 | 0 | 0 | 8.4672 | 0.93877
2019-03-29 14:08 | 884 | 0.057815 | 0 | 0 | 0 | 8.4081 | 0.94855
2019-03-29 14:08 | 885 | 0.048905 | 0 | 0 | 0 | 8.3581 | 0.94259
2019-03-29 14:08 | 886 | 0.02635 | 0 | 0 | 0 | 8.314 | 0.91987
2019-03-29 14:08 | 887 | 0.022066 | 0 | 0 | 0 | 8.3011 | 0.90433
2019-03-29 14:08 | 888 | 0.022194 | 0 | 0 | 0 | 8.2681 | 0.90217
2019-03-29 14:08 | 889 | 0.023164 | 0 | 0 | 0 | 8.2561 | 0.89967
2019-03-29 14:08 | 890 | 0.028875 | 0 | 0 | 0 | 8.2206 | 0.89912
2019-03-29 14:08 | 891 | 0.031651 | 0 | 0 | 0 | 8.1986 | 0.8956
2019-03-29 14:08 | 892 | 0.049896 | 0 | 0 | 0 | 8.1546 | 0.88054
2019-03-29 14:08 | 893 | 0.028211 | 0 | 0 | 0 | 8.1083 | 0.87633
2019-03-29 14:08 | 894 | 0.014146 | 0 | 0 | 0 | 8.0968 | 0.8648
2019-03-29 14:08 | 895 | 0.019826 | 0 | 0 | 0 | 8.0745 | 0.86086
2019-03-29 14:08 | 896 | 0.035903 | 0 | 0 | 0 | 8.0542 | 0.85188
2019-03-29 14:08 | 897 | 0.041451 | 0 | 0 | 0 | 8.0053 | 0.83913
2019-03-29 14:08 | 898 | 0.039111 | 0 | 0 | 0 | 7.9741 | 0.83785
2019-03-29 14:08 | 899 | 0.020993 | 0 | 0 | 0 | 7.9311 | 0.80698
2019-03-29 14:08 | 900 | 0.0097579 | 0 | 0 | 0 | 7.9299 | 0.79677
2019-03-29 14:08 | 901 | 0.029134 | 0 | 0 | 0 | 7.9041 | 0.79022
2019-03-29 14:08 | 902 | 0.031834 | 0 | 0 | 0 | 7.875 | 0.78029
2019-03-29 14:08 | 903 | 0.036544 | 0 | 0 | 0 | 7.841 | 0.78466

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062464 | 0.062685 | 0 | 0.2095 | 2019-03-29 10:39 | Group A | 0.1 | 0.22 | Do Nothing | 0.20949 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025175 | 0.025309 | 0 | 0.0261 | 2019-03-29 10:39 | Group A | -0.01 | 0.1 | Do Nothing | 0.00045503 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002731 | 0.002746 | 0 | 0.0011 | 2019-03-29 10:39 | Group A | -0.01 | 0.1 | Do Nothing | -6.7068e-07 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.7471 | 184.5289 | 0 | 100 | 2019-03-29 10:39 | None | 0 | 1000 | Do Nothing | 101.2189 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.4621 | 210.394 | 0 | 100 | 2019-03-29 10:39 | None | 0 | 1000 | Do Nothing | 101.1444 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -15.579 | 1447 | 0 |  | 2019-03-29 10:24 | None | -10 | 1000 | Do Nothing | 125.7175 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.1251 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0019293 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017963 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -151.7852 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23316 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4023 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0027761 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0041588 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0039794 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 5.5765e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0011442 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 75.0811 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3369 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5473 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 100.21 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.7512 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 37.7266 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 68.1152 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 88.4724 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 64.9177 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.009231 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0443 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.024813 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 23.7605 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN