## Cone Summary
Date: | 2019-03-19
Time: | 15:32
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B3S-Test3
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly TB 117 foam-2013 Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0416 |  | Start Test | 0 | 2019-03-19 15:42
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 6 | 2019-03-19 15:42
Sample Area | 1 | m^2 | Wetting | 251 | 2019-03-19 15:46
Cone Heat Flux | 75 | kW/m^2 | Dripping | 292 | 2019-03-19 15:47
Cone Orientation |  |  | Bottom Ignition | 586 | 2019-03-19 15:52
Sample Grid |  |  | Flame Out | 927 | 2019-03-19 15:58
Sample Thickness | 106 | mm
Sample Initial Mass | 48.9 | g
Cone Separation Distance | 25 | mm
Exhaust Flow Rate | 24.425 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 4.5475 | kW/m2 | Average Heat Release Rate | 0.95877 | kW/m2
Total Heat Release | 0.94692 | MJ/m2 | Average Mass Loss Rate | 0.044219 | g/m2s
Peak Mass Loss Rate | 0.19441 | g/m2s | Effective Heat of Combustion | 21.8018 | MJ/kg
Total Mass Lost | 41.8327 | g | Average Specific Extinction Area | 80.7876 | m2/kg
Time to Ignition | 5 | s | Peak CO Yield | 0.66945 | kg/kg
Time to Peak Heat Release | 25 | s | Peak CO2 Yield | 32.8825 | kg/kg
Time to Flame Out | 927 | s | Average CO Yield | 0.037087 | kg/kg
Average Heat Release Rate | 1.027 | kW/m2 | Average CO2 Yield | 1.7657 | kg/kg
Average Mass Loss Rate | 0.044541 | g/m2s
Effective Heat of Combustion | 22.6358 | MJ/kg
Dripping | 292 | s | Time to dripping
HRRD | 0.46872 | kW | HRR at dripping time
THRD | 0.31368 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.093 | kW | AHRR from Ignition to Dripping
Wetting | 251 | s | Time to wetting
HRRW | 0.48762 | kW | HRR at wetting time
THRW | 0.29376 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.1752 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-19 15:32 | 1553023906.7545 | -623 | 0.20878 | 3.8e-05 | 3e-06 | 100.4734 | 100.6397 | 131.3692 | 21.2612 | 0.001756 | 0.0016 | -152.27 | 0.23591 | 8.403 | 0.002974 | 0.004079 | 0.003761 | 4.3e-05 | 0.001276 | 0.29789 | 1.3381 | 2.5479 | 101.22 | 20.82 | 17.333 | 64.0903 | 93.3685 | 60.3277 | 0.004217 | 1.0587 | 0.02611 | 24.6628 | -0.014789 | 0.14034 | -2.6e-05 | -0.002244 | -0.002244 | -2.6013
2019-03-19 15:32 | 1553023907.7555 | -622 | 0.20878 | 3.8e-05 | 3e-06 | 100.4734 | 100.6397 | 131.3692 | 21.2612 | 0.001756 | 0.0016 | -152.27 | 0.23591 | 8.403 | 0.002974 | 0.004079 | 0.003761 | 4.3e-05 | 0.001276 | 0.29789 | 1.3381 | 2.5479 | 101.22 | 20.82 | 17.332 | 64.0903 | 93.3685 | 60.3277 | 0.004216 | 1.0587 | 0.02611 | 24.6628 | -0.014789 | 0.14034 | -2.6e-05 | -0.002244 | -0.002244 | -2.6019
2019-03-19 15:32 | 1553023908.754 | -621 | 0.20879 | 4.1e-05 | 3e-06 | 100.4665 | 100.3975 | 131.3692 | 21.2612 | 0.001725 | 0.001624 | -152.27 | 0.23627 | 8.4038 | 0.002991 | 0.004521 | 0.004228 | 4.5e-05 | 0.001275 | 0.29806 | 1.3582 | 2.548 | 101.22 | 20.82 | 17.334 | 63.3391 | 93.791 | 60.9069 | 0.004217 | 1.0568 | 0.026087 | 24.6842 | 0.006143 | 0.14055 | -3.1e-05 | -0.002596 | -0.002596 | -2.2524
2019-03-19 15:32 | 1553023909.7542 | -620 | 0.20879 | 3.7e-05 | 3e-06 | 100.4845 | 100.6386 | 131.3692 | 21.2612 | 0.00175 | 0.001667 | -152.27 | 0.23592 | 8.4008 | 0.002972 | 0.00429 | 0.004068 | 5.1e-05 | 0.00127 | 0.28584 | 1.3525 | 2.5482 | 101.22 | 20.82 | 17.337 | 63.9451 | 93.3326 | 61.9463 | 0.004218 | 1.0536 | 0.026047 | 24.7226 | -0.013709 | 0.14035 | -2.5e-05 | -0.002157 | -0.002157 | -2.7071
2019-03-19 15:32 | 1553023910.7556 | -619 | 0.2088 | 3.7e-05 | 3e-06 | 100.5148 | 100.4787 | 131.3692 | 21.2612 | 0.001733 | 0.001656 | -152.27 | 0.23636 | 8.4046 | 0.00296 | 0.003735 | 0.003499 | 4.5e-05 | 0.001272 | 0.28191 | 1.36 | 2.5482 | 101.22 | 20.82 | 17.339 | 63.5211 | 93.0251 | 61.6664 | 0.004218 | 1.0544 | 0.026058 | 24.7123 | 0.003215 | 0.14061 | -2.5e-05 | -0.002143 | -0.002143 | -2.7299
2019-03-19 15:32 | 1553023911.7552 | -618 | 0.20878 | 3.9e-05 | 3e-06 | 100.5304 | 100.6116 | 131.3692 | 21.2612 | 0.001792 | 0.001704 | -152.27 | 0.23609 | 8.4032 | 0.002929 | 0.004117 | 0.003926 | 4.6e-05 | 0.001276 | 0.29015 | 1.3322 | 2.548 | 101.22 | 20.82 | 17.34 | 64.9546 | 92.2985 | 62.8251 | 0.004218 | 1.0508 | 0.026013 | 24.755 | -0.007218 | 0.14045 | -2.8e-05 | -0.002337 | -0.002337 | -2.5
2019-03-19 15:32 | 1553023912.7543 | -617 | 0.20878 | 3.8e-05 | 3e-06 | 100.5787 | 100.603 | 131.3692 | 21.2612 | 0.001754 | 0.001632 | -152.27 | 0.23686 | 8.4021 | 0.002954 | 0.004525 | 0.004322 | 4.4e-05 | 0.00128 | 0.28998 | 1.3078 | 2.548 | 101.22 | 20.82 | 17.34 | 64.0224 | 92.8899 | 61.089 | 0.004218 | 1.0563 | 0.02608 | 24.691 | -0.002159 | 0.14091 | -2.7e-05 | -0.002299 | -0.002299 | -2.5493
2019-03-19 15:32 | 1553023913.7538 | -616 | 0.20877 | 3.5e-05 | 3e-06 | 100.5379 | 100.5157 | 131.3692 | 21.2612 | 0.001743 | 0.001675 | -152.27 | 0.23687 | 8.4038 | 0.002919 | 0.004733 | 0.004459 | 4.7e-05 | 0.001283 | 0.32012 | 1.3505 | 2.5481 | 101.22 | 20.82 | 17.341 | 63.774 | 92.0539 | 62.1213 | 0.004219 | 1.053 | 0.02604 | 24.7291 | 0.001979 | 0.14091 | -2.3e-05 | -0.002044 | -0.002044 | -2.8682
2019-03-19 15:32 | 1553023914.7538 | -615 | 0.20875 | 3.6e-05 | 3e-06 | 100.594 | 100.6219 | 131.3692 | 21.2612 | 0.001756 | 0.001655 | -152.27 | 0.23567 | 8.4023 | 0.002956 | 0.0035 | 0.003352 | 4.5e-05 | 0.001278 | 0.27936 | 1.3398 | 2.5481 | 101.22 | 20.82 | 17.336 | 64.091 | 92.946 | 61.6479 | 0.004217 | 1.0545 | 0.026058 | 24.7116 | -0.002473 | 0.1402 | -2.4e-05 | -0.002048 | -0.002048 | -2.848
2019-03-19 15:32 | 1553023915.7545 | -614 | 0.20877 | 3.8e-05 | 3e-06 | 100.5814 | 100.6777 | 131.3692 | 21.2612 | 0.001724 | 0.00168 | -152.27 | 0.2378 | 8.4021 | 0.002925 | 0.004504 | 0.004232 | 4.7e-05 | 0.001277 | 0.32659 | 1.3541 | 2.5481 | 101.22 | 20.82 | 17.333 | 63.3078 | 92.1954 | 62.2596 | 0.004217 | 1.0526 | 0.026035 | 24.7342 | -0.008555 | 0.14147 | -2.7e-05 | -0.002301 | -0.002301 | -2.558
2019-03-19 15:32 | 1553023916.7545 | -613 | 0.20875 | 3.8e-05 | 3e-06 | 100.6027 | 100.6986 | 131.3692 | 21.2612 | 0.001751 | 0.001641 | -152.27 | 0.23705 | 8.4037 | 0.002996 | 0.004193 | 0.003957 | 4.9e-05 | 0.001278 | 0.28495 | 1.3455 | 2.5481 | 101.22 | 20.82 | 17.326 | 63.9528 | 93.9053 | 61.3071 | 0.004215 | 1.0556 | 0.026072 | 24.699 | -0.008516 | 0.14102 | -2.7e-05 | -0.002343 | -0.002343 | -2.5034
2019-03-19 15:32 | 1553023917.7536 | -612 | 0.20877 | 3.3e-05 | 3e-06 | 100.6385 | 100.7578 | 131.176 | 21.2612 | 0.001803 | 0.001607 | -152.2613 | 0.23741 | 8.4062 | 0.00302 | 0.004256 | 0.004073 | 4.6e-05 | 0.001284 | 0.28031 | 1.3668 | 2.5481 | 101.22 | 20.82 | 17.323 | 65.2113 | 94.49 | 60.4925 | 0.004214 | 1.0582 | 0.026084 | 24.6508 | -0.010603 | 0.14123 | -2.1e-05 | -0.001865 | -0.001865 | -3.1506
2019-03-19 15:32 | 1553023918.756 | -611 | 0.20876 | 3.7e-05 | 3e-06 | 100.667 | 100.8246 | 131.0518 | 21.2612 | 0.001694 | 0.001628 | -152.2616 | 0.23729 | 8.404 | 0.003018 | 0.004299 | 0.004081 | 4.9e-05 | 0.001289 | 0.29391 | 1.3598 | 2.5482 | 101.22 | 20.82 | 17.318 | 62.5923 | 94.422 | 60.9888 | 0.004213 | 1.0566 | 0.026053 | 24.6574 | -0.013996 | 0.14116 | -2.5e-05 | -0.002131 | -0.002131 | -2.7552
2019-03-19 15:32 | 1553023919.755 | -610 | 0.20874 | 3.9e-05 | 3e-06 | 100.6796 | 100.7839 | 130.8251 | 21.2612 | 0.001728 | 0.001644 | -152.2519 | 0.23569 | 8.4034 | 0.00308 | 0.004218 | 0.003995 | 4.5e-05 | 0.001288 | 0.28407 | 1.3393 | 2.5479 | 101.22 | 20.83 | 17.31 | 63.4038 | 95.9181 | 61.3767 | 0.004214 | 1.0554 | 0.026015 | 24.6504 | -0.00926 | 0.14021 | -2.8e-05 | -0.002349 | -0.002349 | -2.4826
2019-03-19 15:32 | 1553023920.7551 | -609 | 0.20875 | 3.7e-05 | 3e-06 | 100.6991 | 100.7523 | 130.8462 | 21.2612 | 0.00176 | 0.001629 | -152.2443 | 0.23705 | 8.4032 | 0.003056 | 0.004084 | 0.003818 | 4.8e-05 | 0.00129 | 0.30431 | 1.3227 | 2.5479 | 101.22 | 20.82 | 17.307 | 64.1745 | 95.3502 | 61.0132 | 0.00421 | 1.0565 | 0.026031 | 24.639 | -0.00473 | 0.14102 | -2.4e-05 | -0.002079 | -0.002079 | -2.8217
2019-03-19 15:32 | 1553023921.7539 | -608 | 0.20874 | 3.6e-05 | 3e-06 | 100.7018 | 100.8763 | 130.3894 | 21.2612 | 0.001775 | 0.001624 | -152.2469 | 0.23744 | 8.4017 | 0.003074 | 0.004695 | 0.004413 | 4.8e-05 | 0.001287 | 0.30879 | 1.3888 | 2.5481 | 101.22 | 20.82 | 17.304 | 64.5352 | 95.7723 | 60.8958 | 0.00421 | 1.0569 | 0.02599 | 24.5916 | -0.01549 | 0.14125 | -2.4e-05 | -0.002069 | -0.002069 | -2.8394
2019-03-19 15:32 | 1553023922.7539 | -607 | 0.20876 | 3.7e-05 | 2e-06 | 100.742 | 100.8047 | 130.051 | 21.2612 | 0.001691 | 0.001659 | -152.2477 | 0.23615 | 8.4017 | 0.003059 | 0.004292 | 0.004149 | 5e-05 | 0.001284 | 0.28622 | 1.3616 | 2.5481 | 101.22 | 20.83 | 17.299 | 62.5158 | 95.4325 | 61.7408 | 0.004211 | 1.0542 | 0.025924 | 24.5907 | -0.00557 | 0.14048 | -2.5e-05 | -0.002068 | -0.002068 | -2.8258
2019-03-19 15:32 | 1553023923.7544 | -606 | 0.20878 | 3.9e-05 | 3e-06 | 100.745 | 100.8698 | 129.5911 | 21.2612 | 0.001736 | 0.001635 | -152.2491 | 0.23697 | 8.4018 | 0.003226 | 0.004153 | 0.003842 | 4.9e-05 | 0.001291 | 0.30469 | 1.3712 | 2.5481 | 101.22 | 20.83 | 17.296 | 63.5882 | 99.4496 | 61.1741 | 0.00421 | 1.056 | 0.0259 | 24.5264 | -0.011075 | 0.14098 | -2.8e-05 | -0.002436 | -0.002436 | -2.4079
2019-03-19 15:32 | 1553023924.754 | -605 | 0.20878 | 3.1e-05 | 3e-06 | 100.7867 | 100.9157 | 129.4104 | 21.2612 | 0.001774 | 0.001666 | -152.2384 | 0.23535 | 8.4035 | 0.003209 | 0.004556 | 0.00437 | 4.8e-05 | 0.001289 | 0.28982 | 1.3372 | 2.548 | 101.22 | 20.83 | 17.293 | 64.5251 | 99.0442 | 61.9229 | 0.004209 | 1.0536 | 0.025853 | 24.5368 | -0.011441 | 0.14001 | -1.8e-05 | -0.001569 | -0.001569 | -3.7128
2019-03-19 15:32 | 1553023925.7554 | -604 | 0.20874 | 3.7e-05 | 3e-06 | 100.7963 | 100.7989 | 129.2628 | 21.2612 | 0.001713 | 0.001587 | -152.2459 | 0.23615 | 8.4047 | 0.003177 | 0.004471 | 0.004285 | 4.8e-05 | 0.001293 | 0.33539 | 1.3353 | 2.548 | 101.22 | 20.83 | 17.292 | 63.0541 | 98.275 | 59.9985 | 0.004209 | 1.0597 | 0.025913 | 24.4522 | -0.000232 | 0.14048 | -2.6e-05 | -0.002195 | -0.002195 | -2.6629
2019-03-19 15:32 | 1553023926.7541 | -603 | 0.20877 | 3.5e-05 | 3e-06 | 100.8368 | 100.9359 | 128.7856 | 21.2612 | 0.001743 | 0.001625 | -152.2422 | 0.23621 | 8.4023 | 0.003152 | 0.004285 | 0.00401 | 4.6e-05 | 0.001295 | 0.29087 | 1.3193 | 2.548 | 101.22 | 20.83 | 17.292 | 63.7747 | 97.6753 | 60.9185 | 0.004209 | 1.0568 | 0.025829 | 24.4407 | -0.008785 | 0.14052 | -2.3e-05 | -0.002013 | -0.002013 | -2.9043
2019-03-19 15:32 | 1553023927.7545 | -602 | 0.20878 | 3.9e-05 | 2e-06 | 100.7978 | 100.8304 | 128.8168 | 21.2612 | 0.001751 | 0.001639 | -152.2441 | 0.23721 | 8.4017 | 0.003044 | 0.004913 | 0.004699 | 4.8e-05 | 0.001296 | 0.31045 | 1.3332 | 2.5481 | 101.22 | 20.83 | 17.294 | 63.9709 | 95.0568 | 61.2688 | 0.00421 | 1.0557 | 0.025819 | 24.4565 | -0.002894 | 0.14112 | -2.7e-05 | -0.00221 | -0.00221 | -2.6563
2019-03-19 15:32 | 1553023928.7551 | -601 | 0.20876 | 4e-05 | 3e-06 | 100.8125 | 100.9263 | 128.7964 | 21.2612 | 0.001744 | 0.001581 | -152.2286 | 0.23595 | 8.4044 | 0.00316 | 0.004941 | 0.004715 | 4.7e-05 | 0.00129 | 0.31996 | 1.3592 | 2.548 | 101.22 | 20.83 | 17.293 | 63.7959 | 97.8524 | 59.871 | 0.004209 | 1.0601 | 0.025871 | 24.4034 | -0.010092 | 0.14037 | -2.9e-05 | -0.002458 | -0.002458 | -2.3759
2019-03-19 15:32 | 1553023929.7543 | -600 | 0.20877 | 3.6e-05 | 3e-06 | 100.8446 | 100.9897 | 128.833 | 21.2612 | 0.001676 | 0.001646 | -152.2373 | 0.23634 | 8.4039 | 0.003141 | 0.004506 | 0.004365 | 4.8e-05 | 0.001295 | 0.3063 | 1.3277 | 2.5481 | 101.22 | 20.83 | 17.292 | 62.1618 | 97.3917 | 61.4203 | 0.004209 | 1.0552 | 0.025814 | 24.4636 | -0.012857 | 0.1406 | -2.4e-05 | -0.002082 | -0.002082 | -2.8098
2019-03-19 15:32 | 1553023930.7547 | -599 | 0.20876 | 3.6e-05 | 2e-06 | 100.8431 | 100.7503 | 128.6485 | 21.2612 | 0.001721 | 0.001647 | -152.2577 | 0.23645 | 8.4033 | 0.003175 | 0.004546 | 0.004351 | 5e-05 | 0.001289 | 0.30021 | 1.3789 | 2.5481 | 101.22 | 20.83 | 17.292 | 63.2414 | 98.2271 | 61.4453 | 0.004209 | 1.0551 | 0.025795 | 24.447 | 0.008238 | 0.14066 | -2.3e-05 | -0.001959 | -0.001959 | -2.9875
2019-03-19 15:32 | 1553023931.7547 | -598 | 0.20875 | 3.8e-05 | 3e-06 | 100.8602 | 100.9476 | 128.5572 | 21.2612 | 0.001712 | 0.001659 | -152.2511 | 0.23587 | 8.4013 | 0.003104 | 0.004018 | 0.003763 | 4.7e-05 | 0.001291 | 0.31509 | 1.3574 | 2.5481 | 101.22 | 20.83 | 17.29 | 63.0172 | 96.518 | 61.7461 | 0.004209 | 1.0542 | 0.025774 | 24.4493 | -0.007741 | 0.14032 | -2.6e-05 | -0.002217 | -0.002217 | -2.633
2019-03-19 15:32 | 1553023932.7534 | -597 | 0.20878 | 3.7e-05 | 2e-06 | 100.9278 | 100.889 | 128.514 | 21.2612 | 0.001788 | 0.001597 | -152.2383 | 0.23621 | 8.4027 | 0.00312 | 0.004408 | 0.004197 | 5.1e-05 | 0.001289 | 0.30901 | 1.3731 | 2.5481 | 101.22 | 20.83 | 17.283 | 64.8468 | 96.9003 | 60.2448 | 0.004207 | 1.0589 | 0.025828 | 24.3903 | 0.003436 | 0.14052 | -2.5e-05 | -0.002112 | -0.002112 | -2.7681
2019-03-19 15:32 | 1553023933.7544 | -596 | 0.20878 | 3.9e-05 | 3e-06 | 100.8734 | 100.9171 | 128.6814 | 21.2612 | 0.001726 | 0.001574 | -152.2262 | 0.23618 | 8.4018 | 0.003135 | 0.003957 | 0.003668 | 4.8e-05 | 0.001285 | 0.31996 | 1.3678 | 2.548 | 101.22 | 20.83 | 17.277 | 63.3579 | 97.2623 | 59.6888 | 0.004206 | 1.0607 | 0.025866 | 24.3858 | -0.003869 | 0.1405 | -2.7e-05 | -0.002309 | -0.002309 | -2.5312
2019-03-19 15:32 | 1553023934.7549 | -595 | 0.20876 | 3.4e-05 | 2e-06 | 100.9248 | 101.0072 | 129.0323 | 21.2612 | 0.001734 | 0.001657 | -152.2334 | 0.23676 | 8.4001 | 0.003162 | 0.004604 | 0.004327 | 4.6e-05 | 0.001282 | 0.31636 | 1.3184 | 2.5481 | 101.22 | 20.83 | 17.27 | 63.5406 | 97.9121 | 61.7057 | 0.004204 | 1.0543 | 0.025823 | 24.4929 | -0.0073 | 0.14085 | -2.1e-05 | -0.001818 | -0.001818 | -3.2228
2019-03-19 15:32 | 1553023935.7536 | -594 | 0.20874 | 3.3e-05 | 2e-06 | 100.9443 | 100.8924 | 128.8271 | 21.2612 | 0.001749 | 0.00166 | -152.2018 | 0.2361 | 8.4054 | 0.003103 | 0.004534 | 0.004309 | 5e-05 | 0.001279 | 0.30596 | 1.3329 | 2.548 | 101.22 | 20.83 | 17.254 | 63.9037 | 96.4835 | 61.7782 | 0.0042 | 1.0541 | 0.0258 | 24.4761 | 0.004595 | 0.14045 | -1.9e-05 | -0.001648 | -0.001648 | -3.5458
2019-03-19 15:32 | 1553023936.755 | -593 | 0.20878 | 3.7e-05 | 3e-06 | 100.9533 | 101.0695 | 128.9316 | 21.2612 | 0.001736 | 0.001603 | -152.1854 | 0.23597 | 8.4041 | 0.003165 | 0.004683 | 0.004495 | 4.6e-05 | 0.001282 | 0.33229 | 1.3474 | 2.5481 | 101.22 | 20.83 | 17.244 | 63.5897 | 97.9841 | 60.405 | 0.004198 | 1.0584 | 0.025864 | 24.4358 | -0.010292 | 0.14038 | -2.5e-05 | -0.002138 | -0.002138 | -2.7315
2019-03-19 15:32 | 1553023937.7547 | -592 | 0.20876 | 3.5e-05 | 3e-06 | 100.9875 | 101.0941 | 129.2557 | 21.2612 | 0.00176 | 0.001654 | -152.196 | 0.23586 | 8.4031 | 0.003124 | 0.004593 | 0.004366 | 4.6e-05 | 0.001268 | 0.30381 | 1.3187 | 2.548 | 101.22 | 20.83 | 17.222 | 64.1668 | 96.9896 | 61.6291 | 0.004192 | 1.0546 | 0.025849 | 24.5113 | -0.009443 | 0.14031 | -2.3e-05 | -0.001956 | -0.001956 | -2.9845
2019-03-19 15:32 | 1553023938.7538 | -591 | 0.20876 | 3.8e-05 | 3e-06 | 100.9797 | 101.2657 | 129.3167 | 21.2612 | 0.001766 | 0.001636 | -152.2074 | 0.23603 | 8.4027 | 0.003052 | 0.004517 | 0.004279 | 5e-05 | 0.001261 | 0.29734 | 1.2842 | 2.548 | 101.22 | 20.83 | 17.21 | 64.3112 | 95.2445 | 61.1844 | 0.004189 | 1.056 | 0.025872 | 24.5008 | -0.025303 | 0.14042 | -2.7e-05 | -0.002319 | -0.002319 | -2.5184
2019-03-19 15:32 | 1553023939.7548 | -590 | 0.2088 | 4e-05 | 3e-06 | 101.0397 | 101.1818 | 129.4696 | 21.2612 | 0.001749 | 0.001601 | -152.198 | 0.23608 | 8.4048 | 0.003068 | 0.004078 | 0.003858 | 5e-05 | 0.001249 | 0.30237 | 1.2926 | 2.548 | 101.22 | 20.83 | 17.198 | 63.9013 | 95.6463 | 60.3487 | 0.004186 | 1.0586 | 0.02592 | 24.4846 | -0.012573 | 0.14044 | -3e-05 | -0.002535 | -0.002535 | -2.3048
2019-03-19 15:32 | 1553023940.7548 | -589 | 0.20877 | 3.8e-05 | 3e-06 | 101.0328 | 101.1123 | 129.7306 | 21.2612 | 0.001717 | 0.001638 | -152.1911 | 0.23689 | 8.4033 | 0.003099 | 0.003618 | 0.003408 | 4.8e-05 | 0.001249 | 0.2855 | 1.3129 | 2.5479 | 101.22 | 20.83 | 17.183 | 63.1383 | 96.3864 | 61.2288 | 0.004183 | 1.0558 | 0.025912 | 24.5416 | -0.007036 | 0.14092 | -2.6e-05 | -0.002217 | -0.002217 | -2.6445
2019-03-19 15:32 | 1553023941.7539 | -588 | 0.20878 | 3.4e-05 | 3e-06 | 101.0421 | 101.2496 | 129.9263 | 21.2612 | 0.001804 | 0.001661 | -152.2056 | 0.23678 | 8.4041 | 0.003121 | 0.004362 | 0.004094 | 4.4e-05 | 0.001252 | 0.32156 | 1.3236 | 2.548 | 101.22 | 20.83 | 17.177 | 65.2307 | 96.9097 | 61.7977 | 0.004181 | 1.054 | 0.025909 | 24.581 | -0.018354 | 0.14086 | -2.2e-05 | -0.001905 | -0.001905 | -3.0767
2019-03-19 15:32 | 1553023942.7543 | -587 | 0.20878 | 3.8e-05 | 3e-06 | 101.0625 | 101.2041 | 129.7486 | 21.2612 | 0.001707 | 0.001636 | -152.2149 | 0.23719 | 8.404 | 0.003066 | 0.004288 | 0.004111 | 4.6e-05 | 0.001248 | 0.30127 | 1.3482 | 2.5481 | 101.22 | 20.83 | 17.17 | 62.9041 | 95.5937 | 61.182 | 0.00418 | 1.056 | 0.025915 | 24.5416 | -0.012522 | 0.14111 | -2.7e-05 | -0.002355 | -0.002355 | -2.493
2019-03-19 15:32 | 1553023943.754 | -586 | 0.20877 | 3.5e-05 | 3e-06 | 101.0823 | 101.2366 | 129.5203 | 21.2612 | 0.001733 | 0.001585 | -152.2263 | 0.23733 | 8.4019 | 0.003108 | 0.004834 | 0.004657 | 4.6e-05 | 0.001253 | 0.33124 | 1.3647 | 2.5481 | 101.22 | 20.83 | 17.169 | 63.5304 | 96.5979 | 59.9532 | 0.004179 | 1.0599 | 0.02594 | 24.4749 | -0.013645 | 0.14119 | -2.3e-05 | -0.002002 | -0.002002 | -2.9339
2019-03-19 15:32 | 1553023944.7544 | -585 | 0.20878 | 3.7e-05 | 3e-06 | 101.157 | 101.1555 | 129.2294 | 21.2612 | 0.001751 | 0.001632 | -152.2231 | 0.23697 | 8.4011 | 0.00313 | 0.004645 | 0.004458 | 4.4e-05 | 0.001254 | 0.33859 | 1.3447 | 2.548 | 101.22 | 20.83 | 17.168 | 63.9584 | 97.1261 | 61.1037 | 0.004179 | 1.0562 | 0.025866 | 24.4896 | 0.00014 | 0.14097 | -2.6e-05 | -0.002204 | -0.002204 | -2.6605
2019-03-19 15:32 | 1553023945.7531 | -584 | 0.20877 | 3.8e-05 | 3e-06 | 101.1507 | 101.3031 | 129.5994 | 21.2612 | 0.001716 | 0.001636 | -152.2301 | 0.23716 | 8.404 | 0.003131 | 0.004158 | 0.003924 | 4.7e-05 | 0.001261 | 0.31885 | 1.3282 | 2.5481 | 101.22 | 20.82 | 17.169 | 63.1109 | 97.1582 | 61.1859 | 0.004177 | 1.056 | 0.0259 | 24.5276 | -0.013461 | 0.14108 | -2.6e-05 | -0.002231 | -0.002231 | -2.6306
2019-03-19 15:32 | 1553023946.7531 | -583 | 0.20875 | 4.2e-05 | 3e-06 | 101.0793 | 101.1274 | 129.7613 | 21.2612 | 0.001744 | 0.001616 | -152.2435 | 0.23704 | 8.4018 | 0.003077 | 0.004307 | 0.004081 | 4.2e-05 | 0.001262 | 0.28639 | 1.3576 | 2.5482 | 101.22 | 20.82 | 17.17 | 63.7865 | 95.8578 | 60.7067 | 0.004177 | 1.0575 | 0.025935 | 24.5254 | -0.004252 | 0.14101 | -3.2e-05 | -0.002698 | -0.002698 | -2.1745
2019-03-19 15:32 | 1553023947.7532 | -582 | 0.20877 | 3.8e-05 | 3e-06 | 101.1054 | 101.0945 | 129.3355 | 21.2612 | 0.001779 | 0.001629 | -152.2136 | 0.23757 | 8.4033 | 0.003175 | 0.00436 | 0.004118 | 4.4e-05 | 0.001258 | 0.30093 | 1.3593 | 2.548 | 101.22 | 20.82 | 17.174 | 64.6329 | 98.2279 | 61.0132 | 0.004178 | 1.0565 | 0.025881 | 24.4963 | 0.000967 | 0.14133 | -2.6e-05 | -0.002262 | -0.002262 | -2.5995
2019-03-19 15:32 | 1553023948.7543 | -581 | 0.20879 | 3.8e-05 | 2e-06 | 101.2117 | 101.3685 | 129.3299 | 21.2612 | 0.00173 | 0.001594 | -152.2228 | 0.23633 | 8.4045 | 0.003161 | 0.004113 | 0.003912 | 4.6e-05 | 0.001253 | 0.29031 | 1.3681 | 2.5481 | 101.22 | 20.82 | 17.176 | 63.4501 | 97.8877 | 60.1703 | 0.004178 | 1.0592 | 0.025913 | 24.4649 | -0.013849 | 0.14059 | -2.6e-05 | -0.002166 | -0.002166 | -2.7003
2019-03-19 15:32 | 1553023949.7543 | -580 | 0.20875 | 4.3e-05 | 3e-06 | 101.151 | 101.3044 | 129.3874 | 21.2612 | 0.001679 | 0.00163 | -152.2428 | 0.23617 | 8.4024 | 0.003131 | 0.004424 | 0.004172 | 4.6e-05 | 0.001255 | 0.31188 | 1.37 | 2.5482 | 101.22 | 20.82 | 17.179 | 62.2206 | 97.1621 | 61.0459 | 0.004179 | 1.0564 | 0.025884 | 24.5024 | -0.013555 | 0.1405 | -3.3e-05 | -0.002709 | -0.002709 | -2.1578
2019-03-19 15:32 | 1553023950.7547 | -579 | 0.20877 | 3.6e-05 | 3e-06 | 101.1483 | 101.3315 | 129.321 | 21.2612 | 0.001744 | 0.001721 | -152.248 | 0.23626 | 8.4018 | 0.003086 | 0.004685 | 0.004436 | 4.6e-05 | 0.001254 | 0.3063 | 1.3129 | 2.5481 | 101.22 | 20.82 | 17.186 | 63.7873 | 96.0629 | 63.2288 | 0.004181 | 1.0495 | 0.025794 | 24.576 | -0.016182 | 0.14055 | -2.4e-05 | -0.002072 | -0.002072 | -2.8224
2019-03-19 15:32 | 1553023951.7548 | -578 | 0.20876 | 3.9e-05 | 3e-06 | 101.1609 | 101.2668 | 129.2064 | 21.2612 | 0.001721 | 0.001599 | -152.2303 | 0.23701 | 8.4018 | 0.003129 | 0.004113 | 0.003909 | 4.7e-05 | 0.00125 | 0.30431 | 1.2994 | 2.5481 | 101.22 | 20.82 | 17.188 | 63.2406 | 97.1135 | 60.2877 | 0.004181 | 1.0588 | 0.025896 | 24.4575 | -0.009352 | 0.141 | -2.8e-05 | -0.002424 | -0.002424 | -2.4196
2019-03-19 15:33 | 1553023952.7545 | -577 | 0.20874 | 4.1e-05 | 3e-06 | 101.2267 | 101.3596 | 129.528 | 21.2612 | 0.001796 | 0.001578 | -152.2369 | 0.23693 | 8.4033 | 0.003119 | 0.004486 | 0.004082 | 4.7e-05 | 0.001249 | 0.30342 | 1.2979 | 2.5481 | 101.22 | 20.82 | 17.188 | 65.0402 | 96.8597 | 59.7874 | 0.004181 | 1.0604 | 0.025947 | 24.4695 | -0.011737 | 0.14095 | -2.9e-05 | -0.002478 | -0.002478 | -2.3659
2019-03-19 15:33 | 1553023953.755 | -576 | 0.20877 | 3.4e-05 | 3e-06 | 101.1294 | 101.2952 | 129.85 | 21.2612 | 0.001713 | 0.001689 | -152.2287 | 0.23649 | 8.4022 | 0.003043 | 0.004844 | 0.004703 | 4.7e-05 | 0.001249 | 0.28838 | 1.2942 | 2.5479 | 101.22 | 20.82 | 17.188 | 63.0392 | 95.0348 | 62.4743 | 0.004181 | 1.0519 | 0.025875 | 24.5986 | -0.014649 | 0.14069 | -2.2e-05 | -0.001875 | -0.001875 | -3.1214
2019-03-19 15:33 | 1553023954.7547 | -575 | 0.20877 | 4.2e-05 | 3e-06 | 101.1189 | 101.2544 | 129.6893 | 21.2612 | 0.001734 | 0.001657 | -152.2248 | 0.23694 | 8.4051 | 0.003078 | 0.004946 | 0.004688 | 4.8e-05 | 0.001254 | 0.2808 | 1.3012 | 2.5479 | 101.22 | 20.82 | 17.185 | 63.542 | 95.871 | 61.6985 | 0.004181 | 1.0543 | 0.025889 | 24.5549 | -0.011979 | 0.14095 | -3.1e-05 | -0.002613 | -0.002613 | -2.244
2019-03-19 15:33 | 1553023955.7537 | -574 | 0.20877 | 3.4e-05 | 2e-06 | 101.1627 | 101.2863 | 129.5557 | 21.2612 | 0.001735 | 0.001616 | -152.2319 | 0.23697 | 8.4027 | 0.003051 | 0.004302 | 0.004103 | 4.3e-05 | 0.001254 | 0.29806 | 1.3398 | 2.548 | 101.22 | 20.82 | 17.18 | 63.5844 | 95.2235 | 60.7154 | 0.004179 | 1.0574 | 0.025914 | 24.5062 | -0.010917 | 0.14098 | -2.1e-05 | -0.001778 | -0.001778 | -3.2983
2019-03-19 15:33 | 1553023956.7538 | -573 | 0.20877 | 3.5e-05 | 3e-06 | 101.1988 | 101.4003 | 129.3064 | 21.2612 | 0.001799 | 0.001622 | -152.2145 | 0.23641 | 8.4056 | 0.003099 | 0.004333 | 0.004016 | 4.7e-05 | 0.001258 | 0.33356 | 1.326 | 2.548 | 101.22 | 20.82 | 17.178 | 65.1114 | 96.3973 | 60.8592 | 0.004179 | 1.057 | 0.025884 | 24.4879 | -0.0178 | 0.14064 | -2.3e-05 | -0.001997 | -0.001997 | -2.93
2019-03-19 15:33 | 1553023957.7547 | -572 | 0.20875 | 3.9e-05 | 3e-06 | 101.1816 | 101.1168 | 129.0987 | 21.2612 | 0.001747 | 0.001614 | -152.2242 | 0.23627 | 8.4014 | 0.00306 | 0.004578 | 0.004384 | 4.5e-05 | 0.001256 | 0.31797 | 1.3557 | 2.5481 | 101.22 | 20.82 | 17.176 | 63.8583 | 95.4436 | 60.6715 | 0.004178 | 1.0576 | 0.02587 | 24.4614 | 0.005738 | 0.14055 | -2.8e-05 | -0.002317 | -0.002317 | -2.5235
2019-03-19 15:33 | 1553023958.7538 | -571 | 0.20875 | 4.2e-05 | 3e-06 | 101.2474 | 101.3363 | 128.8243 | 21.2612 | 0.001727 | 0.001599 | -152.2139 | 0.23601 | 8.4051 | 0.003058 | 0.00451 | 0.004251 | 4.4e-05 | 0.001254 | 0.30989 | 1.3714 | 2.548 | 101.22 | 20.82 | 17.174 | 63.3742 | 95.391 | 60.2995 | 0.004178 | 1.0588 | 0.025857 | 24.4217 | -0.007852 | 0.1404 | -3.1e-05 | -0.002604 | -0.002604 | -2.2434
2019-03-19 15:33 | 1553023959.7532 | -570 | 0.20873 | 3.9e-05 | 3e-06 | 101.1597 | 101.4086 | 128.6168 | 21.2612 | 0.001702 | 0.00162 | -152.2048 | 0.23623 | 8.4038 | 0.003129 | 0.004844 | 0.004651 | 4.7e-05 | 0.001253 | 0.28982 | 1.3422 | 2.5479 | 101.22 | 20.82 | 17.174 | 62.7914 | 97.1205 | 60.806 | 0.004178 | 1.0572 | 0.025817 | 24.4206 | -0.021974 | 0.14053 | -2.8e-05 | -0.002369 | -0.002369 | -2.4681
2019-03-19 15:33 | 1553023960.7539 | -569 | 0.20872 | 3.7e-05 | 4e-06 | 101.2201 | 101.3938 | 128.4077 | 21.2612 | 0.001776 | 0.001604 | -152.2029 | 0.23627 | 8.4018 | 0.003109 | 0.004581 | 0.004485 | 4.8e-05 | 0.001253 | 0.30757 | 1.3283 | 2.548 | 101.22 | 20.82 | 17.175 | 64.5571 | 96.6371 | 60.426 | 0.004178 | 1.0584 | 0.02581 | 24.3868 | -0.015343 | 0.14055 | -2.6e-05 | -0.002264 | -0.002264 | -2.5832
2019-03-19 15:33 | 1553023961.7529 | -568 | 0.20875 | 3.6e-05 | 3e-06 | 101.2891 | 101.5825 | 128.3854 | 21.2612 | 0.00167 | 0.001613 | -152.2181 | 0.23673 | 8.4009 | 0.003126 | 0.003954 | 0.003664 | 4.8e-05 | 0.001256 | 0.3079 | 1.3219 | 2.548 | 101.22 | 20.82 | 17.175 | 61.9985 | 97.0483 | 60.6443 | 0.004178 | 1.0577 | 0.0258 | 24.3927 | -0.025877 | 0.14083 | -2.4e-05 | -0.002044 | -0.002044 | -2.8661
2019-03-19 15:33 | 1553023962.7534 | -567 | 0.20879 | 4.1e-05 | 3e-06 | 101.209 | 101.4555 | 128.3544 | 21.2612 | 0.001761 | 0.001669 | -152.2049 | 0.23738 | 8.4036 | 0.003124 | 0.004409 | 0.004234 | 4.7e-05 | 0.001256 | 0.29286 | 1.3391 | 2.5481 | 101.22 | 20.82 | 17.175 | 64.2096 | 96.9865 | 61.9838 | 0.004178 | 1.0534 | 0.025745 | 24.4387 | -0.02176 | 0.14122 | -3e-05 | -0.002463 | -0.002463 | -2.3852
2019-03-19 15:33 | 1553023963.7534 | -566 | 0.20875 | 3.7e-05 | 2e-06 | 101.2336 | 101.3229 | 128.1812 | 21.2612 | 0.001719 | 0.001593 | -152.2151 | 0.23701 | 8.4017 | 0.003111 | 0.004666 | 0.004399 | 4.7e-05 | 0.00126 | 0.31006 | 1.3297 | 2.5481 | 101.22 | 20.82 | 17.177 | 63.1829 | 96.6685 | 60.1501 | 0.004179 | 1.0592 | 0.025798 | 24.3553 | -0.007893 | 0.141 | -2.5e-05 | -0.002073 | -0.002073 | -2.8289
2019-03-19 15:33 | 1553023964.7541 | -565 | 0.20872 | 4.2e-05 | 3e-06 | 101.2432 | 101.3842 | 128.6722 | 21.2612 | 0.001736 | 0.001669 | -152.2266 | 0.23742 | 8.4021 | 0.00309 | 0.004908 | 0.004614 | 4.5e-05 | 0.001258 | 0.29302 | 1.3518 | 2.548 | 101.22 | 20.82 | 17.177 | 63.599 | 96.182 | 61.9822 | 0.004179 | 1.0535 | 0.025777 | 24.4688 | -0.012455 | 0.14124 | -3.1e-05 | -0.00259 | -0.00259 | -2.2689
2019-03-19 15:33 | 1553023965.7532 | -564 | 0.20879 | 3.3e-05 | 3e-06 | 101.2648 | 101.4555 | 128.7198 | 21.2612 | 0.001691 | 0.001611 | -152.2463 | 0.23699 | 8.4019 | 0.003128 | 0.004536 | 0.004364 | 5.1e-05 | 0.001262 | 0.28783 | 1.323 | 2.5481 | 101.22 | 20.82 | 17.18 | 62.5071 | 97.0954 | 60.5817 | 0.004179 | 1.0579 | 0.025836 | 24.4222 | -0.016829 | 0.14098 | -1.9e-05 | -0.001694 | -0.001694 | -3.4613
2019-03-19 15:33 | 1553023966.7542 | -563 | 0.20878 | 3.9e-05 | 3e-06 | 101.2762 | 101.5572 | 128.8898 | 21.2612 | 0.001796 | 0.001632 | -152.2522 | 0.23614 | 8.4021 | 0.003136 | 0.004687 | 0.004467 | 4.9e-05 | 0.001262 | 0.31166 | 1.3078 | 2.548 | 101.22 | 20.82 | 17.183 | 65.0426 | 97.2797 | 61.1006 | 0.00418 | 1.0562 | 0.025833 | 24.4573 | -0.024785 | 0.14048 | -2.8e-05 | -0.002325 | -0.002325 | -2.5137
2019-03-19 15:33 | 1553023967.7532 | -562 | 0.20877 | 3.8e-05 | 3e-06 | 101.3215 | 101.4031 | 129.0616 | 21.2612 | 0.001752 | 0.001642 | -152.2686 | 0.23598 | 8.4033 | 0.003122 | 0.004902 | 0.004608 | 4.7e-05 | 0.001262 | 0.30486 | 1.3171 | 2.548 | 101.22 | 20.82 | 17.186 | 63.9786 | 96.9457 | 61.3446 | 0.004181 | 1.0555 | 0.02584 | 24.4825 | -0.007199 | 0.14038 | -2.6e-05 | -0.002207 | -0.002207 | -2.6458
2019-03-19 15:33 | 1553023968.7544 | -561 | 0.20877 | 3.9e-05 | 3e-06 | 101.2948 | 101.5976 | 129.2672 | 21.2612 | 0.001703 | 0.001592 | -152.2577 | 0.23622 | 8.4054 | 0.003089 | 0.004127 | 0.003975 | 4.9e-05 | 0.001255 | 0.31277 | 1.3017 | 2.548 | 101.22 | 20.82 | 17.191 | 62.8008 | 96.1547 | 60.1344 | 0.004182 | 1.0593 | 0.025908 | 24.4576 | -0.0267 | 0.14053 | -2.8e-05 | -0.002389 | -0.002389 | -2.4468
2019-03-19 15:33 | 1553023969.7544 | -560 | 0.20879 | 3.4e-05 | 3e-06 | 101.3212 | 101.5935 | 129.3714 | 21.2612 | 0.001741 | 0.001598 | -152.2479 | 0.23686 | 8.4037 | 0.003113 | 0.004366 | 0.004034 | 4.6e-05 | 0.001258 | 0.3115 | 1.3227 | 2.5481 | 101.22 | 20.82 | 17.194 | 63.7256 | 96.7327 | 60.2744 | 0.004183 | 1.0588 | 0.025913 | 24.4726 | -0.024007 | 0.14091 | -2.2e-05 | -0.001889 | -0.001889 | -3.1024
2019-03-19 15:33 | 1553023970.7538 | -559 | 0.20877 | 3.8e-05 | 3e-06 | 101.3152 | 101.4274 | 129.4562 | 21.2612 | 0.001678 | 0.001643 | -152.232 | 0.23706 | 8.4025 | 0.003165 | 0.004454 | 0.004218 | 4.7e-05 | 0.001255 | 0.30541 | 1.3598 | 2.548 | 101.22 | 20.82 | 17.197 | 62.2063 | 97.9841 | 61.3547 | 0.004184 | 1.0554 | 0.025879 | 24.5203 | -0.0099 | 0.14103 | -2.5e-05 | -0.002145 | -0.002145 | -2.735
2019-03-19 15:33 | 1553023971.7535 | -558 | 0.20881 | 3.9e-05 | 3e-06 | 101.3239 | 101.5452 | 129.5088 | 21.2612 | 0.001747 | 0.00162 | -152.2357 | 0.2367 | 8.4044 | 0.003113 | 0.004395 | 0.004187 | 4.9e-05 | 0.001253 | 0.28241 | 1.3458 | 2.548 | 101.22 | 20.82 | 17.203 | 63.8717 | 96.7194 | 60.8067 | 0.004185 | 1.0572 | 0.025906 | 24.5052 | -0.019516 | 0.14081 | -2.7e-05 | -0.002305 | -0.002305 | -2.5412
2019-03-19 15:33 | 1553023972.7545 | -557 | 0.20879 | 3.6e-05 | 3e-06 | 101.4091 | 101.6493 | 129.5059 | 21.2612 | 0.001692 | 0.001579 | -152.249 | 0.23701 | 8.4036 | 0.003117 | 0.004497 | 0.004247 | 4.7e-05 | 0.001255 | 0.31349 | 1.3304 | 2.5481 | 101.22 | 20.81 | 17.205 | 62.5502 | 96.8322 | 59.8163 | 0.004183 | 1.0603 | 0.025944 | 24.4685 | -0.021161 | 0.141 | -2.4e-05 | -0.002077 | -0.002077 | -2.8247
2019-03-19 15:33 | 1553023973.7539 | -556 | 0.20876 | 3.8e-05 | 3e-06 | 101.3635 | 101.5822 | 129.7712 | 21.2612 | 0.001728 | 0.001602 | -152.2427 | 0.23633 | 8.406 | 0.003111 | 0.004856 | 0.004662 | 4.7e-05 | 0.001257 | 0.30127 | 1.3447 | 2.5481 | 101.22 | 20.81 | 17.21 | 63.4007 | 96.6841 | 60.3643 | 0.004184 | 1.0586 | 0.025949 | 24.5137 | -0.019277 | 0.14059 | -2.6e-05 | -0.002226 | -0.002226 | -2.6279
2019-03-19 15:33 | 1553023974.7539 | -555 | 0.20882 | 3.3e-05 | 3e-06 | 101.3629 | 101.4568 | 129.5904 | 21.2612 | 0.001732 | 0.00158 | -152.2429 | 0.2371 | 8.4011 | 0.003142 | 0.004834 | 0.004708 | 5.2e-05 | 0.001259 | 0.31061 | 1.379 | 2.5479 | 101.22 | 20.81 | 17.212 | 63.5085 | 97.4159 | 59.8389 | 0.004185 | 1.0602 | 0.025952 | 24.4773 | -0.008286 | 0.14105 | -2e-05 | -0.001745 | -0.001745 | -3.3625
2019-03-19 15:33 | 1553023975.7537 | -554 | 0.20878 | 3.8e-05 | 3e-06 | 101.4163 | 101.4521 | 129.4374 | 21.2612 | 0.001743 | 0.001624 | -152.2276 | 0.23671 | 8.4009 | 0.003112 | 0.004287 | 0.004007 | 4.9e-05 | 0.001256 | 0.33196 | 1.3597 | 2.5482 | 101.22 | 20.82 | 17.212 | 63.7678 | 96.7052 | 60.8982 | 0.004187 | 1.0569 | 0.025895 | 24.5018 | -0.00315 | 0.14082 | -2.6e-05 | -0.002232 | -0.002232 | -2.624
2019-03-19 15:33 | 1553023976.7528 | -553 | 0.20881 | 3.7e-05 | 3e-06 | 101.4085 | 101.5654 | 129.0851 | 21.2612 | 0.001722 | 0.001598 | -152.2403 | 0.23645 | 8.402 | 0.003093 | 0.004517 | 0.004288 | 4.6e-05 | 0.001255 | 0.30718 | 1.3111 | 2.548 | 101.22 | 20.81 | 17.21 | 63.2625 | 96.2385 | 60.2682 | 0.004184 | 1.0589 | 0.025884 | 24.4453 | -0.013828 | 0.14066 | -2.6e-05 | -0.002214 | -0.002214 | -2.6426
2019-03-19 15:33 | 1553023977.7531 | -552 | 0.20877 | 4e-05 | 3e-06 | 101.3638 | 101.6904 | 128.791 | 21.2612 | 0.00173 | 0.001655 | -152.2388 | 0.23745 | 8.4003 | 0.003077 | 0.004473 | 0.004152 | 5.1e-05 | 0.001256 | 0.31437 | 1.3268 | 2.5482 | 101.22 | 20.82 | 17.207 | 63.4554 | 95.8624 | 61.6602 | 0.004186 | 1.0545 | 0.025801 | 24.4684 | -0.028775 | 0.14126 | -2.9e-05 | -0.002474 | -0.002474 | -2.3749
2019-03-19 15:33 | 1553023978.7529 | -551 | 0.20878 | 4.4e-05 | 3e-06 | 101.4022 | 101.5613 | 128.7668 | 21.2612 | 0.001753 | 0.001649 | -152.2402 | 0.23674 | 8.403 | 0.00313 | 0.004232 | 0.004018 | 5.1e-05 | 0.001252 | 0.30021 | 1.2882 | 2.5479 | 101.22 | 20.82 | 17.194 | 64.0036 | 97.1355 | 61.5111 | 0.004183 | 1.0549 | 0.025804 | 24.4606 | -0.014022 | 0.14084 | -3.4e-05 | -0.00284 | -0.00284 | -2.0626
2019-03-19 15:33 | 1553023979.7539 | -550 | 0.20875 | 3.7e-05 | 3e-06 | 101.395 | 101.3507 | 128.6485 | 21.2612 | 0.001736 | 0.001635 | -152.254 | 0.23595 | 8.4026 | 0.003044 | 0.004511 | 0.004331 | 4.6e-05 | 0.001248 | 0.28279 | 1.3344 | 2.5479 | 101.22 | 20.82 | 17.185 | 63.606 | 95.0708 | 61.1782 | 0.004181 | 1.056 | 0.025805 | 24.4372 | 0.003912 | 0.14037 | -2.5e-05 | -0.00213 | -0.00213 | -2.7409
2019-03-19 15:33 | 1553023980.7534 | -549 | 0.20879 | 3.7e-05 | 3e-06 | 101.3857 | 101.6075 | 128.5572 | 21.2612 | 0.001741 | 0.001649 | -152.2538 | 0.23587 | 8.4042 | 0.003125 | 0.004585 | 0.004299 | 4.9e-05 | 0.001255 | 0.30829 | 1.348 | 2.548 | 101.22 | 20.82 | 17.175 | 63.7155 | 97.0109 | 61.501 | 0.004178 | 1.055 | 0.025784 | 24.4403 | -0.019549 | 0.14032 | -2.5e-05 | -0.002108 | -0.002108 | -2.769
2019-03-19 15:33 | 1553023981.7541 | -548 | 0.2088 | 3.9e-05 | 3e-06 | 101.4091 | 101.7647 | 128.5936 | 21.2612 | 0.0017 | 0.001574 | -152.2521 | 0.2378 | 8.4023 | 0.00314 | 0.004047 | 0.003907 | 4.7e-05 | 0.00126 | 0.3063 | 1.327 | 2.5482 | 101.22 | 20.82 | 17.149 | 62.743 | 97.3729 | 59.6926 | 0.004172 | 1.0607 | 0.025857 | 24.3777 | -0.031312 | 0.14147 | -2.7e-05 | -0.002301 | -0.002301 | -2.5574
2019-03-19 15:33 | 1553023982.7531 | -547 | 0.2088 | 3.9e-05 | 2e-06 | 101.4268 | 101.6264 | 128.8393 | 21.2612 | 0.001764 | 0.001662 | -152.233 | 0.23681 | 8.4023 | 0.003134 | 0.004397 | 0.004134 | 4.9e-05 | 0.001262 | 0.32117 | 1.337 | 2.5481 | 101.22 | 20.82 | 17.135 | 64.2736 | 97.2256 | 61.8275 | 0.004168 | 1.0539 | 0.025799 | 24.4791 | -0.017581 | 0.14088 | -2.7e-05 | -0.00224 | -0.00224 | -2.6167
2019-03-19 15:33 | 1553023983.7531 | -546 | 0.20878 | 3.6e-05 | 3e-06 | 101.431 | 101.5572 | 128.6351 | 21.2612 | 0.001743 | 0.001652 | -152.2333 | 0.23638 | 8.4047 | 0.003039 | 0.004603 | 0.004408 | 4.6e-05 | 0.00126 | 0.32836 | 1.3394 | 2.5479 | 101.22 | 20.82 | 17.104 | 63.767 | 94.9487 | 61.5759 | 0.004161 | 1.0547 | 0.025789 | 24.4505 | -0.011119 | 0.14062 | -2.4e-05 | -0.002047 | -0.002047 | -2.8579
2019-03-19 15:33 | 1553023984.7546 | -545 | 0.2088 | 3.5e-05 | 2e-06 | 101.3896 | 101.5178 | 128.748 | 21.2612 | 0.001772 | 0.001599 | -152.2295 | 0.23694 | 8.4017 | 0.003081 | 0.004155 | 0.003873 | 4.9e-05 | 0.00126 | 0.31133 | 1.3246 | 2.548 | 101.22 | 20.82 | 17.089 | 64.4555 | 95.9635 | 60.3072 | 0.004157 | 1.0587 | 0.025849 | 24.4148 | -0.011302 | 0.14096 | -2.2e-05 | -0.001896 | -0.001896 | -3.0924
2019-03-19 15:33 | 1553023985.7533 | -544 | 0.20875 | 3.6e-05 | 3e-06 | 101.4418 | 101.5445 | 128.7015 | 21.2612 | 0.001756 | 0.001648 | -152.2245 | 0.23651 | 8.4006 | 0.003052 | 0.005182 | 0.00494 | 4.5e-05 | 0.001257 | 0.30845 | 1.3212 | 2.5481 | 101.22 | 20.83 | 17.074 | 64.0739 | 95.2587 | 61.468 | 0.004156 | 1.0551 | 0.025799 | 24.4528 | -0.00905 | 0.1407 | -2.3e-05 | -0.00193 | -0.00193 | -3.0325
2019-03-19 15:33 | 1553023986.7536 | -543 | 0.20878 | 3.5e-05 | 3e-06 | 101.3857 | 101.6962 | 128.7511 | 21.2612 | 0.00169 | 0.00159 | -152.2137 | 0.23672 | 8.4013 | 0.003098 | 0.00442 | 0.004215 | 4.4e-05 | 0.001256 | 0.3027 | 1.3294 | 2.5479 | 101.22 | 20.82 | 17.043 | 62.4813 | 96.3526 | 60.0891 | 0.004146 | 1.0594 | 0.025858 | 24.4071 | -0.027354 | 0.14083 | -2.3e-05 | -0.001942 | -0.001942 | -3.0168
2019-03-19 15:33 | 1553023987.7534 | -542 | 0.20878 | 3.8e-05 | 3e-06 | 101.4289 | 101.6154 | 129.0745 | 21.2612 | 0.001744 | 0.001612 | -152.2058 | 0.23711 | 8.4029 | 0.003055 | 0.005176 | 0.004913 | 4.7e-05 | 0.001258 | 0.32731 | 1.3349 | 2.548 | 101.22 | 20.83 | 17.031 | 63.7904 | 95.3191 | 60.6106 | 0.004146 | 1.0578 | 0.02587 | 24.4569 | -0.016431 | 0.14106 | -2.6e-05 | -0.002165 | -0.002165 | -2.7105
2019-03-19 15:33 | 1553023988.7547 | -541 | 0.20876 | 4.1e-05 | 3e-06 | 101.3827 | 101.6353 | 128.7886 | 21.2612 | 0.001741 | 0.001561 | -152.2216 | 0.23589 | 8.4016 | 0.003095 | 0.004121 | 0.004 | 4.5e-05 | 0.001251 | 0.3235 | 1.3324 | 2.5482 | 101.22 | 20.83 | 17.019 | 63.7281 | 96.301 | 59.387 | 0.004143 | 1.0617 | 0.025889 | 24.3849 | -0.022255 | 0.14033 | -3e-05 | -0.002506 | -0.002506 | -2.3297
2019-03-19 15:33 | 1553023989.7535 | -540 | 0.20876 | 3.5e-05 | 3e-06 | 101.3848 | 101.6736 | 128.7058 | 21.2612 | 0.001685 | 0.001616 | -152.2223 | 0.23716 | 8.4033 | 0.003088 | 0.004021 | 0.003708 | 4.8e-05 | 0.001247 | 0.30702 | 1.2741 | 2.5479 | 101.22 | 20.83 | 17.019 | 62.3633 | 96.1303 | 60.7067 | 0.004143 | 1.0575 | 0.025829 | 24.4254 | -0.025445 | 0.14108 | -2.2e-05 | -0.001911 | -0.001911 | -3.0713
2019-03-19 15:33 | 1553023990.7545 | -539 | 0.2088 | 4e-05 | 3e-06 | 101.3374 | 101.6319 | 128.7924 | 21.2612 | 0.001692 | 0.001624 | -152.225 | 0.23708 | 8.4047 | 0.003035 | 0.004646 | 0.004457 | 4.8e-05 | 0.001251 | 0.31598 | 1.2955 | 2.5481 | 101.22 | 20.83 | 16.998 | 62.5485 | 94.8478 | 60.9052 | 0.004138 | 1.0568 | 0.02583 | 24.4409 | -0.025953 | 0.14104 | -2.8e-05 | -0.002375 | -0.002375 | -2.4704
2019-03-19 15:33 | 1553023991.754 | -538 | 0.20877 | 4.2e-05 | 3e-06 | 101.4442 | 101.6199 | 128.8928 | 21.2612 | 0.001754 | 0.001597 | -152.2239 | 0.2367 | 8.4054 | 0.003083 | 0.00398 | 0.003795 | 5e-05 | 0.001251 | 0.31813 | 1.3162 | 2.548 | 101.22 | 20.82 | 16.995 | 64.0301 | 95.9987 | 60.2455 | 0.004134 | 1.0589 | 0.025866 | 24.4263 | -0.015473 | 0.14081 | -3.1e-05 | -0.002564 | -0.002564 | -2.2847
2019-03-19 15:33 | 1553023992.7526 | -537 | 0.20878 | 3.9e-05 | 3e-06 | 101.4541 | 101.6641 | 128.9229 | 21.2612 | 0.001704 | 0.001683 | -152.2208 | 0.2368 | 8.4059 | 0.003107 | 0.004332 | 0.004158 | 4.6e-05 | 0.001253 | 0.30879 | 1.3152 | 2.548 | 101.22 | 20.82 | 16.994 | 62.8203 | 96.5885 | 62.319 | 0.004134 | 1.0524 | 0.025789 | 24.505 | -0.018488 | 0.14088 | -2.8e-05 | -0.002337 | -0.002337 | -2.5082
2019-03-19 15:33 | 1553023993.7531 | -536 | 0.20876 | 3.3e-05 | 3e-06 | 101.4133 | 101.6688 | 129.0799 | 21.2612 | 0.00174 | 0.001611 | -152.2261 | 0.23735 | 8.4019 | 0.003079 | 0.00435 | 0.004051 | 4.7e-05 | 0.001257 | 0.2943 | 1.3133 | 2.5481 | 101.22 | 20.82 | 16.999 | 63.6919 | 95.914 | 60.5935 | 0.004135 | 1.0578 | 0.025871 | 24.4567 | -0.022508 | 0.1412 | -2e-05 | -0.001774 | -0.001774 | -3.312
2019-03-19 15:33 | 1553023994.7527 | -535 | 0.20877 | 4e-05 | 3e-06 | 101.4073 | 101.536 | 128.8637 | 21.2612 | 0.00171 | 0.001683 | -152.2368 | 0.23587 | 8.402 | 0.003056 | 0.00426 | 0.004035 | 4.5e-05 | 0.001258 | 0.31238 | 1.347 | 2.5482 | 101.22 | 20.82 | 17.005 | 62.9821 | 95.3379 | 62.333 | 0.004137 | 1.0523 | 0.025782 | 24.4998 | -0.011339 | 0.14032 | -2.9e-05 | -0.002398 | -0.002398 | -2.4347
2019-03-19 15:33 | 1553023995.7541 | -534 | 0.20877 | 4e-05 | 3e-06 | 101.4028 | 101.5706 | 128.7675 | 21.2612 | 0.001746 | 0.001605 | -152.2336 | 0.23652 | 8.4029 | 0.00311 | 0.004459 | 0.004237 | 4.3e-05 | 0.001268 | 0.30342 | 1.356 | 2.548 | 101.22 | 20.82 | 17.02 | 63.8318 | 96.6441 | 60.4417 | 0.00414 | 1.0583 | 0.025846 | 24.4216 | -0.014783 | 0.1407 | -2.9e-05 | -0.002456 | -0.002456 | -2.3833
2019-03-19 15:33 | 1553023996.7528 | -533 | 0.20876 | 3.8e-05 | 2e-06 | 101.4667 | 101.4771 | 128.9466 | 21.2612 | 0.001766 | 0.001625 | -152.2412 | 0.23661 | 8.3998 | 0.003109 | 0.004319 | 0.004125 | 4.7e-05 | 0.001267 | 0.31437 | 1.3774 | 2.5482 | 101.22 | 20.82 | 17.029 | 64.3205 | 96.6325 | 60.9358 | 0.004143 | 1.0568 | 0.025845 | 24.4566 | -0.000909 | 0.14076 | -2.6e-05 | -0.002158 | -0.002158 | -2.7131
2019-03-19 15:33 | 1553023997.753 | -532 | 0.20876 | 3.7e-05 | 3e-06 | 101.422 | 101.549 | 129.0135 | 21.2612 | 0.00172 | 0.001593 | -152.2436 | 0.23638 | 8.4037 | 0.003071 | 0.004089 | 0.003862 | 4.4e-05 | 0.001266 | 0.28998 | 1.3513 | 2.5481 | 101.22 | 20.82 | 17.039 | 63.225 | 95.7035 | 60.1525 | 0.004145 | 1.0592 | 0.025882 | 24.4343 | -0.011189 | 0.14062 | -2.5e-05 | -0.002139 | -0.002139 | -2.7346
2019-03-19 15:33 | 1553023998.753 | -531 | 0.20874 | 3.6e-05 | 3e-06 | 101.4913 | 101.5969 | 129.1459 | 21.2612 | 0.001718 | 0.00164 | -152.2358 | 0.23634 | 8.3987 | 0.003089 | 0.004706 | 0.004465 | 4.6e-05 | 0.001272 | 0.3282 | 1.3405 | 2.5481 | 101.22 | 20.82 | 17.063 | 63.1578 | 96.139 | 61.2921 | 0.004151 | 1.0556 | 0.025851 | 24.4886 | -0.0093 | 0.1406 | -2.3e-05 | -0.002005 | -0.002005 | -2.9174
2019-03-19 15:33 | 1553023999.7534 | -530 | 0.20875 | 4e-05 | 3e-06 | 101.4229 | 101.5798 | 129.3243 | 21.2612 | 0.001718 | 0.001654 | -152.239 | 0.23612 | 8.4018 | 0.003058 | 0.004257 | 0.00397 | 4.4e-05 | 0.001262 | 0.33157 | 1.3723 | 2.548 | 101.22 | 20.82 | 17.074 | 63.1586 | 95.3879 | 61.6305 | 0.004154 | 1.0546 | 0.025856 | 24.5179 | -0.013824 | 0.14047 | -2.9e-05 | -0.002468 | -0.002468 | -2.3681
2019-03-19 15:33 | 1553024000.754 | -529 | 0.20875 | 4e-05 | 3e-06 | 101.3878 | 101.6353 | 129.3226 | 21.2612 | 0.001672 | 0.001602 | -152.2544 | 0.23594 | 8.4036 | 0.003161 | 0.004413 | 0.004268 | 4.6e-05 | 0.001258 | 0.33069 | 1.3401 | 2.5481 | 101.22 | 20.82 | 17.095 | 62.0681 | 97.8901 | 60.3814 | 0.004159 | 1.0585 | 0.025904 | 24.4719 | -0.021805 | 0.14036 | -2.9e-05 | -0.002474 | -0.002474 | -2.36
2019-03-19 15:33 | 1553024001.754 | -528 | 0.20873 | 3.7e-05 | 3e-06 | 101.4679 | 101.7353 | 129.3113 | 21.2612 | 0.001755 | 0.001646 | -152.2648 | 0.23683 | 8.4014 | 0.003127 | 0.00474 | 0.004489 | 4.5e-05 | 0.001265 | 0.30287 | 1.3089 | 2.548 | 101.22 | 20.82 | 17.104 | 64.0534 | 97.0594 | 61.4234 | 0.004161 | 1.0552 | 0.025862 | 24.5091 | -0.023536 | 0.14089 | -2.5e-05 | -0.002126 | -0.002126 | -2.7567
2019-03-19 15:33 | 1553024002.7531 | -527 | 0.20874 | 3.7e-05 | 3e-06 | 101.4103 | 101.626 | 129.1031 | 21.2612 | 0.001697 | 0.001642 | -152.272 | 0.23643 | 8.4032 | 0.003103 | 0.004702 | 0.004525 | 4.6e-05 | 0.00127 | 0.30807 | 1.3359 | 2.5481 | 101.22 | 20.82 | 17.113 | 62.6626 | 96.4765 | 61.3454 | 0.004163 | 1.0555 | 0.025844 | 24.4865 | -0.019006 | 0.14065 | -2.5e-05 | -0.002158 | -0.002158 | -2.7108
2019-03-19 15:33 | 1553024003.7545 | -526 | 0.20875 | 3.5e-05 | 3e-06 | 101.4496 | 101.6435 | 129.1447 | 21.2612 | 0.001792 | 0.001642 | -152.2766 | 0.23643 | 8.4037 | 0.003147 | 0.005281 | 0.005024 | 4.8e-05 | 0.001275 | 0.32549 | 1.3099 | 2.5479 | 101.21 | 20.83 | 17.128 | 64.9497 | 97.5586 | 61.3461 | 0.00417 | 1.0555 | 0.025849 | 24.4904 | -0.017077 | 0.14065 | -2.3e-05 | -0.001969 | -0.001969 | -2.9713
2019-03-19 15:33 | 1553024004.7543 | -525 | 0.20877 | 3.9e-05 | 3e-06 | 101.4118 | 101.7452 | 129.2689 | 21.2612 | 0.001732 | 0.001616 | -152.2584 | 0.23601 | 8.4034 | 0.003162 | 0.004589 | 0.004343 | 4.5e-05 | 0.001283 | 0.32947 | 1.3002 | 2.5479 | 101.21 | 20.83 | 17.135 | 63.4953 | 97.9051 | 60.72 | 0.004171 | 1.0574 | 0.025885 | 24.4793 | -0.029358 | 0.1404 | -2.8e-05 | -0.002347 | -0.002347 | -2.4884
2019-03-19 15:33 | 1553024005.7533 | -524 | 0.20873 | 4.1e-05 | 3e-06 | 101.4421 | 101.7096 | 129.5088 | 21.2612 | 0.001751 | 0.001658 | -152.2593 | 0.2355 | 8.4029 | 0.003135 | 0.004488 | 0.00426 | 4.5e-05 | 0.00128 | 0.31708 | 1.3442 | 2.5481 | 101.21 | 20.83 | 17.141 | 63.9504 | 97.2475 | 61.7322 | 0.004173 | 1.0542 | 0.02587 | 24.5391 | -0.023553 | 0.1401 | -3e-05 | -0.002545 | -0.002545 | -2.2902
2019-03-19 15:33 | 1553024006.7524 | -523 | 0.20875 | 3.8e-05 | 2e-06 | 101.4478 | 101.7219 | 129.6827 | 21.2612 | 0.001767 | 0.001617 | -152.2576 | 0.23557 | 8.4032 | 0.00308 | 0.004214 | 0.003972 | 4.9e-05 | 0.001273 | 0.3131 | 1.3197 | 2.548 | 101.21 | 20.82 | 17.151 | 64.3463 | 95.9203 | 60.7318 | 0.004173 | 1.0574 | 0.025926 | 24.5189 | -0.024134 | 0.14014 | -2.5e-05 | -0.002133 | -0.002133 | -2.7329
2019-03-19 15:33 | 1553024007.7533 | -522 | 0.20875 | 4e-05 | 3e-06 | 101.4379 | 101.5366 | 129.5228 | 21.2612 | 0.001787 | 0.001619 | -152.2763 | 0.23571 | 8.4033 | 0.003148 | 0.004693 | 0.004544 | 4.7e-05 | 0.00127 | 0.34435 | 1.3082 | 2.5481 | 101.21 | 20.82 | 17.155 | 64.8357 | 97.5828 | 60.7787 | 0.004174 | 1.0572 | 0.025908 | 24.5055 | -0.0087 | 0.14022 | -2.8e-05 | -0.002348 | -0.002348 | -2.4847
2019-03-19 15:33 | 1553024008.7534 | -521 | 0.20876 | 3.4e-05 | 3e-06 | 101.4256 | 101.8473 | 129.5303 | 21.2612 | 0.001766 | 0.001618 | -152.2742 | 0.2361 | 8.4052 | 0.003109 | 0.004622 | 0.004369 | 4.6e-05 | 0.001262 | 0.29839 | 1.3298 | 2.5481 | 101.21 | 20.82 | 17.163 | 64.3261 | 96.6388 | 60.7607 | 0.004176 | 1.0573 | 0.02591 | 24.5055 | -0.037108 | 0.14046 | -2.1e-05 | -0.001817 | -0.001817 | -3.2156
2019-03-19 15:33 | 1553024009.7535 | -520 | 0.20876 | 3.8e-05 | 2e-06 | 101.4908 | 101.7528 | 129.3804 | 21.2612 | 0.001724 | 0.00162 | -152.2787 | 0.2353 | 8.4032 | 0.003114 | 0.004222 | 0.003955 | 4.6e-05 | 0.001262 | 0.30127 | 1.3119 | 2.548 | 101.21 | 20.82 | 17.167 | 63.3133 | 96.7491 | 60.7951 | 0.004177 | 1.0572 | 0.025893 | 24.4926 | -0.023061 | 0.13998 | -2.5e-05 | -0.002124 | -0.002124 | -2.7415
2019-03-19 15:33 | 1553024010.7526 | -519 | 0.20876 | 3.7e-05 | 3e-06 | 101.4346 | 101.7408 | 129.2428 | 21.2612 | 0.001735 | 0.001634 | -152.2718 | 0.23628 | 8.4035 | 0.00304 | 0.004404 | 0.004197 | 4.8e-05 | 0.001265 | 0.31509 | 1.3257 | 2.5482 | 101.21 | 20.82 | 17.171 | 63.5834 | 94.9588 | 61.1367 | 0.004178 | 1.0561 | 0.025866 | 24.4921 | -0.026955 | 0.14057 | -2.5e-05 | -0.002107 | -0.002107 | -2.7747
2019-03-19 15:34 | 1553024011.7526 | -518 | 0.20878 | 4.1e-05 | 3e-06 | 101.4965 | 101.6849 | 129.0527 | 21.2612 | 0.001793 | 0.0016 | -152.2735 | 0.23569 | 8.4015 | 0.003116 | 0.004425 | 0.00416 | 4.5e-05 | 0.001258 | 0.31188 | 1.3159 | 2.5481 | 101.21 | 20.83 | 17.179 | 64.9678 | 96.786 | 60.3142 | 0.004182 | 1.0587 | 0.025879 | 24.4439 | -0.016596 | 0.14021 | -3e-05 | -0.00249 | -0.00249 | -2.3429
2019-03-19 15:34 | 1553024012.7541 | -517 | 0.20878 | 3.8e-05 | 2e-06 | 101.4727 | 101.7202 | 128.8968 | 21.2612 | 0.001717 | 0.001635 | -152.279 | 0.23704 | 8.403 | 0.003073 | 0.004153 | 0.003967 | 4.9e-05 | 0.001247 | 0.30807 | 1.3361 | 2.548 | 101.21 | 20.82 | 17.184 | 63.1439 | 95.7535 | 61.1772 | 0.004181 | 1.056 | 0.02583 | 24.4608 | -0.021788 | 0.14101 | -2.5e-05 | -0.002053 | -0.002053 | -2.8569
2019-03-19 15:34 | 1553024013.7528 | -516 | 0.20876 | 3.8e-05 | 3e-06 | 101.5166 | 101.614 | 128.8532 | 21.2612 | 0.001666 | 0.001641 | -152.2772 | 0.23625 | 8.4031 | 0.00306 | 0.004539 | 0.0043 | 4.9e-05 | 0.00124 | 0.30431 | 1.347 | 2.5481 | 101.21 | 20.83 | 17.195 | 61.9142 | 95.4506 | 61.3164 | 0.004186 | 1.0555 | 0.02582 | 24.4617 | -0.008586 | 0.14054 | -2.6e-05 | -0.002199 | -0.002199 | -2.6591
2019-03-19 15:34 | 1553024014.7528 | -515 | 0.20879 | 3.6e-05 | 3e-06 | 101.4112 | 101.713 | 128.9449 | 21.2612 | 0.001762 | 0.001633 | -152.2867 | 0.23572 | 8.4031 | 0.003116 | 0.003744 | 0.003536 | 5e-05 | 0.001235 | 0.32748 | 1.3392 | 2.5481 | 101.21 | 20.83 | 17.202 | 64.2277 | 96.8087 | 61.1204 | 0.004188 | 1.0562 | 0.025837 | 24.4632 | -0.02658 | 0.14023 | -2.3e-05 | -0.001981 | -0.001981 | -2.9442
2019-03-19 15:34 | 1553024015.7537 | -514 | 0.20878 | 3.9e-05 | 2e-06 | 101.4616 | 101.7052 | 128.9633 | 21.2612 | 0.001755 | 0.00162 | -152.2833 | 0.23639 | 8.4032 | 0.003165 | 0.004428 | 0.004182 | 5.2e-05 | 0.001232 | 0.32211 | 1.3466 | 2.548 | 101.21 | 20.82 | 17.208 | 64.0659 | 97.9935 | 60.8036 | 0.004187 | 1.0572 | 0.025851 | 24.4534 | -0.021442 | 0.14063 | -2.5e-05 | -0.002079 | -0.002079 | -2.8145
2019-03-19 15:34 | 1553024016.7532 | -513 | 0.20877 | 4.2e-05 | 3e-06 | 101.4625 | 101.5476 | 128.8792 | 21.2612 | 0.001774 | 0.001622 | -152.2806 | 0.23682 | 8.4039 | 0.003118 | 0.005261 | 0.005029 | 5.1e-05 | 0.00123 | 0.30502 | 1.3207 | 2.5479 | 101.21 | 20.83 | 17.223 | 64.522 | 96.8486 | 60.8568 | 0.004193 | 1.057 | 0.025841 | 24.4474 | -0.007496 | 0.14088 | -3.1e-05 | -0.00256 | -0.00256 | -2.2891
2019-03-19 15:34 | 1553024017.7529 | -512 | 0.20876 | 3.9e-05 | 3e-06 | 101.4547 | 101.6456 | 129.063 | 21.2612 | 0.001728 | 0.001633 | -152.2546 | 0.23607 | 8.4017 | 0.003051 | 0.004816 | 0.004595 | 5.2e-05 | 0.00123 | 0.31957 | 1.3113 | 2.5479 | 101.21 | 20.83 | 17.231 | 63.4101 | 95.2406 | 61.1264 | 0.004195 | 1.0561 | 0.025849 | 24.4747 | -0.016808 | 0.14044 | -2.8e-05 | -0.00233 | -0.00233 | -2.5072
2019-03-19 15:34 | 1553024018.7544 | -511 | 0.20875 | 3.6e-05 | 2e-06 | 101.4571 | 101.6531 | 129.3782 | 21.2612 | 0.001771 | 0.001611 | -152.2537 | 0.23699 | 8.4018 | 0.003097 | 0.005246 | 0.005011 | 5e-05 | 0.001232 | 0.31022 | 1.3077 | 2.548 | 101.21 | 20.83 | 17.24 | 64.4471 | 96.3495 | 60.5802 | 0.004197 | 1.0579 | 0.025902 | 24.4845 | -0.017259 | 0.14098 | -2.2e-05 | -0.001871 | -0.001871 | -3.1343
2019-03-19 15:34 | 1553024019.754 | -510 | 0.20879 | 3.8e-05 | 3e-06 | 101.4811 | 101.7315 | 129.6724 | 21.2612 | 0.001752 | 0.001695 | -152.2389 | 0.2365 | 8.405 | 0.003132 | 0.003494 | 0.00324 | 5e-05 | 0.001233 | 0.3011 | 1.3244 | 2.5481 | 101.21 | 20.83 | 17.256 | 63.9957 | 97.1729 | 62.6135 | 0.004201 | 1.0515 | 0.025852 | 24.5869 | -0.022041 | 0.1407 | -2.5e-05 | -0.002161 | -0.002161 | -2.7084
2019-03-19 15:34 | 1553024020.7541 | -509 | 0.20879 | 3.5e-05 | 2e-06 | 101.5211 | 101.7623 | 130.0256 | 21.2612 | 0.001752 | 0.001676 | -152.2422 | 0.23638 | 8.4027 | 0.003051 | 0.004435 | 0.004242 | 5.4e-05 | 0.00123 | 0.335 | 1.3573 | 2.5481 | 101.21 | 20.83 | 17.264 | 63.9918 | 95.2197 | 62.1649 | 0.004203 | 1.0529 | 0.025905 | 24.6039 | -0.021234 | 0.14062 | -2.2e-05 | -0.001849 | -0.001849 | -3.1631
2019-03-19 15:34 | 1553024021.7532 | -508 | 0.20873 | 4.1e-05 | 2e-06 | 101.5088 | 101.6517 | 130.3667 | 21.2612 | 0.001808 | 0.001635 | -152.2395 | 0.23652 | 8.4021 | 0.003091 | 0.00449 | 0.004258 | 5e-05 | 0.001232 | 0.41734 | 1.3299 | 2.548 | 101.21 | 20.83 | 17.279 | 65.3299 | 96.1938 | 61.175 | 0.004206 | 1.056 | 0.025977 | 24.5997 | -0.012589 | 0.1407 | -2.9e-05 | -0.002445 | -0.002445 | -2.3939
2019-03-19 15:34 | 1553024022.7532 | -507 | 0.20875 | 3.9e-05 | 2e-06 | 101.4808 | 101.7202 | 130.6549 | 21.2612 | 0.001729 | 0.001663 | -152.2305 | 0.23559 | 8.4036 | 0.003101 | 0.00467 | 0.004405 | 4.9e-05 | 0.001233 | 0.41375 | 1.3124 | 2.5481 | 101.21 | 20.83 | 17.285 | 63.4421 | 96.4272 | 61.8352 | 0.004208 | 1.0539 | 0.02598 | 24.6512 | -0.021074 | 0.14015 | -2.6e-05 | -0.00219 | -0.00219 | -2.6627
2019-03-19 15:34 | 1553024023.7523 | -506 | 0.20876 | 3.8e-05 | 2e-06 | 101.5433 | 101.6456 | 130.7695 | 21.2612 | 0.001888 | 0.001653 | -152.2275 | 0.23563 | 8.4029 | 0.003048 | 0.004757 | 0.004488 | 4.7e-05 | 0.001227 | 0.40662 | 1.3658 | 2.548 | 101.21 | 20.83 | 17.291 | 67.2566 | 95.1656 | 61.6024 | 0.004209 | 1.0546 | 0.026001 | 24.6535 | -0.009006 | 0.14017 | -2.5e-05 | -0.002113 | -0.002113 | -2.7601
2019-03-19 15:34 | 1553024024.7533 | -505 | 0.20876 | 3.9e-05 | 2e-06 | 101.5793 | 101.8305 | 130.5033 | 21.2612 | 0.00175 | 0.001619 | -152.2169 | 0.2365 | 8.4028 | 0.003102 | 0.004526 | 0.004322 | 4.9e-05 | 0.001224 | 0.5916 | 1.3079 | 2.548 | 101.21 | 20.83 | 17.301 | 63.9317 | 96.4656 | 60.792 | 0.004212 | 1.0572 | 0.026006 | 24.5985 | -0.022094 | 0.14069 | -2.7e-05 | -0.002245 | -0.002245 | -2.6071
2019-03-19 15:34 | 1553024025.7524 | -504 | 0.20878 | 3.3e-05 | 3e-06 | 101.5628 | 101.8343 | 130.4899 | 21.2612 | 0.001676 | 0.001673 | -152.2402 | 0.2352 | 8.4037 | 0.003124 | 0.004101 | 0.003852 | 4.9e-05 | 0.00123 | 1.3859 | 1.3334 | 2.5481 | 101.21 | 20.83 | 17.306 | 62.1433 | 96.9865 | 62.0775 | 0.004213 | 1.0532 | 0.025954 | 24.6446 | -0.023879 | 0.13992 | -2e-05 | -0.001709 | -0.001709 | -3.4069
2019-03-19 15:34 | 1553024026.7527 | -503 | 0.20876 | 3.5e-05 | 2e-06 | 101.5349 | 101.7493 | 130.6035 | 21.2612 | 0.001735 | 0.001584 | -152.2384 | 0.23651 | 8.4033 | 0.00307 | 0.004703 | 0.004456 | 4.7e-05 | 0.001227 | 2.8809 | 1.3243 | 2.5481 | 101.21 | 20.83 | 17.314 | 63.5803 | 95.6784 | 59.9272 | 0.004215 | 1.06 | 0.026049 | 24.5761 | -0.018873 | 0.1407 | -2.2e-05 | -0.001871 | -0.001871 | -3.1283
2019-03-19 15:34 | 1553024027.7525 | -502 | 0.20877 | 3.6e-05 | 2e-06 | 101.5412 | 101.7216 | 130.7204 | 21.2612 | 0.001757 | 0.001604 | -152.2452 | 0.23576 | 8.4046 | 0.003134 | 0.004377 | 0.00412 | 4.6e-05 | 0.001225 | 19.3299 | 1.3372 | 2.5481 | 101.21 | 20.83 | 17.318 | 64.098 | 97.2302 | 60.4106 | 0.004216 | 1.0584 | 0.026042 | 24.6049 | -0.015879 | 0.14025 | -2.3e-05 | -0.00193 | -0.00193 | -3.0233
2019-03-19 15:34 | 1553024028.7536 | -501 | 0.20877 | 3.5e-05 | 2e-06 | 101.5967 | 101.7815 | 130.5237 | 21.2612 | 0.001726 | 0.001668 | -152.2437 | 0.2369 | 8.4041 | 0.003116 | 0.004078 | 0.003968 | 4.8e-05 | 0.001226 | 48.312 | 1.3668 | 2.548 | 101.21 | 20.83 | 17.321 | 63.3648 | 96.7906 | 61.9525 | 0.004217 | 1.0535 | 0.025963 | 24.6432 | -0.016259 | 0.14093 | -2.1e-05 | -0.0018 | -0.0018 | -3.2568
2019-03-19 15:34 | 1553024029.7535 | -500 | 0.20876 | 3.5e-05 | 2e-06 | 101.5565 | 101.85 | 130.0728 | 21.2612 | 0.001693 | 0.001645 | -152.2412 | 0.23641 | 8.4028 | 0.003076 | 0.004474 | 0.004146 | 4.5e-05 | 0.00123 | 62.2875 | 1.3526 | 2.5482 | 101.21 | 20.83 | 17.326 | 62.5752 | 95.8233 | 61.3993 | 0.004218 | 1.0553 | 0.025939 | 24.5802 | -0.025817 | 0.14064 | -2.2e-05 | -0.001844 | -0.001844 | -3.1731
2019-03-19 15:34 | 1553024030.754 | -499 | 0.20876 | 3.5e-05 | 3e-06 | 101.5793 | 101.8713 | 129.6501 | 21.2612 | 0.001781 | 0.001667 | -152.2365 | 0.23666 | 8.4021 | 0.003137 | 0.004162 | 0.003905 | 4.7e-05 | 0.001224 | 70.5313 | 1.3202 | 2.548 | 101.21 | 20.83 | 17.33 | 64.6899 | 97.3031 | 61.9446 | 0.004219 | 1.0536 | 0.025876 | 24.5603 | -0.025673 | 0.14079 | -2.1e-05 | -0.001831 | -0.001831 | -3.1985
2019-03-19 15:34 | 1553024031.7527 | -498 | 0.20876 | 3.6e-05 | 2e-06 | 101.6159 | 102.0004 | 129.2818 | 21.2612 | 0.001749 | 0.001661 | -152.2241 | 0.23508 | 8.4041 | 0.003092 | 0.003959 | 0.003711 | 4.6e-05 | 0.001225 | 73.0334 | 1.3211 | 2.548 | 101.21 | 20.83 | 17.341 | 63.9013 | 96.2165 | 61.7907 | 0.004222 | 1.0541 | 0.025845 | 24.5197 | -0.03378 | 0.13985 | -2.2e-05 | -0.001871 | -0.001871 | -3.1099
2019-03-19 15:34 | 1553024032.7538 | -497 | 0.20877 | 3.6e-05 | 2e-06 | 101.6168 | 101.7603 | 128.9093 | 21.2612 | 0.00176 | 0.001643 | -152.2462 | 0.2354 | 8.4023 | 0.00311 | 0.004655 | 0.004458 | 4.7e-05 | 0.001223 | 74.1462 | 1.3222 | 2.5479 | 101.21 | 20.83 | 17.346 | 64.1738 | 96.6552 | 61.3704 | 0.004223 | 1.0554 | 0.025824 | 24.469 | -0.012622 | 0.14004 | -2.3e-05 | -0.001929 | -0.001929 | -3.0205
2019-03-19 15:34 | 1553024033.7542 | -496 | 0.20878 | 3.7e-05 | 2e-06 | 101.6651 | 101.9689 | 128.5443 | 21.2612 | 0.001801 | 0.001686 | -152.2259 | 0.23617 | 8.402 | 0.003152 | 0.004511 | 0.004312 | 4.2e-05 | 0.001223 | 74.2329 | 1.3256 | 2.548 | 101.21 | 20.83 | 17.352 | 65.1684 | 97.662 | 62.4002 | 0.004224 | 1.0521 | 0.025748 | 24.4719 | -0.026685 | 0.1405 | -2.4e-05 | -0.001964 | -0.001964 | -2.976
2019-03-19 15:34 | 1553024034.7529 | -495 | 0.20875 | 3.4e-05 | 2e-06 | 101.5844 | 101.7832 | 128.6818 | 21.2612 | 0.001786 | 0.001688 | -152.2395 | 0.23673 | 8.4035 | 0.003199 | 0.004824 | 0.004613 | 4.8e-05 | 0.001223 | 74.2919 | 1.3385 | 2.5481 | 101.21 | 20.83 | 17.366 | 64.8006 | 98.809 | 62.4353 | 0.004228 | 1.052 | 0.02576 | 24.4863 | -0.017492 | 0.14083 | -2e-05 | -0.001662 | -0.001662 | -3.5246
2019-03-19 15:34 | 1553024035.7542 | -494 | 0.20875 | 3.5e-05 | 2e-06 | 101.6603 | 101.8175 | 128.3795 | 21.2612 | 0.00177 | 0.001656 | -152.2328 | 0.23679 | 8.4007 | 0.003243 | 0.004505 | 0.004235 | 5e-05 | 0.00122 | 74.2476 | 1.3525 | 2.5482 | 101.21 | 20.83 | 17.373 | 64.4173 | 99.8756 | 61.6799 | 0.004229 | 1.0544 | 0.025759 | 24.43 | -0.013819 | 0.14087 | -2.1e-05 | -0.001759 | -0.001759 | -3.3318
2019-03-19 15:34 | 1553024036.7543 | -493 | 0.20877 | 3.1e-05 | 3e-06 | 101.6522 | 101.8665 | 127.9701 | 21.2612 | 0.00172 | 0.001727 | -152.2299 | 0.23575 | 8.4021 | 0.003238 | 0.004837 | 0.004634 | 5e-05 | 0.001217 | 74.4145 | 1.3754 | 2.548 | 101.21 | 20.83 | 17.381 | 63.2041 | 99.7477 | 63.3757 | 0.004231 | 1.0491 | 0.025653 | 24.4527 | -0.018833 | 0.14025 | -1.7e-05 | -0.001477 | -0.001477 | -3.9499
2019-03-19 15:34 | 1553024037.753 | -492 | 0.20877 | 3.5e-05 | 3e-06 | 101.6234 | 101.8439 | 127.7608 | 21.2612 | 0.001743 | 0.001658 | -152.2272 | 0.23643 | 8.4049 | 0.003215 | 0.004615 | 0.004386 | 4.9e-05 | 0.001214 | 74.4909 | 1.3929 | 2.548 | 101.21 | 20.84 | 17.395 | 63.7622 | 99.1822 | 61.729 | 0.004237 | 1.0542 | 0.025695 | 24.3728 | -0.019384 | 0.14065 | -2.2e-05 | -0.001889 | -0.001889 | -3.0968
2019-03-19 15:34 | 1553024038.7521 | -491 | 0.20876 | 3.4e-05 | 2e-06 | 101.6576 | 102.0466 | 127.6019 | 21.2612 | 0.001814 | 0.00167 | -152.239 | 0.23563 | 8.4012 | 0.003187 | 0.004581 | 0.004416 | 4.8e-05 | 0.001216 | 74.4175 | 1.3761 | 2.548 | 101.21 | 20.84 | 17.404 | 65.4648 | 98.5149 | 62.0048 | 0.00424 | 1.0534 | 0.025668 | 24.3677 | -0.034162 | 0.14017 | -2e-05 | -0.001694 | -0.001694 | -3.4431
2019-03-19 15:34 | 1553024039.7521 | -490 | 0.20877 | 3.5e-05 | 3e-06 | 101.6678 | 101.8041 | 127.7192 | 21.2612 | 0.001757 | 0.001686 | -152.2563 | 0.23595 | 8.4015 | 0.003268 | 0.004723 | 0.00453 | 4.6e-05 | 0.001215 | 74.014 | 1.3591 | 2.5482 | 101.21 | 20.84 | 17.417 | 64.098 | 100.4744 | 62.3838 | 0.004243 | 1.0522 | 0.025666 | 24.3927 | -0.011985 | 0.14036 | -2.3e-05 | -0.001934 | -0.001934 | -3.0199
2019-03-19 15:34 | 1553024040.7521 | -489 | 0.20875 | 3.7e-05 | 3e-06 | 101.5448 | 101.7373 | 127.7556 | 21.2612 | 0.001699 | 0.001649 | -152.254 | 0.23647 | 8.4014 | 0.003367 | 0.004563 | 0.004311 | 4.8e-05 | 0.001221 | 74.4416 | 1.3761 | 2.548 | 101.21 | 20.84 | 17.422 | 62.7141 | 102.8566 | 61.5142 | 0.004244 | 1.0549 | 0.025703 | 24.3645 | -0.016947 | 0.14067 | -2.5e-05 | -0.002084 | -0.002084 | -2.8077
2019-03-19 15:34 | 1553024041.7522 | -488 | 0.20875 | 3.9e-05 | 3e-06 | 101.6444 | 102.0559 | 127.8842 | 21.2612 | 0.001715 | 0.001668 | -152.2549 | 0.23593 | 8.4011 | 0.003365 | 0.004635 | 0.004463 | 4.8e-05 | 0.001219 | 74.4656 | 1.3406 | 2.5481 | 101.21 | 20.84 | 17.427 | 63.0907 | 102.8283 | 61.9658 | 0.004245 | 1.0535 | 0.025698 | 24.3932 | -0.036135 | 0.14035 | -2.7e-05 | -0.002306 | -0.002306 | -2.5321
2019-03-19 15:34 | 1553024042.753 | -487 | 0.20877 | 3.6e-05 | 3e-06 | 101.5865 | 101.712 | 128.0387 | 21.2612 | 0.001763 | 0.001622 | -152.2253 | 0.23654 | 8.4022 | 0.003302 | 0.00522 | 0.004976 | 4.8e-05 | 0.001224 | 74.5119 | 1.3566 | 2.5481 | 101.21 | 20.84 | 17.432 | 64.252 | 101.3055 | 60.8505 | 0.004246 | 1.057 | 0.025757 | 24.3673 | -0.011045 | 0.14072 | -2.3e-05 | -0.001989 | -0.001989 | -2.943
2019-03-19 15:34 | 1553024043.7524 | -486 | 0.20876 | 4.1e-05 | 2e-06 | 101.5691 | 101.7182 | 128.2159 | 21.2612 | 0.001779 | 0.001668 | -152.2517 | 0.2365 | 8.4021 | 0.003331 | 0.004906 | 0.004677 | 4.6e-05 | 0.001222 | 74.4161 | 1.3009 | 2.548 | 101.21 | 20.84 | 17.434 | 64.6343 | 102.0061 | 61.9735 | 0.004247 | 1.0535 | 0.025731 | 24.4251 | -0.01312 | 0.14069 | -2.8e-05 | -0.002333 | -0.002333 | -2.5091
2019-03-19 15:34 | 1553024044.7537 | -485 | 0.20874 | 3.9e-05 | 2e-06 | 101.5517 | 101.6541 | 128.2089 | 21.2612 | 0.001748 | 0.001671 | -152.2556 | 0.23641 | 8.4014 | 0.003327 | 0.005098 | 0.004819 | 5.1e-05 | 0.001222 | 74.3872 | 1.3092 | 2.5481 | 101.21 | 20.84 | 17.439 | 63.885 | 101.896 | 62.0291 | 0.004248 | 1.0533 | 0.025728 | 24.4264 | -0.009019 | 0.14064 | -2.7e-05 | -0.002248 | -0.002248 | -2.6026
2019-03-19 15:34 | 1553024045.7535 | -484 | 0.20876 | 3.6e-05 | 3e-06 | 101.5508 | 101.885 | 128.4225 | 21.2612 | 0.001796 | 0.001682 | -152.253 | 0.23526 | 8.4037 | 0.003262 | 0.004247 | 0.004024 | 5e-05 | 0.001216 | 74.4315 | 1.331 | 2.5481 | 101.21 | 20.84 | 17.442 | 65.0489 | 100.3269 | 62.3041 | 0.004249 | 1.0524 | 0.025739 | 24.4568 | -0.029387 | 0.13996 | -2.3e-05 | -0.001972 | -0.001972 | -2.952
2019-03-19 15:34 | 1553024046.7534 | -483 | 0.20877 | 3.7e-05 | 3e-06 | 101.5919 | 101.7024 | 128.8539 | 21.2612 | 0.001719 | 0.001698 | -152.2599 | 0.23618 | 8.4048 | 0.003308 | 0.004936 | 0.004822 | 5.2e-05 | 0.001219 | 74.4247 | 1.3214 | 2.5481 | 101.21 | 20.83 | 17.445 | 63.1947 | 101.4446 | 62.6776 | 0.004247 | 1.0513 | 0.025768 | 24.5115 | -0.009726 | 0.1405 | -2.4e-05 | -0.002002 | -0.002002 | -2.9199
2019-03-19 15:34 | 1553024047.7536 | -482 | 0.20874 | 3.7e-05 | 3e-06 | 101.5472 | 101.8363 | 128.8651 | 21.2612 | 0.001756 | 0.001687 | -152.2683 | 0.23603 | 8.401 | 0.003245 | 0.004694 | 0.004463 | 5.1e-05 | 0.001221 | 74.4642 | 1.3153 | 2.5481 | 101.21 | 20.83 | 17.455 | 64.0732 | 99.9211 | 62.4103 | 0.004249 | 1.0521 | 0.02578 | 24.5028 | -0.025433 | 0.14042 | -2.4e-05 | -0.002039 | -0.002039 | -2.8644
2019-03-19 15:34 | 1553024048.7542 | -481 | 0.20876 | 3.7e-05 | 3e-06 | 101.5421 | 101.8185 | 128.6743 | 21.2612 | 0.001771 | 0.001628 | -152.2707 | 0.23607 | 8.4033 | 0.003214 | 0.004644 | 0.004356 | 5.2e-05 | 0.001223 | 74.5149 | 1.2969 | 2.548 | 101.21 | 20.83 | 17.46 | 64.4416 | 99.1711 | 61.0006 | 0.004251 | 1.0565 | 0.025815 | 24.4332 | -0.024319 | 0.14044 | -2.4e-05 | -0.00206 | -0.00206 | -2.8364
2019-03-19 15:34 | 1553024049.7528 | -480 | 0.20878 | 3.7e-05 | 2e-06 | 101.4736 | 101.5952 | 128.7227 | 21.2612 | 0.001825 | 0.001647 | -152.256 | 0.23543 | 8.4023 | 0.003229 | 0.004567 | 0.004296 | 5.1e-05 | 0.001226 | 74.5026 | 1.3277 | 2.5479 | 101.21 | 20.83 | 17.47 | 65.7335 | 99.5335 | 61.447 | 0.004253 | 1.0551 | 0.025802 | 24.4541 | -0.010709 | 0.14006 | -2.4e-05 | -0.002062 | -0.002062 | -2.8258
2019-03-19 15:34 | 1553024050.7528 | -479 | 0.20874 | 3.6e-05 | 3e-06 | 101.5124 | 101.6004 | 128.6361 | 21.2612 | 0.001796 | 0.001675 | -152.263 | 0.23531 | 8.4016 | 0.003277 | 0.004364 | 0.004232 | 5.5e-05 | 0.001222 | 74.478 | 1.3194 | 2.548 | 101.21 | 20.83 | 17.475 | 65.0318 | 100.6832 | 62.1182 | 0.004254 | 1.053 | 0.025768 | 24.4704 | -0.00775 | 0.13998 | -2.3e-05 | -0.001945 | -0.001945 | -2.9941
2019-03-19 15:34 | 1553024051.7536 | -478 | 0.20876 | 3.8e-05 | 3e-06 | 101.5016 | 101.5144 | 128.6321 | 21.2612 | 0.001815 | 0.001633 | -152.273 | 0.23673 | 8.4 | 0.003296 | 0.004115 | 0.003884 | 5.3e-05 | 0.001219 | 74.5194 | 1.3053 | 2.5481 | 101.21 | 20.83 | 17.479 | 65.4954 | 101.1392 | 61.1141 | 0.004255 | 1.0562 | 0.025806 | 24.4333 | -0.001131 | 0.14083 | -2.5e-05 | -0.002151 | -0.002151 | -2.7237
2019-03-19 15:34 | 1553024052.7537 | -477 | 0.20877 | 3.6e-05 | 2e-06 | 101.4565 | 101.6363 | 128.6332 | 21.2612 | 0.001762 | 0.001672 | -152.2969 | 0.23593 | 8.4032 | 0.003307 | 0.004116 | 0.003867 | 5.2e-05 | 0.001218 | 74.4936 | 1.332 | 2.5482 | 101.21 | 20.83 | 17.485 | 64.2207 | 101.4093 | 62.0643 | 0.004257 | 1.0532 | 0.02577 | 24.4681 | -0.015835 | 0.14035 | -2.2e-05 | -0.001865 | -0.001865 | -3.1305
2019-03-19 15:34 | 1553024053.7522 | -476 | 0.20875 | 3.4e-05 | 2e-06 | 101.4034 | 101.6137 | 128.6497 | 21.2612 | 0.001742 | 0.001658 | -152.2927 | 0.236 | 8.4061 | 0.003267 | 0.004736 | 0.004541 | 5.5e-05 | 0.001218 | 74.5262 | 1.3924 | 2.5481 | 101.21 | 20.83 | 17.489 | 63.7538 | 100.4367 | 61.7322 | 0.004258 | 1.0542 | 0.025784 | 24.4576 | -0.01853 | 0.1404 | -2e-05 | -0.001717 | -0.001717 | -3.4025
2019-03-19 15:34 | 1553024054.7529 | -475 | 0.20876 | 3.9e-05 | 3e-06 | 101.4466 | 101.7127 | 128.7659 | 21.2612 | 0.001771 | 0.00172 | -152.2887 | 0.23574 | 8.4047 | 0.003329 | 0.004559 | 0.004337 | 5.5e-05 | 0.001218 | 74.488 | 1.3438 | 2.5482 | 101.21 | 20.83 | 17.492 | 64.4486 | 101.9589 | 63.2041 | 0.004258 | 1.0496 | 0.025739 | 24.5223 | -0.023427 | 0.14024 | -2.7e-05 | -0.002284 | -0.002284 | -2.5544
2019-03-19 15:34 | 1553024055.7534 | -474 | 0.20873 | 4.5e-05 | 2e-06 | 101.4406 | 101.6164 | 128.8238 | 21.2612 | 0.001806 | 0.001656 | -152.2842 | 0.23673 | 8.401 | 0.003324 | 0.004992 | 0.004753 | 5.6e-05 | 0.001217 | 74.5155 | 1.3416 | 2.548 | 101.21 | 20.83 | 17.493 | 65.2947 | 101.8354 | 61.6633 | 0.004259 | 1.0545 | 0.025804 | 24.4716 | -0.015489 | 0.14083 | -3.3e-05 | -0.002726 | -0.002726 | -2.1493
2019-03-19 15:34 | 1553024056.7531 | -473 | 0.20877 | 4.2e-05 | 2e-06 | 101.4289 | 101.6051 | 128.686 | 21.2612 | 0.001829 | 0.001669 | -152.2926 | 0.23629 | 8.4012 | 0.003365 | 0.0048 | 0.004606 | 5.4e-05 | 0.001217 | 74.4602 | 1.3514 | 2.5481 | 101.21 | 20.83 | 17.492 | 65.8442 | 102.8198 | 61.9766 | 0.004258 | 1.0535 | 0.025778 | 24.4699 | -0.015526 | 0.14057 | -3.1e-05 | -0.002545 | -0.002545 | -2.2979
2019-03-19 15:34 | 1553024057.7531 | -472 | 0.20875 | 3.8e-05 | 2e-06 | 101.3812 | 101.5647 | 128.7583 | 21.2612 | 0.001833 | 0.001674 | -152.3032 | 0.23639 | 8.404 | 0.003419 | 0.004967 | 0.004741 | 5.5e-05 | 0.001214 | 74.4695 | 1.3914 | 2.5481 | 101.21 | 20.84 | 17.487 | 65.9332 | 104.1361 | 62.1057 | 0.00426 | 1.0531 | 0.025781 | 24.4815 | -0.016176 | 0.14063 | -2.5e-05 | -0.002082 | -0.002082 | -2.81
2019-03-19 15:34 | 1553024058.7531 | -471 | 0.20876 | 3.6e-05 | 2e-06 | 101.398 | 101.5158 | 128.9379 | 21.2612 | 0.001762 | 0.001697 | -152.2889 | 0.23485 | 8.4036 | 0.003359 | 0.005562 | 0.00521 | 5.5e-05 | 0.001219 | 74.4638 | 1.3752 | 2.548 | 101.21 | 20.84 | 17.483 | 64.2245 | 102.6634 | 62.6501 | 0.004259 | 1.0514 | 0.025778 | 24.5185 | -0.01038 | 0.13971 | -2.3e-05 | -0.001961 | -0.001961 | -2.9643
2019-03-19 15:34 | 1553024059.7536 | -470 | 0.20891 | 3.7e-05 | 2e-06 | 101.3542 | 101.7555 | 128.7182 | 21.2612 | 0.001812 | 0.001665 | -152.2855 | 0.23598 | 8.4023 | 0.003311 | 0.005132 | 0.004906 | 5.4e-05 | 0.001214 | 74.5631 | 1.4063 | 2.5481 | 101.21 | 20.84 | 17.477 | 65.4174 | 101.5216 | 61.8774 | 0.004257 | 1.0538 | 0.025785 | 24.4694 | -0.035345 | 0.14038 | -2.3e-05 | -0.001911 | -0.001911 | -3.0564
2019-03-19 15:34 | 1553024060.7523 | -469 | 0.20903 | 3.6e-05 | 3e-06 | 101.3539 | 101.5027 | 128.7177 | 21.2612 | 0.001799 | 0.001675 | -152.2877 | 0.23594 | 8.4027 | 0.003399 | 0.00388 | 0.003613 | 5.5e-05 | 0.001204 | 74.5625 | 1.3905 | 2.5482 | 101.21 | 20.84 | 17.462 | 65.1145 | 103.637 | 62.1408 | 0.004254 | 1.053 | 0.025775 | 24.479 | -0.013125 | 0.14036 | -2.2e-05 | -0.0019 | -0.0019 | -3.0726
2019-03-19 15:34 | 1553024061.7522 | -468 | 0.20911 | 3.5e-05 | 2e-06 | 101.3449 | 101.6603 | 128.4839 | 21.2612 | 0.001785 | 0.001634 | -152.2744 | 0.23632 | 8.4025 | 0.003316 | 0.004647 | 0.004388 | 5.6e-05 | 0.00121 | 74.4819 | 1.3433 | 2.5481 | 101.21 | 20.84 | 17.455 | 64.7703 | 101.6234 | 61.1514 | 0.004252 | 1.0561 | 0.02579 | 24.4206 | -0.027792 | 0.14059 | -2.1e-05 | -0.001763 | -0.001763 | -3.3166
2019-03-19 15:34 | 1553024062.7527 | -467 | 0.20923 | 3.7e-05 | 2e-06 | 101.383 | 101.4955 | 128.3208 | 21.2612 | 0.001789 | 0.001633 | -152.2677 | 0.23716 | 8.4042 | 0.003241 | 0.004178 | 0.004012 | 5.5e-05 | 0.00121 | 74.4916 | 1.3606 | 2.548 | 101.21 | 20.84 | 17.437 | 64.8631 | 99.8176 | 61.1218 | 0.004248 | 1.0562 | 0.025775 | 24.404 | -0.009923 | 0.14109 | -2.4e-05 | -0.002019 | -0.002019 | -2.9076
2019-03-19 15:34 | 1553024063.7524 | -466 | 0.20925 | 2.8e-05 | 2e-06 | 101.2852 | 101.5798 | 128.1023 | 21.2612 | 0.001797 | 0.001659 | -152.2669 | 0.23636 | 8.4046 | 0.003253 | 0.004613 | 0.004365 | 5.4e-05 | 0.001209 | 74.523 | 1.3792 | 2.548 | 101.21 | 20.84 | 17.428 | 65.0645 | 100.1204 | 61.7534 | 0.004245 | 1.0542 | 0.025728 | 24.4062 | -0.02598 | 0.14061 | -1.2e-05 | -0.001057 | -0.001057 | -5.535
2019-03-19 15:34 | 1553024064.7528 | -465 | 0.20931 | 3e-05 | 2e-06 | 101.3611 | 101.563 | 127.7131 | 21.2612 | 0.001726 | 0.00166 | -152.2667 | 0.23658 | 8.3997 | 0.003254 | 0.00491 | 0.004716 | 5.6e-05 | 0.001203 | 74.4841 | 1.3189 | 2.548 | 101.21 | 20.84 | 17.42 | 63.3687 | 100.1245 | 61.7712 | 0.004243 | 1.0541 | 0.025689 | 24.3698 | -0.0178 | 0.14074 | -1.4e-05 | -0.001237 | -0.001237 | -4.7323
2019-03-19 15:34 | 1553024065.7526 | -464 | 0.20935 | 3e-05 | 1e-06 | 101.2684 | 101.5346 | 127.5117 | 21.2612 | 0.001798 | 0.001647 | -152.2788 | 0.2374 | 8.4053 | 0.003261 | 0.004523 | 0.004173 | 5.6e-05 | 0.001208 | 74.4898 | 1.3506 | 2.5481 | 101.21 | 20.84 | 17.402 | 65.0941 | 100.3112 | 61.4617 | 0.004239 | 1.0551 | 0.02568 | 24.3393 | -0.023483 | 0.14123 | -1.4e-05 | -0.001139 | -0.001139 | -5.1573
2019-03-19 15:34 | 1553024066.7526 | -463 | 0.20938 | 2.8e-05 | 2e-06 | 101.2672 | 101.3996 | 127.385 | 21.2612 | 0.001724 | 0.001698 | -152.2648 | 0.23615 | 8.404 | 0.003358 | 0.00477 | 0.004593 | 5.6e-05 | 0.001203 | 74.4913 | 1.2924 | 2.5481 | 101.21 | 20.83 | 17.393 | 63.3179 | 102.6515 | 62.6869 | 0.004234 | 1.0512 | 0.025621 | 24.3717 | -0.011694 | 0.14048 | -1.1e-05 | -0.000997 | -0.000997 | -5.8631
2019-03-19 15:34 | 1553024067.7539 | -462 | 0.20938 | 1.9e-05 | 2e-06 | 101.2387 | 101.5096 | 127.3455 | 21.2612 | 0.001788 | 0.001602 | -152.2454 | 0.236 | 8.4023 | 0.003279 | 0.00493 | 0.004722 | 5.3e-05 | 0.001201 | 72.4473 | 1.366 | 2.5482 | 101.21 | 20.83 | 17.377 | 64.8562 | 100.7263 | 60.3761 | 0.00423 | 1.0585 | 0.025705 | 24.284 | -0.023905 | 0.1404 | 0 | -9.6e-05 | -9.6e-05 | -61.0955
2019-03-19 15:34 | 1553024068.7526 | -461 | 0.20935 | 1.9e-05 | 1e-06 | 101.2135 | 101.401 | 127.3742 | 21.2612 | 0.00181 | 0.00163 | -152.2423 | 0.23636 | 8.4017 | 0.003341 | 0.005334 | 0.005092 | 5.5e-05 | 0.001195 | 46.8554 | 1.3883 | 2.5481 | 101.21 | 20.84 | 17.371 | 65.3876 | 102.2259 | 61.0553 | 0.004231 | 1.0564 | 0.025682 | 24.3114 | -0.01656 | 0.14061 | 0 | -0.0001 | -0.0001 | -58.5036
2019-03-19 15:34 | 1553024069.7527 | -460 | 0.2094 | 2.4e-05 | 2e-06 | 101.1846 | 101.3312 | 127.3852 | 21.2612 | 0.001781 | 0.001655 | -152.2533 | 0.23625 | 8.4011 | 0.003371 | 0.004686 | 0.004525 | 5.4e-05 | 0.001188 | 72.2619 | 1.3962 | 2.548 | 101.21 | 20.83 | 17.365 | 64.6788 | 102.9753 | 61.6602 | 0.004227 | 1.0545 | 0.02566 | 24.3345 | -0.012942 | 0.14055 | -7e-06 | -0.000685 | -0.000685 | -8.5366
2019-03-19 15:34 | 1553024070.7528 | -459 | 0.20938 | 2.1e-05 | 2e-06 | 101.1705 | 101.2599 | 127.4324 | 21.2612 | 0.001777 | 0.001668 | -152.2349 | 0.23688 | 8.4012 | 0.003331 | 0.005148 | 0.004877 | 5.4e-05 | 0.001185 | 74.4493 | 1.4079 | 2.5482 | 101.21 | 20.83 | 17.357 | 64.5867 | 102.0029 | 61.9675 | 0.004225 | 1.0535 | 0.025653 | 24.3501 | -0.007898 | 0.14092 | -3e-06 | -0.000376 | -0.000376 | -15.595
2019-03-19 15:34 | 1553024071.7531 | -458 | 0.2094 | 2e-05 | 2e-06 | 101.1861 | 101.3722 | 127.5831 | 21.2612 | 0.001749 | 0.001628 | -152.2469 | 0.23637 | 8.4017 | 0.003305 | 0.004849 | 0.00462 | 5.5e-05 | 0.001184 | 74.5158 | 1.3365 | 2.548 | 101.21 | 20.83 | 17.355 | 63.9122 | 101.3778 | 61.0038 | 0.004225 | 1.0565 | 0.025705 | 24.3295 | -0.016436 | 0.14062 | -1e-06 | -0.000239 | -0.000239 | -24.4944
2019-03-19 15:35 | 1553024072.7529 | -457 | 0.2094 | 1.7e-05 | 2e-06 | 101.0802 | 101.3113 | 127.6786 | 21.2612 | 0.001827 | 0.001685 | -152.2354 | 0.23671 | 8.4022 | 0.003361 | 0.005221 | 0.004983 | 5e-05 | 0.001193 | 74.4873 | 1.3125 | 2.5481 | 101.21 | 20.83 | 17.354 | 65.7912 | 102.7334 | 62.3807 | 0.004225 | 1.0522 | 0.025662 | 24.3887 | -0.020425 | 0.14082 | 2e-06 | 3.1e-05 | 3.1e-05 | 186.3242
2019-03-19 15:35 | 1553024073.7533 | -456 | 0.20945 | 2.2e-05 | 2e-06 | 101.1618 | 101.2537 | 127.7411 | 21.2612 | 0.001791 | 0.001636 | -152.2423 | 0.23672 | 8.4027 | 0.003353 | 0.004165 | 0.004039 | 5.2e-05 | 0.001199 | 74.5148 | 1.3437 | 2.548 | 101.21 | 20.83 | 17.354 | 64.9185 | 102.5227 | 61.196 | 0.004225 | 1.0559 | 0.025713 | 24.3515 | -0.008123 | 0.14083 | -4e-06 | -0.00044 | -0.00044 | -13.3004
2019-03-19 15:35 | 1553024074.7534 | -455 | 0.20941 | 2.2e-05 | 2e-06 | 101.1126 | 101.263 | 127.649 | 21.2612 | 0.001807 | 0.001638 | -152.2238 | 0.23638 | 8.4023 | 0.003355 | 0.004553 | 0.004175 | 4.9e-05 | 0.001201 | 74.5871 | 1.301 | 2.5481 | 101.21 | 20.83 | 17.356 | 65.3104 | 102.5818 | 61.2343 | 0.004225 | 1.0558 | 0.025703 | 24.3441 | -0.013293 | 0.14062 | -5e-06 | -0.000572 | -0.000572 | -10.2216
2019-03-19 15:35 | 1553024075.7531 | -454 | 0.20943 | 2.1e-05 | 2e-06 | 101.1267 | 101.4051 | 127.5373 | 21.2612 | 0.001794 | 0.001636 | -152.2152 | 0.23582 | 8.4028 | 0.003401 | 0.005009 | 0.004828 | 5.2e-05 | 0.001204 | 74.5925 | 1.3159 | 2.5478 | 101.21 | 20.83 | 17.361 | 65.0051 | 103.6937 | 61.1844 | 0.004226 | 1.056 | 0.025693 | 24.3317 | -0.024591 | 0.14029 | -2e-06 | -0.000297 | -0.000297 | -19.6476
2019-03-19 15:35 | 1553024076.7525 | -453 | 0.20943 | 2e-05 | 2e-06 | 101.1066 | 101.2668 | 127.2594 | 21.2612 | 0.001761 | 0.001695 | -152.2127 | 0.2363 | 8.4022 | 0.003393 | 0.004103 | 0.003835 | 5.2e-05 | 0.001202 | 74.5045 | 1.3276 | 2.548 | 101.21 | 20.83 | 17.364 | 64.2065 | 103.4885 | 62.6063 | 0.004227 | 1.0515 | 0.025611 | 24.3568 | -0.014156 | 0.14057 | -1e-06 | -0.000186 | -0.000186 | -31.4994
2019-03-19 15:35 | 1553024077.7532 | -452 | 0.20942 | 2.2e-05 | 2e-06 | 101.1462 | 101.2637 | 127.088 | 21.2612 | 0.001785 | 0.00164 | -152.2258 | 0.23663 | 8.403 | 0.003283 | 0.004799 | 0.0047 | 4.9e-05 | 0.001202 | 74.5842 | 1.3683 | 2.5481 | 101.21 | 20.84 | 17.369 | 64.7741 | 100.8329 | 61.2813 | 0.004231 | 1.0557 | 0.025644 | 24.2923 | -0.01038 | 0.14077 | -5e-06 | -0.000554 | -0.000554 | -10.5795
2019-03-19 15:35 | 1553024078.7521 | -451 | 0.20943 | 2.2e-05 | 2e-06 | 101.0352 | 101.2051 | 127.1174 | 21.2612 | 0.001746 | 0.001622 | -152.2336 | 0.23662 | 8.4059 | 0.003375 | 0.004732 | 0.004517 | 4.9e-05 | 0.001201 | 74.6167 | 1.379 | 2.548 | 101.21 | 20.84 | 17.38 | 63.8481 | 103.054 | 60.8505 | 0.004234 | 1.057 | 0.025664 | 24.2795 | -0.01503 | 0.14077 | -4e-06 | -0.000404 | -0.000404 | -14.4827
2019-03-19 15:35 | 1553024079.7524 | -450 | 0.20942 | 2e-05 | 2e-06 | 101.1279 | 101.2308 | 127.1404 | 21.2612 | 0.001772 | 0.001656 | -152.2324 | 0.23623 | 8.4019 | 0.003334 | 0.005205 | 0.005004 | 4.8e-05 | 0.001206 | 74.5568 | 1.3613 | 2.548 | 101.21 | 20.84 | 17.38 | 64.4534 | 102.0688 | 61.6845 | 0.004234 | 1.0544 | 0.025634 | 24.3119 | -0.009095 | 0.14053 | -2e-06 | -0.000303 | -0.000303 | -19.2934
2019-03-19 15:35 | 1553024080.7532 | -449 | 0.20946 | 1.8e-05 | 2e-06 | 101.1447 | 101.2757 | 127.469 | 21.2612 | 0.001801 | 0.001664 | -152.2562 | 0.23574 | 8.4035 | 0.003328 | 0.005018 | 0.004928 | 4.4e-05 | 0.001209 | 74.5733 | 1.3236 | 2.5481 | 101.21 | 20.84 | 17.39 | 65.1598 | 101.9313 | 61.8533 | 0.004236 | 1.0539 | 0.025661 | 24.3495 | -0.011572 | 0.14024 | 0 | -0.000146 | -0.000146 | -39.9792
2019-03-19 15:35 | 1553024081.7527 | -448 | 0.20945 | 1.6e-05 | 2e-06 | 101.0358 | 101.0092 | 127.6859 | 21.2612 | 0.001845 | 0.001601 | -152.2503 | 0.23691 | 8.4027 | 0.003315 | 0.004503 | 0.004236 | 4.8e-05 | 0.001215 | 74.6131 | 1.3436 | 2.5481 | 101.21 | 20.84 | 17.395 | 66.2283 | 101.6079 | 60.3361 | 0.004237 | 1.0587 | 0.025741 | 24.3149 | 0.002354 | 0.14094 | 3e-06 | 9e-05 | 9e-05 | 65.1998
2019-03-19 15:35 | 1553024082.7533 | -447 | 0.20941 | 2.4e-05 | 2e-06 | 101.0658 | 101.1489 | 127.8112 | 21.2612 | 0.001747 | 0.001604 | -152.2578 | 0.23639 | 8.4018 | 0.003402 | 0.004619 | 0.004377 | 4.9e-05 | 0.001215 | 74.6274 | 1.3408 | 2.5481 | 101.21 | 20.84 | 17.399 | 63.8693 | 103.7102 | 60.4082 | 0.004238 | 1.0584 | 0.025751 | 24.3295 | -0.007355 | 0.14063 | -7e-06 | -0.000633 | -0.000633 | -9.2492
2019-03-19 15:35 | 1553024083.7524 | -446 | 0.2094 | 2.4e-05 | 2e-06 | 101.0199 | 101.0767 | 128.0028 | 21.2612 | 0.001742 | 0.001704 | -152.2562 | 0.2362 | 8.404 | 0.003389 | 0.004358 | 0.004156 | 5.1e-05 | 0.001208 | 74.6366 | 1.3275 | 2.548 | 101.21 | 20.84 | 17.4 | 63.7483 | 103.3982 | 62.8182 | 0.004239 | 1.0508 | 0.025678 | 24.4355 | -0.005027 | 0.14052 | -7e-06 | -0.000655 | -0.000655 | -8.9288
2019-03-19 15:35 | 1553024084.7526 | -445 | 0.20946 | 1.9e-05 | 1e-06 | 101.0493 | 101.2609 | 128.2603 | 21.2612 | 0.001779 | 0.001654 | -152.2656 | 0.23509 | 8.401 | 0.003309 | 0.00521 | 0.005007 | 4.6e-05 | 0.001208 | 74.6257 | 1.3451 | 2.5481 | 101.21 | 20.84 | 17.399 | 64.6415 | 101.4642 | 61.6354 | 0.004238 | 1.0545 | 0.025749 | 24.417 | -0.018714 | 0.13986 | 0 | -6.2e-05 | -6.2e-05 | -93.7312
2019-03-19 15:35 | 1553024085.7516 | -444 | 0.20947 | 2.4e-05 | 1e-06 | 101.0184 | 101.1942 | 128.5774 | 21.2612 | 0.001758 | 0.001646 | -152.2648 | 0.23664 | 8.4058 | 0.003361 | 0.004177 | 0.003996 | 4.7e-05 | 0.001213 | 74.5989 | 1.3436 | 2.548 | 101.21 | 20.84 | 17.396 | 64.1278 | 102.7162 | 61.4369 | 0.004238 | 1.0552 | 0.025788 | 24.4399 | -0.01555 | 0.14078 | -6e-06 | -0.000549 | -0.000549 | -10.6699
2019-03-19 15:35 | 1553024086.7516 | -443 | 0.20947 | 2.1e-05 | 1e-06 | 101.0589 | 101.1489 | 128.8292 | 21.2612 | 0.001751 | 0.001697 | -152.2892 | 0.23518 | 8.4006 | 0.003335 | 0.005159 | 0.004879 | 4.4e-05 | 0.001218 | 74.6549 | 1.3385 | 2.5481 | 101.21 | 20.84 | 17.393 | 63.9598 | 102.0978 | 62.6556 | 0.004237 | 1.0513 | 0.025767 | 24.5083 | -0.007966 | 0.13991 | -3e-06 | -0.000286 | -0.000286 | -20.3438
2019-03-19 15:35 | 1553024087.7519 | -442 | 0.20949 | 2.5e-05 | 2e-06 | 101.0343 | 101.263 | 128.8074 | 21.2612 | 0.001802 | 0.001695 | -152.2689 | 0.23587 | 8.4018 | 0.003393 | 0.004269 | 0.004092 | 4.3e-05 | 0.00122 | 74.5986 | 1.3434 | 2.5481 | 101.21 | 20.84 | 17.382 | 65.187 | 103.498 | 62.6017 | 0.004234 | 1.0515 | 0.025766 | 24.5043 | -0.020224 | 0.14032 | -7e-06 | -0.000669 | -0.000669 | -8.7281
2019-03-19 15:35 | 1553024088.7527 | -441 | 0.20942 | 2.5e-05 | 2e-06 | 101.0469 | 100.9489 | 128.8464 | 21.2612 | 0.001725 | 0.00167 | -152.2609 | 0.23533 | 8.4013 | 0.003385 | 0.004619 | 0.004378 | 4.9e-05 | 0.001219 | 74.6084 | 1.3309 | 2.548 | 101.21 | 20.84 | 17.377 | 63.3415 | 103.2935 | 62.0204 | 0.004233 | 1.0533 | 0.025793 | 24.4868 | 0.008676 | 0.14 | -8e-06 | -0.000709 | -0.000709 | -8.2166
2019-03-19 15:35 | 1553024089.7528 | -440 | 0.20943 | 2.9e-05 | 2e-06 | 101.0148 | 101.4127 | 128.9269 | 21.2612 | 0.001777 | 0.001685 | -152.244 | 0.23628 | 8.4048 | 0.003448 | 0.00496 | 0.004841 | 4.9e-05 | 0.001215 | 74.6642 | 1.2964 | 2.5479 | 101.21 | 20.84 | 17.37 | 64.5734 | 104.8361 | 62.365 | 0.004231 | 1.0522 | 0.025787 | 24.507 | -0.035162 | 0.14057 | -1.2e-05 | -0.001091 | -0.001091 | -5.358
2019-03-19 15:35 | 1553024090.7522 | -439 | 0.20944 | 2.5e-05 | 2e-06 | 100.9218 | 101.2051 | 128.421 | 21.2612 | 0.001755 | 0.001657 | -152.2392 | 0.23747 | 8.4012 | 0.003439 | 0.004974 | 0.004701 | 4.5e-05 | 0.00121 | 74.7217 | 1.3292 | 2.5481 | 101.21 | 20.84 | 17.355 | 64.0527 | 104.6167 | 61.7009 | 0.004228 | 1.0543 | 0.025762 | 24.4347 | -0.025079 | 0.14127 | -8e-06 | -0.000701 | -0.000701 | -8.3793
2019-03-19 15:35 | 1553024091.7531 | -438 | 0.20942 | 2.5e-05 | 2e-06 | 101.0475 | 101.0359 | 128.0638 | 21.2612 | 0.001789 | 0.001659 | -152.2419 | 0.23591 | 8.4028 | 0.003426 | 0.004896 | 0.004622 | 4.7e-05 | 0.001211 | 74.7974 | 1.3168 | 2.548 | 101.21 | 20.83 | 17.348 | 64.8812 | 104.2919 | 61.7384 | 0.004223 | 1.0542 | 0.025725 | 24.402 | 0.001025 | 0.14034 | -8e-06 | -0.000772 | -0.000772 | -7.5627
2019-03-19 15:35 | 1553024092.7524 | -437 | 0.20947 | 2.3e-05 | 1e-06 | 100.9455 | 100.9332 | 127.7425 | 21.2612 | 0.001778 | 0.001647 | -152.2278 | 0.23639 | 8.4022 | 0.003388 | 0.005015 | 0.004818 | 4.7e-05 | 0.001215 | 74.7955 | 1.3376 | 2.548 | 101.21 | 20.83 | 17.33 | 64.6054 | 103.3706 | 61.4571 | 0.004219 | 1.0551 | 0.025704 | 24.3612 | 0.001089 | 0.14063 | -4e-06 | -0.00038 | -0.00038 | -15.3965
2019-03-19 15:35 | 1553024093.8287 | -436 | 0.20943 | 2.3e-05 | 1e-06 | 101.0508 | 101.2041 | 127.5603 | 21.2612 | 0.001742 | 0.001687 | -152.2231 | 0.23569 | 8.4037 | 0.003442 | 0.004812 | 0.004586 | 4.9e-05 | 0.001211 | 74.914 | 1.3063 | 2.5479 | 101.21 | 20.83 | 17.323 | 63.7365 | 104.6724 | 62.4259 | 0.004217 | 1.0521 | 0.025648 | 24.379 | -0.013558 | 0.14021 | -5e-06 | -0.000441 | -0.000441 | -13.233
2019-03-19 15:35 | 1553024094.7648 | -435 | 0.20944 | 2.5e-05 | 2e-06 | 101.055 | 101.1078 | 127.6751 | 21.2612 | 0.001829 | 0.001681 | -152.2261 | 0.23666 | 8.4031 | 0.003455 | 0.004759 | 0.004467 | 4.8e-05 | 0.001209 | 74.9276 | 1.3448 | 2.5479 | 101.21 | 20.83 | 17.318 | 65.8271 | 105.006 | 62.2697 | 0.004216 | 1.0525 | 0.025666 | 24.3843 | -0.004676 | 0.14079 | -8e-06 | -0.000755 | -0.000755 | -7.7577
2019-03-19 15:35 | 1553024095.7649 | -434 | 0.20944 | 2.2e-05 | 1e-06 | 100.9902 | 101.213 | 127.6333 | 21.2612 | 0.001805 | 0.001673 | -152.2329 | 0.23652 | 8.4046 | 0.003461 | 0.004595 | 0.004365 | 5.1e-05 | 0.001211 | 75.0274 | 1.346 | 2.5479 | 101.21 | 20.83 | 17.309 | 65.2596 | 105.1333 | 62.079 | 0.004214 | 1.0531 | 0.025669 | 24.3734 | -0.019713 | 0.1407 | -4e-06 | -0.000366 | -0.000366 | -16.011
2019-03-19 15:35 | 1553024096.7638 | -433 | 0.20941 | 1.9e-05 | 2e-06 | 100.9224 | 101.0654 | 127.6772 | 21.2612 | 0.00174 | 0.001637 | -152.2143 | 0.23634 | 8.4009 | 0.003471 | 0.004829 | 0.00459 | 4.9e-05 | 0.001219 | 75.106 | 1.3545 | 2.548 | 101.21 | 20.83 | 17.306 | 63.6982 | 105.3835 | 61.2227 | 0.004213 | 1.0558 | 0.025706 | 24.3464 | -0.012667 | 0.1406 | 0 | -0.000108 | -0.000108 | -54.266
2019-03-19 15:35 | 1553024097.7651 | -432 | 0.20944 | 1.8e-05 | 1e-06 | 100.8677 | 101.0123 | 128.0335 | 21.2612 | 0.001681 | 0.001641 | -152.2284 | 0.23702 | 8.4045 | 0.00339 | 0.00462 | 0.004452 | 4.8e-05 | 0.001214 | 75.1917 | 1.331 | 2.5479 | 101.21 | 20.83 | 17.304 | 62.2735 | 103.4139 | 61.3179 | 0.004213 | 1.0555 | 0.025738 | 24.3838 | -0.012809 | 0.141 | 2e-06 | 5.7e-05 | 5.7e-05 | 103.3161
2019-03-19 15:35 | 1553024098.764 | -431 | 0.20946 | 2.7e-05 | 2e-06 | 100.9797 | 100.9664 | 128.1274 | 21.2612 | 0.001825 | 0.001672 | -152.2271 | 0.23719 | 8.4012 | 0.003381 | 0.004593 | 0.00443 | 4.8e-05 | 0.001211 | 75.2315 | 1.3145 | 2.5479 | 101.21 | 20.83 | 17.303 | 65.752 | 103.1976 | 62.0604 | 0.004212 | 1.0532 | 0.025719 | 24.4198 | 0.001176 | 0.14111 | -1e-05 | -0.000857 | -0.000857 | -6.8457
2019-03-19 15:35 | 1553024099.7632 | -430 | 0.20944 | 2.5e-05 | 1e-06 | 100.9113 | 101.0914 | 128.2091 | 21.2612 | 0.001798 | 0.001659 | -152.2346 | 0.23672 | 8.4022 | 0.003427 | 0.004762 | 0.004548 | 5.1e-05 | 0.001212 | 75.2334 | 1.282 | 2.5479 | 101.21 | 20.83 | 17.301 | 65.0794 | 104.3217 | 61.7548 | 0.004212 | 1.0542 | 0.025739 | 24.4165 | -0.015954 | 0.14083 | -7e-06 | -0.000627 | -0.000627 | -9.3426
2019-03-19 15:35 | 1553024100.7643 | -429 | 0.20944 | 2.5e-05 | 2e-06 | 100.9034 | 101.0472 | 128.5497 | 21.2612 | 0.00181 | 0.001588 | -152.2391 | 0.23665 | 8.4037 | 0.003337 | 0.004817 | 0.004592 | 4.8e-05 | 0.001215 | 75.2247 | 1.2799 | 2.5479 | 101.21 | 20.83 | 17.293 | 65.3886 | 102.1419 | 60.0375 | 0.00421 | 1.0596 | 0.02584 | 24.3861 | -0.012736 | 0.14078 | -8e-06 | -0.000703 | -0.000703 | -8.3269
2019-03-19 15:35 | 1553024101.7634 | -428 | 0.20944 | 2.3e-05 | 2e-06 | 100.8425 | 100.9061 | 128.9384 | 21.2612 | 0.001778 | 0.001658 | -152.2372 | 0.23677 | 8.4036 | 0.003388 | 0.004971 | 0.004753 | 4.8e-05 | 0.001215 | 75.2422 | 1.26 | 2.5477 | 101.21 | 20.83 | 17.288 | 64.618 | 103.387 | 61.7151 | 0.004209 | 1.0543 | 0.025814 | 24.4844 | -0.005639 | 0.14086 | -5e-06 | -0.000494 | -0.000494 | -11.873
2019-03-19 15:35 | 1553024102.7644 | -427 | 0.20947 | 2.7e-05 | 1e-06 | 100.874 | 100.9948 | 129.1163 | 21.2612 | 0.001733 | 0.001681 | -152.2523 | 0.23641 | 8.4006 | 0.003327 | 0.004955 | 0.004764 | 4.5e-05 | 0.00121 | 75.313 | 1.316 | 2.5478 | 101.21 | 20.83 | 17.276 | 63.5225 | 101.8984 | 62.2728 | 0.004206 | 1.0525 | 0.02581 | 24.5216 | -0.010704 | 0.14064 | -9e-06 | -0.000755 | -0.000755 | -7.7531
2019-03-19 15:35 | 1553024103.7645 | -426 | 0.20945 | 2.5e-05 | 1e-06 | 100.8119 | 100.8849 | 129.3599 | 21.2612 | 0.001722 | 0.00166 | -152.2483 | 0.23628 | 8.4021 | 0.003308 | 0.004453 | 0.004252 | 4.6e-05 | 0.001203 | 75.3421 | 1.3306 | 2.5479 | 101.21 | 20.83 | 17.269 | 63.2734 | 101.4429 | 61.7681 | 0.004204 | 1.0541 | 0.025854 | 24.5263 | -0.006472 | 0.14056 | -8e-06 | -0.000689 | -0.000689 | -8.4825
2019-03-19 15:35 | 1553024104.7643 | -425 | 0.20945 | 2.6e-05 | 2e-06 | 100.8647 | 100.9356 | 129.2914 | 21.2612 | 0.001822 | 0.001724 | -152.2515 | 0.23614 | 8.4023 | 0.003421 | 0.004564 | 0.004272 | 4.6e-05 | 0.0012 | 75.3301 | 1.3736 | 2.5479 | 101.21 | 20.83 | 17.261 | 65.671 | 104.1865 | 63.3141 | 0.004202 | 1.0493 | 0.025787 | 24.5763 | -0.00628 | 0.14048 | -9e-06 | -0.00078 | -0.00078 | -7.4915
2019-03-19 15:35 | 1553024105.7634 | -424 | 0.20945 | 2.9e-05 | 2e-06 | 100.8659 | 101.0609 | 129.3909 | 21.2612 | 0.001814 | 0.001633 | -152.2501 | 0.23588 | 8.4032 | 0.003344 | 0.00499 | 0.004837 | 4.4e-05 | 0.001193 | 75.3206 | 1.3776 | 2.548 | 101.21 | 20.83 | 17.242 | 65.4773 | 102.313 | 61.111 | 0.004197 | 1.0562 | 0.025882 | 24.5052 | -0.017275 | 0.14033 | -1.3e-05 | -0.001158 | -0.001158 | -5.0395
2019-03-19 15:35 | 1553024106.7636 | -423 | 0.20946 | 2e-05 | 1e-06 | 100.8224 | 100.9106 | 129.5066 | 21.2612 | 0.001799 | 0.001683 | -152.2517 | 0.23651 | 8.4011 | 0.003362 | 0.004941 | 0.004705 | 4.4e-05 | 0.001191 | 75.2901 | 1.4071 | 2.5479 | 101.21 | 20.83 | 17.232 | 65.1059 | 102.7561 | 62.3166 | 0.004195 | 1.0524 | 0.025847 | 24.5603 | -0.007817 | 0.1407 | -1e-06 | -0.00013 | -0.00013 | -45.0815
2019-03-19 15:35 | 1553024107.7643 | -422 | 0.20943 | 2e-05 | 1e-06 | 100.8605 | 100.927 | 129.3543 | 21.2612 | 0.001784 | 0.001657 | -152.255 | 0.23666 | 8.4028 | 0.003401 | 0.004351 | 0.004119 | 4.2e-05 | 0.001197 | 75.2299 | 1.3417 | 2.548 | 101.21 | 20.83 | 17.222 | 64.7532 | 103.6952 | 61.7001 | 0.004193 | 1.0543 | 0.025856 | 24.5233 | -0.005893 | 0.14079 | 0 | -0.000115 | -0.000115 | -50.9079
2019-03-19 15:35 | 1553024108.7645 | -421 | 0.20942 | 2.4e-05 | 1e-06 | 100.8716 | 101.0459 | 129.379 | 21.2612 | 0.001748 | 0.001696 | -152.2785 | 0.2363 | 8.4026 | 0.003308 | 0.004929 | 0.004706 | 4.3e-05 | 0.001201 | 75.147 | 1.3698 | 2.5479 | 101.21 | 20.83 | 17.204 | 63.8944 | 101.4477 | 62.6393 | 0.004188 | 1.0514 | 0.025822 | 24.56 | -0.015435 | 0.14058 | -6e-06 | -0.000595 | -0.000595 | -9.8316
2019-03-19 15:35 | 1553024109.7644 | -420 | 0.20947 | 2.8e-05 | 2e-06 | 100.8428 | 100.9904 | 129.6477 | 21.2612 | 0.001812 | 0.001637 | -152.2965 | 0.2359 | 8.4042 | 0.003264 | 0.005044 | 0.004789 | 4.6e-05 | 0.001204 | 75.1287 | 1.3543 | 2.5478 | 101.21 | 20.83 | 17.197 | 65.4181 | 100.3748 | 61.2203 | 0.004187 | 1.0559 | 0.025904 | 24.5335 | -0.013078 | 0.14034 | -1.2e-05 | -0.001016 | -0.001016 | -5.747
2019-03-19 15:35 | 1553024110.7639 | -419 | 0.20945 | 2.7e-05 | 1e-06 | 100.8269 | 100.9092 | 129.6648 | 21.2612 | 0.00178 | 0.001704 | -152.2984 | 0.23583 | 8.4029 | 0.003238 | 0.005284 | 0.005099 | 4.5e-05 | 0.001205 | 75.1282 | 1.3342 | 2.5479 | 101.21 | 20.83 | 17.188 | 64.6509 | 99.7344 | 62.8266 | 0.004184 | 1.0508 | 0.025843 | 24.594 | -0.007296 | 0.1403 | -1e-05 | -0.000845 | -0.000845 | -6.9034
2019-03-19 15:35 | 1553024111.7638 | -418 | 0.20944 | 3.3e-05 | 2e-06 | 100.8464 | 100.9287 | 129.3588 | 21.2612 | 0.001836 | 0.001674 | -152.3014 | 0.23698 | 8.4038 | 0.00326 | 0.004547 | 0.0043 | 4.3e-05 | 0.001205 | 75.1451 | 1.3399 | 2.5478 | 101.21 | 20.83 | 17.186 | 66.0041 | 100.278 | 62.1088 | 0.004184 | 1.0531 | 0.025841 | 24.5387 | -0.007296 | 0.14098 | -1.8e-05 | -0.001525 | -0.001525 | -3.8446
2019-03-19 15:35 | 1553024112.7638 | -417 | 0.20947 | 3.3e-05 | 1e-06 | 100.8578 | 100.7808 | 129.3895 | 21.2612 | 0.001742 | 0.001644 | -152.2977 | 0.23669 | 8.4037 | 0.003279 | 0.004823 | 0.004647 | 4.2e-05 | 0.001203 | 75.1148 | 1.3398 | 2.5479 | 101.21 | 20.83 | 17.186 | 63.7435 | 100.7412 | 61.3844 | 0.004184 | 1.0553 | 0.025872 | 24.515 | 0.006837 | 0.14081 | -1.8e-05 | -0.001472 | -0.001472 | -3.9792
2019-03-19 15:35 | 1553024113.7637 | -416 | 0.20945 | 3.3e-05 | 2e-06 | 100.8722 | 101.028 | 129.1053 | 21.2612 | 0.001793 | 0.001686 | -152.3101 | 0.23589 | 8.4043 | 0.003208 | 0.004133 | 0.003896 | 4.5e-05 | 0.001196 | 75.0891 | 1.3698 | 2.5479 | 101.21 | 20.83 | 17.189 | 64.9661 | 99.0135 | 62.3925 | 0.004185 | 1.0522 | 0.025804 | 24.525 | -0.013805 | 0.14033 | -1.7e-05 | -0.001456 | -0.001456 | -4.0096
2019-03-19 15:35 | 1553024114.7648 | -415 | 0.20944 | 2.2e-05 | 1e-06 | 100.8107 | 100.9486 | 128.9867 | 21.2612 | 0.001794 | 0.001685 | -152.3225 | 0.23654 | 8.4051 | 0.003211 | 0.004624 | 0.004449 | 4.6e-05 | 0.001194 | 75.0082 | 1.3509 | 2.5479 | 101.21 | 20.83 | 17.194 | 65.0036 | 99.0967 | 62.3689 | 0.004186 | 1.0522 | 0.025793 | 24.5128 | -0.012225 | 0.14072 | -3e-06 | -0.000318 | -0.000318 | -18.4239
2019-03-19 15:35 | 1553024115.7648 | -414 | 0.20945 | 2.1e-05 | 2e-06 | 100.8395 | 100.9054 | 128.7994 | 21.2612 | 0.001745 | 0.001668 | -152.3068 | 0.23548 | 8.4036 | 0.003256 | 0.005167 | 0.004891 | 4.4e-05 | 0.001197 | 75.0453 | 1.3198 | 2.5478 | 101.21 | 20.83 | 17.204 | 63.8154 | 100.1731 | 61.954 | 0.004188 | 1.0535 | 0.025791 | 24.4799 | -0.005844 | 0.14009 | -3e-06 | -0.000326 | -0.000326 | -17.8704
2019-03-19 15:35 | 1553024116.7634 | -413 | 0.2094 | 2.7e-05 | 2e-06 | 100.8449 | 100.8835 | 128.6621 | 21.2612 | 0.001753 | 0.001682 | -152.295 | 0.23703 | 8.404 | 0.003143 | 0.004696 | 0.004472 | 4.1e-05 | 0.001196 | 74.9834 | 1.3408 | 2.5479 | 101.21 | 20.83 | 17.211 | 64.0137 | 97.456 | 62.2916 | 0.00419 | 1.0525 | 0.025764 | 24.4792 | -0.003422 | 0.14101 | -1.1e-05 | -0.000951 | -0.000951 | -6.171
2019-03-19 15:35 | 1553024117.763 | -412 | 0.20948 | 2.4e-05 | 2e-06 | 100.8725 | 101.0315 | 128.6581 | 21.2612 | 0.001756 | 0.001643 | -152.2928 | 0.23635 | 8.4016 | 0.003186 | 0.005009 | 0.004835 | 4.4e-05 | 0.001196 | 74.9528 | 1.3307 | 2.5479 | 101.21 | 20.83 | 17.217 | 64.0809 | 98.4905 | 61.3593 | 0.004191 | 1.0554 | 0.025799 | 24.4447 | -0.014082 | 0.1406 | -7e-06 | -0.000685 | -0.000685 | -8.541
2019-03-19 15:35 | 1553024118.7639 | -411 | 0.20942 | 2.5e-05 | 1e-06 | 100.8434 | 101.0027 | 128.4837 | 21.2612 | 0.0018 | 0.00167 | -152.2781 | 0.23624 | 8.4031 | 0.003209 | 0.004172 | 0.003951 | 4.4e-05 | 0.0012 | 74.8966 | 1.2987 | 2.5478 | 101.21 | 20.82 | 17.228 | 65.1504 | 99.0544 | 62.0055 | 0.004191 | 1.0534 | 0.025757 | 24.4518 | -0.014116 | 0.14054 | -7e-06 | -0.000653 | -0.000653 | -8.9498
2019-03-19 15:35 | 1553024119.763 | -410 | 0.20945 | 2.6e-05 | 1e-06 | 100.8602 | 100.8719 | 128.4433 | 21.2612 | 0.001779 | 0.001676 | -152.2587 | 0.23681 | 8.4037 | 0.003237 | 0.004351 | 0.004177 | 4.6e-05 | 0.001198 | 74.989 | 1.3176 | 2.5477 | 101.21 | 20.82 | 17.233 | 64.6353 | 99.7298 | 62.1502 | 0.004193 | 1.0529 | 0.025747 | 24.4532 | -0.001032 | 0.14088 | -7e-06 | -0.000628 | -0.000628 | -9.3259
2019-03-19 15:35 | 1553024120.7644 | -409 | 0.20936 | 2.4e-05 | 1e-06 | 100.8302 | 100.9503 | 128.4161 | 21.2612 | 0.001851 | 0.001604 | -152.2534 | 0.23614 | 8.4027 | 0.003212 | 0.004352 | 0.004199 | 4.6e-05 | 0.001205 | 74.9384 | 1.3121 | 2.5479 | 101.21 | 20.82 | 17.244 | 66.3719 | 99.1225 | 60.4246 | 0.004195 | 1.0584 | 0.025811 | 24.3876 | -0.010646 | 0.14048 | -6e-06 | -0.000561 | -0.000561 | -10.4096
2019-03-19 15:35 | 1553024121.763 | -408 | 0.2095 | 2.6e-05 | 1e-06 | 100.8077 | 100.8739 | 128.3664 | 21.2612 | 0.001812 | 0.001705 | -152.2514 | 0.2358 | 8.4039 | 0.003207 | 0.004714 | 0.004544 | 4.3e-05 | 0.001203 | 74.9818 | 1.342 | 2.5479 | 101.21 | 20.82 | 17.252 | 65.4213 | 98.9971 | 62.8586 | 0.004197 | 1.0507 | 0.025712 | 24.4717 | -0.005873 | 0.14028 | -8e-06 | -0.000676 | -0.000676 | -8.6329
2019-03-19 15:35 | 1553024122.7632 | -407 | 0.20954 | 2.4e-05 | 1e-06 | 100.8215 | 100.8243 | 128.2466 | 21.2612 | 0.001794 | 0.001678 | -152.2536 | 0.23642 | 8.4038 | 0.00323 | 0.004617 | 0.004431 | 4.3e-05 | 0.001202 | 74.9479 | 1.378 | 2.5478 | 101.21 | 20.82 | 17.26 | 64.9918 | 99.5635 | 62.2136 | 0.004199 | 1.0527 | 0.025725 | 24.4368 | -0.000244 | 0.14065 | -5e-06 | -0.000417 | -0.000417 | -14.0416
2019-03-19 15:35 | 1553024123.7632 | -406 | 0.20953 | 2.4e-05 | 2e-06 | 100.8209 | 100.815 | 128.3272 | 21.2612 | 0.001748 | 0.00172 | -152.2374 | 0.23648 | 8.4029 | 0.003198 | 0.00479 | 0.004496 | 4.3e-05 | 0.001201 | 74.9865 | 1.3715 | 2.5479 | 101.21 | 20.82 | 17.276 | 63.8912 | 98.7791 | 63.2048 | 0.004203 | 1.0496 | 0.025695 | 24.4805 | 0.000524 | 0.14068 | -6e-06 | -0.000526 | -0.000526 | -11.1315
2019-03-19 15:35 | 1553024124.7647 | -405 | 0.20952 | 3e-05 | 1e-06 | 100.8557 | 100.914 | 128.5222 | 21.2612 | 0.001799 | 0.001668 | -152.2385 | 0.23615 | 8.4025 | 0.003127 | 0.004822 | 0.0047 | 4.5e-05 | 0.001204 | 75.0028 | 1.3332 | 2.5479 | 101.21 | 20.82 | 17.284 | 65.1052 | 97.0734 | 61.954 | 0.004205 | 1.0535 | 0.025763 | 24.4535 | -0.005166 | 0.14048 | -1.3e-05 | -0.001079 | -0.001079 | -5.4157
2019-03-19 15:35 | 1553024125.7636 | -404 | 0.20945 | 2.9e-05 | 2e-06 | 100.8551 | 100.8965 | 128.7863 | 21.2612 | 0.001796 | 0.001667 | -152.2454 | 0.23614 | 8.4035 | 0.003211 | 0.0053 | 0.00503 | 4.4e-05 | 0.001208 | 74.9947 | 1.3105 | 2.5478 | 101.21 | 20.82 | 17.292 | 65.0443 | 99.0897 | 61.9282 | 0.004207 | 1.0536 | 0.02579 | 24.4777 | -0.003671 | 0.14048 | -1.3e-05 | -0.001164 | -0.001164 | -5.0189
2019-03-19 15:35 | 1553024126.7647 | -403 | 0.20946 | 2.6e-05 | 1e-06 | 100.8395 | 101.0366 | 128.7363 | 21.2612 | 0.001813 | 0.001716 | -152.2641 | 0.23628 | 8.4027 | 0.0032 | 0.004782 | 0.004607 | 4.5e-05 | 0.001207 | 75.0333 | 1.3488 | 2.5479 | 101.21 | 20.82 | 17.303 | 65.4547 | 98.8339 | 63.1164 | 0.00421 | 1.0499 | 0.02574 | 24.5163 | -0.017464 | 0.14057 | -8e-06 | -0.000642 | -0.000642 | -9.1072
2019-03-19 15:35 | 1553024127.7635 | -402 | 0.20946 | 2.7e-05 | 1e-06 | 100.7813 | 101.0181 | 128.4032 | 21.2612 | 0.001756 | 0.001659 | -152.2805 | 0.2371 | 8.4003 | 0.00315 | 0.004409 | 0.004185 | 4.4e-05 | 0.001213 | 75.0322 | 1.3184 | 2.5478 | 101.21 | 20.82 | 17.309 | 64.0763 | 97.6079 | 61.7399 | 0.004211 | 1.0542 | 0.025759 | 24.4344 | -0.020993 | 0.14105 | -9e-06 | -0.000761 | -0.000761 | -7.7061
2019-03-19 15:35 | 1553024128.7636 | -401 | 0.20946 | 2.4e-05 | 1e-06 | 100.7945 | 100.9667 | 128.6344 | 21.2612 | 0.001786 | 0.001648 | -152.2834 | 0.23575 | 8.4053 | 0.003264 | 0.004889 | 0.004622 | 4.6e-05 | 0.001213 | 75.0331 | 1.3188 | 2.5479 | 101.21 | 20.82 | 17.321 | 64.7943 | 100.3748 | 61.4798 | 0.004214 | 1.055 | 0.025792 | 24.4469 | -0.015271 | 0.14025 | -6e-06 | -0.000524 | -0.000524 | -11.1417
2019-03-19 15:35 | 1553024129.764 | -400 | 0.20946 | 1.7e-05 | 1e-06 | 100.745 | 100.9746 | 128.779 | 21.2612 | 0.001819 | 0.001646 | -152.29 | 0.23673 | 8.4043 | 0.003214 | 0.004944 | 0.004715 | 4.4e-05 | 0.001217 | 75.0362 | 1.3599 | 2.548 | 101.21 | 20.82 | 17.328 | 65.6046 | 99.1711 | 61.4376 | 0.004216 | 1.0552 | 0.025808 | 24.4591 | -0.020364 | 0.14083 | 3e-06 | 0.000156 | 0.000156 | 37.5802
2019-03-19 15:35 | 1553024130.7646 | -399 | 0.20946 | 2.3e-05 | 1e-06 | 100.832 | 101.0822 | 128.7938 | 21.2612 | 0.00175 | 0.001655 | -152.2925 | 0.23638 | 8.4019 | 0.003287 | 0.004638 | 0.004465 | 4.3e-05 | 0.001217 | 75.0748 | 1.354 | 2.548 | 101.21 | 20.82 | 17.335 | 63.9333 | 100.9295 | 61.6462 | 0.004218 | 1.0545 | 0.025802 | 24.4681 | -0.022162 | 0.14062 | -5e-06 | -0.000419 | -0.000419 | -13.9494
2019-03-19 15:36 | 1553024131.7627 | -398 | 0.20947 | 2.2e-05 | 1e-06 | 100.88 | 101.0147 | 128.8778 | 21.2612 | 0.001729 | 0.001677 | -152.3036 | 0.23609 | 8.4013 | 0.003301 | 0.004433 | 0.004286 | 4.6e-05 | 0.00121 | 75.0182 | 1.3165 | 2.5477 | 101.21 | 20.82 | 17.348 | 63.439 | 101.2711 | 62.1799 | 0.004221 | 1.0528 | 0.02579 | 24.4956 | -0.011931 | 0.14045 | -3e-06 | -0.000322 | -0.000322 | -18.1422
2019-03-19 15:36 | 1553024132.7629 | -397 | 0.20947 | 2.1e-05 | 1e-06 | 100.8332 | 100.8294 | 128.9015 | 21.2612 | 0.001727 | 0.001668 | -152.3018 | 0.23605 | 8.4028 | 0.003321 | 0.004639 | 0.004418 | 4.5e-05 | 0.001201 | 75.0426 | 1.3433 | 2.5477 | 101.21 | 20.82 | 17.354 | 63.3704 | 101.7437 | 61.9588 | 0.004222 | 1.0535 | 0.025801 | 24.4898 | 0.000339 | 0.14043 | -1e-06 | -0.000156 | -0.000156 | -37.4108
2019-03-19 15:36 | 1553024133.7628 | -396 | 0.20946 | 2.2e-05 | 1e-06 | 100.956 | 100.8931 | 128.8337 | 21.2612 | 0.001787 | 0.001631 | -152.3208 | 0.237 | 8.4049 | 0.003304 | 0.004942 | 0.004736 | 4.4e-05 | 0.001202 | 75.0326 | 1.3397 | 2.5478 | 101.21 | 20.82 | 17.363 | 64.8256 | 101.3456 | 61.0797 | 0.004224 | 1.0563 | 0.025828 | 24.4512 | 0.005571 | 0.14099 | -3e-06 | -0.000306 | -0.000306 | -19.1445
2019-03-19 15:36 | 1553024134.7638 | -395 | 0.20944 | 2.1e-05 | 1e-06 | 100.8632 | 101.0349 | 128.9757 | 21.2612 | 0.001752 | 0.001678 | -152.3104 | 0.23628 | 8.401 | 0.003383 | 0.004246 | 0.003979 | 4.6e-05 | 0.001206 | 75.0453 | 1.3634 | 2.5481 | 101.21 | 20.82 | 17.368 | 63.9942 | 103.2582 | 62.2136 | 0.004226 | 1.0527 | 0.025798 | 24.5061 | -0.01521 | 0.14057 | -2e-06 | -0.000196 | -0.000196 | -29.7845
2019-03-19 15:36 | 1553024135.7632 | -394 | 0.20944 | 2.2e-05 | 1e-06 | 100.8461 | 100.889 | 128.9217 | 21.2612 | 0.001778 | 0.001664 | -152.3195 | 0.23573 | 8.4029 | 0.003307 | 0.004904 | 0.004697 | 4.4e-05 | 0.001209 | 75.0238 | 1.3706 | 2.5481 | 101.21 | 20.81 | 17.372 | 64.6016 | 101.4107 | 61.8596 | 0.004224 | 1.0538 | 0.025806 | 24.4881 | -0.003801 | 0.14024 | -3e-06 | -0.000338 | -0.000338 | -17.2587
2019-03-19 15:36 | 1553024136.764 | -393 | 0.20943 | 2e-05 | 1e-06 | 100.8803 | 100.925 | 128.7612 | 21.2612 | 0.001797 | 0.001659 | -152.2991 | 0.23679 | 8.4032 | 0.003343 | 0.005278 | 0.005067 | 5e-05 | 0.001208 | 75.0625 | 1.3964 | 2.5482 | 101.21 | 20.82 | 17.383 | 65.0732 | 102.2833 | 61.7541 | 0.004229 | 1.0542 | 0.025794 | 24.469 | -0.003955 | 0.14087 | 0 | -4.3e-05 | -4.3e-05 | -135.8211
2019-03-19 15:36 | 1553024137.7639 | -392 | 0.20943 | 2.5e-05 | 1e-06 | 100.889 | 100.7431 | 128.8337 | 21.2612 | 0.00177 | 0.001704 | -152.3005 | 0.23588 | 8.4039 | 0.003328 | 0.004533 | 0.004369 | 4.3e-05 | 0.00121 | 75.0572 | 1.3308 | 2.5477 | 101.21 | 20.81 | 17.391 | 64.4057 | 101.9323 | 62.8227 | 0.004229 | 1.0508 | 0.025761 | 24.5149 | 0.012949 | 0.14033 | -6e-06 | -0.000537 | -0.000537 | -10.8619
2019-03-19 15:36 | 1553024138.7633 | -391 | 0.20946 | 2.1e-05 | 1e-06 | 100.8995 | 101.0558 | 128.5936 | 21.2612 | 0.001716 | 0.001686 | -152.2884 | 0.23651 | 8.4008 | 0.00327 | 0.004866 | 0.004657 | 4.6e-05 | 0.001207 | 74.9763 | 1.3252 | 2.5478 | 101.21 | 20.81 | 17.413 | 63.1265 | 100.5262 | 62.3915 | 0.004234 | 1.0522 | 0.025753 | 24.4763 | -0.01384 | 0.1407 | -2e-06 | -0.00022 | -0.00022 | -26.6281
2019-03-19 15:36 | 1553024139.7641 | -390 | 0.20943 | 2.9e-05 | 1e-06 | 100.901 | 100.9791 | 128.1908 | 21.2612 | 0.001782 | 0.001619 | -152.2876 | 0.23617 | 8.4014 | 0.003245 | 0.004295 | 0.004125 | 4.6e-05 | 0.001215 | 75.0164 | 1.2941 | 2.548 | 101.21 | 20.81 | 17.426 | 64.7156 | 99.9063 | 60.7857 | 0.004237 | 1.0572 | 0.025775 | 24.3794 | -0.006914 | 0.1405 | -1.1e-05 | -0.000937 | -0.000937 | -6.2347
2019-03-19 15:36 | 1553024140.7644 | -389 | 0.20946 | 2.8e-05 | 1e-06 | 100.9566 | 101.0222 | 128.0028 | 21.2612 | 0.001726 | 0.001603 | -152.2632 | 0.23721 | 8.4026 | 0.003294 | 0.004662 | 0.004382 | 4.6e-05 | 0.001217 | 75.0685 | 1.3296 | 2.5481 | 101.21 | 20.81 | 17.441 | 63.3663 | 101.1085 | 60.3894 | 0.004241 | 1.0585 | 0.025771 | 24.347 | -0.005815 | 0.14112 | -1.1e-05 | -0.00089 | -0.00089 | -6.5978
2019-03-19 15:36 | 1553024141.7645 | -388 | 0.20945 | 2.4e-05 | 1e-06 | 100.9677 | 101.1644 | 128.1072 | 21.2612 | 0.001721 | 0.001666 | -152.2532 | 0.23669 | 8.4015 | 0.00332 | 0.004316 | 0.004055 | 4.5e-05 | 0.001221 | 75.0843 | 1.304 | 2.548 | 101.21 | 20.81 | 17.477 | 63.2382 | 101.7294 | 61.9244 | 0.004249 | 1.0536 | 0.025722 | 24.4129 | -0.017407 | 0.14081 | -5e-06 | -0.000514 | -0.000514 | -11.4024
2019-03-19 15:36 | 1553024142.7643 | -387 | 0.20946 | 2.3e-05 | 1e-06 | 100.9851 | 101.1688 | 128.5677 | 21.2612 | 0.001712 | 0.001689 | -152.2662 | 0.23656 | 8.3988 | 0.003333 | 0.00514 | 0.004882 | 4.5e-05 | 0.001221 | 75.076 | 1.3459 | 2.548 | 101.21 | 20.81 | 17.496 | 63.0141 | 102.0531 | 62.4719 | 0.004254 | 1.0519 | 0.025747 | 24.4768 | -0.016259 | 0.14073 | -5e-06 | -0.000446 | -0.000446 | -13.1216
2019-03-19 15:36 | 1553024143.7638 | -386 | 0.20944 | 3.1e-05 | 1e-06 | 100.9359 | 101.0592 | 128.7316 | 21.2612 | 0.001774 | 0.001653 | -152.2761 | 0.23675 | 8.4034 | 0.003351 | 0.004854 | 0.004625 | 4.3e-05 | 0.001223 | 75.0353 | 1.3611 | 2.5482 | 101.21 | 20.81 | 17.517 | 64.522 | 102.4859 | 61.5898 | 0.004259 | 1.0547 | 0.025798 | 24.4601 | -0.010925 | 0.14084 | -1.4e-05 | -0.001159 | -0.001159 | -5.0568
2019-03-19 15:36 | 1553024144.7633 | -385 | 0.20947 | 2.1e-05 | 1e-06 | 100.9515 | 101.2123 | 128.2314 | 21.2612 | 0.001747 | 0.001655 | -152.2625 | 0.23709 | 8.4019 | 0.003257 | 0.004635 | 0.004418 | 4.7e-05 | 0.001226 | 75.0507 | 1.3315 | 2.5481 | 101.21 | 20.81 | 17.555 | 63.8725 | 100.2067 | 61.6392 | 0.004268 | 1.0545 | 0.025746 | 24.4144 | -0.023082 | 0.14105 | -1e-06 | -0.000175 | -0.000175 | -33.4985
2019-03-19 15:36 | 1553024145.7634 | -384 | 0.20944 | 1.7e-05 | 1e-06 | 101.0238 | 101.1852 | 127.9753 | 21.2612 | 0.001777 | 0.001613 | -152.2591 | 0.23647 | 8.4022 | 0.003313 | 0.004084 | 0.003889 | 4.4e-05 | 0.00123 | 75.0618 | 1.3338 | 2.5482 | 101.21 | 20.81 | 17.574 | 64.5845 | 101.567 | 60.6482 | 0.004273 | 1.0577 | 0.025758 | 24.3539 | -0.014284 | 0.14067 | 4e-06 | 0.000253 | 0.000253 | 23.1329
2019-03-19 15:36 | 1553024146.7638 | -383 | 0.20944 | 2.2e-05 | 1e-06 | 100.9356 | 101.3024 | 128.2342 | 21.2612 | 0.001793 | 0.001608 | -152.2707 | 0.23725 | 8.4025 | 0.003351 | 0.004661 | 0.004341 | 4.5e-05 | 0.001221 | 75.0278 | 1.3209 | 2.548 | 101.21 | 20.81 | 17.61 | 64.963 | 102.4828 | 60.527 | 0.004282 | 1.058 | 0.025789 | 24.3741 | -0.032448 | 0.14114 | -3e-06 | -0.000288 | -0.000288 | -20.3902
2019-03-19 15:36 | 1553024147.7635 | -382 | 0.20946 | 2.1e-05 | 2e-06 | 100.9701 | 101.0575 | 128.333 | 21.2612 | 0.001782 | 0.001676 | -152.2603 | 0.23701 | 8.4026 | 0.003276 | 0.005308 | 0.005027 | 4.5e-05 | 0.001216 | 75.0135 | 1.3342 | 2.5482 | 101.21 | 20.81 | 17.626 | 64.6952 | 100.6728 | 62.1649 | 0.004286 | 1.0529 | 0.025736 | 24.4432 | -0.007742 | 0.141 | -2e-06 | -0.000284 | -0.000284 | -20.6377
2019-03-19 15:36 | 1553024148.7625 | -381 | 0.20946 | 2.2e-05 | 2e-06 | 101.0505 | 101.1613 | 128.4384 | 21.2612 | 0.001738 | 0.001633 | -152.2522 | 0.23688 | 8.4007 | 0.003338 | 0.005107 | 0.004905 | 4.3e-05 | 0.001203 | 75.036 | 1.2996 | 2.5482 | 101.21 | 20.81 | 17.64 | 63.6561 | 102.1678 | 61.1249 | 0.004289 | 1.0562 | 0.025786 | 24.4153 | -0.0098 | 0.14092 | -4e-06 | -0.000433 | -0.000433 | -13.5475
2019-03-19 15:36 | 1553024149.7636 | -380 | 0.20944 | 2.4e-05 | 1e-06 | 101.0334 | 101.1681 | 128.6769 | 21.2612 | 0.001764 | 0.001647 | -152.2485 | 0.23651 | 8.4046 | 0.003348 | 0.004873 | 0.004676 | 4.5e-05 | 0.001192 | 75.0067 | 1.3077 | 2.5479 | 101.21 | 20.81 | 17.66 | 64.284 | 102.4191 | 61.4492 | 0.004294 | 1.0551 | 0.025798 | 24.4498 | -0.01192 | 0.1407 | -6e-06 | -0.000517 | -0.000517 | -11.3207
2019-03-19 15:36 | 1553024150.7639 | -379 | 0.20945 | 2.5e-05 | 1e-06 | 101.0127 | 101.2243 | 128.6706 | 21.2612 | 0.001805 | 0.001654 | -152.2664 | 0.23699 | 8.4026 | 0.003386 | 0.004554 | 0.004365 | 4.6e-05 | 0.001209 | 75.0543 | 1.3533 | 2.5481 | 101.21 | 20.81 | 17.664 | 65.2676 | 103.3296 | 61.6274 | 0.004295 | 1.0546 | 0.02579 | 24.4557 | -0.018718 | 0.14099 | -7e-06 | -0.000617 | -0.000617 | -9.5101
2019-03-19 15:36 | 1553024151.7633 | -378 | 0.20946 | 2.7e-05 | 1e-06 | 101.0256 | 101.1791 | 128.583 | 21.2612 | 0.001785 | 0.001671 | -152.2703 | 0.23679 | 8.4008 | 0.00331 | 0.004956 | 0.004678 | 4.8e-05 | 0.00121 | 74.9991 | 1.3394 | 2.5481 | 101.21 | 20.81 | 17.664 | 64.7671 | 101.4901 | 62.0416 | 0.004295 | 1.0533 | 0.025765 | 24.4625 | -0.01358 | 0.14087 | -1e-05 | -0.000837 | -0.000837 | -7.0005
2019-03-19 15:36 | 1553024152.7629 | -377 | 0.20946 | 3e-05 | 1e-06 | 101.0271 | 101.1339 | 128.2542 | 21.2612 | 0.001788 | 0.001665 | -152.2695 | 0.23712 | 8.3977 | 0.003428 | 0.005126 | 0.004845 | 4.5e-05 | 0.001213 | 75.0116 | 1.3897 | 2.5479 | 101.21 | 20.81 | 17.661 | 64.8429 | 104.3358 | 61.8923 | 0.004294 | 1.0537 | 0.025738 | 24.4258 | -0.009449 | 0.14106 | -1.3e-05 | -0.001111 | -0.001111 | -5.2837
2019-03-19 15:36 | 1553024153.7637 | -376 | 0.20946 | 2.5e-05 | 1e-06 | 101.0202 | 101.2164 | 128.1711 | 21.2612 | 0.001741 | 0.001645 | -152.2611 | 0.23595 | 8.4033 | 0.003467 | 0.004028 | 0.003838 | 4.7e-05 | 0.001209 | 75.1035 | 1.4 | 2.548 | 101.21 | 20.81 | 17.654 | 63.7232 | 105.3001 | 61.4024 | 0.004292 | 1.0553 | 0.025749 | 24.4 | -0.017358 | 0.14037 | -6e-06 | -0.000524 | -0.000524 | -11.1358
2019-03-19 15:36 | 1553024154.7639 | -375 | 0.20946 | 2.5e-05 | 1e-06 | 101.1027 | 101.274 | 128.2957 | 21.2612 | 0.001771 | 0.001666 | -152.2769 | 0.23694 | 8.4026 | 0.003425 | 0.004368 | 0.004114 | 4.5e-05 | 0.0012 | 75.0297 | 1.384 | 2.5482 | 101.21 | 20.81 | 17.637 | 64.4298 | 104.2629 | 61.9094 | 0.004288 | 1.0537 | 0.025742 | 24.4304 | -0.015137 | 0.14096 | -7e-06 | -0.000588 | -0.000588 | -9.9653
2019-03-19 15:36 | 1553024155.764 | -374 | 0.20946 | 2.2e-05 | 1e-06 | 100.9971 | 101.2061 | 128.5133 | 21.2612 | 0.001797 | 0.001695 | -152.2717 | 0.23639 | 8.4026 | 0.003425 | 0.004324 | 0.004049 | 4.6e-05 | 0.001203 | 75.0191 | 1.3697 | 2.5482 | 101.21 | 20.81 | 17.628 | 65.0645 | 104.2801 | 62.6125 | 0.004286 | 1.0515 | 0.025737 | 24.4767 | -0.018496 | 0.14063 | -3e-06 | -0.000263 | -0.000263 | -22.2077
2019-03-19 15:36 | 1553024156.763 | -373 | 0.20947 | 1.9e-05 | 2e-06 | 101.0364 | 101.351 | 128.7581 | 21.2612 | 0.001775 | 0.001632 | -152.2822 | 0.23658 | 8.4027 | 0.003374 | 0.004584 | 0.004327 | 4.7e-05 | 0.0012 | 74.9876 | 1.3522 | 2.5482 | 101.21 | 20.81 | 17.617 | 64.5414 | 103.0453 | 61.1061 | 0.004283 | 1.0562 | 0.025819 | 24.445 | -0.02781 | 0.14074 | 0 | -0.000107 | -0.000107 | -54.6078
2019-03-19 15:36 | 1553024157.7641 | -372 | 0.20947 | 1.8e-05 | 1e-06 | 101.0037 | 101.1486 | 129.0212 | 21.2612 | 0.001739 | 0.00159 | -152.2696 | 0.23664 | 8.4035 | 0.003386 | 0.0053 | 0.005022 | 4.3e-05 | 0.001204 | 75.0294 | 1.3475 | 2.5478 | 101.21 | 20.8 | 17.598 | 63.6725 | 103.3282 | 60.0773 | 0.004276 | 1.0595 | 0.025885 | 24.4323 | -0.012825 | 0.14078 | 3e-06 | 0.000168 | 0.000168 | 34.8962
2019-03-19 15:36 | 1553024158.7641 | -371 | 0.20947 | 2.5e-05 | 1e-06 | 100.9719 | 101.1188 | 128.8076 | 21.2612 | 0.001733 | 0.001674 | -152.2848 | 0.23613 | 8.4012 | 0.003359 | 0.004643 | 0.004432 | 4.4e-05 | 0.001211 | 75.0597 | 1.3454 | 2.5478 | 101.21 | 20.8 | 17.588 | 63.5319 | 102.6634 | 62.0987 | 0.004274 | 1.0531 | 0.025786 | 24.486 | -0.013007 | 0.14047 | -7e-06 | -0.000612 | -0.000612 | -9.5535
2019-03-19 15:36 | 1553024159.7635 | -370 | 0.20946 | 1.8e-05 | 1e-06 | 101.0313 | 101.3219 | 128.5586 | 21.2612 | 0.001822 | 0.001657 | -152.2907 | 0.23578 | 8.4041 | 0.00338 | 0.004673 | 0.004598 | 4.7e-05 | 0.001208 | 75.0765 | 1.3652 | 2.5479 | 101.21 | 20.8 | 17.571 | 65.6601 | 103.186 | 61.6869 | 0.00427 | 1.0544 | 0.025777 | 24.4472 | -0.025692 | 0.14027 | 2e-06 | 6.3e-05 | 6.3e-05 | 92.9397
2019-03-19 15:36 | 1553024160.7643 | -369 | 0.20947 | 2.2e-05 | 1e-06 | 100.986 | 101.1558 | 128.7065 | 21.2612 | 0.001773 | 0.001663 | -152.2985 | 0.23589 | 8.4023 | 0.003254 | 0.004893 | 0.004617 | 4.8e-05 | 0.001211 | 75.0134 | 1.3632 | 2.5481 | 101.21 | 20.8 | 17.564 | 64.4774 | 100.1354 | 61.8321 | 0.004268 | 1.0539 | 0.025786 | 24.4666 | -0.015029 | 0.14033 | -2e-06 | -0.000191 | -0.000191 | -30.543
2019-03-19 15:36 | 1553024161.7636 | -368 | 0.20947 | 2.1e-05 | 1e-06 | 100.98 | 101.0383 | 128.8055 | 21.2612 | 0.001745 | 0.001639 | -152.3058 | 0.23577 | 8.4026 | 0.003259 | 0.004842 | 0.004547 | 4.7e-05 | 0.001211 | 75.0297 | 1.3596 | 2.5479 | 101.21 | 20.8 | 17.556 | 63.8241 | 100.2647 | 61.2719 | 0.004266 | 1.0557 | 0.025817 | 24.4555 | -0.005167 | 0.14026 | -1e-06 | -0.000101 | -0.000101 | -57.4983
2019-03-19 15:36 | 1553024162.7624 | -367 | 0.20946 | 2.1e-05 | 1e-06 | 101.0274 | 101.1109 | 128.9851 | 21.2612 | 0.001731 | 0.00162 | -152.2994 | 0.23735 | 8.4013 | 0.003261 | 0.004764 | 0.004632 | 4.8e-05 | 0.001204 | 75.045 | 1.3401 | 2.5479 | 101.21 | 20.8 | 17.554 | 63.4874 | 100.3017 | 60.8084 | 0.004266 | 1.0572 | 0.025853 | 24.4556 | -0.007393 | 0.1412 | -1e-06 | -0.000154 | -0.000154 | -38.0973
2019-03-19 15:36 | 1553024163.7633 | -366 | 0.20948 | 2.2e-05 | 1e-06 | 101.0337 | 101.0229 | 128.8426 | 21.2612 | 0.001797 | 0.001662 | -152.2828 | 0.23599 | 8.4045 | 0.00331 | 0.004038 | 0.003842 | 4.7e-05 | 0.001201 | 75.0127 | 1.3341 | 2.5479 | 101.21 | 20.8 | 17.553 | 65.0575 | 101.4877 | 61.8112 | 0.004265 | 1.054 | 0.0258 | 24.4788 | 0.000955 | 0.14039 | -3e-06 | -0.000289 | -0.000289 | -20.2369
2019-03-19 15:36 | 1553024164.7638 | -365 | 0.20948 | 1.8e-05 | 1e-06 | 101.0337 | 101.3024 | 128.9149 | 21.2612 | 0.001741 | 0.001668 | -152.293 | 0.23726 | 8.4023 | 0.003239 | 0.004327 | 0.004147 | 4.7e-05 | 0.001196 | 75.1052 | 1.3568 | 2.5479 | 101.21 | 20.8 | 17.556 | 63.7194 | 99.769 | 61.9713 | 0.004266 | 1.0535 | 0.025801 | 24.4915 | -0.023756 | 0.14115 | 2e-06 | 5.5e-05 | 5.5e-05 | 106.621
2019-03-19 15:36 | 1553024165.7625 | -364 | 0.20945 | 2e-05 | 1e-06 | 101.0343 | 101.1616 | 129.0902 | 21.2612 | 0.001738 | 0.001648 | -152.3046 | 0.23653 | 8.4039 | 0.00319 | 0.004407 | 0.004108 | 4.7e-05 | 0.001202 | 75.0893 | 1.3015 | 2.5479 | 101.21 | 20.8 | 17.559 | 63.6561 | 98.591 | 61.4822 | 0.004267 | 1.055 | 0.025838 | 24.4903 | -0.011265 | 0.14071 | -1e-06 | -0.000195 | -0.000195 | -30.0916
2019-03-19 15:36 | 1553024166.7629 | -363 | 0.20944 | 2.2e-05 | 1e-06 | 101.0247 | 101.0969 | 128.6654 | 21.2612 | 0.001803 | 0.001668 | -152.3059 | 0.23713 | 8.4008 | 0.003174 | 0.00465 | 0.004398 | 5e-05 | 0.001201 | 75.0363 | 1.2854 | 2.548 | 101.21 | 20.8 | 17.562 | 65.2113 | 98.203 | 61.9612 | 0.004267 | 1.0535 | 0.025777 | 24.4674 | -0.00639 | 0.14107 | -2e-06 | -0.000225 | -0.000225 | -26.0504
2019-03-19 15:36 | 1553024167.7636 | -362 | 0.20946 | 2.4e-05 | 1e-06 | 101.0418 | 101.1705 | 128.4931 | 21.2612 | 0.001749 | 0.001633 | -152.3018 | 0.23663 | 8.4016 | 0.003175 | 0.004035 | 0.003811 | 4.7e-05 | 0.001202 | 75.008 | 1.3139 | 2.548 | 101.21 | 20.79 | 17.57 | 63.9083 | 98.2208 | 61.1273 | 0.004267 | 1.0561 | 0.025792 | 24.4206 | -0.011388 | 0.14077 | -5e-06 | -0.000475 | -0.000475 | -12.3186
2019-03-19 15:36 | 1553024168.7636 | -361 | 0.20944 | 2.5e-05 | 1e-06 | 101.0478 | 101.2637 | 128.6856 | 21.2612 | 0.001818 | 0.001666 | -152.3097 | 0.23608 | 8.4014 | 0.0032 | 0.004415 | 0.004187 | 4.8e-05 | 0.001197 | 75.0518 | 1.3307 | 2.5481 | 101.21 | 20.79 | 17.575 | 65.5743 | 98.8277 | 61.9073 | 0.004268 | 1.0537 | 0.025781 | 24.4674 | -0.019089 | 0.14045 | -6e-06 | -0.000501 | -0.000501 | -11.6594
2019-03-19 15:36 | 1553024169.7637 | -360 | 0.20944 | 2.4e-05 | 1e-06 | 100.9992 | 101.0044 | 128.621 | 21.2612 | 0.00182 | 0.001664 | -152.3035 | 0.23683 | 8.401 | 0.003277 | 0.004696 | 0.004463 | 5e-05 | 0.001194 | 75.066 | 1.3437 | 2.5481 | 101.21 | 20.79 | 17.575 | 65.6094 | 100.6917 | 61.8697 | 0.004268 | 1.0538 | 0.025776 | 24.4599 | -0.000464 | 0.14089 | -5e-06 | -0.000516 | -0.000516 | -11.357
2019-03-19 15:36 | 1553024170.7637 | -359 | 0.20945 | 3e-05 | 1e-06 | 101.0223 | 101.1246 | 128.5917 | 21.2612 | 0.001761 | 0.001697 | -152.2969 | 0.23623 | 8.4017 | 0.003315 | 0.004632 | 0.004455 | 4.7e-05 | 0.001199 | 75.0686 | 1.3374 | 2.5481 | 101.21 | 20.79 | 17.591 | 64.2074 | 101.6196 | 62.6612 | 0.004272 | 1.0513 | 0.025743 | 24.4859 | -0.009057 | 0.14053 | -1.2e-05 | -0.000977 | -0.000977 | -5.9825
2019-03-19 15:36 | 1553024171.7628 | -358 | 0.20945 | 2.5e-05 | 1e-06 | 100.9812 | 101.2236 | 128.5846 | 21.2612 | 0.001787 | 0.001622 | -152.2767 | 0.23738 | 8.4036 | 0.00325 | 0.005125 | 0.004923 | 4.9e-05 | 0.001206 | 75.0618 | 1.3358 | 2.5479 | 101.21 | 20.79 | 17.596 | 64.828 | 100.042 | 60.8623 | 0.004273 | 1.057 | 0.025811 | 24.4196 | -0.021448 | 0.14122 | -6e-06 | -0.000552 | -0.000552 | -10.6475
2019-03-19 15:36 | 1553024172.7643 | -357 | 0.20944 | 2.5e-05 | 1e-06 | 101.0574 | 101.2877 | 128.1814 | 21.2612 | 0.001764 | 0.001657 | -152.2725 | 0.23616 | 8.4027 | 0.003247 | 0.004556 | 0.004284 | 4.7e-05 | 0.001206 | 75.0331 | 1.3693 | 2.5479 | 101.21 | 20.79 | 17.6 | 64.2777 | 99.969 | 61.69 | 0.004274 | 1.0544 | 0.025739 | 24.4115 | -0.020356 | 0.14049 | -7e-06 | -0.000608 | -0.000608 | -9.6193
2019-03-19 15:36 | 1553024173.7629 | -356 | 0.20944 | 2.4e-05 | 1e-06 | 100.9905 | 101.1541 | 128.2652 | 21.2612 | 0.001791 | 0.001689 | -152.272 | 0.23594 | 8.4038 | 0.003228 | 0.004776 | 0.004511 | 4.6e-05 | 0.001205 | 75.0362 | 1.3191 | 2.548 | 101.21 | 20.79 | 17.605 | 64.9226 | 99.5055 | 62.4768 | 0.004275 | 1.0519 | 0.025717 | 24.4481 | -0.014479 | 0.14036 | -5e-06 | -0.000485 | -0.000485 | -12.0326
2019-03-19 15:36 | 1553024174.764 | -355 | 0.20946 | 2.6e-05 | 1e-06 | 100.9941 | 101.1986 | 127.9922 | 21.2612 | 0.001799 | 0.001644 | -152.2679 | 0.23673 | 8.4037 | 0.003279 | 0.004917 | 0.004739 | 4.6e-05 | 0.001211 | 74.9795 | 1.3175 | 2.5478 | 101.21 | 20.79 | 17.613 | 65.1129 | 100.735 | 61.3875 | 0.004277 | 1.0553 | 0.025731 | 24.3824 | -0.018096 | 0.14083 | -7e-06 | -0.000652 | -0.000652 | -8.9877
2019-03-19 15:36 | 1553024175.7631 | -354 | 0.20943 | 1.5e-05 | 1e-06 | 100.9926 | 101.0739 | 127.6953 | 21.2612 | 0.001778 | 0.001706 | -152.251 | 0.23675 | 8.4024 | 0.00325 | 0.004569 | 0.004424 | 5e-05 | 0.001212 | 75.0055 | 1.3547 | 2.548 | 101.21 | 20.79 | 17.616 | 64.6016 | 100.0294 | 62.8875 | 0.004278 | 1.0506 | 0.025644 | 24.4087 | -0.007203 | 0.14084 | 6e-06 | 0.000395 | 0.000395 | 14.8417
2019-03-19 15:36 | 1553024176.7631 | -353 | 0.20944 | 2.7e-05 | 1e-06 | 100.9509 | 101.0417 | 127.841 | 21.2612 | 0.001735 | 0.001666 | -152.2587 | 0.23673 | 8.4067 | 0.003203 | 0.004606 | 0.004398 | 4.5e-05 | 0.001211 | 74.9764 | 1.3225 | 2.5479 | 101.21 | 20.79 | 17.619 | 63.5718 | 98.8929 | 61.9111 | 0.004279 | 1.0537 | 0.025696 | 24.3871 | -0.008049 | 0.14083 | -9e-06 | -0.000764 | -0.000764 | -7.6675
2019-03-19 15:36 | 1553024177.7644 | -352 | 0.20943 | 2.2e-05 | 1e-06 | 101.0364 | 101.0215 | 127.7241 | 21.2612 | 0.001764 | 0.001637 | -152.2738 | 0.23735 | 8.4013 | 0.003263 | 0.004411 | 0.004154 | 4.9e-05 | 0.001214 | 74.9741 | 1.3502 | 2.5481 | 101.21 | 20.79 | 17.62 | 64.2628 | 100.355 | 61.2061 | 0.004279 | 1.0559 | 0.025711 | 24.3503 | 0.001315 | 0.1412 | -2e-06 | -0.00022 | -0.00022 | -26.7585
2019-03-19 15:36 | 1553024178.7633 | -351 | 0.20944 | 2.4e-05 | 1e-06 | 101.0463 | 101.3692 | 127.7375 | 21.2612 | 0.001733 | 0.00163 | -152.2693 | 0.23661 | 8.4013 | 0.003208 | 0.004565 | 0.004301 | 4.9e-05 | 0.001214 | 74.9744 | 1.3275 | 2.548 | 101.21 | 20.79 | 17.622 | 63.528 | 99.0089 | 61.0397 | 0.004279 | 1.0564 | 0.025719 | 24.3455 | -0.028535 | 0.14076 | -5e-06 | -0.00045 | -0.00045 | -13.0119
2019-03-19 15:36 | 1553024179.7622 | -350 | 0.20944 | 2.9e-05 | 1e-06 | 100.9953 | 101.0589 | 127.6655 | 21.2612 | 0.001793 | 0.00168 | -152.2785 | 0.23705 | 8.4035 | 0.003206 | 0.005009 | 0.004756 | 4.8e-05 | 0.001209 | 74.9967 | 1.331 | 2.548 | 101.21 | 20.79 | 17.623 | 64.9599 | 98.9688 | 62.262 | 0.00428 | 1.0526 | 0.025665 | 24.3831 | -0.00563 | 0.14102 | -1.1e-05 | -0.000916 | -0.000916 | -6.4043
2019-03-19 15:36 | 1553024180.7626 | -349 | 0.20943 | 2.4e-05 | 1e-06 | 101.0178 | 101.1935 | 127.5012 | 21.2612 | 0.001745 | 0.001671 | -152.2741 | 0.23675 | 8.4011 | 0.00311 | 0.004941 | 0.004717 | 4.9e-05 | 0.001212 | 75.0322 | 1.3286 | 2.548 | 101.21 | 20.79 | 17.623 | 63.8137 | 96.6458 | 62.0392 | 0.00428 | 1.0533 | 0.025657 | 24.3593 | -0.015542 | 0.14084 | -4e-06 | -0.000382 | -0.000382 | -15.3359
2019-03-19 15:36 | 1553024181.7626 | -348 | 0.20942 | 2.4e-05 | 1e-06 | 100.986 | 100.9195 | 127.3878 | 21.2612 | 0.001812 | 0.00167 | -152.2814 | 0.23721 | 8.4015 | 0.003244 | 0.004175 | 0.003936 | 4.6e-05 | 0.001214 | 75.0007 | 1.2976 | 2.548 | 101.21 | 20.79 | 17.622 | 65.4321 | 99.9015 | 62.0048 | 0.004279 | 1.0534 | 0.025647 | 24.3472 | 0.005892 | 0.14112 | -4e-06 | -0.000399 | -0.000399 | -14.7266
2019-03-19 15:36 | 1553024182.7635 | -347 | 0.20943 | 2.4e-05 | 1e-06 | 100.9662 | 101.0407 | 127.657 | 21.2612 | 0.00178 | 0.00167 | -152.2874 | 0.23625 | 8.401 | 0.003321 | 0.00465 | 0.004387 | 4.7e-05 | 0.001214 | 74.9962 | 1.3169 | 2.548 | 101.21 | 20.79 | 17.621 | 64.6461 | 101.742 | 62.0204 | 0.004279 | 1.0533 | 0.025673 | 24.3735 | -0.006602 | 0.14054 | -5e-06 | -0.000401 | -0.000401 | -14.5966
2019-03-19 15:36 | 1553024183.7624 | -346 | 0.20941 | 2.4e-05 | 1e-06 | 100.9539 | 101.0404 | 127.5789 | 21.2612 | 0.001782 | 0.001667 | -152.297 | 0.23591 | 8.404 | 0.003319 | 0.004491 | 0.004305 | 4.5e-05 | 0.001211 | 74.9668 | 1.3073 | 2.548 | 101.21 | 20.79 | 17.62 | 64.7031 | 101.7123 | 61.947 | 0.004279 | 1.0536 | 0.025668 | 24.3634 | -0.007662 | 0.14034 | -5e-06 | -0.00045 | -0.00045 | -12.9709
2019-03-19 15:36 | 1553024184.7636 | -345 | 0.20944 | 2.2e-05 | 1e-06 | 100.9647 | 101.2061 | 127.8809 | 21.2612 | 0.001852 | 0.00166 | -152.2811 | 0.23636 | 8.401 | 0.003342 | 0.004924 | 0.004649 | 4.8e-05 | 0.001207 | 74.9655 | 1.3695 | 2.548 | 101.21 | 20.79 | 17.619 | 66.3904 | 102.2557 | 61.7603 | 0.004279 | 1.0541 | 0.025706 | 24.3854 | -0.021367 | 0.14061 | -1e-06 | -0.00015 | -0.00015 | -38.92
2019-03-19 15:36 | 1553024185.7626 | -344 | 0.20943 | 2.2e-05 | 1e-06 | 100.9707 | 100.9387 | 127.9732 | 21.2612 | 0.001738 | 0.001704 | -152.2949 | 0.23666 | 8.4038 | 0.00328 | 0.00516 | 0.004977 | 4.7e-05 | 0.001202 | 75.0199 | 1.3303 | 2.548 | 101.21 | 20.79 | 17.619 | 63.6506 | 100.7633 | 62.8275 | 0.004279 | 1.0508 | 0.025674 | 24.433 | 0.002836 | 0.14079 | -2e-06 | -0.000254 | -0.000254 | -23.0652
2019-03-19 15:36 | 1553024186.763 | -343 | 0.20943 | 2e-05 | 1e-06 | 100.9161 | 101.2575 | 128.1082 | 21.2612 | 0.001748 | 0.001658 | -152.2865 | 0.23668 | 8.4012 | 0.003296 | 0.004341 | 0.004197 | 4.8e-05 | 0.001203 | 74.9796 | 1.3026 | 2.5481 | 101.21 | 20.79 | 17.619 | 63.8777 | 101.1487 | 61.7274 | 0.004279 | 1.0543 | 0.02573 | 24.4059 | -0.030214 | 0.1408 | 0 | -7.2e-05 | -7.2e-05 | -80.8671
2019-03-19 15:36 | 1553024187.7639 | -342 | 0.20943 | 2.5e-05 | 1e-06 | 100.8908 | 101.0979 | 128.0746 | 21.2612 | 0.001754 | 0.001659 | -152.2653 | 0.23744 | 8.4026 | 0.003318 | 0.004952 | 0.004676 | 4.7e-05 | 0.001207 | 75.0563 | 1.3264 | 2.548 | 101.21 | 20.79 | 17.622 | 64.0255 | 101.6856 | 61.7471 | 0.004279 | 1.0542 | 0.025726 | 24.4034 | -0.018339 | 0.14125 | -6e-06 | -0.000561 | -0.000561 | -10.4749
2019-03-19 15:36 | 1553024188.7632 | -341 | 0.20945 | 2.4e-05 | 1e-06 | 100.9137 | 101.0024 | 128.0366 | 21.2612 | 0.001733 | 0.001701 | -152.2693 | 0.23654 | 8.4028 | 0.00329 | 0.004499 | 0.004312 | 5e-05 | 0.0012 | 75.0693 | 1.3509 | 2.5479 | 101.21 | 20.79 | 17.624 | 63.5304 | 101.0064 | 62.7479 | 0.00428 | 1.0511 | 0.025684 | 24.4362 | -0.007859 | 0.14072 | -5e-06 | -0.000427 | -0.000427 | -13.7194
2019-03-19 15:36 | 1553024189.7642 | -340 | 0.20946 | 3e-05 | 1e-06 | 100.8746 | 101.0428 | 128.1079 | 21.2612 | 0.001831 | 0.001689 | -152.2661 | 0.2367 | 8.4013 | 0.003315 | 0.004812 | 0.004593 | 4.5e-05 | 0.001195 | 74.9613 | 1.3372 | 2.5479 | 101.21 | 20.79 | 17.626 | 65.8872 | 101.6181 | 62.4573 | 0.00428 | 1.052 | 0.025702 | 24.4324 | -0.014896 | 0.14081 | -1.2e-05 | -0.000959 | -0.000959 | -6.1066
2019-03-19 15:36 | 1553024190.7641 | -339 | 0.20947 | 2e-05 | 1e-06 | 100.8506 | 100.7886 | 128.4957 | 21.2612 | 0.001782 | 0.00167 | -152.259 | 0.23677 | 8.4012 | 0.003288 | 0.00486 | 0.004682 | 4.7e-05 | 0.001184 | 74.972 | 1.3744 | 2.5481 | 101.21 | 20.79 | 17.628 | 64.6962 | 100.9585 | 62.0159 | 0.004281 | 1.0533 | 0.025758 | 24.4533 | 0.005499 | 0.14085 | -1e-06 | -0.000107 | -0.000107 | -54.8804
2019-03-19 15:36 | 1553024191.7629 | -338 | 0.20944 | 2.3e-05 | 1e-06 | 100.8527 | 101.0897 | 128.6696 | 21.2612 | 0.001792 | 0.001703 | -152.2539 | 0.23753 | 8.3999 | 0.003332 | 0.004419 | 0.004194 | 4.9e-05 | 0.00118 | 75.0028 | 1.3609 | 2.5482 | 101.21 | 20.79 | 17.629 | 64.9396 | 102.0107 | 62.8112 | 0.004281 | 1.0509 | 0.025745 | 24.4988 | -0.020992 | 0.1413 | -4e-06 | -0.000407 | -0.000407 | -14.4299
2019-03-19 15:37 | 1553024192.7639 | -337 | 0.20946 | 2.7e-05 | 1e-06 | 100.8569 | 100.9027 | 128.5856 | 21.2612 | 0.001762 | 0.00166 | -152.2387 | 0.23704 | 8.4016 | 0.003453 | 0.004729 | 0.004434 | 4.7e-05 | 0.001175 | 75.0243 | 1.3541 | 2.5479 | 101.21 | 20.79 | 17.632 | 64.2137 | 104.9541 | 61.7712 | 0.004282 | 1.0541 | 0.025776 | 24.4529 | -0.004058 | 0.14101 | -1e-05 | -0.000816 | -0.000816 | -7.1911
2019-03-19 15:37 | 1553024193.7636 | -336 | 0.20946 | 2.4e-05 | 1e-06 | 100.7717 | 101.062 | 128.6835 | 21.2612 | 0.001754 | 0.001619 | -152.2359 | 0.23682 | 8.4047 | 0.003377 | 0.00477 | 0.004503 | 4.4e-05 | 0.001171 | 75.0096 | 1.3934 | 2.548 | 101.21 | 20.79 | 17.636 | 64.0388 | 103.1097 | 60.7787 | 0.004283 | 1.0572 | 0.025824 | 24.4259 | -0.025726 | 0.14088 | -5e-06 | -0.000456 | -0.000456 | -12.8507
2019-03-19 15:37 | 1553024194.7631 | -335 | 0.20946 | 2.3e-05 | 1e-06 | 100.8311 | 100.9424 | 128.4661 | 21.2612 | 0.00175 | 0.001636 | -152.2504 | 0.23667 | 8.4005 | 0.003303 | 0.004592 | 0.004437 | 4.6e-05 | 0.001169 | 74.9563 | 1.3527 | 2.548 | 101.21 | 20.79 | 17.645 | 63.9365 | 101.3212 | 61.1859 | 0.004285 | 1.056 | 0.025787 | 24.4202 | -0.009868 | 0.14079 | -4e-06 | -0.000357 | -0.000357 | -16.4163
2019-03-19 15:37 | 1553024195.7624 | -334 | 0.20942 | 2.5e-05 | 1e-06 | 100.8317 | 100.9222 | 128.4985 | 21.2612 | 0.001767 | 0.001704 | -152.2507 | 0.23637 | 8.4032 | 0.003327 | 0.004543 | 0.004358 | 4.6e-05 | 0.001175 | 75.0033 | 1.3335 | 2.5479 | 101.21 | 20.79 | 17.648 | 64.3501 | 101.896 | 62.836 | 0.004286 | 1.0508 | 0.025727 | 24.4834 | -0.008024 | 0.14062 | -6e-06 | -0.000552 | -0.000552 | -10.6008
2019-03-19 15:37 | 1553024196.7632 | -333 | 0.20944 | 2.4e-05 | 1e-06 | 100.7567 | 100.9924 | 128.3316 | 21.2612 | 0.001795 | 0.001629 | -152.2262 | 0.23701 | 8.4019 | 0.003321 | 0.004496 | 0.004265 | 4.4e-05 | 0.001178 | 75.0333 | 1.3323 | 2.5479 | 101.21 | 20.79 | 17.652 | 65.0082 | 101.742 | 61.0194 | 0.004287 | 1.0565 | 0.02578 | 24.4013 | -0.020902 | 0.141 | -5e-06 | -0.000435 | -0.000435 | -13.4929
2019-03-19 15:37 | 1553024197.7633 | -332 | 0.20946 | 2.2e-05 | 1e-06 | 100.7372 | 100.9195 | 128.5452 | 21.2612 | 0.001747 | 0.001695 | -152.2375 | 0.23641 | 8.4037 | 0.003341 | 0.003641 | 0.003422 | 4.5e-05 | 0.001177 | 75.0972 | 1.3511 | 2.5479 | 101.21 | 20.79 | 17.656 | 63.8614 | 102.2399 | 62.6205 | 0.004288 | 1.0514 | 0.025739 | 24.48 | -0.016171 | 0.14064 | -3e-06 | -0.000256 | -0.000256 | -22.8443
2019-03-19 15:37 | 1553024198.7633 | -331 | 0.20943 | 2.6e-05 | 1e-06 | 100.7444 | 100.8273 | 128.5844 | 21.2612 | 0.00174 | 0.001697 | -152.2433 | 0.23638 | 8.4027 | 0.003344 | 0.004593 | 0.004432 | 4.5e-05 | 0.001173 | 75.0441 | 1.372 | 2.548 | 101.21 | 20.79 | 17.657 | 63.7021 | 102.313 | 62.6665 | 0.004288 | 1.0513 | 0.025742 | 24.4854 | -0.007362 | 0.14062 | -7e-06 | -0.000634 | -0.000634 | -9.2317
2019-03-19 15:37 | 1553024199.7626 | -330 | 0.20943 | 2.3e-05 | 1e-06 | 100.6961 | 100.8407 | 128.2908 | 21.2612 | 0.001774 | 0.0017 | -152.235 | 0.2374 | 8.4011 | 0.003337 | 0.003591 | 0.00332 | 4.5e-05 | 0.001173 | 75.0757 | 1.3733 | 2.548 | 101.21 | 20.79 | 17.66 | 64.5126 | 102.1433 | 62.7252 | 0.004289 | 1.0511 | 0.02571 | 24.4596 | -0.012838 | 0.14123 | 0.000271 | 0.020998 | 0.020998 | 0.27979
2019-03-19 15:37 | 1553024200.7625 | -329 | 0.20944 | 2.2e-05 | 0 | 100.703 | 100.8938 | 127.9227 | 21.2612 | 0.00181 | 0.001732 | -152.2388 | 0.23706 | 8.4012 | 0.003321 | 0.004656 | 0.00439 | 4.6e-05 | 0.001177 | 75.027 | 1.3159 | 2.548 | 101.21 | 20.79 | 17.662 | 65.3845 | 101.7595 | 63.5093 | 0.004289 | 1.0487 | 0.025643 | 24.453 | -0.016932 | 0.14103 | 0.000172 | 0.013284 | 0.013284 | 0.44166
2019-03-19 15:37 | 1553024201.7636 | -328 | 0.20944 | 2.3e-05 | 1e-06 | 100.7594 | 100.9606 | 127.612 | 21.2612 | 0.001815 | 0.001667 | -152.2404 | 0.2366 | 8.4051 | 0.003383 | 0.004838 | 0.004618 | 4.2e-05 | 0.001178 | 75.0909 | 1.3155 | 2.548 | 101.21 | 20.79 | 17.665 | 65.4992 | 103.2456 | 61.9477 | 0.00429 | 1.0536 | 0.025672 | 24.3666 | -0.017841 | 0.14075 | 0.000228 | 0.017608 | 0.017608 | 0.33254
2019-03-19 15:37 | 1553024202.7629 | -327 | 0.20949 | 2.2e-05 | 1e-06 | 100.7771 | 100.7791 | 127.6242 | 21.2612 | 0.001787 | 0.001625 | -152.2598 | 0.23695 | 8.401 | 0.003371 | 0.00474 | 0.004511 | 4.5e-05 | 0.001181 | 75.0922 | 1.3351 | 2.548 | 101.21 | 20.79 | 17.675 | 64.8326 | 102.9714 | 60.9296 | 0.004292 | 1.0568 | 0.025712 | 24.3307 | -0.000173 | 0.14096 | -9.9e-05 | -0.007751 | -0.007751 | -0.7565
2019-03-19 15:37 | 1553024203.7627 | -326 | 0.20942 | 2.5e-05 | 1e-06 | 100.7069 | 100.7914 | 127.934 | 21.2612 | 0.001805 | 0.001711 | -152.2573 | 0.23608 | 8.4026 | 0.003275 | 0.005132 | 0.004912 | 4.3e-05 | 0.001192 | 75.0925 | 1.3351 | 2.5481 | 101.21 | 20.79 | 17.679 | 65.2471 | 100.6438 | 63.0064 | 0.004293 | 1.0502 | 0.025663 | 24.4358 | -0.007503 | 0.14044 | 0.000325 | 0.025084 | 0.025084 | 0.23291
2019-03-19 15:37 | 1553024204.7626 | -325 | 0.20944 | 1.8e-05 | 1e-06 | 100.7186 | 100.8445 | 128.226 | 21.2612 | 0.001731 | 0.001657 | -152.2419 | 0.237 | 8.4035 | 0.003375 | 0.004496 | 0.004246 | 4.6e-05 | 0.001191 | 75.0586 | 1.3521 | 2.5479 | 101.21 | 20.79 | 17.683 | 63.4821 | 103.0547 | 61.7071 | 0.004294 | 1.0543 | 0.025743 | 24.4163 | -0.011173 | 0.14099 | 0.000219 | 0.016966 | 0.016966 | 0.3457
2019-03-19 15:37 | 1553024205.7639 | -324 | 0.20947 | 2.3e-05 | 1e-06 | 100.6931 | 100.9407 | 128.1129 | 21.2612 | 0.001813 | 0.001645 | -152.2514 | 0.2359 | 8.4025 | 0.003284 | 0.004546 | 0.00434 | 4.7e-05 | 0.001188 | 75.0549 | 1.3976 | 2.548 | 101.21 | 20.79 | 17.686 | 65.447 | 100.8651 | 61.4164 | 0.004295 | 1.0552 | 0.025742 | 24.395 | -0.021971 | 0.14034 | 1e-06 | 2e-06 | 2e-06 | 3691.861
2019-03-19 15:37 | 1553024206.7638 | -323 | 0.20948 | 1.9e-05 | 1e-06 | 100.7198 | 100.802 | 128.1713 | 21.2612 | 0.001741 | 0.001714 | -152.2646 | 0.23735 | 8.4036 | 0.003294 | 0.005077 | 0.004897 | 4.8e-05 | 0.001187 | 75.0252 | 1.3398 | 2.5481 | 101.21 | 20.79 | 17.687 | 63.7124 | 101.1092 | 63.0603 | 0.004295 | 1.0501 | 0.025685 | 24.4604 | -0.007299 | 0.1412 | -3.6e-05 | -0.00291 | -0.00291 | -2.0186
2019-03-19 15:37 | 1553024207.7628 | -322 | 0.20945 | 1.9e-05 | 1e-06 | 100.6697 | 100.7849 | 128.16 | 21.2612 | 0.001773 | 0.001709 | -152.2595 | 0.23677 | 8.4039 | 0.003321 | 0.005041 | 0.004771 | 4.7e-05 | 0.001189 | 75.0455 | 1.3755 | 2.5481 | 101.21 | 20.79 | 17.691 | 64.4938 | 101.7563 | 62.9407 | 0.004296 | 1.0504 | 0.025689 | 24.455 | -0.010231 | 0.14085 | 0.000161 | 0.012436 | 0.012436 | 0.47118
2019-03-19 15:37 | 1553024208.7636 | -321 | 0.20949 | 1.9e-05 | 1e-06 | 100.7015 | 100.7541 | 127.9401 | 21.2612 | 0.001756 | 0.001693 | -152.2453 | 0.23667 | 8.4045 | 0.003319 | 0.004289 | 0.004137 | 4.5e-05 | 0.001189 | 75.0708 | 1.3284 | 2.548 | 101.21 | 20.79 | 17.693 | 64.0878 | 101.7123 | 62.562 | 0.004297 | 1.0516 | 0.025681 | 24.4202 | -0.004669 | 0.1408 | -7.6e-05 | -0.005965 | -0.005965 | -0.982
2019-03-19 15:37 | 1553024209.7629 | -320 | 0.20948 | 2.5e-05 | 0 | 100.7138 | 100.8133 | 128.1905 | 21.2612 | 0.001767 | 0.001705 | -152.2406 | 0.23666 | 8.4035 | 0.003353 | 0.00475 | 0.004572 | 4.8e-05 | 0.001188 | 75.1062 | 1.3152 | 2.5481 | 101.21 | 20.79 | 17.698 | 64.3564 | 102.5331 | 62.861 | 0.004298 | 1.0507 | 0.025695 | 24.455 | -0.008834 | 0.14079 | -2.7e-05 | -0.00211 | -0.00211 | -2.7763
2019-03-19 15:37 | 1553024210.7624 | -319 | 0.20945 | 1.7e-05 | 1e-06 | 100.6913 | 100.7469 | 128.2856 | 21.2612 | 0.001732 | 0.001677 | -152.2421 | 0.23705 | 8.401 | 0.003436 | 0.00454 | 0.00428 | 4.4e-05 | 0.001192 | 75.0825 | 1.3066 | 2.5481 | 101.21 | 20.79 | 17.706 | 63.4999 | 104.5277 | 62.1743 | 0.0043 | 1.0528 | 0.025731 | 24.439 | -0.004937 | 0.14102 | 0.000153 | 0.01184 | 0.01184 | 0.49548
2019-03-19 15:37 | 1553024211.7624 | -318 | 0.20949 | 2.1e-05 | 1e-06 | 100.7009 | 100.7249 | 128.3807 | 21.2612 | 0.00183 | 0.001652 | -152.2519 | 0.23694 | 8.4028 | 0.003456 | 0.004895 | 0.004735 | 4.7e-05 | 0.001197 | 75.0299 | 1.3491 | 2.5479 | 101.21 | 20.79 | 17.71 | 65.8583 | 105.0176 | 61.5682 | 0.004301 | 1.0548 | 0.025763 | 24.426 | -0.002137 | 0.14096 | -0.000108 | -0.00847 | -0.00847 | -0.69231
2019-03-19 15:37 | 1553024212.7631 | -317 | 0.20949 | 2e-05 | 1e-06 | 100.6811 | 100.7338 | 128.6161 | 21.2612 | 0.001806 | 0.001656 | -152.2442 | 0.2366 | 8.4014 | 0.00348 | 0.004768 | 0.004613 | 4.8e-05 | 0.001196 | 74.9944 | 1.3583 | 2.548 | 101.21 | 20.79 | 17.713 | 65.2916 | 105.6173 | 61.6626 | 0.004301 | 1.0545 | 0.025783 | 24.4518 | -0.004687 | 0.14075 | -7.1e-05 | -0.005601 | -0.005601 | -1.0453
2019-03-19 15:37 | 1553024213.7631 | -316 | 0.20952 | 2.1e-05 | 1e-06 | 100.7339 | 100.9013 | 128.4154 | 21.2612 | 0.001729 | 0.001681 | -152.2596 | 0.23595 | 8.4055 | 0.003438 | 0.004083 | 0.003827 | 4.8e-05 | 0.001195 | 74.9492 | 1.4005 | 2.5481 | 101.21 | 20.79 | 17.713 | 63.4265 | 104.5929 | 62.2658 | 0.004301 | 1.0526 | 0.02574 | 24.4547 | -0.014855 | 0.14036 | -0.000284 | -0.022129 | -0.022129 | -0.26387
2019-03-19 15:37 | 1553024214.7622 | -315 | 0.20947 | 1.6e-05 | 1e-06 | 100.742 | 100.8431 | 128.2973 | 21.2612 | 0.001834 | 0.001665 | -152.2488 | 0.23632 | 8.4031 | 0.003445 | 0.004986 | 0.004837 | 4.8e-05 | 0.001197 | 74.9813 | 1.3361 | 2.5479 | 101.21 | 20.79 | 17.713 | 65.9495 | 104.766 | 61.893 | 0.004301 | 1.0537 | 0.025743 | 24.4299 | -0.008973 | 0.14059 | 1.5e-05 | 0.001128 | 0.001128 | 5.1835
2019-03-19 15:37 | 1553024215.7622 | -314 | 0.20944 | 1.6e-05 | 1e-06 | 100.7537 | 100.8006 | 128.1765 | 21.2612 | 0.001787 | 0.001652 | -152.2389 | 0.23708 | 8.4026 | 0.003441 | 0.005326 | 0.005091 | 4.7e-05 | 0.001193 | 75.0582 | 1.3573 | 2.5479 | 101.21 | 20.79 | 17.716 | 64.8326 | 104.6489 | 61.5853 | 0.004302 | 1.0547 | 0.025742 | 24.4072 | -0.004166 | 0.14104 | 0.000241 | 0.018725 | 0.018725 | 0.31335
2019-03-19 15:37 | 1553024216.7623 | -313 | 0.20944 | 2.2e-05 | 1e-06 | 100.7054 | 100.878 | 128.0744 | 21.2612 | 0.001764 | 0.001671 | -152.2321 | 0.23631 | 8.4015 | 0.003338 | 0.005001 | 0.004814 | 4.7e-05 | 0.001185 | 74.9903 | 1.3362 | 2.548 | 101.21 | 20.79 | 17.718 | 64.2729 | 102.1608 | 62.0337 | 0.004303 | 1.0533 | 0.025715 | 24.4138 | -0.015322 | 0.14058 | 0.000197 | 0.015237 | 0.015237 | 0.38383
2019-03-19 15:37 | 1553024217.7634 | -312 | 0.20947 | 1.6e-05 | 1e-06 | 100.6409 | 100.6972 | 127.8183 | 21.2612 | 0.001811 | 0.001625 | -152.232 | 0.23663 | 8.4006 | 0.003336 | 0.004497 | 0.004203 | 5e-05 | 0.001183 | 75.0753 | 1.338 | 2.5481 | 101.21 | 20.79 | 17.719 | 65.3946 | 102.1199 | 60.9154 | 0.004303 | 1.0568 | 0.025732 | 24.3487 | -0.005006 | 0.14077 | 5.6e-05 | 0.00429 | 0.00429 | 1.365
2019-03-19 15:37 | 1553024218.7624 | -311 | 0.2095 | 2.1e-05 | 1e-06 | 100.7153 | 100.964 | 128.0352 | 21.2612 | 0.001815 | 0.0017 | -152.2386 | 0.23632 | 8.4041 | 0.003342 | 0.004579 | 0.004389 | 5.1e-05 | 0.001186 | 75.0441 | 1.3473 | 2.548 | 101.21 | 20.79 | 17.721 | 65.4946 | 102.2721 | 62.7252 | 0.004303 | 1.0511 | 0.025684 | 24.4352 | -0.022061 | 0.14058 | -0.000143 | -0.011201 | -0.011201 | -0.52211
2019-03-19 15:37 | 1553024219.7629 | -310 | 0.20946 | 2.2e-05 | 1e-06 | 100.697 | 100.7061 | 128.4171 | 21.2612 | 0.001813 | 0.001639 | -152.2472 | 0.23539 | 8.4065 | 0.00335 | 0.004771 | 0.004514 | 5e-05 | 0.001187 | 75.0778 | 1.386 | 2.548 | 101.21 | 20.79 | 17.721 | 65.4619 | 102.4576 | 61.2695 | 0.004303 | 1.0557 | 0.025779 | 24.4186 | -0.000811 | 0.14003 | 8.6e-05 | 0.006633 | 0.006633 | 0.87826
2019-03-19 15:37 | 1553024220.7635 | -309 | 0.20947 | 2.6e-05 | 1e-06 | 100.6859 | 100.8229 | 128.6255 | 21.2612 | 0.001762 | 0.001682 | -152.2471 | 0.23593 | 8.4033 | 0.003451 | 0.004488 | 0.004309 | 5.1e-05 | 0.001194 | 75.0059 | 1.3634 | 2.548 | 101.21 | 20.79 | 17.719 | 64.2238 | 104.9139 | 62.2885 | 0.004303 | 1.0525 | 0.02576 | 24.4756 | -0.012164 | 0.14035 | 2.8e-05 | 0.002089 | 0.002089 | 2.7946
2019-03-19 15:37 | 1553024221.7635 | -308 | 0.2095 | 2.3e-05 | 1e-06 | 100.634 | 100.6232 | 128.8459 | 21.2612 | 0.001815 | 0.001654 | -152.2448 | 0.23639 | 8.4035 | 0.003454 | 0.004608 | 0.004345 | 5.1e-05 | 0.001197 | 75.0328 | 1.3219 | 2.5481 | 101.21 | 20.79 | 17.716 | 65.5047 | 104.965 | 61.6134 | 0.004302 | 1.0546 | 0.025808 | 24.4719 | 0.000954 | 0.14063 | -0.000144 | -0.01128 | -0.01128 | -0.51865
2019-03-19 15:37 | 1553024222.7643 | -307 | 0.20949 | 1.8e-05 | 1e-06 | 100.6649 | 100.8643 | 128.6201 | 21.2612 | 0.001768 | 0.001685 | -152.245 | 0.23628 | 8.4033 | 0.003471 | 0.004597 | 0.004391 | 5.1e-05 | 0.001196 | 75.0184 | 1.3321 | 2.548 | 101.21 | 20.79 | 17.714 | 64.3752 | 105.3796 | 62.365 | 0.004302 | 1.0522 | 0.025757 | 24.4778 | -0.017706 | 0.14057 | -9.4e-05 | -0.007381 | -0.007381 | -0.79228
2019-03-19 15:37 | 1553024223.7641 | -306 | 0.20951 | 2e-05 | 0 | 100.6844 | 100.7842 | 128.3243 | 21.2612 | 0.001773 | 0.00165 | -152.2302 | 0.23641 | 8.4033 | 0.00345 | 0.005 | 0.004811 | 4.8e-05 | 0.001203 | 75.0471 | 1.3062 | 2.548 | 101.21 | 20.79 | 17.709 | 64.4837 | 104.8746 | 61.5376 | 0.004301 | 1.0549 | 0.025759 | 24.4195 | -0.008863 | 0.14064 | -0.000195 | -0.015199 | -0.015199 | -0.38493
2019-03-19 15:37 | 1553024224.7617 | -305 | 0.20946 | 1.8e-05 | 1e-06 | 100.6048 | 100.5712 | 128.2647 | 21.2612 | 0.001767 | 0.001697 | -152.2494 | 0.23671 | 8.4009 | 0.00345 | 0.00422 | 0.004037 | 4.9e-05 | 0.001204 | 75.0516 | 1.335 | 2.548 | 101.21 | 20.79 | 17.707 | 64.3376 | 104.8729 | 62.6508 | 0.0043 | 1.0514 | 0.02571 | 24.4544 | 0.002995 | 0.14082 | 0.000102 | 0.007851 | 0.007851 | 0.74615
2019-03-19 15:37 | 1553024225.7618 | -304 | 0.20951 | 2e-05 | 1e-06 | 100.6619 | 100.7284 | 128.529 | 21.2612 | 0.001734 | 0.00168 | -152.2489 | 0.23522 | 8.4049 | 0.003468 | 0.004453 | 0.004239 | 5e-05 | 0.001203 | 74.9967 | 1.3386 | 2.5478 | 101.21 | 20.79 | 17.697 | 63.5555 | 105.3127 | 62.2502 | 0.004298 | 1.0526 | 0.025752 | 24.465 | -0.005907 | 0.13993 | -0.000204 | -0.01591 | -0.01591 | -0.36588
2019-03-19 15:37 | 1553024226.7633 | -303 | 0.20952 | 1.9e-05 | 1e-06 | 100.6814 | 100.6147 | 128.8473 | 21.2612 | 0.001804 | 0.001643 | -152.2541 | 0.23656 | 8.4033 | 0.003421 | 0.004984 | 0.004797 | 4.8e-05 | 0.0012 | 74.9931 | 1.3267 | 2.548 | 101.21 | 20.79 | 17.689 | 65.2346 | 104.1827 | 61.3499 | 0.004296 | 1.0554 | 0.025819 | 24.4624 | 0.005929 | 0.14073 | -0.000284 | -0.022214 | -0.022214 | -0.26354
2019-03-19 15:37 | 1553024227.7631 | -302 | 0.20952 | 1.8e-05 | 1e-06 | 100.6289 | 100.6619 | 129.0656 | 21.2612 | 0.001797 | 0.001586 | -152.2562 | 0.23589 | 8.404 | 0.003474 | 0.004969 | 0.004754 | 4.9e-05 | 0.001206 | 75.0658 | 1.3226 | 2.548 | 101.21 | 20.79 | 17.68 | 65.0583 | 105.4715 | 59.9773 | 0.004293 | 1.0598 | 0.025894 | 24.4328 | -0.002939 | 0.14033 | -0.00026 | -0.020349 | -0.020349 | -0.28687
2019-03-19 15:37 | 1553024228.7633 | -301 | 0.20949 | 2.4e-05 | 1e-06 | 100.6298 | 100.7486 | 128.7663 | 21.2612 | 0.001807 | 0.001636 | -152.2583 | 0.23662 | 8.4036 | 0.003468 | 0.004899 | 0.004692 | 4.9e-05 | 0.001203 | 75.0683 | 1.3135 | 2.5478 | 101.21 | 20.79 | 17.66 | 65.2995 | 105.3134 | 61.1991 | 0.004289 | 1.0559 | 0.025816 | 24.4492 | -0.010555 | 0.14076 | -9.2e-05 | -0.007239 | -0.007239 | -0.80886
2019-03-19 15:37 | 1553024229.763 | -300 | 0.20948 | 2.1e-05 | 1e-06 | 100.6355 | 100.7147 | 128.3436 | 21.2612 | 0.001803 | 0.00167 | -152.2531 | 0.23618 | 8.4022 | 0.003413 | 0.004955 | 0.004736 | 5e-05 | 0.001208 | 75.0394 | 1.3544 | 2.548 | 101.21 | 20.8 | 17.649 | 65.2004 | 103.9893 | 61.9978 | 0.004289 | 1.0534 | 0.025743 | 24.4381 | -0.007038 | 0.1405 | -4.1e-05 | -0.003205 | -0.003205 | -1.8236
2019-03-19 15:37 | 1553024230.7636 | -299 | 0.20955 | 2e-05 | 1e-06 | 100.5868 | 100.6088 | 128.0849 | 21.2612 | 0.001781 | 0.00165 | -152.262 | 0.23633 | 8.4007 | 0.00346 | 0.004805 | 0.004545 | 5.1e-05 | 0.001208 | 75.0469 | 1.3218 | 2.548 | 101.21 | 20.79 | 17.627 | 64.6719 | 105.1199 | 61.5258 | 0.004281 | 1.0549 | 0.025735 | 24.3963 | -0.001956 | 0.14059 | -0.000452 | -0.035196 | -0.035196 | -0.16617
2019-03-19 15:37 | 1553024231.7633 | -298 | 0.2095 | 2e-05 | 2e-06 | 100.6376 | 100.7452 | 127.9208 | 21.2612 | 0.001813 | 0.001699 | -152.2408 | 0.23534 | 8.4021 | 0.003407 | 0.00493 | 0.004699 | 4.9e-05 | 0.001208 | 75.0497 | 1.3004 | 2.5481 | 101.21 | 20.8 | 17.616 | 65.4509 | 103.8461 | 62.7064 | 0.004281 | 1.0512 | 0.025674 | 24.4236 | -0.009558 | 0.14 | -0.000183 | -0.014302 | -0.014302 | -0.40723
2019-03-19 15:37 | 1553024232.7625 | -297 | 0.20951 | 2e-05 | 1e-06 | 100.5847 | 100.5732 | 127.9147 | 21.2612 | 0.001743 | 0.001688 | -152.2453 | 0.23648 | 8.4018 | 0.003325 | 0.00457 | 0.00433 | 4.9e-05 | 0.001207 | 75.0349 | 1.3512 | 2.5481 | 101.21 | 20.8 | 17.605 | 63.7569 | 101.8553 | 62.4368 | 0.004278 | 1.052 | 0.025683 | 24.4132 | 0.001024 | 0.14068 | -0.00021 | -0.016326 | -0.016326 | -0.35847
2019-03-19 15:37 | 1553024233.7637 | -296 | 0.20951 | 2.1e-05 | 1e-06 | 100.5391 | 100.7986 | 128.1812 | 21.2612 | 0.001842 | 0.001638 | -152.2544 | 0.23623 | 8.4004 | 0.003366 | 0.005049 | 0.004822 | 5.2e-05 | 0.001209 | 75.0243 | 1.3635 | 2.5479 | 101.21 | 20.8 | 17.578 | 66.158 | 102.8472 | 61.2391 | 0.004271 | 1.0558 | 0.025756 | 24.395 | -0.023053 | 0.14053 | -0.000229 | -0.017858 | -0.017858 | -0.32737
2019-03-19 15:37 | 1553024234.7623 | -295 | 0.20949 | 1.8e-05 | 0 | 100.6514 | 100.6513 | 128.36 | 21.2612 | 0.001801 | 0.001673 | -152.2607 | 0.23596 | 8.4029 | 0.00337 | 0.004869 | 0.004587 | 5.3e-05 | 0.001199 | 75.0148 | 1.3534 | 2.5481 | 101.21 | 20.8 | 17.563 | 65.1643 | 102.9423 | 62.0705 | 0.004268 | 1.0532 | 0.025742 | 24.4424 | 5e-06 | 0.14037 | -0.00011 | -0.008556 | -0.008556 | -0.68251
2019-03-19 15:37 | 1553024235.7628 | -294 | 0.20954 | 1.8e-05 | 1e-06 | 100.594 | 100.6003 | 127.969 | 21.2612 | 0.001752 | 0.001687 | -152.2454 | 0.23659 | 8.4039 | 0.003482 | 0.004997 | 0.004824 | 4.8e-05 | 0.001196 | 74.98 | 1.3354 | 2.5479 | 101.21 | 20.8 | 17.531 | 63.9918 | 105.6628 | 62.4103 | 0.00426 | 1.0521 | 0.02569 | 24.4175 | -0.000554 | 0.14075 | -0.000382 | -0.029712 | -0.029712 | -0.19706
2019-03-19 15:37 | 1553024236.7628 | -293 | 0.20952 | 1.8e-05 | 1e-06 | 100.5823 | 100.6554 | 127.7018 | 21.2612 | 0.001768 | 0.001669 | -152.2628 | 0.23606 | 8.4013 | 0.003492 | 0.00443 | 0.004216 | 4.5e-05 | 0.001195 | 75.0023 | 1.3185 | 2.548 | 101.21 | 20.8 | 17.517 | 64.3769 | 105.8918 | 61.9829 | 0.004257 | 1.0534 | 0.025679 | 24.3764 | -0.006497 | 0.14043 | -0.000274 | -0.021273 | -0.021273 | -0.27462
2019-03-19 15:37 | 1553024237.7635 | -292 | 0.20951 | 2.3e-05 | 1e-06 | 100.5451 | 100.7219 | 127.7392 | 21.2612 | 0.001817 | 0.001649 | -152.2685 | 0.23731 | 8.4031 | 0.003398 | 0.004515 | 0.00428 | 4.6e-05 | 0.001197 | 74.9716 | 1.339 | 2.548 | 101.21 | 20.8 | 17.503 | 65.5444 | 103.6065 | 61.4961 | 0.004253 | 1.055 | 0.025702 | 24.3623 | -0.015709 | 0.14118 | -0.000192 | -0.014969 | -0.014969 | -0.39234
2019-03-19 15:37 | 1553024238.7634 | -291 | 0.2095 | 2.6e-05 | 1e-06 | 100.4902 | 100.702 | 127.9615 | 21.2612 | 0.001799 | 0.001654 | -152.2723 | 0.23601 | 8.4026 | 0.003437 | 0.004832 | 0.004617 | 4.8e-05 | 0.001203 | 75.0012 | 1.3187 | 2.5481 | 101.21 | 20.8 | 17.477 | 65.1121 | 104.5656 | 61.6329 | 0.004247 | 1.0546 | 0.025719 | 24.3884 | -0.018831 | 0.1404 | -0.000151 | -0.011828 | -0.011828 | -0.49382
2019-03-19 15:37 | 1553024239.7634 | -290 | 0.20952 | 2.7e-05 | 1e-06 | 100.5253 | 100.778 | 128.0671 | 21.2612 | 0.00177 | 0.001664 | -152.2751 | 0.23617 | 8.4016 | 0.003397 | 0.004556 | 0.004365 | 4.8e-05 | 0.001204 | 74.984 | 1.3151 | 2.548 | 101.21 | 20.8 | 17.465 | 64.4057 | 103.5828 | 61.8596 | 0.004244 | 1.0538 | 0.025721 | 24.4068 | -0.022457 | 0.1405 | -0.000292 | -0.022711 | -0.022711 | -0.25735
2019-03-19 15:37 | 1553024240.7616 | -289 | 0.20949 | 2.5e-05 | 1e-06 | 100.5379 | 100.7106 | 128.3112 | 21.2612 | 0.001826 | 0.001638 | -152.2624 | 0.23641 | 8.4024 | 0.003343 | 0.004788 | 0.004492 | 4.6e-05 | 0.001213 | 74.9743 | 1.3185 | 2.548 | 101.21 | 20.8 | 17.454 | 65.7638 | 102.2794 | 61.2485 | 0.004241 | 1.0558 | 0.025769 | 24.4077 | -0.015346 | 0.14064 | -0.000111 | -0.008688 | -0.008688 | -0.67346
2019-03-19 15:37 | 1553024241.7615 | -288 | 0.20949 | 2.4e-05 | 1e-06 | 100.5073 | 100.6129 | 128.2739 | 21.2612 | 0.001755 | 0.001637 | -152.2865 | 0.23628 | 8.4031 | 0.003473 | 0.005065 | 0.004861 | 4.6e-05 | 0.001206 | 74.9881 | 1.328 | 2.548 | 101.21 | 20.8 | 17.431 | 64.0669 | 105.4434 | 61.2203 | 0.004236 | 1.0559 | 0.025766 | 24.4031 | -0.009396 | 0.14057 | -0.000112 | -0.008734 | -0.008734 | -0.66948
2019-03-19 15:37 | 1553024242.7626 | -287 | 0.20949 | 2.7e-05 | 0 | 100.5004 | 100.491 | 128.2689 | 21.2612 | 0.001829 | 0.001687 | -152.2772 | 0.23638 | 8.4021 | 0.003349 | 0.004204 | 0.004076 | 4.3e-05 | 0.001196 | 75.0079 | 1.3467 | 2.548 | 101.21 | 20.8 | 17.42 | 65.842 | 102.4341 | 62.4235 | 0.004233 | 1.0521 | 0.025719 | 24.4465 | 0.000836 | 0.14062 | -0.000122 | -0.009518 | -0.009518 | -0.61462
2019-03-19 15:37 | 1553024243.7626 | -286 | 0.20947 | 2.4e-05 | 1e-06 | 100.6021 | 100.6349 | 128.218 | 21.2612 | 0.001833 | 0.001674 | -152.2764 | 0.23622 | 8.4022 | 0.003377 | 0.005244 | 0.004981 | 4.8e-05 | 0.001189 | 75.0229 | 1.3153 | 2.548 | 101.21 | 20.8 | 17.399 | 65.9269 | 103.1206 | 62.1095 | 0.004228 | 1.0531 | 0.025726 | 24.4302 | -0.002909 | 0.14053 | 5e-06 | 0.000285 | 0.000285 | 20.4999
2019-03-19 15:37 | 1553024244.7621 | -285 | 0.20949 | 1.6e-05 | 1e-06 | 100.5841 | 100.7356 | 128.1326 | 21.2612 | 0.001791 | 0.001661 | -152.2618 | 0.23614 | 8.3997 | 0.003384 | 0.005369 | 0.005189 | 4.9e-05 | 0.001189 | 75.0412 | 1.3344 | 2.548 | 101.21 | 20.8 | 17.389 | 64.9319 | 103.2809 | 61.7917 | 0.004225 | 1.054 | 0.02573 | 24.4105 | -0.013455 | 0.14048 | -8.2e-05 | -0.006416 | -0.006416 | -0.91088
2019-03-19 15:37 | 1553024245.7619 | -284 | 0.20948 | 2.7e-05 | 1e-06 | 100.5583 | 100.7287 | 128.4555 | 21.2612 | 0.001758 | 0.001651 | -152.2683 | 0.23569 | 8.404 | 0.003378 | 0.005006 | 0.004847 | 4.8e-05 | 0.001192 | 75.0231 | 1.3733 | 2.5482 | 101.21 | 20.8 | 17.382 | 64.1372 | 103.1363 | 61.5609 | 0.004224 | 1.0548 | 0.025771 | 24.4329 | -0.015142 | 0.14021 | -2.7e-05 | -0.002205 | -0.002205 | -2.6457
2019-03-19 15:37 | 1553024246.7617 | -283 | 0.20947 | 2.9e-05 | 0 | 100.6346 | 100.6575 | 128.5316 | 21.2612 | 0.001774 | 0.001653 | -152.2585 | 0.23531 | 8.4026 | 0.003357 | 0.005407 | 0.005137 | 4.4e-05 | 0.001198 | 75.0053 | 1.3827 | 2.548 | 101.21 | 20.8 | 17.37 | 64.5056 | 102.6234 | 61.6016 | 0.004221 | 1.0546 | 0.025777 | 24.4416 | -0.002037 | 0.13998 | 2.9e-05 | 0.002265 | 0.002265 | 2.5711
2019-03-19 15:37 | 1553024247.7633 | -282 | 0.20946 | 2.7e-05 | 1e-06 | 100.5598 | 100.7366 | 128.2255 | 21.2612 | 0.00184 | 0.001681 | -152.2534 | 0.23587 | 8.4053 | 0.003373 | 0.004519 | 0.004298 | 4.6e-05 | 0.001196 | 74.9879 | 1.3934 | 2.5482 | 101.21 | 20.8 | 17.365 | 66.1104 | 103.0043 | 62.269 | 0.00422 | 1.0526 | 0.025721 | 24.4368 | -0.015708 | 0.14032 | 7.2e-05 | 0.005544 | 0.005544 | 1.0528
2019-03-19 15:37 | 1553024248.7622 | -281 | 0.20941 | 2.5e-05 | 1e-06 | 100.5937 | 100.7154 | 128.2206 | 21.2612 | 0.001811 | 0.001673 | -152.2568 | 0.23606 | 8.4022 | 0.003329 | 0.004129 | 0.003856 | 4.8e-05 | 0.001186 | 75.0521 | 1.3718 | 2.5481 | 101.21 | 20.8 | 17.361 | 65.4088 | 101.9487 | 62.0821 | 0.004219 | 1.0531 | 0.025728 | 24.4295 | -0.010807 | 0.14043 | 0.000398 | 0.030816 | 0.030816 | 0.18958
2019-03-19 15:37 | 1553024249.7621 | -280 | 0.20948 | 2e-05 | 1e-06 | 100.5775 | 100.6808 | 128.2457 | 21.2612 | 0.001777 | 0.001634 | -152.2665 | 0.23626 | 8.4017 | 0.003358 | 0.004585 | 0.004377 | 4.7e-05 | 0.001182 | 75.0342 | 1.4034 | 2.5481 | 101.21 | 20.8 | 17.355 | 64.5807 | 102.6414 | 61.1367 | 0.004217 | 1.0561 | 0.025766 | 24.3974 | -0.009175 | 0.14055 | -4.2e-05 | -0.003322 | -0.003322 | -1.76
2019-03-19 15:37 | 1553024250.7631 | -279 | 0.20947 | 2.2e-05 | 1e-06 | 100.5604 | 100.7523 | 127.98 | 21.2612 | 0.001739 | 0.001609 | -152.2736 | 0.2357 | 8.403 | 0.003326 | 0.005369 | 0.00516 | 4.9e-05 | 0.001185 | 75.0159 | 1.3927 | 2.5481 | 101.21 | 20.8 | 17.354 | 63.6717 | 101.8851 | 60.5489 | 0.004217 | 1.058 | 0.025762 | 24.3507 | -0.017054 | 0.14022 | 2e-05 | 0.001456 | 0.001456 | 4.0075
2019-03-19 15:38 | 1553024251.7621 | -278 | 0.20941 | 2.6e-05 | 1e-06 | 100.5994 | 100.6068 | 128.2312 | 21.2612 | 0.001752 | 0.001655 | -152.2745 | 0.23611 | 8.4031 | 0.00337 | 0.003917 | 0.003719 | 4.4e-05 | 0.00119 | 74.9946 | 1.342 | 2.5481 | 101.21 | 20.8 | 17.354 | 63.9793 | 102.9511 | 61.6532 | 0.004217 | 1.0545 | 0.025745 | 24.4149 | -0.000652 | 0.14047 | 0.000403 | 0.031222 | 0.031222 | 0.18715
2019-03-19 15:38 | 1553024252.762 | -277 | 0.20943 | 2.5e-05 | 1e-06 | 100.5103 | 100.6951 | 128.4537 | 21.2612 | 0.001783 | 0.001651 | -152.2951 | 0.23618 | 8.4007 | 0.003357 | 0.004686 | 0.004461 | 4.4e-05 | 0.001185 | 75.0281 | 1.3228 | 2.5481 | 101.21 | 20.8 | 17.355 | 64.7303 | 102.6312 | 61.5525 | 0.004217 | 1.0548 | 0.025771 | 24.4324 | -0.016433 | 0.1405 | 0.000265 | 0.020559 | 0.020559 | 0.2843
2019-03-19 15:38 | 1553024253.7636 | -276 | 0.20952 | 2.4e-05 | 1e-06 | 100.5793 | 100.6962 | 128.7551 | 21.2612 | 0.001718 | 0.001686 | -152.2745 | 0.23556 | 8.4 | 0.003346 | 0.004338 | 0.004151 | 4.7e-05 | 0.001183 | 75.0362 | 1.3581 | 2.5479 | 101.21 | 20.8 | 17.358 | 63.1773 | 102.3673 | 62.3845 | 0.004218 | 1.0522 | 0.02577 | 24.4914 | -0.010384 | 0.14013 | -0.000289 | -0.022506 | -0.022506 | -0.25901
2019-03-19 15:38 | 1553024254.7629 | -275 | 0.20949 | 2.7e-05 | 1e-06 | 100.5916 | 100.816 | 128.9809 | 21.2612 | 0.001725 | 0.001674 | -152.2803 | 0.23547 | 8.4024 | 0.003335 | 0.004756 | 0.004495 | 4.5e-05 | 0.001183 | 75.0236 | 1.344 | 2.5481 | 101.21 | 20.8 | 17.368 | 63.3352 | 102.0838 | 62.1009 | 0.00422 | 1.0531 | 0.025803 | 24.5025 | -0.019932 | 0.14008 | -0.000118 | -0.009269 | -0.009269 | -0.62868
2019-03-19 15:38 | 1553024255.7615 | -274 | 0.20944 | 2.5e-05 | 1e-06 | 100.5238 | 100.5609 | 128.7544 | 21.2612 | 0.001796 | 0.001717 | -152.2956 | 0.23674 | 8.4044 | 0.003319 | 0.00492 | 0.004763 | 4.6e-05 | 0.001169 | 75.0193 | 1.3315 | 2.548 | 101.21 | 20.8 | 17.374 | 65.0395 | 101.6934 | 63.1328 | 0.004222 | 1.0498 | 0.025741 | 24.5186 | -0.003298 | 0.14084 | 0.000222 | 0.017167 | 0.017167 | 0.34129
2019-03-19 15:38 | 1553024256.7613 | -273 | 0.20937 | 2.8e-05 | 1e-06 | 100.582 | 100.7143 | 128.4518 | 21.2612 | 0.001808 | 0.001601 | -152.3007 | 0.23692 | 8.4043 | 0.003351 | 0.004396 | 0.004252 | 4.6e-05 | 0.001166 | 75.0083 | 1.3199 | 2.5481 | 101.21 | 20.8 | 17.39 | 65.3253 | 102.4803 | 60.3417 | 0.004226 | 1.0586 | 0.025818 | 24.388 | -0.011756 | 0.14094 | 0.000623 | 0.048422 | 0.048422 | 0.12109
2019-03-19 15:38 | 1553024257.7627 | -272 | 0.20946 | 2.7e-05 | 1e-06 | 100.6262 | 100.7647 | 128.9813 | 21.2612 | 0.001744 | 0.001616 | -152.3109 | 0.23615 | 8.4041 | 0.003244 | 0.004664 | 0.004396 | 4.6e-05 | 0.001162 | 75.003 | 1.3078 | 2.548 | 101.21 | 20.79 | 17.398 | 63.7897 | 99.8851 | 60.7021 | 0.004225 | 1.0575 | 0.025857 | 24.4514 | -0.012304 | 0.14048 | 8.1e-05 | 0.006208 | 0.006208 | 0.94133
2019-03-19 15:38 | 1553024258.7625 | -271 | 0.20941 | 2.4e-05 | 1e-06 | 100.5376 | 100.4742 | 129.0731 | 21.2612 | 0.001819 | 0.001616 | -152.3051 | 0.23735 | 8.4026 | 0.003248 | 0.004917 | 0.004674 | 4.4e-05 | 0.00117 | 75.0627 | 1.3021 | 2.548 | 101.21 | 20.79 | 17.407 | 65.5906 | 99.9791 | 60.699 | 0.004227 | 1.0575 | 0.025866 | 24.46 | 0.005641 | 0.1412 | 0.000401 | 0.031227 | 0.031227 | 0.18811
2019-03-19 15:38 | 1553024259.7629 | -270 | 0.20946 | 2.6e-05 | 1e-06 | 100.5718 | 100.5718 | 128.8069 | 21.2612 | 0.001777 | 0.001716 | -152.3145 | 0.23693 | 8.401 | 0.003257 | 0.004742 | 0.004512 | 4.6e-05 | 0.001175 | 75.0688 | 1.302 | 2.548 | 101.21 | 20.79 | 17.407 | 64.5775 | 100.1941 | 63.107 | 0.004227 | 1.0499 | 0.025747 | 24.5227 | -1e-06 | 0.14095 | 0.0001 | 0.007699 | 0.007699 | 0.7616
2019-03-19 15:38 | 1553024260.7626 | -269 | 0.20945 | 2.5e-05 | 1e-06 | 100.5094 | 100.5236 | 128.8485 | 21.2612 | 0.001725 | 0.001649 | -152.3006 | 0.23673 | 8.4008 | 0.003224 | 0.004121 | 0.003924 | 4.8e-05 | 0.001178 | 75.0143 | 1.3236 | 2.548 | 101.21 | 20.79 | 17.434 | 63.3352 | 99.4059 | 61.5132 | 0.004234 | 1.0549 | 0.025812 | 24.4684 | -0.001259 | 0.14083 | 0.000131 | 0.010133 | 0.010133 | 0.5782
2019-03-19 15:38 | 1553024261.7616 | -268 | 0.2095 | 2.8e-05 | 1e-06 | 100.6057 | 100.8143 | 128.726 | 21.2612 | 0.001745 | 0.001696 | -152.2861 | 0.23682 | 8.403 | 0.0033 | 0.004463 | 0.004264 | 4.7e-05 | 0.001179 | 74.9936 | 1.3398 | 2.5479 | 101.21 | 20.79 | 17.443 | 63.8084 | 101.2546 | 62.6323 | 0.004236 | 1.0514 | 0.025757 | 24.4977 | -0.018526 | 0.14088 | -0.00018 | -0.014074 | -0.014074 | -0.41642
2019-03-19 15:38 | 1553024262.7616 | -267 | 0.20942 | 2.5e-05 | 2e-06 | 100.5805 | 100.7328 | 128.5631 | 21.2612 | 0.001778 | 0.001648 | -152.2785 | 0.23671 | 8.4019 | 0.003229 | 0.004603 | 0.004313 | 5e-05 | 0.001182 | 75.0525 | 1.3337 | 2.548 | 101.21 | 20.79 | 17.453 | 64.5985 | 99.5282 | 61.4704 | 0.004238 | 1.0551 | 0.025785 | 24.4398 | -0.013532 | 0.14082 | 0.000352 | 0.027283 | 0.027283 | 0.21471
2019-03-19 15:38 | 1553024263.7635 | -266 | 0.20947 | 2.7e-05 | 1e-06 | 100.5853 | 100.6294 | 128.0992 | 21.2612 | 0.001795 | 0.001637 | -152.2925 | 0.23572 | 8.4014 | 0.003249 | 0.004382 | 0.004181 | 4.8e-05 | 0.001181 | 75.0259 | 1.3182 | 2.548 | 101.21 | 20.79 | 17.473 | 65.0207 | 100.0144 | 61.2061 | 0.004243 | 1.0559 | 0.025749 | 24.386 | -0.003916 | 0.14023 | 4.7e-05 | 0.003589 | 0.003589 | 1.6253
2019-03-19 15:38 | 1553024264.7617 | -265 | 0.20949 | 2.5e-05 | 1e-06 | 100.643 | 100.6657 | 128.1971 | 21.2612 | 0.001796 | 0.001703 | -152.2924 | 0.23582 | 8.401 | 0.003156 | 0.004435 | 0.004197 | 4.7e-05 | 0.001177 | 75.0461 | 1.3456 | 2.548 | 101.21 | 20.79 | 17.483 | 65.0537 | 97.7686 | 62.8126 | 0.004246 | 1.0508 | 0.025697 | 24.4539 | -0.00202 | 0.14029 | -0.0001 | -0.007801 | -0.007801 | -0.74811
2019-03-19 15:38 | 1553024265.7628 | -264 | 0.20939 | 2.2e-05 | 1e-06 | 100.6087 | 100.7112 | 128.2739 | 21.2612 | 0.001808 | 0.001666 | -152.2916 | 0.23556 | 8.4026 | 0.003246 | 0.004281 | 0.004053 | 4.6e-05 | 0.001173 | 75.1235 | 1.4112 | 2.5481 | 101.21 | 20.79 | 17.501 | 65.3416 | 99.9337 | 61.9056 | 0.00425 | 1.0537 | 0.02574 | 24.4281 | -0.009108 | 0.14014 | 0.000507 | 0.039281 | 0.039281 | 0.14841
2019-03-19 15:38 | 1553024266.7632 | -263 | 0.20953 | 2.4e-05 | 1e-06 | 100.6403 | 100.7277 | 128.4919 | 21.2612 | 0.001759 | 0.001669 | -152.2784 | 0.23574 | 8.4043 | 0.003212 | 0.004819 | 0.004574 | 4.7e-05 | 0.001176 | 75.0868 | 1.3359 | 2.548 | 101.21 | 20.79 | 17.511 | 64.1425 | 99.1155 | 61.9814 | 0.004252 | 1.0535 | 0.025759 | 24.4517 | -0.007767 | 0.14024 | -0.000368 | -0.028651 | -0.028651 | -0.20363
2019-03-19 15:38 | 1553024267.7619 | -262 | 0.20952 | 2.1e-05 | 1e-06 | 100.5763 | 100.7616 | 128.2523 | 21.2612 | 0.001759 | 0.001696 | -152.2824 | 0.23602 | 8.4036 | 0.003244 | 0.004686 | 0.004499 | 4.5e-05 | 0.001175 | 75.1108 | 1.3404 | 2.5481 | 101.21 | 20.79 | 17.521 | 64.141 | 99.8984 | 62.6424 | 0.004255 | 1.0514 | 0.025709 | 24.4529 | -0.01646 | 0.14041 | -0.000281 | -0.021866 | -0.021866 | -0.26713
2019-03-19 15:38 | 1553024268.7634 | -261 | 0.20949 | 2e-05 | 1e-06 | 100.6592 | 100.6571 | 128.0445 | 21.2612 | 0.001789 | 0.001679 | -152.259 | 0.23639 | 8.4032 | 0.003261 | 0.004012 | 0.003836 | 4.7e-05 | 0.001167 | 75.0803 | 1.3694 | 2.5482 | 101.21 | 20.79 | 17.542 | 64.8802 | 100.3102 | 62.2314 | 0.00426 | 1.0527 | 0.025704 | 24.4182 | 0.000181 | 0.14063 | -9.9e-05 | -0.007694 | -0.007694 | -0.76036
2019-03-19 15:38 | 1553024269.7626 | -260 | 0.20946 | 2.2e-05 | 1e-06 | 100.612 | 100.7767 | 128.5485 | 21.2612 | 0.001792 | 0.001702 | -152.2479 | 0.236 | 8.4041 | 0.003157 | 0.004848 | 0.004561 | 4.6e-05 | 0.001167 | 75.0755 | 1.3574 | 2.5481 | 101.21 | 20.8 | 17.552 | 64.942 | 97.7992 | 62.7782 | 0.004265 | 1.051 | 0.025734 | 24.4861 | -0.014622 | 0.14039 | 7.3e-05 | 0.005621 | 0.005621 | 1.039
2019-03-19 15:38 | 1553024270.7631 | -259 | 0.20946 | 2.1e-05 | 1e-06 | 100.6811 | 100.8195 | 128.5569 | 21.2612 | 0.001798 | 0.001652 | -152.2641 | 0.23614 | 8.4036 | 0.003177 | 0.004719 | 0.004555 | 4.6e-05 | 0.001165 | 75.1028 | 1.3434 | 2.548 | 101.21 | 20.8 | 17.562 | 65.0847 | 98.2718 | 61.5828 | 0.004267 | 1.0547 | 0.025781 | 24.4433 | -0.012287 | 0.14048 | 7.1e-05 | 0.005448 | 0.005448 | 1.0726
2019-03-19 15:38 | 1553024271.762 | -258 | 0.20952 | 2.3e-05 | 1e-06 | 100.6102 | 100.6393 | 128.503 | 21.2612 | 0.001847 | 0.001637 | -152.2856 | 0.23629 | 8.4018 | 0.00322 | 0.005059 | 0.00484 | 4.4e-05 | 0.001164 | 75.0751 | 1.3257 | 2.5481 | 101.21 | 20.8 | 17.579 | 66.2756 | 99.3063 | 61.2235 | 0.004272 | 1.0558 | 0.025789 | 24.425 | -0.002584 | 0.14057 | -0.000291 | -0.022754 | -0.022754 | -0.257
2019-03-19 15:38 | 1553024272.7615 | -257 | 0.20944 | 2e-05 | 1e-06 | 100.6346 | 100.8181 | 128.568 | 21.2612 | 0.001801 | 0.001662 | -152.2744 | 0.23653 | 8.4021 | 0.003313 | 0.004595 | 0.004414 | 4.7e-05 | 0.00116 | 75.0349 | 1.3305 | 2.5482 | 101.21 | 20.8 | 17.585 | 65.1542 | 101.5569 | 61.8174 | 0.004273 | 1.054 | 0.025773 | 24.4529 | -0.016299 | 0.14071 | 0.000182 | 0.014119 | 0.014119 | 0.41458
2019-03-19 15:38 | 1553024273.7613 | -256 | 0.20949 | 2.2e-05 | 1e-06 | 100.667 | 100.9191 | 128.718 | 21.2612 | 0.001776 | 0.001593 | -152.277 | 0.23605 | 8.4025 | 0.003268 | 0.004573 | 0.004313 | 4.7e-05 | 0.001161 | 75.0319 | 1.3625 | 2.5481 | 101.21 | 20.8 | 17.596 | 64.5508 | 100.4829 | 60.1525 | 0.004276 | 1.0592 | 0.025852 | 24.4063 | -0.022378 | 0.14043 | -7.5e-05 | -0.005911 | -0.005911 | -0.98826
2019-03-19 15:38 | 1553024274.7617 | -255 | 0.20949 | 2.2e-05 | 1e-06 | 100.6607 | 100.8452 | 128.4889 | 21.2612 | 0.00181 | 0.001667 | -152.2785 | 0.23569 | 8.4014 | 0.003365 | 0.004404 | 0.004175 | 4.6e-05 | 0.001167 | 74.9757 | 1.3741 | 2.5479 | 101.21 | 20.8 | 17.599 | 65.3799 | 102.8261 | 61.9439 | 0.004276 | 1.0536 | 0.02576 | 24.45 | -0.016379 | 0.14021 | -0.000114 | -0.008965 | -0.008965 | -0.6506
2019-03-19 15:38 | 1553024275.7623 | -254 | 0.20947 | 2.2e-05 | 1e-06 | 100.6238 | 100.6599 | 128.7931 | 21.2612 | 0.001779 | 0.001717 | -152.256 | 0.23636 | 8.403 | 0.00337 | 0.005026 | 0.004832 | 4.8e-05 | 0.001168 | 74.9493 | 1.3393 | 2.548 | 101.21 | 20.8 | 17.602 | 64.636 | 102.929 | 63.1501 | 0.004277 | 1.0498 | 0.025744 | 24.5229 | -0.00321 | 0.14061 | 5.6e-05 | 0.004301 | 0.004301 | 1.36
2019-03-19 15:38 | 1553024276.7624 | -253 | 0.20944 | 2e-05 | 1e-06 | 100.5745 | 100.814 | 128.6194 | 21.2612 | 0.001795 | 0.001688 | -152.2636 | 0.23669 | 8.4026 | 0.003293 | 0.004265 | 0.004026 | 4.3e-05 | 0.00117 | 75.0448 | 1.3356 | 2.548 | 101.21 | 20.8 | 17.607 | 65.0161 | 101.0833 | 62.4337 | 0.004278 | 1.052 | 0.025754 | 24.4803 | -0.021271 | 0.14081 | 0.000247 | 0.019135 | 0.019135 | 0.30613
2019-03-19 15:38 | 1553024277.7619 | -252 | 0.20949 | 2.4e-05 | 1e-06 | 100.6397 | 100.7633 | 128.5497 | 21.2612 | 0.001806 | 0.001704 | -152.2463 | 0.23576 | 8.4033 | 0.003308 | 0.004258 | 0.00404 | 4.6e-05 | 0.001168 | 74.989 | 1.318 | 2.5479 | 101.21 | 20.8 | 17.607 | 65.2916 | 101.4327 | 62.8297 | 0.004278 | 1.0508 | 0.025732 | 24.4881 | -0.010982 | 0.14025 | -6.1e-05 | -0.004787 | -0.004787 | -1.2189
2019-03-19 15:38 | 1553024278.7625 | -251 | 0.20946 | 1.9e-05 | 1e-06 | 100.6304 | 100.8003 | 128.5886 | 21.2612 | 0.001756 | 0.001678 | -152.2687 | 0.23683 | 8.4012 | 0.003366 | 0.005085 | 0.004934 | 4.5e-05 | 0.001164 | 75.0549 | 1.3732 | 2.5481 | 101.21 | 20.8 | 17.604 | 64.0722 | 102.8472 | 62.2088 | 0.004278 | 1.0527 | 0.02576 | 24.4692 | -0.015092 | 0.14089 | 8.1e-05 | 0.006215 | 0.006215 | 0.9431
2019-03-19 15:38 | 1553024279.7626 | -250 | 0.20944 | 2.4e-05 | 2e-06 | 100.6562 | 100.7965 | 128.4351 | 21.2612 | 0.001847 | 0.001678 | -152.2612 | 0.23642 | 8.3998 | 0.003321 | 0.004546 | 0.004271 | 4.5e-05 | 0.00117 | 74.9766 | 1.3505 | 2.5481 | 101.21 | 20.8 | 17.601 | 66.2805 | 101.757 | 62.2063 | 0.004277 | 1.0527 | 0.025744 | 24.4545 | -0.012464 | 0.14065 | 0.000182 | 0.014 | 0.014 | 0.41791
2019-03-19 15:38 | 1553024280.7617 | -249 | 0.20948 | 2.2e-05 | 1e-06 | 100.6688 | 100.7013 | 128.5192 | 21.2612 | 0.001788 | 0.001671 | -152.2717 | 0.23641 | 8.4031 | 0.003264 | 0.004748 | 0.004579 | 4.6e-05 | 0.001172 | 75.0457 | 1.2843 | 2.5481 | 101.21 | 20.8 | 17.599 | 64.8381 | 100.3755 | 62.033 | 0.004276 | 1.0533 | 0.025759 | 24.4561 | -0.002891 | 0.14064 | -2.5e-05 | -0.002012 | -0.002012 | -2.9074
2019-03-19 15:38 | 1553024281.763 | -248 | 0.20946 | 2.2e-05 | 1e-06 | 100.6706 | 100.8054 | 128.5208 | 21.2612 | 0.001741 | 0.001623 | -152.2567 | 0.2357 | 8.4033 | 0.00332 | 0.004447 | 0.004285 | 4.6e-05 | 0.001178 | 75.0565 | 1.3375 | 2.5478 | 101.21 | 20.8 | 17.593 | 63.7194 | 101.7311 | 60.8864 | 0.004275 | 1.0569 | 0.025804 | 24.4144 | -0.011974 | 0.14022 | 6.8e-05 | 0.005253 | 0.005253 | 1.1103
2019-03-19 15:38 | 1553024282.7617 | -247 | 0.20945 | 2.3e-05 | 1e-06 | 100.6421 | 100.8486 | 128.6447 | 21.2612 | 0.001802 | 0.001673 | -152.2622 | 0.23537 | 8.4044 | 0.003271 | 0.004908 | 0.004666 | 4.3e-05 | 0.001183 | 75.0765 | 1.3222 | 2.5478 | 101.21 | 20.8 | 17.59 | 65.1831 | 100.5364 | 62.072 | 0.004274 | 1.0532 | 0.025771 | 24.4695 | -0.018336 | 0.14002 | 0.000154 | 0.011957 | 0.011957 | 0.48714
2019-03-19 15:38 | 1553024283.7621 | -246 | 0.20949 | 2.6e-05 | 1e-06 | 100.7129 | 100.8647 | 128.7144 | 21.2612 | 0.00184 | 0.001651 | -152.2642 | 0.23688 | 8.403 | 0.00334 | 0.004619 | 0.004451 | 4.4e-05 | 0.001188 | 75.0281 | 1.2972 | 2.5482 | 101.21 | 20.8 | 17.59 | 66.0941 | 102.2227 | 61.5407 | 0.004274 | 1.0548 | 0.025798 | 24.4567 | -0.013471 | 0.14092 | -6.2e-05 | -0.0049 | -0.0049 | -1.1963
2019-03-19 15:38 | 1553024284.7619 | -245 | 0.2095 | 2e-05 | 1e-06 | 100.6745 | 101.0178 | 128.5802 | 21.2612 | 0.001769 | 0.001674 | -152.2475 | 0.236 | 8.4049 | 0.003248 | 0.004546 | 0.004319 | 4.5e-05 | 0.001189 | 75.0453 | 1.3426 | 2.5479 | 101.21 | 20.81 | 17.591 | 64.3877 | 99.9941 | 62.1002 | 0.004277 | 1.0531 | 0.025763 | 24.4644 | -0.030449 | 0.14039 | -0.000167 | -0.013081 | -0.013081 | -0.4465
2019-03-19 15:38 | 1553024285.7631 | -244 | 0.20949 | 1.9e-05 | 1e-06 | 100.7033 | 100.7318 | 128.4234 | 21.2612 | 0.00174 | 0.001662 | -152.2651 | 0.23617 | 8.4038 | 0.003333 | 0.004131 | 0.00387 | 4.4e-05 | 0.00119 | 75.048 | 1.3575 | 2.5481 | 101.21 | 20.81 | 17.591 | 63.6874 | 102.0436 | 61.8119 | 0.004277 | 1.054 | 0.025758 | 24.439 | -0.002532 | 0.1405 | -0.000102 | -0.007992 | -0.007992 | -0.73137
2019-03-19 15:38 | 1553024286.761 | -243 | 0.20945 | 2.5e-05 | 1e-06 | 100.6727 | 100.8116 | 128.5307 | 21.2612 | 0.00181 | 0.001646 | -152.2525 | 0.2367 | 8.3995 | 0.00339 | 0.004434 | 0.004228 | 4.5e-05 | 0.001183 | 75.0335 | 1.3557 | 2.5481 | 101.21 | 20.81 | 17.592 | 65.3698 | 103.4171 | 61.4391 | 0.004277 | 1.0552 | 0.025783 | 24.4355 | -0.012335 | 0.14081 | 0.000134 | 0.01035 | 0.01035 | 0.56595
2019-03-19 15:38 | 1553024287.7613 | -242 | 0.20948 | 2e-05 | 1e-06 | 100.7699 | 100.8592 | 128.595 | 21.2612 | 0.001715 | 0.001707 | -152.2584 | 0.23605 | 8.4028 | 0.003335 | 0.004712 | 0.004471 | 4.5e-05 | 0.001184 | 75.0428 | 1.3312 | 2.5481 | 101.21 | 20.81 | 17.592 | 63.0931 | 102.0915 | 62.8954 | 0.004277 | 1.0506 | 0.025734 | 24.4948 | -0.007923 | 0.14043 | -1.7e-05 | -0.001381 | -0.001381 | -4.2296
2019-03-19 15:38 | 1553024288.7621 | -241 | 0.20948 | 2.6e-05 | 1e-06 | 100.7537 | 100.7287 | 128.8072 | 21.2612 | 0.001779 | 0.001712 | -152.2516 | 0.23592 | 8.4018 | 0.003382 | 0.004765 | 0.004498 | 4.1e-05 | 0.001182 | 75.0455 | 1.3292 | 2.5482 | 101.21 | 20.81 | 17.593 | 64.6406 | 103.2385 | 63.0252 | 0.004278 | 1.0502 | 0.02575 | 24.5197 | 0.002218 | 0.14035 | -1.6e-05 | -0.001288 | -0.001288 | -4.5322
2019-03-19 15:38 | 1553024289.7621 | -240 | 0.20947 | 1.9e-05 | 1e-06 | 100.697 | 100.8657 | 128.9717 | 21.2612 | 0.001798 | 0.00164 | -152.2524 | 0.23506 | 8.4044 | 0.003422 | 0.004802 | 0.004635 | 4.1e-05 | 0.001182 | 75.0559 | 1.3394 | 2.5481 | 101.21 | 20.81 | 17.598 | 65.0965 | 104.2079 | 61.2837 | 0.004279 | 1.0557 | 0.025834 | 24.4718 | -0.014975 | 0.13984 | 4.1e-05 | 0.003125 | 0.003125 | 1.8617
2019-03-19 15:38 | 1553024290.7626 | -239 | 0.2095 | 2.3e-05 | 1e-06 | 100.6877 | 100.7571 | 128.9295 | 21.2612 | 0.001762 | 0.001659 | -152.2294 | 0.23586 | 8.4029 | 0.003435 | 0.004405 | 0.004162 | 4.3e-05 | 0.001167 | 74.9736 | 1.3163 | 2.5481 | 101.21 | 20.81 | 17.598 | 64.2159 | 104.5033 | 61.7517 | 0.004279 | 1.0542 | 0.025811 | 24.4849 | -0.006169 | 0.14031 | -0.000139 | -0.010919 | -0.010919 | -0.5346
2019-03-19 15:38 | 1553024291.7612 | -238 | 0.20948 | 1.9e-05 | 1e-06 | 100.7279 | 100.7732 | 128.9342 | 21.2612 | 0.001734 | 0.001649 | -152.2236 | 0.23582 | 8.4022 | 0.003456 | 0.005211 | 0.004982 | 4.1e-05 | 0.001158 | 74.9325 | 1.3187 | 2.5481 | 101.21 | 20.81 | 17.597 | 63.542 | 105.0302 | 61.5041 | 0.004279 | 1.055 | 0.025821 | 24.4762 | -0.004026 | 0.14029 | -4.6e-05 | -0.00362 | -0.00362 | -1.612
2019-03-19 15:38 | 1553024292.7627 | -237 | 0.20951 | 2.3e-05 | 1e-06 | 100.7564 | 100.7616 | 128.7191 | 21.2612 | 0.001793 | 0.00169 | -152.2402 | 0.23522 | 8.4039 | 0.003401 | 0.004771 | 0.004504 | 3.9e-05 | 0.001152 | 75.0086 | 1.3486 | 2.5482 | 101.21 | 20.8 | 17.595 | 64.9591 | 103.692 | 62.4837 | 0.004275 | 1.0519 | 0.025762 | 24.4916 | -0.000461 | 0.13993 | -0.000231 | -0.018051 | -0.018051 | -0.32249
2019-03-19 15:38 | 1553024293.7623 | -236 | 0.20949 | 2.1e-05 | 1e-06 | 100.7576 | 100.6417 | 128.7426 | 21.2612 | 0.001821 | 0.001641 | -152.2418 | 0.23604 | 8.4013 | 0.003449 | 0.004998 | 0.004781 | 4e-05 | 0.001163 | 74.9834 | 1.3402 | 2.5482 | 101.21 | 20.8 | 17.592 | 65.6414 | 104.8494 | 61.3227 | 0.004275 | 1.0555 | 0.025809 | 24.4514 | 0.010293 | 0.14042 | -7.6e-05 | -0.005934 | -0.005934 | -0.98436
2019-03-19 15:38 | 1553024294.7628 | -235 | 0.20949 | 1.8e-05 | 1e-06 | 100.7099 | 100.8123 | 128.7612 | 21.2612 | 0.001747 | 0.001682 | -152.2656 | 0.2359 | 8.4039 | 0.003545 | 0.004503 | 0.004338 | 4.2e-05 | 0.001168 | 74.9143 | 1.3109 | 2.548 | 101.21 | 20.8 | 17.58 | 63.8623 | 107.1937 | 62.2978 | 0.004272 | 1.0525 | 0.025773 | 24.4888 | -0.00909 | 0.14034 | -5.8e-05 | -0.004616 | -0.004616 | -1.2647
2019-03-19 15:38 | 1553024295.7624 | -234 | 0.20946 | 2e-05 | 1e-06 | 100.751 | 100.8578 | 128.7316 | 21.2612 | 0.001757 | 0.001693 | -152.2568 | 0.236 | 8.4006 | 0.003488 | 0.004384 | 0.004227 | 3.9e-05 | 0.001169 | 75.003 | 1.3111 | 2.5482 | 101.21 | 20.8 | 17.574 | 64.0934 | 105.8028 | 62.5589 | 0.00427 | 1.0516 | 0.025761 | 24.4955 | -0.009479 | 0.14039 | 8.4e-05 | 0.006465 | 0.006465 | 0.90342
2019-03-19 15:38 | 1553024296.7619 | -233 | 0.2095 | 2.4e-05 | 0 | 100.7084 | 100.9352 | 128.9806 | 21.2612 | 0.001824 | 0.001677 | -152.2557 | 0.23591 | 8.4043 | 0.003394 | 0.004987 | 0.004723 | 4.1e-05 | 0.001166 | 75.0326 | 1.3206 | 2.548 | 101.21 | 20.8 | 17.561 | 65.7085 | 103.519 | 62.1876 | 0.004267 | 1.0528 | 0.0258 | 24.5056 | -0.020125 | 0.14034 | -0.000169 | -0.013227 | -0.013227 | -0.44139
2019-03-19 15:38 | 1553024297.7614 | -232 | 0.2095 | 1.7e-05 | 1e-06 | 100.7192 | 100.7969 | 128.918 | 21.2612 | 0.001761 | 0.001669 | -152.2502 | 0.23733 | 8.4037 | 0.003467 | 0.004269 | 0.00409 | 4.4e-05 | 0.001169 | 75.0432 | 1.3112 | 2.5479 | 101.21 | 20.8 | 17.556 | 64.2043 | 105.2955 | 61.9822 | 0.004266 | 1.0535 | 0.025801 | 24.4922 | -0.006896 | 0.14119 | -0.000158 | -0.012352 | -0.012352 | -0.47549
2019-03-19 15:38 | 1553024298.7619 | -231 | 0.20953 | 2e-05 | 1e-06 | 100.7153 | 101.0274 | 128.8046 | 21.2612 | 0.001824 | 0.001623 | -152.2402 | 0.23622 | 8.4023 | 0.003488 | 0.004851 | 0.004587 | 4e-05 | 0.001173 | 75.0302 | 1.3189 | 2.5481 | 101.21 | 20.8 | 17.551 | 65.7099 | 105.8076 | 60.8857 | 0.004265 | 1.0569 | 0.025832 | 24.4413 | -0.027673 | 0.14053 | -0.000314 | -0.024537 | -0.024537 | -0.23825
2019-03-19 15:38 | 1553024299.7627 | -230 | 0.2095 | 2.5e-05 | 1e-06 | 100.8029 | 100.8383 | 128.8306 | 21.2612 | 0.00177 | 0.00168 | -152.2509 | 0.23597 | 8.402 | 0.003456 | 0.004528 | 0.004313 | 4.1e-05 | 0.001172 | 75.0466 | 1.3315 | 2.5481 | 101.21 | 20.8 | 17.539 | 64.4214 | 105.0249 | 62.247 | 0.004262 | 1.0526 | 0.025782 | 24.4936 | -0.00314 | 0.14038 | -0.000158 | -0.012346 | -0.012346 | -0.47303
2019-03-19 15:38 | 1553024300.7619 | -229 | 0.20949 | 2.1e-05 | 1e-06 | 100.7537 | 100.8486 | 128.9691 | 21.2612 | 0.001753 | 0.001686 | -152.2619 | 0.23675 | 8.4007 | 0.00345 | 0.004827 | 0.004598 | 4.2e-05 | 0.001179 | 75.0193 | 1.3471 | 2.5482 | 101.21 | 20.8 | 17.533 | 64.0082 | 104.8841 | 62.3939 | 0.00426 | 1.0522 | 0.025791 | 24.5121 | -0.00842 | 0.14084 | -6.3e-05 | -0.004966 | -0.004966 | -1.1798
2019-03-19 15:38 | 1553024301.7617 | -228 | 0.20952 | 2.4e-05 | 1e-06 | 100.7792 | 100.8681 | 129.0351 | 21.2612 | 0.001797 | 0.001666 | -152.2806 | 0.23673 | 8.4052 | 0.003422 | 0.004625 | 0.004397 | 4.3e-05 | 0.001176 | 75.1019 | 1.3211 | 2.548 | 101.21 | 20.8 | 17.526 | 65.07 | 104.1921 | 61.9166 | 0.004259 | 1.0537 | 0.025816 | 24.5009 | -0.007887 | 0.14083 | -0.000293 | -0.022866 | -0.022866 | -0.25621
2019-03-19 15:38 | 1553024302.7609 | -227 | 0.20945 | 2.4e-05 | 1e-06 | 100.6991 | 100.7311 | 128.9663 | 21.2612 | 0.001782 | 0.001677 | -152.2771 | 0.23603 | 8.4041 | 0.003418 | 0.005023 | 0.004807 | 4.2e-05 | 0.00118 | 75.0007 | 1.3713 | 2.5482 | 101.21 | 20.8 | 17.524 | 64.7101 | 104.0913 | 62.169 | 0.004258 | 1.0529 | 0.025799 | 24.5036 | -0.002845 | 0.14042 | 0.000172 | 0.01332 | 0.01332 | 0.43854
2019-03-19 15:38 | 1553024303.7622 | -226 | 0.20968 | 2.1e-05 | 1e-06 | 100.7516 | 100.988 | 129.021 | 21.2612 | 0.001724 | 0.00175 | -152.2689 | 0.23692 | 8.398 | 0.003437 | 0.004582 | 0.004355 | 4.1e-05 | 0.00118 | 75.0416 | 1.3666 | 2.5482 | 101.21 | 20.8 | 17.525 | 63.3196 | 104.574 | 63.942 | 0.004258 | 1.0473 | 0.025736 | 24.5735 | -0.020961 | 0.14094 | -0.001251 | -0.097225 | -0.097225 | -0.060307
2019-03-19 15:38 | 1553024304.762 | -225 | 0.20932 | 2.3e-05 | 1e-06 | 100.6874 | 100.9732 | 128.8996 | 21.2612 | 0.001763 | 0.001661 | -152.2664 | 0.23728 | 8.404 | 0.003385 | 0.005309 | 0.004939 | 4.2e-05 | 0.001173 | 75.0317 | 1.373 | 2.5481 | 101.21 | 20.8 | 17.531 | 64.251 | 103.3015 | 61.8008 | 0.00426 | 1.054 | 0.025806 | 24.4838 | -0.02536 | 0.14116 | 0.000946 | 0.073609 | 0.073609 | 0.079775
2019-03-19 15:38 | 1553024305.7622 | -224 | 0.20956 | 2.1e-05 | 1e-06 | 100.7672 | 100.8037 | 128.8694 | 21.2612 | 0.001772 | 0.001627 | -152.2886 | 0.23659 | 8.4036 | 0.003473 | 0.004694 | 0.004432 | 4.6e-05 | 0.00118 | 75.0302 | 1.3519 | 2.5481 | 101.21 | 20.8 | 17.536 | 64.4548 | 105.4424 | 60.9734 | 0.004261 | 1.0566 | 0.025835 | 24.4507 | -0.003241 | 0.14075 | -0.000525 | -0.041024 | -0.041024 | -0.14272
2019-03-19 15:38 | 1553024306.762 | -223 | 0.20949 | 2.5e-05 | 1e-06 | 100.6715 | 100.9592 | 128.6133 | 21.2612 | 0.001749 | 0.001669 | -152.3009 | 0.23636 | 8.404 | 0.0035 | 0.004567 | 0.004413 | 4.3e-05 | 0.001181 | 75.0118 | 1.3272 | 2.5481 | 101.21 | 20.8 | 17.542 | 63.91 | 106.0815 | 61.994 | 0.004263 | 1.0534 | 0.02577 | 24.4637 | -0.025529 | 0.14061 | -6.1e-05 | -0.004824 | -0.004824 | -1.2126
2019-03-19 15:38 | 1553024307.7624 | -222 | 0.20952 | 1.6e-05 | 1e-06 | 100.7585 | 100.8732 | 128.549 | 21.2612 | 0.001726 | 0.001666 | -152.2993 | 0.23648 | 8.4004 | 0.003496 | 0.00491 | 0.004616 | 4.3e-05 | 0.001182 | 75.0778 | 1.3629 | 2.5482 | 101.21 | 20.8 | 17.556 | 63.3485 | 105.99 | 61.9229 | 0.004266 | 1.0536 | 0.025767 | 24.4549 | -0.01018 | 0.14068 | -0.000275 | -0.021478 | -0.021478 | -0.27248
2019-03-19 15:38 | 1553024308.7622 | -221 | 0.20977 | 2.5e-05 | 1e-06 | 100.6988 | 100.8366 | 128.3523 | 21.2612 | 0.00177 | 0.001623 | -152.3135 | 0.23673 | 8.4026 | 0.003488 | 0.004987 | 0.004903 | 4.1e-05 | 0.001183 | 74.9194 | 1.3483 | 2.5483 | 101.21 | 20.8 | 17.564 | 64.411 | 105.8052 | 60.871 | 0.004268 | 1.057 | 0.025787 | 24.3978 | -0.012233 | 0.14083 | -0.001795 | -0.13981 | -0.13981 | -0.041904
2019-03-19 15:38 | 1553024309.7624 | -220 | 0.20969 | 1.9e-05 | 1e-06 | 100.7546 | 100.8602 | 128.1654 | 21.2612 | 0.00173 | 0.001656 | -152.3077 | 0.23656 | 8.4024 | 0.003497 | 0.004523 | 0.004178 | 4.2e-05 | 0.001193 | 74.9382 | 1.3256 | 2.548 | 101.21 | 20.8 | 17.581 | 63.4428 | 106.0279 | 61.6727 | 0.004272 | 1.0544 | 0.025738 | 24.4093 | -0.009372 | 0.14073 | -0.0013 | -0.1011 | -0.1011 | -0.057909
2019-03-19 15:38 | 1553024310.7626 | -219 | 0.20927 | 2.3e-05 | 1e-06 | 100.7366 | 100.7763 | 127.9673 | 21.2612 | 0.001787 | 0.001656 | -152.3232 | 0.23642 | 8.4061 | 0.003509 | 0.004948 | 0.004747 | 4.3e-05 | 0.0012 | 75.033 | 1.3252 | 2.5481 | 101.21 | 20.8 | 17.59 | 64.835 | 106.309 | 61.672 | 0.004274 | 1.0544 | 0.025718 | 24.3904 | -0.003527 | 0.14065 | 0.001217 | 0.094389 | 0.094389 | 0.061987
2019-03-19 15:38 | 1553024311.7612 | -218 | 0.20931 | 1.7e-05 | 1e-06 | 100.7498 | 100.8732 | 127.7124 | 21.2612 | 0.001739 | 0.001645 | -152.3142 | 0.23716 | 8.403 | 0.003503 | 0.004473 | 0.004315 | 4.5e-05 | 0.001203 | 75.0739 | 1.3168 | 2.5481 | 101.21 | 20.8 | 17.599 | 63.6749 | 106.1586 | 61.4157 | 0.004276 | 1.0552 | 0.025702 | 24.3568 | -0.010953 | 0.14109 | 0.000985 | 0.076349 | 0.076349 | 0.076874
2019-03-19 15:39 | 1553024312.7627 | -217 | 0.2093 | 2.1e-05 | 1e-06 | 100.7936 | 100.7616 | 127.9826 | 21.2612 | 0.001789 | 0.001637 | -152.3063 | 0.2363 | 8.4021 | 0.003443 | 0.004249 | 0.004013 | 5.1e-05 | 0.001206 | 75.0207 | 1.3252 | 2.548 | 101.21 | 20.81 | 17.615 | 64.8677 | 104.7149 | 61.2266 | 0.004283 | 1.0558 | 0.025737 | 24.3756 | 0.002842 | 0.14058 | 0.00104 | 0.080665 | 0.080665 | 0.072497
2019-03-19 15:39 | 1553024313.7614 | -216 | 0.20956 | 2.2e-05 | 1e-06 | 100.676 | 100.6297 | 127.9708 | 21.2612 | 0.00179 | 0.001679 | -152.3014 | 0.2357 | 8.4049 | 0.003414 | 0.005079 | 0.004893 | 4.5e-05 | 0.00121 | 75.0292 | 1.3426 | 2.5479 | 101.21 | 20.81 | 17.623 | 64.8906 | 104.0048 | 62.222 | 0.004285 | 1.0527 | 0.025697 | 24.4108 | 0.004109 | 0.14022 | -0.000483 | -0.037575 | -0.037575 | -0.15524
2019-03-19 15:39 | 1553024314.7618 | -215 | 0.20943 | 2e-05 | 1e-06 | 100.7696 | 100.8756 | 128.0568 | 21.2612 | 0.001801 | 0.001672 | -152.2845 | 0.23623 | 8.4032 | 0.003466 | 0.005004 | 0.004729 | 4.6e-05 | 0.001212 | 74.9985 | 1.3466 | 2.5481 | 101.21 | 20.8 | 17.635 | 65.1542 | 105.2545 | 62.0486 | 0.004285 | 1.0532 | 0.025712 | 24.4127 | -0.009407 | 0.14053 | 0.000303 | 0.023434 | 0.023434 | 0.24948
2019-03-19 15:39 | 1553024315.7625 | -214 | 0.20934 | 1.9e-05 | 1e-06 | 100.7105 | 100.8113 | 128.1797 | 21.2612 | 0.001756 | 0.001725 | -152.2671 | 0.23624 | 8.4002 | 0.003433 | 0.004986 | 0.004833 | 4.6e-05 | 0.001211 | 75.0425 | 1.3449 | 2.548 | 101.21 | 20.8 | 17.641 | 64.091 | 104.4657 | 63.3242 | 0.004287 | 1.0493 | 0.025676 | 24.4708 | -0.008945 | 0.14054 | 0.000824 | 0.063798 | 0.063798 | 0.091639
2019-03-19 15:39 | 1553024316.7616 | -213 | 0.20966 | 1.9e-05 | 1e-06 | 100.6844 | 100.7825 | 128.1117 | 21.2612 | 0.001726 | 0.001667 | -152.2834 | 0.23687 | 8.4047 | 0.003426 | 0.00506 | 0.004796 | 4.5e-05 | 0.001198 | 75.0466 | 1.3496 | 2.5481 | 101.21 | 20.8 | 17.646 | 63.3586 | 104.2902 | 61.9393 | 0.004288 | 1.0536 | 0.025722 | 24.4139 | -0.008712 | 0.14091 | -0.001135 | -0.088195 | -0.088195 | -0.066467
2019-03-19 15:39 | 1553024317.7616 | -212 | 0.20978 | 1.8e-05 | 1e-06 | 100.673 | 100.7588 | 128.1591 | 21.2612 | 0.001766 | 0.001708 | -152.3112 | 0.23639 | 8.4012 | 0.003545 | 0.004356 | 0.004144 | 4.5e-05 | 0.00119 | 75.0226 | 1.2904 | 2.5481 | 101.21 | 20.81 | 17.658 | 64.3261 | 107.1891 | 62.9344 | 0.004293 | 1.0505 | 0.025689 | 24.4547 | -0.007627 | 0.14063 | -0.00184 | -0.14274 | -0.14274 | -0.040986
2019-03-19 15:39 | 1553024318.7607 | -211 | 0.20974 | 2.2e-05 | 1e-06 | 100.646 | 100.7756 | 128.4145 | 21.2612 | 0.001776 | 0.00168 | -152.2964 | 0.23673 | 8.4034 | 0.003526 | 0.005211 | 0.004917 | 4.3e-05 | 0.001187 | 74.998 | 1.2997 | 2.548 | 101.21 | 20.81 | 17.664 | 64.5515 | 106.7331 | 62.2439 | 0.004295 | 1.0526 | 0.025741 | 24.4539 | -0.011517 | 0.14083 | -0.001623 | -0.12615 | -0.12615 | -0.046441
2019-03-19 15:39 | 1553024319.7618 | -210 | 0.20953 | 2.1e-05 | 1e-06 | 100.6283 | 100.8369 | 128.4342 | 21.2612 | 0.001779 | 0.00175 | -152.301 | 0.23637 | 8.4031 | 0.00352 | 0.005154 | 0.004997 | 4.1e-05 | 0.001178 | 75.0053 | 1.3234 | 2.5482 | 101.21 | 20.8 | 17.669 | 64.6312 | 106.5693 | 63.9348 | 0.004293 | 1.0473 | 0.025678 | 24.5173 | -0.01853 | 0.14062 | -0.000311 | -0.024153 | -0.024153 | -0.24219
2019-03-19 15:39 | 1553024320.7621 | -209 | 0.20946 | 1.8e-05 | 1e-06 | 100.6388 | 100.7619 | 128.5081 | 21.2612 | 0.001825 | 0.001677 | -152.2828 | 0.23645 | 8.4034 | 0.003464 | 0.004886 | 0.004628 | 4.1e-05 | 0.001182 | 75.0373 | 1.3198 | 2.5481 | 101.21 | 20.81 | 17.677 | 65.7335 | 105.2198 | 62.1893 | 0.004298 | 1.0528 | 0.025752 | 24.4608 | -0.010941 | 0.14067 | 0.000113 | 0.008692 | 0.008692 | 0.67321
2019-03-19 15:39 | 1553024321.7608 | -208 | 0.20954 | 1.8e-05 | 1e-06 | 100.6214 | 100.726 | 128.8067 | 21.2612 | 0.001729 | 0.001704 | -152.2763 | 0.23647 | 8.4029 | 0.003478 | 0.004532 | 0.004298 | 4.4e-05 | 0.001188 | 75.0491 | 1.3366 | 2.5481 | 101.21 | 20.81 | 17.68 | 63.4414 | 105.5612 | 62.8345 | 0.004299 | 1.0508 | 0.025757 | 24.5127 | -0.009295 | 0.14067 | -0.000365 | -0.028398 | -0.028398 | -0.20608
2019-03-19 15:39 | 1553024322.7622 | -207 | 0.20955 | 1.6e-05 | 1e-06 | 100.5994 | 100.702 | 128.6776 | 21.2612 | 0.001768 | 0.001644 | -152.2601 | 0.23676 | 8.4062 | 0.003443 | 0.004958 | 0.004694 | 4.7e-05 | 0.001193 | 74.9981 | 1.3317 | 2.5481 | 101.21 | 20.8 | 17.689 | 64.3581 | 104.7204 | 61.3923 | 0.004298 | 1.0553 | 0.0258 | 24.4478 | -0.009114 | 0.14085 | -0.000412 | -0.032124 | -0.032124 | -0.1824
2019-03-19 15:39 | 1553024323.7619 | -206 | 0.20942 | 2.3e-05 | 1e-06 | 100.625 | 100.6451 | 128.6168 | 21.2612 | 0.001764 | 0.001668 | -152.2798 | 0.23691 | 8.4019 | 0.003468 | 0.005227 | 0.005068 | 4.7e-05 | 0.001198 | 74.987 | 1.3183 | 2.5481 | 101.21 | 20.8 | 17.691 | 64.2729 | 105.3095 | 61.9518 | 0.004299 | 1.0535 | 0.025772 | 24.4625 | -0.001794 | 0.14094 | 0.000346 | 0.026847 | 0.026847 | 0.21839
2019-03-19 15:39 | 1553024324.761 | -205 | 0.20959 | 2.4e-05 | 0 | 100.6015 | 100.7017 | 128.556 | 21.2612 | 0.001783 | 0.001659 | -152.3039 | 0.23627 | 8.4028 | 0.003522 | 0.004692 | 0.004461 | 4.3e-05 | 0.001202 | 74.9935 | 1.3481 | 2.548 | 101.21 | 20.8 | 17.696 | 64.7212 | 106.6152 | 61.7399 | 0.0043 | 1.0542 | 0.025774 | 24.4489 | -0.008896 | 0.14056 | -0.00069 | -0.053733 | -0.053733 | -0.10882
2019-03-19 15:39 | 1553024325.7624 | -204 | 0.2095 | 1.8e-05 | 0 | 100.6136 | 100.763 | 128.2922 | 21.2612 | 0.001832 | 0.001688 | -152.304 | 0.23636 | 8.4048 | 0.003441 | 0.004642 | 0.004385 | 3.9e-05 | 0.001211 | 75.0385 | 1.3699 | 2.5481 | 101.21 | 20.8 | 17.704 | 65.9185 | 104.6511 | 62.4486 | 0.004302 | 1.052 | 0.025721 | 24.4497 | -0.013273 | 0.14061 | -0.000136 | -0.010598 | -0.010598 | -0.55194
2019-03-19 15:39 | 1553024326.7611 | -203 | 0.20951 | 1.6e-05 | 0 | 100.5448 | 100.7154 | 128.5677 | 21.2612 | 0.001776 | 0.001678 | -152.2821 | 0.236 | 8.4021 | 0.003487 | 0.004708 | 0.004459 | 4.4e-05 | 0.001214 | 75.1029 | 1.3279 | 2.5481 | 101.21 | 20.8 | 17.708 | 64.5525 | 105.78 | 62.2136 | 0.004303 | 1.0527 | 0.025757 | 24.4673 | -0.015157 | 0.1404 | -0.000212 | -0.016471 | -0.016471 | -0.3546
2019-03-19 15:39 | 1553024327.7615 | -202 | 0.20947 | 2.2e-05 | 1e-06 | 100.5955 | 100.8428 | 128.6384 | 21.2612 | 0.001823 | 0.001661 | -152.2852 | 0.23761 | 8.4032 | 0.003483 | 0.004964 | 0.004799 | 4.7e-05 | 0.001212 | 75.0615 | 1.3518 | 2.548 | 101.21 | 20.8 | 17.716 | 65.689 | 105.6692 | 61.7868 | 0.004305 | 1.0541 | 0.025781 | 24.4585 | -0.021955 | 0.14136 | 3.9e-05 | 0.002929 | 0.002929 | 2.0076
2019-03-19 15:39 | 1553024328.7612 | -201 | 0.20948 | 1.9e-05 | 0 | 100.5283 | 100.4554 | 128.6818 | 21.2612 | 0.001684 | 0.0017 | -152.2712 | 0.23656 | 8.4022 | 0.003471 | 0.004867 | 0.004618 | 4.3e-05 | 0.001196 | 75.057 | 1.3927 | 2.548 | 101.21 | 20.8 | 17.719 | 62.3556 | 105.3835 | 62.7361 | 0.004306 | 1.0511 | 0.025749 | 24.4973 | 0.00649 | 0.14073 | 3e-06 | 0.0002 | 0.0002 | 29.2815
2019-03-19 15:39 | 1553024329.7622 | -200 | 0.20953 | 1.7e-05 | 1e-06 | 100.5274 | 100.6304 | 128.3495 | 21.2612 | 0.00177 | 0.001618 | -152.2777 | 0.23528 | 8.4044 | 0.003462 | 0.004294 | 0.00417 | 4.7e-05 | 0.001195 | 74.9998 | 1.3421 | 2.548 | 101.21 | 20.8 | 17.722 | 64.4228 | 105.1813 | 60.7513 | 0.004306 | 1.0573 | 0.025792 | 24.3932 | -0.00916 | 0.13997 | -0.000333 | -0.02594 | -0.02594 | -0.22447
2019-03-19 15:39 | 1553024330.7613 | -199 | 0.20943 | 2.1e-05 | 1e-06 | 100.5277 | 100.7188 | 128.3565 | 21.2612 | 0.001758 | 0.001702 | -152.2822 | 0.23652 | 8.4023 | 0.003416 | 0.004779 | 0.004444 | 4.3e-05 | 0.001197 | 74.9691 | 1.3464 | 2.5481 | 101.21 | 20.8 | 17.726 | 64.1331 | 104.0663 | 62.7688 | 0.004307 | 1.051 | 0.025715 | 24.4675 | -0.016984 | 0.1407 | 0.000287 | 0.022213 | 0.022213 | 0.26351
2019-03-19 15:39 | 1553024331.7624 | -198 | 0.20949 | 1.7e-05 | 1e-06 | 100.5358 | 100.6225 | 128.3363 | 21.2612 | 0.001815 | 0.001695 | -152.2777 | 0.23676 | 8.4016 | 0.003453 | 0.004811 | 0.004682 | 4.1e-05 | 0.001204 | 75.0354 | 1.3153 | 2.5482 | 101.21 | 20.8 | 17.728 | 65.504 | 104.9548 | 62.6041 | 0.004308 | 1.0515 | 0.025719 | 24.4595 | -0.007712 | 0.14085 | -0.000105 | -0.008208 | -0.008208 | -0.71383
2019-03-19 15:39 | 1553024332.7605 | -197 | 0.20948 | 2e-05 | 0 | 100.5232 | 100.4088 | 128.2992 | 21.2612 | 0.001745 | 0.001701 | -152.2918 | 0.2365 | 8.3987 | 0.003448 | 0.005247 | 0.004976 | 4.3e-05 | 0.001202 | 74.9933 | 1.33 | 2.5482 | 101.21 | 20.8 | 17.731 | 63.8272 | 104.8312 | 62.7657 | 0.004309 | 1.051 | 0.025709 | 24.4619 | 0.010184 | 0.14069 | -3.2e-05 | -0.002509 | -0.002509 | -2.3324
2019-03-19 15:39 | 1553024333.7607 | -196 | 0.2095 | 2e-05 | 1e-06 | 100.5145 | 100.5482 | 128.5619 | 21.2612 | 0.001765 | 0.001703 | -152.2884 | 0.23585 | 8.4011 | 0.003432 | 0.004505 | 0.004231 | 4.2e-05 | 0.001197 | 75.0125 | 1.3165 | 2.5483 | 101.21 | 20.81 | 17.733 | 64.2871 | 104.4468 | 62.8095 | 0.004312 | 1.0509 | 0.025734 | 24.4885 | -0.002999 | 0.14031 | -0.000144 | -0.0112 | -0.0112 | -0.52113
2019-03-19 15:39 | 1553024334.7606 | -195 | 0.20949 | 1.8e-05 | 1e-06 | 100.5118 | 100.5475 | 128.4081 | 21.2612 | 0.001838 | 0.00168 | -152.2871 | 0.23644 | 8.4054 | 0.003458 | 0.004495 | 0.004323 | 4.4e-05 | 0.001193 | 74.9445 | 1.3196 | 2.548 | 101.21 | 20.81 | 17.734 | 66.0448 | 105.0782 | 62.2509 | 0.004312 | 1.0526 | 0.02574 | 24.4535 | -0.003178 | 0.14066 | -7.9e-05 | -0.00617 | -0.00617 | -0.94832
2019-03-19 15:39 | 1553024335.7619 | -194 | 0.20947 | 2.3e-05 | 1e-06 | 100.5376 | 100.841 | 128.6966 | 21.2612 | 0.001764 | 0.001694 | -152.2886 | 0.23706 | 8.4032 | 0.003403 | 0.003781 | 0.00353 | 4.1e-05 | 0.00118 | 74.956 | 1.3342 | 2.5481 | 101.21 | 20.81 | 17.736 | 64.2777 | 103.7322 | 62.5947 | 0.004312 | 1.0515 | 0.025756 | 24.4935 | -0.026954 | 0.14103 | 3.8e-05 | 0.002924 | 0.002924 | 2.0063
2019-03-19 15:39 | 1553024336.761 | -193 | 0.2095 | 2.2e-05 | 1e-06 | 100.5292 | 100.4571 | 128.5741 | 21.2612 | 0.001802 | 0.001664 | -152.2966 | 0.23578 | 8.401 | 0.00343 | 0.005167 | 0.004987 | 4.1e-05 | 0.001183 | 75.0363 | 1.2878 | 2.548 | 101.21 | 20.81 | 17.735 | 65.1911 | 104.4051 | 61.8634 | 0.004312 | 1.0538 | 0.025771 | 24.4552 | 0.006418 | 0.14026 | -0.000133 | -0.010415 | -0.010415 | -0.56024
2019-03-19 15:39 | 1553024337.7621 | -192 | 0.20945 | 2e-05 | 1e-06 | 100.5088 | 100.576 | 128.468 | 21.2612 | 0.001791 | 0.001683 | -152.2884 | 0.2367 | 8.4034 | 0.00344 | 0.004792 | 0.004498 | 4.6e-05 | 0.001186 | 74.9443 | 1.3213 | 2.5482 | 101.21 | 20.81 | 17.733 | 64.9177 | 104.6465 | 62.3282 | 0.004312 | 1.0524 | 0.025743 | 24.462 | -0.005974 | 0.14081 | 0.000155 | 0.01198 | 0.01198 | 0.48897
2019-03-19 15:39 | 1553024338.7617 | -191 | 0.20951 | 2.1e-05 | 1e-06 | 100.5472 | 100.4595 | 128.5398 | 21.2612 | 0.001803 | 0.001671 | -152.2992 | 0.23731 | 8.401 | 0.003447 | 0.004801 | 0.004561 | 4.1e-05 | 0.00118 | 74.9955 | 1.3529 | 2.5482 | 101.21 | 20.81 | 17.729 | 65.2105 | 104.8084 | 62.0361 | 0.004311 | 1.0533 | 0.025761 | 24.4582 | 0.007806 | 0.14118 | -0.000192 | -0.014986 | -0.014986 | -0.39189
2019-03-19 15:39 | 1553024339.7617 | -190 | 0.20949 | 2e-05 | 1e-06 | 100.4476 | 100.5208 | 128.7912 | 21.2612 | 0.001792 | 0.00166 | -152.3016 | 0.23665 | 8.4062 | 0.00343 | 0.004724 | 0.004497 | 4.2e-05 | 0.001182 | 75.0171 | 1.3564 | 2.5481 | 101.21 | 20.81 | 17.727 | 64.9389 | 104.3957 | 61.7767 | 0.00431 | 1.0541 | 0.025797 | 24.4726 | -0.006518 | 0.14078 | -0.000109 | -0.008589 | -0.008589 | -0.68185
2019-03-19 15:39 | 1553024340.7608 | -189 | 0.20949 | 1.9e-05 | 1e-06 | 100.4494 | 100.6246 | 128.8694 | 21.2612 | 0.001809 | 0.001651 | -152.3094 | 0.23699 | 8.4051 | 0.003362 | 0.004949 | 0.004749 | 4.2e-05 | 0.001186 | 75.0373 | 1.3961 | 2.5482 | 101.21 | 20.81 | 17.722 | 65.351 | 102.7389 | 61.5619 | 0.004309 | 1.0548 | 0.025813 | 24.4722 | -0.015587 | 0.14098 | -6.3e-05 | -0.004956 | -0.004956 | -1.1834
2019-03-19 15:39 | 1553024341.7609 | -188 | 0.20949 | 2.5e-05 | 0 | 100.4335 | 100.5475 | 128.8853 | 21.2612 | 0.001804 | 0.001686 | -152.3288 | 0.23701 | 8.4048 | 0.003386 | 0.005088 | 0.00492 | 4.6e-05 | 0.001185 | 75.0322 | 1.3687 | 2.5481 | 101.21 | 20.81 | 17.719 | 65.2324 | 103.3156 | 62.3908 | 0.004308 | 1.0522 | 0.025782 | 24.504 | -0.010151 | 0.141 | -0.000107 | -0.00836 | -0.00836 | -0.70163
2019-03-19 15:39 | 1553024342.7623 | -187 | 0.20948 | 2e-05 | 0 | 100.4236 | 100.6503 | 128.9834 | 21.2612 | 0.001781 | 0.001709 | -152.3165 | 0.23647 | 8.4028 | 0.003337 | 0.003925 | 0.003691 | 4.3e-05 | 0.001179 | 75.0618 | 1.3431 | 2.5482 | 101.21 | 20.81 | 17.717 | 64.675 | 102.1387 | 62.947 | 0.004308 | 1.0504 | 0.025771 | 24.5336 | -0.020169 | 0.14068 | 0 | 1.5e-05 | 1.5e-05 | 399.4946
2019-03-19 15:39 | 1553024343.761 | -186 | 0.20951 | 1.7e-05 | 0 | 100.4608 | 100.6551 | 128.9093 | 21.2612 | 0.001751 | 0.00168 | -152.317 | 0.23677 | 8.4051 | 0.003338 | 0.005105 | 0.004848 | 4.4e-05 | 0.001189 | 75.0233 | 1.3569 | 2.5481 | 101.21 | 20.81 | 17.713 | 63.956 | 102.1535 | 62.2487 | 0.004307 | 1.0526 | 0.02579 | 24.5011 | -0.017281 | 0.14085 | -0.000175 | -0.013617 | -0.013617 | -0.43032
2019-03-19 15:39 | 1553024344.7621 | -185 | 0.20952 | 2.1e-05 | 1e-06 | 100.4923 | 100.6828 | 129.1651 | 21.2612 | 0.00181 | 0.001671 | -152.3088 | 0.23729 | 8.4027 | 0.003364 | 0.004338 | 0.00422 | 4.3e-05 | 0.001177 | 75.0254 | 1.314 | 2.5478 | 101.21 | 20.81 | 17.711 | 65.3674 | 102.79 | 62.0385 | 0.004306 | 1.0533 | 0.025824 | 24.5177 | -0.016941 | 0.14117 | -0.00024 | -0.018713 | -0.018713 | -0.31381
2019-03-19 15:39 | 1553024345.7621 | -184 | 0.20952 | 1.7e-05 | 1e-06 | 100.4716 | 100.6725 | 128.9534 | 21.2612 | 0.001732 | 0.00163 | -152.3128 | 0.23677 | 8.4026 | 0.00332 | 0.005759 | 0.005554 | 4.1e-05 | 0.00118 | 75.0028 | 1.3516 | 2.5482 | 101.21 | 20.81 | 17.708 | 63.4975 | 101.7202 | 61.0375 | 0.004306 | 1.0564 | 0.025841 | 24.461 | -0.017871 | 0.14086 | -0.000228 | -0.017855 | -0.017855 | -0.32817
2019-03-19 15:39 | 1553024346.7621 | -183 | 0.20951 | 1.5e-05 | 1e-06 | 100.4647 | 100.5643 | 129.1865 | 21.2612 | 0.001724 | 0.001631 | -152.3102 | 0.23628 | 8.4019 | 0.003357 | 0.004586 | 0.004351 | 4.4e-05 | 0.001181 | 74.9398 | 1.3479 | 2.5481 | 101.21 | 20.81 | 17.706 | 63.3109 | 102.6186 | 61.0695 | 0.004305 | 1.0563 | 0.025863 | 24.4843 | -0.008865 | 0.14056 | -0.000206 | -0.016156 | -0.016156 | -0.36194
2019-03-19 15:39 | 1553024347.7621 | -182 | 0.20949 | 2.3e-05 | 1e-06 | 100.4674 | 100.5578 | 129.3045 | 21.2612 | 0.001813 | 0.001667 | -152.2946 | 0.23652 | 8.4012 | 0.003431 | 0.004757 | 0.004543 | 4.1e-05 | 0.00118 | 74.9831 | 1.3768 | 2.548 | 101.21 | 20.81 | 17.705 | 65.447 | 104.4262 | 61.9446 | 0.004305 | 1.0536 | 0.025841 | 24.5275 | -0.008046 | 0.1407 | -9.8e-05 | -0.007684 | -0.007684 | -0.76178
2019-03-19 15:39 | 1553024348.7606 | -181 | 0.2095 | 2.1e-05 | 0 | 100.4992 | 100.7075 | 129.2064 | 21.2612 | 0.001771 | 0.00169 | -152.3081 | 0.23686 | 8.4027 | 0.003366 | 0.00472 | 0.004542 | 4.4e-05 | 0.001175 | 74.9782 | 1.3101 | 2.548 | 101.21 | 20.81 | 17.702 | 64.4329 | 102.8465 | 62.4799 | 0.004304 | 1.0519 | 0.025811 | 24.5378 | -0.018517 | 0.14091 | -0.000152 | -0.011898 | -0.011898 | -0.49264
2019-03-19 15:39 | 1553024349.7603 | -180 | 0.20953 | 1.9e-05 | 0 | 100.561 | 100.5438 | 128.9464 | 21.2612 | 0.001748 | 0.001683 | -152.2831 | 0.23717 | 8.4027 | 0.003363 | 0.004258 | 0.003997 | 4.2e-05 | 0.001172 | 75.0617 | 1.3455 | 2.548 | 101.21 | 20.81 | 17.702 | 63.877 | 102.7585 | 62.3275 | 0.004304 | 1.0524 | 0.025791 | 24.5075 | 0.001536 | 0.14109 | -0.000297 | -0.023155 | -0.023155 | -0.25348
2019-03-19 15:39 | 1553024350.7618 | -179 | 0.20951 | 1.7e-05 | 0 | 100.4884 | 100.6102 | 128.7189 | 21.2612 | 0.001825 | 0.001644 | -152.2725 | 0.2365 | 8.4009 | 0.003304 | 0.004832 | 0.004569 | 4.3e-05 | 0.001172 | 75.0486 | 1.3102 | 2.5482 | 101.21 | 20.81 | 17.696 | 65.7405 | 101.3512 | 61.3743 | 0.004303 | 1.0554 | 0.025805 | 24.4511 | -0.010835 | 0.14069 | -0.000212 | -0.016536 | -0.016536 | -0.35395
2019-03-19 15:39 | 1553024351.7604 | -178 | 0.20949 | 1.9e-05 | 1e-06 | 100.4122 | 100.4444 | 128.4386 | 21.2612 | 0.001801 | 0.001716 | -152.2564 | 0.2369 | 8.403 | 0.003352 | 0.004329 | 0.004094 | 4.3e-05 | 0.001169 | 74.9996 | 1.3001 | 2.548 | 101.21 | 20.81 | 17.692 | 65.1653 | 102.4968 | 63.122 | 0.004302 | 1.0499 | 0.02571 | 24.4882 | -0.002874 | 0.14093 | -0.000102 | -0.007935 | -0.007935 | -0.73885
2019-03-19 15:39 | 1553024352.7615 | -177 | 0.2095 | 1.9e-05 | 0 | 100.4431 | 100.553 | 128.1178 | 21.2612 | 0.001786 | 0.001688 | -152.2615 | 0.23722 | 8.4024 | 0.003342 | 0.00456 | 0.004316 | 4.6e-05 | 0.001166 | 74.9422 | 1.3361 | 2.548 | 101.21 | 20.81 | 17.682 | 64.8015 | 102.2714 | 62.451 | 0.004299 | 1.052 | 0.025703 | 24.4331 | -0.009783 | 0.14112 | -0.000154 | -0.011958 | -0.011958 | -0.49093
2019-03-19 15:39 | 1553024353.7618 | -176 | 0.20953 | 2.4e-05 | 0 | 100.4386 | 100.4184 | 127.9941 | 21.2612 | 0.00176 | 0.001686 | -152.251 | 0.23682 | 8.4024 | 0.003429 | 0.004676 | 0.004507 | 4e-05 | 0.001162 | 74.9644 | 1.3594 | 2.5481 | 101.21 | 20.81 | 17.677 | 64.1754 | 104.3792 | 62.4002 | 0.004298 | 1.0521 | 0.025693 | 24.4195 | 0.001797 | 0.14088 | -0.000332 | -0.02572 | -0.02572 | -0.22787
2019-03-19 15:39 | 1553024354.7605 | -175 | 0.20949 | 1.8e-05 | 1e-06 | 100.4461 | 100.5948 | 127.9948 | 21.2612 | 0.001764 | 0.001644 | -152.239 | 0.23711 | 8.4051 | 0.003466 | 0.004823 | 0.004724 | 4.2e-05 | 0.001156 | 74.9694 | 1.363 | 2.5479 | 101.21 | 20.81 | 17.672 | 64.2808 | 105.2749 | 61.3868 | 0.004297 | 1.0553 | 0.025732 | 24.3826 | -0.013232 | 0.14106 | -0.000102 | -0.007996 | -0.007996 | -0.73389
2019-03-19 15:39 | 1553024355.7606 | -174 | 0.20949 | 1.3e-05 | 1e-06 | 100.4245 | 100.6403 | 128.0215 | 21.2612 | 0.001797 | 0.001651 | -152.2469 | 0.2364 | 8.4037 | 0.003464 | 0.004537 | 0.004242 | 4.3e-05 | 0.001165 | 75.0494 | 1.3609 | 2.5481 | 101.21 | 20.81 | 17.66 | 65.0739 | 105.2135 | 61.5556 | 0.004294 | 1.0548 | 0.025728 | 24.3913 | -0.019206 | 0.14064 | -6.1e-05 | -0.004809 | -0.004809 | -1.2165
2019-03-19 15:39 | 1553024356.7617 | -173 | 0.20953 | 1.8e-05 | 0 | 100.4116 | 100.4119 | 127.9333 | 21.2612 | 0.001767 | 0.001708 | -152.2339 | 0.23636 | 8.4037 | 0.003446 | 0.004065 | 0.003783 | 4.1e-05 | 0.001168 | 74.9475 | 1.3718 | 2.5481 | 101.21 | 20.81 | 17.655 | 64.3463 | 104.7801 | 62.9313 | 0.004293 | 1.0505 | 0.025666 | 24.433 | -2.9e-05 | 0.14061 | -0.000314 | -0.024338 | -0.024338 | -0.24033
2019-03-19 15:39 | 1553024357.7621 | -172 | 0.20952 | 2e-05 | 1e-06 | 100.447 | 100.5342 | 127.6235 | 21.2612 | 0.001788 | 0.00163 | -152.2506 | 0.23682 | 8.4032 | 0.003448 | 0.00505 | 0.004833 | 4.3e-05 | 0.001168 | 74.9052 | 1.3467 | 2.5481 | 101.21 | 20.81 | 17.64 | 64.8576 | 104.8242 | 61.0459 | 0.004289 | 1.0564 | 0.025707 | 24.3348 | -0.00776 | 0.14089 | -0.000234 | -0.018208 | -0.018208 | -0.32189
2019-03-19 15:39 | 1553024358.7607 | -171 | 0.20949 | 2.1e-05 | 0 | 100.3777 | 100.4468 | 127.6577 | 21.2612 | 0.001813 | 0.001681 | -152.2629 | 0.23681 | 8.4005 | 0.003475 | 0.005096 | 0.004821 | 4.5e-05 | 0.001167 | 74.982 | 1.3511 | 2.5481 | 101.21 | 20.81 | 17.633 | 65.4516 | 105.4739 | 62.2752 | 0.004287 | 1.0525 | 0.025664 | 24.3828 | -0.006162 | 0.14088 | -7.7e-05 | -0.005955 | -0.005955 | -0.98417
2019-03-19 15:39 | 1553024359.7612 | -170 | 0.20954 | 2.4e-05 | 1e-06 | 100.4344 | 100.4157 | 127.8763 | 21.2612 | 0.001804 | 0.001692 | -152.2881 | 0.23608 | 8.4047 | 0.003386 | 0.004858 | 0.00477 | 4.4e-05 | 0.001166 | 74.9433 | 1.2941 | 2.5479 | 101.21 | 20.81 | 17.625 | 65.2283 | 103.3296 | 62.5422 | 0.004285 | 1.0517 | 0.025675 | 24.4134 | 0.001667 | 0.14044 | -0.000371 | -0.028812 | -0.028812 | -0.20278
2019-03-19 15:39 | 1553024360.7608 | -169 | 0.20949 | 1.8e-05 | 1e-06 | 100.3828 | 100.3027 | 127.8582 | 21.2612 | 0.001785 | 0.001629 | -152.3097 | 0.2366 | 8.4037 | 0.003416 | 0.004275 | 0.004021 | 4.1e-05 | 0.001161 | 74.984 | 1.2723 | 2.5478 | 101.21 | 20.81 | 17.611 | 64.771 | 104.0482 | 61.0163 | 0.004282 | 1.0565 | 0.025732 | 24.3561 | 0.007142 | 0.14075 | -7.1e-05 | -0.005603 | -0.005603 | -1.045
2019-03-19 15:39 | 1553024361.7619 | -168 | 0.20952 | 2.2e-05 | 1e-06 | 100.4848 | 100.5421 | 128.137 | 21.2612 | 0.001823 | 0.001744 | -152.315 | 0.23641 | 8.4029 | 0.003384 | 0.004692 | 0.00444 | 4.5e-05 | 0.001157 | 74.9454 | 1.3242 | 2.548 | 101.21 | 20.81 | 17.605 | 65.7015 | 103.2865 | 63.7904 | 0.004281 | 1.0478 | 0.025654 | 24.4837 | -0.005094 | 0.14064 | -0.000282 | -0.021933 | -0.021933 | -0.26674
2019-03-19 15:39 | 1553024362.7613 | -167 | 0.20951 | 1.6e-05 | 1e-06 | 100.4311 | 100.5821 | 128.2856 | 21.2612 | 0.001754 | 0.001688 | -152.3111 | 0.23664 | 8.402 | 0.00343 | 0.004901 | 0.004719 | 4.4e-05 | 0.001161 | 74.9725 | 1.3092 | 2.5481 | 101.21 | 20.81 | 17.591 | 64.0325 | 104.3925 | 62.4392 | 0.004277 | 1.052 | 0.02572 | 24.4487 | -0.013442 | 0.14078 | -0.000172 | -0.013437 | -0.013437 | -0.43582
2019-03-19 15:39 | 1553024363.7621 | -166 | 0.20955 | 2.5e-05 | 1e-06 | 100.3591 | 100.5794 | 128.1124 | 21.2612 | 0.001835 | 0.001694 | -152.304 | 0.23662 | 8.4052 | 0.003393 | 0.005182 | 0.00489 | 4.3e-05 | 0.001166 | 74.9933 | 1.3349 | 2.5481 | 101.21 | 20.81 | 17.586 | 65.9832 | 103.5064 | 62.5986 | 0.004276 | 1.0515 | 0.025697 | 24.438 | -0.019615 | 0.14077 | -0.000467 | -0.036288 | -0.036288 | -0.16137
2019-03-19 15:39 | 1553024364.7611 | -165 | 0.2095 | 2.1e-05 | 1e-06 | 100.4296 | 100.4482 | 128.2833 | 21.2612 | 0.001828 | 0.001664 | -152.3005 | 0.23649 | 8.4032 | 0.003554 | 0.004988 | 0.00478 | 4.5e-05 | 0.001169 | 74.9845 | 1.3309 | 2.548 | 101.21 | 20.81 | 17.581 | 65.8131 | 107.4063 | 61.8735 | 0.004275 | 1.0538 | 0.025742 | 24.4279 | -0.001659 | 0.14069 | -0.000109 | -0.008544 | -0.008544 | -0.68503
2019-03-19 15:39 | 1553024365.7602 | -164 | 0.20953 | 2e-05 | 1e-06 | 100.3915 | 100.3451 | 128.2657 | 21.2612 | 0.001809 | 0.001651 | -152.2814 | 0.23615 | 8.4019 | 0.003486 | 0.005214 | 0.005026 | 4.6e-05 | 0.001175 | 74.9734 | 1.3168 | 2.5481 | 101.21 | 20.81 | 17.57 | 65.3479 | 105.7438 | 61.5578 | 0.004272 | 1.0548 | 0.025752 | 24.4147 | 0.004131 | 0.14048 | -0.000292 | -0.022761 | -0.022761 | -0.25677
2019-03-19 15:39 | 1553024366.7602 | -163 | 0.20948 | 1.7e-05 | 1e-06 | 100.3396 | 100.4575 | 128.0995 | 21.2612 | 0.001774 | 0.001649 | -152.2686 | 0.23627 | 8.4029 | 0.003446 | 0.004877 | 0.004665 | 4.7e-05 | 0.001182 | 75.0274 | 1.3394 | 2.5481 | 101.21 | 20.82 | 17.566 | 64.5073 | 104.7716 | 61.4923 | 0.004274 | 1.055 | 0.025738 | 24.3965 | -0.010504 | 0.14055 | -2.2e-05 | -0.001812 | -0.001812 | -3.2271
2019-03-19 15:39 | 1553024367.7613 | -162 | 0.20952 | 2.3e-05 | 1e-06 | 100.3744 | 100.3496 | 128.43 | 21.2612 | 0.0018 | 0.001676 | -152.2554 | 0.23632 | 8.4031 | 0.003526 | 0.004983 | 0.004805 | 4.6e-05 | 0.00118 | 75.0082 | 1.3806 | 2.548 | 101.21 | 20.82 | 17.561 | 65.1364 | 106.7253 | 62.1493 | 0.004272 | 1.0529 | 0.025746 | 24.4519 | 0.00221 | 0.14059 | -0.000266 | -0.020758 | -0.020758 | -0.28175
2019-03-19 15:39 | 1553024368.7607 | -161 | 0.20949 | 2e-05 | 1e-06 | 100.4086 | 100.4492 | 127.9687 | 21.2612 | 0.001758 | 0.001677 | -152.2369 | 0.2365 | 8.4023 | 0.003454 | 0.005398 | 0.005255 | 4.8e-05 | 0.001175 | 74.9935 | 1.3709 | 2.548 | 101.21 | 20.82 | 17.553 | 64.1393 | 104.9769 | 62.19 | 0.004271 | 1.0528 | 0.025698 | 24.4094 | -0.003621 | 0.1407 | -5.3e-05 | -0.004173 | -0.004173 | -1.4025
2019-03-19 15:39 | 1553024369.7614 | -160 | 0.2095 | 2.2e-05 | 1e-06 | 100.3903 | 100.3516 | 128.2293 | 21.2612 | 0.001851 | 0.001651 | -152.2248 | 0.23634 | 8.4025 | 0.003424 | 0.004939 | 0.004727 | 4.5e-05 | 0.001173 | 75.0028 | 1.3388 | 2.5481 | 101.21 | 20.81 | 17.55 | 66.367 | 104.2471 | 61.5508 | 0.004267 | 1.0548 | 0.025749 | 24.411 | 0.003444 | 0.1406 | -0.000136 | -0.010636 | -0.010636 | -0.54989
2019-03-19 15:39 | 1553024370.7603 | -159 | 0.20951 | 2.3e-05 | 0 | 100.4239 | 100.277 | 128.1706 | 21.2612 | 0.001824 | 0.001625 | -152.2136 | 0.23618 | 8.4011 | 0.003395 | 0.004438 | 0.004255 | 4.9e-05 | 0.001175 | 74.9071 | 1.3339 | 2.5481 | 101.21 | 20.82 | 17.546 | 65.7241 | 103.5505 | 60.9194 | 0.004269 | 1.0568 | 0.025767 | 24.3823 | 0.013095 | 0.1405 | -0.000184 | -0.014345 | -0.014345 | -0.40745
2019-03-19 15:40 | 1553024371.7609 | -158 | 0.2095 | 1.8e-05 | 1e-06 | 100.4305 | 100.5397 | 127.9669 | 21.2612 | 0.001756 | 0.001677 | -152.2141 | 0.2362 | 8.4021 | 0.003446 | 0.005171 | 0.004972 | 4.5e-05 | 0.001177 | 74.9827 | 1.408 | 2.5481 | 101.21 | 20.82 | 17.545 | 64.084 | 104.7747 | 62.1743 | 0.004269 | 1.0528 | 0.025699 | 24.4087 | -0.009718 | 0.14052 | -0.000145 | -0.011312 | -0.011312 | -0.51673
2019-03-19 15:40 | 1553024372.7615 | -157 | 0.20947 | 2.2e-05 | 1e-06 | 100.3864 | 100.2503 | 127.8469 | 21.2612 | 0.001792 | 0.001704 | -152.2246 | 0.23676 | 8.4039 | 0.003348 | 0.004452 | 0.004249 | 4.5e-05 | 0.001177 | 75.0206 | 1.3702 | 2.548 | 101.21 | 20.82 | 17.545 | 64.938 | 102.4058 | 62.8328 | 0.004269 | 1.0508 | 0.025661 | 24.4212 | 0.012137 | 0.14085 | 1.5e-05 | 0.001151 | 0.001151 | 5.0923
2019-03-19 15:40 | 1553024373.7609 | -156 | 0.20949 | 2e-05 | 0 | 100.4113 | 100.4414 | 127.6223 | 21.2612 | 0.001789 | 0.001686 | -152.2419 | 0.23674 | 8.4059 | 0.003311 | 0.004732 | 0.004481 | 4.8e-05 | 0.001173 | 75.0848 | 1.327 | 2.5481 | 101.21 | 20.82 | 17.55 | 64.8795 | 101.5065 | 62.4057 | 0.00427 | 1.0521 | 0.025655 | 24.3842 | -0.002679 | 0.14084 | -6.5e-05 | -0.005089 | -0.005089 | -1.1512
2019-03-19 15:40 | 1553024374.7606 | -155 | 0.20948 | 2.1e-05 | 1e-06 | 100.36 | 100.5479 | 127.5613 | 21.2612 | 0.001792 | 0.001703 | -152.248 | 0.23638 | 8.4033 | 0.00333 | 0.004909 | 0.004679 | 4.6e-05 | 0.001178 | 75.0457 | 1.3529 | 2.5481 | 101.21 | 20.82 | 17.553 | 64.9411 | 101.9676 | 62.8064 | 0.004271 | 1.0509 | 0.025634 | 24.3929 | -0.016732 | 0.14062 | -2.3e-05 | -0.001852 | -0.001852 | -3.1589
2019-03-19 15:40 | 1553024375.7616 | -154 | 0.20947 | 2.1e-05 | 1e-06 | 100.3315 | 100.5897 | 127.7411 | 21.2612 | 0.001802 | 0.001664 | -152.2418 | 0.23661 | 8.4034 | 0.003415 | 0.004519 | 0.004322 | 4.9e-05 | 0.00118 | 75.071 | 1.3162 | 2.548 | 101.21 | 20.82 | 17.56 | 65.187 | 104.0215 | 61.8548 | 0.004272 | 1.0539 | 0.025688 | 24.3755 | -0.02299 | 0.14076 | 5.1e-05 | 0.003907 | 0.003907 | 1.4988
2019-03-19 15:40 | 1553024376.7611 | -153 | 0.20942 | 1.9e-05 | 1e-06 | 100.3477 | 100.3657 | 127.7655 | 21.2612 | 0.001847 | 0.001673 | -152.2647 | 0.23614 | 8.4008 | 0.003432 | 0.004871 | 0.004729 | 4.5e-05 | 0.001178 | 75.0439 | 1.3308 | 2.548 | 101.21 | 20.82 | 17.562 | 66.2805 | 104.4335 | 62.0924 | 0.004273 | 1.0531 | 0.025682 | 24.3865 | -0.001606 | 0.14048 | 0.000337 | 0.02606 | 0.02606 | 0.22425
2019-03-19 15:40 | 1553024377.7618 | -152 | 0.20945 | 2.2e-05 | 1e-06 | 100.2927 | 100.3116 | 127.7957 | 21.2612 | 0.001817 | 0.001709 | -152.25 | 0.23628 | 8.4014 | 0.003464 | 0.004903 | 0.004615 | 4.5e-05 | 0.001179 | 74.9766 | 1.3131 | 2.5481 | 101.21 | 20.82 | 17.564 | 65.5476 | 105.2278 | 62.9424 | 0.004273 | 1.0504 | 0.025652 | 24.4203 | -0.001678 | 0.14057 | 0.00016 | 0.012335 | 0.012335 | 0.47405
2019-03-19 15:40 | 1553024378.7622 | -151 | 0.20949 | 2.1e-05 | 1e-06 | 100.2723 | 100.2078 | 128.0591 | 21.2612 | 0.001767 | 0.001659 | -152.2608 | 0.23676 | 8.4026 | 0.003502 | 0.004755 | 0.004581 | 4.7e-05 | 0.001181 | 75.0846 | 1.3358 | 2.5481 | 101.21 | 20.82 | 17.569 | 64.3456 | 106.139 | 61.7478 | 0.004274 | 1.0542 | 0.025724 | 24.4019 | 0.00576 | 0.14085 | -8.2e-05 | -0.006442 | -0.006442 | -0.90948
2019-03-19 15:40 | 1553024379.761 | -150 | 0.20952 | 2.1e-05 | 2e-06 | 100.3026 | 100.3414 | 127.8744 | 21.2612 | 0.00178 | 0.001651 | -152.2666 | 0.23599 | 8.4024 | 0.003425 | 0.004705 | 0.0045 | 4.8e-05 | 0.001186 | 75.1062 | 1.3026 | 2.5479 | 101.21 | 20.82 | 17.571 | 64.6601 | 104.2699 | 61.5407 | 0.004275 | 1.0548 | 0.025714 | 24.3768 | -0.003452 | 0.14039 | -0.000223 | -0.017373 | -0.017373 | -0.33617
2019-03-19 15:40 | 1553024380.761 | -149 | 0.20952 | 2.1e-05 | 1e-06 | 100.2663 | 100.2458 | 127.9828 | 21.2612 | 0.001793 | 0.001679 | -152.2491 | 0.23666 | 8.4025 | 0.003466 | 0.00448 | 0.004209 | 5.2e-05 | 0.001184 | 75.0617 | 1.3034 | 2.5481 | 101.21 | 20.82 | 17.576 | 64.9615 | 105.2703 | 62.2189 | 0.004276 | 1.0527 | 0.025699 | 24.4118 | 0.001832 | 0.14079 | -0.00024 | -0.018729 | -0.018729 | -0.3127
2019-03-19 15:40 | 1553024381.7599 | -148 | 0.20951 | 2.8e-05 | 1e-06 | 100.2882 | 100.1763 | 128.1246 | 21.2612 | 0.001824 | 0.001646 | -152.2392 | 0.23684 | 8.4025 | 0.003518 | 0.004699 | 0.004502 | 4.7e-05 | 0.001181 | 75.057 | 1.323 | 2.5482 | 101.21 | 20.82 | 17.576 | 65.7061 | 106.5191 | 61.4212 | 0.004276 | 1.0552 | 0.025743 | 24.3963 | 0.009992 | 0.14089 | -0.000179 | -0.013967 | -0.013967 | -0.41965
2019-03-19 15:40 | 1553024382.7601 | -147 | 0.20951 | 2.6e-05 | 1e-06 | 100.2588 | 100.3622 | 128.2873 | 21.2612 | 0.001874 | 0.00168 | -152.2398 | 0.23667 | 8.4009 | 0.003541 | 0.004343 | 0.004149 | 4.7e-05 | 0.00118 | 74.9983 | 1.332 | 2.5482 | 101.21 | 20.82 | 17.575 | 66.9163 | 107.0923 | 62.2572 | 0.004276 | 1.0526 | 0.025728 | 24.4422 | -0.009222 | 0.14079 | -0.000216 | -0.016802 | -0.016802 | -0.34858
2019-03-19 15:40 | 1553024383.7612 | -146 | 0.20947 | 2.4e-05 | 1e-06 | 100.2939 | 100.1913 | 128.663 | 21.2612 | 0.001799 | 0.001678 | -152.2281 | 0.23678 | 8.4035 | 0.003473 | 0.005123 | 0.004924 | 4.6e-05 | 0.001181 | 75.059 | 1.3315 | 2.548 | 101.21 | 20.82 | 17.569 | 65.1052 | 105.4402 | 62.2112 | 0.004274 | 1.0527 | 0.025767 | 24.4763 | 0.009155 | 0.14086 | 6.2e-05 | 0.004728 | 0.004728 | 1.2394
2019-03-19 15:40 | 1553024384.7611 | -145 | 0.20944 | 2.2e-05 | 1e-06 | 100.2672 | 100.4277 | 128.3933 | 21.2612 | 0.001786 | 0.001676 | -152.2399 | 0.23641 | 8.4006 | 0.003421 | 0.004903 | 0.00472 | 4.3e-05 | 0.001176 | 75.0288 | 1.3236 | 2.5481 | 101.21 | 20.82 | 17.565 | 64.7974 | 104.1873 | 62.1541 | 0.004273 | 1.0529 | 0.025742 | 24.4486 | -0.0143 | 0.14064 | 0.000207 | 0.016038 | 0.016038 | 0.36481
2019-03-19 15:40 | 1553024385.7601 | -144 | 0.20948 | 2.5e-05 | 1e-06 | 100.2315 | 100.3557 | 128.2204 | 21.2612 | 0.001789 | 0.001703 | -152.2555 | 0.23653 | 8.403 | 0.003423 | 0.004654 | 0.004402 | 4.6e-05 | 0.001168 | 74.9788 | 1.3626 | 2.548 | 101.21 | 20.81 | 17.561 | 64.8756 | 104.2289 | 62.8095 | 0.00427 | 1.0509 | 0.0257 | 24.456 | -0.011078 | 0.14071 | 1e-06 | 1e-06 | 1e-06 | 4945.886
2019-03-19 15:40 | 1553024386.7617 | -143 | 0.20946 | 2.6e-05 | 2e-06 | 100.2759 | 100.2324 | 128.3708 | 21.2612 | 0.001803 | 0.001666 | -152.2518 | 0.23632 | 8.4047 | 0.003537 | 0.005107 | 0.004886 | 5e-05 | 0.001174 | 75.0227 | 1.3397 | 2.5481 | 101.21 | 20.82 | 17.552 | 65.2161 | 106.9843 | 61.9157 | 0.00427 | 1.0537 | 0.025749 | 24.4377 | 0.00388 | 0.14058 | 0.000123 | 0.009444 | 0.009444 | 0.61924
2019-03-19 15:40 | 1553024387.7603 | -142 | 0.2095 | 2.3e-05 | 1e-06 | 100.2447 | 100.4616 | 128.3368 | 21.2612 | 0.001823 | 0.001639 | -152.2548 | 0.23671 | 8.4045 | 0.003492 | 0.004785 | 0.004544 | 4.8e-05 | 0.001178 | 74.9775 | 1.3443 | 2.548 | 101.21 | 20.82 | 17.549 | 65.7022 | 105.8855 | 61.2688 | 0.00427 | 1.0557 | 0.025771 | 24.4109 | -0.019328 | 0.14082 | -0.000125 | -0.009817 | -0.009817 | -0.59674
2019-03-19 15:40 | 1553024388.7604 | -141 | 0.20947 | 3.1e-05 | 1e-06 | 100.2765 | 99.9745 | 128.0697 | 21.2612 | 0.001783 | 0.001615 | -152.2693 | 0.23654 | 8.4046 | 0.003579 | 0.004476 | 0.004306 | 4.8e-05 | 0.001185 | 74.888 | 1.3413 | 2.548 | 101.21 | 20.82 | 17.543 | 64.7226 | 108.0036 | 60.6896 | 0.004268 | 1.0575 | 0.025766 | 24.3644 | 0.026978 | 0.14072 | 3.2e-05 | 0.00239 | 0.00239 | 2.4489
2019-03-19 15:40 | 1553024389.7609 | -140 | 0.20946 | 2.4e-05 | 1e-06 | 100.2327 | 100.3609 | 128.191 | 21.2612 | 0.001825 | 0.00168 | -152.2579 | 0.23623 | 8.4029 | 0.003544 | 0.004852 | 0.004601 | 5e-05 | 0.001187 | 74.8858 | 1.3559 | 2.5482 | 101.21 | 20.82 | 17.541 | 65.7467 | 107.1685 | 62.2557 | 0.004268 | 1.0526 | 0.025718 | 24.433 | -0.011429 | 0.14053 | 0.0001 | 0.007708 | 0.007708 | 0.7585
2019-03-19 15:40 | 1553024390.7615 | -139 | 0.20948 | 2.6e-05 | 1e-06 | 100.2399 | 100.2674 | 128.3725 | 21.2612 | 0.00173 | 0.00169 | -152.2774 | 0.23676 | 8.4018 | 0.003551 | 0.005255 | 0.005004 | 4.7e-05 | 0.001193 | 74.8765 | 1.3175 | 2.548 | 101.21 | 20.82 | 17.538 | 63.4585 | 107.3347 | 62.4924 | 0.004267 | 1.0518 | 0.025727 | 24.4589 | -0.00245 | 0.14085 | -3.7e-05 | -0.002919 | -0.002919 | -2.0075
2019-03-19 15:40 | 1553024391.7606 | -138 | 0.20944 | 2e-05 | 1e-06 | 100.2513 | 100.1838 | 128.1295 | 21.2612 | 0.001775 | 0.001738 | -152.287 | 0.23623 | 8.4045 | 0.003542 | 0.005061 | 0.004854 | 4.9e-05 | 0.001197 | 74.9185 | 1.3128 | 2.5481 | 101.21 | 20.82 | 17.531 | 64.5313 | 107.1119 | 63.6359 | 0.004265 | 1.0483 | 0.025659 | 24.4773 | 0.006026 | 0.14053 | 0.000243 | 0.018773 | 0.018773 | 0.31141
2019-03-19 15:40 | 1553024392.7606 | -137 | 0.20948 | 2.5e-05 | 1e-06 | 100.2009 | 100.0948 | 127.9286 | 21.2612 | 0.001834 | 0.001757 | -152.2922 | 0.236 | 8.4024 | 0.003474 | 0.005163 | 0.004971 | 4.7e-05 | 0.0012 | 74.9493 | 1.2719 | 2.5482 | 101.21 | 20.82 | 17.528 | 65.9567 | 105.4519 | 64.1028 | 0.004264 | 1.0468 | 0.025621 | 24.4751 | 0.00948 | 0.14039 | -4e-05 | -0.003114 | -0.003114 | -1.8753
2019-03-19 15:40 | 1553024393.7607 | -136 | 0.20948 | 2.3e-05 | 1e-06 | 100.2165 | 100.2835 | 127.8643 | 21.2612 | 0.001741 | 0.001644 | -152.2907 | 0.23748 | 8.4013 | 0.003573 | 0.004606 | 0.004428 | 5e-05 | 0.001202 | 74.9306 | 1.305 | 2.548 | 101.21 | 20.82 | 17.523 | 63.7194 | 107.8576 | 61.393 | 0.004263 | 1.0553 | 0.025718 | 24.3704 | -0.005974 | 0.14128 | 1.3e-05 | 0.00091 | 0.00091 | 6.4606
2019-03-19 15:40 | 1553024394.8156 | -135 | 0.20948 | 2.5e-05 | 1e-06 | 100.1937 | 100.0828 | 127.968 | 21.2612 | 0.001855 | 0.001684 | -152.2887 | 0.23666 | 8.4016 | 0.003581 | 0.004956 | 0.00475 | 4.9e-05 | 0.0012 | 74.9936 | 1.333 | 2.5481 | 101.21 | 20.82 | 17.522 | 66.4616 | 108.0587 | 62.3532 | 0.004263 | 1.0523 | 0.025692 | 24.4153 | 0.009909 | 0.14079 | -1e-06 | -0.000155 | -0.000155 | -37.8087
2019-03-19 15:40 | 1553024395.7647 | -134 | 0.2095 | 2.4e-05 | 1e-06 | 100.1265 | 100.0472 | 127.8321 | 21.2612 | 0.001784 | 0.001658 | -152.2733 | 0.23669 | 8.4024 | 0.003434 | 0.004413 | 0.004158 | 4.9e-05 | 0.001203 | 74.9998 | 1.3205 | 2.5481 | 101.21 | 20.82 | 17.52 | 64.7633 | 104.4987 | 61.7095 | 0.004263 | 1.0543 | 0.025703 | 24.3789 | 0.007089 | 0.14081 | -0.000118 | -0.009222 | -0.009222 | -0.63519
2019-03-19 15:40 | 1553024396.7654 | -133 | 0.20948 | 2.6e-05 | 1e-06 | 100.1523 | 100.0475 | 127.9593 | 21.2612 | 0.001789 | 0.001699 | -152.2646 | 0.23752 | 8.4048 | 0.003518 | 0.005204 | 0.005025 | 4.8e-05 | 0.001207 | 75.0269 | 1.3436 | 2.5481 | 101.21 | 20.82 | 17.519 | 64.8646 | 106.5332 | 62.7048 | 0.004262 | 1.0512 | 0.025678 | 24.4272 | 0.009364 | 0.1413 | -4e-06 | -0.000375 | -0.000375 | -15.695
2019-03-19 15:40 | 1553024397.765 | -132 | 0.20946 | 2.2e-05 | 1e-06 | 100.1484 | 100.1944 | 128.0408 | 21.2612 | 0.001854 | 0.001685 | -152.2634 | 0.23605 | 8.4031 | 0.003498 | 0.004549 | 0.004263 | 4.7e-05 | 0.001206 | 74.9888 | 1.3454 | 2.5481 | 101.21 | 20.82 | 17.518 | 66.4491 | 106.0453 | 62.3706 | 0.004262 | 1.0522 | 0.025699 | 24.4229 | -0.004111 | 0.14043 | 9.9e-05 | 0.007608 | 0.007608 | 0.76785
2019-03-19 15:40 | 1553024398.7653 | -131 | 0.20947 | 2.1e-05 | 1e-06 | 100.0836 | 100.2139 | 128.2854 | 21.2612 | 0.001803 | 0.001731 | -152.2282 | 0.23676 | 8.402 | 0.00349 | 0.005348 | 0.005122 | 4.5e-05 | 0.001208 | 74.978 | 1.3244 | 2.5481 | 101.21 | 20.82 | 17.514 | 65.2161 | 105.8445 | 63.4671 | 0.004261 | 1.0488 | 0.025681 | 24.4861 | -0.011645 | 0.14085 | 2.5e-05 | 0.001861 | 0.001861 | 3.149
2019-03-19 15:40 | 1553024399.7652 | -130 | 0.20947 | 2.7e-05 | 1e-06 | 100.095 | 100.1225 | 128.011 | 21.2612 | 0.001803 | 0.001701 | -152.2322 | 0.23638 | 8.4044 | 0.003489 | 0.004766 | 0.004611 | 4.6e-05 | 0.001204 | 75.0848 | 1.3315 | 2.548 | 101.21 | 20.82 | 17.512 | 65.2036 | 105.836 | 62.7604 | 0.004261 | 1.051 | 0.025681 | 24.4342 | -0.00246 | 0.14062 | 3e-05 | 0.002282 | 0.002282 | 2.5632
2019-03-19 15:40 | 1553024400.7653 | -129 | 0.20949 | 2.5e-05 | 1e-06 | 100.1751 | 100.2499 | 127.972 | 21.2612 | 0.001807 | 0.001654 | -152.2181 | 0.23606 | 8.4031 | 0.003461 | 0.004937 | 0.004621 | 4.8e-05 | 0.001207 | 75.0568 | 1.3471 | 2.548 | 101.21 | 20.82 | 17.51 | 65.3104 | 105.1427 | 61.6204 | 0.00426 | 1.0546 | 0.02572 | 24.389 | -0.006677 | 0.14043 | -8.6e-05 | -0.006807 | -0.006807 | -0.85822
2019-03-19 15:40 | 1553024401.7653 | -128 | 0.20949 | 2.6e-05 | 1e-06 | 100.1457 | 100.0581 | 128.0211 | 21.2612 | 0.001804 | 0.001676 | -152.2139 | 0.23666 | 8.4027 | 0.003453 | 0.004609 | 0.004364 | 4.6e-05 | 0.001202 | 74.9946 | 1.3762 | 2.5481 | 101.21 | 20.82 | 17.507 | 65.2449 | 104.9468 | 62.151 | 0.004259 | 1.0529 | 0.025705 | 24.413 | 0.007825 | 0.14079 | -6.6e-05 | -0.005179 | -0.005179 | -1.131
2019-03-19 15:40 | 1553024402.7648 | -127 | 0.20947 | 2.6e-05 | 1e-06 | 100.1034 | 100.2705 | 127.8887 | 21.2612 | 0.001834 | 0.001646 | -152.2055 | 0.23617 | 8.4014 | 0.003384 | 0.004617 | 0.004459 | 4.9e-05 | 0.001203 | 75.0015 | 1.3857 | 2.5481 | 101.21 | 20.82 | 17.506 | 65.9473 | 103.2683 | 61.4227 | 0.004259 | 1.0552 | 0.02572 | 24.3738 | -0.014917 | 0.1405 | 2.5e-05 | 0.001847 | 0.001847 | 3.1637
2019-03-19 15:40 | 1553024403.7654 | -126 | 0.20948 | 2.7e-05 | 1e-06 | 100.0635 | 100.1338 | 127.7521 | 21.2612 | 0.001811 | 0.001749 | -152.2045 | 0.23633 | 8.4037 | 0.003421 | 0.004488 | 0.004319 | 4.7e-05 | 0.001204 | 74.9824 | 1.3789 | 2.5478 | 101.21 | 20.82 | 17.505 | 65.4025 | 104.1873 | 63.9138 | 0.004259 | 1.0474 | 0.025611 | 24.4514 | -0.006286 | 0.14059 | -3.1e-05 | -0.002472 | -0.002472 | -2.3662
2019-03-19 15:40 | 1553024404.7659 | -125 | 0.2095 | 2.3e-05 | 1e-06 | 100.1721 | 100.0996 | 127.5333 | 21.2612 | 0.001754 | 0.001682 | -152.2021 | 0.2365 | 8.4033 | 0.003406 | 0.004431 | 0.004106 | 4.8e-05 | 0.001194 | 74.9791 | 1.3904 | 2.5482 | 101.21 | 20.82 | 17.504 | 64.027 | 103.8131 | 62.2978 | 0.004259 | 1.0525 | 0.02565 | 24.3718 | 0.00648 | 0.14069 | -0.000159 | -0.012408 | -0.012408 | -0.4717
2019-03-19 15:40 | 1553024405.7656 | -124 | 0.20948 | 2.8e-05 | 1e-06 | 100.0812 | 100.1033 | 127.4657 | 21.2612 | 0.001821 | 0.001665 | -152.2145 | 0.2368 | 8.4032 | 0.00347 | 0.004921 | 0.004696 | 4.6e-05 | 0.001193 | 74.9684 | 1.3372 | 2.548 | 101.21 | 20.82 | 17.504 | 65.646 | 105.374 | 61.8837 | 0.004259 | 1.0538 | 0.025659 | 24.3503 | -0.00198 | 0.14088 | -3.9e-05 | -0.003106 | -0.003106 | -1.8868
2019-03-19 15:40 | 1553024406.7656 | -123 | 0.20947 | 2.5e-05 | 1e-06 | 100.1085 | 100.1228 | 127.1608 | 21.2612 | 0.001803 | 0.00166 | -152.2196 | 0.23673 | 8.4031 | 0.003496 | 0.004542 | 0.004325 | 4.9e-05 | 0.001193 | 75.0573 | 1.3263 | 2.5479 | 101.21 | 20.82 | 17.501 | 65.2113 | 105.9973 | 61.7625 | 0.004258 | 1.0541 | 0.025633 | 24.3167 | -0.001284 | 0.14083 | 5.9e-05 | 0.004459 | 0.004459 | 1.3139
2019-03-19 15:40 | 1553024407.765 | -122 | 0.20947 | 2.5e-05 | 1e-06 | 100.1463 | 100.1845 | 126.9268 | 21.2612 | 0.001784 | 0.001677 | -152.2157 | 0.23576 | 8.4006 | 0.003416 | 0.004777 | 0.004553 | 4.7e-05 | 0.00119 | 74.9897 | 1.3192 | 2.548 | 101.21 | 20.82 | 17.499 | 64.7553 | 104.0591 | 62.1861 | 0.004257 | 1.0528 | 0.025594 | 24.3097 | -0.003412 | 0.14025 | 7.8e-05 | 0.005991 | 0.005991 | 0.97392
2019-03-19 15:40 | 1553024408.7648 | -121 | 0.2095 | 2.4e-05 | 1e-06 | 100.0488 | 100.0472 | 126.9219 | 21.2612 | 0.001758 | 0.001636 | -152.2487 | 0.23658 | 8.4013 | 0.003503 | 0.004906 | 0.004695 | 4.9e-05 | 0.001185 | 75.0033 | 1.3609 | 2.5481 | 101.21 | 20.82 | 17.498 | 64.1393 | 106.161 | 61.1921 | 0.004257 | 1.0559 | 0.025631 | 24.2732 | 0.000143 | 0.14074 | -0.000103 | -0.008048 | -0.008048 | -0.72748
2019-03-19 15:40 | 1553024409.7657 | -120 | 0.20949 | 2.3e-05 | 1e-06 | 100.0944 | 100.1157 | 126.8531 | 21.2612 | 0.001775 | 0.001708 | -152.2528 | 0.23604 | 8.4038 | 0.003447 | 0.004781 | 0.004602 | 5.2e-05 | 0.001188 | 74.982 | 1.3233 | 2.548 | 101.21 | 20.82 | 17.499 | 64.533 | 104.8172 | 62.9313 | 0.004257 | 1.0505 | 0.025558 | 24.3296 | -0.001902 | 0.14042 | -7.7e-05 | -0.006028 | -0.006028 | -0.96907
2019-03-19 15:40 | 1553024410.7648 | -119 | 0.20947 | 2.8e-05 | 1e-06 | 100.1382 | 100.116 | 126.6932 | 21.2612 | 0.001775 | 0.001709 | -152.2478 | 0.2366 | 8.4015 | 0.003492 | 0.004599 | 0.004442 | 5e-05 | 0.001188 | 74.9983 | 1.3202 | 2.5479 | 101.21 | 20.82 | 17.499 | 64.5446 | 105.8911 | 62.9455 | 0.004257 | 1.0504 | 0.025541 | 24.3148 | 0.001982 | 0.14075 | 3.6e-05 | 0.002757 | 0.002757 | 2.1242
2019-03-19 15:40 | 1553024411.7649 | -118 | 0.20948 | 2.4e-05 | 1e-06 | 100.0401 | 100.08 | 126.9521 | 21.2612 | 0.001773 | 0.00168 | -152.2551 | 0.23637 | 8.4013 | 0.003409 | 0.00546 | 0.005211 | 5e-05 | 0.001191 | 74.9854 | 1.3433 | 2.5479 | 101.21 | 20.82 | 17.5 | 64.4798 | 103.8752 | 62.2401 | 0.004258 | 1.0526 | 0.025594 | 24.3141 | -0.003574 | 0.14062 | -3.5e-05 | -0.002755 | -0.002755 | -2.123
2019-03-19 15:40 | 1553024412.7639 | -117 | 0.20947 | 2.5e-05 | 1e-06 | 100.0548 | 100.0345 | 127.3465 | 21.2612 | 0.001861 | 0.001713 | -152.2525 | 0.23648 | 8.4013 | 0.003489 | 0.005108 | 0.004956 | 4.8e-05 | 0.001197 | 75.033 | 1.3337 | 2.548 | 101.21 | 20.82 | 17.501 | 66.5958 | 105.8297 | 63.0555 | 0.004258 | 1.0501 | 0.025603 | 24.3814 | 0.001813 | 0.14068 | 7.1e-05 | 0.00543 | 0.00543 | 1.0778
2019-03-19 15:40 | 1553024413.7653 | -116 | 0.20947 | 2.3e-05 | 0 | 100.0148 | 99.8512 | 127.2338 | 21.2612 | 0.001775 | 0.001657 | -152.2654 | 0.23587 | 8.4048 | 0.003487 | 0.004599 | 0.004442 | 5e-05 | 0.001192 | 75.0223 | 1.3195 | 2.548 | 101.21 | 20.82 | 17.5 | 64.5345 | 105.7785 | 61.697 | 0.004258 | 1.0543 | 0.025643 | 24.3213 | 0.014643 | 0.14032 | 5e-05 | 0.003887 | 0.003887 | 1.5017
2019-03-19 15:40 | 1553024414.7641 | -115 | 0.20949 | 3.1e-05 | 1e-06 | 100.0746 | 100.0184 | 127.4054 | 21.2612 | 0.001846 | 0.001715 | -152.273 | 0.23575 | 8.3988 | 0.003543 | 0.004645 | 0.004348 | 4.9e-05 | 0.001193 | 75.033 | 1.3275 | 2.548 | 101.21 | 20.82 | 17.502 | 66.2453 | 107.1228 | 63.0931 | 0.004258 | 1.05 | 0.025607 | 24.3884 | 0.005023 | 0.14025 | -6.3e-05 | -0.004973 | -0.004973 | -1.1733
2019-03-19 15:40 | 1553024415.7651 | -114 | 0.2095 | 3.1e-05 | 0 | 100.0286 | 100.0605 | 127.546 | 21.2612 | 0.001825 | 0.001676 | -152.267 | 0.23572 | 8.4028 | 0.003486 | 0.004891 | 0.004635 | 5.4e-05 | 0.001193 | 74.9903 | 1.3226 | 2.5479 | 101.21 | 20.82 | 17.504 | 65.7412 | 105.7485 | 62.1486 | 0.004259 | 1.0529 | 0.025657 | 24.3676 | -0.002849 | 0.14023 | -0.000132 | -0.010251 | -0.010251 | -0.56911
2019-03-19 15:40 | 1553024416.7655 | -113 | 0.20952 | 2.9e-05 | 1e-06 | 100.1046 | 100.2376 | 127.6803 | 21.2612 | 0.001838 | 0.001729 | -152.2913 | 0.23637 | 8.4012 | 0.003421 | 0.004761 | 0.004597 | 5.2e-05 | 0.001197 | 75.0033 | 1.3198 | 2.5481 | 101.21 | 20.82 | 17.506 | 66.0486 | 104.1708 | 63.439 | 0.004259 | 1.0489 | 0.025622 | 24.4273 | -0.011876 | 0.14062 | -0.000236 | -0.018255 | -0.018255 | -0.32044
2019-03-19 15:40 | 1553024417.7652 | -112 | 0.20948 | 2.8e-05 | 1e-06 | 100.071 | 100.0335 | 127.8049 | 21.2612 | 0.001802 | 0.001705 | -152.3044 | 0.23626 | 8.4067 | 0.003491 | 0.004621 | 0.004374 | 5.2e-05 | 0.001197 | 74.996 | 1.3283 | 2.5479 | 101.21 | 20.82 | 17.513 | 65.1973 | 105.861 | 62.8485 | 0.004261 | 1.0507 | 0.025657 | 24.4177 | 0.003353 | 0.14055 | -1.1e-05 | -0.000886 | -0.000886 | -6.5961
2019-03-19 15:40 | 1553024418.7657 | -111 | 0.2095 | 2.8e-05 | 0 | 100.053 | 100.0109 | 127.8861 | 21.2612 | 0.001807 | 0.001676 | -152.2956 | 0.23552 | 8.404 | 0.003424 | 0.004749 | 0.004531 | 5e-05 | 0.001201 | 75.0188 | 1.327 | 2.548 | 101.21 | 20.82 | 17.517 | 65.3048 | 104.2587 | 62.1596 | 0.004262 | 1.0529 | 0.025691 | 24.4004 | 0.003765 | 0.14011 | -0.000106 | -0.008217 | -0.008217 | -0.7093
2019-03-19 15:40 | 1553024419.7645 | -110 | 0.20947 | 2.6e-05 | 0 | 100.0313 | 100.0954 | 127.773 | 21.2612 | 0.001775 | 0.00168 | -152.2683 | 0.23699 | 8.4039 | 0.003438 | 0.004626 | 0.004425 | 5.2e-05 | 0.001195 | 75.0258 | 1.3176 | 2.5481 | 101.21 | 20.82 | 17.528 | 64.5369 | 104.5908 | 62.2415 | 0.004264 | 1.0526 | 0.025677 | 24.3926 | -0.005729 | 0.14098 | 2.6e-05 | 0.002011 | 0.002011 | 2.917
2019-03-19 15:40 | 1553024420.7644 | -109 | 0.20949 | 2.5e-05 | 0 | 100.0307 | 100.0718 | 128.111 | 21.2612 | 0.001834 | 0.00168 | -152.2707 | 0.23559 | 8.4002 | 0.003446 | 0.004761 | 0.004512 | 5.2e-05 | 0.001188 | 75.0669 | 1.3317 | 2.5481 | 101.21 | 20.82 | 17.532 | 65.9675 | 104.7779 | 62.2564 | 0.004265 | 1.0526 | 0.02571 | 24.4254 | -0.003671 | 0.14016 | -6.1e-05 | -0.00479 | -0.00479 | -1.2172
2019-03-19 15:40 | 1553024421.7657 | -108 | 0.20949 | 2.9e-05 | 1e-06 | 100.0398 | 100.0448 | 128.1225 | 21.2612 | 0.001858 | 0.001701 | -152.2723 | 0.23621 | 8.4002 | 0.003346 | 0.004767 | 0.004568 | 5.1e-05 | 0.00118 | 75.0382 | 1.3643 | 2.548 | 101.21 | 20.82 | 17.536 | 66.5443 | 102.3554 | 62.7486 | 0.004266 | 1.051 | 0.025692 | 24.4444 | -0.000447 | 0.14052 | -8.5e-05 | -0.00664 | -0.00664 | -0.88042
2019-03-19 15:40 | 1553024422.7654 | -107 | 0.20947 | 3.1e-05 | 1e-06 | 100.0326 | 99.9691 | 128.0732 | 21.2612 | 0.001841 | 0.001707 | -152.2736 | 0.236 | 8.4036 | 0.003392 | 0.004669 | 0.00441 | 5.3e-05 | 0.001176 | 75.0301 | 1.3473 | 2.5479 | 101.21 | 20.82 | 17.547 | 66.1212 | 103.4837 | 62.8892 | 0.004269 | 1.0506 | 0.025682 | 24.4448 | 0.005678 | 0.1404 | 7e-05 | 0.005371 | 0.005371 | 1.0874
2019-03-19 15:40 | 1553024423.7655 | -106 | 0.20946 | 2.3e-05 | 1e-06 | 100.1169 | 100.1787 | 128.042 | 21.2612 | 0.001777 | 0.001711 | -152.2706 | 0.23601 | 8.3994 | 0.003337 | 0.004904 | 0.004667 | 5.2e-05 | 0.001174 | 75.0148 | 1.3317 | 2.5481 | 101.21 | 20.82 | 17.553 | 64.5953 | 102.1293 | 63.0064 | 0.004271 | 1.0502 | 0.025674 | 24.4461 | -0.005519 | 0.1404 | 0.000138 | 0.010645 | 0.010645 | 0.54867
2019-03-19 15:40 | 1553024424.7647 | -105 | 0.20949 | 2.6e-05 | 1e-06 | 100.0845 | 100.2458 | 128.033 | 21.2612 | 0.001794 | 0.001693 | -152.2549 | 0.23617 | 8.4032 | 0.00344 | 0.00506 | 0.00485 | 4.7e-05 | 0.001174 | 74.9884 | 1.3645 | 2.548 | 101.21 | 20.82 | 17.568 | 65.0036 | 104.6402 | 62.5651 | 0.004274 | 1.0516 | 0.02569 | 24.4292 | -0.014407 | 0.1405 | -9.9e-05 | -0.00776 | -0.00776 | -0.75319
2019-03-19 15:40 | 1553024425.765 | -104 | 0.20947 | 3.3e-05 | 1e-06 | 100.0353 | 100.0109 | 128.1035 | 21.2612 | 0.001759 | 0.001638 | -152.27 | 0.23694 | 8.4059 | 0.003379 | 0.004495 | 0.004361 | 5.3e-05 | 0.001178 | 74.9756 | 1.3601 | 2.548 | 101.21 | 20.82 | 17.577 | 64.1497 | 103.156 | 61.2312 | 0.004276 | 1.0558 | 0.025749 | 24.3873 | 0.002182 | 0.14095 | 2.2e-05 | 0.00168 | 0.00168 | 3.4901
2019-03-19 15:40 | 1553024426.7637 | -103 | 0.20954 | 2.9e-05 | 1e-06 | 100.0395 | 100.0509 | 128.0798 | 21.2612 | 0.001817 | 0.001701 | -152.2555 | 0.2363 | 8.4027 | 0.003348 | 0.005084 | 0.004896 | 5.3e-05 | 0.00118 | 75.0008 | 1.3342 | 2.5481 | 101.21 | 20.82 | 17.586 | 65.5399 | 102.4041 | 62.7493 | 0.004279 | 1.051 | 0.025688 | 24.4404 | -0.001025 | 0.14057 | -0.000375 | -0.029111 | -0.029111 | -0.20088
2019-03-19 15:40 | 1553024427.7641 | -102 | 0.20953 | 3.3e-05 | 1e-06 | 100.0557 | 100.1129 | 127.8439 | 21.2612 | 0.001753 | 0.001688 | -152.2549 | 0.23646 | 8.4061 | 0.003293 | 0.00449 | 0.004302 | 5.3e-05 | 0.001183 | 72.3491 | 1.3074 | 2.5479 | 101.21 | 20.82 | 17.605 | 64.0082 | 101.0654 | 62.4322 | 0.004283 | 1.052 | 0.025676 | 24.4063 | -0.005117 | 0.14067 | -0.000307 | -0.023828 | -0.023828 | -0.24558
2019-03-19 15:40 | 1553024428.7648 | -101 | 0.20952 | 3e-05 | 1e-06 | 100.0416 | 100.0776 | 127.527 | 21.2612 | 0.001844 | 0.001665 | -152.2528 | 0.23627 | 8.4026 | 0.003207 | 0.004875 | 0.00471 | 5e-05 | 0.001181 | 51.3717 | 1.3346 | 2.5481 | 101.21 | 20.82 | 17.615 | 66.2054 | 98.9978 | 61.8971 | 0.004286 | 1.0537 | 0.025665 | 24.3566 | -0.003226 | 0.14055 | -0.000283 | -0.021942 | -0.021942 | -0.26648
2019-03-19 15:40 | 1553024429.7638 | -100 | 0.20952 | 2.8e-05 | 0 | 100.0172 | 99.9845 | 127.6746 | 21.2612 | 0.001765 | 0.00164 | -152.292 | 0.23681 | 8.402 | 0.00325 | 0.005134 | 0.004957 | 5.1e-05 | 0.001183 | 32.1335 | 1.3202 | 2.548 | 101.21 | 20.82 | 17.634 | 64.3073 | 100.0263 | 61.2851 | 0.00429 | 1.0556 | 0.025703 | 24.3484 | 0.002931 | 0.14088 | -0.00026 | -0.020215 | -0.020215 | -0.28991
2019-03-19 15:40 | 1553024430.7653 | -99 | 0.20952 | 2.5e-05 | 1e-06 | 100.0512 | 99.9088 | 127.3594 | 21.2612 | 0.001792 | 0.001706 | -152.2966 | 0.23652 | 8.3993 | 0.003281 | 0.0047 | 0.00447 | 4.9e-05 | 0.001181 | 8.0553 | 1.327 | 2.5481 | 101.21 | 20.82 | 17.645 | 64.9466 | 100.7812 | 62.8649 | 0.004293 | 1.0507 | 0.025611 | 24.3757 | 0.012737 | 0.1407 | -0.000253 | -0.019614 | -0.019614 | -0.29842
2019-03-19 15:40 | 1553024431.7649 | -98 | 0.20953 | 3e-05 | 1e-06 | 100.0332 | 100.1451 | 127.2261 | 21.2612 | 0.001833 | 0.001736 | -152.2908 | 0.23661 | 8.402 | 0.003303 | 0.005012 | 0.004752 | 5.3e-05 | 0.001188 | 1.4015 | 1.3299 | 2.548 | 101.21 | 20.82 | 17.655 | 65.9394 | 101.3117 | 63.5966 | 0.004295 | 1.0484 | 0.02557 | 24.3895 | -0.010005 | 0.14076 | -0.000291 | -0.022554 | -0.022554 | -0.25962
2019-03-19 15:41 | 1553024432.765 | -97 | 0.20952 | 3.1e-05 | 1e-06 | 100.0608 | 100.0441 | 127.3526 | 21.2612 | 0.001773 | 0.001703 | -152.3011 | 0.23616 | 8.4023 | 0.003338 | 0.004758 | 0.004534 | 5.3e-05 | 0.001188 | 0.73605 | 1.3365 | 2.548 | 101.21 | 20.82 | 17.676 | 64.4955 | 102.1576 | 62.8157 | 0.0043 | 1.0508 | 0.025612 | 24.3733 | 0.001492 | 0.14049 | -0.000247 | -0.019157 | -0.019157 | -0.30509
2019-03-19 15:41 | 1553024433.765 | -96 | 0.20953 | 2.5e-05 | 1e-06 | 100.0211 | 100.1129 | 127.6561 | 21.2612 | 0.001842 | 0.001679 | -152.3002 | 0.23595 | 8.4045 | 0.003302 | 0.00464 | 0.004364 | 5.5e-05 | 0.001192 | 0.6307 | 1.3678 | 2.5482 | 101.21 | 20.82 | 17.685 | 66.1518 | 101.2992 | 62.2307 | 0.004303 | 1.0527 | 0.025665 | 24.3811 | -0.008203 | 0.14037 | -0.000315 | -0.024465 | -0.024465 | -0.23868
2019-03-19 15:41 | 1553024434.7645 | -95 | 0.20952 | 3.3e-05 | 1e-06 | 100.01 | 99.8872 | 127.638 | 21.2612 | 0.001831 | 0.001661 | -152.3225 | 0.23673 | 8.4027 | 0.003307 | 0.0051 | 0.004908 | 5e-05 | 0.001201 | 0.56987 | 1.3447 | 2.5481 | 101.21 | 20.82 | 17.703 | 65.8956 | 101.4093 | 61.7938 | 0.004307 | 1.054 | 0.02568 | 24.3635 | 0.010992 | 0.14083 | -0.00026 | -0.020203 | -0.020203 | -0.28998
2019-03-19 15:41 | 1553024435.7642 | -94 | 0.20952 | 3.3e-05 | 1e-06 | 100.1292 | 100.0896 | 127.7725 | 21.2612 | 0.0019 | 0.001672 | -152.3116 | 0.23671 | 8.4018 | 0.003221 | 0.004669 | 0.004442 | 5.4e-05 | 0.001183 | 0.49593 | 1.338 | 2.5479 | 101.21 | 20.82 | 17.71 | 67.5405 | 99.3391 | 62.0604 | 0.004309 | 1.0532 | 0.025683 | 24.386 | 0.003535 | 0.14082 | -0.000244 | -0.018964 | -0.018964 | -0.3089
2019-03-19 15:41 | 1553024436.7645 | -93 | 0.20953 | 2.7e-05 | 1e-06 | 100.1091 | 100.0859 | 128.0788 | 21.2612 | 0.001793 | 0.001667 | -152.3203 | 0.23582 | 8.4022 | 0.003307 | 0.004673 | 0.004536 | 5.5e-05 | 0.00118 | 0.35778 | 1.3513 | 2.548 | 101.21 | 20.82 | 17.717 | 64.9678 | 101.4054 | 61.9369 | 0.00431 | 1.0536 | 0.025719 | 24.4107 | 0.002075 | 0.14029 | -0.000313 | -0.024393 | -0.024393 | -0.23925
2019-03-19 15:41 | 1553024437.7653 | -92 | 0.20952 | 3.1e-05 | 1e-06 | 100.0737 | 100.0468 | 128.5065 | 21.2612 | 0.001804 | 0.001643 | -152.3123 | 0.23653 | 8.4046 | 0.003285 | 0.004322 | 0.004124 | 5.4e-05 | 0.001173 | 0.39367 | 1.3676 | 2.5481 | 101.21 | 20.82 | 17.727 | 65.244 | 100.8707 | 61.3656 | 0.004313 | 1.0554 | 0.025784 | 24.4306 | 0.0024 | 0.14071 | -0.000224 | -0.017502 | -0.017502 | -0.33445
2019-03-19 15:41 | 1553024438.7656 | -91 | 0.20951 | 3.2e-05 | 1e-06 | 100.1091 | 100.0722 | 128.8623 | 21.2612 | 0.001802 | 0.001734 | -152.3079 | 0.23708 | 8.4046 | 0.00331 | 0.004548 | 0.004301 | 5.5e-05 | 0.001174 | 0.39672 | 1.3389 | 2.5479 | 101.21 | 20.82 | 17.731 | 65.1848 | 101.4761 | 63.5507 | 0.004314 | 1.0485 | 0.025736 | 24.5441 | 0.003299 | 0.14104 | -0.000218 | -0.016964 | -0.016964 | -0.34587
2019-03-19 15:41 | 1553024439.7654 | -90 | 0.20952 | 3.3e-05 | 1e-06 | 100.0344 | 100.0636 | 129.0194 | 21.2612 | 0.00181 | 0.001632 | -152.2954 | 0.23684 | 8.404 | 0.003296 | 0.004563 | 0.00434 | 5.3e-05 | 0.001182 | 0.40496 | 1.3228 | 2.5482 | 101.21 | 20.82 | 17.731 | 65.3893 | 101.1431 | 61.0881 | 0.004314 | 1.0563 | 0.025846 | 24.4691 | -0.002614 | 0.1409 | -0.00027 | -0.021153 | -0.021153 | -0.2771
2019-03-19 15:41 | 1553024440.7654 | -89 | 0.20947 | 3.2e-05 | 1e-06 | 100.1247 | 100.1098 | 129.048 | 21.2612 | 0.001837 | 0.001664 | -152.3005 | 0.23632 | 8.4018 | 0.00335 | 0.00492 | 0.004701 | 5.5e-05 | 0.001194 | 0.39295 | 1.3275 | 2.548 | 101.21 | 20.82 | 17.744 | 66.0315 | 102.4677 | 61.8721 | 0.004317 | 1.0538 | 0.025819 | 24.5005 | 0.001327 | 0.14059 | 4.4e-05 | 0.003352 | 0.003352 | 1.7448
2019-03-19 15:41 | 1553024441.7649 | -88 | 0.20947 | 3.3e-05 | 0 | 100.1394 | 100.1403 | 129.0738 | 21.2612 | 0.001839 | 0.001689 | -152.2959 | 0.23613 | 8.4018 | 0.003376 | 0.004941 | 0.004733 | 5.6e-05 | 0.001197 | 0.36674 | 1.3071 | 2.548 | 101.21 | 20.83 | 17.75 | 66.0847 | 103.0877 | 62.4604 | 0.004321 | 1.0519 | 0.025799 | 24.5245 | -8.3e-05 | 0.14047 | 6.7e-05 | 0.005154 | 0.005154 | 1.1339
2019-03-19 15:41 | 1553024442.7638 | -87 | 0.20946 | 3e-05 | 1e-06 | 100.1028 | 100.1311 | 129.2522 | 21.2612 | 0.001773 | 0.001667 | -152.2849 | 0.23648 | 8.4032 | 0.003426 | 0.004526 | 0.004277 | 5.5e-05 | 0.001192 | 0.36121 | 1.3444 | 2.5481 | 101.21 | 20.83 | 17.754 | 64.4854 | 104.2864 | 61.9362 | 0.004322 | 1.0536 | 0.025837 | 24.5222 | -0.002527 | 0.14068 | 8.4e-05 | 0.006464 | 0.006464 | 0.90539
2019-03-19 15:41 | 1553024443.7648 | -86 | 0.20944 | 3.3e-05 | 1e-06 | 100.0346 | 100.1359 | 129.1297 | 21.2612 | 0.001813 | 0.001649 | -152.3018 | 0.23608 | 8.4016 | 0.003376 | 0.004625 | 0.004414 | 5.4e-05 | 0.001195 | 0.37609 | 1.302 | 2.5479 | 101.21 | 20.83 | 17.76 | 65.4516 | 103.0831 | 61.4937 | 0.004324 | 1.055 | 0.025841 | 24.4944 | -0.009045 | 0.14044 | 0.000203 | 0.015773 | 0.015773 | 0.37041
2019-03-19 15:41 | 1553024444.7636 | -85 | 0.20947 | 3.6e-05 | 1e-06 | 100.1529 | 100.0773 | 129.1614 | 21.2612 | 0.001767 | 0.001673 | -152.2858 | 0.23645 | 8.4031 | 0.003382 | 0.005173 | 0.00494 | 4.9e-05 | 0.001201 | 0.4091 | 1.315 | 2.548 | 101.21 | 20.82 | 17.773 | 64.3417 | 103.2424 | 62.0869 | 0.004324 | 1.0531 | 0.025822 | 24.5191 | 0.006754 | 0.14067 | 2.7e-05 | 0.002071 | 0.002071 | 2.8258
2019-03-19 15:41 | 1553024445.7636 | -84 | 0.20946 | 3.2e-05 | 2e-06 | 100.0647 | 100.2092 | 129.0743 | 21.2612 | 0.001762 | 0.001723 | -152.2839 | 0.23594 | 8.4031 | 0.00332 | 0.004589 | 0.004413 | 5.2e-05 | 0.001206 | 0.35402 | 1.3229 | 2.548 | 101.21 | 20.82 | 17.78 | 64.2262 | 101.7287 | 63.2977 | 0.004326 | 1.0493 | 0.025766 | 24.5551 | -0.012906 | 0.14036 | 0.000103 | 0.007903 | 0.007903 | 0.73884
2019-03-19 15:41 | 1553024446.7647 | -83 | 0.20947 | 3.4e-05 | 1e-06 | 100.0911 | 100.0948 | 128.7001 | 21.2612 | 0.001765 | 0.001676 | -152.2649 | 0.23572 | 8.4028 | 0.003335 | 0.004431 | 0.004268 | 5.2e-05 | 0.001213 | 0.35026 | 1.31 | 2.5481 | 101.21 | 20.82 | 17.793 | 64.3042 | 102.0828 | 62.1628 | 0.004329 | 1.0529 | 0.025773 | 24.4781 | -0.00033 | 0.14023 | 1.7e-05 | 0.001309 | 0.001309 | 4.4581
2019-03-19 15:41 | 1553024447.7642 | -82 | 0.20945 | 3.8e-05 | 2e-06 | 100.143 | 100.3314 | 128.5677 | 21.2612 | 0.001847 | 0.001717 | -152.2815 | 0.23625 | 8.4032 | 0.003425 | 0.004848 | 0.004644 | 5.4e-05 | 0.001218 | 0.36674 | 1.3231 | 2.548 | 101.21 | 20.83 | 17.8 | 66.2781 | 104.265 | 63.1345 | 0.004333 | 1.0498 | 0.025722 | 24.5009 | -0.016815 | 0.14054 | 0.000144 | 0.011049 | 0.011049 | 0.52915
2019-03-19 15:41 | 1553024448.7638 | -81 | 0.20947 | 3.6e-05 | 1e-06 | 100.1775 | 100.0636 | 128.5178 | 21.2612 | 0.001781 | 0.001655 | -152.2815 | 0.2366 | 8.4038 | 0.003409 | 0.004859 | 0.00461 | 5.3e-05 | 0.001227 | 0.34827 | 1.2899 | 2.548 | 101.21 | 20.83 | 17.806 | 64.693 | 103.8863 | 61.6438 | 0.004335 | 1.0545 | 0.025774 | 24.4418 | 0.010176 | 0.14075 | 6.2e-05 | 0.004725 | 0.004725 | 1.2393
2019-03-19 15:41 | 1553024449.7643 | -80 | 0.20948 | 4.1e-05 | 1e-06 | 100.1049 | 100.3379 | 128.4856 | 21.2612 | 0.001816 | 0.001694 | -152.2857 | 0.23628 | 8.4008 | 0.003368 | 0.004942 | 0.004774 | 5.2e-05 | 0.001228 | 0.35258 | 1.3181 | 2.5481 | 101.21 | 20.83 | 17.819 | 65.5165 | 102.8888 | 62.5798 | 0.004338 | 1.0516 | 0.025735 | 24.4729 | -0.0208 | 0.14056 | -3.8e-05 | -0.003053 | -0.003053 | -1.9152
2019-03-19 15:41 | 1553024450.7651 | -79 | 0.20949 | 3.6e-05 | 1e-06 | 100.1268 | 100.1722 | 128.6321 | 21.2612 | 0.001824 | 0.001706 | -152.2739 | 0.2371 | 8.4001 | 0.003373 | 0.004398 | 0.004204 | 5.1e-05 | 0.001242 | 0.34634 | 1.3617 | 2.5481 | 101.21 | 20.83 | 17.826 | 65.7273 | 103.024 | 62.8836 | 0.00434 | 1.0506 | 0.025738 | 24.4979 | -0.004053 | 0.14105 | -5.5e-05 | -0.004345 | -0.004345 | -1.3505
2019-03-19 15:41 | 1553024451.7642 | -78 | 0.20945 | 4.3e-05 | 1e-06 | 100.1472 | 100.0715 | 128.6555 | 21.2612 | 0.001781 | 0.001692 | -152.2731 | 0.23613 | 8.4041 | 0.003359 | 0.00434 | 0.004125 | 5.4e-05 | 0.001245 | 0.323 | 1.393 | 2.548 | 101.21 | 20.82 | 17.839 | 64.6851 | 102.6808 | 62.5447 | 0.00434 | 1.0517 | 0.025753 | 24.4878 | 0.006765 | 0.14047 | 0.000143 | 0.011046 | 0.011046 | 0.52905
2019-03-19 15:41 | 1553024452.765 | -77 | 0.20947 | 3.8e-05 | 1e-06 | 100.2336 | 100.2372 | 128.0636 | 21.2612 | 0.001821 | 0.001661 | -152.2661 | 0.23643 | 8.4004 | 0.00337 | 0.004728 | 0.004493 | 5.3e-05 | 0.001236 | 0.32084 | 1.3852 | 2.5481 | 101.21 | 20.82 | 17.846 | 65.6522 | 102.9368 | 61.7813 | 0.004342 | 1.0541 | 0.025723 | 24.4036 | -0.000323 | 0.14066 | 5.1e-05 | 0.003899 | 0.003899 | 1.5007
2019-03-19 15:41 | 1553024453.7642 | -76 | 0.20952 | 4e-05 | 1e-06 | 100.2948 | 100.253 | 128.2126 | 21.2612 | 0.001776 | 0.001619 | -152.2473 | 0.23578 | 8.4044 | 0.003394 | 0.005272 | 0.005069 | 5.2e-05 | 0.00124 | 0.34882 | 1.3654 | 2.5479 | 101.21 | 20.82 | 17.853 | 64.5665 | 103.5207 | 60.7756 | 0.004344 | 1.0573 | 0.025777 | 24.3811 | 0.003733 | 0.14027 | -0.000255 | -0.019949 | -0.019949 | -0.29251
2019-03-19 15:41 | 1553024454.7645 | -75 | 0.2095 | 3.9e-05 | 2e-06 | 100.3303 | 100.4753 | 128.3544 | 21.2612 | 0.001787 | 0.001722 | -152.26 | 0.23625 | 8.4035 | 0.003335 | 0.004498 | 0.004311 | 5.5e-05 | 0.00124 | 0.3131 | 1.3302 | 2.548 | 101.21 | 20.82 | 17.865 | 64.8234 | 102.0947 | 63.2719 | 0.004346 | 1.0494 | 0.025695 | 24.4856 | -0.012919 | 0.14055 | -0.000155 | -0.012109 | -0.012109 | -0.48282
2019-03-19 15:41 | 1553024455.7654 | -74 | 0.2095 | 4e-05 | 2e-06 | 100.2672 | 100.2859 | 128.6163 | 21.2612 | 0.001763 | 0.001685 | -152.2484 | 0.23594 | 8.4013 | 0.003397 | 0.004322 | 0.004044 | 5.1e-05 | 0.001241 | 0.33035 | 1.3187 | 2.5481 | 101.21 | 20.82 | 17.87 | 64.2488 | 103.5961 | 62.3713 | 0.004348 | 1.0522 | 0.025756 | 24.4777 | -0.001663 | 0.14036 | -0.000175 | -0.013714 | -0.013714 | -0.42577
2019-03-19 15:41 | 1553024456.7643 | -73 | 0.2095 | 4.1e-05 | 1e-06 | 100.3516 | 100.3979 | 128.6342 | 21.2612 | 0.001793 | 0.001709 | -152.2371 | 0.23546 | 8.4057 | 0.003371 | 0.005281 | 0.005048 | 5.4e-05 | 0.001238 | 0.33539 | 1.3263 | 2.5481 | 101.21 | 20.82 | 17.875 | 64.9654 | 102.9581 | 62.9563 | 0.004349 | 1.0504 | 0.025736 | 24.5007 | -0.004126 | 0.14007 | -0.000163 | -0.012708 | -0.012708 | -0.45853
2019-03-19 15:41 | 1553024457.7647 | -72 | 0.20951 | 3.8e-05 | 1e-06 | 100.3621 | 100.4242 | 128.3588 | 21.2612 | 0.001854 | 0.001637 | -152.2347 | 0.23589 | 8.4047 | 0.003394 | 0.004735 | 0.004632 | 5.3e-05 | 0.001238 | 0.37697 | 1.334 | 2.5481 | 101.21 | 20.82 | 17.882 | 66.4466 | 103.5278 | 61.2109 | 0.004351 | 1.0559 | 0.025775 | 24.4109 | -0.005539 | 0.14033 | -0.000182 | -0.014208 | -0.014208 | -0.41089
2019-03-19 15:41 | 1553024458.7648 | -71 | 0.20947 | 3.4e-05 | 1e-06 | 100.4008 | 100.4458 | 128.2743 | 21.2612 | 0.001814 | 0.001684 | -152.2415 | 0.23638 | 8.4011 | 0.00336 | 0.005041 | 0.004796 | 4.8e-05 | 0.00123 | 0.34219 | 1.3286 | 2.548 | 101.21 | 20.82 | 17.886 | 65.4679 | 102.6886 | 62.347 | 0.004352 | 1.0523 | 0.025723 | 24.4443 | -0.004011 | 0.14062 | 5.3e-05 | 0.004055 | 0.004055 | 1.4428
2019-03-19 15:41 | 1553024459.7634 | -70 | 0.20951 | 3.7e-05 | 1e-06 | 100.4011 | 100.3951 | 128.4591 | 21.2612 | 0.001772 | 0.001691 | -152.2394 | 0.23658 | 8.4032 | 0.003379 | 0.004317 | 0.004166 | 5e-05 | 0.001234 | 0.347 | 1.3222 | 2.5481 | 101.21 | 20.82 | 17.894 | 64.4565 | 103.1497 | 62.5088 | 0.004354 | 1.0518 | 0.025735 | 24.4678 | 0.00053 | 0.14074 | -0.000215 | -0.016753 | -0.016753 | -0.34947
2019-03-19 15:41 | 1553024460.7648 | -69 | 0.20951 | 3.3e-05 | 1e-06 | 100.4974 | 100.6537 | 128.645 | 21.2612 | 0.001819 | 0.001732 | -152.261 | 0.23619 | 8.4022 | 0.00331 | 0.004913 | 0.004689 | 5.1e-05 | 0.001236 | 0.35579 | 1.3654 | 2.5481 | 101.21 | 20.82 | 17.898 | 65.5962 | 101.4957 | 63.4905 | 0.004354 | 1.0487 | 0.025716 | 24.5212 | -0.0139 | 0.14051 | -0.000207 | -0.016116 | -0.016116 | -0.3627
2019-03-19 15:41 | 1553024461.7646 | -68 | 0.20951 | 3.1e-05 | 1e-06 | 100.4605 | 100.5095 | 128.567 | 21.2612 | 0.001846 | 0.001684 | -152.2594 | 0.23656 | 8.4037 | 0.003421 | 0.004754 | 0.004585 | 5.2e-05 | 0.00124 | 0.36049 | 1.3956 | 2.5479 | 101.21 | 20.82 | 17.902 | 66.2367 | 104.1693 | 62.3494 | 0.004355 | 1.0523 | 0.025752 | 24.4722 | -0.004364 | 0.14073 | -0.000164 | -0.012861 | -0.012861 | -0.45518
2019-03-19 15:41 | 1553024462.7646 | -67 | 0.20949 | 3.4e-05 | 1e-06 | 100.6304 | 100.7862 | 128.7224 | 21.2612 | 0.00184 | 0.00165 | -152.2765 | 0.23589 | 8.4019 | 0.00335 | 0.004665 | 0.004485 | 4.9e-05 | 0.001236 | 0.35651 | 1.348 | 2.548 | 101.21 | 20.82 | 17.91 | 66.1104 | 102.445 | 61.5226 | 0.004357 | 1.0549 | 0.025799 | 24.4568 | -0.013846 | 0.14033 | -6.9e-05 | -0.005419 | -0.005419 | -1.0773
2019-03-19 15:41 | 1553024463.7637 | -66 | 0.20951 | 3e-05 | 1e-06 | 100.6568 | 100.8623 | 128.526 | 21.2612 | 0.001797 | 0.001682 | -152.2805 | 0.23576 | 8.4047 | 0.003383 | 0.004526 | 0.004318 | 4.8e-05 | 0.001236 | 0.3314 | 1.3103 | 2.548 | 101.21 | 20.82 | 17.913 | 65.07 | 103.2589 | 62.3056 | 0.004358 | 1.0524 | 0.02575 | 24.4667 | -0.018244 | 0.14026 | -0.000202 | -0.015766 | -0.015766 | -0.37008
2019-03-19 15:41 | 1553024464.7638 | -65 | 0.20951 | 3.3e-05 | 1e-06 | 100.6757 | 100.6568 | 128.6973 | 21.2612 | 0.001826 | 0.001659 | -152.2826 | 0.23668 | 8.4018 | 0.003353 | 0.004952 | 0.004729 | 5.2e-05 | 0.001232 | 0.33682 | 1.3227 | 2.5482 | 101.21 | 20.82 | 17.914 | 65.7677 | 102.5276 | 61.7517 | 0.004358 | 1.0542 | 0.025788 | 24.4628 | 0.001678 | 0.1408 | -0.000187 | -0.014622 | -0.014622 | -0.40059
2019-03-19 15:41 | 1553024465.7642 | -64 | 0.2095 | 2.7e-05 | 1e-06 | 100.7069 | 100.8311 | 128.2281 | 21.2612 | 0.001767 | 0.001672 | -152.2978 | 0.23598 | 8.4022 | 0.003309 | 0.004475 | 0.004266 | 4.9e-05 | 0.001232 | 0.3246 | 1.3481 | 2.5481 | 101.21 | 20.82 | 17.915 | 64.34 | 101.4611 | 62.0619 | 0.004359 | 1.0532 | 0.025729 | 24.4295 | -0.011027 | 0.14038 | -0.000113 | -0.008851 | -0.008851 | -0.65984
2019-03-19 15:41 | 1553024466.7641 | -63 | 0.20952 | 3.1e-05 | 1e-06 | 100.787 | 101.0839 | 128.015 | 21.2612 | 0.001749 | 0.001632 | -152.3234 | 0.23648 | 8.4058 | 0.003362 | 0.005445 | 0.005278 | 5.2e-05 | 0.00123 | 0.32189 | 1.3372 | 2.5481 | 101.21 | 20.82 | 17.915 | 63.9037 | 102.7389 | 61.0859 | 0.004359 | 1.0563 | 0.025745 | 24.3736 | -0.026307 | 0.14068 | -0.000271 | -0.021148 | -0.021148 | -0.27673
2019-03-19 15:41 | 1553024467.7638 | -62 | 0.20949 | 2.8e-05 | 0 | 100.8266 | 101.1133 | 128.1206 | 21.2612 | 0.001792 | 0.00168 | -152.3323 | 0.23583 | 8.4036 | 0.003342 | 0.005379 | 0.005173 | 5e-05 | 0.001232 | 0.33085 | 1.3393 | 2.548 | 101.21 | 20.82 | 17.915 | 64.9428 | 102.2581 | 62.2557 | 0.004359 | 1.0526 | 0.025711 | 24.4263 | -0.025398 | 0.1403 | -8.8e-05 | -0.006841 | -0.006841 | -0.85317
2019-03-19 15:41 | 1553024468.765 | -61 | 0.20951 | 3.2e-05 | 1e-06 | 100.8491 | 101.0846 | 128.1915 | 21.2612 | 0.001787 | 0.001649 | -152.3186 | 0.23623 | 8.4019 | 0.003402 | 0.004771 | 0.004633 | 5.1e-05 | 0.001227 | 0.34379 | 1.3835 | 2.5481 | 101.21 | 20.82 | 17.916 | 64.8193 | 103.7267 | 61.5101 | 0.004359 | 1.0549 | 0.025747 | 24.4059 | -0.020857 | 0.14053 | -0.000167 | -0.013042 | -0.013042 | -0.44827
2019-03-19 15:41 | 1553024469.7654 | -60 | 0.20949 | 2.4e-05 | 1e-06 | 100.9554 | 101.1692 | 128.0952 | 21.2612 | 0.001757 | 0.001686 | -152.3158 | 0.23604 | 8.401 | 0.003245 | 0.004969 | 0.004708 | 4.8e-05 | 0.001212 | 0.31614 | 1.3569 | 2.5481 | 101.21 | 20.82 | 17.919 | 64.0989 | 99.9172 | 62.3893 | 0.00436 | 1.0522 | 0.025703 | 24.4287 | -0.018921 | 0.14042 | -4.5e-05 | -0.003522 | -0.003522 | -1.6584
2019-03-19 15:41 | 1553024470.765 | -59 | 0.2095 | 2.7e-05 | 1e-06 | 100.9935 | 101.1192 | 127.6821 | 21.2612 | 0.001801 | 0.001651 | -152.3059 | 0.23709 | 8.4052 | 0.003289 | 0.00478 | 0.00459 | 5.1e-05 | 0.001202 | 0.34562 | 1.3303 | 2.5479 | 101.21 | 20.82 | 17.921 | 65.1518 | 100.9774 | 61.5571 | 0.00436 | 1.0548 | 0.025694 | 24.359 | -0.011123 | 0.14105 | -0.000143 | -0.011157 | -0.011157 | -0.52591
2019-03-19 15:41 | 1553024471.7651 | -58 | 0.20951 | 3.1e-05 | 1e-06 | 101.1045 | 101.2637 | 127.6108 | 21.2612 | 0.001852 | 0.001676 | -152.3121 | 0.23645 | 8.4034 | 0.003267 | 0.004522 | 0.004249 | 4.9e-05 | 0.001186 | 0.33627 | 1.3192 | 2.5482 | 101.21 | 20.82 | 17.922 | 66.3795 | 100.4461 | 62.1541 | 0.00436 | 1.0529 | 0.025664 | 24.3739 | -0.01407 | 0.14067 | -0.000196 | -0.015259 | -0.015259 | -0.38349
2019-03-19 15:41 | 1553024472.7655 | -57 | 0.20948 | 2.8e-05 | 1e-06 | 101.13 | 101.3099 | 127.6335 | 21.2612 | 0.001786 | 0.001703 | -152.3061 | 0.23632 | 8.4033 | 0.003353 | 0.004502 | 0.004292 | 5.3e-05 | 0.001186 | 0.34468 | 1.3562 | 2.548 | 101.21 | 20.82 | 17.925 | 64.8085 | 102.5244 | 62.7994 | 0.004361 | 1.0509 | 0.025641 | 24.3996 | -0.015896 | 0.14059 | 5e-06 | 0.000335 | 0.000335 | 17.4532
2019-03-19 15:41 | 1553024473.7635 | -56 | 0.20949 | 2.7e-05 | 1e-06 | 101.2663 | 101.4236 | 127.4542 | 21.2612 | 0.001805 | 0.001659 | -152.2846 | 0.23576 | 8.4019 | 0.003304 | 0.004211 | 0.003985 | 5e-05 | 0.00119 | 0.33356 | 1.3363 | 2.5481 | 101.21 | 20.83 | 17.925 | 65.2627 | 101.3369 | 61.7565 | 0.004364 | 1.0542 | 0.025663 | 24.3445 | -0.013888 | 0.14026 | -6.2e-05 | -0.004867 | -0.004867 | -1.1988
2019-03-19 15:41 | 1553024474.7641 | -55 | 0.20947 | 2.7e-05 | 1e-06 | 101.2516 | 101.4031 | 127.1554 | 21.2612 | 0.001823 | 0.001717 | -152.2674 | 0.23688 | 8.4023 | 0.003302 | 0.004389 | 0.004196 | 5e-05 | 0.001195 | 0.30143 | 1.3412 | 2.548 | 101.21 | 20.83 | 17.926 | 65.6984 | 101.29 | 63.1345 | 0.004364 | 1.0498 | 0.02558 | 24.366 | -0.013374 | 0.14092 | 6.7e-05 | 0.005104 | 0.005104 | 1.1486
2019-03-19 15:41 | 1553024475.7632 | -54 | 0.20955 | 2.2e-05 | 1e-06 | 101.3572 | 101.663 | 127.5744 | 21.2612 | 0.001829 | 0.001631 | -152.2685 | 0.23541 | 8.405 | 0.003342 | 0.004372 | 0.004176 | 5e-05 | 0.001198 | 0.33876 | 1.3739 | 2.5481 | 101.21 | 20.83 | 17.926 | 65.8482 | 102.2731 | 61.0678 | 0.004364 | 1.0563 | 0.025702 | 24.331 | -0.026947 | 0.14004 | -0.000391 | -0.030395 | -0.030395 | -0.19167
2019-03-19 15:41 | 1553024476.7637 | -53 | 0.20951 | 2.6e-05 | 1e-06 | 101.3362 | 101.5712 | 127.5559 | 21.2612 | 0.001809 | 0.001654 | -152.2711 | 0.23634 | 8.4045 | 0.003346 | 0.005183 | 0.004915 | 5e-05 | 0.001192 | 0.3298 | 1.3443 | 2.548 | 101.21 | 20.83 | 17.927 | 65.3667 | 102.3634 | 61.6337 | 0.004364 | 1.0545 | 0.025678 | 24.3498 | -0.020722 | 0.1406 | -0.000186 | -0.014475 | -0.014475 | -0.40408
2019-03-19 15:41 | 1553024477.7634 | -52 | 0.20952 | 2.7e-05 | 1e-06 | 101.4457 | 101.5041 | 127.4263 | 21.2612 | 0.001796 | 0.001665 | -152.2553 | 0.23589 | 8.4005 | 0.003287 | 0.004191 | 0.003979 | 4.7e-05 | 0.001183 | 0.32084 | 1.3184 | 2.548 | 101.21 | 20.83 | 17.934 | 65.0332 | 100.928 | 61.8791 | 0.004366 | 1.0538 | 0.025656 | 24.3463 | -0.005146 | 0.14033 | -0.00022 | -0.017103 | -0.017103 | -0.34134
2019-03-19 15:41 | 1553024478.7648 | -51 | 0.20951 | 2.2e-05 | 1e-06 | 101.5178 | 101.7082 | 127.5115 | 21.2612 | 0.001751 | 0.001668 | -152.2576 | 0.23606 | 8.4052 | 0.003351 | 0.004228 | 0.004093 | 5e-05 | 0.001178 | 0.36066 | 1.3798 | 2.5481 | 101.21 | 20.83 | 17.942 | 63.9521 | 102.4733 | 61.9651 | 0.004368 | 1.0535 | 0.025661 | 24.3576 | -0.016767 | 0.14043 | -0.000186 | -0.014511 | -0.014511 | -0.40258
2019-03-19 15:41 | 1553024479.7635 | -50 | 0.20951 | 2.6e-05 | 1e-06 | 101.4926 | 101.7281 | 127.2817 | 21.2612 | 0.001828 | 0.00167 | -152.2653 | 0.23646 | 8.406 | 0.003312 | 0.004408 | 0.00416 | 5.4e-05 | 0.001179 | 0.32211 | 1.3544 | 2.5479 | 101.21 | 20.83 | 17.951 | 65.8076 | 101.5349 | 62.0142 | 0.00437 | 1.0534 | 0.025636 | 24.3374 | -0.020735 | 0.14067 | -0.000171 | -0.01329 | -0.01329 | -0.44032
2019-03-19 15:41 | 1553024480.7649 | -49 | 0.2095 | 2.7e-05 | 0 | 101.5679 | 101.7343 | 127.6331 | 21.2612 | 0.001787 | 0.001673 | -152.2682 | 0.23651 | 8.4 | 0.003291 | 0.004705 | 0.004413 | 5.1e-05 | 0.001183 | 0.33196 | 1.3289 | 2.548 | 101.21 | 20.83 | 17.975 | 64.8193 | 101.0354 | 62.09 | 0.004376 | 1.0531 | 0.025668 | 24.3738 | -0.014641 | 0.1407 | -0.000141 | -0.010959 | -0.010959 | -0.53409
2019-03-19 15:41 | 1553024481.7636 | -48 | 0.20949 | 2.7e-05 | 1e-06 | 101.606 | 101.8463 | 127.5772 | 21.2612 | 0.001834 | 0.001632 | -152.2594 | 0.23669 | 8.4066 | 0.003235 | 0.004556 | 0.004371 | 5.2e-05 | 0.001187 | 0.33915 | 1.3122 | 2.5481 | 101.21 | 20.82 | 17.987 | 65.9505 | 99.6841 | 61.0852 | 0.004376 | 1.0563 | 0.025701 | 24.3319 | -0.021126 | 0.14081 | -6.3e-05 | -0.004932 | -0.004932 | -1.1877
2019-03-19 15:41 | 1553024482.7637 | -47 | 0.20954 | 1.9e-05 | 1e-06 | 101.726 | 101.9302 | 127.365 | 21.2612 | 0.001808 | 0.001718 | -152.2518 | 0.23621 | 8.401 | 0.003313 | 0.004453 | 0.004236 | 5.2e-05 | 0.001192 | 0.29734 | 1.3061 | 2.5481 | 101.21 | 20.82 | 18.007 | 65.3431 | 101.5639 | 63.161 | 0.004381 | 1.0498 | 0.0256 | 24.387 | -0.017932 | 0.14052 | -0.000366 | -0.028342 | -0.028342 | -0.20625
2019-03-19 15:41 | 1553024483.7641 | -46 | 0.20952 | 2.2e-05 | 1e-06 | 101.7254 | 101.8928 | 127.5967 | 21.2612 | 0.001833 | 0.001735 | -152.2555 | 0.23718 | 8.406 | 0.003292 | 0.004877 | 0.004734 | 5.2e-05 | 0.00119 | 0.30342 | 1.3518 | 2.5482 | 101.21 | 20.83 | 18.016 | 65.9231 | 101.0403 | 63.5858 | 0.004386 | 1.0484 | 0.025608 | 24.4246 | -0.014708 | 0.1411 | -0.000243 | -0.018829 | -0.018829 | -0.31174
2019-03-19 15:41 | 1553024484.7649 | -45 | 0.20951 | 2e-05 | 1e-06 | 101.7281 | 102.0247 | 127.6922 | 21.2612 | 0.00182 | 0.001706 | -152.2643 | 0.23626 | 8.4018 | 0.003271 | 0.00496 | 0.00469 | 5.1e-05 | 0.001193 | 0.33699 | 1.3329 | 2.5479 | 101.21 | 20.83 | 18.023 | 65.6289 | 100.5497 | 62.869 | 0.004388 | 1.0507 | 0.025644 | 24.4077 | -0.026038 | 0.14055 | -0.000192 | -0.014881 | -0.014881 | -0.39291
2019-03-19 15:41 | 1553024485.7639 | -44 | 0.20952 | 2.5e-05 | 1e-06 | 101.7062 | 101.8668 | 127.4207 | 21.2612 | 0.001843 | 0.001737 | -152.2592 | 0.23604 | 8.4021 | 0.003184 | 0.005015 | 0.004794 | 5.2e-05 | 0.001192 | 0.31022 | 1.3514 | 2.5482 | 101.21 | 20.82 | 18.035 | 66.1838 | 98.4342 | 63.6327 | 0.004388 | 1.0483 | 0.025588 | 24.4094 | -0.014111 | 0.14042 | -0.000232 | -0.017966 | -0.017966 | -0.32513
2019-03-19 15:41 | 1553024486.7638 | -43 | 0.20947 | 2.5e-05 | 1e-06 | 101.8146 | 101.9959 | 127.6526 | 21.2612 | 0.001778 | 0.001717 | -152.2424 | 0.23567 | 8.4041 | 0.003247 | 0.004832 | 0.004591 | 5.5e-05 | 0.001195 | 0.31116 | 1.4173 | 2.548 | 101.21 | 20.82 | 18.04 | 64.6211 | 99.9612 | 63.1453 | 0.004389 | 1.0498 | 0.02563 | 24.4139 | -0.015919 | 0.1402 | 4.8e-05 | 0.003659 | 0.003659 | 1.594
2019-03-19 15:41 | 1553024487.7649 | -42 | 0.2095 | 2.5e-05 | 1e-06 | 101.8224 | 102.075 | 127.7993 | 21.2612 | 0.001791 | 0.001656 | -152.2424 | 0.23646 | 8.4036 | 0.003198 | 0.004173 | 0.004024 | 5.3e-05 | 0.001195 | 0.30773 | 1.3869 | 2.5482 | 101.21 | 20.82 | 18.049 | 64.9341 | 98.7689 | 61.6768 | 0.004391 | 1.0544 | 0.025701 | 24.3746 | -0.022169 | 0.14067 | -9.5e-05 | -0.007413 | -0.007413 | -0.78942
2019-03-19 15:41 | 1553024488.764 | -41 | 0.20949 | 2.4e-05 | 1e-06 | 101.8725 | 102.2326 | 127.8087 | 21.2612 | 0.00184 | 0.001668 | -152.2574 | 0.23583 | 8.4033 | 0.003239 | 0.004403 | 0.004163 | 5.2e-05 | 0.001199 | 0.32693 | 1.3303 | 2.5481 | 101.21 | 20.82 | 18.05 | 66.0955 | 99.7806 | 61.9564 | 0.004391 | 1.0535 | 0.025691 | 24.3856 | -0.031562 | 0.1403 | -8.4e-05 | -0.006601 | -0.006601 | -0.88419
2019-03-19 15:41 | 1553024489.7634 | -40 | 0.20948 | 2.5e-05 | 1e-06 | 101.8773 | 102.2285 | 127.9866 | 21.2612 | 0.001833 | 0.00174 | -152.2616 | 0.23719 | 8.4061 | 0.003267 | 0.004803 | 0.004516 | 5.2e-05 | 0.001192 | 0.335 | 1.3244 | 2.5481 | 101.21 | 20.83 | 18.048 | 65.9418 | 100.4524 | 63.6919 | 0.004394 | 1.0481 | 0.025643 | 24.4657 | -0.030781 | 0.1411 | -1.6e-05 | -0.001336 | -0.001336 | -4.3933
2019-03-19 15:41 | 1553024490.7641 | -39 | 0.20949 | 2.7e-05 | 2e-06 | 101.9256 | 102.1898 | 127.842 | 21.2612 | 0.001848 | 0.001712 | -152.2526 | 0.2363 | 8.4014 | 0.003295 | 0.005039 | 0.004814 | 5.4e-05 | 0.001184 | 0.32621 | 1.3224 | 2.5481 | 101.21 | 20.83 | 18.037 | 66.2836 | 101.1141 | 63.0095 | 0.004391 | 1.0502 | 0.025654 | 24.4271 | -0.023153 | 0.14058 | -8.2e-05 | -0.006463 | -0.006463 | -0.90482
2019-03-19 15:42 | 1553024491.7635 | -38 | 0.20949 | 2.5e-05 | 1e-06 | 101.9925 | 102.3778 | 127.8903 | 21.2612 | 0.001797 | 0.001674 | -152.2514 | 0.23629 | 8.4033 | 0.003292 | 0.004327 | 0.004173 | 5.3e-05 | 0.001182 | 0.31885 | 1.345 | 2.548 | 101.21 | 20.83 | 18.029 | 65.0739 | 101.063 | 62.1103 | 0.004389 | 1.053 | 0.025693 | 24.399 | -0.033725 | 0.14057 | -3.8e-05 | -0.003077 | -0.003077 | -1.9008
2019-03-19 15:42 | 1553024492.7643 | -37 | 0.20952 | 2.8e-05 | 1e-06 | 102.0066 | 102.1833 | 127.9978 | 21.2612 | 0.001854 | 0.001725 | -152.2567 | 0.23655 | 8.4041 | 0.003352 | 0.00421 | 0.004034 | 5.5e-05 | 0.00118 | 0.29894 | 1.2932 | 2.5481 | 101.21 | 20.83 | 18.021 | 66.4404 | 102.4992 | 63.329 | 0.004387 | 1.0492 | 0.025658 | 24.4536 | -0.015476 | 0.14072 | -0.000209 | -0.016232 | -0.016232 | -0.36066
2019-03-19 15:42 | 1553024493.7639 | -36 | 0.20949 | 2.6e-05 | 1e-06 | 101.9787 | 102.1322 | 127.9802 | 21.2612 | 0.001851 | 0.001727 | -152.2742 | 0.23697 | 8.4015 | 0.003326 | 0.004208 | 0.003993 | 5.3e-05 | 0.00118 | 0.30685 | 1.3602 | 2.5479 | 101.21 | 20.83 | 18.003 | 66.3584 | 101.8851 | 63.3875 | 0.004383 | 1.0491 | 0.025654 | 24.4541 | -0.013455 | 0.14098 | -4.5e-05 | -0.003514 | -0.003514 | -1.669
2019-03-19 15:42 | 1553024494.7648 | -35 | 0.20948 | 2.3e-05 | 1e-06 | 102.0087 | 102.3394 | 128.0133 | 21.2612 | 0.001782 | 0.001741 | -152.2574 | 0.23678 | 8.403 | 0.003299 | 0.00476 | 0.004512 | 5.3e-05 | 0.001185 | 0.31813 | 1.382 | 2.5481 | 101.21 | 20.83 | 17.994 | 64.6976 | 101.2239 | 63.7256 | 0.004381 | 1.048 | 0.025644 | 24.4695 | -0.028952 | 0.14086 | 2.4e-05 | 0.001777 | 0.001777 | 3.2979
2019-03-19 15:42 | 1553024495.7634 | -34 | 0.2095 | 2.6e-05 | 1e-06 | 102.0582 | 102.2357 | 127.9105 | 21.2612 | 0.001789 | 0.001707 | -152.2524 | 0.2362 | 8.4027 | 0.003315 | 0.004894 | 0.004708 | 5.3e-05 | 0.001187 | 0.35723 | 1.3039 | 2.5482 | 101.21 | 20.83 | 17.982 | 64.8858 | 101.6188 | 62.8986 | 0.004378 | 1.0506 | 0.025665 | 24.4296 | -0.015536 | 0.14052 | -0.000126 | -0.009782 | -0.009782 | -0.59758
2019-03-19 15:42 | 1553024496.7647 | -33 | 0.20949 | 2.5e-05 | 1e-06 | 102.0561 | 102.2747 | 127.7758 | 21.2612 | 0.00184 | 0.001756 | -152.2673 | 0.23575 | 8.4026 | 0.003359 | 0.004628 | 0.004445 | 5.3e-05 | 0.001191 | 0.34324 | 1.2943 | 2.5481 | 101.21 | 20.83 | 17.978 | 66.1126 | 102.676 | 64.0708 | 0.004377 | 1.0469 | 0.025607 | 24.4593 | -0.019136 | 0.14025 | -8.5e-05 | -0.006643 | -0.006643 | -0.87833
2019-03-19 15:42 | 1553024497.7644 | -32 | 0.2095 | 2.2e-05 | 1e-06 | 102.0489 | 102.3275 | 127.6591 | 21.2612 | 0.00176 | 0.001723 | -152.2615 | 0.23649 | 8.4013 | 0.003319 | 0.005024 | 0.004798 | 5.2e-05 | 0.001197 | 0.33787 | 1.3353 | 2.548 | 101.21 | 20.83 | 17.975 | 64.1824 | 101.7004 | 63.2828 | 0.004376 | 1.0494 | 0.025625 | 24.4196 | -0.024379 | 0.14069 | -9.1e-05 | -0.007075 | -0.007075 | -0.82721
2019-03-19 15:42 | 1553024498.7639 | -31 | 0.20951 | 2.3e-05 | 1e-06 | 102.079 | 102.3025 | 127.8194 | 21.2612 | 0.001859 | 0.001636 | -152.2585 | 0.235 | 8.405 | 0.003334 | 0.00478 | 0.004525 | 5e-05 | 0.0012 | 0.29535 | 1.3156 | 2.5481 | 101.21 | 20.83 | 17.971 | 66.5669 | 102.0586 | 61.1914 | 0.004375 | 1.0559 | 0.025722 | 24.3588 | -0.019563 | 0.1398 | -0.000185 | -0.014446 | -0.014446 | -0.40258
2019-03-19 15:42 | 1553024499.7636 | -30 | 0.20952 | 2.5e-05 | 1e-06 | 102.0402 | 102.3299 | 127.9701 | 21.2612 | 0.001828 | 0.00167 | -152.255 | 0.23462 | 8.4057 | 0.003372 | 0.005031 | 0.004898 | 5e-05 | 0.001204 | 0.27831 | 1.366 | 2.548 | 101.21 | 20.83 | 17.971 | 65.8153 | 103.0005 | 62.0204 | 0.004375 | 1.0533 | 0.025705 | 24.4034 | -0.025351 | 0.13958 | -0.000212 | -0.016506 | -0.016506 | -0.35177
2019-03-19 15:42 | 1553024500.764 | -29 | 0.20949 | 2.5e-05 | 1e-06 | 101.9976 | 102.2559 | 127.6826 | 21.2612 | 0.00185 | 0.001674 | -152.2474 | 0.23603 | 8.4041 | 0.003335 | 0.004662 | 0.004422 | 5.1e-05 | 0.001215 | 0.30685 | 1.3105 | 2.5479 | 101.21 | 20.83 | 17.972 | 66.3507 | 102.0932 | 62.1141 | 0.004375 | 1.053 | 0.025672 | 24.3794 | -0.022619 | 0.14042 | -7.1e-05 | -0.00556 | -0.00556 | -1.0506
2019-03-19 15:42 | 1553024501.765 | -28 | 0.20949 | 2.4e-05 | 0 | 102.0009 | 102.4097 | 127.6056 | 21.2612 | 0.001802 | 0.001681 | -152.2639 | 0.23638 | 8.3999 | 0.003397 | 0.004565 | 0.00441 | 5.4e-05 | 0.001206 | 0.31653 | 1.3906 | 2.5481 | 101.21 | 20.83 | 17.972 | 65.1848 | 103.6024 | 62.2861 | 0.004375 | 1.0525 | 0.025658 | 24.3783 | -0.03577 | 0.14062 | -4.4e-05 | -0.003429 | -0.003429 | -1.7059
2019-03-19 15:42 | 1553024502.7647 | -27 | 0.20949 | 2.6e-05 | 1e-06 | 102.0492 | 102.2894 | 127.9427 | 21.2612 | 0.001854 | 0.001702 | -152.2687 | 0.2355 | 8.4024 | 0.00333 | 0.004755 | 0.004534 | 5.6e-05 | 0.001198 | 0.2855 | 1.3336 | 2.5481 | 101.21 | 20.83 | 17.972 | 66.4296 | 101.9809 | 62.7852 | 0.004375 | 1.0509 | 0.025673 | 24.4286 | -0.021029 | 0.1401 | -6.2e-05 | -0.004818 | -0.004818 | -1.2096
2019-03-19 15:42 | 1553024503.7652 | -26 | 0.20949 | 2.5e-05 | 1e-06 | 101.9511 | 102.2596 | 128.152 | 21.2612 | 0.00175 | 0.001676 | -152.2588 | 0.23639 | 8.4019 | 0.003323 | 0.004788 | 0.004535 | 5.4e-05 | 0.00119 | 0.32997 | 1.3809 | 2.548 | 101.21 | 20.83 | 17.969 | 63.9379 | 101.8081 | 62.1548 | 0.004374 | 1.0529 | 0.025718 | 24.4256 | -0.027028 | 0.14063 | -4.5e-05 | -0.003555 | -0.003555 | -1.6455
2019-03-19 15:42 | 1553024504.7638 | -25 | 0.20947 | 2.5e-05 | 1e-06 | 101.9163 | 102.15 | 127.9392 | 21.2612 | 0.001859 | 0.001684 | -152.2748 | 0.23602 | 8.4015 | 0.003371 | 0.004715 | 0.004558 | 4.9e-05 | 0.001183 | 0.30973 | 1.3621 | 2.5481 | 101.21 | 20.83 | 17.969 | 66.5575 | 102.9697 | 62.3462 | 0.004374 | 1.0523 | 0.025689 | 24.4123 | -0.020491 | 0.14041 | 3.8e-05 | 0.002917 | 0.002917 | 2.0024
2019-03-19 15:42 | 1553024505.7629 | -24 | 0.2095 | 2e-05 | 1e-06 | 101.9262 | 102.2829 | 128.0957 | 21.2612 | 0.001817 | 0.001697 | -152.2892 | 0.23537 | 8.4024 | 0.003432 | 0.005186 | 0.004991 | 5.4e-05 | 0.001191 | 0.33555 | 1.3322 | 2.5482 | 101.21 | 20.83 | 17.969 | 65.5531 | 104.4325 | 62.6494 | 0.004374 | 1.0514 | 0.025693 | 24.4382 | -0.031251 | 0.14002 | -8.3e-05 | -0.006532 | -0.006532 | -0.89174
2019-03-19 15:42 | 1553024506.764 | -23 | 0.20951 | 2.3e-05 | 1e-06 | 101.7806 | 101.9812 | 128.6924 | 21.2612 | 0.001854 | 0.001688 | -152.2817 | 0.2355 | 8.4019 | 0.003379 | 0.004601 | 0.004358 | 5.1e-05 | 0.001191 | 0.30613 | 1.3196 | 2.548 | 101.21 | 20.83 | 17.97 | 66.4264 | 103.1482 | 62.4346 | 0.004375 | 1.052 | 0.025761 | 24.4873 | -0.017607 | 0.1401 | -0.00016 | -0.012438 | -0.012438 | -0.46858
2019-03-19 15:42 | 1553024507.7643 | -22 | 0.20947 | 2.1e-05 | 1e-06 | 101.8821 | 102.0381 | 129.0097 | 21.2612 | 0.001818 | 0.001673 | -152.2752 | 0.23549 | 8.4044 | 0.003465 | 0.004396 | 0.004182 | 5.4e-05 | 0.001193 | 0.33107 | 1.3689 | 2.548 | 101.21 | 20.83 | 17.972 | 65.5788 | 105.2404 | 62.0705 | 0.004375 | 1.0532 | 0.025807 | 24.5041 | -0.013682 | 0.14009 | 8.2e-05 | 0.006327 | 0.006327 | 0.92107
2019-03-19 15:42 | 1553024508.7641 | -21 | 0.20947 | 2.5e-05 | 1e-06 | 101.7731 | 101.9103 | 128.5091 | 21.2612 | 0.001765 | 0.001705 | -152.2573 | 0.23586 | 8.4049 | 0.003466 | 0.004124 | 0.003987 | 4.9e-05 | 0.00119 | 0.34379 | 1.3403 | 2.548 | 101.21 | 20.84 | 17.974 | 64.3035 | 105.264 | 62.8478 | 0.004378 | 1.0507 | 0.025727 | 24.4849 | -0.012046 | 0.14031 | 6.1e-05 | 0.004704 | 0.004704 | 1.241
2019-03-19 15:42 | 1553024509.7632 | -20 | 0.2095 | 2.1e-05 | 1e-06 | 101.69 | 101.9531 | 128.4786 | 21.2612 | 0.0018 | 0.001642 | -152.2544 | 0.23521 | 8.4035 | 0.00339 | 0.004665 | 0.004399 | 5.4e-05 | 0.001193 | 0.30879 | 1.2935 | 2.5479 | 101.21 | 20.84 | 17.974 | 65.1473 | 103.4255 | 61.3282 | 0.004378 | 1.0555 | 0.025782 | 24.4265 | -0.023112 | 0.13993 | -0.000108 | -0.008447 | -0.008447 | -0.68912
2019-03-19 15:42 | 1553024510.7632 | -19 | 0.20952 | 2.5e-05 | 1e-06 | 101.6084 | 101.8709 | 128.9912 | 21.2612 | 0.001813 | 0.001614 | -152.268 | 0.23652 | 8.404 | 0.0034 | 0.004517 | 0.004314 | 5.4e-05 | 0.001194 | 0.31758 | 1.3157 | 2.5481 | 101.21 | 20.84 | 17.974 | 65.4408 | 103.6709 | 60.6629 | 0.004378 | 1.0576 | 0.02586 | 24.4509 | -0.02308 | 0.1407 | -0.000221 | -0.01734 | -0.01734 | -0.33756
2019-03-19 15:42 | 1553024511.7645 | -18 | 0.20952 | 2.5e-05 | 0 | 101.561 | 101.6589 | 129.5311 | 21.2612 | 0.001752 | 0.001659 | -152.2677 | 0.23607 | 8.3982 | 0.003417 | 0.005012 | 0.004784 | 5.2e-05 | 0.001193 | 0.29607 | 1.3118 | 2.548 | 101.21 | 20.84 | 17.973 | 63.9832 | 104.0891 | 61.7565 | 0.004378 | 1.0542 | 0.025871 | 24.5421 | -0.008622 | 0.14044 | -0.000255 | -0.019933 | -0.019933 | -0.2931
2019-03-19 15:42 | 1553024512.7633 | -17 | 0.20952 | 2.5e-05 | 1e-06 | 101.4793 | 101.6603 | 129.3423 | 21.2612 | 0.001765 | 0.001617 | -152.2815 | 0.23611 | 8.4058 | 0.003487 | 0.005038 | 0.004873 | 5.2e-05 | 0.001197 | 0.3194 | 1.3212 | 2.5479 | 101.21 | 20.84 | 17.973 | 64.2986 | 105.7705 | 60.7436 | 0.004378 | 1.0574 | 0.025892 | 24.4871 | -0.015934 | 0.14046 | -0.000255 | -0.01999 | -0.01999 | -0.2923
2019-03-19 15:42 | 1553024513.7633 | -16 | 0.20951 | 2.7e-05 | 0 | 101.476 | 101.7863 | 129.2677 | 21.2612 | 0.00181 | 0.001675 | -152.2752 | 0.23502 | 8.4035 | 0.003474 | 0.005191 | 0.004885 | 5.1e-05 | 0.001198 | 0.28672 | 1.3447 | 2.548 | 101.21 | 20.84 | 17.972 | 65.3837 | 105.4621 | 62.1197 | 0.004378 | 1.053 | 0.025831 | 24.5304 | -0.027308 | 0.13981 | -0.000194 | -0.015145 | -0.015145 | -0.38402
2019-03-19 15:42 | 1553024514.7634 | -15 | 0.20952 | 1.9e-05 | 0 | 101.4181 | 101.5709 | 129.5827 | 21.2612 | 0.001797 | 0.001642 | -152.2683 | 0.2355 | 8.4023 | 0.003444 | 0.004555 | 0.004446 | 5.3e-05 | 0.001198 | 0.31094 | 1.3503 | 2.5481 | 101.21 | 20.84 | 17.97 | 65.0715 | 104.7282 | 61.329 | 0.004377 | 1.0555 | 0.025893 | 24.5313 | -0.013463 | 0.1401 | -0.000244 | -0.019085 | -0.019085 | -0.30538
2019-03-19 15:42 | 1553024515.7643 | -14 | 0.2095 | 2.5e-05 | 0 | 101.3368 | 101.474 | 129.6639 | 21.2612 | 0.001768 | 0.001676 | -148.0785 | 0.23583 | 8.4013 | 0.003464 | 0.004929 | 0.004697 | 5.5e-05 | 0.001197 | 0.28351 | 1.3436 | 2.548 | 101.21 | 20.84 | 17.968 | 64.3595 | 105.2058 | 62.1555 | 0.004377 | 1.0529 | 0.025869 | 24.5693 | -0.0121 | 0.1403 | -0.000104 | -0.008172 | -0.008172 | -0.71416
2019-03-19 15:42 | 1553024516.7635 | -13 | 0.20949 | 2.3e-05 | 1e-06 | 101.2108 | 101.1366 | 129.466 | 21.2612 | 0.001788 | 0.001653 | -116.1051 | 0.23549 | 8.4045 | 0.003514 | 0.005349 | 0.005058 | 4.9e-05 | 0.001191 | 0.29142 | 1.351 | 2.5481 | 101.21 | 20.84 | 17.962 | 64.8381 | 106.4293 | 61.6055 | 0.004375 | 1.0546 | 0.025871 | 24.5304 | 0.006554 | 0.14009 | -4.1e-05 | -0.003325 | -0.003325 | -1.7526
2019-03-19 15:42 | 1553024517.7638 | -12 | 0.20953 | 2.3e-05 | 1e-06 | 101.196 | 101.4493 | 129.5404 | 21.2612 | 0.001782 | 0.001652 | -73.2191 | 0.23657 | 8.402 | 0.003467 | 0.004532 | 0.004336 | 4.9e-05 | 0.001199 | 0.32565 | 1.3231 | 2.548 | 101.21 | 20.84 | 17.957 | 64.7039 | 105.2882 | 61.5696 | 0.004374 | 1.0548 | 0.025879 | 24.5361 | -0.022357 | 0.14074 | -0.000266 | -0.02087 | -0.02087 | -0.28052
2019-03-19 15:42 | 1553024518.7636 | -11 | 0.20954 | 2.6e-05 | 1e-06 | 101.1192 | 101.1267 | 129.885 | 21.2612 | 0.001718 | 0.001619 | -31.6964 | 0.23557 | 8.4029 | 0.003523 | 0.004994 | 0.004806 | 5.1e-05 | 0.001194 | 0.31741 | 1.3325 | 2.5481 | 101.21 | 20.84 | 17.944 | 63.1689 | 106.6608 | 60.7913 | 0.004371 | 1.0572 | 0.025944 | 24.5402 | -0.00066 | 0.14014 | -0.00034 | -0.026666 | -0.026666 | -0.21862
2019-03-19 15:42 | 1553024519.7636 | -10 | 0.20954 | 2.8e-05 | 1e-06 | 101.0406 | 101.1626 | 129.7353 | 21.2612 | 0.001728 | 0.001605 | 9.837 | 0.23636 | 8.4024 | 0.003651 | 0.005438 | 0.005204 | 5e-05 | 0.001194 | 0.31725 | 1.3197 | 2.548 | 101.21 | 20.84 | 17.938 | 63.4014 | 109.7612 | 60.4479 | 0.00437 | 1.0583 | 0.025942 | 24.5134 | -0.010798 | 0.14061 | -0.000372 | -0.02918 | -0.02918 | -0.20046
2019-03-19 15:42 | 1553024520.7636 | -9 | 0.20952 | 2.4e-05 | 1e-06 | 101.0625 | 101.0712 | 129.5681 | 21.2612 | 0.001794 | 0.001651 | 51.3913 | 0.23925 | 8.4042 | 0.003666 | 0.004653 | 0.004389 | 5e-05 | 0.001188 | 0.30934 | 1.3447 | 2.548 | 101.21 | 20.84 | 17.932 | 64.9974 | 110.1309 | 61.5532 | 0.004368 | 1.0548 | 0.025883 | 24.5382 | -0.00077 | 0.14233 | -0.0002 | -0.01567 | -0.01567 | -0.37785
2019-03-19 15:42 | 1553024521.7631 | -8 | 0.20953 | 2.1e-05 | 1e-06 | 100.9536 | 100.8681 | 129.3191 | 21.2612 | 0.001764 | 0.001637 | 92.9367 | 0.25949 | 8.4019 | 0.003762 | 0.005173 | 0.00506 | 5.1e-05 | 0.001187 | 0.30735 | 1.357 | 2.5481 | 101.21 | 20.85 | 17.921 | 64.283 | 112.4521 | 61.2054 | 0.004368 | 1.0559 | 0.025871 | 24.5018 | 0.007575 | 0.15437 | -0.000287 | -0.022475 | -0.022475 | -0.28573
2019-03-19 15:42 | 1553024522.7629 | -7 | 0.20952 | 2.2e-05 | 1e-06 | 100.8116 | 100.8852 | 129.2949 | 21.2612 | 0.001783 | 0.001616 | 134.5044 | 0.26904 | 8.4015 | 0.003799 | 0.005867 | 0.005719 | 4.9e-05 | 0.001185 | 0.31166 | 1.3646 | 2.5481 | 101.21 | 20.85 | 17.917 | 64.7327 | 113.3559 | 60.7161 | 0.004367 | 1.0574 | 0.025888 | 24.4816 | -0.006529 | 0.16005 | -0.000254 | -0.019906 | -0.019906 | -0.33449
2019-03-19 15:42 | 1553024523.7628 | -6 | 0.20951 | 2.8e-05 | 1e-06 | 100.712 | 100.6136 | 129.3444 | 21.2612 | 0.001715 | 0.00168 | 176.0578 | 0.26811 | 8.404 | 0.003782 | 0.005979 | 0.005805 | 5.1e-05 | 0.001177 | 0.30309 | 1.3615 | 2.548 | 101.21 | 20.85 | 17.913 | 63.0899 | 112.9394 | 62.2439 | 0.004366 | 1.0526 | 0.025834 | 24.5422 | 0.008739 | 0.1595 | -0.000172 | -0.013452 | -0.013452 | -0.49327
2019-03-19 15:42 | 1553024524.7638 | -5 | 0.20952 | 2.3e-05 | 1e-06 | 100.6051 | 100.6143 | 129.4484 | 21.2612 | 0.001727 | 0.00167 | 263.0048 | 0.26818 | 8.4037 | 0.00393 | 0.00578 | 0.005757 | 5.1e-05 | 0.001169 | 0.30093 | 1.3519 | 2.548 | 101.21 | 20.85 | 17.911 | 63.3819 | 116.5487 | 62.0048 | 0.004366 | 1.0534 | 0.025853 | 24.5434 | -0.000815 | 0.15954 | -0.000201 | -0.015777 | -0.015777 | -0.42066
2019-03-19 15:42 | 1553024525.7629 | -4 | 0.20952 | 1.5e-05 | 1e-06 | 100.5622 | 100.6232 | 129.129 | 21.2612 | 0.001719 | 0.001647 | 263.225 | 0.26796 | 8.4015 | 0.003829 | 0.005813 | 0.005754 | 5.4e-05 | 0.001168 | 0.32189 | 1.3841 | 2.5482 | 101.21 | 20.85 | 17.91 | 63.2 | 114.0964 | 61.4439 | 0.004365 | 1.0551 | 0.025843 | 24.4925 | -0.005423 | 0.15941 | -0.000228 | -0.017872 | -0.017872 | -0.37105
2019-03-19 15:42 | 1553024526.7645 | -3 | 0.20952 | 2.7e-05 | 1e-06 | 100.4287 | 100.4232 | 128.8797 | 21.2612 | 0.001766 | 0.001603 | 263.1183 | 0.26765 | 8.4048 | 0.003843 | 0.005014 | 0.004861 | 4.8e-05 | 0.001161 | 0.323 | 1.3797 | 2.5481 | 101.21 | 20.85 | 17.914 | 64.3174 | 114.4331 | 60.4074 | 0.004366 | 1.0584 | 0.025858 | 24.4309 | 0.000488 | 0.15922 | -0.000252 | -0.019722 | -0.019722 | -0.33584
2019-03-19 15:42 | 1553024527.764 | -2 | 0.20955 | 2.6e-05 | 1e-06 | 100.3465 | 100.4866 | 128.9313 | 21.2612 | 0.001811 | 0.001709 | 263.2151 | 0.268 | 8.402 | 0.003872 | 0.005111 | 0.005022 | 5.2e-05 | 0.001151 | 0.29999 | 1.3346 | 2.5481 | 101.21 | 20.85 | 17.917 | 65.4001 | 115.1441 | 62.9501 | 0.004367 | 1.0504 | 0.025765 | 24.5288 | -0.01248 | 0.15943 | -0.000435 | -0.033917 | -0.033917 | -0.19555
2019-03-19 15:42 | 1553024528.7631 | -1 | 0.20952 | 2.8e-05 | 1e-06 | 100.2528 | 100.4092 | 128.4793 | 21.2612 | 0.001699 | 0.001615 | 263.3705 | 0.26847 | 8.4021 | 0.003898 | 0.005992 | 0.005952 | 5.1e-05 | 0.00114 | 0.27096 | 1.33 | 2.5482 | 101.21 | 20.85 | 17.921 | 62.7173 | 115.773 | 60.6903 | 0.004368 | 1.0575 | 0.025807 | 24.4033 | -0.013938 | 0.15971 | -0.000233 | -0.018183 | -0.018183 | -0.36538
2019-03-19 15:42 | 1553024529.7631 | 0 | 0.20952 | 3.2e-05 | 1e-06 | 100.2132 | 100.0119 | 127.8307 | 21.2612 | 0.001792 | 0.001651 | 263.4744 | 0.26837 | 8.4027 | 0.003912 | 0.006065 | 0.011234 | 0.000345 | 0.001133 | 0.32764 | 1.3515 | 2.6674 | 101.21 | 20.85 | 17.921 | 64.9466 | 116.1189 | 61.5431 | 0.004368 | 1.0548 | 0.025709 | 24.3727 | 0.017988 | 0.15965 | -0.000246 | -0.019146 | -0.019146 | -0.3469
2019-03-19 15:42 | 1553024530.7631 | 1 | 0.20951 | 3.1e-05 | 2e-06 | 100.1172 | 99.9208 | 126.8918 | 21.2612 | 0.001764 | 0.001571 | 263.5516 | 0.26875 | 8.406 | 0.003913 | 0.005304 | 0.004988 | 5.6e-05 | 0.001137 | 0.35098 | 1.3414 | 2.5481 | 101.21 | 20.85 | 17.935 | 64.2746 | 116.1315 | 59.6324 | 0.004372 | 1.0609 | 0.025688 | 24.2136 | 0.017564 | 0.15988 | -0.000172 | -0.013465 | -0.013465 | -0.49392
2019-03-19 15:42 | 1553024531.7632 | 2 | 0.20951 | 3.4e-05 | 3e-06 | 100.0088 | 99.842 | 125.5581 | 21.2611 | 0.001779 | 0.001696 | 263.4023 | 0.26778 | 8.4015 | 0.004 | 0.00593 | 0.006112 | 6.4e-05 | 0.001118 | 0.36481 | 1.3179 | 2.5479 | 101.21 | 20.85 | 17.937 | 64.6328 | 118.2538 | 62.6274 | 0.004372 | 1.0514 | 0.025438 | 24.1942 | 0.014934 | 0.1593 | -0.000188 | -0.014587 | -0.014587 | -0.45432
2019-03-19 15:42 | 1553024532.7633 | 3 | 0.20949 | 3.4e-05 | 5e-06 | 99.9392 | 99.6601 | 124.1754 | 21.2611 | 0.00177 | 0.001747 | 263.618 | 0.26724 | 8.401 | 0.003883 | 0.005396 | 0.005933 | 8.1e-05 | 0.001141 | 0.55068 | 1.3491 | 2.5479 | 101.21 | 20.85 | 17.936 | 64.4189 | 115.3964 | 63.8707 | 0.004372 | 1.0475 | 0.025251 | 24.1051 | 0.025012 | 0.15898 | -7.2e-05 | -0.005807 | -0.005807 | -1.1389
2019-03-19 15:42 | 1553024533.7633 | 4 | 0.20946 | 3.6e-05 | 7e-06 | 99.7033 | 99.5166 | 122.4105 | 21.2612 | 0.001794 | 0.001786 | 263.5584 | 0.26857 | 8.4044 | 0.003843 | 0.006064 | 0.00585 | 5.1e-05 | 0.00115 | 0.37034 | 1.3379 | 2.5482 | 101.21 | 20.86 | 17.934 | 65.0036 | 114.4285 | 64.8015 | 0.004374 | 1.0447 | 0.025037 | 23.9662 | 0.016765 | 0.15977 | 0.000102 | 0.007215 | 0.007215 | 0.92122
2019-03-19 15:42 | 1553024534.7634 | 5 | 0.20941 | 3.7e-05 | 1e-05 | 99.5608 | 98.6669 | 120.0118 | 21.2612 | 0.001791 | 0.001837 | 263.3188 | 0.26965 | 8.4033 | 0.00386 | 0.005567 | 0.005493 | 5.3e-05 | 0.001151 | 0.36481 | 1.3277 | 2.5478 | 101.21 | 20.85 | 17.932 | 64.9153 | 114.8516 | 66.0385 | 0.004371 | 1.0409 | 0.024745 | 23.7736 | 0.080668 | 0.16042 | 0.000384 | 0.028006 | 0.028006 | 0.23828
2019-03-19 15:42 | 1553024535.7638 | 6 | 0.20895 | 4.7e-05 | 1.5e-05 | 99.6508 | 97.2148 | 117.8695 | 21.2612 | 0.0019 | 0.001887 | 263.1586 | 0.26781 | 8.4029 | 0.004019 | 0.005299 | 0.005213 | 4.5e-05 | 0.001149 | 0.2896 | 1.3061 | 1.8242 | 101.21 | 20.86 | 17.924 | 67.5357 | 118.7128 | 67.2393 | 0.004372 | 1.0372 | 0.02448 | 23.6021 | 0.22137 | 0.15932 | 0.003184 | 0.23422 | 0.23422 | 0.028297
2019-03-19 15:42 | 1553024536.7639 | 7 | 0.20817 | 7.6e-05 | 2.4e-05 | 99.4443 | 97.5244 | 115.8058 | 21.2612 | 0.001946 | 0.001954 | 263.0854 | 0.26748 | 8.4026 | 0.003834 | 0.005496 | 0.005415 | 4.6e-05 | 0.001147 | 0.25155 | 1.3368 | 2.548 | 101.21 | 20.86 | 17.92 | 68.6457 | 114.2133 | 68.8375 | 0.004371 | 1.0323 | 0.024208 | 23.4495 | 0.17438 | 0.15912 | 0.007831 | 0.57021 | 0.57021 | 0.011609
2019-03-19 15:42 | 1553024537.7639 | 8 | 0.2069 | 0.000153 | 3.4e-05 | 99.2985 | 96.7849 | 113.2915 | 21.2612 | 0.001971 | 0.002059 | 262.9656 | 0.26737 | 8.4043 | 0.003803 | 0.005884 | 0.005818 | 5.2e-05 | 0.001142 | 0.31006 | 1.3466 | 2.548 | 101.21 | 20.86 | 17.915 | 69.2572 | 113.4569 | 71.3639 | 0.004369 | 1.0248 | 0.023856 | 23.279 | 0.22933 | 0.15906 | 0.01536 | 1.102 | 1.102 | 0.006004
2019-03-19 15:42 | 1553024538.7637 | 9 | 0.20511 | 0.000381 | 4.7e-05 | 99.2118 | 92.5274 | 111.5787 | 21.2612 | 0.002124 | 0.002188 | 262.849 | 0.26714 | 8.4046 | 0.003821 | 0.005498 | 0.005464 | 4.9e-05 | 0.001135 | 0.26593 | 1.3556 | 2.548 | 101.21 | 20.86 | 17.903 | 72.9217 | 113.9005 | 74.4596 | 0.004366 | 1.0156 | 0.023569 | 23.2059 | 0.6239 | 0.15892 | 0.025736 | 1.8229 | 1.8229 | 0.003627
2019-03-19 15:42 | 1553024539.7626 | 10 | 0.20345 | 0.00088 | 6.2e-05 | 99.1121 | 87.4178 | 110.1135 | 21.2612 | 0.002155 | 0.002268 | 262.7437 | 0.26811 | 8.402 | 0.003859 | 0.005485 | 0.005444 | 4.7e-05 | 0.001143 | 0.29679 | 1.3377 | 2.548 | 101.21 | 20.86 | 17.896 | 73.6785 | 114.8214 | 76.3897 | 0.004365 | 1.01 | 0.023349 | 23.117 | 1.123 | 0.1595 | 0.035055 | 2.4576 | 2.4576 | 0.0027
2019-03-19 15:42 | 1553024540.7632 | 11 | 0.20211 | 0.001534 | 7.5e-05 | 99.0092 | 81.3931 | 108.5679 | 21.2612 | 0.002307 | 0.002282 | 262.6273 | 0.2677 | 8.4008 | 0.003817 | 0.005543 | 0.005428 | 4.7e-05 | 0.001137 | 0.30718 | 1.3224 | 2.5481 | 101.21 | 20.86 | 17.888 | 77.3296 | 113.7882 | 76.7263 | 0.004363 | 1.0091 | 0.023173 | 22.9652 | 1.7524 | 0.15925 | 0.042194 | 2.9336 | 2.9336 | 0.002258
2019-03-19 15:42 | 1553024541.7641 | 12 | 0.20137 | 0.002256 | 8.5e-05 | 98.9696 | 75.1998 | 107.1788 | 21.2612 | 0.00233 | 0.002363 | 262.5185 | 0.26707 | 8.3998 | 0.003825 | 0.005519 | 0.005494 | 4.7e-05 | 0.001132 | 0.30469 | 1.3066 | 2.5481 | 101.21 | 20.86 | 17.872 | 77.8652 | 114.0009 | 78.665 | 0.004359 | 1.0035 | 0.022961 | 22.881 | 2.4567 | 0.15888 | 0.045708 | 3.1474 | 3.1474 | 0.0021
2019-03-19 15:42 | 1553024542.7629 | 13 | 0.2006 | 0.003018 | 9.4e-05 | 98.8207 | 71.8595 | 106.2643 | 21.2612 | 0.002392 | 0.002475 | 262.3971 | 0.26835 | 8.4023 | 0.003806 | 0.005216 | 0.005089 | 4.9e-05 | 0.001132 | 0.32587 | 1.3326 | 2.5481 | 101.21 | 20.87 | 17.864 | 79.3573 | 113.5203 | 81.3636 | 0.00436 | 0.99586 | 0.022776 | 22.8704 | 2.8497 | 0.15964 | 0.049396 | 3.3725 | 3.3725 | 0.001969
2019-03-19 15:42 | 1553024543.7639 | 14 | 0.20023 | 0.003772 | 0.000102 | 98.7691 | 71.1478 | 105.3268 | 21.2612 | 0.002463 | 0.002528 | 262.2689 | 0.26736 | 8.4044 | 0.003957 | 0.005543 | 0.005481 | 5e-05 | 0.001134 | 0.3246 | 1.347 | 2.548 | 101.21 | 20.87 | 17.853 | 81.0775 | 117.2186 | 82.6426 | 0.004357 | 0.99228 | 0.022634 | 22.8103 | 2.934 | 0.15905 | 0.050644 | 3.4355 | 3.4355 | 0.001926
2019-03-19 15:42 | 1553024544.7631 | 15 | 0.19973 | 0.004369 | 0.000107 | 98.7064 | 71.5923 | 104.7269 | 21.2612 | 0.002551 | 0.00262 | 262.15 | 0.26786 | 8.4031 | 0.003909 | 0.005214 | 0.005102 | 5.2e-05 | 0.00113 | 0.32211 | 1.364 | 2.5482 | 101.21 | 20.87 | 17.848 | 83.1771 | 116.0423 | 84.8418 | 0.004356 | 0.98618 | 0.0225 | 22.8154 | 2.8726 | 0.15935 | 0.052922 | 3.5679 | 3.5679 | 0.001858
2019-03-19 15:42 | 1553024545.7633 | 16 | 0.19909 | 0.004846 | 0.000112 | 98.7589 | 71.2413 | 104.5975 | 21.2612 | 0.002627 | 0.002686 | 262.0193 | 0.2671 | 8.4042 | 0.003926 | 0.005864 | 0.005745 | 5e-05 | 0.001131 | 0.27616 | 1.3479 | 2.548 | 101.21 | 20.87 | 17.842 | 85.0073 | 116.4521 | 86.4317 | 0.004354 | 0.98182 | 0.022437 | 22.8519 | 2.9214 | 0.1589 | 0.056103 | 3.7705 | 3.7705 | 0.001753
2019-03-19 15:42 | 1553024546.7641 | 17 | 0.19833 | 0.005282 | 0.000116 | 98.6821 | 70.8508 | 104.2448 | 21.2612 | 0.002724 | 0.002752 | 261.8803 | 0.26697 | 8.4017 | 0.00387 | 0.005627 | 0.005525 | 5.2e-05 | 0.001138 | 0.30038 | 1.336 | 2.548 | 101.21 | 20.87 | 17.831 | 87.3538 | 115.0944 | 88.0308 | 0.004352 | 0.97747 | 0.022349 | 22.864 | 2.9636 | 0.15882 | 0.06012 | 4.0232 | 4.0232 | 0.001642
2019-03-19 15:42 | 1553024547.7631 | 18 | 0.19785 | 0.005697 | 0.00012 | 98.5422 | 71.3396 | 103.8384 | 21.2612 | 0.002809 | 0.002837 | 261.7511 | 0.26806 | 8.4039 | 0.004003 | 0.005673 | 0.005605 | 4.9e-05 | 0.001142 | 0.31293 | 1.3358 | 2.548 | 101.21 | 20.87 | 17.826 | 89.4048 | 118.3418 | 90.0753 | 0.00435 | 0.97197 | 0.022243 | 22.8839 | 2.8894 | 0.15947 | 0.062456 | 4.1587 | 4.1587 | 0.001595
2019-03-19 15:42 | 1553024548.7636 | 19 | 0.19748 | 0.006109 | 0.000123 | 98.4813 | 71.953 | 103.2371 | 21.2612 | 0.00286 | 0.002838 | 261.6123 | 0.26722 | 8.4037 | 0.004069 | 0.005618 | 0.005519 | 5.2e-05 | 0.001143 | 0.28727 | 1.3462 | 2.548 | 101.21 | 20.87 | 17.815 | 90.6317 | 119.956 | 90.0924 | 0.004348 | 0.97193 | 0.022178 | 22.8181 | 2.8073 | 0.15897 | 0.064126 | 4.2566 | 4.2566 | 0.001554
2019-03-19 15:42 | 1553024549.7636 | 20 | 0.19697 | 0.006499 | 0.000126 | 98.4402 | 71.8523 | 102.8796 | 21.2612 | 0.002976 | 0.002936 | 261.4595 | 0.26735 | 8.4012 | 0.004092 | 0.005547 | 0.005446 | 5.1e-05 | 0.001138 | 0.31564 | 1.3691 | 2.5482 | 101.21 | 20.87 | 17.809 | 93.4172 | 120.5123 | 92.4637 | 0.004346 | 0.96562 | 0.022067 | 22.8528 | 2.8161 | 0.15905 | 0.06674 | 4.4071 | 4.4071 | 0.001501
2019-03-19 15:42 | 1553024550.7632 | 21 | 0.19676 | 0.006892 | 0.000128 | 98.3661 | 71.5358 | 102.3383 | 21.2612 | 0.002987 | 0.003033 | 261.3369 | 0.26729 | 8.4014 | 0.004068 | 0.005932 | 0.005884 | 5.2e-05 | 0.00113 | 0.3282 | 1.3353 | 2.5482 | 101.21 | 20.87 | 17.802 | 93.6902 | 119.9316 | 94.7991 | 0.004344 | 0.95949 | 0.021939 | 22.8652 | 2.8488 | 0.15901 | 0.067494 | 4.4306 | 4.4306 | 0.001493
2019-03-19 15:42 | 1553024551.7623 | 22 | 0.19667 | 0.007137 | 0.00013 | 98.3574 | 71.606 | 102.586 | 21.2612 | 0.003141 | 0.003067 | 261.1758 | 0.26694 | 8.4036 | 0.004132 | 0.005277 | 0.005174 | 5.3e-05 | 0.001126 | 0.27848 | 1.3288 | 2.548 | 101.21 | 20.87 | 17.785 | 97.4026 | 121.4837 | 95.6181 | 0.00434 | 0.95736 | 0.021941 | 22.9184 | 2.8392 | 0.1588 | 0.067717 | 4.4455 | 4.4455 | 0.001486
2019-03-19 15:43 | 1553024552.7623 | 23 | 0.19653 | 0.007377 | 0.000131 | 98.255 | 71.9047 | 102.659 | 21.2612 | 0.003233 | 0.003125 | 261.0221 | 0.26777 | 8.402 | 0.004194 | 0.005714 | 0.00562 | 5.1e-05 | 0.00112 | 0.28926 | 1.3606 | 2.548 | 101.21 | 20.87 | 17.775 | 99.6341 | 123.0142 | 97.0116 | 0.004338 | 0.95376 | 0.021908 | 22.9698 | 2.7927 | 0.1593 | 0.068295 | 4.4764 | 4.4764 | 0.00148
2019-03-19 15:43 | 1553024553.7624 | 24 | 0.19635 | 0.00762 | 0.000133 | 98.1116 | 71.0351 | 102.9864 | 21.2612 | 0.003254 | 0.003151 | 260.8683 | 0.2672 | 8.4021 | 0.004184 | 0.00599 | 0.005916 | 5.3e-05 | 0.001124 | 0.32228 | 1.3392 | 2.548 | 101.21 | 20.87 | 17.754 | 100.1259 | 122.7537 | 97.6432 | 0.004333 | 0.95214 | 0.021924 | 23.026 | 2.8885 | 0.15896 | 0.069036 | 4.5279 | 4.5279 | 0.00146
2019-03-19 15:43 | 1553024554.7625 | 25 | 0.19627 | 0.007837 | 0.000134 | 98.0612 | 70.8454 | 103.3462 | 21.2612 | 0.003246 | 0.003168 | 260.7202 | 0.26765 | 8.4017 | 0.004288 | 0.005785 | 0.005629 | 5.5e-05 | 0.001123 | 0.30574 | 1.3688 | 2.5481 | 101.21 | 20.87 | 17.742 | 99.9455 | 125.3082 | 98.0437 | 0.00433 | 0.95111 | 0.02195 | 23.0787 | 2.9078 | 0.15922 | 0.069253 | 4.5475 | 4.5475 | 0.001457
2019-03-19 15:43 | 1553024555.7636 | 26 | 0.19628 | 0.008044 | 0.000134 | 97.9192 | 71.9088 | 103.1754 | 21.2612 | 0.003359 | 0.003242 | 260.5786 | 0.26833 | 8.4023 | 0.004247 | 0.006528 | 0.006412 | 5.1e-05 | 0.001119 | 0.3079 | 1.3429 | 2.548 | 101.21 | 20.87 | 17.729 | 102.6791 | 124.313 | 99.8529 | 0.004327 | 0.94649 | 0.021879 | 23.1157 | 2.7616 | 0.15963 | 0.068994 | 4.5158 | 4.5158 | 0.001471
2019-03-19 15:43 | 1553024556.7643 | 27 | 0.19626 | 0.008236 | 0.000135 | 97.7395 | 71.1361 | 103.2887 | 21.2612 | 0.003371 | 0.0033 | 260.4377 | 0.26763 | 8.4023 | 0.004326 | 0.006099 | 0.005967 | 4.6e-05 | 0.001116 | 0.33843 | 1.4134 | 2.5481 | 101.21 | 20.87 | 17.703 | 102.9644 | 126.2386 | 101.2505 | 0.00432 | 0.94296 | 0.02185 | 23.1717 | 2.8418 | 0.15921 | 0.068838 | 4.4997 | 4.4997 | 0.001472
2019-03-19 15:43 | 1553024557.7627 | 28 | 0.19637 | 0.008404 | 0.000135 | 97.6473 | 70.8313 | 103.5631 | 21.2612 | 0.003439 | 0.003356 | 260.2974 | 0.2676 | 8.4024 | 0.004349 | 0.005595 | 0.005522 | 4.9e-05 | 0.001114 | 0.31022 | 1.3649 | 2.5482 | 101.21 | 20.87 | 17.689 | 104.6094 | 126.7952 | 102.6116 | 0.004317 | 0.93955 | 0.021839 | 23.2446 | 2.8717 | 0.1592 | 0.068005 | 4.4434 | 4.4434 | 0.00149
2019-03-19 15:43 | 1553024558.7627 | 29 | 0.19653 | 0.008554 | 0.000134 | 97.5588 | 71.2379 | 104.2619 | 21.2612 | 0.003407 | 0.00333 | 260.135 | 0.26773 | 8.4023 | 0.004374 | 0.005937 | 0.005762 | 5e-05 | 0.001116 | 0.30613 | 1.3307 | 2.5482 | 101.21 | 20.87 | 17.675 | 103.8352 | 127.4109 | 101.9613 | 0.004314 | 0.94117 | 0.021932 | 23.3027 | 2.8124 | 0.15927 | 0.066873 | 4.3884 | 4.3884 | 0.00151
2019-03-19 15:43 | 1553024559.7627 | 30 | 0.19671 | 0.008667 | 0.000132 | 97.5045 | 71.4317 | 104.7269 | 21.2612 | 0.003525 | 0.003376 | 259.9945 | 0.26709 | 8.4028 | 0.004425 | 0.005235 | 0.005267 | 5.4e-05 | 0.001128 | 0.31669 | 1.3113 | 2.5481 | 101.21 | 20.87 | 17.648 | 106.7047 | 128.6756 | 103.0869 | 0.004307 | 0.93836 | 0.021948 | 23.3896 | 2.7832 | 0.15889 | 0.065676 | 4.3135 | 4.3135 | 0.001532
2019-03-19 15:43 | 1553024560.7638 | 31 | 0.19695 | 0.008745 | 0.00013 | 97.3961 | 71.0492 | 105.713 | 21.2612 | 0.00356 | 0.003401 | 259.8548 | 0.26814 | 8.4035 | 0.004412 | 0.005842 | 0.005765 | 5.2e-05 | 0.001132 | 0.33804 | 1.3084 | 2.5481 | 101.21 | 20.87 | 17.633 | 107.553 | 128.3638 | 103.6952 | 0.004303 | 0.93685 | 0.022033 | 23.5184 | 2.8212 | 0.15952 | 0.064159 | 4.2309 | 4.2309 | 0.001568
2019-03-19 15:43 | 1553024561.7628 | 32 | 0.19741 | 0.008809 | 0.000127 | 97.2629 | 71.2433 | 105.977 | 21.2612 | 0.003511 | 0.00342 | 259.744 | 0.26843 | 8.4037 | 0.004449 | 0.005949 | 0.005877 | 5.3e-05 | 0.001128 | 0.27383 | 1.342 | 2.548 | 101.21 | 20.87 | 17.609 | 106.3568 | 129.2618 | 104.1613 | 0.004297 | 0.93569 | 0.022047 | 23.5623 | 2.7846 | 0.15969 | 0.06137 | 4.0506 | 4.0506 | 0.00164
2019-03-19 15:43 | 1553024562.7639 | 33 | 0.19772 | 0.008835 | 0.000125 | 97.2254 | 72.7431 | 105.82 | 21.2612 | 0.003578 | 0.003387 | 259.6266 | 0.26844 | 8.3988 | 0.004404 | 0.00589 | 0.005791 | 5e-05 | 0.001133 | 0.31293 | 1.3827 | 2.548 | 101.21 | 20.87 | 17.601 | 107.9744 | 128.1568 | 103.3439 | 0.004295 | 0.93772 | 0.022055 | 23.5194 | 2.5948 | 0.1597 | 0.05953 | 3.9311 | 3.9311 | 0.00169
2019-03-19 15:43 | 1553024563.763 | 34 | 0.19805 | 0.008842 | 0.000123 | 97.1936 | 73.7 | 105.6036 | 21.2612 | 0.003643 | 0.003454 | 259.5162 | 0.26763 | 8.4048 | 0.004384 | 0.005639 | 0.005543 | 5.1e-05 | 0.001136 | 0.32565 | 1.3283 | 2.548 | 101.21 | 20.87 | 17.596 | 109.5529 | 127.6658 | 104.9665 | 0.004294 | 0.9337 | 0.021985 | 23.5459 | 2.475 | 0.15921 | 0.057544 | 3.7887 | 3.7887 | 0.001748
2019-03-19 15:43 | 1553024564.763 | 35 | 0.19844 | 0.008813 | 0.00012 | 97.0054 | 74.074 | 105.855 | 21.2612 | 0.003626 | 0.003473 | 259.4058 | 0.26841 | 8.4001 | 0.004383 | 0.005654 | 0.005506 | 5.2e-05 | 0.001141 | 0.29192 | 1.3142 | 2.548 | 101.21 | 20.87 | 17.591 | 109.1485 | 127.6363 | 105.4441 | 0.004293 | 0.93252 | 0.021997 | 23.5888 | 2.4124 | 0.15968 | 0.055292 | 3.6432 | 3.6432 | 0.001823
2019-03-19 15:43 | 1553024565.7641 | 36 | 0.19875 | 0.008753 | 0.000118 | 96.8593 | 73.9518 | 106.6263 | 21.2612 | 0.003668 | 0.003461 | 259.267 | 0.26771 | 8.4046 | 0.004454 | 0.005784 | 0.005758 | 5.1e-05 | 0.001145 | 0.33588 | 1.3395 | 2.5479 | 101.21 | 20.87 | 17.589 | 110.1822 | 129.3844 | 105.157 | 0.004293 | 0.93322 | 0.022085 | 23.6656 | 2.4137 | 0.15926 | 0.053515 | 3.5408 | 3.5408 | 0.001871
2019-03-19 15:43 | 1553024566.7631 | 37 | 0.19905 | 0.008666 | 0.000116 | 96.8046 | 73.6367 | 107.2922 | 21.2612 | 0.003702 | 0.003488 | 259.1313 | 0.26638 | 8.4023 | 0.00438 | 0.005526 | 0.005464 | 5.1e-05 | 0.001143 | 0.30342 | 1.3223 | 2.548 | 101.21 | 20.86 | 17.594 | 110.9853 | 127.5685 | 105.8059 | 0.004291 | 0.93163 | 0.022135 | 23.7597 | 2.4468 | 0.15847 | 0.051791 | 3.435 | 3.435 | 0.001919
2019-03-19 15:43 | 1553024567.7622 | 38 | 0.19915 | 0.008547 | 0.000112 | 96.7377 | 73.4243 | 107.8799 | 21.2612 | 0.003735 | 0.003503 | 259.0147 | 0.26794 | 8.4032 | 0.0044 | 0.006514 | 0.006429 | 4.7e-05 | 0.001148 | 0.32731 | 1.3297 | 2.5479 | 101.21 | 20.86 | 17.601 | 111.8023 | 128.0651 | 106.1569 | 0.004293 | 0.93076 | 0.022185 | 23.8357 | 2.4664 | 0.1594 | 0.051351 | 3.4138 | 3.4138 | 0.001942
2019-03-19 15:43 | 1553024568.7633 | 39 | 0.19931 | 0.008421 | 0.000108 | 96.636 | 73.4247 | 107.8546 | 21.2612 | 0.00376 | 0.003527 | 258.9104 | 0.26681 | 8.4038 | 0.004435 | 0.00598 | 0.005871 | 4.8e-05 | 0.001146 | 0.29103 | 1.3202 | 2.5479 | 101.21 | 20.86 | 17.609 | 112.4143 | 128.9292 | 106.7341 | 0.004295 | 0.92935 | 0.022166 | 23.8511 | 2.457 | 0.15873 | 0.050582 | 3.3602 | 3.3602 | 0.001965
2019-03-19 15:43 | 1553024569.7636 | 40 | 0.19949 | 0.008285 | 0.000103 | 96.5715 | 74.5083 | 108.0745 | 21.2612 | 0.003743 | 0.003519 | 258.8064 | 0.26746 | 8.4012 | 0.004373 | 0.006697 | 0.00658 | 4.9e-05 | 0.001148 | 0.30326 | 1.2935 | 2.5478 | 101.21 | 20.87 | 17.63 | 111.9911 | 127.3918 | 106.5521 | 0.004303 | 0.9298 | 0.022194 | 23.8696 | 2.32 | 0.15911 | 0.049646 | 3.3026 | 3.3026 | 0.002004
2019-03-19 15:43 | 1553024570.7624 | 41 | 0.19977 | 0.008155 | 9.8e-05 | 96.5826 | 76.217 | 108.1128 | 21.2612 | 0.003764 | 0.003519 | 258.7049 | 0.26697 | 8.4019 | 0.00443 | 0.00593 | 0.005874 | 4.8e-05 | 0.001148 | 0.30055 | 1.3174 | 2.5481 | 101.21 | 20.87 | 17.642 | 112.5066 | 128.7904 | 106.5623 | 0.004305 | 0.92977 | 0.022198 | 23.8742 | 2.1182 | 0.15882 | 0.04816 | 3.2049 | 3.2049 | 0.002062
2019-03-19 15:43 | 1553024571.7623 | 42 | 0.20015 | 0.008024 | 9.3e-05 | 96.515 | 78.0247 | 108.5322 | 21.2612 | 0.003802 | 0.003442 | 258.6013 | 0.26733 | 8.4022 | 0.004355 | 0.005885 | 0.005766 | 4.7e-05 | 0.001149 | 0.30646 | 1.3256 | 2.548 | 101.21 | 20.87 | 17.655 | 113.4287 | 126.9584 | 104.6787 | 0.004309 | 0.93441 | 0.022296 | 23.8611 | 1.9023 | 0.15904 | 0.046043 | 3.0783 | 3.0783 | 0.002149
2019-03-19 15:43 | 1553024572.7625 | 43 | 0.20061 | 0.007887 | 8.8e-05 | 96.3812 | 79.2319 | 109.4037 | 21.2612 | 0.003733 | 0.003502 | 258.4983 | 0.26705 | 8.4037 | 0.004369 | 0.006147 | 0.006049 | 4.9e-05 | 0.001146 | 0.3147 | 1.3198 | 2.5479 | 101.21 | 20.87 | 17.682 | 111.7589 | 127.2931 | 106.1491 | 0.004315 | 0.93078 | 0.022342 | 24.0032 | 1.7525 | 0.15887 | 0.04345 | 2.9116 | 2.9116 | 0.00227
2019-03-19 15:43 | 1553024573.7636 | 44 | 0.20094 | 0.007745 | 8.4e-05 | 96.2792 | 80.3878 | 110.6417 | 21.2612 | 0.003764 | 0.003455 | 258.3925 | 0.26707 | 8.3999 | 0.004274 | 0.005241 | 0.005098 | 4.9e-05 | 0.001146 | 0.31581 | 1.348 | 2.548 | 101.21 | 20.87 | 17.693 | 112.4956 | 124.9764 | 105.0075 | 0.004318 | 0.93359 | 0.022502 | 24.1023 | 1.6135 | 0.15888 | 0.041707 | 2.8154 | 2.8154 | 0.002348
2019-03-19 15:43 | 1553024574.7637 | 45 | 0.20139 | 0.007585 | 8e-05 | 96.2119 | 81.5999 | 111.5036 | 21.2612 | 0.003788 | 0.003465 | 258.2768 | 0.2665 | 8.4039 | 0.004283 | 0.005739 | 0.005619 | 4.7e-05 | 0.001148 | 0.32908 | 1.3244 | 2.5479 | 101.21 | 20.87 | 17.715 | 113.0809 | 125.188 | 105.2482 | 0.004323 | 0.933 | 0.022582 | 24.2037 | 1.4734 | 0.15854 | 0.039212 | 2.657 | 2.657 | 0.002482
2019-03-19 15:43 | 1553024575.7636 | 46 | 0.20181 | 0.007407 | 7.6e-05 | 96.1507 | 82.6743 | 112.1427 | 21.2612 | 0.003763 | 0.003464 | 258.1932 | 0.26699 | 8.4013 | 0.004393 | 0.006291 | 0.006203 | 4.7e-05 | 0.00115 | 0.32444 | 1.3692 | 2.548 | 101.21 | 20.87 | 17.725 | 112.4766 | 127.888 | 105.223 | 0.004326 | 0.93306 | 0.022647 | 24.2722 | 1.3507 | 0.15883 | 0.03687 | 2.5061 | 2.5061 | 0.002636
2019-03-19 15:43 | 1553024576.764 | 47 | 0.20201 | 0.007223 | 7.3e-05 | 96.0124 | 83.2778 | 112.5758 | 21.2612 | 0.00376 | 0.003429 | 258.112 | 0.2671 | 8.4028 | 0.004366 | 0.00614 | 0.00599 | 4.8e-05 | 0.001145 | 0.30021 | 1.3187 | 2.548 | 101.21 | 20.87 | 17.745 | 112.407 | 127.2165 | 104.3712 | 0.004331 | 0.93517 | 0.022717 | 24.2916 | 1.2728 | 0.1589 | 0.035912 | 2.4488 | 2.4488 | 0.002699
2019-03-19 15:43 | 1553024577.7637 | 48 | 0.20237 | 0.007029 | 7e-05 | 95.9364 | 83.3754 | 113.0511 | 21.2612 | 0.003737 | 0.00343 | 257.9993 | 0.26738 | 8.4051 | 0.004358 | 0.005795 | 0.005791 | 4.7e-05 | 0.00115 | 0.31094 | 1.3415 | 2.5479 | 101.21 | 20.86 | 17.754 | 111.8378 | 127.0149 | 104.4051 | 0.00433 | 0.93508 | 0.022764 | 24.3439 | 1.2552 | 0.15906 | 0.034019 | 2.3249 | 2.3249 | 0.002846
2019-03-19 15:43 | 1553024578.7641 | 49 | 0.20271 | 0.006826 | 6.7e-05 | 95.8128 | 83.7867 | 114.0477 | 21.2612 | 0.003802 | 0.00333 | 257.9307 | 0.26763 | 8.4044 | 0.004416 | 0.005381 | 0.005277 | 4.9e-05 | 0.001153 | 0.30071 | 1.3487 | 2.548 | 101.21 | 20.86 | 17.761 | 113.4411 | 128.4437 | 101.9831 | 0.004332 | 0.94112 | 0.022937 | 24.3724 | 1.1997 | 0.15921 | 0.03221 | 2.2184 | 2.2184 | 0.002986
2019-03-19 15:43 | 1553024579.7629 | 50 | 0.20289 | 0.00662 | 6.5e-05 | 95.7444 | 84.1433 | 114.9317 | 21.2612 | 0.003791 | 0.00335 | 257.8598 | 0.26715 | 8.4029 | 0.004381 | 0.006188 | 0.006022 | 4.9e-05 | 0.001153 | 0.31454 | 1.3871 | 2.548 | 101.21 | 20.86 | 17.773 | 113.1661 | 127.583 | 102.4481 | 0.004335 | 0.93995 | 0.023012 | 24.4819 | 1.1553 | 0.15893 | 0.031378 | 2.1684 | 2.1684 | 0.003049
2019-03-19 15:43 | 1553024580.7632 | 51 | 0.20322 | 0.006397 | 6.2e-05 | 95.6745 | 83.9227 | 115.7737 | 21.2612 | 0.003753 | 0.003337 | 257.7891 | 0.26795 | 8.4033 | 0.004398 | 0.006092 | 0.005937 | 4.8e-05 | 0.001149 | 0.30469 | 1.3643 | 2.548 | 101.21 | 20.86 | 17.777 | 112.2272 | 128.0116 | 102.1293 | 0.004336 | 0.94075 | 0.023106 | 24.5609 | 1.1722 | 0.1594 | 0.029687 | 2.0602 | 2.0602 | 0.003219
2019-03-19 15:43 | 1553024581.7629 | 52 | 0.20371 | 0.006157 | 6e-05 | 95.639 | 83.5597 | 116.5149 | 21.2612 | 0.003659 | 0.003319 | 257.7112 | 0.26703 | 8.405 | 0.004299 | 0.005971 | 0.005865 | 4.7e-05 | 0.001147 | 0.33987 | 1.3509 | 2.548 | 101.21 | 20.86 | 17.782 | 109.9464 | 125.574 | 101.6997 | 0.004337 | 0.94183 | 0.023193 | 24.6253 | 1.2077 | 0.15886 | 0.027041 | 1.884 | 1.884 | 0.003508
2019-03-19 15:43 | 1553024582.7643 | 53 | 0.20426 | 0.005786 | 5.7e-05 | 95.6615 | 84.5162 | 117.2196 | 21.2612 | 0.003627 | 0.003304 | 257.6405 | 0.26684 | 8.402 | 0.004289 | 0.005709 | 0.005557 | 5e-05 | 0.001139 | 0.27992 | 1.3566 | 2.548 | 101.21 | 20.86 | 17.787 | 109.1847 | 125.3393 | 101.3316 | 0.004338 | 0.94276 | 0.023274 | 24.6876 | 1.108 | 0.15875 | 0.0242 | 1.6923 | 1.6923 | 0.003902
2019-03-19 15:43 | 1553024583.762 | 54 | 0.20469 | 0.00556 | 5.5e-05 | 95.5997 | 85.2516 | 117.7076 | 21.2612 | 0.003729 | 0.003181 | 257.5913 | 0.26753 | 8.4051 | 0.004298 | 0.005765 | 0.005672 | 4.9e-05 | 0.001141 | 0.31509 | 1.3545 | 2.548 | 101.21 | 20.86 | 17.789 | 111.6516 | 125.551 | 98.3668 | 0.004339 | 0.95028 | 0.023416 | 24.6408 | 1.0247 | 0.15915 | 0.021865 | 1.5384 | 1.5384 | 0.004304
2019-03-19 15:43 | 1553024584.7631 | 55 | 0.20505 | 0.005215 | 5.2e-05 | 95.6279 | 85.7266 | 118.2455 | 21.2612 | 0.003664 | 0.003312 | 257.5205 | 0.26751 | 8.4025 | 0.004221 | 0.005653 | 0.005482 | 4.7e-05 | 0.001143 | 0.31564 | 1.3287 | 2.5481 | 101.21 | 20.86 | 17.793 | 110.0757 | 123.6668 | 101.5412 | 0.00434 | 0.94223 | 0.023369 | 24.8023 | 0.97766 | 0.15914 | 0.020161 | 1.4159 | 1.4159 | 0.004676
2019-03-19 15:43 | 1553024585.7622 | 56 | 0.20534 | 0.004987 | 5.1e-05 | 95.6177 | 85.9708 | 118.7443 | 21.2612 | 0.003607 | 0.00322 | 257.4686 | 0.26726 | 8.4028 | 0.004204 | 0.005605 | 0.005563 | 4.8e-05 | 0.001148 | 0.32012 | 1.3331 | 2.548 | 101.21 | 20.86 | 17.795 | 108.6921 | 123.2518 | 99.3133 | 0.00434 | 0.94787 | 0.023489 | 24.7806 | 0.95126 | 0.15899 | 0.018712 | 1.3209 | 1.3209 | 0.005007
2019-03-19 15:43 | 1553024586.7633 | 57 | 0.20558 | 0.004773 | 4.8e-05 | 95.5388 | 86.2499 | 119.2398 | 21.2612 | 0.003554 | 0.003221 | 257.4206 | 0.26753 | 8.4029 | 0.004217 | 0.006141 | 0.006031 | 5e-05 | 0.001151 | 0.31293 | 1.3279 | 2.5481 | 101.21 | 20.86 | 17.796 | 107.4017 | 123.561 | 99.3297 | 0.00434 | 0.94782 | 0.023537 | 24.8327 | 0.91488 | 0.15915 | 0.017552 | 1.2417 | 1.2417 | 0.005332
2019-03-19 15:43 | 1553024587.7632 | 58 | 0.2058 | 0.004551 | 4.6e-05 | 95.3359 | 86.5852 | 119.8236 | 21.2612 | 0.003644 | 0.003147 | 257.3948 | 0.26755 | 8.4046 | 0.004157 | 0.005266 | 0.005177 | 4.9e-05 | 0.001148 | 0.31349 | 1.2826 | 2.5479 | 101.21 | 20.86 | 17.794 | 109.5813 | 122.1078 | 97.5461 | 0.00434 | 0.95239 | 0.023651 | 24.8338 | 0.86116 | 0.15917 | 0.016498 | 1.1729 | 1.1729 | 0.005645
2019-03-19 15:43 | 1553024588.7632 | 59 | 0.20594 | 0.004334 | 4.4e-05 | 95.257 | 86.791 | 120.0076 | 21.2612 | 0.003558 | 0.003159 | 257.3397 | 0.2673 | 8.4052 | 0.004197 | 0.00566 | 0.005587 | 4.9e-05 | 0.001151 | 0.30326 | 1.3124 | 2.5478 | 101.21 | 20.86 | 17.793 | 107.5095 | 123.0787 | 97.8439 | 0.00434 | 0.95162 | 0.02366 | 24.8628 | 0.83251 | 0.15902 | 0.015883 | 1.1296 | 1.1296 | 0.005856
2019-03-19 15:43 | 1553024589.7627 | 60 | 0.20612 | 0.004126 | 4.3e-05 | 95.2171 | 86.8065 | 120.1313 | 21.2612 | 0.003576 | 0.003165 | 257.2856 | 0.26747 | 8.4009 | 0.004131 | 0.005336 | 0.005131 | 4.8e-05 | 0.00115 | 0.2912 | 1.337 | 2.5482 | 101.21 | 20.86 | 17.792 | 107.946 | 121.4639 | 97.9756 | 0.004339 | 0.95128 | 0.023668 | 24.8801 | 0.82718 | 0.15912 | 0.01506 | 1.0716 | 1.0716 | 0.006177
2019-03-19 15:43 | 1553024590.7624 | 61 | 0.20624 | 0.003916 | 4.1e-05 | 95.0703 | 87.1404 | 120.3069 | 21.2612 | 0.003482 | 0.0031 | 257.2131 | 0.26801 | 8.4011 | 0.004109 | 0.005803 | 0.005716 | 4.9e-05 | 0.001148 | 0.2912 | 1.307 | 2.5481 | 101.21 | 20.86 | 17.789 | 105.6439 | 120.9313 | 96.4154 | 0.004339 | 0.9553 | 0.023735 | 24.8458 | 0.77904 | 0.15944 | 0.014645 | 1.045 | 1.045 | 0.006347
2019-03-19 15:43 | 1553024591.7625 | 62 | 0.20636 | 0.003731 | 3.9e-05 | 95.0082 | 87.4329 | 120.5949 | 21.2612 | 0.003511 | 0.003117 | 257.1641 | 0.26775 | 8.4044 | 0.004172 | 0.005907 | 0.005891 | 4.7e-05 | 0.001142 | 0.29358 | 1.3278 | 2.5479 | 101.21 | 20.86 | 17.788 | 106.3648 | 122.4748 | 96.8172 | 0.004338 | 0.95426 | 0.023751 | 24.8891 | 0.74322 | 0.15928 | 0.014112 | 1.0078 | 1.0078 | 0.006575
2019-03-19 15:43 | 1553024592.7634 | 63 | 0.20648 | 0.003551 | 3.8e-05 | 94.9116 | 87.6439 | 120.5817 | 21.2612 | 0.003509 | 0.003062 | 257.1212 | 0.26746 | 8.4004 | 0.004161 | 0.005556 | 0.005426 | 4.7e-05 | 0.001148 | 0.3366 | 1.3494 | 2.548 | 101.21 | 20.86 | 17.784 | 106.3206 | 122.2079 | 95.489 | 0.004337 | 0.9577 | 0.023792 | 24.843 | 0.71256 | 0.15911 | 0.013665 | 0.97761 | 0.97761 | 0.006771
2019-03-19 15:43 | 1553024593.764 | 64 | 0.20657 | 0.003374 | 3.7e-05 | 94.7946 | 87.9726 | 120.5606 | 21.2612 | 0.003512 | 0.003017 | 257.0649 | 0.26798 | 8.4009 | 0.004067 | 0.005601 | 0.005491 | 4.9e-05 | 0.001147 | 0.29679 | 1.3833 | 2.5481 | 101.21 | 20.86 | 17.78 | 106.3813 | 119.8854 | 94.4094 | 0.004336 | 0.96051 | 0.023825 | 24.8044 | 0.66803 | 0.15942 | 0.01331 | 0.95349 | 0.95349 | 0.006956
2019-03-19 15:43 | 1553024594.7637 | 65 | 0.20663 | 0.003212 | 3.6e-05 | 94.7519 | 88.1038 | 120.5728 | 21.2612 | 0.003424 | 0.003047 | 257.038 | 0.26739 | 8.4001 | 0.003984 | 0.005478 | 0.005315 | 5e-05 | 0.001152 | 0.30453 | 1.3701 | 2.5481 | 101.21 | 20.86 | 17.775 | 104.2493 | 117.8783 | 95.1255 | 0.004335 | 0.95864 | 0.023803 | 24.8298 | 0.65068 | 0.15907 | 0.013178 | 0.94326 | 0.94326 | 0.007015
2019-03-19 15:43 | 1553024595.7627 | 66 | 0.20666 | 0.003074 | 3.5e-05 | 94.5728 | 87.7045 | 120.6461 | 21.2612 | 0.003427 | 0.00301 | 256.9772 | 0.26796 | 8.4024 | 0.004013 | 0.005464 | 0.005378 | 5.4e-05 | 0.001154 | 0.33301 | 1.339 | 2.5481 | 101.21 | 20.86 | 17.763 | 104.3193 | 118.5875 | 94.2387 | 0.004332 | 0.96096 | 0.023839 | 24.8075 | 0.67439 | 0.15941 | 0.013141 | 0.94207 | 0.94207 | 0.007039
2019-03-19 15:43 | 1553024596.7642 | 67 | 0.20674 | 0.002941 | 3.5e-05 | 94.3726 | 87.5678 | 120.5583 | 21.2612 | 0.003386 | 0.00303 | 256.9212 | 0.26735 | 8.404 | 0.004054 | 0.005071 | 0.004967 | 5.2e-05 | 0.001151 | 0.2912 | 1.3662 | 2.5481 | 101.21 | 20.86 | 17.756 | 103.3376 | 119.5881 | 94.7272 | 0.004331 | 0.95968 | 0.023814 | 24.8149 | 0.66939 | 0.15905 | 0.012853 | 0.92044 | 0.92044 | 0.007188
2019-03-19 15:43 | 1553024597.7628 | 68 | 0.20675 | 0.002822 | 3.4e-05 | 94.2625 | 87.4175 | 120.6555 | 21.2612 | 0.003415 | 0.002991 | 256.8737 | 0.26708 | 8.4058 | 0.004033 | 0.005842 | 0.005648 | 4.6e-05 | 0.00115 | 0.29336 | 1.3681 | 2.5481 | 101.21 | 20.86 | 17.738 | 104.0387 | 119.0774 | 93.7801 | 0.004326 | 0.96216 | 0.023855 | 24.7929 | 0.67431 | 0.15889 | 0.012939 | 0.92825 | 0.92825 | 0.007121
2019-03-19 15:43 | 1553024598.762 | 69 | 0.20679 | 0.002719 | 3.4e-05 | 94.0731 | 87.1397 | 120.524 | 21.2612 | 0.003412 | 0.002951 | 256.8379 | 0.26743 | 8.4018 | 0.003929 | 0.005457 | 0.005445 | 4.9e-05 | 0.00115 | 0.31741 | 1.3441 | 2.5481 | 101.21 | 20.86 | 17.732 | 103.9498 | 116.5265 | 92.8282 | 0.004325 | 0.96466 | 0.023873 | 24.7473 | 0.68479 | 0.15909 | 0.012799 | 0.9189 | 0.9189 | 0.007202
2019-03-19 15:43 | 1553024599.763 | 70 | 0.20679 | 0.002633 | 3.4e-05 | 94.0098 | 87.4151 | 120.3132 | 21.2612 | 0.00338 | 0.002943 | 256.7808 | 0.26679 | 8.4021 | 0.003946 | 0.005581 | 0.005464 | 4.9e-05 | 0.001137 | 0.3095 | 1.3601 | 2.5481 | 101.21 | 20.86 | 17.727 | 103.1727 | 116.9382 | 92.6285 | 0.004324 | 0.96519 | 0.023858 | 24.7189 | 0.65055 | 0.15871 | 0.012932 | 0.9279 | 0.9279 | 0.007115
2019-03-19 15:43 | 1553024600.7633 | 71 | 0.20681 | 0.002553 | 3.3e-05 | 93.8282 | 87.1558 | 120.2048 | 21.2612 | 0.003413 | 0.002944 | 256.7333 | 0.26764 | 8.4032 | 0.003853 | 0.005645 | 0.005528 | 4.6e-05 | 0.001134 | 0.3126 | 1.3501 | 2.548 | 101.21 | 20.86 | 17.718 | 103.9767 | 114.6705 | 92.6411 | 0.004321 | 0.96516 | 0.023847 | 24.7081 | 0.65982 | 0.15922 | 0.012913 | 0.92616 | 0.92616 | 0.007152
2019-03-19 15:43 | 1553024601.7619 | 72 | 0.20681 | 0.002486 | 3.4e-05 | 93.6302 | 87.0561 | 119.9356 | 21.2612 | 0.003349 | 0.002962 | 256.6968 | 0.26789 | 8.4015 | 0.003901 | 0.005726 | 0.005668 | 5e-05 | 0.001129 | 0.31382 | 1.3829 | 2.5479 | 101.21 | 20.86 | 17.715 | 102.4293 | 115.8435 | 93.0854 | 0.004321 | 0.96398 | 0.023806 | 24.6954 | 0.65116 | 0.15937 | 0.013013 | 0.93168 | 0.93168 | 0.007116
2019-03-19 15:43 | 1553024602.7634 | 73 | 0.20682 | 0.00242 | 3.3e-05 | 93.5179 | 86.9106 | 119.86 | 21.2612 | 0.003333 | 0.00287 | 256.654 | 0.26648 | 8.4021 | 0.00387 | 0.005353 | 0.005269 | 4.6e-05 | 0.001132 | 0.34987 | 1.3793 | 2.5479 | 101.21 | 20.86 | 17.709 | 102.0547 | 115.0966 | 90.8755 | 0.004319 | 0.96984 | 0.023871 | 24.6131 | 0.6554 | 0.15853 | 0.013032 | 0.93561 | 0.93561 | 0.007049
2019-03-19 15:43 | 1553024603.762 | 74 | 0.20677 | 0.002377 | 3.3e-05 | 93.2763 | 86.6582 | 119.6966 | 21.2612 | 0.003286 | 0.00291 | 256.6252 | 0.26699 | 8.4045 | 0.003858 | 0.004399 | 0.004259 | 4.7e-05 | 0.001133 | 0.31133 | 1.3501 | 2.5481 | 101.21 | 20.86 | 17.705 | 100.9099 | 114.8002 | 91.8293 | 0.004318 | 0.9673 | 0.023823 | 24.6285 | 0.65828 | 0.15884 | 0.013341 | 0.9559 | 0.9559 | 0.006912
2019-03-19 15:43 | 1553024604.7622 | 75 | 0.20681 | 0.002325 | 3.4e-05 | 93.1881 | 86.4314 | 119.6342 | 21.2612 | 0.003285 | 0.002929 | 256.578 | 0.26719 | 8.4031 | 0.003875 | 0.005632 | 0.005576 | 5e-05 | 0.001135 | 0.30381 | 1.3772 | 2.548 | 101.21 | 20.86 | 17.703 | 100.8895 | 115.2185 | 92.2773 | 0.004318 | 0.96612 | 0.023802 | 24.6372 | 0.67324 | 0.15895 | 0.013197 | 0.94476 | 0.94476 | 0.006999
2019-03-19 15:43 | 1553024605.7622 | 76 | 0.20678 | 0.002283 | 3.4e-05 | 93.0002 | 86.0845 | 119.6882 | 21.2612 | 0.003305 | 0.002861 | 256.5486 | 0.26785 | 8.4005 | 0.003969 | 0.004939 | 0.004859 | 4.9e-05 | 0.001137 | 0.30198 | 1.3446 | 2.5481 | 101.21 | 20.86 | 17.7 | 101.3565 | 117.5108 | 90.6481 | 0.004317 | 0.97044 | 0.023861 | 24.5877 | 0.69117 | 0.15934 | 0.013395 | 0.96127 | 0.96127 | 0.006896
2019-03-19 15:43 | 1553024606.7626 | 77 | 0.20676 | 0.002256 | 3.4e-05 | 92.8529 | 86.0376 | 119.9417 | 21.2612 | 0.003329 | 0.002872 | 256.5049 | 0.26741 | 8.4026 | 0.003934 | 0.005289 | 0.005127 | 4.6e-05 | 0.001134 | 0.32493 | 1.3189 | 2.548 | 101.21 | 20.86 | 17.701 | 101.9439 | 116.6429 | 90.9076 | 0.004317 | 0.96975 | 0.023878 | 24.6225 | 0.68186 | 0.15908 | 0.013567 | 0.97428 | 0.97428 | 0.006792
2019-03-19 15:43 | 1553024607.7637 | 78 | 0.20679 | 0.002228 | 3.5e-05 | 92.734 | 85.8523 | 119.6532 | 21.2612 | 0.003262 | 0.002907 | 256.4495 | 0.26684 | 8.4031 | 0.003905 | 0.005022 | 0.004935 | 4.9e-05 | 0.001129 | 0.2959 | 1.3078 | 2.548 | 101.21 | 20.86 | 17.701 | 100.3165 | 115.933 | 91.748 | 0.004317 | 0.96752 | 0.023822 | 24.6213 | 0.68969 | 0.15875 | 0.013439 | 0.96278 | 0.96278 | 0.006859
2019-03-19 15:43 | 1553024608.7623 | 79 | 0.20678 | 0.00221 | 3.5e-05 | 92.5585 | 85.9276 | 119.8086 | 21.2612 | 0.003274 | 0.002846 | 256.4248 | 0.26747 | 8.4032 | 0.00385 | 0.005608 | 0.005532 | 4.9e-05 | 0.001131 | 0.2896 | 1.297 | 2.548 | 101.21 | 20.86 | 17.705 | 100.6053 | 114.6057 | 90.2877 | 0.004318 | 0.9714 | 0.023885 | 24.5879 | 0.6649 | 0.15912 | 0.013536 | 0.97235 | 0.97235 | 0.006808
2019-03-19 15:43 | 1553024609.7634 | 80 | 0.20683 | 0.002183 | 3.5e-05 | 92.4192 | 85.742 | 120.0297 | 21.2612 | 0.003233 | 0.002821 | 256.4016 | 0.26817 | 8.4022 | 0.003869 | 0.005246 | 0.005118 | 5.1e-05 | 0.001137 | 0.32388 | 1.305 | 2.548 | 101.21 | 20.86 | 17.707 | 99.6261 | 115.0683 | 89.6885 | 0.004319 | 0.97301 | 0.023927 | 24.5903 | 0.67077 | 0.15954 | 0.013274 | 0.95514 | 0.95514 | 0.006948
2019-03-19 15:43 | 1553024610.7635 | 81 | 0.20678 | 0.00217 | 3.6e-05 | 92.1578 | 85.2303 | 120.083 | 21.2612 | 0.003241 | 0.002769 | 256.3748 | 0.26811 | 8.4028 | 0.003865 | 0.005358 | 0.005211 | 4.6e-05 | 0.001141 | 0.33771 | 1.3313 | 2.5481 | 101.21 | 20.86 | 17.71 | 99.8145 | 114.9574 | 88.4425 | 0.004319 | 0.97636 | 0.023973 | 24.5535 | 0.69898 | 0.1595 | 0.013565 | 0.97797 | 0.97797 | 0.006785
2019-03-19 15:43 | 1553024611.7635 | 82 | 0.20681 | 0.002153 | 3.6e-05 | 91.976 | 85.1039 | 119.8299 | 21.2612 | 0.003209 | 0.002794 | 256.3313 | 0.26647 | 8.4028 | 0.003899 | 0.00566 | 0.005599 | 5e-05 | 0.001141 | 0.3063 | 1.3202 | 2.5481 | 101.21 | 20.86 | 17.711 | 99.052 | 115.7913 | 89.0439 | 0.00432 | 0.97474 | 0.023928 | 24.548 | 0.69458 | 0.15852 | 0.013372 | 0.96215 | 0.96215 | 0.006854
2019-03-19 15:44 | 1553024612.7636 | 83 | 0.2068 | 0.00214 | 3.7e-05 | 91.7995 | 85.0375 | 119.9344 | 21.2612 | 0.003212 | 0.002793 | 256.2797 | 0.26686 | 8.4015 | 0.003786 | 0.005357 | 0.005223 | 4.7e-05 | 0.00114 | 0.31045 | 1.3218 | 2.5482 | 101.21 | 20.86 | 17.712 | 99.1211 | 113.0295 | 89.0118 | 0.00432 | 0.97483 | 0.023939 | 24.5576 | 0.68439 | 0.15875 | 0.013447 | 0.96802 | 0.96802 | 0.006822
2019-03-19 15:44 | 1553024613.763 | 84 | 0.20681 | 0.002131 | 3.7e-05 | 91.5687 | 84.7813 | 119.9346 | 21.2612 | 0.003151 | 0.002796 | 256.2311 | 0.26676 | 8.4039 | 0.003864 | 0.005561 | 0.005522 | 5e-05 | 0.001144 | 0.30646 | 1.3177 | 2.5479 | 101.21 | 20.86 | 17.712 | 97.6456 | 114.9503 | 89.072 | 0.00432 | 0.97467 | 0.023937 | 24.5597 | 0.68886 | 0.1587 | 0.013435 | 0.96709 | 0.96709 | 0.006826
2019-03-19 15:44 | 1553024614.7627 | 85 | 0.20682 | 0.002124 | 3.7e-05 | 91.4028 | 84.6388 | 119.9064 | 21.2612 | 0.003232 | 0.002781 | 256.2104 | 0.26748 | 8.4017 | 0.003797 | 0.005406 | 0.005293 | 4.5e-05 | 0.001144 | 0.34938 | 1.3544 | 2.548 | 101.21 | 20.86 | 17.711 | 99.5956 | 113.303 | 88.7157 | 0.00432 | 0.97562 | 0.023946 | 24.5447 | 0.68768 | 0.15913 | 0.013351 | 0.9614 | 0.9614 | 0.006885
2019-03-19 15:44 | 1553024615.7618 | 86 | 0.20685 | 0.002115 | 3.7e-05 | 91.2227 | 84.4676 | 119.9752 | 21.2612 | 0.00312 | 0.002785 | 256.1602 | 0.26703 | 8.4014 | 0.003805 | 0.005248 | 0.00515 | 4.6e-05 | 0.001145 | 0.31741 | 1.3311 | 2.5478 | 101.21 | 20.86 | 17.71 | 96.8901 | 113.5061 | 88.8149 | 0.004319 | 0.97536 | 0.02395 | 24.5551 | 0.68816 | 0.15885 | 0.013214 | 0.95167 | 0.95167 | 0.006944
2019-03-19 15:44 | 1553024616.7628 | 87 | 0.20689 | 0.002105 | 3.7e-05 | 90.9802 | 84.4635 | 119.9116 | 21.2612 | 0.003146 | 0.002722 | 256.1292 | 0.26691 | 8.4026 | 0.003818 | 0.005357 | 0.005306 | 4.5e-05 | 0.001146 | 0.30989 | 1.3165 | 2.5482 | 101.21 | 20.85 | 17.711 | 97.521 | 113.8094 | 87.2922 | 0.004317 | 0.97948 | 0.023994 | 24.4969 | 0.66479 | 0.15879 | 0.012964 | 0.93538 | 0.93538 | 0.007062
2019-03-19 15:44 | 1553024617.7622 | 88 | 0.20686 | 0.002098 | 3.8e-05 | 90.7969 | 84.3186 | 120.0407 | 21.2612 | 0.003196 | 0.002755 | 256.0872 | 0.2672 | 8.4033 | 0.003742 | 0.004957 | 0.004913 | 4.8e-05 | 0.001143 | 0.33035 | 1.368 | 2.548 | 101.21 | 20.85 | 17.711 | 98.7312 | 111.9775 | 88.0987 | 0.004317 | 0.97729 | 0.02398 | 24.5375 | 0.6621 | 0.15896 | 0.013137 | 0.94724 | 0.94724 | 0.006981
2019-03-19 15:44 | 1553024618.7619 | 89 | 0.20686 | 0.002087 | 3.7e-05 | 90.6171 | 84.3707 | 120.0536 | 21.2612 | 0.003149 | 0.002723 | 256.0194 | 0.26765 | 8.4018 | 0.003733 | 0.004753 | 0.004653 | 4.3e-05 | 0.001139 | 0.30165 | 1.3556 | 2.5481 | 101.21 | 20.85 | 17.714 | 97.5922 | 111.7589 | 87.3203 | 0.004318 | 0.9794 | 0.024007 | 24.5124 | 0.63885 | 0.15922 | 0.013157 | 0.94974 | 0.94974 | 0.006974
2019-03-19 15:44 | 1553024619.762 | 90 | 0.20688 | 0.002083 | 3.7e-05 | 90.6387 | 84.4751 | 120.0952 | 21.2612 | 0.003129 | 0.002796 | 255.9607 | 0.2673 | 8.4025 | 0.003759 | 0.004941 | 0.004852 | 4.5e-05 | 0.001142 | 0.32139 | 1.32 | 2.5481 | 101.21 | 20.85 | 17.714 | 97.108 | 112.3929 | 89.0829 | 0.004318 | 0.97464 | 0.023953 | 24.5765 | 0.62991 | 0.15902 | 0.013068 | 0.94121 | 0.94121 | 0.007028
2019-03-19 15:44 | 1553024620.7629 | 91 | 0.2069 | 0.00207 | 3.6e-05 | 90.5325 | 84.2128 | 120.1215 | 21.2612 | 0.003177 | 0.002669 | 255.9095 | 0.26702 | 8.4045 | 0.003649 | 0.004718 | 0.004618 | 4.7e-05 | 0.001144 | 0.30917 | 1.3201 | 2.5479 | 101.21 | 20.85 | 17.716 | 98.2655 | 109.706 | 86.0345 | 0.004318 | 0.98291 | 0.024057 | 24.4755 | 0.64725 | 0.15885 | 0.012975 | 0.9386 | 0.9386 | 0.00704
2019-03-19 15:44 | 1553024621.7634 | 92 | 0.20693 | 0.002058 | 3.5e-05 | 90.5076 | 84.4059 | 120.3719 | 21.2612 | 0.003094 | 0.002759 | 255.8695 | 0.26673 | 8.404 | 0.003688 | 0.004904 | 0.004858 | 4.8e-05 | 0.001148 | 0.30718 | 1.2759 | 2.5481 | 101.21 | 20.85 | 17.717 | 96.2626 | 110.65 | 88.187 | 0.004318 | 0.97705 | 0.02401 | 24.5743 | 0.62429 | 0.15868 | 0.012797 | 0.92403 | 0.92403 | 0.007144
2019-03-19 15:44 | 1553024622.7622 | 93 | 0.20688 | 0.002041 | 3.5e-05 | 90.591 | 84.709 | 120.2092 | 21.2612 | 0.003104 | 0.002685 | 255.8368 | 0.26783 | 8.4013 | 0.003672 | 0.004369 | 0.004229 | 4.3e-05 | 0.00115 | 0.29662 | 1.319 | 2.548 | 101.21 | 20.85 | 17.719 | 96.5086 | 110.2573 | 86.4192 | 0.004319 | 0.98186 | 0.024053 | 24.4976 | 0.60047 | 0.15933 | 0.013144 | 0.9508 | 0.9508 | 0.006971
2019-03-19 15:44 | 1553024623.7622 | 94 | 0.20688 | 0.002036 | 3.5e-05 | 90.712 | 84.9806 | 120.078 | 21.2612 | 0.003059 | 0.002735 | 255.7756 | 0.26769 | 8.4038 | 0.003593 | 0.004677 | 0.00461 | 4.3e-05 | 0.00115 | 0.3183 | 1.343 | 2.5479 | 101.21 | 20.85 | 17.718 | 95.4332 | 108.3423 | 87.6161 | 0.004319 | 0.9786 | 0.024 | 24.5249 | 0.58377 | 0.15925 | 0.013148 | 0.94898 | 0.94898 | 0.006981
2019-03-19 15:44 | 1553024624.7622 | 95 | 0.20697 | 0.002025 | 3.4e-05 | 90.7897 | 85.4886 | 120.1125 | 21.2612 | 0.003095 | 0.002738 | 255.7279 | 0.26705 | 8.4023 | 0.003558 | 0.005023 | 0.004918 | 4.8e-05 | 0.001152 | 0.32228 | 1.3208 | 2.548 | 101.21 | 20.85 | 17.716 | 96.294 | 107.4891 | 87.6975 | 0.004318 | 0.97838 | 0.024001 | 24.5312 | 0.53813 | 0.15887 | 0.012623 | 0.91107 | 0.91107 | 0.007254
2019-03-19 15:44 | 1553024625.7637 | 96 | 0.20687 | 0.002005 | 3.4e-05 | 91.0319 | 85.6417 | 119.7999 | 21.2612 | 0.003055 | 0.002713 | 255.6897 | 0.26733 | 8.4027 | 0.00358 | 0.004732 | 0.004682 | 4.5e-05 | 0.001156 | 0.3314 | 1.3489 | 2.5481 | 101.21 | 20.85 | 17.715 | 95.3369 | 108.0215 | 87.079 | 0.004318 | 0.98006 | 0.02399 | 24.4782 | 0.54595 | 0.15903 | 0.013258 | 0.95656 | 0.95656 | 0.006916
2019-03-19 15:44 | 1553024626.7622 | 97 | 0.20713 | 0.001994 | 3.4e-05 | 91.1924 | 85.6146 | 119.4436 | 21.2612 | 0.003076 | 0.002701 | 255.6462 | 0.26766 | 8.4004 | 0.003543 | 0.004852 | 0.004805 | 4.7e-05 | 0.001154 | 0.29159 | 1.3383 | 2.548 | 101.21 | 20.85 | 17.714 | 95.8428 | 107.1386 | 86.7861 | 0.004318 | 0.98086 | 0.023964 | 24.4319 | 0.56454 | 0.15923 | 0.011697 | 0.84293 | 0.84293 | 0.007858
2019-03-19 15:44 | 1553024627.7627 | 98 | 0.20711 | 0.00198 | 3.4e-05 | 91.4025 | 86.0783 | 119.1948 | 21.2612 | 0.003067 | 0.002689 | 255.6338 | 0.26716 | 8.4012 | 0.003597 | 0.004837 | 0.00467 | 4.4e-05 | 0.001157 | 0.33411 | 1.3205 | 2.548 | 101.21 | 20.85 | 17.709 | 95.6041 | 108.4353 | 86.5043 | 0.004316 | 0.98162 | 0.023949 | 24.3969 | 0.53681 | 0.15894 | 0.011843 | 0.85285 | 0.85285 | 0.007752
2019-03-19 15:44 | 1553024628.7628 | 99 | 0.20715 | 0.001971 | 3.4e-05 | 91.6054 | 86.6229 | 119.091 | 21.2612 | 0.003032 | 0.002663 | 255.583 | 0.26739 | 8.4029 | 0.003503 | 0.005155 | 0.005022 | 4.7e-05 | 0.001151 | 0.32084 | 1.3301 | 2.548 | 101.21 | 20.85 | 17.706 | 94.7687 | 106.1688 | 85.8807 | 0.004316 | 0.98333 | 0.023959 | 24.3651 | 0.50023 | 0.15907 | 0.011568 | 0.83347 | 0.83347 | 0.00794
2019-03-19 15:44 | 1553024629.7615 | 100 | 0.20717 | 0.001957 | 3.4e-05 | 91.7671 | 87.1452 | 119.1826 | 21.2612 | 0.00303 | 0.002625 | 255.533 | 0.26782 | 8.401 | 0.003598 | 0.005134 | 0.005097 | 4.6e-05 | 0.001146 | 0.3147 | 1.3009 | 2.5481 | 101.21 | 20.85 | 17.702 | 94.7202 | 108.4684 | 84.9621 | 0.004315 | 0.98585 | 0.023999 | 24.3433 | 0.46224 | 0.15933 | 0.011453 | 0.82655 | 0.82655 | 0.008019
2019-03-19 15:44 | 1553024630.762 | 101 | 0.20701 | 0.001939 | 3.3e-05 | 91.8929 | 87.2397 | 119.2793 | 21.2612 | 0.00309 | 0.002658 | 255.4859 | 0.26723 | 8.4015 | 0.003566 | 0.004983 | 0.004928 | 4.8e-05 | 0.001144 | 0.31061 | 1.286 | 2.548 | 101.21 | 20.85 | 17.696 | 96.175 | 107.6867 | 85.7722 | 0.004313 | 0.98363 | 0.023981 | 24.3807 | 0.46479 | 0.15898 | 0.012479 | 0.90005 | 0.90005 | 0.007348
2019-03-19 15:44 | 1553024631.7617 | 102 | 0.20713 | 0.001928 | 3.2e-05 | 92.1161 | 87.7141 | 119.495 | 21.2612 | 0.003016 | 0.00267 | 255.4413 | 0.26627 | 8.4039 | 0.003559 | 0.005412 | 0.005327 | 4.7e-05 | 0.001147 | 0.33196 | 1.3122 | 2.5479 | 101.21 | 20.85 | 17.693 | 94.3882 | 107.5301 | 86.047 | 0.004313 | 0.98287 | 0.023994 | 24.412 | 0.438 | 0.15841 | 0.011777 | 0.84989 | 0.84989 | 0.007754
2019-03-19 15:44 | 1553024632.7617 | 103 | 0.20711 | 0.001908 | 3.1e-05 | 92.2602 | 88.1274 | 119.9198 | 21.2612 | 0.002989 | 0.002628 | 255.4043 | 0.2669 | 8.4034 | 0.003573 | 0.005247 | 0.005154 | 5e-05 | 0.00115 | 0.31022 | 1.3531 | 2.5481 | 101.21 | 20.85 | 17.687 | 93.7285 | 107.8727 | 85.0292 | 0.004311 | 0.98567 | 0.024071 | 24.4207 | 0.40992 | 0.15878 | 0.011927 | 0.86355 | 0.86355 | 0.007649
2019-03-19 15:44 | 1553024633.7627 | 104 | 0.20703 | 0.001895 | 3e-05 | 92.4495 | 88.4641 | 120.1304 | 21.2612 | 0.003031 | 0.002686 | 255.398 | 0.2679 | 8.4031 | 0.003491 | 0.004641 | 0.004541 | 4.8e-05 | 0.00115 | 0.31188 | 1.3476 | 2.548 | 101.21 | 20.85 | 17.684 | 94.7397 | 105.8641 | 86.4271 | 0.00431 | 0.98184 | 0.024045 | 24.4898 | 0.39415 | 0.15937 | 0.01239 | 0.89616 | 0.89616 | 0.007398
2019-03-19 15:44 | 1553024634.7631 | 105 | 0.2073 | 0.001873 | 2.9e-05 | 92.6191 | 88.9169 | 120.1327 | 21.2612 | 0.003021 | 0.002651 | 255.3639 | 0.26634 | 8.4029 | 0.003445 | 0.005218 | 0.005122 | 4.9e-05 | 0.001154 | 0.31957 | 1.3247 | 2.548 | 101.21 | 20.85 | 17.68 | 94.5096 | 104.7583 | 85.5911 | 0.004309 | 0.98412 | 0.024073 | 24.4616 | 0.36488 | 0.15845 | 0.010785 | 0.78098 | 0.78098 | 0.00844
2019-03-19 15:44 | 1553024635.7629 | 106 | 0.20716 | 0.001857 | 2.9e-05 | 92.7713 | 89.0871 | 120.2731 | 21.2612 | 0.003068 | 0.002619 | 255.3413 | 0.26684 | 8.403 | 0.003479 | 0.004786 | 0.004732 | 4.8e-05 | 0.001153 | 0.32477 | 1.3192 | 2.5479 | 101.21 | 20.85 | 17.676 | 95.6439 | 105.5889 | 84.8248 | 0.004308 | 0.98623 | 0.024113 | 24.4497 | 0.36245 | 0.15874 | 0.011709 | 0.84936 | 0.84936 | 0.007775
2019-03-19 15:44 | 1553024636.7619 | 107 | 0.20724 | 0.001846 | 2.9e-05 | 92.9264 | 89.3649 | 120.6449 | 21.2612 | 0.003035 | 0.002677 | 255.3123 | 0.26772 | 8.4002 | 0.003526 | 0.005218 | 0.005117 | 4.8e-05 | 0.001156 | 0.32012 | 1.3325 | 2.5481 | 101.21 | 20.85 | 17.673 | 94.8399 | 106.7198 | 86.2279 | 0.004308 | 0.98238 | 0.024103 | 24.5354 | 0.34956 | 0.15926 | 0.011206 | 0.81248 | 0.81248 | 0.008155
2019-03-19 15:44 | 1553024637.762 | 108 | 0.20727 | 0.001824 | 3e-05 | 93.0933 | 89.4899 | 120.9268 | 21.2612 | 0.002912 | 0.002595 | 255.2655 | 0.26772 | 8.4025 | 0.003511 | 0.005288 | 0.005162 | 5.1e-05 | 0.001159 | 0.29927 | 1.3297 | 2.548 | 101.21 | 20.85 | 17.669 | 91.8794 | 106.3599 | 84.2341 | 0.004307 | 0.98786 | 0.024198 | 24.4958 | 0.3531 | 0.15927 | 0.011043 | 0.80375 | 0.80375 | 0.008243
2019-03-19 15:44 | 1553024638.7623 | 109 | 0.20725 | 0.001803 | 3e-05 | 93.2088 | 89.7334 | 121.1829 | 21.2612 | 0.002952 | 0.002661 | 255.2478 | 0.26761 | 8.4022 | 0.003444 | 0.004975 | 0.004944 | 5.2e-05 | 0.001155 | 0.30829 | 1.3707 | 2.5479 | 101.21 | 20.85 | 17.667 | 92.8383 | 104.733 | 85.822 | 0.004306 | 0.98349 | 0.02417 | 24.5761 | 0.33989 | 0.1592 | 0.011223 | 0.81593 | 0.81593 | 0.008117
2019-03-19 15:44 | 1553024639.762 | 110 | 0.20727 | 0.001796 | 3e-05 | 93.4714 | 90.0327 | 121.2308 | 21.2612 | 0.002956 | 0.002629 | 255.2415 | 0.268 | 8.4014 | 0.003469 | 0.004877 | 0.004731 | 5.5e-05 | 0.00116 | 0.31022 | 1.3619 | 2.5479 | 101.21 | 20.85 | 17.665 | 92.9352 | 105.3331 | 85.0557 | 0.004306 | 0.98559 | 0.024201 | 24.5547 | 0.33526 | 0.15943 | 0.011096 | 0.8077 | 0.8077 | 0.008212
2019-03-19 15:44 | 1553024640.763 | 111 | 0.20732 | 0.001775 | 3e-05 | 93.5902 | 90.2416 | 121.4145 | 21.2612 | 0.002957 | 0.002601 | 255.2114 | 0.26731 | 8.4008 | 0.003348 | 0.00501 | 0.004925 | 5.1e-05 | 0.001161 | 0.30287 | 1.3045 | 2.5478 | 101.21 | 20.85 | 17.661 | 92.9641 | 102.3987 | 84.3986 | 0.004305 | 0.9874 | 0.024242 | 24.5508 | 0.32589 | 0.15902 | 0.010838 | 0.7902 | 0.7902 | 0.008372
2019-03-19 15:44 | 1553024641.7631 | 112 | 0.20728 | 0.001762 | 3e-05 | 93.7841 | 90.6759 | 121.2812 | 21.2612 | 0.00298 | 0.002605 | 255.1629 | 0.26656 | 8.4015 | 0.003381 | 0.004545 | 0.004418 | 5.1e-05 | 0.001159 | 0.30182 | 1.3232 | 2.5479 | 101.21 | 20.85 | 17.66 | 93.5267 | 103.2039 | 84.4775 | 0.004305 | 0.98719 | 0.024226 | 24.54 | 0.30146 | 0.15858 | 0.011047 | 0.80494 | 0.80494 | 0.008195
2019-03-19 15:44 | 1553024642.7636 | 113 | 0.20734 | 0.001744 | 3e-05 | 93.8657 | 90.9797 | 120.8575 | 21.2612 | 0.002938 | 0.002617 | 255.1335 | 0.2671 | 8.4048 | 0.003526 | 0.00485 | 0.004792 | 5.2e-05 | 0.001161 | 0.29214 | 1.3173 | 2.548 | 101.21 | 20.85 | 17.655 | 92.4957 | 106.7324 | 84.7747 | 0.004303 | 0.98637 | 0.024173 | 24.5073 | 0.27933 | 0.1589 | 0.010748 | 0.78149 | 0.78149 | 0.008458
2019-03-19 15:44 | 1553024643.7633 | 114 | 0.20733 | 0.001735 | 2.9e-05 | 93.9687 | 91.0948 | 120.9101 | 21.2612 | 0.002931 | 0.002588 | 255.0746 | 0.26788 | 8.4047 | 0.003447 | 0.005121 | 0.004926 | 4.8e-05 | 0.001162 | 0.31526 | 1.2737 | 2.548 | 101.21 | 20.85 | 17.653 | 92.3313 | 104.807 | 84.0687 | 0.004303 | 0.98832 | 0.024202 | 24.4884 | 0.27782 | 0.15936 | 0.010781 | 0.78483 | 0.78483 | 0.008447
2019-03-19 15:44 | 1553024644.7616 | 115 | 0.20736 | 0.001717 | 2.9e-05 | 93.9825 | 91.2397 | 121.0531 | 21.2612 | 0.002905 | 0.002593 | 255.0541 | 0.26759 | 8.4028 | 0.003435 | 0.005 | 0.004983 | 5.2e-05 | 0.00116 | 0.31116 | 1.3032 | 2.5481 | 101.21 | 20.85 | 17.65 | 91.7049 | 104.5081 | 84.206 | 0.004302 | 0.98794 | 0.024212 | 24.5076 | 0.26493 | 0.15919 | 0.01062 | 0.77345 | 0.77345 | 0.008562
2019-03-19 15:44 | 1553024645.7617 | 116 | 0.20736 | 0.001701 | 2.9e-05 | 94.1208 | 91.4993 | 120.9561 | 21.2612 | 0.002942 | 0.002605 | 255.0308 | 0.26796 | 8.4013 | 0.003485 | 0.00536 | 0.005246 | 5.3e-05 | 0.001159 | 0.30215 | 1.3177 | 2.548 | 101.21 | 20.85 | 17.645 | 92.6013 | 105.7368 | 84.4782 | 0.004301 | 0.98719 | 0.024193 | 24.5071 | 0.25267 | 0.15941 | 0.010694 | 0.77824 | 0.77824 | 0.008521
2019-03-19 15:44 | 1553024646.7624 | 117 | 0.20738 | 0.001685 | 2.8e-05 | 94.231 | 91.7082 | 120.7841 | 21.2612 | 0.002903 | 0.002623 | 255.0021 | 0.26771 | 8.3988 | 0.003362 | 0.004868 | 0.004836 | 5.3e-05 | 0.001166 | 0.32084 | 1.2984 | 2.5481 | 101.21 | 20.84 | 17.642 | 91.6644 | 102.7537 | 84.9246 | 0.004297 | 0.98596 | 0.024161 | 24.505 | 0.24273 | 0.15926 | 0.010567 | 0.76799 | 0.76799 | 0.008627
2019-03-19 15:44 | 1553024647.7618 | 118 | 0.20741 | 0.001679 | 2.8e-05 | 94.426 | 92.0216 | 120.846 | 21.2612 | 0.002896 | 0.002602 | 254.9686 | 0.26727 | 8.402 | 0.003414 | 0.004587 | 0.004479 | 5.4e-05 | 0.001157 | 0.31996 | 1.3006 | 2.548 | 101.21 | 20.85 | 17.634 | 91.5032 | 104.0072 | 84.4121 | 0.004298 | 0.98737 | 0.024184 | 24.4937 | 0.23071 | 0.159 | 0.010367 | 0.75422 | 0.75422 | 0.00877
2019-03-19 15:44 | 1553024648.7624 | 119 | 0.2074 | 0.001662 | 2.7e-05 | 94.5077 | 92.348 | 120.9221 | 21.2612 | 0.00291 | 0.00255 | 254.9195 | 0.26733 | 8.4016 | 0.003446 | 0.004947 | 0.004758 | 5.3e-05 | 0.001156 | 0.32117 | 1.3231 | 2.548 | 101.21 | 20.85 | 17.629 | 91.8199 | 104.7864 | 83.1708 | 0.004297 | 0.99081 | 0.024234 | 24.4589 | 0.20677 | 0.15903 | 0.010476 | 0.76378 | 0.76378 | 0.008662
2019-03-19 15:44 | 1553024649.763 | 120 | 0.20747 | 0.001648 | 2.7e-05 | 94.541 | 92.5682 | 120.9923 | 21.2612 | 0.002873 | 0.002604 | 254.9015 | 0.26784 | 8.4039 | 0.003354 | 0.005322 | 0.005255 | 5.2e-05 | 0.001156 | 0.30901 | 1.3126 | 2.5479 | 101.21 | 20.85 | 17.624 | 90.9476 | 102.5472 | 84.4744 | 0.004296 | 0.9872 | 0.024197 | 24.5107 | 0.18862 | 0.15934 | 0.010062 | 0.73246 | 0.73246 | 0.00905
2019-03-19 15:44 | 1553024650.7626 | 121 | 0.20747 | 0.001632 | 2.7e-05 | 94.6646 | 92.76 | 121.02 | 21.2612 | 0.002842 | 0.002582 | 254.8893 | 0.26756 | 8.4024 | 0.003448 | 0.005713 | 0.005637 | 5.4e-05 | 0.001155 | 0.33915 | 1.3138 | 2.548 | 101.21 | 20.84 | 17.615 | 90.1972 | 104.8375 | 83.9453 | 0.004291 | 0.98866 | 0.024218 | 24.4953 | 0.1818 | 0.15917 | 0.010087 | 0.7349 | 0.7349 | 0.00901
2019-03-19 15:44 | 1553024651.7626 | 122 | 0.20746 | 0.001622 | 2.7e-05 | 94.7913 | 92.9281 | 121.2031 | 21.2612 | 0.002837 | 0.002556 | 254.8887 | 0.26785 | 8.4031 | 0.003347 | 0.004226 | 0.004131 | 5.4e-05 | 0.001142 | 0.29806 | 1.362 | 2.5481 | 101.21 | 20.84 | 17.609 | 90.0659 | 102.3862 | 83.3165 | 0.004289 | 0.9904 | 0.024257 | 24.4923 | 0.17755 | 0.15934 | 0.010148 | 0.74051 | 0.74051 | 0.008951
2019-03-19 15:44 | 1553024652.7617 | 123 | 0.20748 | 0.001606 | 2.7e-05 | 94.9083 | 93.0583 | 121.4922 | 21.2612 | 0.002891 | 0.002551 | 254.866 | 0.26778 | 8.4018 | 0.003337 | 0.004636 | 0.004531 | 5.6e-05 | 0.001149 | 0.3131 | 1.3235 | 2.548 | 101.21 | 20.84 | 17.604 | 91.3781 | 102.1293 | 83.1833 | 0.004288 | 0.99077 | 0.024291 | 24.5169 | 0.17607 | 0.1593 | 0.010053 | 0.73462 | 0.73462 | 0.009021
2019-03-19 15:44 | 1553024653.7617 | 124 | 0.2075 | 0.001594 | 2.8e-05 | 95.0106 | 93.1086 | 121.9145 | 21.2612 | 0.002897 | 0.002566 | 254.8322 | 0.26762 | 8.401 | 0.003355 | 0.004947 | 0.004903 | 5.5e-05 | 0.001151 | 0.32803 | 1.3428 | 2.5482 | 101.21 | 20.84 | 17.591 | 91.5198 | 102.5832 | 83.5529 | 0.004285 | 0.98975 | 0.02432 | 24.5722 | 0.18088 | 0.15921 | 0.00996 | 0.72869 | 0.72869 | 0.009089
2019-03-19 15:44 | 1553024654.7628 | 125 | 0.2075 | 0.001583 | 2.7e-05 | 95.1121 | 93.1785 | 121.8586 | 21.2612 | 0.002857 | 0.002526 | 254.7756 | 0.26698 | 8.4023 | 0.003333 | 0.004769 | 0.004708 | 5.5e-05 | 0.001147 | 0.31166 | 1.3118 | 2.5481 | 101.21 | 20.84 | 17.585 | 90.5628 | 102.0327 | 82.5755 | 0.004284 | 0.99247 | 0.024348 | 24.5329 | 0.18371 | 0.15883 | 0.009948 | 0.72861 | 0.72861 | 0.009068
2019-03-19 15:44 | 1553024655.7632 | 126 | 0.2075 | 0.001571 | 2.7e-05 | 95.3287 | 93.5453 | 121.857 | 21.2612 | 0.002891 | 0.0026 | 254.7325 | 0.26699 | 8.4033 | 0.003292 | 0.004567 | 0.004363 | 5.6e-05 | 0.001156 | 0.30829 | 1.3212 | 2.5478 | 101.21 | 20.83 | 17.574 | 91.3593 | 101.0473 | 84.3635 | 0.004278 | 0.9875 | 0.024287 | 24.5943 | 0.16892 | 0.15883 | 0.010007 | 0.73111 | 0.73111 | 0.009037
2019-03-19 15:44 | 1553024656.763 | 127 | 0.20752 | 0.00156 | 2.7e-05 | 95.384 | 93.4703 | 121.7772 | 21.2612 | 0.002824 | 0.002514 | 254.6948 | 0.2674 | 8.4008 | 0.003238 | 0.005086 | 0.004984 | 5.4e-05 | 0.001158 | 0.29286 | 1.3202 | 2.5478 | 101.21 | 20.83 | 17.57 | 89.755 | 99.7431 | 82.2871 | 0.004277 | 0.99327 | 0.02435 | 24.5147 | 0.18127 | 0.15907 | 0.009916 | 0.72638 | 0.72638 | 0.00911
2019-03-19 15:44 | 1553024657.762 | 128 | 0.20754 | 0.00155 | 2.7e-05 | 95.5631 | 93.771 | 121.78 | 21.2612 | 0.002825 | 0.002567 | 254.6663 | 0.26762 | 8.4045 | 0.003256 | 0.004538 | 0.004508 | 5.8e-05 | 0.001157 | 0.29159 | 1.3289 | 2.548 | 101.21 | 20.84 | 17.566 | 89.7822 | 100.191 | 83.5827 | 0.004279 | 0.98966 | 0.024306 | 24.5597 | 0.16933 | 0.15921 | 0.009788 | 0.71571 | 0.71571 | 0.009254
2019-03-19 15:44 | 1553024658.7621 | 129 | 0.20756 | 0.001538 | 2.6e-05 | 95.6033 | 93.9957 | 121.6523 | 21.2612 | 0.002839 | 0.002539 | 254.6447 | 0.26736 | 8.403 | 0.003186 | 0.004598 | 0.004518 | 5.6e-05 | 0.001159 | 0.29878 | 1.3617 | 2.548 | 101.21 | 20.83 | 17.561 | 90.1143 | 98.4883 | 82.889 | 0.004275 | 0.99159 | 0.024317 | 24.5229 | 0.15169 | 0.15905 | 0.009646 | 0.70569 | 0.70569 | 0.009376
2019-03-19 15:44 | 1553024659.7634 | 130 | 0.20753 | 0.00153 | 2.5e-05 | 95.7663 | 94.3478 | 121.626 | 21.2612 | 0.00281 | 0.002565 | 254.5736 | 0.26804 | 8.403 | 0.003226 | 0.004424 | 0.004264 | 5.8e-05 | 0.001157 | 0.32228 | 1.348 | 2.5481 | 101.21 | 20.83 | 17.56 | 89.4212 | 99.4489 | 83.5365 | 0.004275 | 0.98979 | 0.024292 | 24.5425 | 0.13348 | 0.15945 | 0.009847 | 0.71966 | 0.71966 | 0.009217
2019-03-19 15:44 | 1553024660.7622 | 131 | 0.20758 | 0.001529 | 2.5e-05 | 95.9508 | 94.6564 | 121.3831 | 21.2612 | 0.002811 | 0.002498 | 254.5611 | 0.26732 | 8.4036 | 0.003176 | 0.005072 | 0.005022 | 5.5e-05 | 0.001164 | 0.3063 | 1.3337 | 2.5482 | 101.21 | 20.83 | 17.56 | 89.4556 | 98.2382 | 81.9041 | 0.004275 | 0.99434 | 0.024323 | 24.4618 | 0.1215 | 0.15903 | 0.009591 | 0.70189 | 0.70189 | 0.009425
2019-03-19 15:44 | 1553024661.7611 | 132 | 0.20761 | 0.001517 | 2.4e-05 | 95.9962 | 94.9064 | 121.4918 | 21.2612 | 0.002838 | 0.002527 | 254.5111 | 0.26798 | 8.4019 | 0.003255 | 0.004612 | 0.0044 | 5.8e-05 | 0.001158 | 0.3465 | 1.3333 | 2.5482 | 101.21 | 20.83 | 17.561 | 90.0869 | 100.1472 | 82.6207 | 0.004275 | 0.99234 | 0.02431 | 24.4975 | 0.10212 | 0.15942 | 0.009385 | 0.68646 | 0.68646 | 0.009661
2019-03-19 15:44 | 1553024662.7616 | 133 | 0.2076 | 0.00151 | 2.4e-05 | 96.0601 | 94.8252 | 121.4603 | 21.2612 | 0.002836 | 0.002522 | 254.473 | 0.2673 | 8.4036 | 0.003229 | 0.004678 | 0.004621 | 5.6e-05 | 0.00115 | 0.32516 | 1.3372 | 2.548 | 101.21 | 20.83 | 17.564 | 90.044 | 99.5163 | 82.4977 | 0.004276 | 0.99268 | 0.024311 | 24.4901 | 0.11573 | 0.15902 | 0.009449 | 0.69115 | 0.69115 | 0.009571
2019-03-19 15:44 | 1553024663.7615 | 134 | 0.20767 | 0.001503 | 2.4e-05 | 96.1165 | 94.9238 | 121.2343 | 21.2612 | 0.002807 | 0.002526 | 254.4525 | 0.26794 | 8.4029 | 0.003159 | 0.004808 | 0.004607 | 5.2e-05 | 0.001144 | 0.33516 | 1.3412 | 2.548 | 101.21 | 20.84 | 17.573 | 89.3564 | 97.8273 | 82.5974 | 0.004281 | 0.9924 | 0.024285 | 24.4707 | 0.11168 | 0.1594 | 0.00908 | 0.66344 | 0.66344 | 0.009995
2019-03-19 15:44 | 1553024664.7614 | 135 | 0.20769 | 0.00149 | 2.3e-05 | 96.169 | 94.9851 | 121.1801 | 21.2612 | 0.002745 | 0.002548 | 254.438 | 0.26784 | 8.4034 | 0.003213 | 0.004985 | 0.004836 | 5.5e-05 | 0.001141 | 0.32477 | 1.3239 | 2.5479 | 101.21 | 20.84 | 17.578 | 87.8638 | 99.1368 | 83.1131 | 0.004282 | 0.99097 | 0.024262 | 24.483 | 0.1108 | 0.15934 | 0.008938 | 0.65251 | 0.65251 | 0.010158
2019-03-19 15:44 | 1553024665.7623 | 136 | 0.20767 | 0.001481 | 2.4e-05 | 96.223 | 95.1711 | 121.3599 | 21.2612 | 0.00288 | 0.002513 | 254.4261 | 0.26729 | 8.4031 | 0.00316 | 0.004602 | 0.004577 | 5.3e-05 | 0.001141 | 0.30381 | 1.3348 | 2.5479 | 101.21 | 20.84 | 17.59 | 91.1047 | 97.858 | 82.2683 | 0.004285 | 0.99332 | 0.024309 | 24.472 | 0.098322 | 0.15901 | 0.009099 | 0.66551 | 0.66551 | 0.009939
2019-03-19 15:44 | 1553024666.7616 | 137 | 0.20767 | 0.001472 | 2.4e-05 | 96.2993 | 95.2146 | 121.5263 | 21.2612 | 0.002837 | 0.002526 | 254.3943 | 0.26739 | 8.4026 | 0.003225 | 0.004676 | 0.004591 | 5.4e-05 | 0.001138 | 0.31852 | 1.3809 | 2.5481 | 101.21 | 20.84 | 17.596 | 90.0611 | 99.4419 | 82.5818 | 0.004286 | 0.99245 | 0.024315 | 24.4996 | 0.10132 | 0.15907 | 0.009095 | 0.66534 | 0.66534 | 0.009946
2019-03-19 15:44 | 1553024667.7625 | 138 | 0.20763 | 0.00146 | 2.4e-05 | 96.3584 | 95.5722 | 121.8401 | 21.2612 | 0.002802 | 0.002455 | 254.3742 | 0.26645 | 8.4016 | 0.003189 | 0.005247 | 0.005099 | 5.4e-05 | 0.001143 | 0.33964 | 1.3336 | 2.5481 | 101.21 | 20.84 | 17.601 | 89.2189 | 98.5596 | 80.8809 | 0.004288 | 0.99722 | 0.024404 | 24.4725 | 0.073281 | 0.15851 | 0.009362 | 0.68747 | 0.68747 | 0.009592
2019-03-19 15:44 | 1553024668.7629 | 139 | 0.20772 | 0.001446 | 2.4e-05 | 96.4763 | 95.3143 | 121.9347 | 21.2612 | 0.002739 | 0.002573 | 254.3553 | 0.26671 | 8.402 | 0.003168 | 0.004117 | 0.004007 | 5.2e-05 | 0.001138 | 0.31133 | 1.3148 | 2.5481 | 101.21 | 20.83 | 17.613 | 87.7052 | 98.0609 | 83.7284 | 0.004288 | 0.98926 | 0.024316 | 24.5803 | 0.10839 | 0.15867 | 0.008848 | 0.64735 | 0.64735 | 0.010196
2019-03-19 15:44 | 1553024669.7616 | 140 | 0.20767 | 0.001441 | 2.4e-05 | 96.546 | 95.6003 | 122.0131 | 21.2612 | 0.002777 | 0.00249 | 254.347 | 0.2675 | 8.4051 | 0.003125 | 0.003936 | 0.00388 | 5.3e-05 | 0.001139 | 0.32549 | 1.2636 | 2.5481 | 101.21 | 20.83 | 17.619 | 88.6269 | 97.0234 | 81.7176 | 0.004289 | 0.99486 | 0.024393 | 24.5188 | 0.088047 | 0.15914 | 0.009158 | 0.67212 | 0.67212 | 0.00985
2019-03-19 15:44 | 1553024670.762 | 141 | 0.20768 | 0.001432 | 2.3e-05 | 96.6801 | 95.7297 | 122.3692 | 21.2612 | 0.002757 | 0.002448 | 254.289 | 0.26698 | 8.4025 | 0.003129 | 0.004553 | 0.004479 | 5.3e-05 | 0.001139 | 0.30596 | 1.2834 | 2.548 | 101.21 | 20.83 | 17.626 | 88.1355 | 97.1009 | 80.7132 | 0.004291 | 0.99769 | 0.024463 | 24.5198 | 0.088361 | 0.15883 | 0.009109 | 0.6705 | 0.6705 | 0.009854
2019-03-19 15:44 | 1553024671.7617 | 142 | 0.20769 | 0.001428 | 2.4e-05 | 96.7674 | 95.9852 | 122.1258 | 21.2612 | 0.002775 | 0.002501 | 254.2549 | 0.26738 | 8.4033 | 0.003124 | 0.004957 | 0.004872 | 5.3e-05 | 0.001138 | 0.29502 | 1.3289 | 2.5481 | 101.21 | 20.83 | 17.638 | 88.5713 | 96.9795 | 81.982 | 0.004294 | 0.99412 | 0.024395 | 24.5393 | 0.072597 | 0.15906 | 0.009039 | 0.66348 | 0.66348 | 0.009973
2019-03-19 15:45 | 1553024672.7627 | 143 | 0.20773 | 0.001413 | 2.3e-05 | 96.8761 | 96.0958 | 122.1114 | 21.2612 | 0.00277 | 0.002515 | 254.2264 | 0.26696 | 8.4058 | 0.003195 | 0.004953 | 0.004812 | 5.2e-05 | 0.001134 | 0.30807 | 1.3227 | 2.548 | 101.21 | 20.83 | 17.643 | 88.4581 | 98.7172 | 82.3181 | 0.004295 | 0.99318 | 0.024382 | 24.5494 | 0.07233 | 0.15881 | 0.008814 | 0.64658 | 0.64658 | 0.010218
2019-03-19 15:45 | 1553024673.7629 | 144 | 0.20774 | 0.001411 | 2.3e-05 | 97.0363 | 96.1575 | 121.8108 | 21.2612 | 0.002759 | 0.002496 | 254.1781 | 0.26715 | 8.4017 | 0.003276 | 0.004594 | 0.004536 | 5.1e-05 | 0.001129 | 0.31332 | 1.2985 | 2.548 | 101.21 | 20.83 | 17.652 | 88.1895 | 100.6706 | 81.8722 | 0.004297 | 0.99443 | 0.024367 | 24.5038 | 0.081378 | 0.15893 | 0.00874 | 0.64076 | 0.64076 | 0.010318
2019-03-19 15:45 | 1553024674.7619 | 145 | 0.20771 | 0.001407 | 2.4e-05 | 97.1089 | 96.2393 | 121.8136 | 21.2612 | 0.002765 | 0.002497 | 254.1493 | 0.26798 | 8.4034 | 0.003171 | 0.004907 | 0.004903 | 5e-05 | 0.001131 | 0.31006 | 1.3393 | 2.5481 | 101.21 | 20.83 | 17.656 | 88.3409 | 98.1157 | 81.8892 | 0.004298 | 0.99438 | 0.024367 | 24.5047 | 0.080458 | 0.15942 | 0.008932 | 0.65482 | 0.65482 | 0.010128
2019-03-19 15:45 | 1553024675.7619 | 146 | 0.20776 | 0.001396 | 2.3e-05 | 97.1597 | 96.601 | 121.7598 | 21.2612 | 0.002743 | 0.002516 | 254.1063 | 0.26675 | 8.4024 | 0.003146 | 0.004871 | 0.00468 | 5.1e-05 | 0.001136 | 0.30198 | 1.3488 | 2.548 | 101.21 | 20.83 | 17.658 | 87.8091 | 97.5147 | 82.3525 | 0.004299 | 0.99309 | 0.024346 | 24.5152 | 0.051578 | 0.15869 | 0.008645 | 0.63323 | 0.63323 | 0.010425
2019-03-19 15:45 | 1553024676.7623 | 147 | 0.20774 | 0.001395 | 2.3e-05 | 97.286 | 96.5044 | 121.8345 | 21.2612 | 0.002784 | 0.002462 | 254.0711 | 0.26655 | 8.403 | 0.003142 | 0.005022 | 0.005045 | 4.8e-05 | 0.001135 | 0.32803 | 1.3397 | 2.5481 | 101.21 | 20.83 | 17.661 | 88.8048 | 97.4277 | 81.054 | 0.004299 | 0.99673 | 0.024398 | 24.4779 | 0.072149 | 0.15857 | 0.008795 | 0.64565 | 0.64565 | 0.010217
2019-03-19 15:45 | 1553024677.761 | 148 | 0.20771 | 0.001388 | 2.3e-05 | 97.3058 | 96.7192 | 121.4894 | 21.2612 | 0.002719 | 0.002436 | 254.0276 | 0.26714 | 8.4031 | 0.003135 | 0.004332 | 0.004209 | 5e-05 | 0.001139 | 0.31526 | 1.3553 | 2.5481 | 101.21 | 20.83 | 17.663 | 87.2327 | 97.2459 | 80.4217 | 0.0043 | 0.99851 | 0.024385 | 24.4214 | 0.054088 | 0.15892 | 0.008982 | 0.65907 | 0.65907 | 0.010031
2019-03-19 15:45 | 1553024678.7614 | 149 | 0.20775 | 0.001382 | 2.2e-05 | 97.349 | 96.6746 | 121.3805 | 21.2612 | 0.002745 | 0.00246 | 254.0115 | 0.26719 | 8.4019 | 0.003122 | 0.005214 | 0.005138 | 4.7e-05 | 0.001138 | 0.31437 | 1.3299 | 2.5481 | 101.21 | 20.83 | 17.668 | 87.8513 | 96.9379 | 80.9963 | 0.004301 | 0.99689 | 0.024354 | 24.4303 | 0.062178 | 0.15895 | 0.00871 | 0.63831 | 0.63831 | 0.010359
2019-03-19 15:45 | 1553024679.7622 | 150 | 0.20777 | 0.001375 | 2.2e-05 | 97.4099 | 96.8253 | 121.2153 | 21.2612 | 0.002748 | 0.002498 | 253.9984 | 0.26771 | 8.4017 | 0.00315 | 0.005316 | 0.00519 | 4.8e-05 | 0.001138 | 0.29767 | 1.3267 | 2.5481 | 101.21 | 20.84 | 17.672 | 87.9199 | 97.6096 | 81.9094 | 0.004305 | 0.99433 | 0.024306 | 24.4451 | 0.053842 | 0.15926 | 0.008612 | 0.62983 | 0.62983 | 0.010519
2019-03-19 15:45 | 1553024680.7612 | 151 | 0.20781 | 0.001373 | 2.3e-05 | 97.5453 | 97.0572 | 121.1575 | 21.2612 | 0.002759 | 0.002485 | 253.9735 | 0.26803 | 8.4031 | 0.003162 | 0.004705 | 0.004653 | 4.6e-05 | 0.001136 | 0.32189 | 1.317 | 2.548 | 101.21 | 20.84 | 17.676 | 88.1863 | 97.9198 | 81.6061 | 0.004306 | 0.99518 | 0.024311 | 24.4288 | 0.044868 | 0.15945 | 0.008354 | 0.61103 | 0.61103 | 0.010856
2019-03-19 15:45 | 1553024681.7619 | 152 | 0.2078 | 0.001367 | 2.3e-05 | 97.6659 | 96.986 | 121.4129 | 21.2612 | 0.002719 | 0.00246 | 253.9643 | 0.26766 | 8.4033 | 0.003114 | 0.004513 | 0.0044 | 4.6e-05 | 0.001136 | 0.33019 | 1.3052 | 2.548 | 101.21 | 20.84 | 17.684 | 87.2407 | 96.7609 | 81.0095 | 0.004308 | 0.99685 | 0.024357 | 24.434 | 0.062491 | 0.15923 | 0.008436 | 0.61821 | 0.61821 | 0.010715
2019-03-19 15:45 | 1553024682.7613 | 153 | 0.20779 | 0.001356 | 2.2e-05 | 97.7007 | 97.0832 | 121.5289 | 21.2612 | 0.002734 | 0.002433 | 253.9034 | 0.2677 | 8.4004 | 0.003088 | 0.004867 | 0.004837 | 4.9e-05 | 0.001137 | 0.29336 | 1.3079 | 2.5481 | 101.21 | 20.84 | 17.689 | 87.5928 | 96.1279 | 80.3467 | 0.004309 | 0.99872 | 0.024392 | 24.4228 | 0.056713 | 0.15925 | 0.008524 | 0.62558 | 0.62558 | 0.01059
2019-03-19 15:45 | 1553024683.7624 | 154 | 0.20783 | 0.001353 | 2.2e-05 | 97.78 | 97.2521 | 121.6124 | 21.2612 | 0.002783 | 0.002489 | 253.8696 | 0.26659 | 8.4032 | 0.00315 | 0.004531 | 0.004353 | 4.9e-05 | 0.001138 | 0.35026 | 1.3269 | 2.5479 | 101.21 | 20.84 | 17.692 | 88.7651 | 97.6181 | 81.6981 | 0.00431 | 0.99492 | 0.024353 | 24.4778 | 0.04842 | 0.15859 | 0.008259 | 0.60521 | 0.60521 | 0.010901
2019-03-19 15:45 | 1553024684.7624 | 155 | 0.20782 | 0.001343 | 2.2e-05 | 97.8541 | 97.2743 | 121.7375 | 21.2612 | 0.002744 | 0.002439 | 253.8181 | 0.26749 | 8.4032 | 0.003063 | 0.004837 | 0.00468 | 4.6e-05 | 0.001137 | 0.31741 | 1.3241 | 2.548 | 101.21 | 20.84 | 17.702 | 87.8419 | 95.5117 | 80.495 | 0.004312 | 0.9983 | 0.024407 | 24.4489 | 0.053151 | 0.15913 | 0.008386 | 0.61589 | 0.61589 | 0.010748
2019-03-19 15:45 | 1553024685.7629 | 156 | 0.20784 | 0.001337 | 2.2e-05 | 97.8979 | 97.3983 | 121.903 | 21.2612 | 0.002732 | 0.002457 | 253.7917 | 0.2671 | 8.4006 | 0.003059 | 0.004734 | 0.004707 | 4.7e-05 | 0.001138 | 0.29927 | 1.306 | 2.5481 | 101.21 | 20.84 | 17.706 | 87.5319 | 95.4175 | 80.9393 | 0.004313 | 0.99705 | 0.024409 | 24.4808 | 0.045762 | 0.1589 | 0.008229 | 0.60439 | 0.60439 | 0.010937
2019-03-19 15:45 | 1553024686.7616 | 157 | 0.20785 | 0.001327 | 2.2e-05 | 97.9 | 97.6483 | 121.9288 | 21.2612 | 0.00273 | 0.002486 | 253.7657 | 0.26733 | 8.4039 | 0.003043 | 0.004693 | 0.004607 | 4.9e-05 | 0.001137 | 0.29607 | 1.3291 | 2.548 | 101.21 | 20.84 | 17.717 | 87.4835 | 95.0247 | 81.6203 | 0.004316 | 0.99514 | 0.024388 | 24.507 | 0.023023 | 0.15903 | 0.008202 | 0.60189 | 0.60189 | 0.010992
2019-03-19 15:45 | 1553024687.7616 | 158 | 0.20787 | 0.001321 | 2.1e-05 | 98.0111 | 97.4377 | 122.2584 | 21.2612 | 0.002754 | 0.002496 | 253.7365 | 0.26824 | 8.4023 | 0.003021 | 0.004376 | 0.004275 | 4.5e-05 | 0.001139 | 0.32803 | 1.3179 | 2.5481 | 101.21 | 20.84 | 17.722 | 88.0792 | 94.5002 | 81.8635 | 0.004317 | 0.99446 | 0.024412 | 24.5485 | 0.052477 | 0.15958 | 0.008098 | 0.59484 | 0.59484 | 0.01116
2019-03-19 15:45 | 1553024688.7617 | 159 | 0.20789 | 0.001318 | 2.1e-05 | 98.0021 | 97.61 | 122.2551 | 21.2612 | 0.002707 | 0.002469 | 253.6959 | 0.26731 | 8.4014 | 0.003071 | 0.005253 | 0.005133 | 4.6e-05 | 0.001142 | 0.31492 | 1.3154 | 2.5481 | 101.21 | 20.84 | 17.727 | 86.9346 | 95.7199 | 81.2249 | 0.004318 | 0.99625 | 0.024434 | 24.5261 | 0.035855 | 0.15902 | 0.007967 | 0.58575 | 0.58575 | 0.011294
2019-03-19 15:45 | 1553024689.762 | 160 | 0.20789 | 0.001311 | 2.1e-05 | 98.0948 | 97.5196 | 122.226 | 21.2612 | 0.002726 | 0.00248 | 253.6578 | 0.26788 | 8.4011 | 0.003083 | 0.004952 | 0.004856 | 4.7e-05 | 0.001144 | 0.31332 | 1.3533 | 2.5481 | 101.21 | 20.83 | 17.734 | 87.4022 | 96.0082 | 81.4799 | 0.004317 | 0.99553 | 0.024422 | 24.532 | 0.052604 | 0.15936 | 0.007998 | 0.58776 | 0.58776 | 0.011279
2019-03-19 15:45 | 1553024690.7608 | 161 | 0.20788 | 0.001302 | 2.1e-05 | 98.0972 | 97.8511 | 122.1889 | 21.2612 | 0.002695 | 0.002468 | 253.6194 | 0.26785 | 8.4043 | 0.003152 | 0.004366 | 0.004209 | 4.4e-05 | 0.001139 | 0.30453 | 1.3271 | 2.548 | 101.21 | 20.83 | 17.736 | 86.6465 | 97.6582 | 81.1835 | 0.004318 | 0.99636 | 0.024429 | 24.518 | 0.022466 | 0.15934 | 0.008039 | 0.59094 | 0.59094 | 0.011217
2019-03-19 15:45 | 1553024691.7619 | 162 | 0.20789 | 0.001294 | 2.1e-05 | 98.1773 | 97.7042 | 122.1725 | 21.2612 | 0.002664 | 0.002356 | 253.5843 | 0.26724 | 8.4031 | 0.003115 | 0.004614 | 0.004529 | 4.6e-05 | 0.001132 | 0.31382 | 1.3428 | 2.5481 | 101.21 | 20.83 | 17.736 | 85.9142 | 96.7718 | 78.5067 | 0.004318 | 1.0039 | 0.02452 | 24.4236 | 0.04321 | 0.15898 | 0.007985 | 0.5892 | 0.5892 | 0.011225
2019-03-19 15:45 | 1553024692.7608 | 163 | 0.20788 | 0.001293 | 2.1e-05 | 98.2412 | 97.7597 | 122.3767 | 21.2612 | 0.002708 | 0.002442 | 253.5802 | 0.26744 | 8.4027 | 0.003089 | 0.004906 | 0.004756 | 4.5e-05 | 0.001131 | 0.32029 | 1.3552 | 2.5482 | 101.21 | 20.83 | 17.734 | 86.9673 | 96.1359 | 80.5769 | 0.004317 | 0.99807 | 0.024469 | 24.5158 | 0.043954 | 0.1591 | 0.008059 | 0.59335 | 0.59335 | 0.011155
2019-03-19 15:45 | 1553024693.7613 | 164 | 0.20789 | 0.00128 | 2.1e-05 | 98.3289 | 97.9854 | 122.6072 | 21.2612 | 0.002682 | 0.002483 | 253.5646 | 0.26688 | 8.4055 | 0.003161 | 0.004344 | 0.004308 | 4.5e-05 | 0.001132 | 0.32173 | 1.3324 | 2.5481 | 101.21 | 20.83 | 17.732 | 86.328 | 97.8792 | 81.557 | 0.004317 | 0.99531 | 0.024458 | 24.5729 | 0.031302 | 0.15877 | 0.008032 | 0.59113 | 0.59113 | 0.011173
2019-03-19 15:45 | 1553024694.762 | 165 | 0.20788 | 0.001273 | 2.1e-05 | 98.4009 | 98.0902 | 122.4311 | 21.2612 | 0.002699 | 0.002471 | 253.5581 | 0.26692 | 8.4035 | 0.003167 | 0.004881 | 0.004814 | 4.5e-05 | 0.001129 | 0.3178 | 1.3478 | 2.548 | 101.21 | 20.83 | 17.726 | 86.7589 | 98.0375 | 81.2652 | 0.004315 | 0.99613 | 0.02445 | 24.5451 | 0.028289 | 0.15879 | 0.008066 | 0.59348 | 0.59348 | 0.01113
2019-03-19 15:45 | 1553024695.8164 | 166 | 0.20789 | 0.001272 | 2.1e-05 | 98.4987 | 98.0131 | 122.3171 | 21.2612 | 0.002705 | 0.002366 | 253.5289 | 0.26675 | 8.4043 | 0.003029 | 0.004815 | 0.004662 | 4.4e-05 | 0.001129 | 0.30287 | 1.3323 | 2.5481 | 101.21 | 20.83 | 17.723 | 86.8939 | 94.6895 | 78.7484 | 0.004315 | 1.0033 | 0.024526 | 24.4464 | 0.044208 | 0.15869 | 0.008009 | 0.59108 | 0.59108 | 0.011168
2019-03-19 15:45 | 1553024696.7681 | 167 | 0.20791 | 0.001271 | 2.2e-05 | 98.5701 | 98.2032 | 122.4558 | 21.2612 | 0.002644 | 0.002412 | 253.5217 | 0.26746 | 8.4013 | 0.003095 | 0.004412 | 0.004354 | 4.4e-05 | 0.001133 | 0.32117 | 1.3245 | 2.5482 | 101.21 | 20.83 | 17.72 | 85.428 | 96.3003 | 79.8564 | 0.004314 | 1.0001 | 0.024501 | 24.4987 | 0.033362 | 0.15911 | 0.007927 | 0.58435 | 0.58435 | 0.011327
2019-03-19 15:45 | 1553024697.7682 | 168 | 0.2079 | 0.001263 | 2.2e-05 | 98.6992 | 98.2357 | 122.2567 | 21.2612 | 0.002619 | 0.002451 | 253.5138 | 0.26791 | 8.4026 | 0.003161 | 0.004564 | 0.00447 | 4.6e-05 | 0.00113 | 0.29983 | 1.3192 | 2.5482 | 101.21 | 20.83 | 17.713 | 84.8349 | 97.8768 | 80.7927 | 0.004312 | 0.99747 | 0.024449 | 24.5113 | 0.042101 | 0.15938 | 0.007995 | 0.58815 | 0.58815 | 0.011273
2019-03-19 15:45 | 1553024698.7687 | 169 | 0.20791 | 0.001259 | 2.2e-05 | 98.7283 | 98.3463 | 122.3283 | 21.2612 | 0.002651 | 0.002385 | 253.4854 | 0.26772 | 8.4024 | 0.003034 | 0.004861 | 0.004734 | 4.5e-05 | 0.001127 | 0.3079 | 1.3308 | 2.5481 | 101.21 | 20.83 | 17.71 | 85.5981 | 94.8148 | 79.2093 | 0.004311 | 1.0019 | 0.024511 | 24.4635 | 0.034671 | 0.15927 | 0.007955 | 0.58662 | 0.58662 | 0.011295
2019-03-19 15:45 | 1553024699.7688 | 170 | 0.2079 | 0.001249 | 2.2e-05 | 98.7682 | 98.5032 | 122.21 | 21.2612 | 0.002651 | 0.002432 | 253.4357 | 0.2675 | 8.4044 | 0.00308 | 0.005263 | 0.005264 | 4.3e-05 | 0.001127 | 0.31078 | 1.3023 | 2.548 | 101.21 | 20.83 | 17.707 | 85.6005 | 95.9321 | 80.3366 | 0.004311 | 0.99875 | 0.02446 | 24.4908 | 0.024032 | 0.15914 | 0.007996 | 0.58846 | 0.58846 | 0.01125
2019-03-19 15:45 | 1553024700.7679 | 171 | 0.20791 | 0.001251 | 2.2e-05 | 98.8336 | 98.6368 | 122.1208 | 21.2612 | 0.002635 | 0.002419 | 253.3898 | 0.268 | 8.4028 | 0.00309 | 0.004881 | 0.004726 | 4.2e-05 | 0.001121 | 0.29623 | 1.3216 | 2.5481 | 101.21 | 20.83 | 17.706 | 85.211 | 96.1617 | 80.027 | 0.00431 | 0.99963 | 0.024462 | 24.4711 | 0.017834 | 0.15943 | 0.007923 | 0.5831 | 0.5831 | 0.011375
2019-03-19 15:45 | 1553024701.7689 | 172 | 0.20793 | 0.001251 | 2.2e-05 | 98.8907 | 98.5114 | 122.1542 | 21.2612 | 0.00264 | 0.00249 | 253.3473 | 0.26815 | 8.4016 | 0.003108 | 0.004669 | 0.004578 | 4.5e-05 | 0.001124 | 0.26753 | 1.3344 | 2.5481 | 101.21 | 20.84 | 17.708 | 85.3272 | 96.6112 | 81.7239 | 0.004314 | 0.99485 | 0.024407 | 24.5332 | 0.034367 | 0.15952 | 0.007801 | 0.57282 | 0.57282 | 0.011585
2019-03-19 15:45 | 1553024702.7694 | 173 | 0.20794 | 0.001243 | 2.2e-05 | 98.9408 | 98.7686 | 121.8596 | 21.2612 | 0.002641 | 0.002394 | 253.3207 | 0.26693 | 8.4021 | 0.003167 | 0.004568 | 0.004536 | 4.2e-05 | 0.001128 | 0.30486 | 1.3165 | 2.5479 | 101.21 | 20.84 | 17.711 | 85.357 | 98.0234 | 79.4244 | 0.004314 | 1.0013 | 0.024457 | 24.4241 | 0.015576 | 0.1588 | 0.007795 | 0.57358 | 0.57358 | 0.011517
2019-03-19 15:45 | 1553024703.7681 | 174 | 0.20796 | 0.001236 | 2.1e-05 | 98.9663 | 98.806 | 121.5831 | 21.2612 | 0.002669 | 0.002401 | 253.286 | 0.26754 | 8.3999 | 0.003132 | 0.004963 | 0.004797 | 4.4e-05 | 0.001131 | 0.32101 | 1.3163 | 2.548 | 101.21 | 20.84 | 17.714 | 86.0171 | 97.1792 | 79.5935 | 0.004315 | 1.0009 | 0.024423 | 24.4022 | 0.014501 | 0.15916 | 0.007653 | 0.56239 | 0.56239 | 0.011773
2019-03-19 15:45 | 1553024704.7692 | 175 | 0.20798 | 0.001237 | 2e-05 | 99.0632 | 98.7902 | 121.5871 | 21.2612 | 0.002679 | 0.002412 | 253.2317 | 0.26805 | 8.4006 | 0.003147 | 0.004532 | 0.004522 | 4.1e-05 | 0.001121 | 0.32836 | 1.3363 | 2.5482 | 101.21 | 20.84 | 17.724 | 86.2717 | 97.5437 | 79.8432 | 0.004317 | 1.0001 | 0.024415 | 24.4112 | 0.024684 | 0.15947 | 0.007559 | 0.55536 | 0.55536 | 0.011945
2019-03-19 15:45 | 1553024705.7682 | 176 | 0.20799 | 0.001228 | 2e-05 | 99.1412 | 98.8447 | 121.5864 | 21.2612 | 0.00263 | 0.002443 | 253.1977 | 0.2675 | 8.4021 | 0.003161 | 0.004215 | 0.004048 | 4.3e-05 | 0.001121 | 0.32532 | 1.329 | 2.5481 | 101.21 | 20.84 | 17.73 | 85.097 | 97.8855 | 80.5822 | 0.004319 | 0.99806 | 0.024389 | 24.4367 | 0.026797 | 0.15914 | 0.007509 | 0.55105 | 0.55105 | 0.012014
2019-03-19 15:45 | 1553024706.7686 | 177 | 0.20796 | 0.001225 | 2.1e-05 | 99.1752 | 99.055 | 121.6227 | 21.2612 | 0.002691 | 0.002438 | 253.1547 | 0.26763 | 8.4014 | 0.003183 | 0.004409 | 0.004289 | 4e-05 | 0.001121 | 0.31526 | 1.3017 | 2.5481 | 101.21 | 20.84 | 17.736 | 86.5659 | 98.4146 | 80.4613 | 0.00432 | 0.9984 | 0.024397 | 24.4362 | 0.010847 | 0.15921 | 0.007645 | 0.56125 | 0.56125 | 0.011801
2019-03-19 15:45 | 1553024707.7676 | 178 | 0.20799 | 0.00122 | 2.1e-05 | 99.2178 | 99.0258 | 121.7218 | 21.2612 | 0.002645 | 0.002409 | 253.1091 | 0.2677 | 8.4071 | 0.0032 | 0.00524 | 0.005187 | 4.5e-05 | 0.001126 | 0.31365 | 1.2859 | 2.548 | 101.21 | 20.84 | 17.749 | 85.4552 | 98.8349 | 79.7839 | 0.004324 | 1.0003 | 0.02443 | 24.4227 | 0.017319 | 0.15925 | 0.007527 | 0.55328 | 0.55328 | 0.011974
2019-03-19 15:45 | 1553024708.7684 | 179 | 0.208 | 0.001213 | 2.1e-05 | 99.3342 | 99.0752 | 121.6953 | 21.2612 | 0.002708 | 0.002427 | 253.0659 | 0.26731 | 8.4019 | 0.003224 | 0.004874 | 0.004746 | 4.4e-05 | 0.001121 | 0.3282 | 1.3211 | 2.548 | 101.21 | 20.84 | 17.755 | 86.9743 | 99.4121 | 80.2032 | 0.004325 | 0.99913 | 0.024413 | 24.4345 | 0.023356 | 0.15902 | 0.00746 | 0.54796 | 0.54796 | 0.012073
2019-03-19 15:45 | 1553024709.7688 | 180 | 0.20799 | 0.001211 | 2e-05 | 99.3579 | 99.2033 | 121.7359 | 21.2612 | 0.002656 | 0.002414 | 253.049 | 0.26766 | 8.4015 | 0.003283 | 0.004678 | 0.004622 | 4.5e-05 | 0.001119 | 0.323 | 1.3584 | 2.5482 | 101.21 | 20.84 | 17.755 | 85.7222 | 100.8228 | 79.8999 | 0.004325 | 0.99999 | 0.024428 | 24.4281 | 0.013934 | 0.15923 | 0.007501 | 0.55131 | 0.55131 | 0.012015
2019-03-19 15:45 | 1553024710.7684 | 181 | 0.20801 | 0.00121 | 2.1e-05 | 99.4398 | 99.1378 | 121.7932 | 21.2612 | 0.002657 | 0.002404 | 253.0282 | 0.26861 | 8.4046 | 0.003299 | 0.004899 | 0.004731 | 4.3e-05 | 0.00112 | 0.32278 | 1.3665 | 2.5481 | 101.21 | 20.84 | 17.773 | 85.7402 | 101.2106 | 79.6661 | 0.004329 | 1.0007 | 0.024442 | 24.4258 | 0.027206 | 0.1598 | 0.00738 | 0.5427 | 0.5427 | 0.012249
2019-03-19 15:45 | 1553024711.7685 | 182 | 0.20801 | 0.001204 | 2.1e-05 | 99.5146 | 99.3372 | 121.9232 | 21.2612 | 0.002658 | 0.002407 | 253.0083 | 0.26737 | 8.404 | 0.003337 | 0.004268 | 0.00426 | 4.4e-05 | 0.001124 | 0.31509 | 1.3798 | 2.5481 | 101.21 | 20.84 | 17.778 | 85.758 | 102.1458 | 79.7207 | 0.004331 | 1.0005 | 0.024453 | 24.4407 | 0.015959 | 0.15906 | 0.007379 | 0.54286 | 0.54286 | 0.012189
2019-03-19 15:45 | 1553024712.7685 | 183 | 0.208 | 0.001204 | 2.1e-05 | 99.5041 | 99.3139 | 122.2229 | 21.2612 | 0.002663 | 0.002431 | 252.9761 | 0.26808 | 8.4024 | 0.003386 | 0.004532 | 0.004399 | 4.6e-05 | 0.001125 | 0.29839 | 1.3407 | 2.5479 | 101.21 | 20.84 | 17.783 | 85.8815 | 103.3187 | 80.3125 | 0.004332 | 0.99882 | 0.024462 | 24.4912 | 0.017112 | 0.15948 | 0.007426 | 0.54658 | 0.54658 | 0.012138
2019-03-19 15:45 | 1553024713.7689 | 184 | 0.20801 | 0.001198 | 2.1e-05 | 99.6064 | 99.3981 | 122.4844 | 21.2612 | 0.002672 | 0.002442 | 252.9458 | 0.26777 | 8.4059 | 0.003349 | 0.004314 | 0.004333 | 4.5e-05 | 0.001124 | 0.31852 | 1.3654 | 2.548 | 101.21 | 20.84 | 17.79 | 86.1009 | 102.4198 | 80.5651 | 0.004334 | 0.99811 | 0.02448 | 24.5262 | 0.018722 | 0.1593 | 0.007429 | 0.54714 | 0.54714 | 0.012112
2019-03-19 15:45 | 1553024714.7677 | 185 | 0.208 | 0.001194 | 2e-05 | 99.6775 | 99.5567 | 122.656 | 21.2612 | 0.002637 | 0.002424 | 252.9136 | 0.26718 | 8.4035 | 0.003282 | 0.004889 | 0.004813 | 4.6e-05 | 0.001132 | 0.31437 | 1.3148 | 2.5479 | 101.21 | 20.84 | 17.792 | 85.2476 | 100.8046 | 80.1362 | 0.004334 | 0.99932 | 0.024512 | 24.5285 | 0.010848 | 0.15894 | 0.007489 | 0.55233 | 0.55233 | 0.011971
2019-03-19 15:45 | 1553024715.7677 | 186 | 0.20801 | 0.001199 | 2.1e-05 | 99.6514 | 99.6999 | 122.5422 | 21.2612 | 0.002616 | 0.002393 | 252.8856 | 0.26759 | 8.4017 | 0.003329 | 0.004702 | 0.004537 | 4.5e-05 | 0.001136 | 0.33931 | 1.3421 | 2.5482 | 101.21 | 20.84 | 17.794 | 84.7483 | 101.9582 | 79.3986 | 0.004335 | 1.0014 | 0.024526 | 24.4915 | -0.004348 | 0.15919 | 0.007407 | 0.54662 | 0.54662 | 0.012115
2019-03-19 15:45 | 1553024716.7678 | 187 | 0.20802 | 0.00119 | 2e-05 | 99.758 | 99.7047 | 122.4027 | 21.2612 | 0.002634 | 0.002413 | 252.8464 | 0.26726 | 8.4037 | 0.003384 | 0.004749 | 0.0046 | 4.4e-05 | 0.00113 | 0.32836 | 1.3291 | 2.5479 | 101.21 | 20.84 | 17.796 | 85.1908 | 103.2904 | 79.8627 | 0.004335 | 1.0001 | 0.024496 | 24.4936 | 0.00478 | 0.15899 | 0.007331 | 0.54038 | 0.54038 | 0.012239
2019-03-19 15:45 | 1553024717.7692 | 188 | 0.20805 | 0.001188 | 1.9e-05 | 99.7381 | 99.6619 | 122.426 | 21.2612 | 0.002582 | 0.002365 | 252.8167 | 0.26759 | 8.4023 | 0.003279 | 0.004784 | 0.004776 | 5e-05 | 0.001127 | 0.29839 | 1.3277 | 2.548 | 101.21 | 20.84 | 17.795 | 83.9289 | 100.7381 | 78.722 | 0.004335 | 1.0033 | 0.024538 | 24.4564 | 0.006845 | 0.15919 | 0.00721 | 0.5324 | 0.5324 | 0.012438
2019-03-19 15:45 | 1553024718.7679 | 189 | 0.20805 | 0.001183 | 1.9e-05 | 99.9716 | 99.8458 | 122.5243 | 21.2612 | 0.002632 | 0.00238 | 252.7877 | 0.26752 | 8.4058 | 0.003326 | 0.004757 | 0.004644 | 4.8e-05 | 0.001129 | 0.30663 | 1.3888 | 2.5482 | 101.21 | 20.84 | 17.794 | 85.1283 | 101.8718 | 79.0759 | 0.004335 | 1.0023 | 0.024535 | 24.4785 | 0.011268 | 0.15915 | 0.007199 | 0.5315 | 0.5315 | 0.012456
2019-03-19 15:45 | 1553024719.768 | 190 | 0.20805 | 0.00118 | 1.9e-05 | 99.9737 | 99.9958 | 122.7177 | 21.2612 | 0.002592 | 0.002422 | 252.7659 | 0.26664 | 8.4027 | 0.003387 | 0.005065 | 0.004924 | 4.5e-05 | 0.001125 | 0.30127 | 1.3635 | 2.548 | 101.21 | 20.84 | 17.788 | 84.1709 | 103.3587 | 80.091 | 0.004333 | 0.99945 | 0.024519 | 24.5331 | -0.001973 | 0.15862 | 0.007171 | 0.52907 | 0.52907 | 0.012472
2019-03-19 15:45 | 1553024720.769 | 191 | 0.20806 | 0.001172 | 2e-05 | 99.9659 | 99.8328 | 122.6968 | 21.2612 | 0.002605 | 0.002449 | 252.7648 | 0.26787 | 8.4037 | 0.003319 | 0.005566 | 0.00542 | 4.3e-05 | 0.001131 | 0.33699 | 1.3641 | 2.548 | 101.21 | 20.84 | 17.784 | 84.4948 | 101.7098 | 80.743 | 0.004332 | 0.9976 | 0.024495 | 24.5536 | 0.011924 | 0.15936 | 0.007121 | 0.52479 | 0.52479 | 0.012632
2019-03-19 15:45 | 1553024721.7676 | 192 | 0.20805 | 0.001165 | 2e-05 | 100.0551 | 99.9787 | 122.395 | 21.2612 | 0.002591 | 0.002359 | 252.7265 | 0.26755 | 8.4028 | 0.003368 | 0.005156 | 0.005103 | 4.9e-05 | 0.001129 | 0.31205 | 1.3132 | 2.548 | 101.21 | 20.84 | 17.776 | 84.1406 | 102.8881 | 78.5769 | 0.00433 | 1.0037 | 0.02454 | 24.4482 | 0.006833 | 0.15916 | 0.007181 | 0.53018 | 0.53018 | 0.012489
2019-03-19 15:45 | 1553024722.7685 | 193 | 0.20806 | 0.001167 | 2.1e-05 | 100.0737 | 100.0824 | 122.1643 | 21.2612 | 0.002635 | 0.002406 | 252.6929 | 0.26802 | 8.4039 | 0.003265 | 0.005008 | 0.004934 | 4.7e-05 | 0.001134 | 0.32012 | 1.4055 | 2.548 | 101.21 | 20.84 | 17.772 | 85.211 | 100.3958 | 79.6966 | 0.004329 | 1.0006 | 0.024478 | 24.464 | -0.000784 | 0.15944 | 0.007169 | 0.52797 | 0.52797 | 0.012563
2019-03-19 15:45 | 1553024723.7671 | 194 | 0.20804 | 0.001155 | 2.1e-05 | 100.2171 | 100.1663 | 121.919 | 21.2612 | 0.002615 | 0.00236 | 252.6861 | 0.26722 | 8.4013 | 0.003343 | 0.004431 | 0.00431 | 5e-05 | 0.001132 | 0.31741 | 1.3754 | 2.5481 | 101.21 | 20.84 | 17.768 | 84.7389 | 102.2879 | 78.5949 | 0.004328 | 1.0037 | 0.024491 | 24.4013 | 0.004533 | 0.15897 | 0.007287 | 0.53695 | 0.53695 | 0.012316
2019-03-19 15:45 | 1553024724.7682 | 195 | 0.20801 | 0.001144 | 2.1e-05 | 100.1691 | 100.0126 | 121.9751 | 21.2612 | 0.002633 | 0.002373 | 252.6635 | 0.2679 | 8.4016 | 0.003433 | 0.00495 | 0.004974 | 4.7e-05 | 0.001131 | 0.3246 | 1.3582 | 2.548 | 101.21 | 20.84 | 17.761 | 85.1548 | 104.4555 | 78.9185 | 0.004326 | 1.0028 | 0.024486 | 24.4181 | 0.013989 | 0.15937 | 0.007485 | 0.5514 | 0.5514 | 0.012024
2019-03-19 15:45 | 1553024725.7672 | 196 | 0.208 | 0.001149 | 2.2e-05 | 100.2276 | 100.1842 | 122.1924 | 21.2612 | 0.002603 | 0.002402 | 252.6183 | 0.26709 | 8.4012 | 0.003474 | 0.004487 | 0.004366 | 4.9e-05 | 0.001124 | 0.3246 | 1.3497 | 2.5481 | 101.21 | 20.84 | 17.759 | 84.4431 | 105.4519 | 79.5967 | 0.004326 | 1.0008 | 0.024484 | 24.4634 | 0.00388 | 0.15889 | 0.007539 | 0.5553 | 0.5553 | 0.011903
2019-03-19 15:45 | 1553024726.7674 | 197 | 0.20801 | 0.001145 | 2.3e-05 | 100.2357 | 100.018 | 122.2661 | 21.2612 | 0.002586 | 0.002372 | 252.5746 | 0.26794 | 8.4019 | 0.003355 | 0.004125 | 0.003979 | 4.7e-05 | 0.001124 | 0.35817 | 1.3113 | 2.5481 | 101.21 | 20.84 | 17.753 | 84.0412 | 102.5685 | 78.8842 | 0.004325 | 1.0029 | 0.024516 | 24.446 | 0.019445 | 0.1594 | 0.00745 | 0.5494 | 0.5494 | 0.01207
2019-03-19 15:45 | 1553024727.7673 | 198 | 0.208 | 0.00115 | 2.3e-05 | 100.2915 | 100.1345 | 122.3422 | 21.2612 | 0.002654 | 0.002383 | 252.5493 | 0.26796 | 8.403 | 0.003468 | 0.004609 | 0.004509 | 5.2e-05 | 0.001124 | 0.32101 | 1.3007 | 2.5481 | 101.21 | 20.84 | 17.751 | 85.6722 | 105.3246 | 79.16 | 0.004324 | 1.0021 | 0.024514 | 24.4632 | 0.014017 | 0.15941 | 0.007507 | 0.55356 | 0.55356 | 0.01198
2019-03-19 15:45 | 1553024728.7678 | 199 | 0.20804 | 0.001148 | 2.3e-05 | 100.3167 | 100.264 | 122.3666 | 21.2612 | 0.00262 | 0.002408 | 252.5222 | 0.26776 | 8.4014 | 0.003403 | 0.004897 | 0.004805 | 4.9e-05 | 0.001127 | 0.32709 | 1.321 | 2.5481 | 101.21 | 20.84 | 17.749 | 84.8519 | 103.747 | 79.7401 | 0.004324 | 1.0004 | 0.024497 | 24.4858 | 0.004709 | 0.15929 | 0.007302 | 0.53804 | 0.53804 | 0.012316
2019-03-19 15:45 | 1553024729.7685 | 200 | 0.20804 | 0.001152 | 2.4e-05 | 100.2969 | 100.3496 | 122.2265 | 21.2612 | 0.002576 | 0.002362 | 252.4727 | 0.26798 | 8.4028 | 0.00347 | 0.00503 | 0.004897 | 4.7e-05 | 0.001129 | 0.33251 | 1.3274 | 2.5481 | 101.21 | 20.84 | 17.749 | 83.7965 | 105.3709 | 78.647 | 0.004324 | 1.0035 | 0.024521 | 24.4338 | -0.004693 | 0.15942 | 0.007296 | 0.53805 | 0.53805 | 0.012326
2019-03-19 15:45 | 1553024730.7686 | 201 | 0.20805 | 0.001152 | 2.3e-05 | 100.375 | 100.4047 | 122.079 | 21.2612 | 0.002557 | 0.002352 | 252.4287 | 0.26773 | 8.4021 | 0.003503 | 0.005042 | 0.004919 | 5.2e-05 | 0.00113 | 0.31437 | 1.3228 | 2.5481 | 101.21 | 20.84 | 17.75 | 83.3338 | 106.1554 | 78.4077 | 0.004324 | 1.0042 | 0.024514 | 24.4108 | -0.00265 | 0.15927 | 0.007229 | 0.53299 | 0.53299 | 0.012431
2019-03-19 15:46 | 1553024731.7686 | 202 | 0.20805 | 0.001151 | 2.3e-05 | 100.4173 | 100.3664 | 122.3586 | 21.2612 | 0.002538 | 0.002355 | 252.4056 | 0.26684 | 8.4048 | 0.003544 | 0.005072 | 0.005002 | 5.2e-05 | 0.001126 | 0.32156 | 1.3182 | 2.5481 | 101.21 | 20.84 | 17.752 | 82.865 | 107.1575 | 78.4872 | 0.004324 | 1.004 | 0.024539 | 24.4415 | 0.004537 | 0.15875 | 0.007234 | 0.53394 | 0.53394 | 0.012368
2019-03-19 15:46 | 1553024732.768 | 203 | 0.20802 | 0.001146 | 2.3e-05 | 100.4191 | 100.5047 | 122.4161 | 21.2612 | 0.002594 | 0.00239 | 252.3623 | 0.26653 | 8.4029 | 0.003508 | 0.004972 | 0.004826 | 5.2e-05 | 0.001118 | 0.33035 | 1.3358 | 2.5481 | 101.21 | 20.84 | 17.766 | 84.223 | 106.2845 | 79.3261 | 0.004328 | 1.0016 | 0.024516 | 24.4764 | -0.007625 | 0.15856 | 0.00742 | 0.54717 | 0.54717 | 0.012055
2019-03-19 15:46 | 1553024733.7686 | 204 | 0.20804 | 0.001156 | 2.2e-05 | 100.4998 | 100.5383 | 122.864 | 21.2612 | 0.002586 | 0.00238 | 252.3201 | 0.26646 | 8.4036 | 0.003424 | 0.005037 | 0.004968 | 4.9e-05 | 0.001117 | 0.3282 | 1.4103 | 2.5482 | 101.21 | 20.84 | 17.774 | 84.0297 | 104.2369 | 79.0884 | 0.00433 | 1.0023 | 0.024569 | 24.5128 | -0.003424 | 0.15852 | 0.007293 | 0.539 | 0.539 | 0.012234
2019-03-19 15:46 | 1553024734.7676 | 205 | 0.20804 | 0.001151 | 2.3e-05 | 100.4971 | 100.4869 | 122.81 | 21.2612 | 0.002572 | 0.002335 | 252.2822 | 0.26699 | 8.4008 | 0.003368 | 0.005188 | 0.00501 | 5e-05 | 0.001114 | 0.30558 | 1.3499 | 2.548 | 101.21 | 20.84 | 17.792 | 83.6818 | 102.8835 | 78.0016 | 0.004334 | 1.0054 | 0.024602 | 24.4696 | 0.000908 | 0.15883 | 0.007284 | 0.539 | 0.539 | 0.012259
2019-03-19 15:46 | 1553024735.7687 | 206 | 0.20805 | 0.001156 | 2.3e-05 | 100.6415 | 100.6958 | 122.9771 | 21.2612 | 0.00263 | 0.002362 | 252.2401 | 0.26703 | 8.4012 | 0.003388 | 0.004961 | 0.004857 | 5.1e-05 | 0.001113 | 0.3095 | 1.3608 | 2.5482 | 101.21 | 20.84 | 17.802 | 85.0908 | 103.3706 | 78.647 | 0.004336 | 1.0035 | 0.024596 | 24.5087 | -0.004831 | 0.15886 | 0.007227 | 0.53465 | 0.53465 | 0.01236
2019-03-19 15:46 | 1553024736.7682 | 207 | 0.20803 | 0.001156 | 2.3e-05 | 100.6189 | 100.6109 | 122.9323 | 21.2612 | 0.002541 | 0.002422 | 252.2325 | 0.26862 | 8.4034 | 0.003342 | 0.005593 | 0.005456 | 5.2e-05 | 0.001114 | 0.32892 | 1.3088 | 2.5479 | 101.21 | 20.84 | 17.811 | 82.9484 | 102.2588 | 80.0989 | 0.004339 | 0.99942 | 0.024541 | 24.5548 | 0.000716 | 0.1598 | 0.007316 | 0.54 | 0.54 | 0.012311
2019-03-19 15:46 | 1553024737.7679 | 208 | 0.20804 | 0.001157 | 2.3e-05 | 100.6865 | 100.5999 | 122.6757 | 21.2612 | 0.002575 | 0.002394 | 252.1911 | 0.26684 | 8.3998 | 0.003331 | 0.00516 | 0.005201 | 4.9e-05 | 0.001111 | 0.3178 | 1.3059 | 2.5481 | 101.21 | 20.84 | 17.829 | 83.7666 | 101.9925 | 79.4056 | 0.004343 | 1.0014 | 0.024539 | 24.5051 | 0.007691 | 0.15875 | 0.007291 | 0.53811 | 0.53811 | 0.012272
2019-03-19 15:46 | 1553024738.7683 | 209 | 0.20803 | 0.001155 | 2.4e-05 | 100.6457 | 100.814 | 122.6851 | 21.2612 | 0.002561 | 0.002345 | 252.1601 | 0.26731 | 8.4005 | 0.003347 | 0.004667 | 0.004418 | 5.4e-05 | 0.001112 | 0.32853 | 1.38 | 2.5481 | 101.21 | 20.84 | 17.838 | 83.4211 | 102.3939 | 78.234 | 0.004345 | 1.0047 | 0.024581 | 24.4653 | -0.014947 | 0.15902 | 0.00732 | 0.54119 | 0.54119 | 0.012224
2019-03-19 15:46 | 1553024739.768 | 210 | 0.20803 | 0.001152 | 2.4e-05 | 100.7267 | 100.7047 | 122.9048 | 21.2612 | 0.002579 | 0.002361 | 252.1378 | 0.26746 | 8.4002 | 0.003314 | 0.004432 | 0.004317 | 5.2e-05 | 0.001107 | 0.33124 | 1.3763 | 2.548 | 101.21 | 20.84 | 17.852 | 83.8604 | 101.5905 | 78.6112 | 0.004349 | 1.0037 | 0.02459 | 24.5003 | 0.00195 | 0.15911 | 0.00735 | 0.54357 | 0.54357 | 0.012177
2019-03-19 15:46 | 1553024740.7671 | 211 | 0.20805 | 0.001153 | 2.4e-05 | 100.7834 | 100.7421 | 123.2118 | 21.2612 | 0.002574 | 0.002386 | 252.0987 | 0.26754 | 8.4006 | 0.003286 | 0.005074 | 0.004952 | 5.5e-05 | 0.0011 | 0.33301 | 1.3752 | 2.548 | 101.21 | 20.84 | 17.859 | 83.7457 | 100.899 | 79.2295 | 0.00435 | 1.0019 | 0.024599 | 24.5524 | 0.00367 | 0.15916 | 0.007244 | 0.53593 | 0.53593 | 0.012354
2019-03-19 15:46 | 1553024741.7681 | 212 | 0.20805 | 0.001149 | 2.5e-05 | 100.7996 | 100.8873 | 123.0257 | 21.2612 | 0.002566 | 0.002399 | 252.07 | 0.26688 | 8.4019 | 0.003217 | 0.004885 | 0.004773 | 5.2e-05 | 0.001095 | 0.31974 | 1.3579 | 2.5483 | 101.21 | 20.84 | 17.866 | 83.5396 | 99.2357 | 79.5289 | 0.004352 | 1.001 | 0.02457 | 24.5443 | -0.007776 | 0.15877 | 0.007253 | 0.53594 | 0.53594 | 0.012323
2019-03-19 15:46 | 1553024742.7676 | 213 | 0.20806 | 0.001151 | 2.5e-05 | 100.8776 | 100.9688 | 122.9557 | 21.2612 | 0.002601 | 0.002383 | 252.0503 | 0.26786 | 8.4028 | 0.003165 | 0.004542 | 0.004404 | 5.2e-05 | 0.001095 | 0.31045 | 1.3773 | 2.5481 | 101.21 | 20.84 | 17.875 | 84.3854 | 97.9725 | 79.1593 | 0.004354 | 1.0021 | 0.024576 | 24.5244 | -0.008079 | 0.15935 | 0.00714 | 0.52763 | 0.52763 | 0.012564
2019-03-19 15:46 | 1553024743.7676 | 214 | 0.20809 | 0.001151 | 2.5e-05 | 100.8923 | 100.9715 | 122.7457 | 21.2612 | 0.002528 | 0.002391 | 252.0132 | 0.26705 | 8.4024 | 0.003175 | 0.004579 | 0.004413 | 5.2e-05 | 0.001097 | 0.33843 | 1.3758 | 2.5481 | 101.21 | 20.84 | 17.875 | 82.6381 | 98.2257 | 79.3496 | 0.004354 | 1.0015 | 0.024548 | 24.5101 | -0.007018 | 0.15887 | 0.00698 | 0.51519 | 0.51519 | 0.012828
2019-03-19 15:46 | 1553024744.7672 | 215 | 0.2081 | 0.001149 | 2.6e-05 | 101.0148 | 100.9476 | 122.7407 | 21.2612 | 0.002577 | 0.002369 | 252.0089 | 0.26811 | 8.4026 | 0.003295 | 0.005197 | 0.005126 | 5.4e-05 | 0.001093 | 0.32421 | 1.4036 | 2.5481 | 101.21 | 20.84 | 17.87 | 83.8111 | 101.1148 | 78.821 | 0.004353 | 1.0031 | 0.024566 | 24.4912 | 0.005955 | 0.1595 | 0.006945 | 0.51292 | 0.51292 | 0.012936
2019-03-19 15:46 | 1553024745.7673 | 216 | 0.2081 | 0.001148 | 2.6e-05 | 101.0274 | 101.1065 | 122.736 | 21.2612 | 0.002593 | 0.002389 | 251.9867 | 0.26816 | 8.4019 | 0.003242 | 0.005223 | 0.005053 | 5.2e-05 | 0.001086 | 0.3027 | 1.3785 | 2.548 | 101.21 | 20.84 | 17.866 | 84.2098 | 99.8348 | 79.2965 | 0.004352 | 1.0017 | 0.024549 | 24.5073 | -0.006999 | 0.15953 | 0.006952 | 0.51313 | 0.51313 | 0.012933
2019-03-19 15:46 | 1553024746.7684 | 217 | 0.20811 | 0.00114 | 2.5e-05 | 101.0013 | 100.9626 | 122.8729 | 21.2612 | 0.002584 | 0.002389 | 251.9444 | 0.26779 | 8.403 | 0.0032 | 0.004323 | 0.004234 | 5.1e-05 | 0.001084 | 0.32477 | 1.3755 | 2.5481 | 101.21 | 20.84 | 17.859 | 83.9797 | 98.8325 | 79.284 | 0.00435 | 1.0017 | 0.024563 | 24.5205 | 0.003424 | 0.15931 | 0.006903 | 0.5098 | 0.5098 | 0.013
2019-03-19 15:46 | 1553024747.7684 | 218 | 0.20811 | 0.001139 | 2.5e-05 | 101.0709 | 101.0503 | 123.0043 | 21.2612 | 0.002529 | 0.002376 | 251.8993 | 0.26787 | 8.4034 | 0.003225 | 0.005224 | 0.005097 | 5.2e-05 | 0.00108 | 0.31996 | 1.3653 | 2.5481 | 101.21 | 20.84 | 17.843 | 82.6652 | 99.4293 | 78.9918 | 0.004346 | 1.0026 | 0.024586 | 24.5235 | 0.001823 | 0.15935 | 0.006897 | 0.50988 | 0.50988 | 0.013001
2019-03-19 15:46 | 1553024748.7685 | 219 | 0.20812 | 0.001132 | 2.6e-05 | 101.1378 | 100.9517 | 122.9766 | 21.2612 | 0.002568 | 0.002386 | 251.8717 | 0.26681 | 8.4024 | 0.003187 | 0.005451 | 0.005274 | 5.3e-05 | 0.001082 | 0.32565 | 1.3014 | 2.5479 | 101.21 | 20.84 | 17.834 | 83.5959 | 98.5125 | 79.2256 | 0.004344 | 1.0019 | 0.024575 | 24.5288 | 0.016479 | 0.15873 | 0.00685 | 0.50616 | 0.50616 | 0.013045
2019-03-19 15:46 | 1553024749.7679 | 220 | 0.20812 | 0.001128 | 2.5e-05 | 101.0889 | 101.1575 | 122.7593 | 21.2612 | 0.002582 | 0.002365 | 251.8468 | 0.26732 | 8.4039 | 0.003136 | 0.004405 | 0.004345 | 5.3e-05 | 0.001082 | 0.3011 | 1.3309 | 2.5481 | 101.21 | 20.84 | 17.824 | 83.9431 | 97.2717 | 78.7064 | 0.004342 | 1.0034 | 0.024572 | 24.4891 | -0.006068 | 0.15903 | 0.006841 | 0.50545 | 0.50545 | 0.013089
2019-03-19 15:46 | 1553024750.7685 | 221 | 0.20811 | 0.001127 | 2.5e-05 | 101.1144 | 101.2561 | 122.4696 | 21.2612 | 0.00257 | 0.00239 | 251.8136 | 0.26792 | 8.4053 | 0.003148 | 0.004909 | 0.004772 | 5.3e-05 | 0.001079 | 0.34578 | 1.39 | 2.548 | 101.21 | 20.84 | 17.803 | 83.6529 | 97.5618 | 79.3066 | 0.004337 | 1.0017 | 0.024522 | 24.481 | -0.012528 | 0.15938 | 0.006886 | 0.5077 | 0.5077 | 0.01306
2019-03-19 15:46 | 1553024751.7677 | 222 | 0.20813 | 0.001119 | 2.5e-05 | 101.1951 | 101.302 | 122.3628 | 21.2612 | 0.002548 | 0.002388 | 251.7854 | 0.26713 | 8.4032 | 0.00317 | 0.00446 | 0.004367 | 5.1e-05 | 0.00108 | 0.30486 | 1.3346 | 2.548 | 101.21 | 20.84 | 17.791 | 83.1122 | 98.1095 | 79.2771 | 0.004334 | 1.0018 | 0.024512 | 24.4693 | -0.009443 | 0.15892 | 0.006805 | 0.50152 | 0.50152 | 0.013182
2019-03-19 15:46 | 1553024752.7678 | 223 | 0.20814 | 0.001116 | 2.5e-05 | 101.2096 | 101.2599 | 122.6283 | 21.2612 | 0.002539 | 0.00237 | 251.7521 | 0.26709 | 8.4031 | 0.003142 | 0.004791 | 0.004685 | 5.1e-05 | 0.001081 | 0.32516 | 1.3303 | 2.5481 | 101.21 | 20.84 | 17.766 | 82.9001 | 97.4355 | 78.8318 | 0.004328 | 1.003 | 0.024554 | 24.4804 | -0.00445 | 0.15889 | 0.006731 | 0.49691 | 0.49691 | 0.013302
2019-03-19 15:46 | 1553024753.7679 | 224 | 0.20815 | 0.001111 | 2.6e-05 | 101.2006 | 101.1661 | 122.7499 | 21.2612 | 0.002574 | 0.002346 | 251.7393 | 0.26711 | 8.4022 | 0.003059 | 0.004886 | 0.004791 | 5.1e-05 | 0.001076 | 0.31598 | 1.3716 | 2.548 | 101.21 | 20.84 | 17.753 | 83.7512 | 95.4105 | 78.2652 | 0.004325 | 1.0046 | 0.024586 | 24.4728 | 0.003048 | 0.1589 | 0.006691 | 0.49451 | 0.49451 | 0.013368
2019-03-19 15:46 | 1553024754.7673 | 225 | 0.20814 | 0.001104 | 2.6e-05 | 101.284 | 101.4616 | 122.8548 | 21.2612 | 0.002507 | 0.002352 | 251.6971 | 0.26788 | 8.4037 | 0.003113 | 0.004093 | 0.004025 | 5.2e-05 | 0.001075 | 0.33428 | 1.3772 | 2.548 | 101.21 | 20.83 | 17.741 | 82.12 | 96.7279 | 78.4063 | 0.004319 | 1.0042 | 0.024592 | 24.4882 | -0.015676 | 0.15936 | 0.006745 | 0.49868 | 0.49868 | 0.013294
2019-03-19 15:46 | 1553024755.7682 | 226 | 0.20813 | 0.001108 | 2.7e-05 | 101.3083 | 101.3175 | 122.8062 | 21.2612 | 0.002533 | 0.002353 | 251.6909 | 0.26759 | 8.4041 | 0.003124 | 0.005097 | 0.005008 | 5.5e-05 | 0.00108 | 0.31653 | 1.369 | 2.5481 | 101.21 | 20.83 | 17.72 | 82.7496 | 97.0007 | 78.4274 | 0.004314 | 1.0042 | 0.024586 | 24.4841 | -0.000809 | 0.15919 | 0.006777 | 0.50086 | 0.50086 | 0.013222
2019-03-19 15:46 | 1553024756.7669 | 227 | 0.20814 | 0.001095 | 2.7e-05 | 101.2819 | 101.2859 | 122.549 | 21.2612 | 0.002552 | 0.002323 | 251.6591 | 0.26742 | 8.4027 | 0.003055 | 0.004721 | 0.004627 | 5.4e-05 | 0.001083 | 0.29645 | 1.3824 | 2.548 | 101.21 | 20.83 | 17.71 | 83.2081 | 95.3307 | 77.7133 | 0.004311 | 1.0062 | 0.024586 | 24.4336 | -0.000359 | 0.15909 | 0.006766 | 0.50007 | 0.50007 | 0.013234
2019-03-19 15:46 | 1553024757.7682 | 228 | 0.20814 | 0.001097 | 2.7e-05 | 101.3128 | 101.438 | 122.5945 | 21.2612 | 0.002526 | 0.00233 | 251.6361 | 0.26757 | 8.4024 | 0.003063 | 0.004839 | 0.004769 | 5.3e-05 | 0.001077 | 0.28655 | 1.3124 | 2.548 | 101.21 | 20.83 | 17.694 | 82.5967 | 95.5264 | 77.8839 | 0.004308 | 1.0057 | 0.024584 | 24.444 | -0.011048 | 0.15918 | 0.006737 | 0.49782 | 0.49782 | 0.013301
2019-03-19 15:46 | 1553024758.767 | 229 | 0.20813 | 0.00109 | 2.7e-05 | 101.3302 | 101.4339 | 122.76 | 21.2612 | 0.00254 | 0.002368 | 251.6039 | 0.26771 | 8.4023 | 0.003119 | 0.005 | 0.004959 | 5.2e-05 | 0.001079 | 0.32803 | 1.3448 | 2.5481 | 101.21 | 20.83 | 17.688 | 82.9335 | 96.8597 | 78.7936 | 0.004306 | 1.0031 | 0.024569 | 24.4922 | -0.00915 | 0.15926 | 0.006789 | 0.50139 | 0.50139 | 0.013214
2019-03-19 15:46 | 1553024759.7671 | 230 | 0.20817 | 0.001092 | 2.8e-05 | 101.3671 | 101.4452 | 123.2327 | 21.2612 | 0.002491 | 0.002395 | 251.5649 | 0.26795 | 8.4029 | 0.003153 | 0.00473 | 0.004671 | 5.7e-05 | 0.001079 | 0.32621 | 1.3303 | 2.5481 | 101.21 | 20.83 | 17.682 | 81.7565 | 97.6879 | 79.4361 | 0.004305 | 1.0013 | 0.024594 | 24.5617 | -0.006888 | 0.1594 | 0.006547 | 0.48391 | 0.48391 | 0.013704
2019-03-19 15:46 | 1553024760.7681 | 231 | 0.20818 | 0.001091 | 2.6e-05 | 101.3278 | 101.5753 | 123.5318 | 21.2612 | 0.002513 | 0.002368 | 251.5532 | 0.26694 | 8.4049 | 0.003152 | 0.004102 | 0.003952 | 5.3e-05 | 0.001082 | 0.32637 | 1.3258 | 2.548 | 101.21 | 20.83 | 17.674 | 82.2729 | 97.6739 | 78.7866 | 0.004303 | 1.0031 | 0.024646 | 24.5688 | -0.021826 | 0.1588 | 0.006533 | 0.48401 | 0.48401 | 0.013649
2019-03-19 15:46 | 1553024761.7673 | 232 | 0.20819 | 0.001093 | 2.6e-05 | 101.4397 | 101.6171 | 123.9212 | 21.2612 | 0.002504 | 0.002383 | 251.5105 | 0.2675 | 8.4012 | 0.003124 | 0.004808 | 0.004689 | 5.2e-05 | 0.001082 | 0.33301 | 1.3247 | 2.5481 | 101.21 | 20.82 | 17.671 | 82.0498 | 96.9928 | 79.1547 | 0.004299 | 1.0021 | 0.024672 | 24.6204 | -0.015629 | 0.15913 | 0.00648 | 0.48054 | 0.48054 | 0.013776
2019-03-19 15:46 | 1553024762.7683 | 233 | 0.20819 | 0.001083 | 2.6e-05 | 101.4469 | 101.4962 | 123.7573 | 21.2612 | 0.002543 | 0.002329 | 251.4686 | 0.26718 | 8.4023 | 0.003161 | 0.004719 | 0.004742 | 5.5e-05 | 0.001073 | 0.3027 | 1.3198 | 2.5481 | 101.21 | 20.82 | 17.67 | 82.9929 | 97.8918 | 77.8481 | 0.004299 | 1.0058 | 0.024702 | 24.5584 | -0.004346 | 0.15894 | 0.006448 | 0.47879 | 0.47879 | 0.01381
2019-03-19 15:46 | 1553024763.7687 | 234 | 0.20819 | 0.001081 | 2.6e-05 | 101.4247 | 101.476 | 123.8118 | 21.2612 | 0.002597 | 0.002331 | 251.4342 | 0.26774 | 8.4023 | 0.003111 | 0.004736 | 0.004533 | 5.3e-05 | 0.00107 | 0.33301 | 1.3234 | 2.548 | 101.21 | 20.82 | 17.673 | 84.3005 | 96.674 | 77.9072 | 0.0043 | 1.0057 | 0.024705 | 24.5659 | -0.004523 | 0.15928 | 0.006442 | 0.47844 | 0.47844 | 0.013849
2019-03-19 15:46 | 1553024764.7674 | 235 | 0.20819 | 0.001074 | 2.7e-05 | 101.4115 | 101.5654 | 123.8357 | 21.2612 | 0.002556 | 0.002298 | 251.3879 | 0.26712 | 8.4026 | 0.003089 | 0.004372 | 0.004409 | 5.6e-05 | 0.001069 | 0.31382 | 1.3456 | 2.5481 | 101.21 | 20.82 | 17.675 | 83.3127 | 96.1554 | 77.0997 | 0.0043 | 1.008 | 0.024736 | 24.54 | -0.013563 | 0.15891 | 0.006467 | 0.48086 | 0.48086 | 0.013748
2019-03-19 15:46 | 1553024765.7684 | 236 | 0.20819 | 0.00107 | 2.6e-05 | 101.5016 | 101.7093 | 124.0268 | 21.2612 | 0.0026 | 0.002389 | 251.3413 | 0.26679 | 8.4025 | 0.00301 | 0.004921 | 0.0048 | 5.9e-05 | 0.001072 | 0.34003 | 1.342 | 2.5481 | 101.21 | 20.82 | 17.683 | 84.3565 | 94.2466 | 79.3004 | 0.004302 | 1.0017 | 0.024678 | 24.636 | -0.018285 | 0.15871 | 0.006468 | 0.47976 | 0.47976 | 0.013762
2019-03-19 15:46 | 1553024766.7684 | 237 | 0.2082 | 0.001062 | 2.6e-05 | 101.4938 | 101.6058 | 123.9503 | 21.2612 | 0.002527 | 0.002366 | 251.3181 | 0.26777 | 8.4013 | 0.00312 | 0.004248 | 0.004103 | 5.5e-05 | 0.00107 | 0.29336 | 1.3358 | 2.5482 | 101.21 | 20.82 | 17.687 | 82.6176 | 96.8918 | 78.7508 | 0.004303 | 1.0033 | 0.024689 | 24.6092 | -0.009872 | 0.1593 | 0.006418 | 0.47632 | 0.47632 | 0.013912
2019-03-19 15:46 | 1553024767.7676 | 238 | 0.20819 | 0.00106 | 2.7e-05 | 101.5241 | 101.6438 | 123.7282 | 21.2612 | 0.002558 | 0.002338 | 251.314 | 0.26703 | 8.4019 | 0.003054 | 0.004307 | 0.004229 | 5.5e-05 | 0.001074 | 0.31614 | 1.3482 | 2.5482 | 101.21 | 20.82 | 17.692 | 83.3603 | 95.2971 | 78.0756 | 0.004304 | 1.0052 | 0.024691 | 24.5635 | -0.010547 | 0.15886 | 0.006477 | 0.48073 | 0.48073 | 0.013747
2019-03-19 15:46 | 1553024768.7676 | 239 | 0.20819 | 0.001053 | 2.6e-05 | 101.534 | 101.4634 | 123.4865 | 21.2612 | 0.002591 | 0.002344 | 251.2768 | 0.2675 | 8.4047 | 0.003173 | 0.004456 | 0.004327 | 5.4e-05 | 0.001077 | 0.32604 | 1.3274 | 2.548 | 101.21 | 20.82 | 17.704 | 84.1396 | 98.174 | 78.2215 | 0.004307 | 1.0048 | 0.024662 | 24.5446 | 0.006223 | 0.15913 | 0.006493 | 0.48134 | 0.48134 | 0.013753
2019-03-19 15:46 | 1553024769.7666 | 240 | 0.20819 | 0.001052 | 2.7e-05 | 101.5109 | 101.6045 | 123.1686 | 21.2612 | 0.002499 | 0.002383 | 251.2462 | 0.26766 | 8.403 | 0.003138 | 0.004919 | 0.004834 | 5.8e-05 | 0.00108 | 0.31437 | 1.3469 | 2.5481 | 101.21 | 20.82 | 17.71 | 81.9361 | 97.3352 | 79.1562 | 0.004309 | 1.0021 | 0.024597 | 24.5456 | -0.008244 | 0.15923 | 0.006536 | 0.48323 | 0.48323 | 0.013708
2019-03-19 15:46 | 1553024770.768 | 241 | 0.20817 | 0.001049 | 2.7e-05 | 101.57 | 101.8041 | 123.4071 | 21.2612 | 0.002487 | 0.002331 | 251.2187 | 0.26796 | 8.4027 | 0.0031 | 0.004146 | 0.004033 | 5.7e-05 | 0.001074 | 0.31509 | 1.3242 | 2.548 | 101.21 | 20.82 | 17.723 | 81.638 | 96.4178 | 77.9113 | 0.004312 | 1.0057 | 0.024664 | 24.5259 | -0.020597 | 0.15941 | 0.006601 | 0.48941 | 0.48941 | 0.01355
2019-03-19 15:46 | 1553024771.7677 | 242 | 0.20817 | 0.001053 | 2.7e-05 | 101.5811 | 101.8397 | 123.1022 | 21.2612 | 0.002547 | 0.002367 | 251.1795 | 0.26692 | 8.4025 | 0.003076 | 0.004927 | 0.004806 | 5.5e-05 | 0.001069 | 0.31116 | 1.3363 | 2.548 | 101.21 | 20.82 | 17.729 | 83.0833 | 95.8257 | 78.7602 | 0.004313 | 1.0032 | 0.024604 | 24.5251 | -0.022748 | 0.15879 | 0.006608 | 0.48869 | 0.48869 | 0.013517
2019-03-19 15:46 | 1553024772.7668 | 243 | 0.20816 | 0.001049 | 2.7e-05 | 101.5511 | 101.7072 | 123.0386 | 21.2612 | 0.002567 | 0.002373 | 251.1657 | 0.26651 | 8.4037 | 0.003117 | 0.004067 | 0.004034 | 5.5e-05 | 0.001069 | 0.35258 | 1.3042 | 2.548 | 101.21 | 20.82 | 17.735 | 83.5639 | 96.8126 | 78.9169 | 0.004315 | 1.0028 | 0.024592 | 24.5243 | -0.013742 | 0.15854 | 0.006669 | 0.49294 | 0.49294 | 0.01338
2019-03-19 15:46 | 1553024773.7678 | 244 | 0.20817 | 0.001051 | 2.7e-05 | 101.6117 | 101.6661 | 122.8844 | 21.2612 | 0.00254 | 0.002304 | 251.1245 | 0.26683 | 8.4023 | 0.003101 | 0.005005 | 0.004865 | 5.5e-05 | 0.001073 | 0.33268 | 1.401 | 2.5481 | 101.21 | 20.82 | 17.747 | 82.921 | 96.4335 | 77.2604 | 0.004318 | 1.0075 | 0.024635 | 24.4512 | -0.004789 | 0.15874 | 0.006655 | 0.4928 | 0.4928 | 0.0134
2019-03-19 15:46 | 1553024774.7668 | 245 | 0.20817 | 0.001048 | 2.8e-05 | 101.6438 | 101.6706 | 122.9689 | 21.2612 | 0.002524 | 0.002336 | 251.1074 | 0.2675 | 8.4047 | 0.003196 | 0.004707 | 0.004649 | 5.6e-05 | 0.001065 | 0.30237 | 1.3473 | 2.5481 | 101.21 | 20.82 | 17.753 | 82.5429 | 98.7361 | 78.0173 | 0.004319 | 1.0053 | 0.024617 | 24.486 | -0.002354 | 0.15914 | 0.00661 | 0.48902 | 0.48902 | 0.013537
2019-03-19 15:46 | 1553024775.767 | 246 | 0.20818 | 0.001048 | 2.7e-05 | 101.6036 | 101.7528 | 122.7407 | 21.2612 | 0.002547 | 0.002366 | 251.0866 | 0.26754 | 8.4035 | 0.003196 | 0.004635 | 0.004504 | 5.4e-05 | 0.001061 | 0.31454 | 1.3488 | 2.5481 | 101.21 | 20.82 | 17.762 | 83.0833 | 98.7298 | 78.7484 | 0.004321 | 1.0033 | 0.024569 | 24.4887 | -0.013122 | 0.15916 | 0.00658 | 0.48591 | 0.48591 | 0.013626
2019-03-19 15:46 | 1553024776.767 | 247 | 0.20819 | 0.001045 | 2.8e-05 | 101.5712 | 101.7661 | 122.6222 | 21.2612 | 0.002467 | 0.00229 | 251.0447 | 0.26787 | 8.4018 | 0.003082 | 0.004333 | 0.004132 | 5.6e-05 | 0.001061 | 0.34772 | 1.3425 | 2.5481 | 101.21 | 20.82 | 17.767 | 81.178 | 95.9886 | 76.9205 | 0.004323 | 1.0085 | 0.024621 | 24.4132 | -0.01715 | 0.15936 | 0.006544 | 0.48423 | 0.48423 | 0.01369
2019-03-19 15:46 | 1553024777.7681 | 248 | 0.20817 | 0.00105 | 2.8e-05 | 101.5793 | 101.7134 | 122.7011 | 21.2612 | 0.002539 | 0.002344 | 250.9861 | 0.26725 | 8.4004 | 0.003165 | 0.003903 | 0.00382 | 5.7e-05 | 0.001062 | 0.31166 | 1.3431 | 2.5481 | 101.21 | 20.82 | 17.772 | 82.8907 | 97.9739 | 78.2027 | 0.004324 | 1.0048 | 0.024584 | 24.4658 | -0.0118 | 0.15899 | 0.006599 | 0.48762 | 0.48762 | 0.013563
2019-03-19 15:46 | 1553024778.7672 | 249 | 0.20817 | 0.001044 | 2.7e-05 | 101.6021 | 101.7391 | 122.667 | 21.2612 | 0.002484 | 0.002356 | 250.9336 | 0.26731 | 8.4032 | 0.003188 | 0.004779 | 0.004712 | 5.7e-05 | 0.001063 | 0.33157 | 1.3441 | 2.548 | 101.21 | 20.82 | 17.783 | 81.5695 | 98.5471 | 78.5122 | 0.004326 | 1.0039 | 0.024569 | 24.4731 | -0.01205 | 0.15902 | 0.006638 | 0.49018 | 0.49018 | 0.013496
2019-03-19 15:46 | 1553024779.7681 | 250 | 0.20818 | 0.001046 | 2.7e-05 | 101.6174 | 101.7774 | 122.8137 | 21.2612 | 0.002537 | 0.002297 | 250.9083 | 0.26717 | 8.4032 | 0.00309 | 0.00413 | 0.004039 | 5.8e-05 | 0.001061 | 0.29103 | 1.3377 | 2.548 | 101.21 | 20.82 | 17.788 | 82.84 | 96.1687 | 77.0882 | 0.004328 | 1.008 | 0.024634 | 24.4381 | -0.014074 | 0.15894 | 0.006561 | 0.48583 | 0.48583 | 0.01361
2019-03-19 15:46 | 1553024780.7682 | 251 | 0.20818 | 0.001044 | 2.7e-05 | 101.6579 | 101.7969 | 122.4257 | 21.2612 | 0.002529 | 0.002302 | 250.8534 | 0.26752 | 8.4043 | 0.003125 | 0.004556 | 0.004397 | 5.7e-05 | 0.00107 | 0.33107 | 1.3535 | 2.5479 | 101.21 | 20.82 | 17.792 | 82.652 | 97.0203 | 77.1965 | 0.004329 | 1.0077 | 0.024591 | 24.4033 | -0.012225 | 0.15915 | 0.006597 | 0.48762 | 0.48762 | 0.013577
2019-03-19 15:46 | 1553024781.7686 | 252 | 0.20816 | 0.00104 | 2.7e-05 | 101.6546 | 101.7322 | 122.4546 | 21.2612 | 0.002514 | 0.002385 | 250.8347 | 0.26673 | 8.4046 | 0.003163 | 0.005076 | 0.005033 | 5.6e-05 | 0.001067 | 0.35081 | 1.3377 | 2.5479 | 101.21 | 20.82 | 17.8 | 82.2948 | 97.9292 | 79.1867 | 0.004331 | 1.002 | 0.024525 | 24.4754 | -0.006826 | 0.15868 | 0.006704 | 0.49423 | 0.49423 | 0.013356
2019-03-19 15:46 | 1553024782.7673 | 253 | 0.20817 | 0.001037 | 2.7e-05 | 101.663 | 101.9202 | 122.5983 | 21.2612 | 0.002527 | 0.002336 | 250.8009 | 0.267 | 8.4009 | 0.003111 | 0.004672 | 0.004496 | 6e-05 | 0.001066 | 0.34418 | 1.3201 | 2.5481 | 101.21 | 20.82 | 17.803 | 82.6217 | 96.6692 | 78.0211 | 0.004331 | 1.0053 | 0.02458 | 24.4492 | -0.022603 | 0.15884 | 0.006626 | 0.48953 | 0.48953 | 0.013498
2019-03-19 15:46 | 1553024783.7674 | 254 | 0.20815 | 0.001039 | 2.8e-05 | 101.7059 | 101.7904 | 122.9111 | 21.2612 | 0.002506 | 0.002341 | 250.7801 | 0.26713 | 8.4025 | 0.003195 | 0.004838 | 0.004774 | 5.7e-05 | 0.001063 | 0.32676 | 1.3657 | 2.5481 | 101.21 | 20.82 | 17.809 | 82.0984 | 98.7095 | 78.1436 | 0.004333 | 1.005 | 0.024607 | 24.4846 | -0.007429 | 0.15892 | 0.006781 | 0.50152 | 0.50152 | 0.013182
2019-03-19 15:46 | 1553024784.7685 | 255 | 0.20815 | 0.001033 | 2.8e-05 | 101.6405 | 101.8504 | 122.6764 | 21.2612 | 0.002519 | 0.002319 | 250.7374 | 0.26752 | 8.4027 | 0.00317 | 0.004654 | 0.004482 | 5.5e-05 | 0.001063 | 0.32676 | 1.3402 | 2.5482 | 101.21 | 20.82 | 17.811 | 82.4101 | 98.1025 | 77.622 | 0.004333 | 1.0065 | 0.024601 | 24.4431 | -0.01845 | 0.15915 | 0.00679 | 0.50205 | 0.50205 | 0.013187
2019-03-19 15:46 | 1553024785.7678 | 256 | 0.20814 | 0.001035 | 2.8e-05 | 101.6663 | 101.888 | 122.5647 | 21.2612 | 0.002535 | 0.002326 | 250.7038 | 0.26607 | 8.4031 | 0.003219 | 0.00489 | 0.004733 | 5.6e-05 | 0.001068 | 0.33085 | 1.3534 | 2.5481 | 101.21 | 20.82 | 17.814 | 82.7994 | 99.2896 | 77.7763 | 0.004334 | 1.006 | 0.024585 | 24.4373 | -0.019487 | 0.15829 | 0.006837 | 0.50525 | 0.50525 | 0.013033
2019-03-19 15:46 | 1553024786.7665 | 257 | 0.20817 | 0.001038 | 2.8e-05 | 101.6342 | 101.676 | 123.0029 | 21.2612 | 0.002504 | 0.002399 | 250.6747 | 0.2669 | 8.4026 | 0.003192 | 0.00479 | 0.004727 | 5.5e-05 | 0.001062 | 0.28816 | 1.4084 | 2.5481 | 101.21 | 20.82 | 17.814 | 82.0662 | 98.6341 | 79.5265 | 0.004334 | 1.001 | 0.024568 | 24.5419 | -0.003682 | 0.15878 | 0.006622 | 0.48895 | 0.48895 | 0.013509
2019-03-19 15:46 | 1553024787.7676 | 258 | 0.20819 | 0.001033 | 2.7e-05 | 101.5601 | 101.8394 | 123.0712 | 21.2612 | 0.002605 | 0.002279 | 250.657 | 0.26676 | 8.4026 | 0.003177 | 0.005023 | 0.004885 | 5.6e-05 | 0.001066 | 0.32908 | 1.3727 | 2.5481 | 101.21 | 20.82 | 17.814 | 84.4799 | 98.2774 | 76.6569 | 0.004334 | 1.0093 | 0.024675 | 24.4487 | -0.024568 | 0.1587 | 0.006559 | 0.48646 | 0.48646 | 0.013571
2019-03-19 15:46 | 1553024788.7666 | 259 | 0.20819 | 0.001035 | 2.8e-05 | 101.657 | 101.749 | 123.2506 | 21.2612 | 0.002551 | 0.002331 | 250.6293 | 0.2661 | 8.4027 | 0.003143 | 0.004742 | 0.004606 | 5.7e-05 | 0.001065 | 0.30502 | 1.3667 | 2.5481 | 101.21 | 20.82 | 17.813 | 83.1809 | 97.448 | 77.9058 | 0.004334 | 1.0057 | 0.024649 | 24.5101 | -0.00809 | 0.1583 | 0.006527 | 0.48349 | 0.48349 | 0.013621
2019-03-19 15:46 | 1553024789.7672 | 260 | 0.20817 | 0.001033 | 2.8e-05 | 101.636 | 101.6606 | 123.0229 | 21.2612 | 0.002537 | 0.002356 | 250.5816 | 0.26658 | 8.4045 | 0.003248 | 0.005175 | 0.005041 | 5.6e-05 | 0.001064 | 0.29247 | 1.3286 | 2.5481 | 101.21 | 20.82 | 17.812 | 82.858 | 99.9934 | 78.5053 | 0.004334 | 1.004 | 0.024605 | 24.5084 | -0.002167 | 0.15859 | 0.00664 | 0.49106 | 0.49106 | 0.013435
2019-03-19 15:46 | 1553024790.7677 | 261 | 0.20817 | 0.001039 | 2.7e-05 | 101.6249 | 101.8617 | 123.0247 | 21.2612 | 0.002494 | 0.002354 | 250.5418 | 0.26604 | 8.4047 | 0.003166 | 0.004611 | 0.004519 | 5.5e-05 | 0.001062 | 0.32964 | 1.3552 | 2.5481 | 101.21 | 20.82 | 17.811 | 81.8065 | 98.0029 | 78.4562 | 0.004333 | 1.0041 | 0.024607 | 24.5068 | -0.020816 | 0.15827 | 0.006657 | 0.4924 | 0.4924 | 0.013371
2019-03-19 15:47 | 1553024791.7678 | 262 | 0.20816 | 0.001027 | 2.7e-05 | 101.6162 | 101.8291 | 123.1801 | 21.2612 | 0.002531 | 0.002342 | 250.5202 | 0.26662 | 8.4029 | 0.003157 | 0.005896 | 0.005823 | 5.4e-05 | 0.001061 | 0.28335 | 1.3321 | 2.5481 | 101.21 | 20.82 | 17.812 | 82.7097 | 97.7797 | 78.1607 | 0.004334 | 1.0049 | 0.024633 | 24.512 | -0.018724 | 0.15862 | 0.006741 | 0.49914 | 0.49914 | 0.01322
2019-03-19 15:47 | 1553024792.7668 | 263 | 0.20815 | 0.001028 | 2.7e-05 | 101.6336 | 101.7144 | 123.1574 | 21.2612 | 0.002457 | 0.002348 | 250.4904 | 0.26625 | 8.4026 | 0.00318 | 0.004818 | 0.004717 | 5.3e-05 | 0.001067 | 0.33069 | 1.3263 | 2.5481 | 101.21 | 20.82 | 17.813 | 80.9316 | 98.3392 | 78.3065 | 0.004334 | 1.0045 | 0.024626 | 24.5148 | -0.007108 | 0.15839 | 0.00676 | 0.50042 | 0.50042 | 0.013167
2019-03-19 15:47 | 1553024793.7671 | 264 | 0.20814 | 0.001024 | 2.7e-05 | 101.6207 | 101.6504 | 122.8116 | 21.2612 | 0.002569 | 0.002325 | 250.4463 | 0.26736 | 8.4009 | 0.003179 | 0.004826 | 0.004766 | 5.5e-05 | 0.001075 | 0.32156 | 1.3417 | 2.548 | 101.21 | 20.82 | 17.817 | 83.6091 | 98.3173 | 77.7474 | 0.004335 | 1.0061 | 0.024611 | 24.4609 | -0.00261 | 0.15905 | 0.00683 | 0.50533 | 0.50533 | 0.013094
2019-03-19 15:47 | 1553024794.7679 | 265 | 0.20813 | 0.001029 | 2.6e-05 | 101.5466 | 101.6651 | 122.7809 | 21.2612 | 0.00248 | 0.002354 | 250.422 | 0.26627 | 8.403 | 0.00318 | 0.004606 | 0.004508 | 5.5e-05 | 0.001078 | 0.31957 | 1.3288 | 2.5482 | 101.21 | 20.82 | 17.818 | 81.4852 | 98.3487 | 78.4639 | 0.004335 | 1.0041 | 0.024582 | 24.4828 | -0.010433 | 0.15841 | 0.006882 | 0.5086 | 0.5086 | 0.012956
2019-03-19 15:47 | 1553024795.7669 | 266 | 0.20814 | 0.001026 | 2.6e-05 | 101.564 | 101.6329 | 122.9128 | 21.2612 | 0.002533 | 0.002386 | 250.3815 | 0.26677 | 8.4042 | 0.003172 | 0.004606 | 0.004555 | 5.3e-05 | 0.001075 | 0.30773 | 1.2922 | 2.548 | 101.21 | 20.82 | 17.82 | 82.7486 | 98.1409 | 79.2312 | 0.004336 | 1.0019 | 0.024569 | 24.5227 | -0.006067 | 0.1587 | 0.006844 | 0.50554 | 0.50554 | 0.013059
2019-03-19 15:47 | 1553024796.7681 | 267 | 0.20814 | 0.001025 | 2.6e-05 | 101.5793 | 101.7863 | 122.5328 | 21.2612 | 0.002526 | 0.002329 | 250.3481 | 0.26628 | 8.4007 | 0.003231 | 0.005102 | 0.004893 | 5.3e-05 | 0.001082 | 0.30934 | 1.3354 | 2.548 | 101.21 | 20.82 | 17.823 | 82.5929 | 99.5862 | 77.8637 | 0.004336 | 1.0058 | 0.024579 | 24.4372 | -0.018212 | 0.15841 | 0.006861 | 0.507 | 0.507 | 0.012998
2019-03-19 15:47 | 1553024797.7671 | 268 | 0.20815 | 0.001025 | 2.5e-05 | 101.5295 | 101.7654 | 122.4778 | 21.2612 | 0.002468 | 0.002356 | 250.3035 | 0.2665 | 8.4042 | 0.00325 | 0.004321 | 0.004304 | 5.4e-05 | 0.001087 | 0.30038 | 1.3191 | 2.548 | 101.21 | 20.82 | 17.824 | 81.1951 | 100.0403 | 78.5021 | 0.004336 | 1.004 | 0.024551 | 24.4539 | -0.020764 | 0.15854 | 0.006814 | 0.50301 | 0.50301 | 0.013112
2019-03-19 15:47 | 1553024798.7682 | 269 | 0.20815 | 0.001033 | 2.5e-05 | 101.5787 | 101.4825 | 122.3189 | 21.2612 | 0.002541 | 0.002286 | 250.2896 | 0.26646 | 8.4013 | 0.003247 | 0.005241 | 0.005128 | 5.3e-05 | 0.001089 | 0.3147 | 1.3119 | 2.548 | 101.21 | 20.83 | 17.834 | 82.9547 | 99.9627 | 76.809 | 0.004342 | 1.0088 | 0.024594 | 24.3791 | 0.00847 | 0.15852 | 0.006764 | 0.50017 | 0.50017 | 0.013184
2019-03-19 15:47 | 1553024799.7672 | 270 | 0.20815 | 0.001032 | 2.5e-05 | 101.5622 | 101.7555 | 122.3079 | 21.2612 | 0.002538 | 0.002284 | 250.2766 | 0.26603 | 8.4044 | 0.003266 | 0.004654 | 0.004528 | 5.3e-05 | 0.001091 | 0.2844 | 1.3344 | 2.5481 | 101.21 | 20.83 | 17.834 | 82.8664 | 100.4319 | 76.7677 | 0.004342 | 1.0089 | 0.024594 | 24.3766 | -0.01701 | 0.15826 | 0.006753 | 0.49941 | 0.49941 | 0.013183
2019-03-19 15:47 | 1553024800.7676 | 271 | 0.20815 | 0.00103 | 2.5e-05 | 101.4637 | 101.4397 | 122.4105 | 21.2612 | 0.002563 | 0.002356 | 250.2702 | 0.26636 | 8.4045 | 0.003199 | 0.004935 | 0.004785 | 5.4e-05 | 0.001098 | 0.34506 | 1.3412 | 2.5481 | 101.21 | 20.83 | 17.839 | 83.4875 | 98.7916 | 78.4997 | 0.004343 | 1.004 | 0.024544 | 24.4471 | 0.002118 | 0.15846 | 0.006777 | 0.50011 | 0.50011 | 0.013181
2019-03-19 15:47 | 1553024801.7664 | 272 | 0.20819 | 0.001033 | 2.4e-05 | 101.5088 | 101.461 | 122.3222 | 21.2612 | 0.002534 | 0.002333 | 250.2549 | 0.26652 | 8.401 | 0.003187 | 0.004786 | 0.004657 | 5.3e-05 | 0.001095 | 0.31636 | 1.3342 | 2.5481 | 101.21 | 20.83 | 17.841 | 82.7712 | 98.5016 | 77.9478 | 0.004343 | 1.0055 | 0.024555 | 24.4191 | 0.004213 | 0.15855 | 0.006562 | 0.48447 | 0.48447 | 0.013615
2019-03-19 15:47 | 1553024802.7673 | 273 | 0.2082 | 0.001035 | 2.5e-05 | 101.4655 | 101.6027 | 122.2121 | 21.2612 | 0.002518 | 0.002307 | 250.2339 | 0.26662 | 8.401 | 0.003275 | 0.004374 | 0.004253 | 5.5e-05 | 0.001094 | 0.323 | 1.306 | 2.548 | 101.21 | 20.83 | 17.844 | 82.3931 | 100.6447 | 77.3171 | 0.004344 | 1.0074 | 0.024566 | 24.3862 | -0.012087 | 0.15862 | 0.006474 | 0.47809 | 0.47809 | 0.013801
2019-03-19 15:47 | 1553024803.7664 | 274 | 0.20818 | 0.001032 | 2.5e-05 | 101.4445 | 101.463 | 122.4255 | 21.2612 | 0.002528 | 0.002301 | 250.19 | 0.26661 | 8.4047 | 0.003259 | 0.004464 | 0.004427 | 5.5e-05 | 0.001088 | 0.3027 | 1.328 | 2.5479 | 101.21 | 20.83 | 17.843 | 82.6287 | 100.2514 | 77.1768 | 0.004344 | 1.0078 | 0.024592 | 24.4026 | -0.001629 | 0.15861 | 0.006573 | 0.486 | 0.486 | 0.013576
2019-03-19 15:47 | 1553024804.7674 | 275 | 0.2082 | 0.001033 | 2.5e-05 | 101.4163 | 101.5774 | 122.5267 | 21.2612 | 0.002432 | 0.002353 | 250.1716 | 0.26569 | 8.4016 | 0.003261 | 0.004992 | 0.004828 | 5.5e-05 | 0.00109 | 0.30309 | 1.3023 | 2.5481 | 101.21 | 20.83 | 17.842 | 80.3226 | 100.2947 | 78.4375 | 0.004344 | 1.0041 | 0.024558 | 24.4565 | -0.014195 | 0.15806 | 0.006481 | 0.47851 | 0.47851 | 0.013741
2019-03-19 15:47 | 1553024805.7675 | 276 | 0.20819 | 0.001032 | 2.5e-05 | 101.3776 | 101.4866 | 122.3288 | 21.2612 | 0.002495 | 0.002331 | 250.1358 | 0.26645 | 8.4035 | 0.003231 | 0.004553 | 0.004479 | 5.7e-05 | 0.001089 | 0.31526 | 1.3066 | 2.548 | 101.21 | 20.83 | 17.837 | 81.8464 | 99.5782 | 77.8971 | 0.004342 | 1.0057 | 0.024557 | 24.418 | -0.009615 | 0.15851 | 0.006519 | 0.48131 | 0.48131 | 0.0137
2019-03-19 15:47 | 1553024806.7676 | 277 | 0.20822 | 0.00103 | 2.6e-05 | 101.3494 | 101.4219 | 122.4239 | 21.2612 | 0.002497 | 0.002274 | 250.1105 | 0.26703 | 8.4047 | 0.00326 | 0.004776 | 0.004763 | 5.9e-05 | 0.001084 | 0.30071 | 1.3007 | 2.5481 | 101.21 | 20.83 | 17.836 | 81.8806 | 100.2727 | 76.5416 | 0.004342 | 1.0096 | 0.024614 | 24.3803 | -0.006397 | 0.15886 | 0.006374 | 0.47159 | 0.47159 | 0.014013
2019-03-19 15:47 | 1553024807.768 | 278 | 0.20822 | 0.00103 | 2.6e-05 | 101.3644 | 101.4366 | 122.786 | 21.2612 | 0.002532 | 0.002312 | 250.0873 | 0.2675 | 8.4011 | 0.003183 | 0.004814 | 0.004671 | 5.9e-05 | 0.001079 | 0.30773 | 1.3296 | 2.5481 | 101.21 | 20.83 | 17.833 | 82.7229 | 98.4067 | 77.4341 | 0.004341 | 1.007 | 0.024619 | 24.4474 | -0.006372 | 0.15914 | 0.006367 | 0.47118 | 0.47118 | 0.01405
2019-03-19 15:47 | 1553024808.7676 | 279 | 0.20821 | 0.001029 | 2.7e-05 | 101.3704 | 101.4969 | 122.7806 | 21.2612 | 0.002437 | 0.002361 | 250.0564 | 0.26655 | 8.4025 | 0.003212 | 0.005473 | 0.005316 | 5.7e-05 | 0.001076 | 0.29662 | 1.294 | 2.5481 | 101.21 | 20.83 | 17.831 | 80.4428 | 99.1242 | 78.6105 | 0.004341 | 1.0037 | 0.024577 | 24.4879 | -0.011156 | 0.15857 | 0.006445 | 0.4761 | 0.4761 | 0.013855
2019-03-19 15:47 | 1553024809.7667 | 280 | 0.20822 | 0.001025 | 2.6e-05 | 101.3539 | 101.2955 | 122.7785 | 21.2612 | 0.002515 | 0.002327 | 250.0337 | 0.26678 | 8.4016 | 0.003207 | 0.004574 | 0.004567 | 5.7e-05 | 0.001075 | 0.31022 | 1.3005 | 2.5481 | 101.21 | 20.83 | 17.831 | 82.3229 | 98.9877 | 77.7943 | 0.004341 | 1.006 | 0.024606 | 24.4592 | 0.005153 | 0.15871 | 0.006371 | 0.47119 | 0.47119 | 0.014012
2019-03-19 15:47 | 1553024810.7678 | 281 | 0.20822 | 0.001029 | 2.6e-05 | 101.2378 | 101.3175 | 122.706 | 21.2612 | 0.00246 | 0.002343 | 250.0039 | 0.26595 | 8.4034 | 0.003222 | 0.005059 | 0.004918 | 5.7e-05 | 0.001077 | 0.31238 | 1.3001 | 2.5479 | 101.21 | 20.83 | 17.828 | 80.9886 | 99.3587 | 78.1818 | 0.00434 | 1.0049 | 0.024585 | 24.4655 | -0.007038 | 0.15822 | 0.006358 | 0.46984 | 0.46984 | 0.014009
2019-03-19 15:47 | 1553024811.7672 | 282 | 0.20821 | 0.001024 | 2.6e-05 | 101.2921 | 101.4421 | 122.7703 | 21.2612 | 0.002558 | 0.002304 | 249.949 | 0.26679 | 8.4 | 0.003292 | 0.004469 | 0.00438 | 5.9e-05 | 0.001076 | 0.29839 | 1.3172 | 2.548 | 101.21 | 20.83 | 17.829 | 83.3641 | 101.0574 | 77.2518 | 0.00434 | 1.0075 | 0.024624 | 24.4395 | -0.013239 | 0.15871 | 0.006413 | 0.47468 | 0.47468 | 0.013909
2019-03-19 15:47 | 1553024812.7668 | 283 | 0.20821 | 0.00102 | 2.6e-05 | 101.3047 | 101.3065 | 122.9482 | 21.2612 | 0.002482 | 0.002288 | 249.9141 | 0.26717 | 8.4021 | 0.003206 | 0.004497 | 0.004426 | 5.9e-05 | 0.001075 | 0.32278 | 1.2937 | 2.5479 | 101.21 | 20.83 | 17.829 | 81.5407 | 98.9657 | 76.8777 | 0.00434 | 1.0086 | 0.024655 | 24.4442 | -0.00016 | 0.15894 | 0.006464 | 0.47902 | 0.47902 | 0.013803
2019-03-19 15:47 | 1553024813.7669 | 284 | 0.20822 | 0.001016 | 2.6e-05 | 101.236 | 101.2524 | 122.6593 | 21.2612 | 0.002461 | 0.002265 | 249.9064 | 0.26648 | 8.4014 | 0.003195 | 0.004313 | 0.004193 | 5.7e-05 | 0.001079 | 0.32493 | 1.327 | 2.5482 | 101.21 | 20.83 | 17.833 | 81.0321 | 98.7046 | 76.3085 | 0.004341 | 1.0103 | 0.024646 | 24.3956 | -0.00145 | 0.15853 | 0.0064 | 0.47412 | 0.47412 | 0.01391
2019-03-19 15:47 | 1553024814.7679 | 285 | 0.20822 | 0.001014 | 2.5e-05 | 101.2741 | 101.2198 | 122.6816 | 21.2612 | 0.002476 | 0.002287 | 249.8823 | 0.26627 | 8.3991 | 0.003251 | 0.00467 | 0.004613 | 5.7e-05 | 0.001074 | 0.32659 | 1.3508 | 2.5481 | 101.21 | 20.83 | 17.837 | 81.3941 | 100.0553 | 76.8347 | 0.004342 | 1.0087 | 0.02463 | 24.4161 | 0.004791 | 0.1584 | 0.006388 | 0.473 | 0.473 | 0.013931
2019-03-19 15:47 | 1553024815.767 | 286 | 0.20821 | 0.001015 | 2.5e-05 | 101.1768 | 101.162 | 122.9616 | 21.2612 | 0.002527 | 0.002362 | 249.8707 | 0.26739 | 8.4045 | 0.003276 | 0.005522 | 0.005407 | 6.1e-05 | 0.001073 | 0.32908 | 1.2916 | 2.548 | 101.21 | 20.83 | 17.84 | 82.62 | 100.6532 | 78.648 | 0.004343 | 1.0035 | 0.024594 | 24.5072 | 0.001316 | 0.15907 | 0.006467 | 0.47819 | 0.47819 | 0.013838
2019-03-19 15:47 | 1553024816.766 | 287 | 0.20822 | 0.001011 | 2.6e-05 | 101.1573 | 101.1876 | 123.1686 | 21.2612 | 0.00252 | 0.0023 | 249.842 | 0.26659 | 8.402 | 0.003279 | 0.004559 | 0.004354 | 5.8e-05 | 0.001071 | 0.33251 | 1.3038 | 2.548 | 101.21 | 20.83 | 17.85 | 82.4313 | 100.7364 | 77.1573 | 0.004345 | 1.0078 | 0.024667 | 24.4758 | -0.00268 | 0.15859 | 0.006376 | 0.47275 | 0.47275 | 0.013956
2019-03-19 15:47 | 1553024817.767 | 288 | 0.20822 | 0.00101 | 2.6e-05 | 101.2018 | 101.1325 | 122.7421 | 21.2612 | 0.002489 | 0.002334 | 249.8188 | 0.2663 | 8.4011 | 0.00331 | 0.004499 | 0.004491 | 6e-05 | 0.001072 | 0.31382 | 1.3136 | 2.548 | 101.21 | 20.83 | 17.854 | 81.6864 | 101.4782 | 77.9752 | 0.004346 | 1.0055 | 0.024596 | 24.4619 | 0.006122 | 0.15842 | 0.006358 | 0.47006 | 0.47006 | 0.01402
2019-03-19 15:47 | 1553024818.7661 | 289 | 0.20823 | 0.00101 | 2.6e-05 | 101.1333 | 101.1969 | 122.7954 | 21.2612 | 0.002483 | 0.002311 | 249.8121 | 0.26728 | 8.4015 | 0.003322 | 0.004643 | 0.00459 | 6e-05 | 0.001068 | 0.32587 | 1.3231 | 2.5481 | 101.21 | 20.83 | 17.866 | 81.546 | 101.7798 | 77.4185 | 0.004349 | 1.0071 | 0.024621 | 24.4478 | -0.00562 | 0.15901 | 0.00635 | 0.46991 | 0.46991 | 0.014076
2019-03-19 15:47 | 1553024819.7672 | 290 | 0.20823 | 0.001008 | 2.6e-05 | 101.1156 | 101.1835 | 122.8248 | 21.2612 | 0.002501 | 0.002274 | 249.7819 | 0.26738 | 8.4033 | 0.00325 | 0.004692 | 0.004481 | 5.7e-05 | 0.001061 | 0.3194 | 1.2701 | 2.548 | 101.21 | 20.83 | 17.872 | 81.9976 | 100.0413 | 76.5228 | 0.004351 | 1.0096 | 0.024655 | 24.4195 | -0.006006 | 0.15906 | 0.006317 | 0.46813 | 0.46813 | 0.014135
2019-03-19 15:47 | 1553024820.7673 | 291 | 0.20824 | 0.001005 | 2.7e-05 | 101.0427 | 101.1246 | 122.8316 | 21.2612 | 0.002501 | 0.002342 | 249.7416 | 0.26603 | 8.4049 | 0.003262 | 0.004617 | 0.004497 | 5.8e-05 | 0.001056 | 0.34092 | 1.325 | 2.548 | 101.21 | 20.83 | 17.877 | 81.9914 | 100.315 | 78.1607 | 0.004352 | 1.0049 | 0.024598 | 24.4773 | -0.00725 | 0.15826 | 0.006273 | 0.46375 | 0.46375 | 0.014197
2019-03-19 15:47 | 1553024821.7668 | 292 | 0.20823 | 0.001002 | 2.6e-05 | 101.0118 | 101.1274 | 122.5081 | 21.2612 | 0.002514 | 0.002334 | 249.7217 | 0.2671 | 8.4047 | 0.00325 | 0.004187 | 0.004079 | 5.9e-05 | 0.001049 | 0.27704 | 1.3155 | 2.5481 | 101.21 | 20.83 | 17.885 | 82.3073 | 100.0319 | 77.9783 | 0.004354 | 1.0055 | 0.024572 | 24.4387 | -0.010229 | 0.1589 | 0.006346 | 0.46872 | 0.46872 | 0.014103
2019-03-19 15:47 | 1553024822.7674 | 293 | 0.20823 | 0.001002 | 2.6e-05 | 101.0097 | 101.0123 | 122.3973 | 21.2612 | 0.002498 | 0.002327 | 249.6874 | 0.26663 | 8.4017 | 0.003229 | 0.005623 | 0.005406 | 5.6e-05 | 0.001051 | 0.35441 | 1.3496 | 2.548 | 101.21 | 20.83 | 17.887 | 81.9104 | 99.5391 | 77.8082 | 0.004354 | 1.0059 | 0.024567 | 24.4217 | -0.000231 | 0.15862 | 0.006322 | 0.46682 | 0.46682 | 0.014135
2019-03-19 15:47 | 1553024823.7668 | 294 | 0.20823 | 0.000992 | 2.7e-05 | 101.0283 | 101.0753 | 122.6485 | 21.2612 | 0.002489 | 0.00231 | 249.6541 | 0.26584 | 8.4057 | 0.00323 | 0.004747 | 0.004654 | 5.7e-05 | 0.001056 | 0.30541 | 1.3182 | 2.548 | 101.21 | 20.83 | 17.89 | 81.702 | 99.5422 | 77.3952 | 0.004355 | 1.0071 | 0.024607 | 24.4324 | -0.004162 | 0.15815 | 0.006323 | 0.46759 | 0.46759 | 0.01407
2019-03-19 15:47 | 1553024824.7679 | 295 | 0.20823 | 0.000994 | 2.6e-05 | 100.8515 | 100.9147 | 122.5131 | 21.2612 | 0.00254 | 0.002334 | 249.6193 | 0.26639 | 8.4031 | 0.003186 | 0.004896 | 0.004819 | 5.5e-05 | 0.001063 | 0.33804 | 1.3372 | 2.5481 | 101.21 | 20.83 | 17.889 | 82.914 | 98.4915 | 77.9627 | 0.004355 | 1.0055 | 0.024573 | 24.4387 | -0.005599 | 0.15848 | 0.006354 | 0.46931 | 0.46931 | 0.014048
2019-03-19 15:47 | 1553024825.7679 | 296 | 0.20822 | 0.000993 | 2.6e-05 | 100.8668 | 100.8859 | 122.1823 | 21.2612 | 0.002511 | 0.002314 | 249.5739 | 0.26575 | 8.4028 | 0.003198 | 0.004442 | 0.004377 | 5.4e-05 | 0.00106 | 0.31852 | 1.3483 | 2.5481 | 101.21 | 20.83 | 17.887 | 82.2294 | 98.7846 | 77.4918 | 0.004354 | 1.0069 | 0.024556 | 24.3893 | -0.001692 | 0.1581 | 0.006378 | 0.4708 | 0.4708 | 0.01397
2019-03-19 15:47 | 1553024826.7665 | 297 | 0.20822 | 0.000992 | 2.6e-05 | 100.8803 | 100.8558 | 122.0772 | 21.2612 | 0.00247 | 0.002342 | 249.5529 | 0.26567 | 8.4025 | 0.003098 | 0.004652 | 0.004557 | 5.4e-05 | 0.001063 | 0.32156 | 1.4116 | 2.5481 | 101.21 | 20.83 | 17.883 | 81.2357 | 96.3599 | 78.16 | 0.004354 | 1.0049 | 0.024523 | 24.402 | 0.002178 | 0.15805 | 0.006406 | 0.47222 | 0.47222 | 0.013923
2019-03-19 15:47 | 1553024827.7667 | 298 | 0.20825 | 0.000985 | 2.5e-05 | 100.7951 | 100.8349 | 122.3938 | 21.2612 | 0.002522 | 0.002356 | 249.5279 | 0.26615 | 8.4022 | 0.003154 | 0.004901 | 0.004803 | 5.3e-05 | 0.001066 | 0.3235 | 1.3625 | 2.548 | 101.21 | 20.83 | 17.88 | 82.4803 | 97.7207 | 78.5067 | 0.004353 | 1.0039 | 0.024542 | 24.4457 | -0.003529 | 0.15833 | 0.006249 | 0.461 | 0.461 | 0.014288
2019-03-19 15:47 | 1553024828.7681 | 299 | 0.20825 | 0.000989 | 2.6e-05 | 100.8128 | 100.6198 | 122.3187 | 21.2612 | 0.002529 | 0.002331 | 249.4893 | 0.26665 | 8.4043 | 0.003186 | 0.005042 | 0.004951 | 5.4e-05 | 0.001066 | 0.32101 | 1.3262 | 2.548 | 101.21 | 20.83 | 17.877 | 82.6582 | 98.4789 | 77.8909 | 0.004352 | 1.0057 | 0.024556 | 24.4168 | 0.017142 | 0.15863 | 0.006207 | 0.45816 | 0.45816 | 0.014403
2019-03-19 15:47 | 1553024829.7667 | 300 | 0.20825 | 0.000988 | 2.6e-05 | 100.7585 | 100.7475 | 122.1572 | 21.2612 | 0.002523 | 0.002298 | 249.4548 | 0.26643 | 8.4024 | 0.003147 | 0.004998 | 0.004894 | 5.4e-05 | 0.001066 | 0.31741 | 1.2965 | 2.548 | 101.21 | 20.83 | 17.871 | 82.5241 | 97.5524 | 77.1115 | 0.004351 | 1.0079 | 0.024567 | 24.3735 | 0.000972 | 0.1585 | 0.006264 | 0.46254 | 0.46254 | 0.014255
2019-03-19 15:47 | 1553024830.7682 | 301 | 0.20827 | 0.000985 | 2.5e-05 | 100.745 | 100.5455 | 122.2823 | 21.2612 | 0.002557 | 0.002319 | 249.44 | 0.26628 | 8.4039 | 0.003094 | 0.004845 | 0.00476 | 5.5e-05 | 0.001071 | 0.3361 | 1.3218 | 2.548 | 101.21 | 20.83 | 17.866 | 83.3237 | 96.2611 | 77.6134 | 0.004349 | 1.0065 | 0.024562 | 24.4035 | 0.017731 | 0.15841 | 0.006124 | 0.45216 | 0.45216 | 0.014574
2019-03-19 15:47 | 1553024831.7659 | 302 | 0.20827 | 0.000978 | 2.6e-05 | 100.7513 | 100.7273 | 122.2556 | 21.2612 | 0.002476 | 0.002331 | 249.4086 | 0.26586 | 8.4045 | 0.003121 | 0.004479 | 0.004377 | 5.4e-05 | 0.001075 | 0.3194 | 1.3159 | 2.5481 | 101.21 | 20.83 | 17.858 | 81.3777 | 96.9278 | 77.8957 | 0.004347 | 1.0057 | 0.02455 | 24.4106 | 0.002127 | 0.15816 | 0.006154 | 0.45412 | 0.45412 | 0.014488
2019-03-19 15:47 | 1553024832.766 | 303 | 0.20825 | 0.000979 | 2.6e-05 | 100.6631 | 100.6157 | 122.2534 | 21.2612 | 0.002493 | 0.002284 | 249.3742 | 0.26621 | 8.4054 | 0.003153 | 0.004826 | 0.004686 | 5.2e-05 | 0.001077 | 0.32333 | 1.2979 | 2.548 | 101.21 | 20.83 | 17.855 | 81.7854 | 97.7004 | 76.7825 | 0.004347 | 1.0089 | 0.024588 | 24.3717 | 0.004211 | 0.15837 | 0.006233 | 0.46066 | 0.46066 | 0.014301
2019-03-19 15:47 | 1553024833.767 | 304 | 0.20828 | 0.00098 | 2.6e-05 | 100.7078 | 100.728 | 122.1267 | 21.2612 | 0.002505 | 0.002354 | 249.3479 | 0.26587 | 8.4011 | 0.003146 | 0.004407 | 0.004283 | 5.1e-05 | 0.001074 | 0.33716 | 1.3032 | 2.548 | 101.21 | 20.83 | 17.85 | 82.071 | 97.5318 | 78.4601 | 0.004345 | 1.0041 | 0.024517 | 24.4174 | -0.001798 | 0.15817 | 0.006046 | 0.44549 | 0.44549 | 0.014769
2019-03-19 15:47 | 1553024834.767 | 305 | 0.20827 | 0.000974 | 2.6e-05 | 100.6457 | 100.3575 | 122.0931 | 21.2612 | 0.002481 | 0.002315 | 249.3294 | 0.26667 | 8.4032 | 0.003213 | 0.005201 | 0.005139 | 5.4e-05 | 0.001077 | 0.32388 | 1.3076 | 2.5481 | 101.21 | 20.83 | 17.848 | 81.5039 | 99.1382 | 77.5067 | 0.004345 | 1.0068 | 0.024547 | 24.3809 | 0.02565 | 0.15864 | 0.006164 | 0.4548 | 0.4548 | 0.014511
2019-03-19 15:47 | 1553024835.7661 | 306 | 0.20827 | 0.000978 | 2.6e-05 | 100.6667 | 100.638 | 122.199 | 21.2612 | 0.002478 | 0.002267 | 249.3229 | 0.26666 | 8.4051 | 0.003186 | 0.004891 | 0.004764 | 5.8e-05 | 0.001075 | 0.30541 | 1.3522 | 2.5481 | 101.21 | 20.83 | 17.847 | 81.4393 | 98.4852 | 76.3607 | 0.004345 | 1.0101 | 0.024598 | 24.3516 | 0.002552 | 0.15864 | 0.006112 | 0.45186 | 0.45186 | 0.014605
2019-03-19 15:47 | 1553024836.7672 | 307 | 0.20829 | 0.000971 | 2.6e-05 | 100.5988 | 100.6082 | 122.1647 | 21.2612 | 0.002463 | 0.002306 | 249.3154 | 0.2676 | 8.4048 | 0.003214 | 0.00496 | 0.004894 | 5.5e-05 | 0.001079 | 0.3126 | 1.3238 | 2.5481 | 101.21 | 20.83 | 17.845 | 81.0797 | 99.1721 | 77.3102 | 0.004344 | 1.0074 | 0.024561 | 24.3812 | -0.000828 | 0.1592 | 0.006034 | 0.44539 | 0.44539 | 0.014869
2019-03-19 15:47 | 1553024837.7672 | 308 | 0.20826 | 0.00097 | 2.6e-05 | 100.5991 | 100.6369 | 122.0692 | 21.2612 | 0.002425 | 0.002338 | 249.2839 | 0.2672 | 8.4056 | 0.003192 | 0.005436 | 0.005286 | 5.2e-05 | 0.001071 | 0.31188 | 1.3326 | 2.5481 | 101.21 | 20.83 | 17.843 | 80.1621 | 98.6278 | 78.078 | 0.004344 | 1.0052 | 0.024525 | 24.3983 | -0.003358 | 0.15896 | 0.006181 | 0.45563 | 0.45563 | 0.014513
2019-03-19 15:47 | 1553024838.7673 | 309 | 0.20827 | 0.00097 | 2.7e-05 | 100.6141 | 100.5403 | 122.3725 | 21.2612 | 0.002467 | 0.002256 | 249.2561 | 0.2658 | 8.4003 | 0.003125 | 0.004769 | 0.004641 | 5.5e-05 | 0.001064 | 0.31868 | 1.3384 | 2.5481 | 101.21 | 20.83 | 17.843 | 81.178 | 97.0273 | 76.1074 | 0.004344 | 1.0108 | 0.024624 | 24.36 | 0.006564 | 0.15812 | 0.006108 | 0.45203 | 0.45203 | 0.014552
2019-03-19 15:47 | 1553024839.7673 | 310 | 0.20827 | 0.000968 | 2.6e-05 | 100.5898 | 100.5612 | 122.5201 | 21.2612 | 0.002505 | 0.002278 | 249.2387 | 0.26687 | 8.4039 | 0.003218 | 0.004177 | 0.00404 | 5.3e-05 | 0.001053 | 0.29502 | 1.3115 | 2.548 | 101.21 | 20.83 | 17.841 | 82.0779 | 99.2724 | 76.6187 | 0.004343 | 1.0094 | 0.024621 | 24.3925 | 0.002544 | 0.15876 | 0.006112 | 0.45223 | 0.45223 | 0.014604
2019-03-19 15:47 | 1553024840.7674 | 311 | 0.20827 | 0.000969 | 2.7e-05 | 100.5688 | 100.5061 | 122.4757 | 21.2612 | 0.002458 | 0.002276 | 249.2409 | 0.2666 | 8.4057 | 0.003152 | 0.004758 | 0.004609 | 5.5e-05 | 0.00105 | 0.30165 | 1.3756 | 2.5481 | 101.21 | 20.83 | 17.84 | 80.9542 | 97.662 | 76.5822 | 0.004343 | 1.0095 | 0.024618 | 24.3868 | 0.005582 | 0.1586 | 0.006124 | 0.4531 | 0.4531 | 0.014561
2019-03-19 15:47 | 1553024841.768 | 312 | 0.20827 | 0.00096 | 2.7e-05 | 100.4935 | 100.4407 | 122.7034 | 21.2612 | 0.002461 | 0.002321 | 249.2343 | 0.26713 | 8.4038 | 0.003118 | 0.004518 | 0.004325 | 5.3e-05 | 0.001035 | 0.29927 | 1.3204 | 2.5481 | 101.21 | 20.83 | 17.836 | 81.0158 | 96.8479 | 77.6588 | 0.004342 | 1.0064 | 0.024603 | 24.447 | 0.004703 | 0.15892 | 0.006168 | 0.45604 | 0.45604 | 0.014496
2019-03-19 15:47 | 1553024842.7674 | 313 | 0.20827 | 0.000965 | 2.7e-05 | 100.4284 | 100.3386 | 122.5424 | 21.2612 | 0.002512 | 0.002329 | 249.2341 | 0.26646 | 8.404 | 0.003203 | 0.004568 | 0.004524 | 5.5e-05 | 0.001039 | 0.31725 | 1.2726 | 2.5479 | 101.21 | 20.83 | 17.834 | 82.2479 | 98.9069 | 77.8613 | 0.004342 | 1.0058 | 0.02458 | 24.438 | 0.007999 | 0.15852 | 0.006122 | 0.45223 | 0.45223 | 0.014582
2019-03-19 15:47 | 1553024843.7676 | 314 | 0.20828 | 0.000961 | 2.8e-05 | 100.4479 | 100.4314 | 122.4858 | 21.2612 | 0.002485 | 0.002358 | 249.2123 | 0.26675 | 8.4038 | 0.003194 | 0.004977 | 0.004867 | 5.2e-05 | 0.001039 | 0.3454 | 1.3133 | 2.548 | 101.21 | 20.83 | 17.83 | 81.6061 | 98.6834 | 78.5418 | 0.004341 | 1.0038 | 0.02455 | 24.4561 | 0.001466 | 0.15869 | 0.006088 | 0.44912 | 0.44912 | 0.014699
2019-03-19 15:47 | 1553024844.7669 | 315 | 0.20829 | 0.000959 | 2.8e-05 | 100.4803 | 100.3016 | 122.6938 | 21.2612 | 0.002521 | 0.002316 | 249.1816 | 0.26599 | 8.4014 | 0.003079 | 0.004846 | 0.004728 | 5.3e-05 | 0.001043 | 0.34722 | 1.3135 | 2.5481 | 101.21 | 20.83 | 17.82 | 82.4633 | 95.8961 | 77.5504 | 0.004338 | 1.0067 | 0.024606 | 24.4423 | 0.01592 | 0.15824 | 0.006037 | 0.44632 | 0.44632 | 0.014749
2019-03-19 15:47 | 1553024845.7676 | 316 | 0.2083 | 0.000955 | 2.8e-05 | 100.3489 | 100.1937 | 122.5619 | 21.2612 | 0.002514 | 0.00234 | 249.1434 | 0.26636 | 8.4066 | 0.00308 | 0.004937 | 0.004819 | 5.7e-05 | 0.001046 | 0.32189 | 1.3167 | 2.5479 | 101.21 | 20.83 | 17.813 | 82.3065 | 95.9283 | 78.1249 | 0.004336 | 1.005 | 0.024572 | 24.4492 | 0.013837 | 0.15846 | 0.005967 | 0.44055 | 0.44055 | 0.014962
2019-03-19 15:47 | 1553024846.7667 | 317 | 0.2083 | 0.000959 | 2.9e-05 | 100.4134 | 100.3578 | 122.8065 | 21.2612 | 0.002507 | 0.002352 | 249.1184 | 0.2667 | 8.4023 | 0.003165 | 0.004867 | 0.004794 | 5.1e-05 | 0.001042 | 0.35601 | 1.3413 | 2.5481 | 101.21 | 20.83 | 17.805 | 82.1318 | 97.9778 | 78.3993 | 0.004335 | 1.0043 | 0.024587 | 24.4831 | 0.004953 | 0.15866 | 0.005973 | 0.44119 | 0.44119 | 0.01496
2019-03-19 15:47 | 1553024847.7668 | 318 | 0.20829 | 0.000957 | 2.8e-05 | 100.2948 | 100.1215 | 122.8205 | 21.2612 | 0.002486 | 0.002301 | 249.1071 | 0.26629 | 8.4005 | 0.00321 | 0.004718 | 0.004571 | 5.3e-05 | 0.001037 | 0.31852 | 1.3666 | 2.5481 | 101.21 | 20.83 | 17.789 | 81.6162 | 99.0592 | 77.1785 | 0.004331 | 1.0078 | 0.024632 | 24.442 | 0.015475 | 0.15841 | 0.00603 | 0.44622 | 0.44622 | 0.014769
2019-03-19 15:47 | 1553024848.7661 | 319 | 0.20829 | 0.00095 | 2.8e-05 | 100.26 | 100.291 | 122.7015 | 21.2612 | 0.002532 | 0.002284 | 249.0899 | 0.26617 | 8.4041 | 0.003213 | 0.004211 | 0.004167 | 5.2e-05 | 0.001035 | 0.2943 | 1.3725 | 2.548 | 101.21 | 20.83 | 17.78 | 82.7316 | 99.1515 | 76.7677 | 0.004328 | 1.0089 | 0.024634 | 24.4158 | -0.002763 | 0.15834 | 0.006061 | 0.4486 | 0.4486 | 0.014684
2019-03-19 15:47 | 1553024849.7669 | 320 | 0.2083 | 0.000953 | 2.8e-05 | 100.2828 | 100.0742 | 122.6724 | 21.2612 | 0.002511 | 0.002335 | 249.0639 | 0.26593 | 8.4031 | 0.003113 | 0.004558 | 0.004418 | 5.3e-05 | 0.001043 | 0.33931 | 1.3978 | 2.548 | 101.21 | 20.83 | 17.759 | 82.2356 | 96.7341 | 77.986 | 0.004323 | 1.0054 | 0.024588 | 24.4554 | 0.018627 | 0.1582 | 0.005984 | 0.44208 | 0.44208 | 0.014887
2019-03-19 15:47 | 1553024850.7659 | 321 | 0.2083 | 0.000947 | 2.8e-05 | 100.2237 | 100.2622 | 122.5929 | 21.2612 | 0.002447 | 0.00233 | 249.0159 | 0.26626 | 8.4029 | 0.003118 | 0.003941 | 0.003808 | 5.3e-05 | 0.001048 | 0.29878 | 1.3276 | 2.5481 | 101.21 | 20.83 | 17.749 | 80.6781 | 96.8471 | 77.8846 | 0.004321 | 1.0057 | 0.024584 | 24.4439 | -0.003437 | 0.1584 | 0.006019 | 0.44455 | 0.44455 | 0.014822
2019-03-19 15:48 | 1553024851.766 | 322 | 0.20833 | 0.000945 | 2.8e-05 | 100.1871 | 100.105 | 122.7156 | 21.2612 | 0.002499 | 0.002318 | 248.9712 | 0.26599 | 8.4043 | 0.003112 | 0.004706 | 0.004559 | 5.4e-05 | 0.001048 | 0.31045 | 1.3331 | 2.548 | 101.21 | 20.82 | 17.728 | 81.9445 | 96.7124 | 77.5977 | 0.004313 | 1.0066 | 0.024606 | 24.4461 | 0.00733 | 0.15824 | 0.005831 | 0.43108 | 0.43108 | 0.01527
2019-03-19 15:48 | 1553024852.7674 | 323 | 0.20833 | 0.000941 | 2.9e-05 | 100.1946 | 99.9797 | 123.1874 | 21.2612 | 0.002449 | 0.002302 | 248.9455 | 0.26641 | 8.405 | 0.003106 | 0.004855 | 0.004775 | 5.7e-05 | 0.001051 | 0.3569 | 1.3484 | 2.548 | 101.21 | 20.82 | 17.719 | 80.7295 | 96.5627 | 77.2104 | 0.004311 | 1.0077 | 0.024667 | 24.4796 | 0.019207 | 0.15849 | 0.005796 | 0.42948 | 0.42948 | 0.015351
2019-03-19 15:48 | 1553024853.7671 | 324 | 0.20834 | 0.000942 | 2.8e-05 | 100.1679 | 99.943 | 123.3137 | 21.2612 | 0.002465 | 0.002336 | 248.9388 | 0.2666 | 8.4015 | 0.003118 | 0.004563 | 0.004514 | 5.5e-05 | 0.001055 | 0.29391 | 1.3369 | 2.5481 | 101.21 | 20.83 | 17.709 | 81.1242 | 96.8377 | 78.0204 | 0.004311 | 1.0053 | 0.024651 | 24.5204 | 0.020101 | 0.1586 | 0.005767 | 0.42707 | 0.42707 | 0.015449
2019-03-19 15:48 | 1553024854.767 | 325 | 0.20833 | 0.000938 | 2.9e-05 | 100.1412 | 100.1276 | 123.3055 | 21.2612 | 0.002508 | 0.002305 | 248.9289 | 0.26613 | 8.4045 | 0.003104 | 0.004841 | 0.004735 | 5.2e-05 | 0.001053 | 0.3246 | 1.3203 | 2.5481 | 101.21 | 20.82 | 17.69 | 82.152 | 96.5156 | 77.2768 | 0.004304 | 1.0075 | 0.024677 | 24.4936 | 0.00121 | 0.15832 | 0.005797 | 0.42966 | 0.42966 | 0.015329
2019-03-19 15:48 | 1553024855.7661 | 326 | 0.20832 | 0.000935 | 2.9e-05 | 100.1382 | 99.9862 | 123.5402 | 21.2612 | 0.002518 | 0.00232 | 248.9119 | 0.26658 | 8.4022 | 0.003113 | 0.00469 | 0.004554 | 5.3e-05 | 0.001048 | 0.31349 | 1.3334 | 2.548 | 101.21 | 20.82 | 17.684 | 82.3883 | 96.7138 | 77.6376 | 0.004302 | 1.0064 | 0.024687 | 24.5295 | 0.013587 | 0.15859 | 0.005902 | 0.43768 | 0.43768 | 0.015073
2019-03-19 15:48 | 1553024856.7662 | 327 | 0.20834 | 0.000932 | 3e-05 | 100.1511 | 99.8673 | 123.2057 | 21.2612 | 0.002513 | 0.002323 | 248.8792 | 0.26638 | 8.401 | 0.003175 | 0.004783 | 0.004688 | 5.7e-05 | 0.001042 | 0.3011 | 1.3473 | 2.5481 | 101.21 | 20.82 | 17.677 | 82.269 | 98.2288 | 77.7054 | 0.004301 | 1.0062 | 0.024652 | 24.4987 | 0.025377 | 0.15847 | 0.005749 | 0.42563 | 0.42563 | 0.015488
2019-03-19 15:48 | 1553024857.7672 | 328 | 0.20832 | 0.000928 | 3e-05 | 100.107 | 100.028 | 123.2182 | 21.2612 | 0.002517 | 0.002333 | 248.8494 | 0.2656 | 8.4025 | 0.003166 | 0.004436 | 0.004296 | 5.8e-05 | 0.00104 | 0.32964 | 1.3166 | 2.5482 | 101.21 | 20.82 | 17.669 | 82.3681 | 98.0061 | 77.9384 | 0.004299 | 1.0056 | 0.024645 | 24.508 | 0.007061 | 0.15801 | 0.005895 | 0.43641 | 0.43641 | 0.015062
2019-03-19 15:48 | 1553024858.7676 | 329 | 0.20833 | 0.000922 | 3e-05 | 100.0911 | 99.9081 | 123.0287 | 21.2612 | 0.002498 | 0.002267 | 248.8116 | 0.26724 | 8.4006 | 0.003207 | 0.004517 | 0.004419 | 5.5e-05 | 0.001034 | 0.32637 | 1.3492 | 2.548 | 101.21 | 20.82 | 17.665 | 81.9181 | 99.0058 | 76.3717 | 0.004298 | 1.0101 | 0.024681 | 24.4345 | 0.016366 | 0.15898 | 0.005855 | 0.43408 | 0.43408 | 0.015236
2019-03-19 15:48 | 1553024859.7674 | 330 | 0.20835 | 0.00092 | 3e-05 | 100.0791 | 100.0109 | 122.8687 | 21.2612 | 0.002458 | 0.002298 | 248.798 | 0.26631 | 8.4025 | 0.003177 | 0.004648 | 0.004481 | 5.4e-05 | 0.001032 | 0.31686 | 1.2898 | 2.5481 | 101.21 | 20.82 | 17.666 | 80.9526 | 98.2694 | 77.1059 | 0.004298 | 1.008 | 0.024639 | 24.4442 | 0.006099 | 0.15843 | 0.005735 | 0.42439 | 0.42439 | 0.015529
2019-03-19 15:48 | 1553024860.7665 | 331 | 0.20831 | 0.000918 | 3.1e-05 | 100.1061 | 99.9961 | 122.8194 | 21.2612 | 0.002451 | 0.002351 | 248.7864 | 0.26612 | 8.4033 | 0.00321 | 0.004477 | 0.004398 | 5.2e-05 | 0.001029 | 0.32908 | 1.3654 | 2.5481 | 101.21 | 20.82 | 17.667 | 80.7834 | 99.0662 | 78.3844 | 0.004298 | 1.0043 | 0.024589 | 24.4839 | 0.00983 | 0.15831 | 0.005944 | 0.439 | 0.439 | 0.015002
2019-03-19 15:48 | 1553024861.7675 | 332 | 0.20835 | 0.000919 | 3e-05 | 100.1313 | 100.0043 | 122.7424 | 21.2612 | 0.002499 | 0.002305 | 248.7792 | 0.26589 | 8.4052 | 0.00323 | 0.004446 | 0.004285 | 5.8e-05 | 0.001032 | 0.31708 | 1.3276 | 2.5482 | 101.21 | 20.82 | 17.672 | 81.9445 | 99.5492 | 77.2727 | 0.004299 | 1.0075 | 0.02462 | 24.4375 | 0.011347 | 0.15818 | 0.005738 | 0.42423 | 0.42423 | 0.015511
2019-03-19 15:48 | 1553024862.7676 | 333 | 0.20836 | 0.000916 | 3e-05 | 100.116 | 99.8646 | 122.6553 | 21.2612 | 0.002499 | 0.002278 | 248.7605 | 0.26617 | 8.4012 | 0.003217 | 0.005151 | 0.005121 | 5.4e-05 | 0.001027 | 0.32731 | 1.3133 | 2.5481 | 101.21 | 20.82 | 17.681 | 81.9306 | 99.2284 | 76.6305 | 0.004302 | 1.0093 | 0.024634 | 24.4064 | 0.022486 | 0.15834 | 0.005676 | 0.41993 | 0.41993 | 0.015686
2019-03-19 15:48 | 1553024863.767 | 334 | 0.20836 | 0.000911 | 3e-05 | 100.0818 | 99.8971 | 122.713 | 21.2612 | 0.002482 | 0.002284 | 248.7271 | 0.26633 | 8.4026 | 0.00319 | 0.0045 | 0.004427 | 5.4e-05 | 0.001027 | 0.28241 | 1.3564 | 2.5481 | 101.21 | 20.82 | 17.687 | 81.5258 | 98.591 | 76.7756 | 0.004303 | 1.0089 | 0.024635 | 24.4172 | 0.016516 | 0.15844 | 0.005675 | 0.41986 | 0.41986 | 0.015698
2019-03-19 15:48 | 1553024864.7667 | 335 | 0.20838 | 0.000905 | 3e-05 | 100.1073 | 99.7879 | 122.5506 | 21.2612 | 0.00244 | 0.002311 | 248.6951 | 0.2662 | 8.4025 | 0.00325 | 0.004543 | 0.004446 | 5.3e-05 | 0.001028 | 0.32444 | 1.3752 | 2.5481 | 101.21 | 20.82 | 17.699 | 80.5284 | 100.0256 | 77.4178 | 0.004306 | 1.0071 | 0.024596 | 24.4234 | 0.028583 | 0.15836 | 0.005571 | 0.41152 | 0.41152 | 0.016009
2019-03-19 15:48 | 1553024865.7667 | 336 | 0.20836 | 0.000905 | 3e-05 | 100.0674 | 99.7766 | 122.5553 | 21.2612 | 0.002408 | 0.002231 | 248.6644 | 0.26565 | 8.4049 | 0.003275 | 0.004894 | 0.004766 | 5.3e-05 | 0.001031 | 0.33843 | 1.3254 | 2.548 | 101.21 | 20.82 | 17.707 | 79.7565 | 100.6399 | 75.5073 | 0.004308 | 1.0126 | 0.024664 | 24.3572 | 0.026029 | 0.15804 | 0.005653 | 0.41868 | 0.41868 | 0.015703
2019-03-19 15:48 | 1553024866.7671 | 337 | 0.20836 | 0.000898 | 3e-05 | 100.0917 | 100.0139 | 122.6243 | 21.2612 | 0.002531 | 0.002325 | 248.6447 | 0.26608 | 8.4029 | 0.003192 | 0.00485 | 0.004727 | 5.4e-05 | 0.001038 | 0.31902 | 1.3467 | 2.548 | 101.21 | 20.82 | 17.715 | 82.7051 | 98.6316 | 77.7546 | 0.00431 | 1.0061 | 0.024592 | 24.4425 | 0.00695 | 0.15829 | 0.005694 | 0.42052 | 0.42052 | 0.015659
2019-03-19 15:48 | 1553024867.7672 | 338 | 0.20836 | 0.000898 | 3e-05 | 100.0491 | 99.7783 | 122.5074 | 21.2612 | 0.002457 | 0.002272 | 248.6188 | 0.26585 | 8.4021 | 0.003114 | 0.0045 | 0.004393 | 5.4e-05 | 0.001044 | 0.32084 | 1.3433 | 2.5481 | 101.21 | 20.82 | 17.73 | 80.9355 | 96.7578 | 76.4832 | 0.004314 | 1.0098 | 0.024625 | 24.3865 | 0.024239 | 0.15815 | 0.005687 | 0.42055 | 0.42055 | 0.015644
2019-03-19 15:48 | 1553024868.7659 | 339 | 0.20834 | 0.000893 | 3.1e-05 | 100.0662 | 99.9571 | 122.7513 | 21.2612 | 0.002518 | 0.002285 | 248.5898 | 0.26703 | 8.4036 | 0.003193 | 0.005307 | 0.005174 | 5.2e-05 | 0.001047 | 0.3402 | 1.3439 | 2.5482 | 101.21 | 20.82 | 17.739 | 82.3883 | 98.649 | 76.795 | 0.004316 | 1.0089 | 0.024638 | 24.4217 | 0.009756 | 0.15886 | 0.005824 | 0.43095 | 0.43095 | 0.015335
2019-03-19 15:48 | 1553024869.7659 | 340 | 0.20836 | 0.00089 | 3.1e-05 | 100.1121 | 99.8461 | 123.0792 | 21.2612 | 0.002505 | 0.00227 | 248.5674 | 0.2671 | 8.4002 | 0.0032 | 0.004264 | 0.004165 | 5.2e-05 | 0.001051 | 0.32853 | 1.358 | 2.5481 | 101.21 | 20.82 | 17.746 | 82.0866 | 98.8349 | 76.4426 | 0.004317 | 1.0099 | 0.024683 | 24.442 | 0.023794 | 0.1589 | 0.005686 | 0.42148 | 0.42148 | 0.015683
2019-03-19 15:48 | 1553024870.7669 | 341 | 0.20842 | 0.000887 | 3.1e-05 | 100.092 | 99.8211 | 123.0487 | 21.2612 | 0.002495 | 0.002274 | 248.5116 | 0.26562 | 8.4037 | 0.003154 | 0.004258 | 0.004138 | 5.1e-05 | 0.001052 | 0.31564 | 1.3762 | 2.548 | 101.21 | 20.82 | 17.759 | 81.8378 | 97.7176 | 76.5423 | 0.004321 | 1.0096 | 0.024677 | 24.4424 | 0.024238 | 0.15801 | 0.005336 | 0.39526 | 0.39526 | 0.01663
2019-03-19 15:48 | 1553024871.7674 | 342 | 0.20841 | 0.000885 | 3.2e-05 | 100.1277 | 99.9759 | 122.9525 | 21.2612 | 0.002431 | 0.002312 | 248.4859 | 0.26632 | 8.4033 | 0.003228 | 0.005186 | 0.005084 | 5.2e-05 | 0.00106 | 0.32156 | 1.3577 | 2.5482 | 101.21 | 20.82 | 17.765 | 80.3125 | 99.5108 | 77.4421 | 0.004322 | 1.007 | 0.024635 | 24.4643 | 0.013568 | 0.15844 | 0.005391 | 0.39865 | 0.39865 | 0.016533
2019-03-19 15:48 | 1553024872.766 | 343 | 0.20842 | 0.000884 | 3.2e-05 | 100.1538 | 100.0109 | 122.6966 | 21.2612 | 0.002469 | 0.002279 | 248.432 | 0.26572 | 8.4025 | 0.003143 | 0.004805 | 0.004642 | 5.5e-05 | 0.001055 | 0.31365 | 1.3217 | 2.5481 | 101.21 | 20.82 | 17.773 | 81.2263 | 97.4456 | 76.6562 | 0.004324 | 1.0093 | 0.024637 | 24.4114 | 0.012775 | 0.15808 | 0.005334 | 0.39449 | 0.39449 | 0.016669
2019-03-19 15:48 | 1553024873.7661 | 344 | 0.2084 | 0.00088 | 3.1e-05 | 100.1325 | 99.9365 | 122.5903 | 21.2612 | 0.002485 | 0.002289 | 248.4116 | 0.26707 | 8.402 | 0.003181 | 0.004642 | 0.004521 | 5.4e-05 | 0.001057 | 0.31598 | 1.3113 | 2.5481 | 101.21 | 20.82 | 17.776 | 81.5929 | 98.3581 | 76.8878 | 0.004325 | 1.0086 | 0.024619 | 24.4089 | 0.017521 | 0.15888 | 0.005458 | 0.40337 | 0.40337 | 0.016386
2019-03-19 15:48 | 1553024874.7672 | 345 | 0.2084 | 0.000879 | 3.1e-05 | 100.1505 | 99.8098 | 122.7013 | 21.2612 | 0.002455 | 0.002396 | 248.3882 | 0.26573 | 8.4018 | 0.003194 | 0.004979 | 0.004936 | 5.3e-05 | 0.001061 | 0.35153 | 1.3319 | 2.548 | 101.21 | 20.82 | 17.777 | 80.8716 | 98.6756 | 79.4556 | 0.004325 | 1.0012 | 0.02454 | 24.5094 | 0.030478 | 0.15809 | 0.005477 | 0.40348 | 0.40348 | 0.016299
2019-03-19 15:48 | 1553024875.7662 | 346 | 0.20843 | 0.000876 | 3.2e-05 | 100.1523 | 100.0037 | 122.5666 | 21.2612 | 0.002492 | 0.00232 | 248.3646 | 0.26568 | 8.4006 | 0.003257 | 0.005128 | 0.005012 | 5.2e-05 | 0.00106 | 0.29048 | 1.2862 | 2.5481 | 101.21 | 20.82 | 17.776 | 81.7731 | 100.1982 | 77.6251 | 0.004325 | 1.0065 | 0.02459 | 24.4322 | 0.013284 | 0.15805 | 0.005309 | 0.39186 | 0.39186 | 0.016779
2019-03-19 15:48 | 1553024876.7666 | 347 | 0.20837 | 0.000879 | 3.2e-05 | 100.1373 | 99.9208 | 122.6076 | 21.2612 | 0.002486 | 0.002341 | 248.3392 | 0.26699 | 8.4021 | 0.003157 | 0.005308 | 0.005269 | 5.4e-05 | 0.001069 | 0.31166 | 1.281 | 2.5481 | 101.21 | 20.82 | 17.774 | 81.6349 | 97.7983 | 78.1302 | 0.004324 | 1.005 | 0.024577 | 24.4539 | 0.01936 | 0.15883 | 0.005651 | 0.41696 | 0.41696 | 0.015847
2019-03-19 15:48 | 1553024877.7663 | 348 | 0.20836 | 0.00087 | 3.3e-05 | 100.1943 | 99.9554 | 122.4389 | 21.2612 | 0.002515 | 0.002296 | 248.3218 | 0.26545 | 8.4023 | 0.003165 | 0.004355 | 0.004184 | 5.3e-05 | 0.001072 | 0.31382 | 1.3334 | 2.5482 | 101.21 | 20.82 | 17.768 | 82.3284 | 97.9865 | 77.0562 | 0.004323 | 1.0081 | 0.024598 | 24.3997 | 0.021356 | 0.15792 | 0.00573 | 0.42315 | 0.42315 | 0.015525
2019-03-19 15:48 | 1553024878.7654 | 349 | 0.20837 | 0.000865 | 3.3e-05 | 100.2246 | 100.203 | 122.1844 | 21.2612 | 0.002468 | 0.002259 | 248.3102 | 0.26543 | 8.4048 | 0.003255 | 0.004942 | 0.004881 | 5.3e-05 | 0.001078 | 0.34379 | 1.3109 | 2.5481 | 101.21 | 20.82 | 17.763 | 81.1999 | 100.1612 | 76.1644 | 0.004322 | 1.0107 | 0.024603 | 24.3433 | 0.00193 | 0.15791 | 0.005683 | 0.41971 | 0.41971 | 0.015651
2019-03-19 15:48 | 1553024879.7655 | 350 | 0.20836 | 0.00087 | 3.3e-05 | 100.2663 | 99.9509 | 122.2478 | 21.2612 | 0.00244 | 0.002306 | 248.2606 | 0.26746 | 8.4016 | 0.00323 | 0.004937 | 0.004802 | 5.3e-05 | 0.00108 | 0.29878 | 1.3506 | 2.5481 | 101.21 | 20.82 | 17.758 | 80.5315 | 99.5439 | 77.3087 | 0.00432 | 1.0074 | 0.024569 | 24.3894 | 0.028182 | 0.15911 | 0.005691 | 0.41971 | 0.41971 | 0.01577
2019-03-19 15:48 | 1553024880.7668 | 351 | 0.20838 | 0.00087 | 3.4e-05 | 100.2579 | 100.1674 | 122.5623 | 21.2612 | 0.002469 | 0.002311 | 248.2493 | 0.26734 | 8.402 | 0.003244 | 0.00458 | 0.004466 | 5.6e-05 | 0.001077 | 0.30934 | 1.3178 | 2.5481 | 101.21 | 20.82 | 17.743 | 81.2186 | 99.9022 | 77.4161 | 0.004317 | 1.0071 | 0.024597 | 24.4245 | 0.008083 | 0.15904 | 0.005587 | 0.4124 | 0.4124 | 0.016043
2019-03-19 15:48 | 1553024881.7656 | 352 | 0.20839 | 0.000864 | 3.4e-05 | 100.2831 | 100.1146 | 122.6224 | 21.2612 | 0.002422 | 0.00233 | 248.223 | 0.26645 | 8.4018 | 0.003262 | 0.005124 | 0.00495 | 5.3e-05 | 0.001067 | 0.31492 | 1.3102 | 2.5482 | 101.21 | 20.82 | 17.735 | 80.0948 | 100.3354 | 77.8793 | 0.004315 | 1.0057 | 0.024587 | 24.4467 | 0.015043 | 0.15851 | 0.005523 | 0.40754 | 0.40754 | 0.01618
2019-03-19 15:48 | 1553024882.7666 | 353 | 0.20838 | 0.000864 | 3.5e-05 | 100.2786 | 99.9108 | 122.8684 | 21.2612 | 0.002434 | 0.002265 | 248.2107 | 0.2662 | 8.4062 | 0.00322 | 0.004838 | 0.004817 | 5.4e-05 | 0.00106 | 0.32764 | 1.313 | 2.5482 | 101.21 | 20.82 | 17.716 | 80.3765 | 99.3154 | 76.3054 | 0.00431 | 1.0103 | 0.024667 | 24.4162 | 0.032866 | 0.15836 | 0.005612 | 0.41542 | 0.41542 | 0.015858
2019-03-19 15:48 | 1553024883.767 | 354 | 0.20843 | 0.000868 | 3.5e-05 | 100.348 | 100.1256 | 123.0839 | 21.2612 | 0.002454 | 0.002282 | 248.184 | 0.26684 | 8.403 | 0.003278 | 0.004694 | 0.004538 | 5.5e-05 | 0.001058 | 0.32637 | 1.3351 | 2.5482 | 101.21 | 20.82 | 17.707 | 80.8622 | 100.7098 | 76.7186 | 0.004308 | 1.0091 | 0.024674 | 24.4521 | 0.019843 | 0.15874 | 0.005299 | 0.39219 | 0.39219 | 0.016838
2019-03-19 15:48 | 1553024884.7657 | 355 | 0.20839 | 0.000859 | 3.5e-05 | 100.3342 | 100.2975 | 123.1414 | 21.2612 | 0.00242 | 0.002281 | 248.1652 | 0.26631 | 8.4009 | 0.003252 | 0.004575 | 0.004467 | 5.6e-05 | 0.00106 | 0.30702 | 1.3127 | 2.5481 | 101.21 | 20.82 | 17.697 | 80.0371 | 100.086 | 76.6975 | 0.004306 | 1.0091 | 0.024681 | 24.4571 | 0.003267 | 0.15843 | 0.005559 | 0.41166 | 0.41166 | 0.01601
2019-03-19 15:48 | 1553024885.7668 | 356 | 0.20841 | 0.000862 | 3.6e-05 | 100.3549 | 100.3215 | 123.1085 | 21.2612 | 0.002476 | 0.00228 | 248.1349 | 0.26696 | 8.4021 | 0.003177 | 0.00522 | 0.005097 | 5.6e-05 | 0.001059 | 0.3314 | 1.3242 | 2.548 | 101.21 | 20.82 | 17.679 | 81.3768 | 98.2711 | 76.6843 | 0.004301 | 1.0092 | 0.024678 | 24.4533 | 0.002975 | 0.15882 | 0.005442 | 0.4029 | 0.4029 | 0.016398
2019-03-19 15:48 | 1553024886.7658 | 357 | 0.20841 | 0.000859 | 3.6e-05 | 100.4038 | 100.2119 | 123.3264 | 21.2612 | 0.002424 | 0.002286 | 248.0931 | 0.26614 | 8.4036 | 0.003298 | 0.004727 | 0.004634 | 5.2e-05 | 0.001055 | 0.30502 | 1.3056 | 2.5481 | 101.21 | 20.82 | 17.671 | 80.127 | 101.2082 | 76.8183 | 0.004299 | 1.0088 | 0.024695 | 24.4796 | 0.01711 | 0.15832 | 0.005459 | 0.40442 | 0.40442 | 0.016286
2019-03-19 15:48 | 1553024887.7662 | 358 | 0.2084 | 0.000861 | 3.6e-05 | 100.3897 | 100.2054 | 123.4998 | 21.2612 | 0.002425 | 0.002289 | 248.0604 | 0.26712 | 8.4068 | 0.003309 | 0.005288 | 0.005126 | 5.5e-05 | 0.001053 | 0.33931 | 1.3415 | 2.5481 | 101.21 | 20.82 | 17.663 | 80.1698 | 101.4642 | 76.8947 | 0.004297 | 1.0086 | 0.02471 | 24.4995 | 0.016435 | 0.15891 | 0.005477 | 0.40597 | 0.40597 | 0.016284
2019-03-19 15:48 | 1553024888.766 | 359 | 0.20843 | 0.000861 | 3.6e-05 | 100.417 | 100.3194 | 123.5019 | 21.2612 | 0.002468 | 0.002263 | 248.0305 | 0.26683 | 8.4022 | 0.003348 | 0.004852 | 0.004804 | 5.5e-05 | 0.001055 | 0.29336 | 1.3422 | 2.5481 | 101.21 | 20.82 | 17.647 | 81.1828 | 102.3956 | 76.261 | 0.004293 | 1.0104 | 0.024732 | 24.4776 | 0.008693 | 0.15874 | 0.005283 | 0.39191 | 0.39191 | 0.016849
2019-03-19 15:48 | 1553024889.767 | 360 | 0.20849 | 0.000854 | 3.6e-05 | 100.4248 | 100.1992 | 123.3754 | 21.2612 | 0.002509 | 0.002259 | 248.0204 | 0.26679 | 8.4029 | 0.003367 | 0.004859 | 0.004755 | 5.4e-05 | 0.001062 | 0.32604 | 1.3278 | 2.548 | 101.21 | 20.82 | 17.647 | 82.1693 | 102.8559 | 76.1745 | 0.004293 | 1.0107 | 0.024723 | 24.462 | 0.020113 | 0.15871 | 0.004949 | 0.36685 | 0.36685 | 0.017997
2019-03-19 15:48 | 1553024890.7666 | 361 | 0.20845 | 0.000855 | 3.7e-05 | 100.4686 | 100.403 | 123.2956 | 21.2612 | 0.002504 | 0.002297 | 248.0031 | 0.26537 | 8.4042 | 0.003329 | 0.004626 | 0.004548 | 5.3e-05 | 0.001059 | 0.31045 | 1.3278 | 2.5481 | 101.21 | 20.82 | 17.637 | 82.057 | 101.9432 | 77.0934 | 0.004291 | 1.008 | 0.024682 | 24.4862 | 0.005842 | 0.15787 | 0.005188 | 0.38397 | 0.38397 | 0.017104
2019-03-19 15:48 | 1553024891.7674 | 362 | 0.20844 | 0.000856 | 3.7e-05 | 100.4887 | 100.4167 | 123.6601 | 21.2612 | 0.002415 | 0.002322 | 247.9795 | 0.26655 | 8.4019 | 0.003282 | 0.00465 | 0.00445 | 5.5e-05 | 0.001061 | 0.31382 | 1.3954 | 2.548 | 101.21 | 20.82 | 17.633 | 79.9163 | 100.8189 | 77.6773 | 0.00429 | 1.0063 | 0.024698 | 24.5428 | 0.006411 | 0.15857 | 0.005245 | 0.38844 | 0.38844 | 0.016982
2019-03-19 15:48 | 1553024892.7661 | 363 | 0.20838 | 0.000855 | 3.8e-05 | 100.5223 | 100.4085 | 123.5604 | 21.2612 | 0.002496 | 0.002298 | 247.936 | 0.26516 | 8.4038 | 0.003309 | 0.005066 | 0.004968 | 5.5e-05 | 0.001062 | 0.33229 | 1.3832 | 2.548 | 101.21 | 20.82 | 17.629 | 81.8604 | 101.4618 | 77.0958 | 0.004289 | 1.008 | 0.024709 | 24.5126 | 0.010135 | 0.15774 | 0.005639 | 0.41795 | 0.41795 | 0.015701
2019-03-19 15:48 | 1553024893.7652 | 364 | 0.20845 | 0.000847 | 3.8e-05 | 100.5181 | 100.2955 | 123.4827 | 21.2612 | 0.002452 | 0.002286 | 247.9106 | 0.26636 | 8.4033 | 0.003237 | 0.005366 | 0.005148 | 5.4e-05 | 0.001064 | 0.29607 | 1.3792 | 2.5481 | 101.21 | 20.82 | 17.627 | 80.8069 | 99.7226 | 76.8191 | 0.004289 | 1.0088 | 0.024711 | 24.4952 | 0.019835 | 0.15846 | 0.005206 | 0.38573 | 0.38573 | 0.017089
2019-03-19 15:48 | 1553024894.7663 | 365 | 0.20847 | 0.000849 | 3.8e-05 | 100.6156 | 100.4123 | 123.2775 | 21.2612 | 0.002428 | 0.002291 | 247.8657 | 0.26574 | 8.4027 | 0.003291 | 0.004875 | 0.004775 | 5.3e-05 | 0.001063 | 0.30397 | 1.3741 | 2.548 | 101.21 | 20.82 | 17.625 | 80.2267 | 101.0175 | 76.943 | 0.004288 | 1.0084 | 0.024686 | 24.4792 | 0.0181 | 0.15809 | 0.005081 | 0.37604 | 0.37604 | 0.017489
2019-03-19 15:48 | 1553024895.7657 | 366 | 0.20841 | 0.000846 | 3.8e-05 | 100.6406 | 100.6249 | 123.1553 | 21.2612 | 0.002475 | 0.00226 | 247.8614 | 0.26585 | 8.4021 | 0.003287 | 0.005116 | 0.00501 | 5.4e-05 | 0.001065 | 0.3449 | 1.3427 | 2.5482 | 101.21 | 20.82 | 17.621 | 81.3611 | 100.9319 | 76.1831 | 0.004287 | 1.0106 | 0.0247 | 24.4404 | 0.001388 | 0.15815 | 0.005467 | 0.40497 | 0.40497 | 0.016246
2019-03-19 15:48 | 1553024896.7654 | 367 | 0.20844 | 0.000845 | 3.8e-05 | 100.6877 | 100.516 | 123.0055 | 21.2612 | 0.002511 | 0.002272 | 247.8692 | 0.26599 | 8.4029 | 0.003371 | 0.004447 | 0.004293 | 5.4e-05 | 0.001064 | 0.31708 | 1.2945 | 2.5481 | 101.21 | 20.82 | 17.619 | 82.2294 | 102.9532 | 76.4801 | 0.004287 | 1.0098 | 0.024675 | 24.436 | 0.015261 | 0.15823 | 0.005263 | 0.38937 | 0.38937 | 0.016906
2019-03-19 15:48 | 1553024897.7664 | 368 | 0.20842 | 0.000841 | 3.9e-05 | 100.6712 | 100.5222 | 122.9719 | 21.2612 | 0.002406 | 0.002263 | 247.8468 | 0.26582 | 8.4024 | 0.003374 | 0.004573 | 0.004516 | 5.1e-05 | 0.001062 | 0.32947 | 1.3064 | 2.5481 | 101.21 | 20.82 | 17.615 | 79.7043 | 103.0288 | 76.2703 | 0.004286 | 1.0104 | 0.024679 | 24.4253 | 0.013247 | 0.15813 | 0.005413 | 0.40056 | 0.40056 | 0.016423
2019-03-19 15:48 | 1553024898.7668 | 369 | 0.20842 | 0.000842 | 3.9e-05 | 100.6841 | 100.6434 | 123.2649 | 21.2612 | 0.002462 | 0.002291 | 247.8165 | 0.26556 | 8.4033 | 0.003378 | 0.005331 | 0.005265 | 5.1e-05 | 0.001065 | 0.29878 | 1.2994 | 2.5479 | 101.21 | 20.83 | 17.614 | 81.0376 | 103.1237 | 76.9416 | 0.004288 | 1.0084 | 0.024684 | 24.4779 | 0.003611 | 0.15798 | 0.005404 | 0.39998 | 0.39998 | 0.016431
2019-03-19 15:48 | 1553024899.7655 | 370 | 0.20843 | 0.000841 | 4e-05 | 100.7012 | 100.5171 | 123.4081 | 21.2612 | 0.002395 | 0.002284 | 247.8012 | 0.26599 | 8.4016 | 0.003417 | 0.005205 | 0.005099 | 5.1e-05 | 0.001064 | 0.30342 | 1.2875 | 2.548 | 101.21 | 20.83 | 17.613 | 79.4292 | 104.0881 | 76.7662 | 0.004288 | 1.0089 | 0.024705 | 24.4859 | 0.01637 | 0.15823 | 0.005354 | 0.39655 | 0.39655 | 0.0166
2019-03-19 15:48 | 1553024900.766 | 371 | 0.20844 | 0.000844 | 4e-05 | 100.7585 | 100.6246 | 123.4174 | 21.2612 | 0.002513 | 0.002235 | 247.7863 | 0.26518 | 8.4014 | 0.003413 | 0.004882 | 0.004778 | 5.1e-05 | 0.001067 | 0.29607 | 1.3133 | 2.5482 | 101.21 | 20.83 | 17.612 | 82.2645 | 103.9893 | 75.5844 | 0.004288 | 1.0124 | 0.024748 | 24.4455 | 0.011895 | 0.15776 | 0.005283 | 0.39193 | 0.39193 | 0.016744
2019-03-19 15:48 | 1553024901.7665 | 372 | 0.20842 | 0.000838 | 4e-05 | 100.7993 | 100.6893 | 122.8391 | 21.2612 | 0.002388 | 0.002277 | 247.7538 | 0.26607 | 8.4037 | 0.003413 | 0.00539 | 0.005277 | 5e-05 | 0.001065 | 0.31564 | 1.288 | 2.5481 | 101.21 | 20.83 | 17.609 | 79.2732 | 103.9837 | 76.6141 | 0.004287 | 1.0094 | 0.024653 | 24.4241 | 0.009765 | 0.15829 | 0.005391 | 0.39845 | 0.39845 | 0.016526
2019-03-19 15:48 | 1553024902.7668 | 373 | 0.20841 | 0.000837 | 4.1e-05 | 100.7891 | 100.8154 | 122.5882 | 21.2612 | 0.002428 | 0.002284 | 247.7556 | 0.26535 | 8.3997 | 0.003342 | 0.005181 | 0.005108 | 4.8e-05 | 0.00107 | 0.32444 | 1.2851 | 2.548 | 101.21 | 20.83 | 17.604 | 80.2306 | 102.2668 | 76.7621 | 0.004286 | 1.009 | 0.024623 | 24.4043 | -0.002329 | 0.15786 | 0.005463 | 0.4032 | 0.4032 | 0.016287
2019-03-19 15:48 | 1553024903.7658 | 374 | 0.20845 | 0.000837 | 4.1e-05 | 100.7984 | 100.5486 | 122.6666 | 21.2612 | 0.002479 | 0.00234 | 247.7238 | 0.26605 | 8.4027 | 0.003337 | 0.004935 | 0.004773 | 5.1e-05 | 0.001068 | 0.31581 | 1.331 | 2.5481 | 101.21 | 20.83 | 17.601 | 81.458 | 102.1308 | 78.1179 | 0.004285 | 1.0051 | 0.024583 | 24.4594 | 0.022199 | 0.15827 | 0.005193 | 0.38256 | 0.38256 | 0.017211
2019-03-19 15:48 | 1553024904.7661 | 375 | 0.20868 | 0.000833 | 4.2e-05 | 100.8656 | 100.8267 | 123.0189 | 21.2612 | 0.002474 | 0.00231 | 247.7218 | 0.26632 | 8.4033 | 0.003343 | 0.004961 | 0.004922 | 5.2e-05 | 0.001068 | 0.30879 | 1.3462 | 2.548 | 101.21 | 20.83 | 17.598 | 81.3364 | 102.2959 | 77.3966 | 0.004284 | 1.0071 | 0.024644 | 24.4693 | 0.003457 | 0.15844 | 0.00382 | 0.28145 | 0.28145 | 0.023418
2019-03-19 15:48 | 1553024905.7661 | 376 | 0.20832 | 0.000828 | 4.3e-05 | 100.8821 | 100.7462 | 123.3759 | 21.2612 | 0.002434 | 0.002298 | 247.7027 | 0.26679 | 8.4027 | 0.003343 | 0.005062 | 0.004912 | 5.1e-05 | 0.001071 | 0.31022 | 1.3244 | 2.5481 | 101.21 | 20.83 | 17.595 | 80.3695 | 102.2959 | 77.0973 | 0.004283 | 1.008 | 0.02469 | 24.4943 | 0.012063 | 0.15871 | 0.00602 | 0.44567 | 0.44567 | 0.014815
2019-03-19 15:48 | 1553024906.766 | 377 | 0.20843 | 0.000827 | 4.3e-05 | 100.8524 | 100.5773 | 123.3646 | 21.2612 | 0.002457 | 0.002284 | 247.6771 | 0.26731 | 8.4037 | 0.003377 | 0.004804 | 0.004732 | 5.2e-05 | 0.001074 | 0.29999 | 1.3222 | 2.5481 | 101.21 | 20.83 | 17.594 | 80.9175 | 103.1153 | 76.7809 | 0.004283 | 1.0089 | 0.0247 | 24.4821 | 0.024432 | 0.15902 | 0.00532 | 0.39372 | 0.39372 | 0.016802
2019-03-19 15:48 | 1553024907.7659 | 378 | 0.20843 | 0.000826 | 4.3e-05 | 100.8887 | 100.7164 | 123.5367 | 21.2612 | 0.00248 | 0.002312 | 247.6604 | 0.2665 | 8.4025 | 0.003339 | 0.004741 | 0.004583 | 5.1e-05 | 0.001076 | 0.28783 | 1.3224 | 2.5481 | 101.21 | 20.83 | 17.596 | 81.4931 | 102.193 | 77.438 | 0.004284 | 1.007 | 0.024694 | 24.5222 | 0.015294 | 0.15854 | 0.005364 | 0.3969 | 0.3969 | 0.016617
2019-03-19 15:48 | 1553024908.7673 | 379 | 0.2085 | 0.000821 | 4.4e-05 | 100.9103 | 100.816 | 123.3566 | 21.2612 | 0.002395 | 0.002287 | 247.6528 | 0.26727 | 8.4021 | 0.003388 | 0.004579 | 0.004441 | 4.9e-05 | 0.001073 | 0.31094 | 1.3131 | 2.5482 | 101.21 | 20.83 | 17.596 | 79.4469 | 103.388 | 76.8409 | 0.004284 | 1.0087 | 0.024697 | 24.4834 | 0.008363 | 0.159 | 0.004904 | 0.36262 | 0.36262 | 0.018241
2019-03-19 15:48 | 1553024909.7661 | 380 | 0.20847 | 0.000822 | 4.4e-05 | 100.9626 | 100.7832 | 123.5318 | 21.2612 | 0.002448 | 0.002261 | 247.6165 | 0.26698 | 8.4032 | 0.003338 | 0.004476 | 0.004411 | 4.7e-05 | 0.001068 | 0.3147 | 1.3153 | 2.5481 | 101.21 | 20.83 | 17.597 | 80.7156 | 102.1702 | 76.2127 | 0.004284 | 1.0105 | 0.024737 | 24.4788 | 0.015908 | 0.15883 | 0.005132 | 0.3802 | 0.3802 | 0.017378
2019-03-19 15:48 | 1553024910.7651 | 381 | 0.2085 | 0.000816 | 4.4e-05 | 100.9296 | 100.9208 | 123.4327 | 21.2612 | 0.002426 | 0.002208 | 247.5848 | 0.2661 | 8.4053 | 0.003328 | 0.005303 | 0.005149 | 5e-05 | 0.001069 | 0.34147 | 1.3154 | 2.548 | 101.21 | 20.82 | 17.598 | 80.1893 | 101.9173 | 74.9381 | 0.004281 | 1.0142 | 0.024772 | 24.4243 | 0.000772 | 0.15831 | 0.004947 | 0.36694 | 0.36694 | 0.017947
2019-03-19 15:48 | 1553024911.7665 | 382 | 0.20843 | 0.000821 | 4.5e-05 | 101.0169 | 100.9924 | 123.5529 | 21.2612 | 0.002454 | 0.002252 | 247.575 | 0.26594 | 8.4031 | 0.003247 | 0.005186 | 0.005067 | 5e-05 | 0.001061 | 0.3063 | 1.2806 | 2.548 | 101.21 | 20.82 | 17.598 | 80.8528 | 99.955 | 76.0123 | 0.004281 | 1.0111 | 0.024746 | 24.4739 | 0.002166 | 0.15821 | 0.00537 | 0.39808 | 0.39808 | 0.016533
2019-03-19 15:49 | 1553024912.7652 | 383 | 0.20855 | 0.000814 | 4.4e-05 | 101.0655 | 100.7489 | 123.6883 | 21.2612 | 0.002455 | 0.002252 | 247.5631 | 0.26603 | 8.406 | 0.003328 | 0.004584 | 0.004558 | 5.1e-05 | 0.001056 | 0.30773 | 1.3026 | 2.5481 | 101.21 | 20.82 | 17.598 | 80.874 | 101.9282 | 75.9967 | 0.004281 | 1.0112 | 0.02476 | 24.4867 | 0.028063 | 0.15826 | 0.004614 | 0.34187 | 0.34187 | 0.019258
2019-03-19 15:49 | 1553024913.7653 | 384 | 0.20847 | 0.000807 | 4.5e-05 | 101.0217 | 100.877 | 123.75 | 21.2612 | 0.002451 | 0.002276 | 247.5464 | 0.26639 | 8.405 | 0.003312 | 0.004653 | 0.004559 | 5.1e-05 | 0.00105 | 0.29927 | 1.318 | 2.548 | 101.21 | 20.82 | 17.597 | 80.7874 | 101.5387 | 76.5728 | 0.004281 | 1.0095 | 0.024746 | 24.5131 | 0.012819 | 0.15848 | 0.005152 | 0.38178 | 0.38178 | 0.017268
2019-03-19 15:49 | 1553024914.7663 | 385 | 0.20855 | 0.000799 | 4.6e-05 | 101.0799 | 101.0414 | 123.4367 | 21.2612 | 0.002439 | 0.002219 | 247.4956 | 0.26603 | 8.4041 | 0.003293 | 0.005125 | 0.004991 | 5e-05 | 0.001054 | 0.33196 | 1.3175 | 2.5482 | 101.21 | 20.82 | 17.595 | 80.4911 | 101.0819 | 75.2188 | 0.004281 | 1.0134 | 0.024763 | 24.4346 | 0.003408 | 0.15826 | 0.004682 | 0.34695 | 0.34695 | 0.018976
2019-03-19 15:49 | 1553024915.7657 | 386 | 0.20852 | 0.000798 | 4.6e-05 | 101.0844 | 101.0774 | 123.1236 | 21.2612 | 0.002412 | 0.002234 | 247.4488 | 0.26638 | 8.4028 | 0.00334 | 0.004489 | 0.004404 | 5e-05 | 0.001056 | 0.28313 | 1.3535 | 2.5481 | 101.21 | 20.83 | 17.588 | 79.8533 | 102.2077 | 75.5618 | 0.004282 | 1.0124 | 0.024719 | 24.4156 | 0.000623 | 0.15847 | 0.004843 | 0.35829 | 0.35829 | 0.018399
2019-03-19 15:49 | 1553024916.7664 | 387 | 0.2085 | 0.000798 | 4.7e-05 | 101.1264 | 101.2555 | 123.1081 | 21.2612 | 0.002425 | 0.002284 | 247.4033 | 0.2674 | 8.4013 | 0.003198 | 0.005106 | 0.004937 | 5.2e-05 | 0.001062 | 0.33539 | 1.3757 | 2.5481 | 101.21 | 20.83 | 17.584 | 80.1559 | 98.78 | 76.7621 | 0.004281 | 1.009 | 0.024675 | 24.456 | -0.011406 | 0.15907 | 0.004945 | 0.3652 | 0.3652 | 0.01812
2019-03-19 15:49 | 1553024917.7666 | 388 | 0.20853 | 0.000792 | 4.8e-05 | 101.1522 | 100.9989 | 123.0647 | 21.2612 | 0.002431 | 0.00224 | 247.3912 | 0.26705 | 8.4045 | 0.00326 | 0.004323 | 0.004275 | 5.4e-05 | 0.001055 | 0.30453 | 1.3467 | 2.548 | 101.21 | 20.83 | 17.572 | 80.3046 | 100.2749 | 75.7084 | 0.004278 | 1.012 | 0.024708 | 24.4148 | 0.013566 | 0.15887 | 0.004755 | 0.35148 | 0.35148 | 0.018803
2019-03-19 15:49 | 1553024918.7666 | 389 | 0.20853 | 0.000793 | 4.8e-05 | 101.1981 | 101.0298 | 123.1733 | 21.2612 | 0.002421 | 0.002289 | 247.3923 | 0.26625 | 8.4011 | 0.003339 | 0.005389 | 0.00532 | 5.3e-05 | 0.001059 | 0.33356 | 1.3327 | 2.5481 | 101.21 | 20.83 | 17.564 | 80.0746 | 102.1787 | 76.9017 | 0.004276 | 1.0086 | 0.024677 | 24.4674 | 0.014896 | 0.15839 | 0.004775 | 0.35246 | 0.35246 | 0.018695
2019-03-19 15:49 | 1553024919.7656 | 390 | 0.20853 | 0.00079 | 4.8e-05 | 101.2096 | 101.1904 | 122.9203 | 21.2612 | 0.002431 | 0.002211 | 247.3812 | 0.26677 | 8.4021 | 0.00337 | 0.004672 | 0.004523 | 5.3e-05 | 0.001056 | 0.32493 | 1.3773 | 2.5481 | 101.21 | 20.82 | 17.555 | 80.3125 | 102.9423 | 75.0246 | 0.004271 | 1.014 | 0.024718 | 24.3766 | 0.001694 | 0.1587 | 0.004767 | 0.35249 | 0.35249 | 0.01873
2019-03-19 15:49 | 1553024920.7661 | 391 | 0.2085 | 0.000785 | 4.8e-05 | 101.2312 | 101.2144 | 123.0952 | 21.2612 | 0.002478 | 0.002309 | 247.3539 | 0.26684 | 8.4032 | 0.003324 | 0.004605 | 0.004533 | 5.2e-05 | 0.001056 | 0.31492 | 1.3833 | 2.5479 | 101.21 | 20.83 | 17.538 | 81.4244 | 101.8214 | 77.3709 | 0.00427 | 1.0072 | 0.024652 | 24.476 | 0.001484 | 0.15874 | 0.004946 | 0.36488 | 0.36488 | 0.018098
2019-03-19 15:49 | 1553024921.7657 | 392 | 0.20853 | 0.000781 | 4.9e-05 | 101.1966 | 101.3133 | 123.5069 | 21.2612 | 0.002458 | 0.002237 | 247.3284 | 0.26682 | 8.4012 | 0.00336 | 0.005348 | 0.005253 | 5e-05 | 0.001052 | 0.32709 | 1.3076 | 2.5481 | 101.21 | 20.83 | 17.528 | 80.9456 | 102.7082 | 75.65 | 0.004267 | 1.0122 | 0.024754 | 24.4566 | -0.010308 | 0.15873 | 0.004765 | 0.35282 | 0.35282 | 0.018715
2019-03-19 15:49 | 1553024922.7668 | 393 | 0.20855 | 0.00078 | 4.9e-05 | 101.1873 | 101.2236 | 123.566 | 21.2612 | 0.002381 | 0.002273 | 247.2986 | 0.26635 | 8.4045 | 0.003318 | 0.004617 | 0.004537 | 5.1e-05 | 0.001055 | 0.30646 | 1.3194 | 2.5481 | 101.21 | 20.83 | 17.517 | 79.1033 | 101.6871 | 76.5089 | 0.004264 | 1.0097 | 0.02473 | 24.4926 | -0.003205 | 0.15845 | 0.004672 | 0.34557 | 0.34557 | 0.019075
2019-03-19 15:49 | 1553024923.7668 | 394 | 0.20851 | 0.000783 | 5e-05 | 101.2723 | 101.2695 | 123.2907 | 21.2612 | 0.002362 | 0.002293 | 247.2807 | 0.26723 | 8.4035 | 0.003269 | 0.00506 | 0.004975 | 5.5e-05 | 0.001059 | 0.31924 | 1.3132 | 2.5481 | 101.2 | 20.83 | 17.496 | 78.6511 | 100.4948 | 76.993 | 0.00426 | 1.0083 | 0.024685 | 24.4822 | 0.000245 | 0.15898 | 0.004918 | 0.36317 | 0.36317 | 0.01821
2019-03-19 15:49 | 1553024924.7662 | 395 | 0.20849 | 0.000781 | 5e-05 | 101.2405 | 101.0513 | 123.5735 | 21.2612 | 0.002404 | 0.00231 | 247.2973 | 0.26606 | 8.4018 | 0.003298 | 0.004638 | 0.004582 | 5.3e-05 | 0.001056 | 0.32565 | 1.3121 | 2.5481 | 101.2 | 20.83 | 17.484 | 79.6481 | 101.191 | 77.4077 | 0.004257 | 1.0071 | 0.024699 | 24.5248 | 0.016725 | 0.15828 | 0.005023 | 0.37116 | 0.37116 | 0.01774
2019-03-19 15:49 | 1553024925.7663 | 396 | 0.20855 | 0.000777 | 5.1e-05 | 101.2378 | 101.2544 | 123.7803 | 21.2612 | 0.002463 | 0.002275 | 247.2873 | 0.26659 | 8.4036 | 0.003236 | 0.005019 | 0.004853 | 5e-05 | 0.001054 | 0.29551 | 1.3214 | 2.5482 | 101.21 | 20.83 | 17.459 | 81.0775 | 99.7054 | 76.5447 | 0.00425 | 1.0096 | 0.02475 | 24.5151 | -0.001473 | 0.1586 | 0.004666 | 0.34529 | 0.34529 | 0.019108
2019-03-19 15:49 | 1553024926.765 | 397 | 0.20848 | 0.000776 | 5.1e-05 | 101.3011 | 101.0911 | 123.755 | 21.2612 | 0.002452 | 0.0023 | 247.2591 | 0.26633 | 8.4043 | 0.00326 | 0.00459 | 0.004616 | 5.5e-05 | 0.001056 | 0.2995 | 1.325 | 2.548 | 101.21 | 20.83 | 17.446 | 80.8115 | 100.279 | 77.1441 | 0.004247 | 1.0079 | 0.024726 | 24.5336 | 0.018563 | 0.15844 | 0.005079 | 0.37572 | 0.37572 | 0.017543
2019-03-19 15:49 | 1553024927.766 | 398 | 0.20853 | 0.000776 | 5.2e-05 | 101.2756 | 101.0962 | 123.8306 | 21.2612 | 0.002429 | 0.00226 | 247.2341 | 0.26697 | 8.4029 | 0.003273 | 0.004967 | 0.004825 | 5.2e-05 | 0.001054 | 0.30414 | 1.3147 | 2.5481 | 101.21 | 20.83 | 17.433 | 80.2517 | 100.5857 | 76.1862 | 0.004244 | 1.0106 | 0.024768 | 24.5075 | 0.015856 | 0.15882 | 0.004785 | 0.35433 | 0.35433 | 0.018646
2019-03-19 15:49 | 1553024928.7668 | 399 | 0.20856 | 0.000772 | 5.2e-05 | 101.2843 | 101.1589 | 123.9646 | 21.2612 | 0.002437 | 0.002269 | 247.2035 | 0.26627 | 8.4037 | 0.003265 | 0.004283 | 0.004198 | 5.2e-05 | 0.001059 | 0.33179 | 1.3061 | 2.5481 | 101.21 | 20.82 | 17.408 | 80.4519 | 100.3966 | 76.4154 | 0.004235 | 1.01 | 0.024773 | 24.5288 | 0.011081 | 0.15841 | 0.004597 | 0.34029 | 0.34029 | 0.019365
2019-03-19 15:49 | 1553024929.7661 | 400 | 0.20854 | 0.000768 | 5.3e-05 | 101.3884 | 101.3459 | 123.8897 | 21.2612 | 0.002409 | 0.00223 | 247.1767 | 0.26634 | 8.3997 | 0.003215 | 0.004737 | 0.004634 | 5.1e-05 | 0.001055 | 0.33428 | 1.3034 | 2.5481 | 101.21 | 20.82 | 17.396 | 79.7863 | 99.1955 | 75.4722 | 0.004232 | 1.0127 | 0.024799 | 24.4883 | 0.003753 | 0.15845 | 0.004748 | 0.35197 | 0.35197 | 0.018727
2019-03-19 15:49 | 1553024930.7661 | 401 | 0.20856 | 0.000769 | 5.3e-05 | 101.3527 | 101.2637 | 123.7442 | 21.2612 | 0.002445 | 0.002264 | 247.1865 | 0.26587 | 8.4028 | 0.003209 | 0.004986 | 0.004886 | 5.3e-05 | 0.00106 | 0.31868 | 1.3261 | 2.5482 | 101.21 | 20.82 | 17.384 | 80.6291 | 99.0362 | 76.2821 | 0.004229 | 1.0103 | 0.024756 | 24.5023 | 0.00786 | 0.15817 | 0.004644 | 0.34353 | 0.34353 | 0.019153
2019-03-19 15:49 | 1553024931.7652 | 402 | 0.20852 | 0.000768 | 5.3e-05 | 101.3731 | 101.4497 | 123.5203 | 21.2612 | 0.002454 | 0.002253 | 247.1858 | 0.26638 | 8.4028 | 0.003232 | 0.004318 | 0.004202 | 5.2e-05 | 0.001062 | 0.31598 | 1.3481 | 2.5482 | 101.21 | 20.82 | 17.362 | 80.8629 | 99.6002 | 76.0241 | 0.004224 | 1.0111 | 0.024742 | 24.4711 | -0.006751 | 0.15847 | 0.004844 | 0.35828 | 0.35828 | 0.0184
2019-03-19 15:49 | 1553024932.7656 | 403 | 0.20855 | 0.000766 | 5.3e-05 | 101.4172 | 101.1462 | 123.471 | 21.2612 | 0.002444 | 0.002236 | 247.1663 | 0.26634 | 8.4016 | 0.003282 | 0.004534 | 0.004443 | 5.3e-05 | 0.001065 | 0.29856 | 1.346 | 2.548 | 101.21 | 20.83 | 17.35 | 80.6211 | 100.822 | 75.6063 | 0.004224 | 1.0123 | 0.024752 | 24.4515 | 0.023934 | 0.15845 | 0.004685 | 0.34654 | 0.34654 | 0.01902
2019-03-19 15:49 | 1553024933.7664 | 404 | 0.20856 | 0.00076 | 5.4e-05 | 101.4304 | 101.6072 | 123.5418 | 21.2612 | 0.002434 | 0.002312 | 247.159 | 0.2659 | 8.4015 | 0.003231 | 0.00494 | 0.004882 | 5.4e-05 | 0.001069 | 0.33323 | 1.3212 | 2.5481 | 101.2 | 20.82 | 17.333 | 80.3873 | 99.5886 | 77.4521 | 0.004217 | 1.007 | 0.024694 | 24.5232 | -0.015575 | 0.15818 | 0.004653 | 0.34336 | 0.34336 | 0.019165
2019-03-19 15:49 | 1553024934.7654 | 405 | 0.20855 | 0.000759 | 5.4e-05 | 101.3944 | 101.2966 | 123.5247 | 21.2612 | 0.002457 | 0.002328 | 247.1211 | 0.26633 | 8.4036 | 0.003215 | 0.005186 | 0.00499 | 5.4e-05 | 0.001067 | 0.31669 | 1.3165 | 2.5482 | 101.2 | 20.82 | 17.326 | 80.9237 | 99.1791 | 77.82 | 0.004216 | 1.0059 | 0.024679 | 24.5344 | 0.008637 | 0.15844 | 0.004693 | 0.34612 | 0.34612 | 0.019043
2019-03-19 15:49 | 1553024935.7654 | 406 | 0.20856 | 0.000756 | 5.4e-05 | 101.482 | 101.2805 | 123.4196 | 21.2612 | 0.002427 | 0.002201 | 247.1139 | 0.26632 | 8.4017 | 0.003236 | 0.004621 | 0.004579 | 5.4e-05 | 0.001061 | 0.31741 | 1.3689 | 2.5481 | 101.2 | 20.83 | 17.32 | 80.1986 | 99.7015 | 74.787 | 0.004217 | 1.0147 | 0.024776 | 24.4177 | 0.017785 | 0.15843 | 0.004624 | 0.34228 | 0.34228 | 0.019256
2019-03-19 15:49 | 1553024936.7655 | 407 | 0.20854 | 0.000754 | 5.4e-05 | 101.4808 | 101.3609 | 123.6118 | 21.2612 | 0.002465 | 0.002237 | 247.1196 | 0.2671 | 8.4051 | 0.00335 | 0.005507 | 0.005405 | 5.4e-05 | 0.001055 | 0.31542 | 1.3291 | 2.548 | 101.2 | 20.83 | 17.315 | 81.1172 | 102.4506 | 75.6358 | 0.004216 | 1.0122 | 0.024765 | 24.4665 | 0.010574 | 0.1589 | 0.004748 | 0.35136 | 0.35136 | 0.018813
2019-03-19 15:49 | 1553024937.7655 | 408 | 0.20854 | 0.000752 | 5.5e-05 | 101.4643 | 101.2606 | 123.5153 | 21.2612 | 0.002386 | 0.002227 | 247.1012 | 0.26647 | 8.404 | 0.003274 | 0.005081 | 0.005053 | 5.3e-05 | 0.001059 | 0.33892 | 1.3266 | 2.5481 | 101.2 | 20.83 | 17.316 | 79.2163 | 100.6164 | 75.391 | 0.004216 | 1.0129 | 0.024764 | 24.4484 | 0.017979 | 0.15853 | 0.004768 | 0.35279 | 0.35279 | 0.018693
2019-03-19 15:49 | 1553024938.7656 | 409 | 0.20854 | 0.000751 | 5.4e-05 | 101.5286 | 101.4469 | 123.1546 | 21.2612 | 0.002423 | 0.002214 | 247.0712 | 0.26617 | 8.4025 | 0.003277 | 0.0046 | 0.004453 | 5.4e-05 | 0.00106 | 0.32029 | 1.3277 | 2.548 | 101.2 | 20.83 | 17.323 | 80.1027 | 100.6941 | 75.0979 | 0.004218 | 1.0138 | 0.024739 | 24.4024 | 0.007196 | 0.15835 | 0.004744 | 0.35065 | 0.35065 | 0.018786
2019-03-19 15:49 | 1553024939.7657 | 410 | 0.20855 | 0.000752 | 5.4e-05 | 101.516 | 101.4294 | 123.501 | 21.2612 | 0.002376 | 0.002269 | 247.0597 | 0.26688 | 8.4045 | 0.003306 | 0.004272 | 0.004249 | 5.2e-05 | 0.001059 | 0.29662 | 1.3368 | 2.548 | 101.2 | 20.83 | 17.33 | 78.9769 | 101.3841 | 76.4061 | 0.004219 | 1.01 | 0.024727 | 24.4825 | 0.007626 | 0.15877 | 0.004695 | 0.34687 | 0.34687 | 0.019041
2019-03-19 15:49 | 1553024940.7647 | 411 | 0.20854 | 0.000748 | 5.5e-05 | 101.5577 | 101.5264 | 123.4329 | 21.2612 | 0.002366 | 0.002281 | 247.0238 | 0.26599 | 8.4006 | 0.00334 | 0.004905 | 0.00481 | 5.6e-05 | 0.001062 | 0.34418 | 1.3304 | 2.5482 | 101.2 | 20.83 | 17.34 | 78.7314 | 102.2172 | 76.692 | 0.004222 | 1.0092 | 0.02471 | 24.4858 | 0.002757 | 0.15823 | 0.004773 | 0.35243 | 0.35243 | 0.018677
2019-03-19 15:49 | 1553024941.7664 | 412 | 0.20856 | 0.000741 | 5.5e-05 | 101.5481 | 101.4658 | 123.171 | 21.2612 | 0.002435 | 0.002218 | 246.9935 | 0.26586 | 8.4035 | 0.003244 | 0.005001 | 0.004866 | 5.4e-05 | 0.001068 | 0.3402 | 1.3125 | 2.5482 | 101.2 | 20.83 | 17.366 | 80.406 | 99.9001 | 75.1837 | 0.004228 | 1.0135 | 0.024737 | 24.407 | 0.007254 | 0.15816 | 0.004643 | 0.34312 | 0.34312 | 0.019175
2019-03-19 15:49 | 1553024942.7648 | 413 | 0.20858 | 0.00074 | 5.5e-05 | 101.6054 | 101.6343 | 123.6224 | 21.2612 | 0.002477 | 0.002291 | 246.9734 | 0.26681 | 8.4024 | 0.003281 | 0.005659 | 0.005556 | 5.6e-05 | 0.001068 | 0.33539 | 1.3194 | 2.548 | 101.2 | 20.83 | 17.381 | 81.4119 | 100.7773 | 76.9469 | 0.004232 | 1.0084 | 0.02472 | 24.5135 | -0.002541 | 0.15873 | 0.004521 | 0.33381 | 0.33381 | 0.019781
2019-03-19 15:49 | 1553024943.7662 | 414 | 0.20856 | 0.000742 | 5.5e-05 | 101.6549 | 101.576 | 123.8045 | 21.2612 | 0.002342 | 0.002222 | 246.9444 | 0.26553 | 8.4019 | 0.003169 | 0.005001 | 0.004848 | 5.7e-05 | 0.00106 | 0.32372 | 1.3185 | 2.548 | 101.2 | 20.83 | 17.415 | 78.1638 | 98.086 | 75.2726 | 0.00424 | 1.0133 | 0.024798 | 24.4728 | 0.006943 | 0.15796 | 0.00467 | 0.34596 | 0.34596 | 0.018994
2019-03-19 15:49 | 1553024944.765 | 415 | 0.20861 | 0.000737 | 5.5e-05 | 101.6501 | 101.5603 | 123.7867 | 21.2612 | 0.002412 | 0.002257 | 246.9213 | 0.26655 | 8.4052 | 0.003168 | 0.005559 | 0.005545 | 5.5e-05 | 0.001058 | 0.3178 | 1.3202 | 2.5482 | 101.2 | 20.83 | 17.433 | 79.8533 | 98.0421 | 76.1324 | 0.004244 | 1.0108 | 0.024765 | 24.5012 | 0.007908 | 0.15857 | 0.004363 | 0.32255 | 0.32255 | 0.020451
2019-03-19 15:49 | 1553024945.765 | 416 | 0.20857 | 0.000735 | 5.6e-05 | 101.6189 | 101.45 | 124.058 | 21.2612 | 0.002441 | 0.002204 | 246.8991 | 0.26571 | 8.4028 | 0.003232 | 0.004942 | 0.004786 | 5.4e-05 | 0.001052 | 0.30558 | 1.2977 | 2.5481 | 101.2 | 20.83 | 17.451 | 80.544 | 99.6128 | 74.8384 | 0.004249 | 1.0145 | 0.024838 | 24.4826 | 0.014879 | 0.15807 | 0.004625 | 0.34309 | 0.34309 | 0.019166
2019-03-19 15:49 | 1553024946.766 | 417 | 0.20861 | 0.000731 | 5.7e-05 | 101.6522 | 101.9117 | 123.9601 | 21.2612 | 0.002398 | 0.002244 | 246.8752 | 0.26585 | 8.4021 | 0.003235 | 0.005017 | 0.004916 | 5.4e-05 | 0.00105 | 0.32925 | 1.3096 | 2.5481 | 101.2 | 20.83 | 17.482 | 79.5164 | 99.6624 | 75.8199 | 0.004256 | 1.0117 | 0.024794 | 24.5074 | -0.022802 | 0.15815 | 0.00434 | 0.32114 | 0.32114 | 0.020487
2019-03-19 15:49 | 1553024947.7651 | 418 | 0.20861 | 0.00073 | 5.7e-05 | 101.6357 | 101.5784 | 123.8331 | 21.2612 | 0.002379 | 0.002253 | 246.8436 | 0.26611 | 8.405 | 0.003305 | 0.004833 | 0.004725 | 5.2e-05 | 0.001051 | 0.32925 | 1.3717 | 2.5482 | 101.2 | 20.83 | 17.495 | 79.0495 | 101.3778 | 76.0279 | 0.004259 | 1.0111 | 0.024774 | 24.5022 | 0.005041 | 0.15831 | 0.004355 | 0.32197 | 0.32197 | 0.020454
2019-03-19 15:49 | 1553024948.7652 | 419 | 0.20859 | 0.000728 | 5.7e-05 | 101.6861 | 101.6665 | 124.147 | 21.2612 | 0.002464 | 0.002293 | 246.8216 | 0.26515 | 8.4023 | 0.003279 | 0.004766 | 0.004721 | 5.4e-05 | 0.00105 | 0.31758 | 1.3283 | 2.5481 | 101.2 | 20.82 | 17.508 | 81.0929 | 100.7458 | 76.9921 | 0.00426 | 1.0083 | 0.024771 | 24.567 | 0.00173 | 0.15774 | 0.004475 | 0.33092 | 0.33092 | 0.019829
2019-03-19 15:49 | 1553024949.7652 | 420 | 0.20861 | 0.000726 | 5.7e-05 | 101.7098 | 101.6966 | 124.0855 | 21.2612 | 0.002416 | 0.002238 | 246.8013 | 0.26499 | 8.4032 | 0.00329 | 0.004811 | 0.004773 | 5.3e-05 | 0.001055 | 0.29391 | 1.3268 | 2.5481 | 101.2 | 20.83 | 17.526 | 79.9319 | 100.9987 | 75.6615 | 0.004267 | 1.0121 | 0.024812 | 24.5143 | 0.001164 | 0.15764 | 0.00438 | 0.32432 | 0.32432 | 0.02022
2019-03-19 15:49 | 1553024950.7663 | 421 | 0.2086 | 0.000725 | 5.8e-05 | 101.6966 | 101.6319 | 123.9099 | 21.2612 | 0.002422 | 0.002216 | 246.7929 | 0.26575 | 8.4009 | 0.003222 | 0.004781 | 0.004675 | 5e-05 | 0.001054 | 0.33948 | 1.3519 | 2.548 | 101.2 | 20.83 | 17.534 | 80.0965 | 99.3524 | 75.1496 | 0.004269 | 1.0136 | 0.024812 | 24.4789 | 0.005698 | 0.1581 | 0.004449 | 0.32948 | 0.32948 | 0.019961
2019-03-19 15:49 | 1553024951.7666 | 422 | 0.20857 | 0.00072 | 5.8e-05 | 101.7864 | 101.5812 | 124.3603 | 21.2612 | 0.002422 | 0.002232 | 246.7622 | 0.2661 | 8.4015 | 0.003222 | 0.004461 | 0.004372 | 5.4e-05 | 0.001051 | 0.30414 | 1.3507 | 2.5482 | 101.2 | 20.82 | 17.543 | 80.0871 | 99.3689 | 75.5198 | 0.004269 | 1.0125 | 0.024844 | 24.5364 | 0.018049 | 0.1583 | 0.004604 | 0.3415 | 0.3415 | 0.019284
2019-03-19 15:49 | 1553024952.7664 | 423 | 0.20857 | 0.000719 | 5.8e-05 | 101.8008 | 101.5986 | 124.0667 | 21.2612 | 0.002376 | 0.002203 | 246.7429 | 0.26513 | 8.4031 | 0.003344 | 0.005418 | 0.005344 | 5.1e-05 | 0.001056 | 0.34827 | 1.3621 | 2.5481 | 101.2 | 20.83 | 17.545 | 78.9894 | 102.3186 | 74.826 | 0.004272 | 1.0146 | 0.02484 | 24.483 | 0.017776 | 0.15773 | 0.004625 | 0.34295 | 0.34295 | 0.019132
2019-03-19 15:49 | 1553024953.7654 | 424 | 0.20861 | 0.000717 | 5.8e-05 | 101.8188 | 101.7538 | 123.7827 | 21.2612 | 0.002419 | 0.002287 | 246.7268 | 0.26549 | 8.4027 | 0.003381 | 0.004884 | 0.004739 | 5.4e-05 | 0.001046 | 0.3131 | 1.3841 | 2.5482 | 101.2 | 20.83 | 17.546 | 80.0232 | 103.2039 | 76.8527 | 0.004272 | 1.0087 | 0.024739 | 24.5261 | 0.00571 | 0.15794 | 0.004373 | 0.32275 | 0.32275 | 0.020357
2019-03-19 15:49 | 1553024954.7668 | 425 | 0.2086 | 0.000712 | 5.8e-05 | 101.8263 | 101.6284 | 123.709 | 21.2612 | 0.00238 | 0.00219 | 246.7099 | 0.26603 | 8.4036 | 0.003349 | 0.005052 | 0.004927 | 5.5e-05 | 0.001047 | 0.29999 | 1.4003 | 2.5481 | 101.2 | 20.82 | 17.542 | 79.0682 | 102.4293 | 74.5204 | 0.004268 | 1.0155 | 0.024815 | 24.437 | 0.017395 | 0.15826 | 0.004461 | 0.33036 | 0.33036 | 0.019929
2019-03-19 15:49 | 1553024955.7669 | 426 | 0.2086 | 0.000716 | 5.9e-05 | 101.8419 | 101.786 | 123.508 | 21.2612 | 0.00244 | 0.00222 | 246.6927 | 0.26603 | 8.4062 | 0.003378 | 0.005054 | 0.00498 | 5.4e-05 | 0.00105 | 0.30845 | 1.3574 | 2.5481 | 101.2 | 20.82 | 17.538 | 80.5207 | 103.1405 | 75.2298 | 0.004267 | 1.0134 | 0.024769 | 24.442 | 0.00491 | 0.15826 | 0.004456 | 0.32934 | 0.32934 | 0.019991
2019-03-19 15:49 | 1553024956.7656 | 427 | 0.20857 | 0.000709 | 5.8e-05 | 101.8911 | 101.6839 | 123.2259 | 21.2612 | 0.002412 | 0.002292 | 246.6731 | 0.26518 | 8.4032 | 0.003294 | 0.004984 | 0.004885 | 5.2e-05 | 0.001048 | 0.31094 | 1.3589 | 2.5482 | 101.2 | 20.82 | 17.532 | 79.8432 | 101.0983 | 76.9687 | 0.004266 | 1.0084 | 0.02468 | 24.4749 | 0.018205 | 0.15776 | 0.004626 | 0.34081 | 0.34081 | 0.019256
2019-03-19 15:49 | 1553024957.7646 | 428 | 0.20861 | 0.000708 | 5.8e-05 | 101.9379 | 101.6969 | 123.466 | 21.2612 | 0.00239 | 0.002228 | 246.6647 | 0.26517 | 8.4013 | 0.003368 | 0.00469 | 0.004644 | 5.2e-05 | 0.001053 | 0.31813 | 1.3037 | 2.5482 | 101.2 | 20.82 | 17.53 | 79.3302 | 102.8951 | 75.4208 | 0.004265 | 1.0128 | 0.024758 | 24.4445 | 0.021169 | 0.15775 | 0.004404 | 0.32537 | 0.32537 | 0.020169
2019-03-19 15:49 | 1553024958.765 | 429 | 0.2086 | 0.000704 | 5.8e-05 | 101.8809 | 101.863 | 123.1984 | 21.2612 | 0.002445 | 0.002252 | 246.6299 | 0.26705 | 8.4018 | 0.003309 | 0.005171 | 0.005029 | 5.6e-05 | 0.001059 | 0.34667 | 1.3224 | 2.5481 | 101.2 | 20.82 | 17.531 | 80.636 | 101.4603 | 75.9984 | 0.004266 | 1.0112 | 0.024711 | 24.4383 | 0.001567 | 0.15887 | 0.004439 | 0.32734 | 0.32734 | 0.02019
2019-03-19 15:49 | 1553024959.7658 | 430 | 0.20863 | 0.000697 | 5.8e-05 | 101.9238 | 101.7315 | 123.0086 | 21.2612 | 0.002429 | 0.002256 | 246.604 | 0.26616 | 8.4019 | 0.003312 | 0.005699 | 0.005585 | 5.4e-05 | 0.001054 | 0.30254 | 1.3136 | 2.5482 | 101.2 | 20.82 | 17.533 | 80.2493 | 101.545 | 76.0942 | 0.004266 | 1.0109 | 0.024689 | 24.4228 | 0.016889 | 0.15834 | 0.00431 | 0.31742 | 0.31742 | 0.020752
2019-03-19 15:49 | 1553024960.7648 | 431 | 0.20862 | 0.000696 | 5.8e-05 | 101.9238 | 102.0319 | 123.3062 | 21.2612 | 0.002386 | 0.002219 | 246.5772 | 0.26607 | 8.4031 | 0.003378 | 0.005046 | 0.004988 | 5.4e-05 | 0.001059 | 0.33339 | 1.3353 | 2.5481 | 101.2 | 20.82 | 17.537 | 79.2333 | 103.131 | 75.2197 | 0.004267 | 1.0134 | 0.024749 | 24.4217 | -0.009481 | 0.15829 | 0.004328 | 0.31953 | 0.31953 | 0.020607
2019-03-19 15:49 | 1553024961.7648 | 432 | 0.2086 | 0.000697 | 6e-05 | 101.9463 | 101.7819 | 123.1999 | 21.2612 | 0.002397 | 0.002279 | 246.5593 | 0.26631 | 8.4009 | 0.003354 | 0.004522 | 0.004373 | 5.4e-05 | 0.00106 | 0.30326 | 1.3498 | 2.5483 | 101.2 | 20.82 | 17.545 | 79.4924 | 102.5622 | 76.6547 | 0.004269 | 1.0093 | 0.024688 | 24.4614 | 0.014439 | 0.15843 | 0.004457 | 0.3283 | 0.3283 | 0.020075
2019-03-19 15:49 | 1553024962.7658 | 433 | 0.20861 | 0.000692 | 6e-05 | 101.9742 | 101.9261 | 122.9691 | 21.2612 | 0.002399 | 0.002189 | 246.5384 | 0.26604 | 8.4033 | 0.003392 | 0.00442 | 0.00437 | 5.5e-05 | 0.001066 | 0.3147 | 1.353 | 2.5481 | 101.2 | 20.82 | 17.55 | 79.5445 | 103.4728 | 74.4994 | 0.00427 | 1.0155 | 0.024741 | 24.363 | 0.004225 | 0.15827 | 0.004447 | 0.32825 | 0.32825 | 0.020058
2019-03-19 15:49 | 1553024963.765 | 434 | 0.2086 | 0.000695 | 6e-05 | 101.9031 | 101.7545 | 122.8114 | 21.2612 | 0.002368 | 0.002213 | 246.5035 | 0.26583 | 8.4046 | 0.003331 | 0.005506 | 0.005316 | 5.3e-05 | 0.00106 | 0.29192 | 1.3334 | 2.5481 | 101.2 | 20.82 | 17.555 | 78.7914 | 101.9855 | 75.0676 | 0.004271 | 1.0139 | 0.024705 | 24.3673 | 0.013055 | 0.15814 | 0.004467 | 0.32925 | 0.32925 | 0.019981
2019-03-19 15:49 | 1553024964.7659 | 435 | 0.2086 | 0.000693 | 5.9e-05 | 102.0243 | 102.1476 | 122.6788 | 21.2612 | 0.00238 | 0.002165 | 246.4643 | 0.26608 | 8.4027 | 0.003299 | 0.004667 | 0.004536 | 4.9e-05 | 0.001056 | 0.28584 | 1.3233 | 2.548 | 101.2 | 20.82 | 17.566 | 79.0697 | 101.2246 | 73.9188 | 0.004274 | 1.0172 | 0.024733 | 24.3139 | -0.010805 | 0.15829 | 0.004497 | 0.33183 | 0.33183 | 0.019845
2019-03-19 15:49 | 1553024965.7664 | 436 | 0.20861 | 0.000695 | 6e-05 | 102.0165 | 101.972 | 122.6661 | 21.2612 | 0.002405 | 0.002251 | 246.4254 | 0.2656 | 8.4016 | 0.003369 | 0.004841 | 0.004713 | 5.2e-05 | 0.001057 | 0.28174 | 1.3445 | 2.5482 | 101.2 | 20.82 | 17.571 | 79.6856 | 102.9189 | 75.9835 | 0.004275 | 1.0112 | 0.024658 | 24.3849 | 0.003909 | 0.15801 | 0.004439 | 0.32654 | 0.32654 | 0.020129
2019-03-19 15:49 | 1553024966.7661 | 437 | 0.20862 | 0.000696 | 6e-05 | 102.027 | 101.8576 | 123.0039 | 21.2612 | 0.002359 | 0.002278 | 246.3938 | 0.26638 | 8.4023 | 0.003369 | 0.004788 | 0.004775 | 5.1e-05 | 0.001054 | 0.31686 | 1.3449 | 2.5481 | 101.2 | 20.82 | 17.576 | 78.5636 | 102.9235 | 76.6228 | 0.004277 | 1.0094 | 0.024669 | 24.4408 | 0.01487 | 0.15847 | 0.004331 | 0.31862 | 0.31862 | 0.02069
2019-03-19 15:49 | 1553024967.7652 | 438 | 0.20865 | 0.000697 | 6e-05 | 102.0225 | 101.9082 | 122.6607 | 21.2612 | 0.002365 | 0.002243 | 246.3633 | 0.26575 | 8.4063 | 0.003375 | 0.004075 | 0.003992 | 4.8e-05 | 0.001051 | 0.28744 | 1.3369 | 2.548 | 101.2 | 20.82 | 17.585 | 78.7189 | 103.0617 | 75.7855 | 0.004279 | 1.0118 | 0.024665 | 24.3774 | 0.010025 | 0.15809 | 0.004161 | 0.30588 | 0.30588 | 0.0215
2019-03-19 15:49 | 1553024968.7662 | 439 | 0.20858 | 0.000696 | 6e-05 | 102.0702 | 101.898 | 122.6102 | 21.2612 | 0.002386 | 0.002236 | 246.3541 | 0.26624 | 8.4019 | 0.003371 | 0.005028 | 0.004894 | 5.2e-05 | 0.001051 | 0.34722 | 1.3323 | 2.548 | 101.2 | 20.82 | 17.589 | 79.2131 | 102.9714 | 75.607 | 0.00428 | 1.0123 | 0.024666 | 24.3662 | 0.01511 | 0.15839 | 0.004577 | 0.3369 | 0.3369 | 0.019557
2019-03-19 15:49 | 1553024969.7652 | 440 | 0.20865 | 0.000698 | 6.1e-05 | 102.0997 | 101.9747 | 122.7403 | 21.2612 | 0.002408 | 0.002198 | 246.3311 | 0.26597 | 8.4028 | 0.003298 | 0.005075 | 0.004961 | 5.2e-05 | 0.001055 | 0.31365 | 1.2938 | 2.5482 | 101.2 | 20.82 | 17.592 | 79.7627 | 101.2036 | 74.7169 | 0.00428 | 1.0149 | 0.02471 | 24.348 | 0.010955 | 0.15823 | 0.004192 | 0.30872 | 0.30872 | 0.021321
2019-03-19 15:49 | 1553024970.7656 | 441 | 0.20863 | 0.000694 | 6.1e-05 | 102.0663 | 101.9476 | 122.7365 | 21.2612 | 0.002473 | 0.002215 | 246.2888 | 0.26649 | 8.4014 | 0.003377 | 0.004861 | 0.004743 | 5e-05 | 0.001056 | 0.31332 | 1.3471 | 2.5481 | 101.2 | 20.82 | 17.594 | 81.309 | 103.1167 | 75.1058 | 0.004281 | 1.0138 | 0.024696 | 24.3612 | 0.010409 | 0.15853 | 0.004323 | 0.31836 | 0.31836 | 0.020715
2019-03-19 15:50 | 1553024971.7644 | 442 | 0.2086 | 0.000692 | 6e-05 | 102.039 | 101.9089 | 122.5548 | 21.2612 | 0.002429 | 0.002241 | 246.2684 | 0.26543 | 8.4012 | 0.003269 | 0.005371 | 0.005249 | 5.3e-05 | 0.001056 | 0.31686 | 1.3583 | 2.5482 | 101.2 | 20.82 | 17.595 | 80.2525 | 100.4877 | 75.7317 | 0.004281 | 1.0119 | 0.024656 | 24.365 | 0.011412 | 0.15791 | 0.004458 | 0.32784 | 0.32784 | 0.020037
2019-03-19 15:50 | 1553024972.7654 | 443 | 0.20858 | 0.000694 | 6.1e-05 | 102.0486 | 102.05 | 122.7278 | 21.2612 | 0.00236 | 0.002194 | 246.2389 | 0.26648 | 8.401 | 0.00328 | 0.004264 | 0.004194 | 5.3e-05 | 0.001056 | 0.3585 | 1.3442 | 2.548 | 101.2 | 20.82 | 17.597 | 78.5949 | 100.7616 | 74.6109 | 0.004282 | 1.0152 | 0.024713 | 24.343 | -0.000123 | 0.15853 | 0.004593 | 0.33861 | 0.33861 | 0.019476
2019-03-19 15:50 | 1553024973.7644 | 444 | 0.20865 | 0.000694 | 6.2e-05 | 102.1159 | 102.0593 | 123.4536 | 21.2612 | 0.002457 | 0.00224 | 246.2174 | 0.26581 | 8.4063 | 0.003362 | 0.004729 | 0.004659 | 5.3e-05 | 0.001054 | 0.31542 | 1.3205 | 2.5481 | 101.2 | 20.82 | 17.596 | 80.9247 | 102.7443 | 75.7036 | 0.004281 | 1.012 | 0.024747 | 24.4532 | 0.004957 | 0.15813 | 0.004189 | 0.30891 | 0.30891 | 0.021295
2019-03-19 15:50 | 1553024974.7648 | 445 | 0.20861 | 0.000693 | 6.2e-05 | 102.0579 | 102.1394 | 123.7536 | 21.2612 | 0.002382 | 0.002274 | 246.2147 | 0.26606 | 8.4025 | 0.003351 | 0.005031 | 0.004902 | 4.9e-05 | 0.001052 | 0.32261 | 1.32 | 2.5481 | 101.2 | 20.83 | 17.595 | 79.136 | 102.4702 | 76.5315 | 0.004284 | 1.0096 | 0.024748 | 24.512 | -0.007139 | 0.15828 | 0.004421 | 0.32623 | 0.32623 | 0.020184
2019-03-19 15:50 | 1553024975.7655 | 446 | 0.20859 | 0.000692 | 6.3e-05 | 102.0838 | 102.0589 | 123.7329 | 21.2612 | 0.002363 | 0.002234 | 246.2004 | 0.26546 | 8.4034 | 0.003315 | 0.004336 | 0.004211 | 5.3e-05 | 0.001049 | 0.31924 | 1.341 | 2.5481 | 101.2 | 20.82 | 17.593 | 78.6636 | 101.6133 | 75.5695 | 0.004281 | 1.0124 | 0.02478 | 24.4762 | 0.002174 | 0.15792 | 0.004512 | 0.33338 | 0.33338 | 0.019706
2019-03-19 15:50 | 1553024976.7646 | 447 | 0.20861 | 0.000688 | 6.3e-05 | 102.085 | 101.9716 | 123.559 | 21.2612 | 0.002418 | 0.002265 | 246.1896 | 0.26658 | 8.4034 | 0.003398 | 0.005164 | 0.004988 | 5.2e-05 | 0.001047 | 0.32947 | 1.3481 | 2.5481 | 101.2 | 20.82 | 17.591 | 79.9811 | 103.622 | 76.3217 | 0.00428 | 1.0102 | 0.024736 | 24.4853 | 0.009937 | 0.15859 | 0.004395 | 0.3241 | 0.3241 | 0.020356
2019-03-19 15:50 | 1553024977.7661 | 448 | 0.20857 | 0.000687 | 6.3e-05 | 102.063 | 101.7956 | 123.7937 | 21.2612 | 0.00234 | 0.002246 | 246.1636 | 0.26595 | 8.4036 | 0.003302 | 0.004482 | 0.004403 | 5.2e-05 | 0.001043 | 0.30381 | 1.3272 | 2.5481 | 101.2 | 20.82 | 17.589 | 78.1086 | 101.2931 | 75.8595 | 0.00428 | 1.0116 | 0.024776 | 24.4924 | 0.023472 | 0.15822 | 0.004665 | 0.34475 | 0.34475 | 0.019092
2019-03-19 15:50 | 1553024978.7662 | 449 | 0.20861 | 0.000692 | 6.4e-05 | 102.0832 | 102.1165 | 124.3265 | 21.2612 | 0.002405 | 0.002314 | 246.1383 | 0.26612 | 8.4032 | 0.003342 | 0.00464 | 0.004606 | 5.2e-05 | 0.00104 | 0.32549 | 1.3327 | 2.5481 | 101.2 | 20.82 | 17.587 | 79.6904 | 102.2535 | 77.4841 | 0.004279 | 1.0069 | 0.024771 | 24.6021 | -0.00292 | 0.15831 | 0.004383 | 0.32357 | 0.32357 | 0.020353
2019-03-19 15:50 | 1553024979.7658 | 450 | 0.20861 | 0.000689 | 6.4e-05 | 102.0495 | 102.0651 | 124.2796 | 21.2612 | 0.00242 | 0.002182 | 246.1197 | 0.26695 | 8.4051 | 0.003308 | 0.005138 | 0.005011 | 5.2e-05 | 0.001041 | 0.29983 | 1.3163 | 2.5481 | 101.2 | 20.82 | 17.587 | 80.0403 | 101.4327 | 74.3185 | 0.004279 | 1.016 | 0.024879 | 24.4861 | -0.001364 | 0.15881 | 0.004434 | 0.32881 | 0.32881 | 0.020092
2019-03-19 15:50 | 1553024980.7648 | 451 | 0.2086 | 0.000689 | 6.5e-05 | 102.1345 | 102.1035 | 124.2301 | 21.2612 | 0.002464 | 0.002234 | 246.0985 | 0.26625 | 8.4053 | 0.003284 | 0.005128 | 0.0051 | 5.4e-05 | 0.00104 | 0.33699 | 1.3682 | 2.5481 | 101.2 | 20.82 | 17.582 | 81.1064 | 100.862 | 75.5743 | 0.004278 | 1.0124 | 0.024829 | 24.5255 | 0.002715 | 0.15839 | 0.004511 | 0.3339 | 0.3339 | 0.019734
2019-03-19 15:50 | 1553024981.7659 | 452 | 0.2086 | 0.000686 | 6.5e-05 | 102.0603 | 102.0788 | 124.3106 | 21.2612 | 0.002379 | 0.002243 | 246.0636 | 0.26547 | 8.4033 | 0.003275 | 0.00473 | 0.00459 | 5.5e-05 | 0.001038 | 0.29983 | 1.3176 | 2.5481 | 101.2 | 20.82 | 17.577 | 79.0603 | 100.6479 | 75.7949 | 0.004277 | 1.0118 | 0.02483 | 24.5412 | -0.001618 | 0.15793 | 0.004503 | 0.33327 | 0.33327 | 0.019713
2019-03-19 15:50 | 1553024982.7663 | 453 | 0.2086 | 0.000687 | 6.6e-05 | 102.0871 | 102.0833 | 124.3955 | 21.2612 | 0.002329 | 0.002263 | 246.0634 | 0.26598 | 8.4016 | 0.003381 | 0.004986 | 0.004925 | 5.4e-05 | 0.001024 | 0.28031 | 1.3657 | 2.5482 | 101.2 | 20.82 | 17.571 | 77.8544 | 103.1969 | 76.2586 | 0.004275 | 1.0104 | 0.024822 | 24.5659 | 0.000332 | 0.15823 | 0.00449 | 0.33211 | 0.33211 | 0.01982
2019-03-19 15:50 | 1553024983.7659 | 454 | 0.20859 | 0.000686 | 6.6e-05 | 102.0579 | 102.1093 | 123.9634 | 21.2612 | 0.002354 | 0.002242 | 246.0582 | 0.26657 | 8.406 | 0.003281 | 0.004772 | 0.004632 | 5.3e-05 | 0.001022 | 0.30574 | 1.3911 | 2.548 | 101.2 | 20.82 | 17.564 | 78.4459 | 100.7787 | 75.756 | 0.004274 | 1.0119 | 0.024796 | 24.5055 | -0.004499 | 0.15858 | 0.004544 | 0.3358 | 0.3358 | 0.019646
2019-03-19 15:50 | 1553024984.765 | 455 | 0.20862 | 0.000689 | 6.6e-05 | 102.0669 | 102.1415 | 123.7223 | 21.2612 | 0.002406 | 0.002245 | 246.0373 | 0.26603 | 8.4041 | 0.003362 | 0.005397 | 0.005347 | 5.5e-05 | 0.001013 | 0.30453 | 1.3788 | 2.5482 | 101.2 | 20.82 | 17.545 | 79.7089 | 102.7491 | 75.84 | 0.004269 | 1.0116 | 0.024769 | 24.4846 | -0.00653 | 0.15826 | 0.004377 | 0.32299 | 0.32299 | 0.020384
2019-03-19 15:50 | 1553024985.7652 | 456 | 0.20863 | 0.000678 | 6.7e-05 | 102.0384 | 102.0555 | 123.6038 | 21.2612 | 0.002367 | 0.002229 | 246.0138 | 0.26585 | 8.4043 | 0.003298 | 0.004793 | 0.004673 | 5.6e-05 | 0.001004 | 0.29502 | 1.3542 | 2.5481 | 101.2 | 20.82 | 17.534 | 78.7539 | 101.1995 | 75.4559 | 0.004266 | 1.0127 | 0.024771 | 24.4594 | -0.001497 | 0.15815 | 0.004335 | 0.31979 | 0.31979 | 0.020573
2019-03-19 15:50 | 1553024986.7662 | 457 | 0.20862 | 0.000683 | 6.7e-05 | 102.1312 | 101.9329 | 123.7301 | 21.2612 | 0.002458 | 0.002202 | 245.9827 | 0.26687 | 8.405 | 0.00353 | 0.005098 | 0.004943 | 5.4e-05 | 0.001006 | 0.31526 | 1.3312 | 2.5481 | 101.2 | 20.82 | 17.513 | 80.9487 | 106.8127 | 74.8058 | 0.004261 | 1.0146 | 0.024807 | 24.4491 | 0.01738 | 0.15876 | 0.00438 | 0.32359 | 0.32359 | 0.02041
2019-03-19 15:50 | 1553024987.7646 | 458 | 0.20864 | 0.000686 | 6.7e-05 | 102.0504 | 102.2048 | 123.8533 | 21.2612 | 0.002406 | 0.002209 | 245.9671 | 0.26609 | 8.4034 | 0.003352 | 0.004917 | 0.004814 | 5.7e-05 | 0.001009 | 0.33539 | 1.3413 | 2.548 | 101.2 | 20.82 | 17.502 | 79.6957 | 102.5125 | 74.9763 | 0.004259 | 1.0141 | 0.024813 | 24.4672 | -0.013523 | 0.1583 | 0.004242 | 0.31337 | 0.31337 | 0.021014
2019-03-19 15:50 | 1553024988.7653 | 459 | 0.20864 | 0.000681 | 6.7e-05 | 102.0919 | 102.1069 | 123.2355 | 21.2612 | 0.002402 | 0.002295 | 245.9389 | 0.26738 | 8.4017 | 0.003403 | 0.005155 | 0.005122 | 5.4e-05 | 0.001014 | 0.30807 | 1.3284 | 2.548 | 101.2 | 20.83 | 17.478 | 79.6178 | 103.7281 | 77.0295 | 0.004255 | 1.0082 | 0.024678 | 24.478 | -0.001317 | 0.15907 | 0.004269 | 0.31365 | 0.31365 | 0.021097
2019-03-19 15:50 | 1553024989.7643 | 460 | 0.20864 | 0.000674 | 6.7e-05 | 102.1069 | 101.9891 | 123.1799 | 21.2612 | 0.0024 | 0.002248 | 245.9408 | 0.26645 | 8.403 | 0.003443 | 0.005097 | 0.005004 | 5.4e-05 | 0.00102 | 0.28368 | 1.3461 | 2.5482 | 101.2 | 20.83 | 17.464 | 79.5522 | 104.6991 | 75.9126 | 0.004252 | 1.0114 | 0.024712 | 24.4334 | 0.010324 | 0.15851 | 0.004274 | 0.31452 | 0.31452 | 0.020965
2019-03-19 15:50 | 1553024990.7644 | 461 | 0.20864 | 0.000672 | 6.8e-05 | 102.0408 | 101.8812 | 123.3003 | 21.2612 | 0.002388 | 0.002197 | 245.9396 | 0.26547 | 8.4042 | 0.003399 | 0.005222 | 0.005152 | 5.3e-05 | 0.001023 | 0.33069 | 1.3387 | 2.548 | 101.2 | 20.83 | 17.451 | 79.2787 | 103.6331 | 74.6919 | 0.004249 | 1.015 | 0.024768 | 24.4026 | 0.014005 | 0.15793 | 0.004236 | 0.31231 | 0.31231 | 0.021036
2019-03-19 15:50 | 1553024991.7648 | 462 | 0.20863 | 0.000674 | 6.7e-05 | 102.0811 | 102.1983 | 123.5299 | 21.2612 | 0.00244 | 0.00225 | 245.9389 | 0.26589 | 8.4018 | 0.003402 | 0.004944 | 0.004869 | 5.4e-05 | 0.001021 | 0.3126 | 1.3334 | 2.5481 | 101.2 | 20.82 | 17.425 | 80.5144 | 103.7109 | 75.947 | 0.00424 | 1.0113 | 0.024746 | 24.4693 | -0.010271 | 0.15818 | 0.004316 | 0.31804 | 0.31804 | 0.02069
2019-03-19 15:50 | 1553024992.7655 | 463 | 0.20867 | 0.000673 | 6.7e-05 | 102.0826 | 102.0302 | 123.4475 | 21.2612 | 0.002343 | 0.00225 | 245.9044 | 0.26673 | 8.404 | 0.003532 | 0.004919 | 0.004766 | 5.4e-05 | 0.001024 | 0.33212 | 1.3216 | 2.5481 | 101.2 | 20.83 | 17.415 | 78.1794 | 106.86 | 75.9508 | 0.00424 | 1.0113 | 0.024738 | 24.4613 | 0.00459 | 0.15868 | 0.004099 | 0.30172 | 0.30172 | 0.021878
2019-03-19 15:50 | 1553024993.7655 | 464 | 0.20866 | 0.000675 | 6.7e-05 | 102.082 | 101.9754 | 123.3372 | 21.2612 | 0.002417 | 0.002249 | 245.8597 | 0.26589 | 8.401 | 0.003435 | 0.004904 | 0.004845 | 5.5e-05 | 0.001025 | 0.30038 | 1.3405 | 2.548 | 101.2 | 20.83 | 17.407 | 79.9655 | 104.5239 | 75.9366 | 0.004238 | 1.0113 | 0.024727 | 24.4499 | 0.009343 | 0.15818 | 0.004121 | 0.30325 | 0.30325 | 0.0217
2019-03-19 15:50 | 1553024994.7646 | 465 | 0.20864 | 0.000673 | 6.6e-05 | 102.0153 | 102.0815 | 123.47 | 21.2612 | 0.002395 | 0.002214 | 245.8273 | 0.26714 | 8.4033 | 0.003452 | 0.005304 | 0.005074 | 5.7e-05 | 0.00103 | 0.32421 | 1.3046 | 2.5481 | 101.2 | 20.83 | 17.394 | 79.4323 | 104.9352 | 75.0809 | 0.004235 | 1.0138 | 0.024771 | 24.433 | -0.005804 | 0.15892 | 0.004242 | 0.31292 | 0.31292 | 0.021127
2019-03-19 15:50 | 1553024995.7657 | 466 | 0.20867 | 0.000671 | 6.7e-05 | 102.082 | 101.8086 | 123.6885 | 21.2612 | 0.002371 | 0.002217 | 245.8197 | 0.26489 | 8.4037 | 0.00345 | 0.005656 | 0.005624 | 5.3e-05 | 0.001035 | 0.3366 | 1.318 | 2.5481 | 101.2 | 20.83 | 17.39 | 78.8577 | 104.8785 | 75.158 | 0.004234 | 1.0136 | 0.02479 | 24.4573 | 0.023985 | 0.15759 | 0.004099 | 0.30242 | 0.30242 | 0.021677
2019-03-19 15:50 | 1553024996.8236 | 467 | 0.20866 | 0.000668 | 6.7e-05 | 102.1186 | 101.9788 | 123.6301 | 21.2612 | 0.002359 | 0.002252 | 245.7938 | 0.26588 | 8.402 | 0.003462 | 0.005107 | 0.005022 | 5.4e-05 | 0.001038 | 0.33124 | 1.3272 | 2.5482 | 101.2 | 20.83 | 17.384 | 78.58 | 105.1798 | 76.014 | 0.004232 | 1.0111 | 0.024754 | 24.4816 | 0.01225 | 0.15817 | 0.004165 | 0.30686 | 0.30686 | 0.021443
2019-03-19 15:50 | 1553024997.7662 | 468 | 0.20863 | 0.000673 | 6.8e-05 | 102.039 | 102.1408 | 123.4674 | 21.2612 | 0.002403 | 0.002218 | 245.7635 | 0.26555 | 8.4024 | 0.00338 | 0.004987 | 0.004965 | 5.8e-05 | 0.001039 | 0.3063 | 1.304 | 2.548 | 101.2 | 20.83 | 17.381 | 79.6317 | 103.1899 | 75.193 | 0.004232 | 1.0135 | 0.024767 | 24.4367 | -0.008916 | 0.15798 | 0.004346 | 0.32052 | 0.32052 | 0.020503
2019-03-19 15:50 | 1553024998.7672 | 469 | 0.20862 | 0.000667 | 6.8e-05 | 102.018 | 101.9106 | 123.7003 | 21.2612 | 0.002379 | 0.002257 | 245.7547 | 0.26633 | 8.4037 | 0.003409 | 0.004814 | 0.004603 | 5.4e-05 | 0.001042 | 0.36033 | 1.3442 | 2.548 | 101.2 | 20.83 | 17.377 | 79.0526 | 103.8933 | 76.1182 | 0.004231 | 1.0108 | 0.024757 | 24.4922 | 0.00942 | 0.15844 | 0.004371 | 0.32225 | 0.32225 | 0.020453
2019-03-19 15:50 | 1553024999.7659 | 470 | 0.20863 | 0.000667 | 6.8e-05 | 102.103 | 101.9589 | 123.8087 | 21.2612 | 0.002387 | 0.002233 | 245.7283 | 0.26687 | 8.4026 | 0.003419 | 0.005252 | 0.005111 | 5.5e-05 | 0.001048 | 0.29374 | 1.3215 | 2.548 | 101.2 | 20.83 | 17.37 | 79.2437 | 104.1322 | 75.557 | 0.004229 | 1.0124 | 0.024788 | 24.4832 | 0.012626 | 0.15876 | 0.004345 | 0.32076 | 0.32076 | 0.02059
2019-03-19 15:50 | 1553025000.766 | 471 | 0.20864 | 0.000672 | 6.7e-05 | 102.1126 | 101.9463 | 123.4219 | 21.2612 | 0.002425 | 0.002241 | 245.6865 | 0.26629 | 8.4045 | 0.003353 | 0.005086 | 0.00493 | 5.3e-05 | 0.001046 | 0.31492 | 1.3435 | 2.5482 | 101.2 | 20.83 | 17.366 | 80.1691 | 102.5401 | 75.7379 | 0.004228 | 1.0119 | 0.024743 | 24.4513 | 0.014579 | 0.15841 | 0.004259 | 0.31375 | 0.31375 | 0.021004
2019-03-19 15:50 | 1553025001.766 | 472 | 0.20863 | 0.000668 | 6.8e-05 | 102.1732 | 102.0144 | 123.1224 | 21.2612 | 0.002401 | 0.002261 | 245.651 | 0.26629 | 8.4025 | 0.003355 | 0.004026 | 0.003973 | 5.4e-05 | 0.001047 | 0.33107 | 1.3352 | 2.5481 | 101.2 | 20.83 | 17.36 | 79.5717 | 102.5755 | 76.2103 | 0.004227 | 1.0105 | 0.024696 | 24.4381 | 0.013909 | 0.15842 | 0.004334 | 0.31873 | 0.31873 | 0.020676
2019-03-19 15:50 | 1553025002.7655 | 473 | 0.20865 | 0.000668 | 6.7e-05 | 102.0651 | 101.9634 | 123.2538 | 21.2612 | 0.002369 | 0.002224 | 245.6187 | 0.26647 | 8.402 | 0.003404 | 0.005035 | 0.00495 | 5e-05 | 0.001045 | 0.32477 | 1.343 | 2.5481 | 101.2 | 20.83 | 17.35 | 78.8123 | 103.7526 | 75.335 | 0.004224 | 1.0131 | 0.02474 | 24.4205 | 0.008922 | 0.15852 | 0.004218 | 0.31067 | 0.31067 | 0.021227
2019-03-19 15:50 | 1553025003.7661 | 474 | 0.20863 | 0.000667 | 6.7e-05 | 102.0618 | 101.9583 | 123.2531 | 21.2612 | 0.002409 | 0.002219 | 245.6094 | 0.26661 | 8.4025 | 0.003389 | 0.00547 | 0.005286 | 5.5e-05 | 0.001047 | 0.31598 | 1.3381 | 2.548 | 101.2 | 20.83 | 17.345 | 79.7853 | 103.3989 | 75.2063 | 0.004223 | 1.0135 | 0.024745 | 24.4159 | 0.009083 | 0.15861 | 0.004293 | 0.31628 | 0.31628 | 0.020862
2019-03-19 15:50 | 1553025004.7665 | 475 | 0.20866 | 0.000661 | 6.7e-05 | 102.0991 | 102.0093 | 123.2693 | 21.2612 | 0.002425 | 0.002265 | 245.5855 | 0.26589 | 8.4022 | 0.003387 | 0.00475 | 0.004717 | 5.2e-05 | 0.001044 | 0.3517 | 1.3489 | 2.5482 | 101.2 | 20.83 | 17.335 | 80.1635 | 103.3541 | 76.3133 | 0.004221 | 1.0103 | 0.024707 | 24.4563 | 0.007868 | 0.15818 | 0.004151 | 0.30525 | 0.30525 | 0.021557
2019-03-19 15:50 | 1553025005.7653 | 476 | 0.20863 | 0.000664 | 6.7e-05 | 102.0372 | 101.9024 | 123.1348 | 21.2612 | 0.002407 | 0.002225 | 245.5553 | 0.26685 | 8.4019 | 0.003419 | 0.004672 | 0.004552 | 5.5e-05 | 0.001035 | 0.30901 | 1.334 | 2.5481 | 101.2 | 20.83 | 17.33 | 79.7255 | 104.1245 | 75.3552 | 0.004219 | 1.013 | 0.024727 | 24.4094 | 0.011825 | 0.15875 | 0.004341 | 0.3197 | 0.3197 | 0.020657
2019-03-19 15:50 | 1553025006.7652 | 477 | 0.20866 | 0.000661 | 6.7e-05 | 102.1063 | 101.9065 | 123.2268 | 21.2612 | 0.002382 | 0.00227 | 245.5436 | 0.26696 | 8.4031 | 0.00335 | 0.005015 | 0.004933 | 5.3e-05 | 0.00103 | 0.33682 | 1.3601 | 2.5481 | 101.2 | 20.83 | 17.326 | 79.1353 | 102.4607 | 76.4371 | 0.004218 | 1.0099 | 0.024698 | 24.4564 | 0.017513 | 0.15881 | 0.004144 | 0.30465 | 0.30465 | 0.021686
2019-03-19 15:50 | 1553025007.7664 | 478 | 0.20862 | 0.000663 | 6.8e-05 | 102.0853 | 101.9641 | 123.3503 | 21.2612 | 0.002347 | 0.002228 | 245.5144 | 0.26655 | 8.4 | 0.003343 | 0.005239 | 0.005134 | 5.2e-05 | 0.001031 | 0.32516 | 1.3449 | 2.5481 | 101.2 | 20.82 | 17.316 | 78.2892 | 102.284 | 75.434 | 0.004213 | 1.0128 | 0.024746 | 24.4336 | 0.010624 | 0.15857 | 0.004381 | 0.32284 | 0.32284 | 0.020433
2019-03-19 15:50 | 1553025008.7654 | 479 | 0.20865 | 0.000666 | 6.8e-05 | 102.094 | 101.9651 | 123.4071 | 21.2612 | 0.002453 | 0.002236 | 245.4968 | 0.26605 | 8.4017 | 0.003319 | 0.005057 | 0.005006 | 5.7e-05 | 0.001034 | 0.31974 | 1.3355 | 2.5481 | 101.2 | 20.82 | 17.311 | 80.8302 | 101.6989 | 75.6087 | 0.004212 | 1.0123 | 0.024746 | 24.4453 | 0.011296 | 0.15827 | 0.004189 | 0.30851 | 0.30851 | 0.021342
2019-03-19 15:50 | 1553025009.7659 | 480 | 0.20864 | 0.000664 | 6.8e-05 | 102.1288 | 102.0473 | 123.132 | 21.2612 | 0.002356 | 0.002194 | 245.4546 | 0.26637 | 8.4015 | 0.00327 | 0.00492 | 0.004758 | 5.4e-05 | 0.001038 | 0.31669 | 1.3876 | 2.5482 | 101.2 | 20.82 | 17.305 | 78.5053 | 100.5189 | 74.6068 | 0.004211 | 1.0152 | 0.024754 | 24.3829 | 0.007138 | 0.15846 | 0.004286 | 0.31585 | 0.31585 | 0.020871
2019-03-19 15:50 | 1553025010.7665 | 481 | 0.20861 | 0.000669 | 6.8e-05 | 102.0898 | 102.1199 | 123.4184 | 21.2612 | 0.002401 | 0.002233 | 245.4291 | 0.26731 | 8.4038 | 0.003245 | 0.005241 | 0.005179 | 5.6e-05 | 0.001033 | 0.31581 | 1.3408 | 2.5482 | 101.2 | 20.82 | 17.304 | 79.5882 | 99.9141 | 75.5417 | 0.00421 | 1.0125 | 0.024749 | 24.4441 | -0.002641 | 0.15902 | 0.00443 | 0.32657 | 0.32657 | 0.020257
2019-03-19 15:50 | 1553025011.7655 | 482 | 0.20863 | 0.000667 | 6.9e-05 | 102.0051 | 102.1117 | 123.4566 | 21.2612 | 0.002371 | 0.002218 | 245.4124 | 0.26606 | 8.4009 | 0.003293 | 0.004491 | 0.004471 | 5.7e-05 | 0.001035 | 0.30127 | 1.3329 | 2.5481 | 101.2 | 20.82 | 17.304 | 78.8585 | 101.085 | 75.1885 | 0.00421 | 1.0135 | 0.024766 | 24.4355 | -0.009339 | 0.15828 | 0.004302 | 0.31715 | 0.31715 | 0.020761
2019-03-19 15:50 | 1553025012.7666 | 483 | 0.20861 | 0.000667 | 6.9e-05 | 102.1102 | 102.0456 | 123.3975 | 21.2612 | 0.00239 | 0.002182 | 245.4312 | 0.26694 | 8.4013 | 0.003293 | 0.004844 | 0.00469 | 5.6e-05 | 0.001037 | 0.32228 | 1.3016 | 2.5481 | 101.2 | 20.82 | 17.308 | 79.3261 | 101.0812 | 74.3161 | 0.004211 | 1.0161 | 0.024791 | 24.399 | 0.005658 | 0.1588 | 0.004476 | 0.33049 | 0.33049 | 0.019989
2019-03-19 15:50 | 1553025013.7661 | 484 | 0.20863 | 0.000669 | 6.9e-05 | 102.06 | 101.7788 | 123.4034 | 21.2612 | 0.002369 | 0.002285 | 245.4254 | 0.26628 | 8.4022 | 0.003295 | 0.005194 | 0.005067 | 5.4e-05 | 0.001035 | 0.33826 | 1.3228 | 2.548 | 101.2 | 20.82 | 17.31 | 78.8039 | 101.1155 | 76.8044 | 0.004212 | 1.0088 | 0.024703 | 24.4868 | 0.024683 | 0.15841 | 0.004333 | 0.31868 | 0.31868 | 0.020678
2019-03-19 15:50 | 1553025014.7657 | 485 | 0.20865 | 0.000671 | 6.9e-05 | 101.991 | 101.8493 | 123.1557 | 21.2612 | 0.002406 | 0.002268 | 245.4171 | 0.26694 | 8.4031 | 0.00334 | 0.005079 | 0.005012 | 5.4e-05 | 0.001038 | 0.30326 | 1.3097 | 2.5481 | 101.2 | 20.82 | 17.316 | 79.705 | 102.2077 | 76.3756 | 0.004213 | 1.0101 | 0.024693 | 24.4472 | 0.012434 | 0.1588 | 0.004216 | 0.30986 | 0.30986 | 0.02132
2019-03-19 15:50 | 1553025015.7657 | 486 | 0.20862 | 0.000667 | 6.9e-05 | 102.1138 | 101.9257 | 123.4057 | 21.2612 | 0.002466 | 0.002268 | 245.3961 | 0.26642 | 8.4028 | 0.003307 | 0.005299 | 0.00513 | 5.5e-05 | 0.001037 | 0.31902 | 1.3042 | 2.548 | 101.2 | 20.82 | 17.318 | 81.1436 | 101.424 | 76.3818 | 0.004214 | 1.0101 | 0.024718 | 24.4722 | 0.016487 | 0.15849 | 0.004394 | 0.32343 | 0.32343 | 0.020386
2019-03-19 15:50 | 1553025016.7672 | 487 | 0.20871 | 0.000671 | 6.8e-05 | 102.0781 | 101.9973 | 123.5698 | 21.2612 | 0.002407 | 0.002219 | 245.3638 | 0.26667 | 8.4007 | 0.003396 | 0.005008 | 0.004977 | 5.7e-05 | 0.001032 | 0.3131 | 1.3747 | 2.5481 | 101.2 | 20.82 | 17.321 | 79.7207 | 103.5772 | 75.2072 | 0.004215 | 1.0135 | 0.024776 | 24.4473 | 0.007079 | 0.15864 | 0.003814 | 0.28091 | 0.28091 | 0.023493
2019-03-19 15:50 | 1553025017.7663 | 488 | 0.20864 | 0.000674 | 6.9e-05 | 102.136 | 101.8572 | 123.4353 | 21.2612 | 0.002376 | 0.002249 | 245.353 | 0.26659 | 8.4016 | 0.003391 | 0.004582 | 0.004515 | 5.7e-05 | 0.001036 | 0.32875 | 1.393 | 2.5483 | 101.2 | 20.82 | 17.328 | 78.9877 | 103.4406 | 75.9282 | 0.004216 | 1.0114 | 0.024737 | 24.4593 | 0.024445 | 0.1586 | 0.004238 | 0.31201 | 0.31201 | 0.021145
2019-03-19 15:50 | 1553025018.765 | 489 | 0.20864 | 0.000673 | 6.9e-05 | 102.0342 | 101.7987 | 123.4179 | 21.2612 | 0.002368 | 0.002265 | 245.3365 | 0.26594 | 8.4029 | 0.003302 | 0.005315 | 0.005135 | 5.5e-05 | 0.001042 | 0.3246 | 1.3558 | 2.5481 | 101.21 | 20.82 | 17.333 | 78.7998 | 101.2868 | 76.3234 | 0.004217 | 1.0102 | 0.024722 | 24.4714 | 0.020676 | 0.15821 | 0.004257 | 0.31325 | 0.31325 | 0.02101
2019-03-19 15:50 | 1553025019.7659 | 490 | 0.20863 | 0.000676 | 6.9e-05 | 102.0285 | 101.8562 | 123.7017 | 21.2612 | 0.002465 | 0.002251 | 245.3189 | 0.26672 | 8.402 | 0.003323 | 0.005196 | 0.005136 | 5.5e-05 | 0.00104 | 0.27704 | 1.3302 | 2.548 | 101.21 | 20.82 | 17.344 | 81.1095 | 101.812 | 75.9835 | 0.00422 | 1.0112 | 0.024762 | 24.4876 | 0.015122 | 0.15867 | 0.00434 | 0.31991 | 0.31991 | 0.020633
2019-03-19 15:50 | 1553025020.7664 | 491 | 0.2086 | 0.00067 | 7e-05 | 102.06 | 102.0055 | 123.7796 | 21.2612 | 0.002402 | 0.002226 | 245.3007 | 0.2665 | 8.4049 | 0.003309 | 0.005187 | 0.005059 | 5.6e-05 | 0.001036 | 0.31293 | 1.3229 | 2.5481 | 101.21 | 20.82 | 17.35 | 79.6116 | 101.4744 | 75.3857 | 0.004221 | 1.0129 | 0.024791 | 24.4743 | 0.00478 | 0.15854 | 0.004534 | 0.33483 | 0.33483 | 0.019698
2019-03-19 15:50 | 1553025021.7652 | 492 | 0.2086 | 0.000679 | 6.9e-05 | 102.0696 | 101.9343 | 123.7627 | 21.2612 | 0.002405 | 0.002241 | 245.2905 | 0.26543 | 8.4026 | 0.003193 | 0.005034 | 0.004864 | 5.5e-05 | 0.001039 | 0.29927 | 1.3242 | 2.5482 | 101.21 | 20.82 | 17.355 | 79.6731 | 98.6553 | 75.7286 | 0.004222 | 1.0119 | 0.024777 | 24.4847 | 0.011871 | 0.1579 | 0.004479 | 0.33053 | 0.33053 | 0.019873
2019-03-19 15:50 | 1553025022.7661 | 493 | 0.20866 | 0.000675 | 7e-05 | 102.076 | 101.7935 | 123.8942 | 21.2612 | 0.002394 | 0.002214 | 245.2317 | 0.26685 | 8.4029 | 0.003301 | 0.005047 | 0.005027 | 5.7e-05 | 0.00104 | 0.29806 | 1.3251 | 2.5481 | 101.21 | 20.82 | 17.368 | 79.4128 | 101.2757 | 75.084 | 0.004226 | 1.0138 | 0.024813 | 24.4751 | 0.024783 | 0.15875 | 0.004124 | 0.30438 | 0.30438 | 0.021697
2019-03-19 15:50 | 1553025023.7652 | 494 | 0.20857 | 0.00068 | 7e-05 | 102.0654 | 101.8336 | 123.7533 | 21.2612 | 0.002399 | 0.002226 | 245.1728 | 0.26584 | 8.4043 | 0.003279 | 0.005568 | 0.005406 | 5.6e-05 | 0.001041 | 0.32925 | 1.3427 | 2.5481 | 101.21 | 20.82 | 17.375 | 79.5445 | 100.7287 | 75.3754 | 0.004227 | 1.013 | 0.024789 | 24.4714 | 0.020342 | 0.15815 | 0.004697 | 0.34693 | 0.34693 | 0.018964
2019-03-19 15:50 | 1553025024.7652 | 495 | 0.20858 | 0.000686 | 7e-05 | 102.0585 | 101.8774 | 123.8627 | 21.2612 | 0.002386 | 0.002257 | 245.1356 | 0.2655 | 8.405 | 0.003345 | 0.004616 | 0.004603 | 5.5e-05 | 0.001042 | 0.32875 | 1.3158 | 2.548 | 101.21 | 20.82 | 17.382 | 79.2201 | 102.3382 | 76.1168 | 0.004229 | 1.0108 | 0.024773 | 24.5082 | 0.015888 | 0.15795 | 0.004628 | 0.3415 | 0.3415 | 0.01924
2019-03-19 15:50 | 1553025025.7653 | 496 | 0.20858 | 0.000685 | 6.9e-05 | 102.0051 | 102.0545 | 123.5921 | 21.2612 | 0.002415 | 0.002231 | 245.1272 | 0.26537 | 8.3996 | 0.003262 | 0.005071 | 0.004968 | 5.8e-05 | 0.001045 | 0.30397 | 1.3065 | 2.5482 | 101.21 | 20.82 | 17.397 | 79.9235 | 100.3213 | 75.5025 | 0.004233 | 1.0126 | 0.024768 | 24.4599 | -0.004328 | 0.15787 | 0.004583 | 0.33815 | 0.33815 | 0.019421
2019-03-19 15:50 | 1553025026.7664 | 497 | 0.2086 | 0.000684 | 7e-05 | 102.0288 | 101.9942 | 123.3904 | 21.2612 | 0.002405 | 0.002232 | 245.1049 | 0.26609 | 8.4003 | 0.003277 | 0.004944 | 0.004811 | 5.5e-05 | 0.00105 | 0.32997 | 1.3029 | 2.548 | 101.21 | 20.82 | 17.403 | 79.6731 | 100.6909 | 75.5251 | 0.004234 | 1.0125 | 0.024747 | 24.4407 | 0.003035 | 0.1583 | 0.004502 | 0.33174 | 0.33174 | 0.01985
2019-03-19 15:50 | 1553025027.7665 | 498 | 0.20862 | 0.000681 | 6.9e-05 | 102.0102 | 101.8089 | 123.5465 | 21.2612 | 0.002393 | 0.002239 | 245.091 | 0.26605 | 8.4019 | 0.003285 | 0.005089 | 0.004979 | 5.5e-05 | 0.001051 | 0.34252 | 1.2928 | 2.5482 | 101.2 | 20.82 | 17.414 | 79.3994 | 100.8794 | 75.7014 | 0.004237 | 1.012 | 0.024756 | 24.4624 | 0.017668 | 0.15827 | 0.004392 | 0.32372 | 0.32372 | 0.020339
2019-03-19 15:50 | 1553025028.7655 | 499 | 0.20858 | 0.00069 | 7e-05 | 101.9991 | 101.9891 | 123.5325 | 21.2612 | 0.002401 | 0.002148 | 245.051 | 0.26592 | 8.401 | 0.003242 | 0.00424 | 0.004149 | 5.5e-05 | 0.001049 | 0.31022 | 1.2893 | 2.5481 | 101.2 | 20.82 | 17.418 | 79.5897 | 99.8326 | 73.5118 | 0.004238 | 1.0184 | 0.024833 | 24.3841 | 0.000881 | 0.1582 | 0.004571 | 0.33811 | 0.33811 | 0.019464
2019-03-19 15:50 | 1553025029.7655 | 500 | 0.2086 | 0.000687 | 6.9e-05 | 101.985 | 101.8493 | 123.3055 | 21.2612 | 0.002448 | 0.002239 | 245.0459 | 0.26645 | 8.4022 | 0.003188 | 0.004673 | 0.00455 | 5.5e-05 | 0.001044 | 0.28783 | 1.3539 | 2.548 | 101.2 | 20.82 | 17.421 | 80.718 | 98.5352 | 75.7021 | 0.004239 | 1.012 | 0.024732 | 24.4385 | 0.011908 | 0.15851 | 0.004464 | 0.32882 | 0.32882 | 0.020054
2019-03-19 15:50 | 1553025030.7665 | 501 | 0.20861 | 0.000684 | 6.9e-05 | 101.8926 | 101.637 | 123.0435 | 21.2612 | 0.002403 | 0.002248 | 245.0281 | 0.26604 | 8.4011 | 0.003222 | 0.005259 | 0.005086 | 5.7e-05 | 0.001045 | 0.33019 | 1.3425 | 2.5481 | 101.2 | 20.81 | 17.424 | 79.6286 | 99.3486 | 75.8977 | 0.004237 | 1.0115 | 0.024699 | 24.4194 | 0.022465 | 0.15827 | 0.004426 | 0.32554 | 0.32554 | 0.020224
2019-03-19 15:50 | 1553025031.767 | 502 | 0.20861 | 0.000683 | 6.9e-05 | 101.9826 | 101.7545 | 122.9053 | 21.2612 | 0.002434 | 0.00223 | 244.9904 | 0.26515 | 8.4018 | 0.00318 | 0.00484 | 0.004777 | 5.5e-05 | 0.001044 | 0.31094 | 1.3342 | 2.5481 | 101.2 | 20.81 | 17.424 | 80.3717 | 98.3368 | 75.4722 | 0.004237 | 1.0127 | 0.0247 | 24.3908 | 0.020032 | 0.15774 | 0.00442 | 0.32512 | 0.32512 | 0.020183
2019-03-19 15:51 | 1553025032.7666 | 503 | 0.20862 | 0.000688 | 6.9e-05 | 101.8707 | 101.7428 | 122.8243 | 21.2612 | 0.002365 | 0.002272 | 244.9855 | 0.26713 | 8.402 | 0.00328 | 0.004769 | 0.004718 | 5.2e-05 | 0.001042 | 0.33572 | 1.316 | 2.5481 | 101.2 | 20.81 | 17.423 | 78.7119 | 100.7671 | 76.4721 | 0.004237 | 1.0098 | 0.024657 | 24.4177 | 0.011233 | 0.15892 | 0.004357 | 0.31987 | 0.31987 | 0.020667
2019-03-19 15:51 | 1553025033.7667 | 504 | 0.20862 | 0.000681 | 7e-05 | 101.8602 | 101.7784 | 122.9851 | 21.2612 | 0.002394 | 0.002248 | 244.9937 | 0.26616 | 8.405 | 0.003145 | 0.003827 | 0.003717 | 5.3e-05 | 0.001042 | 0.32892 | 1.3336 | 2.5482 | 101.2 | 20.81 | 17.422 | 79.4159 | 97.5007 | 75.8955 | 0.004236 | 1.0115 | 0.024693 | 24.4135 | 0.00718 | 0.15834 | 0.00436 | 0.32053 | 0.32053 | 0.020549
2019-03-19 15:51 | 1553025034.7648 | 505 | 0.20862 | 0.00068 | 6.9e-05 | 101.9049 | 101.6945 | 122.9562 | 21.2612 | 0.00242 | 0.002232 | 244.9837 | 0.2656 | 8.4007 | 0.003202 | 0.005076 | 0.004965 | 5.1e-05 | 0.001042 | 0.29645 | 1.3146 | 2.5481 | 101.2 | 20.81 | 17.421 | 80.0506 | 98.8835 | 75.526 | 0.004236 | 1.0125 | 0.024703 | 24.3977 | 0.018483 | 0.158 | 0.004394 | 0.32323 | 0.32323 | 0.020335
2019-03-19 15:51 | 1553025035.7652 | 506 | 0.20861 | 0.00068 | 6.9e-05 | 101.7885 | 101.5867 | 123.0323 | 21.2612 | 0.002428 | 0.002221 | 244.9528 | 0.26635 | 8.4033 | 0.003213 | 0.005271 | 0.005274 | 5.2e-05 | 0.001047 | 0.29856 | 1.3015 | 2.5481 | 101.2 | 20.81 | 17.418 | 80.2407 | 99.1414 | 75.2579 | 0.004236 | 1.0133 | 0.024721 | 24.3959 | 0.017751 | 0.15845 | 0.004436 | 0.32651 | 0.32651 | 0.020188
2019-03-19 15:51 | 1553025036.7649 | 507 | 0.2086 | 0.000676 | 6.9e-05 | 101.8317 | 101.7189 | 122.8802 | 21.2612 | 0.002427 | 0.002215 | 244.94 | 0.26608 | 8.4048 | 0.003147 | 0.004773 | 0.004688 | 5.3e-05 | 0.001047 | 0.28639 | 1.328 | 2.548 | 101.2 | 20.81 | 17.417 | 80.2181 | 97.5555 | 75.1097 | 0.004235 | 1.0137 | 0.024711 | 24.3756 | 0.009915 | 0.15829 | 0.004487 | 0.33023 | 0.33023 | 0.01994
2019-03-19 15:51 | 1553025037.766 | 508 | 0.20861 | 0.000679 | 6.9e-05 | 101.9202 | 101.6171 | 122.691 | 21.2612 | 0.002424 | 0.00226 | 244.9148 | 0.26568 | 8.4028 | 0.003179 | 0.005349 | 0.005277 | 5.3e-05 | 0.001051 | 0.33931 | 1.3603 | 2.548 | 101.2 | 20.81 | 17.415 | 80.1239 | 98.311 | 76.1862 | 0.004235 | 1.0106 | 0.024653 | 24.3944 | 0.026637 | 0.15805 | 0.004451 | 0.32682 | 0.32682 | 0.020118
2019-03-19 15:51 | 1553025038.7661 | 509 | 0.20861 | 0.000674 | 7e-05 | 101.7722 | 101.7048 | 122.9238 | 21.2612 | 0.002413 | 0.002256 | 244.903 | 0.26567 | 8.4018 | 0.003058 | 0.004369 | 0.00421 | 5.4e-05 | 0.00105 | 0.32228 | 1.3095 | 2.5481 | 101.2 | 20.81 | 17.413 | 79.8696 | 95.4004 | 76.0935 | 0.004234 | 1.0109 | 0.02468 | 24.4143 | 0.005929 | 0.15805 | 0.004432 | 0.32569 | 0.32569 | 0.020187
2019-03-19 15:51 | 1553025039.7664 | 510 | 0.20861 | 0.000677 | 7e-05 | 101.8098 | 101.537 | 123.1494 | 21.2612 | 0.002417 | 0.002231 | 244.8557 | 0.26614 | 8.4037 | 0.003052 | 0.004806 | 0.004699 | 5.2e-05 | 0.001051 | 0.31669 | 1.3172 | 2.548 | 101.2 | 20.81 | 17.413 | 79.9648 | 95.2539 | 75.4941 | 0.004234 | 1.0126 | 0.024724 | 24.4157 | 0.023997 | 0.15833 | 0.004446 | 0.32727 | 0.32727 | 0.020125
2019-03-19 15:51 | 1553025040.7651 | 511 | 0.20861 | 0.000682 | 7.1e-05 | 101.8704 | 101.4932 | 123.3595 | 21.2612 | 0.002392 | 0.002232 | 244.8172 | 0.26602 | 8.4008 | 0.003108 | 0.004707 | 0.004591 | 5.4e-05 | 0.001054 | 0.30414 | 1.3142 | 2.5482 | 101.2 | 20.81 | 17.412 | 79.3604 | 96.6168 | 75.5292 | 0.004234 | 1.0125 | 0.024744 | 24.4378 | 0.033183 | 0.15826 | 0.0044 | 0.32409 | 0.32409 | 0.020314
2019-03-19 15:51 | 1553025041.7655 | 512 | 0.20862 | 0.000679 | 7.1e-05 | 101.8281 | 101.4548 | 123.4562 | 21.2612 | 0.002384 | 0.002245 | 244.817 | 0.2652 | 8.4027 | 0.003104 | 0.004673 | 0.004569 | 5.4e-05 | 0.001056 | 0.30525 | 1.3244 | 2.548 | 101.2 | 20.81 | 17.411 | 79.1663 | 96.5072 | 75.8283 | 0.004234 | 1.0117 | 0.024743 | 24.4579 | 0.032849 | 0.15777 | 0.004368 | 0.32165 | 0.32165 | 0.020405
2019-03-19 15:51 | 1553025042.7653 | 513 | 0.2086 | 0.000677 | 7.1e-05 | 101.759 | 101.2524 | 123.2205 | 21.2612 | 0.002458 | 0.002259 | 244.7856 | 0.26676 | 8.4023 | 0.003149 | 0.004979 | 0.00488 | 5.2e-05 | 0.001046 | 0.3011 | 1.337 | 2.5481 | 101.2 | 20.8 | 17.412 | 80.9598 | 97.5922 | 76.1658 | 0.004231 | 1.0107 | 0.024707 | 24.4463 | 0.044647 | 0.1587 | 0.004521 | 0.33259 | 0.33259 | 0.019849
2019-03-19 15:51 | 1553025043.7663 | 514 | 0.20861 | 0.000673 | 7.1e-05 | 101.8644 | 101.4068 | 123.2095 | 21.2612 | 0.002423 | 0.002271 | 244.7454 | 0.26714 | 8.4044 | 0.003248 | 0.004553 | 0.004472 | 5.2e-05 | 0.001043 | 0.29247 | 1.341 | 2.5481 | 101.2 | 20.8 | 17.414 | 80.102 | 99.9924 | 76.4544 | 0.004232 | 1.0098 | 0.024696 | 24.4553 | 0.040266 | 0.15892 | 0.004438 | 0.32627 | 0.32627 | 0.020263
2019-03-19 15:51 | 1553025044.7663 | 515 | 0.20859 | 0.000676 | 7.1e-05 | 101.753 | 101.177 | 123.2919 | 21.2612 | 0.002391 | 0.002263 | 244.7194 | 0.2665 | 8.404 | 0.003262 | 0.004912 | 0.004773 | 5.1e-05 | 0.001041 | 0.29302 | 1.3271 | 2.548 | 101.2 | 20.8 | 17.415 | 79.3403 | 100.3221 | 76.2641 | 0.004232 | 1.0104 | 0.024711 | 24.4568 | 0.050778 | 0.15854 | 0.004577 | 0.33683 | 0.33683 | 0.019581
2019-03-19 15:51 | 1553025045.7663 | 516 | 0.2086 | 0.000676 | 7.1e-05 | 101.6993 | 101.2572 | 123.366 | 21.2612 | 0.002384 | 0.002244 | 244.6986 | 0.26742 | 8.4034 | 0.003312 | 0.005003 | 0.004849 | 5.5e-05 | 0.001043 | 0.30359 | 1.3164 | 2.5481 | 101.2 | 20.8 | 17.417 | 79.1656 | 101.5458 | 75.8066 | 0.004233 | 1.0117 | 0.024735 | 24.4482 | 0.038972 | 0.15909 | 0.004492 | 0.33081 | 0.33081 | 0.020006
2019-03-19 15:51 | 1553025046.7655 | 517 | 0.20859 | 0.000675 | 7.1e-05 | 101.7689 | 101.5274 | 123.4118 | 21.2612 | 0.00239 | 0.002275 | 244.6839 | 0.26607 | 8.4028 | 0.003251 | 0.004933 | 0.004856 | 5.4e-05 | 0.001045 | 0.30326 | 1.3049 | 2.548 | 101.2 | 20.8 | 17.418 | 79.3302 | 100.071 | 76.5526 | 0.004233 | 1.0096 | 0.024713 | 24.4788 | 0.021255 | 0.15829 | 0.00456 | 0.33559 | 0.33559 | 0.019621
2019-03-19 15:51 | 1553025047.7655 | 518 | 0.20857 | 0.000676 | 7.1e-05 | 101.7041 | 101.3699 | 123.4355 | 21.2612 | 0.002409 | 0.002288 | 244.658 | 0.26658 | 8.4021 | 0.003217 | 0.005039 | 0.004978 | 5.2e-05 | 0.001041 | 0.28512 | 1.3596 | 2.5482 | 101.2 | 20.8 | 17.419 | 79.7644 | 99.227 | 76.8558 | 0.004233 | 1.0087 | 0.024704 | 24.4918 | 0.029446 | 0.15859 | 0.004667 | 0.3434 | 0.3434 | 0.019211
2019-03-19 15:51 | 1553025048.7669 | 519 | 0.20857 | 0.000679 | 7.1e-05 | 101.7647 | 101.4154 | 123.6451 | 21.2612 | 0.002403 | 0.002224 | 244.6278 | 0.26725 | 8.4046 | 0.003226 | 0.004703 | 0.004551 | 5.3e-05 | 0.001038 | 0.30326 | 1.3304 | 2.5481 | 101.2 | 20.8 | 17.424 | 79.6373 | 99.467 | 75.3225 | 0.004234 | 1.0131 | 0.02478 | 24.4588 | 0.030758 | 0.15899 | 0.004707 | 0.34748 | 0.34748 | 0.019034
2019-03-19 15:51 | 1553025049.765 | 520 | 0.2086 | 0.000676 | 7.2e-05 | 101.7299 | 101.6223 | 123.7167 | 21.2612 | 0.002434 | 0.002195 | 244.5969 | 0.26534 | 8.4053 | 0.003223 | 0.005038 | 0.004983 | 5.5e-05 | 0.001041 | 0.29175 | 1.3428 | 2.5482 | 101.2 | 20.8 | 17.426 | 80.3702 | 99.3892 | 74.6325 | 0.004235 | 1.0151 | 0.024812 | 24.4417 | 0.009471 | 0.15785 | 0.004492 | 0.33175 | 0.33175 | 0.019794
2019-03-19 15:51 | 1553025050.7657 | 521 | 0.20859 | 0.000671 | 7.2e-05 | 101.747 | 101.6195 | 123.6702 | 21.2612 | 0.002376 | 0.002275 | 244.6124 | 0.26636 | 8.4022 | 0.00328 | 0.005104 | 0.004978 | 5.5e-05 | 0.001034 | 0.31293 | 1.3485 | 2.5481 | 101.2 | 20.8 | 17.431 | 78.9877 | 100.7577 | 76.5627 | 0.004236 | 1.0095 | 0.024738 | 24.5048 | 0.011216 | 0.15846 | 0.004586 | 0.33782 | 0.33782 | 0.019513
2019-03-19 15:51 | 1553025051.7648 | 522 | 0.20857 | 0.000672 | 7.2e-05 | 101.6801 | 101.425 | 124.0059 | 21.2612 | 0.002373 | 0.002253 | 244.5677 | 0.26592 | 8.4057 | 0.003284 | 0.004643 | 0.00457 | 5.4e-05 | 0.001033 | 0.31222 | 1.3546 | 2.548 | 101.2 | 20.8 | 17.435 | 78.9207 | 100.8518 | 76.0327 | 0.004237 | 1.0111 | 0.024791 | 24.5194 | 0.022471 | 0.1582 | 0.004666 | 0.34448 | 0.34448 | 0.019104
2019-03-19 15:51 | 1553025052.7651 | 523 | 0.20861 | 0.000672 | 7.2e-05 | 101.6858 | 101.3099 | 124.1791 | 21.2612 | 0.002417 | 0.002264 | 244.5444 | 0.26625 | 8.4036 | 0.003194 | 0.004598 | 0.004494 | 5.9e-05 | 0.001033 | 0.2891 | 1.3131 | 2.5481 | 101.2 | 20.8 | 17.438 | 79.9617 | 98.6834 | 76.2828 | 0.004238 | 1.0103 | 0.024799 | 24.5453 | 0.033126 | 0.15839 | 0.004448 | 0.32832 | 0.32832 | 0.020069
2019-03-19 15:51 | 1553025053.7648 | 524 | 0.20858 | 0.000678 | 7.1e-05 | 101.6507 | 101.1654 | 123.8165 | 21.2612 | 0.002384 | 0.002238 | 244.5308 | 0.26667 | 8.4021 | 0.003191 | 0.004964 | 0.004842 | 5.6e-05 | 0.001036 | 0.3095 | 1.3014 | 2.5482 | 101.2 | 20.8 | 17.446 | 79.1648 | 98.6138 | 75.6764 | 0.00424 | 1.0121 | 0.024784 | 24.4882 | 0.042807 | 0.15864 | 0.004643 | 0.34274 | 0.34274 | 0.019255
2019-03-19 15:51 | 1553025054.7659 | 525 | 0.20858 | 0.000672 | 7e-05 | 101.6672 | 101.3592 | 123.6238 | 21.2612 | 0.002423 | 0.002269 | 244.5075 | 0.26742 | 8.4029 | 0.003285 | 0.005395 | 0.005293 | 5.7e-05 | 0.001041 | 0.34435 | 1.339 | 2.5481 | 101.2 | 20.8 | 17.45 | 80.1208 | 100.8777 | 76.42 | 0.004241 | 1.0099 | 0.024739 | 24.4952 | 0.027136 | 0.15909 | 0.004594 | 0.3385 | 0.3385 | 0.019552
2019-03-19 15:51 | 1553025055.7653 | 526 | 0.20858 | 0.000676 | 7e-05 | 101.6297 | 101.3003 | 123.4217 | 21.2612 | 0.002449 | 0.0022 | 244.4642 | 0.26666 | 8.4043 | 0.003263 | 0.00451 | 0.00443 | 5.4e-05 | 0.001044 | 0.31725 | 1.3508 | 2.5481 | 101.2 | 20.8 | 17.455 | 80.7257 | 100.3465 | 74.7471 | 0.004242 | 1.0148 | 0.024778 | 24.4165 | 0.029035 | 0.15864 | 0.004635 | 0.34218 | 0.34218 | 0.019286
2019-03-19 15:51 | 1553025056.766 | 527 | 0.20855 | 0.000671 | 6.9e-05 | 101.6327 | 101.3825 | 123.3407 | 21.2612 | 0.002375 | 0.00224 | 244.4328 | 0.26652 | 8.4048 | 0.003247 | 0.005029 | 0.004979 | 5.4e-05 | 0.001045 | 0.28119 | 1.3483 | 2.5481 | 101.2 | 20.8 | 17.455 | 78.9613 | 99.9557 | 75.7091 | 0.004242 | 1.012 | 0.024736 | 24.4422 | 0.022044 | 0.15855 | 0.004784 | 0.35277 | 0.35277 | 0.018697
2019-03-19 15:51 | 1553025057.7649 | 528 | 0.2086 | 0.000675 | 6.9e-05 | 101.6702 | 101.2219 | 123.3698 | 21.2612 | 0.002349 | 0.002237 | 244.4093 | 0.26625 | 8.4052 | 0.003187 | 0.004457 | 0.004366 | 5.4e-05 | 0.00105 | 0.29878 | 1.351 | 2.5482 | 101.2 | 20.8 | 17.455 | 78.3236 | 98.5212 | 75.6303 | 0.004242 | 1.0122 | 0.024741 | 24.4424 | 0.039528 | 0.15839 | 0.004485 | 0.33049 | 0.33049 | 0.019938
2019-03-19 15:51 | 1553025058.7665 | 529 | 0.2086 | 0.000679 | 6.9e-05 | 101.6021 | 101.3133 | 123.0607 | 21.2612 | 0.002405 | 0.002254 | 244.3877 | 0.26645 | 8.4034 | 0.00326 | 0.004765 | 0.004617 | 5.7e-05 | 0.00105 | 0.31365 | 1.3137 | 2.5482 | 101.2 | 20.8 | 17.453 | 79.6856 | 100.2679 | 76.0404 | 0.004241 | 1.011 | 0.024696 | 24.4261 | 0.025456 | 0.15851 | 0.004516 | 0.33215 | 0.33215 | 0.019853
2019-03-19 15:51 | 1553025059.7664 | 530 | 0.20858 | 0.000675 | 6.8e-05 | 101.66 | 101.1517 | 122.9736 | 21.2612 | 0.002472 | 0.002247 | 244.3878 | 0.26589 | 8.4037 | 0.003268 | 0.005378 | 0.005281 | 5.6e-05 | 0.001051 | 0.3079 | 1.3233 | 2.5481 | 101.2 | 20.8 | 17.449 | 81.3003 | 100.4815 | 75.8821 | 0.00424 | 1.0115 | 0.024693 | 24.4119 | 0.044837 | 0.15818 | 0.004655 | 0.34257 | 0.34257 | 0.019209
2019-03-19 15:51 | 1553025060.7661 | 531 | 0.20855 | 0.000676 | 6.9e-05 | 101.612 | 101.2103 | 122.9776 | 21.2612 | 0.002412 | 0.002296 | 244.3596 | 0.26645 | 8.4037 | 0.003262 | 0.004521 | 0.004546 | 5.5e-05 | 0.001044 | 0.29319 | 1.3257 | 2.5479 | 101.2 | 20.8 | 17.444 | 79.8415 | 100.3346 | 77.0694 | 0.004239 | 1.0081 | 0.024651 | 24.4538 | 0.035434 | 0.15851 | 0.004802 | 0.35287 | 0.35287 | 0.018687
2019-03-19 15:51 | 1553025061.7664 | 532 | 0.20855 | 0.000672 | 6.8e-05 | 101.6123 | 101.1616 | 122.9503 | 21.2612 | 0.00238 | 0.002255 | 244.3284 | 0.26672 | 8.401 | 0.00322 | 0.005475 | 0.005237 | 5.2e-05 | 0.001046 | 0.32997 | 1.3019 | 2.548 | 101.2 | 20.8 | 17.432 | 79.0774 | 99.3179 | 76.0825 | 0.004236 | 1.0109 | 0.024683 | 24.4166 | 0.039759 | 0.15867 | 0.004836 | 0.3559 | 0.3559 | 0.018547
2019-03-19 15:51 | 1553025062.7653 | 533 | 0.20853 | 0.000683 | 6.8e-05 | 101.5928 | 101.0818 | 123.0987 | 21.2612 | 0.002434 | 0.002255 | 244.3169 | 0.26666 | 8.4035 | 0.003241 | 0.004708 | 0.004687 | 5.6e-05 | 0.001041 | 0.27936 | 1.3157 | 2.5481 | 101.2 | 20.8 | 17.425 | 80.3717 | 99.8295 | 76.0709 | 0.004235 | 1.011 | 0.024698 | 24.4309 | 0.0451 | 0.15864 | 0.004929 | 0.36301 | 0.36301 | 0.01818
2019-03-19 15:51 | 1553025063.7667 | 534 | 0.20856 | 0.000684 | 6.9e-05 | 101.5106 | 100.913 | 123.6001 | 21.2612 | 0.002439 | 0.002242 | 244.2789 | 0.26681 | 8.4023 | 0.003319 | 0.004589 | 0.004425 | 5.3e-05 | 0.001039 | 0.29695 | 1.3409 | 2.5482 | 101.2 | 20.8 | 17.411 | 80.4849 | 101.6934 | 75.7511 | 0.004231 | 1.0119 | 0.02476 | 24.4694 | 0.052812 | 0.15872 | 0.004745 | 0.35018 | 0.35018 | 0.018856
2019-03-19 15:51 | 1553025064.7654 | 535 | 0.20854 | 0.000682 | 6.9e-05 | 101.4487 | 100.9191 | 123.9043 | 21.2612 | 0.002429 | 0.002238 | 244.2676 | 0.26749 | 8.4033 | 0.003379 | 0.004944 | 0.004835 | 5.2e-05 | 0.001034 | 0.29407 | 1.3164 | 2.5481 | 101.2 | 20.8 | 17.403 | 80.2469 | 103.1678 | 75.6577 | 0.004229 | 1.0121 | 0.024794 | 24.4962 | 0.046816 | 0.15913 | 0.004845 | 0.35813 | 0.35813 | 0.018484
2019-03-19 15:51 | 1553025065.7645 | 536 | 0.20856 | 0.00069 | 7e-05 | 101.4829 | 101.0054 | 124.1974 | 21.2612 | 0.002436 | 0.002258 | 244.2615 | 0.26679 | 8.4021 | 0.003352 | 0.00486 | 0.004772 | 5.1e-05 | 0.00103 | 0.31974 | 1.2934 | 2.5482 | 101.2 | 20.8 | 17.393 | 80.4356 | 102.4978 | 76.137 | 0.004227 | 1.0108 | 0.024806 | 24.542 | 0.042186 | 0.15871 | 0.004709 | 0.34809 | 0.34809 | 0.018968
2019-03-19 15:51 | 1553025066.7655 | 537 | 0.20857 | 0.000688 | 7e-05 | 101.434 | 100.8191 | 124.2714 | 21.2612 | 0.002385 | 0.002259 | 244.2352 | 0.26603 | 8.4029 | 0.003346 | 0.004447 | 0.004383 | 5e-05 | 0.001031 | 0.32029 | 1.3032 | 2.548 | 101.2 | 20.8 | 17.373 | 79.2054 | 102.361 | 76.1745 | 0.004222 | 1.0107 | 0.024812 | 24.5506 | 0.054388 | 0.15826 | 0.004695 | 0.34709 | 0.34709 | 0.018968
2019-03-19 15:51 | 1553025067.7657 | 538 | 0.20855 | 0.000693 | 7e-05 | 101.4535 | 100.6239 | 124.2467 | 21.2612 | 0.002409 | 0.002273 | 244.2064 | 0.2661 | 8.4023 | 0.003325 | 0.004894 | 0.004858 | 5.1e-05 | 0.001028 | 0.30342 | 1.3483 | 2.5481 | 101.2 | 20.8 | 17.363 | 79.769 | 101.8403 | 76.5159 | 0.00422 | 1.0097 | 0.024798 | 24.5602 | 0.073444 | 0.1583 | 0.004775 | 0.35289 | 0.35289 | 0.018661
2019-03-19 15:51 | 1553025068.7656 | 539 | 0.20858 | 0.000695 | 7.1e-05 | 101.3845 | 100.6321 | 123.9383 | 21.2612 | 0.002428 | 0.002299 | 244.1886 | 0.26696 | 8.4042 | 0.003327 | 0.004985 | 0.00491 | 5e-05 | 0.00103 | 0.34108 | 1.3283 | 2.5481 | 101.2 | 20.8 | 17.346 | 80.2282 | 101.9054 | 77.134 | 0.004215 | 1.0079 | 0.024745 | 24.5514 | 0.066626 | 0.15882 | 0.004625 | 0.34088 | 0.34088 | 0.019381
2019-03-19 15:51 | 1553025069.7656 | 540 | 0.20857 | 0.000695 | 7.1e-05 | 101.4472 | 100.8061 | 123.5874 | 21.2612 | 0.002407 | 0.00223 | 244.1638 | 0.26677 | 8.3996 | 0.003288 | 0.00461 | 0.004474 | 4.4e-05 | 0.001031 | 0.33483 | 1.3307 | 2.548 | 101.2 | 20.8 | 17.346 | 79.7293 | 100.9648 | 75.4636 | 0.004215 | 1.0127 | 0.024769 | 24.4581 | 0.056707 | 0.1587 | 0.004661 | 0.3439 | 0.3439 | 0.019197
2019-03-19 15:51 | 1553025070.7653 | 541 | 0.20857 | 0.000695 | 7.1e-05 | 101.4379 | 100.8119 | 123.2707 | 21.2612 | 0.00239 | 0.002248 | 244.1245 | 0.26652 | 8.4036 | 0.00331 | 0.004643 | 0.004617 | 4.7e-05 | 0.001035 | 0.3178 | 1.3106 | 2.5482 | 101.2 | 20.8 | 17.331 | 79.323 | 101.4947 | 75.9032 | 0.004212 | 1.0114 | 0.024722 | 24.4421 | 0.05537 | 0.15855 | 0.004649 | 0.34232 | 0.34232 | 0.019268
2019-03-19 15:51 | 1553025071.7659 | 542 | 0.20855 | 0.000696 | 7.3e-05 | 101.4352 | 100.7263 | 122.7062 | 21.2612 | 0.002391 | 0.002246 | 244.1071 | 0.26603 | 8.4047 | 0.003277 | 0.004243 | 0.004129 | 4.7e-05 | 0.001035 | 0.2896 | 1.3276 | 2.548 | 101.2 | 20.8 | 17.326 | 79.3479 | 100.6871 | 75.8665 | 0.004211 | 1.0115 | 0.024666 | 24.3848 | 0.062732 | 0.15826 | 0.004794 | 0.35224 | 0.35224 | 0.01869
2019-03-19 15:51 | 1553025072.7648 | 543 | 0.20855 | 0.000696 | 7.3e-05 | 101.3824 | 100.6308 | 122.3084 | 21.2612 | 0.002437 | 0.002289 | 244.0673 | 0.2661 | 8.4024 | 0.003259 | 0.004949 | 0.004893 | 4.8e-05 | 0.001036 | 0.31614 | 1.3227 | 2.548 | 101.2 | 20.8 | 17.318 | 80.4457 | 100.264 | 76.901 | 0.004209 | 1.0086 | 0.02459 | 24.3813 | 0.066562 | 0.15831 | 0.004751 | 0.34797 | 0.34797 | 0.018926
2019-03-19 15:51 | 1553025073.7662 | 544 | 0.20854 | 0.000698 | 7.4e-05 | 101.3641 | 100.5927 | 122.1354 | 21.2612 | 0.002416 | 0.002313 | 244.0627 | 0.2668 | 8.4027 | 0.0033 | 0.004827 | 0.004677 | 4.8e-05 | 0.001043 | 0.32565 | 1.3088 | 2.5481 | 101.2 | 20.8 | 17.316 | 79.9398 | 101.2529 | 77.4646 | 0.004208 | 1.0069 | 0.024553 | 24.3836 | 0.068327 | 0.15872 | 0.004819 | 0.35243 | 0.35243 | 0.018735
2019-03-19 15:51 | 1553025074.765 | 545 | 0.20855 | 0.000704 | 7.5e-05 | 101.3653 | 100.4407 | 121.9938 | 21.2612 | 0.002431 | 0.002336 | 244.0179 | 0.26627 | 8.4028 | 0.003298 | 0.005134 | 0.005116 | 5.1e-05 | 0.001043 | 0.33124 | 1.3189 | 2.5479 | 101.2 | 20.8 | 17.316 | 80.3 | 101.1925 | 78.0242 | 0.004208 | 1.0053 | 0.024519 | 24.3889 | 0.081964 | 0.1584 | 0.004752 | 0.34694 | 0.34694 | 0.018993
2019-03-19 15:51 | 1553025075.7659 | 546 | 0.20854 | 0.000702 | 7.4e-05 | 101.2963 | 100.5287 | 122.022 | 21.2612 | 0.002453 | 0.002238 | 243.9928 | 0.26668 | 8.4028 | 0.003261 | 0.005149 | 0.005016 | 4.9e-05 | 0.001043 | 0.32211 | 1.3308 | 2.548 | 101.2 | 20.8 | 17.318 | 80.8357 | 100.2962 | 75.6615 | 0.004209 | 1.0121 | 0.024605 | 24.3096 | 0.068037 | 0.15865 | 0.004848 | 0.35527 | 0.35527 | 0.018577
2019-03-19 15:51 | 1553025076.7655 | 547 | 0.2085 | 0.000704 | 7.5e-05 | 101.248 | 100.3801 | 122.2657 | 21.2612 | 0.002381 | 0.002308 | 243.9638 | 0.26659 | 8.4038 | 0.003266 | 0.004634 | 0.004556 | 5.1e-05 | 0.001037 | 0.28064 | 1.3664 | 2.548 | 101.2 | 20.8 | 17.321 | 79.0923 | 100.4249 | 77.3414 | 0.004209 | 1.0073 | 0.02457 | 24.3924 | 0.077004 | 0.15859 | 0.005035 | 0.36862 | 0.36862 | 0.017897
2019-03-19 15:51 | 1553025077.7665 | 548 | 0.20852 | 0.000702 | 7.6e-05 | 101.1852 | 100.6075 | 122.3495 | 21.2612 | 0.002388 | 0.002247 | 243.948 | 0.26716 | 8.4049 | 0.003257 | 0.004929 | 0.00484 | 4.8e-05 | 0.001038 | 0.32676 | 1.3544 | 2.5481 | 101.2 | 20.8 | 17.326 | 79.259 | 100.2106 | 75.883 | 0.004211 | 1.0115 | 0.02463 | 24.3499 | 0.05122 | 0.15894 | 0.004978 | 0.3652 | 0.3652 | 0.018104
2019-03-19 15:51 | 1553025078.7661 | 549 | 0.2085 | 0.000712 | 7.5e-05 | 101.1054 | 100.6143 | 122.5976 | 21.2612 | 0.002433 | 0.00236 | 243.9203 | 0.26629 | 8.4049 | 0.003359 | 0.004416 | 0.004348 | 4.9e-05 | 0.001044 | 0.32139 | 1.379 | 2.5481 | 101.2 | 20.8 | 17.338 | 80.3452 | 102.6641 | 78.6067 | 0.004213 | 1.0037 | 0.024559 | 24.4695 | 0.043552 | 0.15841 | 0.005045 | 0.36914 | 0.36914 | 0.017852
2019-03-19 15:51 | 1553025079.7655 | 550 | 0.2085 | 0.000713 | 7.6e-05 | 101.1426 | 100.4811 | 122.8691 | 21.2612 | 0.002452 | 0.002275 | 243.8997 | 0.26742 | 8.4025 | 0.003284 | 0.004506 | 0.004447 | 5.2e-05 | 0.00104 | 0.2975 | 1.3505 | 2.548 | 101.2 | 20.8 | 17.347 | 80.8184 | 100.8683 | 76.5495 | 0.004216 | 1.0096 | 0.024659 | 24.4248 | 0.058695 | 0.15909 | 0.005025 | 0.36912 | 0.36912 | 0.017929
2019-03-19 15:51 | 1553025080.7653 | 551 | 0.20849 | 0.000715 | 7.6e-05 | 101.109 | 100.3225 | 122.8968 | 21.2612 | 0.002479 | 0.002283 | 243.8656 | 0.26591 | 8.4029 | 0.003277 | 0.00454 | 0.004441 | 4.7e-05 | 0.001036 | 0.30397 | 1.3354 | 2.548 | 101.2 | 20.8 | 17.363 | 81.4549 | 100.6839 | 76.7482 | 0.00422 | 1.009 | 0.024654 | 24.4345 | 0.069849 | 0.15819 | 0.005086 | 0.37358 | 0.37358 | 0.017615
2019-03-19 15:51 | 1553025081.7643 | 552 | 0.20851 | 0.000716 | 7.6e-05 | 101.0295 | 100.217 | 123.0409 | 21.2612 | 0.002424 | 0.002298 | 243.8305 | 0.26704 | 8.4012 | 0.003265 | 0.004809 | 0.004694 | 4.8e-05 | 0.001031 | 0.34938 | 1.3286 | 2.5482 | 101.2 | 20.8 | 17.372 | 80.1393 | 100.3888 | 77.116 | 0.004222 | 1.0079 | 0.024656 | 24.4617 | 0.072221 | 0.15886 | 0.005011 | 0.36803 | 0.36803 | 0.017957
2019-03-19 15:51 | 1553025082.7653 | 553 | 0.20851 | 0.000718 | 7.5e-05 | 101.0223 | 100.0694 | 122.9464 | 21.2612 | 0.002445 | 0.002282 | 243.8067 | 0.26656 | 8.4003 | 0.003269 | 0.00382 | 0.003754 | 4.8e-05 | 0.001026 | 0.31957 | 1.3179 | 2.548 | 101.2 | 20.8 | 17.381 | 80.6399 | 100.4979 | 76.7326 | 0.004224 | 1.009 | 0.02466 | 24.4389 | 0.084768 | 0.15858 | 0.004997 | 0.36714 | 0.36714 | 0.017968
2019-03-19 15:51 | 1553025083.7657 | 554 | 0.2085 | 0.000722 | 7.5e-05 | 100.9521 | 99.8242 | 122.8501 | 21.2612 | 0.002433 | 0.002248 | 243.7834 | 0.26601 | 8.4031 | 0.003245 | 0.004765 | 0.004645 | 4.7e-05 | 0.001026 | 0.33915 | 1.3406 | 2.5481 | 101.2 | 20.8 | 17.398 | 80.3397 | 99.9078 | 75.9126 | 0.004228 | 1.0114 | 0.024679 | 24.4007 | 0.10049 | 0.15825 | 0.005031 | 0.36996 | 0.36996 | 0.017794
2019-03-19 15:51 | 1553025084.7655 | 555 | 0.2085 | 0.000721 | 7.5e-05 | 100.9344 | 100.1787 | 122.6748 | 21.2612 | 0.002442 | 0.002286 | 243.7599 | 0.26601 | 8.4047 | 0.003279 | 0.004456 | 0.004414 | 4.8e-05 | 0.001024 | 0.30127 | 1.3377 | 2.5481 | 101.2 | 20.8 | 17.407 | 80.5707 | 100.7301 | 76.8176 | 0.00423 | 1.0088 | 0.02463 | 24.4149 | 0.067219 | 0.15825 | 0.005027 | 0.36893 | 0.36893 | 0.017844
2019-03-19 15:51 | 1553025085.7645 | 556 | 0.20851 | 0.000721 | 7.4e-05 | 100.9311 | 100.0773 | 122.5271 | 21.2612 | 0.002421 | 0.002262 | 243.7311 | 0.2669 | 8.4029 | 0.003214 | 0.005061 | 0.004982 | 5e-05 | 0.001025 | 0.31526 | 1.3094 | 2.5481 | 101.2 | 20.8 | 17.424 | 80.0746 | 99.1697 | 76.2547 | 0.004234 | 1.0104 | 0.024635 | 24.3805 | 0.075983 | 0.15878 | 0.004985 | 0.36592 | 0.36592 | 0.018051
2019-03-19 15:51 | 1553025086.7659 | 557 | 0.20851 | 0.000724 | 7.5e-05 | 100.8953 | 99.8228 | 122.2135 | 21.2612 | 0.002389 | 0.002282 | 243.6915 | 0.26655 | 8.4024 | 0.003295 | 0.005153 | 0.005022 | 5e-05 | 0.001026 | 0.30431 | 1.312 | 2.5481 | 101.2 | 20.8 | 17.431 | 79.2857 | 101.1235 | 76.7256 | 0.004236 | 1.0091 | 0.024586 | 24.3657 | 0.09559 | 0.15857 | 0.004987 | 0.36529 | 0.36529 | 0.018058
2019-03-19 15:51 | 1553025087.7648 | 558 | 0.2085 | 0.000726 | 7.5e-05 | 100.7966 | 99.6276 | 122.234 | 21.2612 | 0.002418 | 0.002232 | 243.6685 | 0.26681 | 8.4031 | 0.003231 | 0.004732 | 0.004678 | 5e-05 | 0.00103 | 0.32244 | 1.3604 | 2.5483 | 101.2 | 20.8 | 17.439 | 79.9958 | 99.5845 | 75.5181 | 0.004238 | 1.0126 | 0.024631 | 24.3257 | 0.10434 | 0.15873 | 0.005013 | 0.36789 | 0.36789 | 0.017948
2019-03-19 15:51 | 1553025088.7656 | 559 | 0.20851 | 0.00073 | 7.6e-05 | 100.7987 | 99.7851 | 122.2889 | 21.2612 | 0.002475 | 0.002294 | 243.6356 | 0.2664 | 8.4016 | 0.003255 | 0.005201 | 0.005128 | 5.1e-05 | 0.001035 | 0.32444 | 1.3044 | 2.548 | 101.2 | 20.8 | 17.452 | 81.3542 | 100.1574 | 77.018 | 0.004241 | 1.0082 | 0.024584 | 24.3834 | 0.090396 | 0.15848 | 0.004975 | 0.36431 | 0.36431 | 0.018096
2019-03-19 15:51 | 1553025089.7661 | 560 | 0.20853 | 0.000727 | 7.6e-05 | 100.838 | 99.9276 | 122.2077 | 21.2612 | 0.002417 | 0.002199 | 243.5719 | 0.26578 | 8.4005 | 0.003209 | 0.00495 | 0.004836 | 5.2e-05 | 0.001036 | 0.30663 | 1.335 | 2.5481 | 101.2 | 20.8 | 17.458 | 79.96 | 99.0372 | 74.7193 | 0.004243 | 1.0149 | 0.024657 | 24.2952 | 0.081121 | 0.15812 | 0.004871 | 0.35761 | 0.35761 | 0.018393
2019-03-19 15:51 | 1553025090.7652 | 561 | 0.20846 | 0.000724 | 7.6e-05 | 100.8287 | 99.4564 | 122.4335 | 21.2612 | 0.002427 | 0.002314 | 243.5478 | 0.26604 | 8.4025 | 0.003195 | 0.004712 | 0.004652 | 5.1e-05 | 0.001037 | 0.33931 | 1.3004 | 2.548 | 101.2 | 20.8 | 17.469 | 80.2063 | 98.6959 | 77.5004 | 0.004245 | 1.0068 | 0.024581 | 24.4146 | 0.12258 | 0.15827 | 0.005277 | 0.3866 | 0.3866 | 0.017031
2019-03-19 15:52 | 1553025091.7659 | 562 | 0.20847 | 0.000726 | 7.7e-05 | 100.7762 | 99.5331 | 122.445 | 21.2612 | 0.002388 | 0.002305 | 243.5328 | 0.2664 | 8.3998 | 0.003194 | 0.005256 | 0.005158 | 5.1e-05 | 0.001031 | 0.31332 | 1.3265 | 2.548 | 101.2 | 20.79 | 17.475 | 79.2653 | 98.6875 | 77.2775 | 0.004244 | 1.0075 | 0.02459 | 24.408 | 0.11102 | 0.15848 | 0.005224 | 0.38281 | 0.38281 | 0.017223
2019-03-19 15:52 | 1553025092.7652 | 563 | 0.20847 | 0.000732 | 7.7e-05 | 100.748 | 99.6317 | 122.587 | 21.2612 | 0.002431 | 0.00227 | 243.4916 | 0.26612 | 8.4044 | 0.003284 | 0.004137 | 0.004042 | 4.9e-05 | 0.001029 | 0.32731 | 1.3268 | 2.548 | 101.2 | 20.79 | 17.481 | 80.2976 | 100.8644 | 76.4443 | 0.004246 | 1.0099 | 0.024634 | 24.3931 | 0.099658 | 0.15832 | 0.00522 | 0.38317 | 0.38317 | 0.017188
2019-03-19 15:52 | 1553025093.7663 | 564 | 0.20848 | 0.000733 | 7.8e-05 | 100.6748 | 99.6208 | 122.5452 | 21.2612 | 0.002451 | 0.002291 | 243.4402 | 0.2667 | 8.4028 | 0.00324 | 0.004781 | 0.004634 | 5.1e-05 | 0.001032 | 0.32012 | 1.3447 | 2.5481 | 101.2 | 20.8 | 17.494 | 80.7944 | 99.7949 | 76.9392 | 0.004251 | 1.0084 | 0.024612 | 24.4062 | 0.094139 | 0.15866 | 0.005152 | 0.37771 | 0.37771 | 0.017474
2019-03-19 15:52 | 1553025094.766 | 565 | 0.20846 | 0.000734 | 7.9e-05 | 100.6937 | 99.4122 | 122.6006 | 21.2612 | 0.002419 | 0.002275 | 243.4109 | 0.26656 | 8.4016 | 0.003294 | 0.004682 | 0.004604 | 5.3e-05 | 0.001035 | 0.33843 | 1.2943 | 2.548 | 101.2 | 20.79 | 17.501 | 80.017 | 101.1008 | 76.5454 | 0.00425 | 1.0096 | 0.024632 | 24.398 | 0.11457 | 0.15858 | 0.005266 | 0.38646 | 0.38646 | 0.01707
2019-03-19 15:52 | 1553025095.7661 | 566 | 0.20848 | 0.00074 | 8e-05 | 100.6739 | 99.3478 | 122.4788 | 21.2612 | 0.00247 | 0.002255 | 243.3702 | 0.26626 | 8.4011 | 0.003258 | 0.005023 | 0.004914 | 4.9e-05 | 0.001034 | 0.33732 | 1.3206 | 2.5479 | 101.2 | 20.79 | 17.508 | 81.2374 | 100.2273 | 76.0748 | 0.004252 | 1.0109 | 0.024636 | 24.3695 | 0.1186 | 0.1584 | 0.005141 | 0.37717 | 0.37717 | 0.017471
2019-03-19 15:52 | 1553025096.7642 | 567 | 0.20847 | 0.000745 | 8e-05 | 100.5892 | 99.3341 | 122.5464 | 21.2612 | 0.002479 | 0.00226 | 243.3513 | 0.26682 | 8.4027 | 0.003307 | 0.004634 | 0.004582 | 5.1e-05 | 0.00103 | 0.31022 | 1.3444 | 2.5482 | 101.2 | 20.79 | 17.521 | 81.4494 | 101.4187 | 76.205 | 0.004255 | 1.0106 | 0.024638 | 24.3807 | 0.11231 | 0.15873 | 0.005217 | 0.38278 | 0.38278 | 0.017251
2019-03-19 15:52 | 1553025097.7645 | 568 | 0.20844 | 0.000745 | 8.1e-05 | 100.6592 | 99.4516 | 122.4426 | 21.2612 | 0.002474 | 0.002275 | 243.2999 | 0.26596 | 8.4021 | 0.003227 | 0.005123 | 0.005023 | 5.2e-05 | 0.001031 | 0.34506 | 1.3234 | 2.548 | 101.2 | 20.79 | 17.526 | 81.3301 | 99.4702 | 76.5471 | 0.004257 | 1.0096 | 0.024616 | 24.3823 | 0.10796 | 0.15822 | 0.005342 | 0.39174 | 0.39174 | 0.016802
2019-03-19 15:52 | 1553025098.7642 | 569 | 0.20841 | 0.000746 | 8.1e-05 | 100.5505 | 99.1663 | 122.257 | 21.2612 | 0.002455 | 0.002302 | 243.2726 | 0.26696 | 8.4055 | 0.003242 | 0.004794 | 0.004745 | 5.1e-05 | 0.00103 | 0.3235 | 1.3526 | 2.5481 | 101.2 | 20.79 | 17.533 | 80.8809 | 99.8536 | 77.2004 | 0.004258 | 1.0077 | 0.024574 | 24.3866 | 0.12399 | 0.15881 | 0.005521 | 0.40427 | 0.40427 | 0.016342
2019-03-19 15:52 | 1553025099.7643 | 570 | 0.20844 | 0.000754 | 8.2e-05 | 100.5631 | 99.1553 | 122.1758 | 21.2612 | 0.002404 | 0.002298 | 243.2569 | 0.26616 | 8.4028 | 0.003258 | 0.004296 | 0.004229 | 5.4e-05 | 0.00102 | 0.31094 | 1.3266 | 2.5481 | 101.2 | 20.79 | 17.536 | 79.6591 | 100.2396 | 77.116 | 0.004259 | 1.0079 | 0.024569 | 24.3755 | 0.1261 | 0.15834 | 0.005384 | 0.39403 | 0.39403 | 0.016717
2019-03-19 15:52 | 1553025100.7657 | 571 | 0.20841 | 0.00075 | 8.2e-05 | 100.5472 | 99.0145 | 122.1361 | 21.2612 | 0.002428 | 0.002275 | 243.2272 | 0.26647 | 8.4032 | 0.003245 | 0.005135 | 0.00499 | 5e-05 | 0.001023 | 0.33229 | 1.3074 | 2.5481 | 101.2 | 20.8 | 17.539 | 80.2243 | 99.9047 | 76.5596 | 0.004262 | 1.0095 | 0.024585 | 24.3522 | 0.13739 | 0.15853 | 0.005546 | 0.40624 | 0.40624 | 0.016233
2019-03-19 15:52 | 1553025101.7654 | 572 | 0.20842 | 0.000757 | 8.3e-05 | 100.507 | 99.0197 | 122.2743 | 21.2612 | 0.002525 | 0.002287 | 243.1694 | 0.26757 | 8.4022 | 0.003275 | 0.004288 | 0.004166 | 5.2e-05 | 0.00102 | 0.33085 | 1.2917 | 2.548 | 101.2 | 20.8 | 17.541 | 82.5724 | 100.6525 | 76.8479 | 0.004263 | 1.0087 | 0.024588 | 24.376 | 0.13335 | 0.15918 | 0.005496 | 0.4025 | 0.4025 | 0.016452
2019-03-19 15:52 | 1553025102.7644 | 573 | 0.20843 | 0.000758 | 8.4e-05 | 100.5349 | 98.9495 | 122.4018 | 21.2612 | 0.00243 | 0.002298 | 243.1326 | 0.26744 | 8.403 | 0.003271 | 0.004695 | 0.00457 | 5.2e-05 | 0.00102 | 0.30254 | 1.2834 | 2.5481 | 101.2 | 20.8 | 17.542 | 80.2868 | 100.5535 | 77.1067 | 0.004263 | 1.008 | 0.024592 | 24.3978 | 0.14218 | 0.1591 | 0.005393 | 0.39491 | 0.39491 | 0.016759
2019-03-19 15:52 | 1553025103.7655 | 574 | 0.20842 | 0.000764 | 8.5e-05 | 100.3966 | 98.7477 | 122.814 | 21.2612 | 0.002418 | 0.002284 | 243.088 | 0.26728 | 8.4029 | 0.003224 | 0.00497 | 0.004846 | 5.1e-05 | 0.001019 | 0.31509 | 1.3154 | 2.5482 | 101.2 | 20.8 | 17.539 | 80.003 | 99.4034 | 76.7777 | 0.004262 | 1.0089 | 0.024645 | 24.4273 | 0.14812 | 0.15901 | 0.005469 | 0.40133 | 0.40133 | 0.016482
2019-03-19 15:52 | 1553025104.7644 | 575 | 0.20843 | 0.000765 | 8.7e-05 | 100.4233 | 98.4474 | 122.8379 | 21.2612 | 0.002443 | 0.002299 | 243.0497 | 0.26665 | 8.4049 | 0.003198 | 0.004414 | 0.004407 | 5.3e-05 | 0.001018 | 0.33467 | 1.3478 | 2.548 | 101.2 | 20.8 | 17.535 | 80.5853 | 98.7745 | 77.1201 | 0.004261 | 1.0079 | 0.024635 | 24.4417 | 0.17774 | 0.15863 | 0.005425 | 0.39776 | 0.39776 | 0.016591
2019-03-19 15:52 | 1553025105.7656 | 576 | 0.20841 | 0.000773 | 8.9e-05 | 100.3684 | 98.7953 | 122.8536 | 21.2612 | 0.002499 | 0.002293 | 243.0331 | 0.26634 | 8.4016 | 0.003172 | 0.004348 | 0.004221 | 5.1e-05 | 0.001017 | 0.30989 | 1.3206 | 2.5481 | 101.2 | 20.8 | 17.529 | 81.94 | 98.1426 | 76.9812 | 0.00426 | 1.0083 | 0.024642 | 24.4384 | 0.14129 | 0.15845 | 0.00552 | 0.40477 | 0.40477 | 0.016285
2019-03-19 15:52 | 1553025106.7646 | 577 | 0.20842 | 0.000775 | 9.1e-05 | 100.4125 | 98.7183 | 122.9848 | 21.2612 | 0.002458 | 0.002307 | 243.0102 | 0.26604 | 8.4 | 0.003114 | 0.004617 | 0.004511 | 5.2e-05 | 0.001018 | 0.3314 | 1.3276 | 2.548 | 101.2 | 20.8 | 17.515 | 80.9636 | 96.7397 | 77.3351 | 0.004256 | 1.0073 | 0.024642 | 24.4638 | 0.1522 | 0.15827 | 0.005433 | 0.3982 | 0.3982 | 0.016534
2019-03-19 15:52 | 1553025107.7647 | 578 | 0.2084 | 0.00078 | 9.4e-05 | 100.3273 | 98.634 | 122.9196 | 21.2612 | 0.002392 | 0.0023 | 242.993 | 0.26622 | 8.4001 | 0.003219 | 0.004732 | 0.004549 | 4.9e-05 | 0.001021 | 0.30359 | 1.3018 | 2.5479 | 101.2 | 20.8 | 17.506 | 79.3768 | 99.2801 | 77.1458 | 0.004254 | 1.0078 | 0.024643 | 24.4507 | 0.15224 | 0.15837 | 0.005537 | 0.40574 | 0.40574 | 0.016238
2019-03-19 15:52 | 1553025108.7647 | 579 | 0.20843 | 0.000779 | 9.5e-05 | 100.3197 | 98.5371 | 123.0665 | 21.2612 | 0.002416 | 0.002285 | 242.9749 | 0.26667 | 8.4027 | 0.003221 | 0.004828 | 0.00473 | 5.3e-05 | 0.001023 | 0.30862 | 1.3167 | 2.548 | 101.2 | 20.8 | 17.488 | 79.9492 | 99.3353 | 76.7847 | 0.00425 | 1.0089 | 0.02467 | 24.4527 | 0.16037 | 0.15864 | 0.005363 | 0.39322 | 0.39322 | 0.016783
2019-03-19 15:52 | 1553025109.7648 | 580 | 0.2084 | 0.000783 | 9.7e-05 | 100.2087 | 98.4796 | 123.078 | 21.2612 | 0.002542 | 0.002287 | 242.9222 | 0.2669 | 8.4015 | 0.0033 | 0.00439 | 0.004314 | 5.2e-05 | 0.001023 | 0.30021 | 1.3218 | 2.5481 | 101.2 | 20.8 | 17.479 | 82.9828 | 101.2474 | 76.8417 | 0.004248 | 1.0087 | 0.024669 | 24.4558 | 0.15569 | 0.15878 | 0.005557 | 0.40751 | 0.40751 | 0.016208
2019-03-19 15:52 | 1553025110.7658 | 581 | 0.20835 | 0.000787 | 0.0001 | 100.1931 | 98.5662 | 122.8445 | 21.2612 | 0.002542 | 0.002341 | 242.8647 | 0.26653 | 8.403 | 0.003262 | 0.005314 | 0.005189 | 5.2e-05 | 0.001021 | 0.30845 | 1.3423 | 2.5481 | 101.2 | 20.8 | 17.47 | 82.9641 | 100.3354 | 78.1427 | 0.004246 | 1.005 | 0.0246 | 24.478 | 0.14643 | 0.15856 | 0.005857 | 0.42846 | 0.42846 | 0.015395
2019-03-19 15:52 | 1553025111.7649 | 582 | 0.20833 | 0.00079 | 0.000102 | 100.1667 | 98.4686 | 122.6797 | 21.2612 | 0.002477 | 0.002287 | 242.8236 | 0.26592 | 8.4039 | 0.003301 | 0.004992 | 0.004883 | 5.5e-05 | 0.001021 | 0.32228 | 1.306 | 2.5481 | 101.2 | 20.79 | 17.453 | 81.4018 | 101.2813 | 76.8517 | 0.004239 | 1.0087 | 0.024629 | 24.4165 | 0.15293 | 0.1582 | 0.005944 | 0.4353 | 0.4353 | 0.015118
2019-03-19 15:52 | 1553025112.7639 | 583 | 0.20829 | 0.000797 | 0.000105 | 100.1745 | 98.3645 | 122.249 | 21.2612 | 0.00251 | 0.002333 | 242.7709 | 0.26637 | 8.404 | 0.003355 | 0.004341 | 0.004297 | 5.2e-05 | 0.001021 | 0.3178 | 1.3116 | 2.5482 | 101.2 | 20.79 | 17.444 | 82.2027 | 102.5762 | 77.9401 | 0.004237 | 1.0056 | 0.024547 | 24.4115 | 0.16309 | 0.15846 | 0.006192 | 0.45205 | 0.45205 | 0.014583
2019-03-19 15:52 | 1553025113.7644 | 584 | 0.20825 | 0.000801 | 0.000107 | 100.0827 | 98.0025 | 121.8483 | 21.2612 | 0.002444 | 0.002382 | 242.7327 | 0.26593 | 8.4037 | 0.003318 | 0.00419 | 0.0041 | 5.2e-05 | 0.001019 | 0.29535 | 1.2968 | 2.5482 | 101.2 | 20.79 | 17.435 | 80.6182 | 101.6745 | 79.1274 | 0.004234 | 1.0022 | 0.024466 | 24.4126 | 0.18787 | 0.1582 | 0.006412 | 0.46664 | 0.46664 | 0.014103
2019-03-19 15:52 | 1553025114.764 | 585 | 0.20793 | 0.0008 | 0.000109 | 99.956 | 98.1693 | 121.6577 | 21.2612 | 0.002488 | 0.002329 | 242.6977 | 0.26703 | 8.4041 | 0.003351 | 0.004818 | 0.00466 | 5.3e-05 | 0.001017 | 0.32068 | 1.3259 | 2.5481 | 101.2 | 20.79 | 17.416 | 81.6779 | 102.4772 | 77.8606 | 0.00423 | 1.0058 | 0.024491 | 24.3497 | 0.16133 | 0.15886 | 0.008351 | 0.61016 | 0.61016 | 0.010831
2019-03-19 15:52 | 1553025115.7651 | 586 | 0.20766 | 0.000814 | 0.000109 | 100.0124 | 98.3354 | 121.0043 | 21.2612 | 0.002524 | 0.002383 | 242.6209 | 0.26599 | 8.4025 | 0.003439 | 0.004839 | 0.00473 | 5.6e-05 | 0.001018 | 0.34811 | 1.328 | 2.548 | 101.2 | 20.79 | 17.407 | 82.5474 | 104.615 | 79.1617 | 0.004228 | 1.0021 | 0.02438 | 24.3291 | 0.15126 | 0.15824 | 0.009947 | 0.72458 | 0.72458 | 0.009085
2019-03-19 15:52 | 1553025116.7651 | 587 | 0.20746 | 0.000838 | 0.00011 | 99.9023 | 98.6536 | 120.4961 | 21.2612 | 0.002498 | 0.00237 | 242.5344 | 0.26729 | 8.4021 | 0.003378 | 0.004618 | 0.004515 | 5.5e-05 | 0.00101 | 0.29247 | 1.3055 | 2.5482 | 101.2 | 20.79 | 17.392 | 81.908 | 103.1332 | 78.8328 | 0.004224 | 1.003 | 0.02434 | 24.2667 | 0.11251 | 0.15901 | 0.011145 | 0.81123 | 0.81123 | 0.008154
2019-03-19 15:52 | 1553025117.7642 | 588 | 0.20719 | 0.000883 | 0.000111 | 99.8498 | 98.9118 | 120.0233 | 21.2612 | 0.002522 | 0.002392 | 242.4401 | 0.26688 | 8.4013 | 0.003325 | 0.004587 | 0.004509 | 5.8e-05 | 0.001011 | 0.30414 | 1.3418 | 2.5482 | 101.2 | 20.79 | 17.383 | 82.489 | 101.8466 | 79.3722 | 0.004222 | 1.0015 | 0.024274 | 24.2376 | 0.084422 | 0.15877 | 0.012693 | 0.92204 | 0.92204 | 0.007163
2019-03-19 15:52 | 1553025118.7652 | 589 | 0.20688 | 0.000939 | 0.000111 | 99.8582 | 99.1944 | 119.2267 | 21.2612 | 0.002511 | 0.002349 | 242.3246 | 0.2667 | 8.4032 | 0.003388 | 0.004694 | 0.004562 | 5.3e-05 | 0.001013 | 0.33339 | 1.333 | 2.5482 | 101.2 | 20.79 | 17.376 | 82.2301 | 103.3681 | 78.3212 | 0.00422 | 1.0045 | 0.024229 | 24.121 | 0.05966 | 0.15866 | 0.014471 | 1.0501 | 1.0501 | 0.006286
2019-03-19 15:52 | 1553025119.7653 | 590 | 0.20661 | 0.001024 | 0.000112 | 99.7631 | 99.392 | 118.5211 | 21.2612 | 0.002519 | 0.0024 | 242.2686 | 0.26713 | 8.4023 | 0.003399 | 0.00513 | 0.005022 | 5.5e-05 | 0.001012 | 0.30917 | 1.2987 | 2.5481 | 101.2 | 20.79 | 17.361 | 82.4164 | 103.6339 | 79.5476 | 0.004216 | 1.001 | 0.024115 | 24.0914 | 0.033333 | 0.15891 | 0.015972 | 1.154 | 1.154 | 0.005729
2019-03-19 15:52 | 1553025120.7643 | 591 | 0.20638 | 0.001128 | 0.000113 | 99.806 | 99.307 | 118.0939 | 21.2612 | 0.00257 | 0.002446 | 242.206 | 0.26712 | 8.4017 | 0.003475 | 0.004439 | 0.004289 | 5.4e-05 | 0.00101 | 0.30038 | 1.347 | 2.548 | 101.2 | 20.79 | 17.355 | 83.6529 | 105.4778 | 80.6711 | 0.004215 | 0.99781 | 0.024033 | 24.0862 | 0.044827 | 0.15891 | 0.01726 | 1.2431 | 1.2431 | 0.005318
2019-03-19 15:52 | 1553025121.7644 | 592 | 0.20606 | 0.001253 | 0.000114 | 99.691 | 99.2183 | 117.151 | 21.2612 | 0.002601 | 0.00244 | 242.1448 | 0.26651 | 8.4018 | 0.003421 | 0.005195 | 0.005018 | 5.6e-05 | 0.00101 | 0.30254 | 1.3431 | 2.5482 | 101.2 | 20.79 | 17.343 | 84.4017 | 104.1802 | 80.513 | 0.004212 | 0.99825 | 0.023943 | 23.9845 | 0.042513 | 0.15855 | 0.018991 | 1.363 | 1.363 | 0.004839
2019-03-19 15:52 | 1553025122.7654 | 593 | 0.2057 | 0.001391 | 0.000117 | 99.6679 | 99.1844 | 116.5091 | 21.2612 | 0.002602 | 0.002453 | 242.0635 | 0.26728 | 8.4019 | 0.00346 | 0.004386 | 0.004329 | 5.5e-05 | 0.001006 | 0.28799 | 1.374 | 2.5481 | 101.2 | 20.79 | 17.339 | 84.4253 | 105.127 | 80.8264 | 0.004211 | 0.99737 | 0.023866 | 23.9293 | 0.043497 | 0.15901 | 0.020975 | 1.5008 | 1.5008 | 0.004407
2019-03-19 15:52 | 1553025123.7645 | 594 | 0.20522 | 0.001546 | 0.00012 | 99.6343 | 99.0618 | 115.7406 | 21.2612 | 0.002659 | 0.002535 | 241.9703 | 0.26681 | 8.4032 | 0.003412 | 0.004771 | 0.004612 | 5.6e-05 | 0.001006 | 0.32029 | 1.3522 | 2.5482 | 101.2 | 20.79 | 17.337 | 85.7799 | 103.9609 | 82.8056 | 0.004211 | 0.99182 | 0.023721 | 23.9168 | 0.051544 | 0.15873 | 0.023705 | 1.6861 | 1.6861 | 0.003916
2019-03-19 15:52 | 1553025124.7645 | 595 | 0.20456 | 0.001808 | 0.000125 | 99.6412 | 98.7217 | 114.5428 | 21.2612 | 0.002627 | 0.002622 | 241.8502 | 0.26629 | 8.4045 | 0.003477 | 0.004768 | 0.004723 | 5.7e-05 | 0.001005 | 0.32516 | 1.3833 | 2.5483 | 101.2 | 20.79 | 17.339 | 85.016 | 105.529 | 84.902 | 0.004211 | 0.98602 | 0.023529 | 23.8627 | 0.082924 | 0.15841 | 0.027317 | 1.9273 | 1.9273 | 0.003419
2019-03-19 15:52 | 1553025125.7655 | 596 | 0.20421 | 0.00208 | 0.000132 | 99.5008 | 98.5758 | 113.8489 | 21.2612 | 0.002655 | 0.002568 | 241.7306 | 0.26696 | 8.4023 | 0.003439 | 0.005036 | 0.004905 | 5.5e-05 | 0.001006 | 0.2844 | 1.3683 | 2.5483 | 101.2 | 20.8 | 17.34 | 85.695 | 104.6017 | 83.6084 | 0.004214 | 0.98959 | 0.0235 | 23.7473 | 0.083537 | 0.15882 | 0.029062 | 2.0476 | 2.0476 | 0.003227
2019-03-19 15:52 | 1553025126.7656 | 597 | 0.2038 | 0.002377 | 0.00014 | 99.5182 | 98.5066 | 112.7666 | 21.2612 | 0.002763 | 0.002643 | 241.6217 | 0.26719 | 8.4016 | 0.003479 | 0.004562 | 0.00442 | 5.6e-05 | 0.001004 | 0.32211 | 1.3464 | 2.5482 | 101.2 | 20.79 | 17.347 | 88.2792 | 105.5896 | 85.3975 | 0.004213 | 0.98465 | 0.02333 | 23.6933 | 0.091381 | 0.15895 | 0.031106 | 2.1753 | 2.1753 | 0.00304
2019-03-19 15:52 | 1553025127.7637 | 598 | 0.2036 | 0.002692 | 0.000149 | 99.5029 | 98.3165 | 111.8202 | 21.2612 | 0.002755 | 0.002719 | 241.5301 | 0.2669 | 8.403 | 0.003465 | 0.004643 | 0.004548 | 5.7e-05 | 0.000999 | 0.29103 | 1.3173 | 2.5481 | 101.2 | 20.8 | 17.352 | 88.1067 | 105.238 | 87.2195 | 0.004217 | 0.97968 | 0.023173 | 23.6536 | 0.10728 | 0.15878 | 0.031938 | 2.218 | 2.218 | 0.002978
2019-03-19 15:52 | 1553025128.7637 | 599 | 0.20291 | 0.003069 | 0.000159 | 99.4648 | 98.2635 | 111.0149 | 21.2612 | 0.002793 | 0.002787 | 241.4204 | 0.26772 | 8.4048 | 0.003486 | 0.004581 | 0.004471 | 5.6e-05 | 0.000996 | 0.29103 | 1.2934 | 2.548 | 101.2 | 20.79 | 17.357 | 89.0032 | 105.7431 | 88.8751 | 0.004215 | 0.9752 | 0.023036 | 23.6223 | 0.10869 | 0.15926 | 0.035611 | 2.458 | 2.458 | 0.002695
2019-03-19 15:52 | 1553025129.7648 | 600 | 0.20236 | 0.003424 | 0.000169 | 99.3888 | 98.0939 | 110.1633 | 21.2612 | 0.002822 | 0.002832 | 241.2892 | 0.26755 | 8.4021 | 0.003529 | 0.004989 | 0.004815 | 5.9e-05 | 0.000998 | 0.29878 | 1.3229 | 2.5481 | 101.2 | 20.79 | 17.371 | 89.6979 | 106.7843 | 89.9455 | 0.004219 | 0.97232 | 0.022914 | 23.5663 | 0.1173 | 0.15916 | 0.0384 | 2.6358 | 2.6358 | 0.002512
2019-03-19 15:52 | 1553025130.7639 | 601 | 0.20166 | 0.003849 | 0.000177 | 99.306 | 97.9234 | 109.3615 | 21.2612 | 0.002887 | 0.002779 | 241.1349 | 0.26637 | 8.4041 | 0.003523 | 0.004841 | 0.004749 | 5.7e-05 | 0.001002 | 0.31022 | 1.3007 | 2.5481 | 101.2 | 20.79 | 17.38 | 91.282 | 106.6552 | 88.6736 | 0.004221 | 0.97574 | 0.022871 | 23.4392 | 0.12541 | 0.15846 | 0.042049 | 2.8802 | 2.8802 | 0.002289
2019-03-19 15:52 | 1553025131.764 | 602 | 0.20092 | 0.004272 | 0.000185 | 99.2694 | 97.6247 | 108.9801 | 21.2612 | 0.003 | 0.002894 | 240.9541 | 0.26621 | 8.4016 | 0.003652 | 0.005077 | 0.004952 | 5.6e-05 | 0.001006 | 0.3194 | 1.3137 | 2.548 | 101.2 | 20.79 | 17.402 | 93.9907 | 109.777 | 91.4314 | 0.004226 | 0.96836 | 0.022744 | 23.4873 | 0.14943 | 0.15837 | 0.045922 | 3.1272 | 3.1272 | 0.002107
2019-03-19 15:52 | 1553025132.765 | 603 | 0.2 | 0.004699 | 0.000191 | 99.2589 | 97.4853 | 108.2583 | 21.2612 | 0.003038 | 0.002976 | 240.7885 | 0.2676 | 8.4009 | 0.003685 | 0.005224 | 0.005107 | 5.7e-05 | 0.001006 | 0.32278 | 1.3351 | 2.548 | 101.2 | 20.79 | 17.414 | 94.9159 | 110.5744 | 93.4186 | 0.004229 | 0.96311 | 0.022607 | 23.4731 | 0.16127 | 0.1592 | 0.050893 | 3.4439 | 3.4439 | 0.001923
2019-03-19 15:52 | 1553025133.765 | 604 | 0.19934 | 0.005168 | 0.000196 | 99.1508 | 97.199 | 107.3246 | 21.2612 | 0.003087 | 0.002969 | 240.6039 | 0.26682 | 8.4042 | 0.003659 | 0.005611 | 0.005476 | 5.5e-05 | 0.00101 | 0.32659 | 1.3294 | 2.548 | 101.2 | 20.79 | 17.426 | 96.1076 | 109.9629 | 93.2489 | 0.004232 | 0.96356 | 0.022515 | 23.3663 | 0.17783 | 0.15873 | 0.05423 | 3.6538 | 3.6538 | 0.001807
2019-03-19 15:52 | 1553025134.765 | 605 | 0.19871 | 0.00558 | 0.0002 | 99.0884 | 97.1521 | 106.8579 | 21.2612 | 0.003143 | 0.003013 | 240.441 | 0.26673 | 8.4009 | 0.003716 | 0.004771 | 0.004643 | 5.5e-05 | 0.001012 | 0.31116 | 1.3577 | 2.548 | 101.2 | 20.79 | 17.448 | 97.4543 | 111.3271 | 94.3216 | 0.004238 | 0.96074 | 0.022433 | 23.3495 | 0.17652 | 0.15868 | 0.05749 | 3.8585 | 3.8585 | 0.001711
2019-03-19 15:52 | 1553025135.7651 | 606 | 0.19852 | 0.005995 | 0.000202 | 98.9747 | 96.9349 | 106.5922 | 21.2612 | 0.003143 | 0.003064 | 240.2962 | 0.26722 | 8.4032 | 0.003766 | 0.00499 | 0.00488 | 5.6e-05 | 0.001017 | 0.34075 | 1.3691 | 2.5481 | 101.2 | 20.79 | 17.458 | 97.4487 | 112.5492 | 95.5546 | 0.00424 | 0.95753 | 0.022367 | 23.3596 | 0.18626 | 0.15897 | 0.05812 | 3.8891 | 3.8891 | 0.0017
2019-03-19 15:52 | 1553025136.7652 | 607 | 0.19835 | 0.006368 | 0.000202 | 99.0011 | 96.6031 | 106.3183 | 21.2612 | 0.003184 | 0.003111 | 240.1382 | 0.26696 | 8.4028 | 0.003771 | 0.005106 | 0.005029 | 5.3e-05 | 0.001015 | 0.30862 | 1.3206 | 2.5479 | 101.19 | 20.79 | 17.468 | 98.4436 | 112.667 | 96.6866 | 0.004243 | 0.9546 | 0.022304 | 23.3653 | 0.21932 | 0.15881 | 0.058698 | 3.9166 | 3.9166 | 0.001687
2019-03-19 15:52 | 1553025137.7642 | 608 | 0.19806 | 0.006625 | 0.000201 | 98.9384 | 96.5767 | 106.2575 | 21.2612 | 0.003264 | 0.00315 | 239.9419 | 0.26661 | 8.4027 | 0.003782 | 0.005583 | 0.005473 | 5.7e-05 | 0.001021 | 0.33555 | 1.3422 | 2.5481 | 101.2 | 20.79 | 17.489 | 100.3787 | 112.938 | 97.6188 | 0.004248 | 0.9522 | 0.02227 | 23.3881 | 0.2161 | 0.15861 | 0.060126 | 4.0054 | 4.0054 | 0.001647
2019-03-19 15:52 | 1553025138.7643 | 609 | 0.19767 | 0.006903 | 0.0002 | 98.8258 | 96.3541 | 106.3082 | 21.2612 | 0.003276 | 0.003172 | 239.7547 | 0.26653 | 8.4032 | 0.003797 | 0.005617 | 0.005577 | 5.4e-05 | 0.001019 | 0.31669 | 1.3411 | 2.548 | 101.2 | 20.79 | 17.5 | 100.6675 | 113.3188 | 98.1614 | 0.00425 | 0.95081 | 0.022259 | 23.4108 | 0.22656 | 0.15856 | 0.062127 | 4.1362 | 4.1362 | 0.001595
2019-03-19 15:52 | 1553025139.7643 | 610 | 0.19729 | 0.007257 | 0.0002 | 98.815 | 96.2147 | 106.289 | 21.2612 | 0.003309 | 0.00319 | 239.5611 | 0.26715 | 8.4003 | 0.003818 | 0.005029 | 0.004906 | 5.5e-05 | 0.001026 | 0.30989 | 1.3307 | 2.5481 | 101.2 | 20.79 | 17.522 | 101.4579 | 113.8199 | 98.5855 | 0.004256 | 0.94972 | 0.022244 | 23.422 | 0.23853 | 0.15893 | 0.063929 | 4.2529 | 4.2529 | 0.001555
2019-03-19 15:52 | 1553025140.7644 | 611 | 0.19709 | 0.007486 | 0.000201 | 98.7592 | 96.1212 | 106.0268 | 21.2612 | 0.003373 | 0.003242 | 239.3935 | 0.26619 | 8.4037 | 0.003861 | 0.004882 | 0.004665 | 5.4e-05 | 0.00103 | 0.32189 | 1.3446 | 2.5482 | 101.2 | 20.79 | 17.535 | 103.0075 | 114.8611 | 99.8333 | 0.004259 | 0.94654 | 0.02218 | 23.4323 | 0.24217 | 0.15835 | 0.064839 | 4.3006 | 4.3006 | 0.001532
2019-03-19 15:52 | 1553025141.7644 | 612 | 0.19717 | 0.007725 | 0.000202 | 98.6422 | 96.0739 | 105.8669 | 21.2612 | 0.003386 | 0.003299 | 239.2273 | 0.26704 | 8.4048 | 0.003994 | 0.00546 | 0.005419 | 5.5e-05 | 0.001033 | 0.34468 | 1.3574 | 2.5481 | 101.2 | 20.79 | 17.545 | 103.3328 | 118.1098 | 101.2239 | 0.004261 | 0.94303 | 0.022122 | 23.4583 | 0.23597 | 0.15886 | 0.064084 | 4.2395 | 4.2395 | 0.001559
2019-03-19 15:52 | 1553025142.7645 | 613 | 0.19716 | 0.007954 | 0.000204 | 98.617 | 95.9383 | 105.6338 | 21.2612 | 0.003472 | 0.003292 | 239.0691 | 0.26729 | 8.4035 | 0.003929 | 0.005717 | 0.005598 | 5.4e-05 | 0.001031 | 0.35723 | 1.3435 | 2.5479 | 101.2 | 20.79 | 17.564 | 105.4213 | 116.5219 | 101.0529 | 0.004266 | 0.94346 | 0.022102 | 23.4271 | 0.24632 | 0.15901 | 0.06387 | 4.2216 | 4.2216 | 0.001567
2019-03-19 15:52 | 1553025143.7639 | 614 | 0.1972 | 0.008176 | 0.000205 | 98.5026 | 95.6239 | 105.9639 | 21.2612 | 0.003439 | 0.003397 | 238.9029 | 0.26613 | 8.3994 | 0.004007 | 0.005638 | 0.005583 | 5.3e-05 | 0.001031 | 0.30215 | 1.3553 | 2.5482 | 101.19 | 20.79 | 17.573 | 104.6143 | 118.4427 | 103.589 | 0.004268 | 0.93711 | 0.022062 | 23.543 | 0.2653 | 0.15832 | 0.06336 | 4.1803 | 4.1803 | 0.001576
2019-03-19 15:52 | 1553025144.7647 | 615 | 0.19721 | 0.008372 | 0.000206 | 98.4102 | 95.6068 | 106.2793 | 21.2612 | 0.003519 | 0.00335 | 238.7454 | 0.26698 | 8.403 | 0.004008 | 0.005648 | 0.005447 | 5.4e-05 | 0.001025 | 0.32029 | 1.3276 | 2.548 | 101.19 | 20.79 | 17.58 | 106.5514 | 118.4576 | 102.4537 | 0.00427 | 0.93994 | 0.022129 | 23.5425 | 0.2585 | 0.15883 | 0.063102 | 4.1758 | 4.1758 | 0.001582
2019-03-19 15:52 | 1553025145.7647 | 616 | 0.19727 | 0.008536 | 0.000207 | 98.2974 | 95.7023 | 106.5779 | 21.2612 | 0.003543 | 0.00339 | 238.5665 | 0.26696 | 8.4021 | 0.003991 | 0.005553 | 0.005511 | 5.5e-05 | 0.00103 | 0.31188 | 1.3377 | 2.5481 | 101.19 | 20.79 | 17.593 | 107.1284 | 118.0472 | 103.4272 | 0.004273 | 0.93751 | 0.022131 | 23.6061 | 0.23931 | 0.15881 | 0.062539 | 4.1391 | 4.1391 | 0.001596
2019-03-19 15:52 | 1553025146.7651 | 617 | 0.19733 | 0.008659 | 0.000207 | 98.2409 | 95.7129 | 106.8016 | 21.2612 | 0.003643 | 0.00343 | 238.3968 | 0.26646 | 8.4041 | 0.004022 | 0.006241 | 0.006088 | 5.4e-05 | 0.001027 | 0.28407 | 1.3406 | 2.548 | 101.19 | 20.79 | 17.6 | 109.5648 | 118.7993 | 104.4003 | 0.004275 | 0.9351 | 0.022126 | 23.6613 | 0.23318 | 0.15852 | 0.062018 | 4.1037 | 4.1037 | 0.001607
2019-03-19 15:52 | 1553025147.7648 | 618 | 0.19734 | 0.008772 | 0.000208 | 98.2034 | 95.6773 | 107.1194 | 21.2612 | 0.003646 | 0.003411 | 238.2419 | 0.26675 | 8.401 | 0.004062 | 0.005766 | 0.005677 | 5.3e-05 | 0.001029 | 0.29623 | 1.3473 | 2.5481 | 101.19 | 20.78 | 17.609 | 109.6397 | 119.769 | 103.9263 | 0.004274 | 0.93627 | 0.022172 | 23.6816 | 0.23309 | 0.15869 | 0.061852 | 4.1014 | 4.1014 | 0.00161
2019-03-19 15:52 | 1553025148.7638 | 619 | 0.19736 | 0.008863 | 0.00021 | 98.0957 | 95.4359 | 107.3354 | 21.2612 | 0.003608 | 0.003487 | 238.1026 | 0.26708 | 8.403 | 0.004093 | 0.005339 | 0.005182 | 5.3e-05 | 0.001034 | 0.30829 | 1.3296 | 2.5481 | 101.19 | 20.78 | 17.612 | 108.704 | 120.5448 | 105.7848 | 0.004275 | 0.93168 | 0.02214 | 23.7638 | 0.24588 | 0.15889 | 0.061623 | 4.0802 | 4.0802 | 0.00162
2019-03-19 15:52 | 1553025149.764 | 620 | 0.19755 | 0.008939 | 0.000212 | 98.0765 | 95.3965 | 107.5431 | 21.2612 | 0.003639 | 0.003451 | 237.9704 | 0.26586 | 8.4018 | 0.004172 | 0.005576 | 0.005529 | 5.5e-05 | 0.001037 | 0.31045 | 1.3322 | 2.5481 | 101.19 | 20.78 | 17.614 | 109.4686 | 122.4557 | 104.8957 | 0.004276 | 0.93387 | 0.022188 | 23.7589 | 0.24782 | 0.15816 | 0.060393 | 4.0075 | 4.0075 | 0.001642
2019-03-19 15:52 | 1553025150.7649 | 621 | 0.19776 | 0.009001 | 0.000213 | 97.9585 | 95.4338 | 108.0247 | 21.2612 | 0.003695 | 0.003498 | 237.8324 | 0.2665 | 8.4002 | 0.004023 | 0.005733 | 0.005605 | 5e-05 | 0.001036 | 0.30757 | 1.3299 | 2.5482 | 101.19 | 20.78 | 17.615 | 110.8197 | 118.823 | 106.0429 | 0.004276 | 0.93104 | 0.022204 | 23.8481 | 0.23356 | 0.15854 | 0.059092 | 3.9242 | 3.9242 | 0.001681
2019-03-19 15:52 | 1553025151.7645 | 622 | 0.19786 | 0.009035 | 0.000215 | 97.8394 | 94.8197 | 108.2639 | 21.2612 | 0.003672 | 0.003487 | 237.707 | 0.26677 | 8.4035 | 0.004062 | 0.0057 | 0.005654 | 5.1e-05 | 0.001036 | 0.30359 | 1.3268 | 2.5481 | 101.19 | 20.78 | 17.612 | 110.2619 | 119.7744 | 105.7839 | 0.004275 | 0.93168 | 0.022236 | 23.8664 | 0.28041 | 0.1587 | 0.058434 | 3.8862 | 3.8862 | 0.001699
2019-03-19 15:53 | 1553025152.7641 | 623 | 0.19797 | 0.009047 | 0.000216 | 97.8295 | 94.9149 | 109.303 | 21.2612 | 0.003749 | 0.003459 | 237.5836 | 0.26692 | 8.4064 | 0.004065 | 0.005672 | 0.00559 | 5.3e-05 | 0.001042 | 0.33483 | 1.338 | 2.5481 | 101.19 | 20.78 | 17.603 | 112.1459 | 119.8458 | 105.0845 | 0.004273 | 0.9334 | 0.022363 | 23.9585 | 0.27053 | 0.15879 | 0.05775 | 3.8627 | 3.8627 | 0.00171
2019-03-19 15:53 | 1553025153.7652 | 624 | 0.19838 | 0.009031 | 0.000215 | 97.6899 | 94.9851 | 109.8784 | 21.2612 | 0.003719 | 0.003469 | 237.4706 | 0.26668 | 8.4045 | 0.004022 | 0.004963 | 0.004807 | 5.3e-05 | 0.001041 | 0.32316 | 1.3515 | 2.5482 | 101.19 | 20.78 | 17.596 | 111.3998 | 118.792 | 105.3372 | 0.004271 | 0.93278 | 0.022414 | 24.0295 | 0.25115 | 0.15865 | 0.055337 | 3.7103 | 3.7103 | 0.001779
2019-03-19 15:53 | 1553025154.7652 | 625 | 0.19891 | 0.008997 | 0.000214 | 97.6881 | 94.8646 | 110.6572 | 21.2612 | 0.003765 | 0.00348 | 237.3482 | 0.26652 | 8.4042 | 0.004082 | 0.005829 | 0.005784 | 5.3e-05 | 0.001038 | 0.28871 | 1.3267 | 2.5481 | 101.19 | 20.78 | 17.589 | 112.5185 | 120.2656 | 105.6163 | 0.00427 | 0.93209 | 0.022485 | 24.1234 | 0.26234 | 0.15855 | 0.05225 | 3.515 | 3.515 | 0.001876
2019-03-19 15:53 | 1553025155.7643 | 626 | 0.19921 | 0.008937 | 0.000211 | 97.6236 | 94.9372 | 111.1672 | 21.2612 | 0.00379 | 0.00345 | 237.2515 | 0.26623 | 8.404 | 0.004077 | 0.005735 | 0.005689 | 5.3e-05 | 0.001042 | 0.30165 | 1.3215 | 2.5479 | 101.19 | 20.78 | 17.573 | 113.1323 | 120.1497 | 104.8879 | 0.004266 | 0.93389 | 0.022559 | 24.1557 | 0.24959 | 0.15838 | 0.050507 | 3.4092 | 3.4092 | 0.001933
2019-03-19 15:53 | 1553025156.7652 | 627 | 0.19963 | 0.008849 | 0.000207 | 97.5456 | 94.9444 | 111.7531 | 21.2612 | 0.003725 | 0.003416 | 237.1567 | 0.26645 | 8.404 | 0.004085 | 0.005928 | 0.005739 | 5e-05 | 0.001044 | 0.31542 | 1.3514 | 2.548 | 101.19 | 20.78 | 17.564 | 111.5655 | 120.3424 | 104.0632 | 0.004264 | 0.93593 | 0.022643 | 24.1928 | 0.24176 | 0.15851 | 0.048099 | 3.2593 | 3.2593 | 0.002023
2019-03-19 15:53 | 1553025157.7657 | 628 | 0.20012 | 0.008733 | 0.000202 | 97.5075 | 95.2773 | 111.8653 | 21.2612 | 0.003734 | 0.00346 | 237.0742 | 0.26636 | 8.4029 | 0.004097 | 0.005342 | 0.005284 | 5.2e-05 | 0.001041 | 0.3079 | 1.3558 | 2.5482 | 101.19 | 20.78 | 17.541 | 111.7857 | 120.6258 | 105.1129 | 0.004258 | 0.93333 | 0.022623 | 24.2386 | 0.20696 | 0.15846 | 0.045318 | 3.0686 | 3.0686 | 0.002148
2019-03-19 15:53 | 1553025158.7644 | 629 | 0.20037 | 0.008585 | 0.000195 | 97.4117 | 95.1831 | 112.4256 | 21.2612 | 0.003777 | 0.003408 | 236.9867 | 0.26626 | 8.4034 | 0.004049 | 0.005434 | 0.005325 | 5.1e-05 | 0.001037 | 0.28622 | 1.29 | 2.5481 | 101.19 | 20.78 | 17.529 | 112.8335 | 119.4691 | 103.8555 | 0.004255 | 0.93645 | 0.022717 | 24.2588 | 0.20702 | 0.1584 | 0.043998 | 2.9921 | 2.9921 | 0.002202
2019-03-19 15:53 | 1553025159.7634 | 630 | 0.20112 | 0.008397 | 0.000188 | 97.3736 | 95.5749 | 113.1455 | 21.2612 | 0.003694 | 0.003389 | 236.8913 | 0.26639 | 8.4039 | 0.004139 | 0.005676 | 0.005479 | 5.1e-05 | 0.001029 | 0.29623 | 1.3143 | 2.548 | 101.19 | 20.78 | 17.529 | 110.7905 | 121.6575 | 103.4035 | 0.004255 | 0.93757 | 0.022803 | 24.3217 | 0.16677 | 0.15848 | 0.039785 | 2.7164 | 2.7164 | 0.002427
2019-03-19 15:53 | 1553025160.7635 | 631 | 0.20169 | 0.008181 | 0.000181 | 97.3448 | 95.5352 | 113.5922 | 21.2612 | 0.003709 | 0.003387 | 236.7992 | 0.26617 | 8.4025 | 0.004071 | 0.005758 | 0.005697 | 5.4e-05 | 0.001027 | 0.30525 | 1.3288 | 2.5481 | 101.19 | 20.78 | 17.506 | 111.1558 | 119.9839 | 103.3548 | 0.004249 | 0.93769 | 0.02285 | 24.3681 | 0.16784 | 0.15834 | 0.036645 | 2.5075 | 2.5075 | 0.002627
2019-03-19 15:53 | 1553025161.7639 | 632 | 0.20223 | 0.007867 | 0.000174 | 97.1816 | 95.7571 | 114.0956 | 21.2612 | 0.00363 | 0.003398 | 236.7217 | 0.26632 | 8.4057 | 0.004005 | 0.005759 | 0.005539 | 5.1e-05 | 0.001021 | 0.30613 | 1.3794 | 2.548 | 101.19 | 20.78 | 17.484 | 109.2462 | 118.3855 | 103.6174 | 0.004244 | 0.93704 | 0.022892 | 24.4306 | 0.13208 | 0.15843 | 0.033781 | 2.3161 | 2.3161 | 0.002846
2019-03-19 15:53 | 1553025162.7636 | 633 | 0.20256 | 0.007456 | 0.000167 | 97.2245 | 95.7215 | 114.4684 | 21.2612 | 0.003666 | 0.003385 | 236.6433 | 0.26627 | 8.4026 | 0.00398 | 0.00531 | 0.005318 | 5.1e-05 | 0.001018 | 0.30829 | 1.4025 | 2.548 | 101.19 | 20.78 | 17.473 | 110.1097 | 117.7808 | 103.303 | 0.004241 | 0.93782 | 0.022939 | 24.4603 | 0.13935 | 0.1584 | 0.032299 | 2.2193 | 2.2193 | 0.002969
2019-03-19 15:53 | 1553025163.7647 | 634 | 0.20285 | 0.007124 | 0.000161 | 97.1167 | 95.5595 | 114.7796 | 21.2612 | 0.003665 | 0.003362 | 236.5795 | 0.26653 | 8.402 | 0.004 | 0.005698 | 0.00561 | 5.6e-05 | 0.001023 | 0.28783 | 1.4141 | 2.5482 | 101.19 | 20.78 | 17.463 | 110.0883 | 118.2517 | 102.7474 | 0.004239 | 0.93921 | 0.022987 | 24.4754 | 0.14459 | 0.15856 | 0.030994 | 2.1344 | 2.1344 | 0.00309
2019-03-19 15:53 | 1553025164.7647 | 635 | 0.20309 | 0.006746 | 0.000155 | 97.0915 | 95.8927 | 115.0035 | 21.2612 | 0.003651 | 0.003298 | 236.5187 | 0.26625 | 8.4044 | 0.003957 | 0.005873 | 0.005768 | 5.5e-05 | 0.001025 | 0.2891 | 1.3561 | 2.5481 | 101.19 | 20.78 | 17.446 | 109.7675 | 117.2193 | 101.198 | 0.004235 | 0.94309 | 0.023057 | 24.4487 | 0.11113 | 0.15839 | 0.030045 | 2.0756 | 2.0756 | 0.003175
2019-03-19 15:53 | 1553025165.7652 | 636 | 0.20336 | 0.006513 | 0.00015 | 97.0315 | 95.9126 | 114.9045 | 21.2612 | 0.003643 | 0.003375 | 236.431 | 0.26628 | 8.402 | 0.003888 | 0.005956 | 0.00587 | 5e-05 | 0.001027 | 0.31725 | 1.295 | 2.5482 | 101.19 | 20.78 | 17.44 | 109.5648 | 115.5389 | 103.0666 | 0.004233 | 0.93841 | 0.02299 | 24.4991 | 0.10374 | 0.15841 | 0.028675 | 1.9754 | 1.9754 | 0.003336
2019-03-19 15:53 | 1553025166.7649 | 637 | 0.20357 | 0.00628 | 0.000144 | 96.9202 | 95.816 | 115.2754 | 21.2612 | 0.003661 | 0.003292 | 236.337 | 0.2671 | 8.4028 | 0.003897 | 0.005655 | 0.005526 | 5.3e-05 | 0.001031 | 0.29535 | 1.2932 | 2.5481 | 101.19 | 20.78 | 17.436 | 109.9984 | 115.7374 | 101.055 | 0.004232 | 0.94345 | 0.023089 | 24.4729 | 0.10249 | 0.1589 | 0.027718 | 1.9179 | 1.9179 | 0.003447
2019-03-19 15:53 | 1553025167.7638 | 638 | 0.20383 | 0.006053 | 0.00014 | 96.8184 | 95.6842 | 115.426 | 21.2612 | 0.003713 | 0.003368 | 236.2591 | 0.26634 | 8.4015 | 0.003875 | 0.005486 | 0.005441 | 5.4e-05 | 0.001034 | 0.30254 | 1.3298 | 2.5481 | 101.19 | 20.78 | 17.433 | 111.2743 | 115.209 | 102.8811 | 0.004232 | 0.93887 | 0.023048 | 24.5486 | 0.10541 | 0.15845 | 0.026463 | 1.8279 | 1.8279 | 0.003606
2019-03-19 15:53 | 1553025168.7649 | 639 | 0.20401 | 0.005824 | 0.000137 | 96.7728 | 95.6044 | 115.5359 | 21.2612 | 0.003656 | 0.003286 | 236.17 | 0.2663 | 8.4048 | 0.003835 | 0.005801 | 0.00566 | 5.3e-05 | 0.001036 | 0.31797 | 1.2785 | 2.548 | 101.19 | 20.78 | 17.433 | 109.8888 | 114.226 | 100.9169 | 0.004232 | 0.9438 | 0.023119 | 24.496 | 0.10866 | 0.15842 | 0.02568 | 1.7794 | 1.7794 | 0.003704
2019-03-19 15:53 | 1553025169.7651 | 640 | 0.20403 | 0.005602 | 0.000133 | 96.5297 | 95.3934 | 115.6767 | 21.2612 | 0.003658 | 0.003295 | 236.0733 | 0.26663 | 8.4026 | 0.003865 | 0.005091 | 0.005016 | 5.5e-05 | 0.001038 | 0.29031 | 1.341 | 2.5481 | 101.19 | 20.79 | 17.437 | 109.9189 | 114.9742 | 101.1337 | 0.004235 | 0.94326 | 0.023127 | 24.5181 | 0.10592 | 0.15862 | 0.025845 | 1.7917 | 1.7917 | 0.003683
2019-03-19 15:53 | 1553025170.765 | 641 | 0.20406 | 0.005405 | 0.000129 | 96.5541 | 95.1954 | 115.7478 | 21.2612 | 0.003694 | 0.003299 | 236.0009 | 0.26692 | 8.4031 | 0.003824 | 0.005477 | 0.005311 | 5.7e-05 | 0.001041 | 0.28423 | 1.3134 | 2.548 | 101.19 | 20.78 | 17.441 | 110.8071 | 113.9707 | 101.2239 | 0.004234 | 0.94303 | 0.023131 | 24.5286 | 0.12675 | 0.15879 | 0.025861 | 1.7933 | 1.7933 | 0.003683
2019-03-19 15:53 | 1553025171.7651 | 642 | 0.20419 | 0.005213 | 0.000126 | 96.4502 | 95.0122 | 115.7542 | 21.2612 | 0.003665 | 0.003264 | 235.9273 | 0.26631 | 8.4032 | 0.0038 | 0.00501 | 0.004933 | 5.1e-05 | 0.001041 | 0.32676 | 1.3037 | 2.5479 | 101.19 | 20.79 | 17.445 | 110.0947 | 113.3907 | 100.3857 | 0.004237 | 0.94515 | 0.023158 | 24.5018 | 0.13436 | 0.15843 | 0.025356 | 1.7605 | 1.7605 | 0.003744
2019-03-19 15:53 | 1553025172.7651 | 643 | 0.20421 | 0.005043 | 0.000124 | 96.3239 | 95.1502 | 115.5685 | 21.2612 | 0.003596 | 0.003246 | 235.852 | 0.26638 | 8.4026 | 0.003811 | 0.005047 | 0.004938 | 5.2e-05 | 0.001029 | 0.3235 | 1.3516 | 2.548 | 101.19 | 20.79 | 17.454 | 108.4242 | 113.6474 | 99.9417 | 0.004239 | 0.94627 | 0.023153 | 24.4675 | 0.10965 | 0.15847 | 0.025435 | 1.7657 | 1.7657 | 0.003734
2019-03-19 15:53 | 1553025173.7653 | 644 | 0.20427 | 0.004888 | 0.000122 | 96.278 | 95.0814 | 115.2824 | 21.2612 | 0.003604 | 0.003235 | 235.7647 | 0.26646 | 8.4049 | 0.003701 | 0.005693 | 0.005537 | 5.4e-05 | 0.001028 | 0.30901 | 1.3192 | 2.5481 | 101.19 | 20.79 | 17.458 | 108.6267 | 110.9751 | 99.682 | 0.00424 | 0.94693 | 0.023132 | 24.4287 | 0.11186 | 0.15852 | 0.025276 | 1.7532 | 1.7532 | 0.003761
2019-03-19 15:53 | 1553025174.7633 | 645 | 0.20435 | 0.004749 | 0.00012 | 96.3443 | 95.1523 | 115.3784 | 21.2612 | 0.003612 | 0.003272 | 235.6841 | 0.26592 | 8.4005 | 0.00377 | 0.005373 | 0.005346 | 5.3e-05 | 0.00103 | 0.31365 | 1.3461 | 2.5481 | 101.19 | 20.79 | 17.468 | 108.8055 | 112.6473 | 100.5685 | 0.004243 | 0.94468 | 0.023114 | 24.4679 | 0.11135 | 0.15819 | 0.02495 | 1.7293 | 1.7293 | 0.003805
2019-03-19 15:53 | 1553025175.7644 | 646 | 0.20442 | 0.004627 | 0.000118 | 96.3077 | 94.9975 | 115.6169 | 21.2612 | 0.003626 | 0.003212 | 235.6147 | 0.26619 | 8.404 | 0.003727 | 0.005237 | 0.005127 | 5.3e-05 | 0.001029 | 0.31437 | 1.2829 | 2.548 | 101.19 | 20.79 | 17.474 | 109.1405 | 111.6168 | 99.1281 | 0.004244 | 0.94834 | 0.023183 | 24.446 | 0.12252 | 0.15835 | 0.024672 | 1.7152 | 1.7152 | 0.003841
2019-03-19 15:53 | 1553025176.7648 | 647 | 0.20441 | 0.004528 | 0.000117 | 96.2975 | 95.2235 | 115.4364 | 21.2612 | 0.003618 | 0.003181 | 235.5585 | 0.26602 | 8.402 | 0.003647 | 0.005354 | 0.005249 | 5.3e-05 | 0.001033 | 0.30215 | 1.3121 | 2.548 | 101.19 | 20.79 | 17.48 | 108.9656 | 109.6674 | 98.3588 | 0.004246 | 0.9503 | 0.023189 | 24.4016 | 0.10031 | 0.15826 | 0.024855 | 1.7284 | 1.7284 | 0.003809
2019-03-19 15:53 | 1553025177.7634 | 648 | 0.20445 | 0.004424 | 0.000116 | 96.2489 | 95.0417 | 115.7978 | 21.2612 | 0.003608 | 0.003191 | 235.4955 | 0.26597 | 8.4024 | 0.003681 | 0.00475 | 0.004611 | 5.1e-05 | 0.001033 | 0.3115 | 1.3343 | 2.548 | 101.19 | 20.79 | 17.491 | 108.7054 | 110.4756 | 98.6169 | 0.004248 | 0.94964 | 0.023217 | 24.4483 | 0.1129 | 0.15822 | 0.024731 | 1.722 | 1.722 | 0.003822
2019-03-19 15:53 | 1553025178.7645 | 649 | 0.2045 | 0.004343 | 0.000114 | 96.2681 | 95.0417 | 115.8164 | 21.2612 | 0.00364 | 0.003276 | 235.4391 | 0.26634 | 8.4026 | 0.003654 | 0.004973 | 0.004804 | 5.4e-05 | 0.001034 | 0.2907 | 1.3034 | 2.548 | 101.19 | 20.79 | 17.496 | 109.4985 | 109.8346 | 100.6573 | 0.00425 | 0.94446 | 0.023156 | 24.5172 | 0.11468 | 0.15845 | 0.024569 | 1.7062 | 1.7062 | 0.003863
2019-03-19 15:53 | 1553025179.7645 | 650 | 0.20452 | 0.004265 | 0.000113 | 96.2047 | 95.0889 | 115.6852 | 21.2612 | 0.003557 | 0.003212 | 235.3941 | 0.26585 | 8.404 | 0.003701 | 0.005305 | 0.005169 | 5e-05 | 0.001036 | 0.31094 | 1.3249 | 2.5481 | 101.19 | 20.79 | 17.508 | 107.4789 | 110.9839 | 99.1257 | 0.004253 | 0.94834 | 0.02319 | 24.4531 | 0.10435 | 0.15815 | 0.024524 | 1.7057 | 1.7057 | 0.003857
2019-03-19 15:53 | 1553025180.764 | 651 | 0.2045 | 0.004201 | 0.000112 | 96.136 | 95.0232 | 115.8964 | 21.2612 | 0.003512 | 0.00318 | 235.3316 | 0.26632 | 8.4021 | 0.003648 | 0.004888 | 0.004906 | 5e-05 | 0.001038 | 0.3115 | 1.351 | 2.5481 | 101.19 | 20.79 | 17.515 | 106.3859 | 109.6807 | 98.3472 | 0.004254 | 0.95033 | 0.023235 | 24.4498 | 0.10414 | 0.15844 | 0.024749 | 1.7248 | 1.7248 | 0.003821
2019-03-19 15:53 | 1553025181.7636 | 652 | 0.20451 | 0.004146 | 0.00011 | 96.0985 | 94.8786 | 116.2243 | 21.2612 | 0.003594 | 0.003161 | 235.2772 | 0.2663 | 8.4023 | 0.003609 | 0.004994 | 0.004864 | 5.1e-05 | 0.00104 | 0.34578 | 1.3508 | 2.5481 | 101.19 | 20.79 | 17.523 | 108.3768 | 108.7409 | 97.8744 | 0.004256 | 0.95154 | 0.023283 | 24.4688 | 0.11427 | 0.15842 | 0.024778 | 1.7305 | 1.7305 | 0.003808
2019-03-19 15:53 | 1553025182.7643 | 653 | 0.20457 | 0.004091 | 0.000109 | 96.0247 | 94.6505 | 116.1964 | 21.2612 | 0.003602 | 0.003199 | 235.1902 | 0.26658 | 8.4039 | 0.003595 | 0.005383 | 0.005265 | 5e-05 | 0.001042 | 0.3183 | 1.341 | 2.5481 | 101.19 | 20.79 | 17.54 | 108.5692 | 108.3894 | 98.7972 | 0.00426 | 0.94918 | 0.023251 | 24.4963 | 0.12892 | 0.15859 | 0.024477 | 1.7072 | 1.7072 | 0.003864
2019-03-19 15:53 | 1553025183.7648 | 654 | 0.20458 | 0.004047 | 0.000107 | 95.9103 | 94.742 | 116.1948 | 21.2612 | 0.003552 | 0.003188 | 235.1388 | 0.26553 | 8.4033 | 0.003605 | 0.004689 | 0.004521 | 5.1e-05 | 0.001044 | 0.30071 | 1.3203 | 2.5482 | 101.19 | 20.79 | 17.549 | 107.352 | 108.6292 | 98.5376 | 0.004263 | 0.94984 | 0.023259 | 24.4875 | 0.10963 | 0.15796 | 0.024456 | 1.7064 | 1.7064 | 0.003851
2019-03-19 15:53 | 1553025184.7641 | 655 | 0.20461 | 0.004006 | 0.000106 | 95.7648 | 94.5468 | 115.8833 | 21.2612 | 0.003535 | 0.003162 | 235.0714 | 0.26589 | 8.4038 | 0.003674 | 0.005247 | 0.005182 | 4.8e-05 | 0.001039 | 0.30989 | 1.3038 | 2.5481 | 101.19 | 20.79 | 17.558 | 106.9466 | 110.3115 | 97.9072 | 0.004265 | 0.95146 | 0.023248 | 24.4339 | 0.11449 | 0.15818 | 0.024308 | 1.6953 | 1.6953 | 0.003881
2019-03-19 15:53 | 1553025185.7648 | 656 | 0.20465 | 0.003968 | 0.000105 | 95.7345 | 94.3056 | 115.792 | 21.2612 | 0.003562 | 0.003136 | 235.0044 | 0.26634 | 8.4055 | 0.003665 | 0.004724 | 0.004643 | 4.9e-05 | 0.001039 | 0.32477 | 1.3427 | 2.5481 | 101.19 | 20.79 | 17.577 | 107.6017 | 110.1034 | 97.2804 | 0.004269 | 0.95307 | 0.023258 | 24.4037 | 0.1345 | 0.15845 | 0.024139 | 1.6843 | 1.6843 | 0.003913
2019-03-19 15:53 | 1553025186.7638 | 657 | 0.20469 | 0.00394 | 0.000103 | 95.7087 | 94.4728 | 115.8727 | 21.2612 | 0.003513 | 0.003141 | 234.9391 | 0.26645 | 8.4011 | 0.003599 | 0.00525 | 0.005202 | 4.9e-05 | 0.001036 | 0.30862 | 1.3311 | 2.548 | 101.19 | 20.79 | 17.588 | 106.3978 | 108.5031 | 97.3995 | 0.004272 | 0.95276 | 0.023263 | 24.4161 | 0.11625 | 0.15851 | 0.023906 | 1.6684 | 1.6684 | 0.003952
2019-03-19 15:53 | 1553025187.7654 | 658 | 0.20472 | 0.003912 | 0.000102 | 95.5886 | 94.4426 | 115.9661 | 21.2612 | 0.003499 | 0.003134 | 234.8676 | 0.26597 | 8.4037 | 0.003536 | 0.004909 | 0.004766 | 4.8e-05 | 0.001035 | 0.32139 | 1.3119 | 2.5482 | 101.19 | 20.79 | 17.604 | 106.0689 | 106.9726 | 97.2224 | 0.004276 | 0.95322 | 0.023278 | 24.4201 | 0.10788 | 0.15822 | 0.023766 | 1.6598 | 1.6598 | 0.003966
2019-03-19 15:53 | 1553025188.764 | 659 | 0.20473 | 0.003879 | 0.000102 | 95.4992 | 94.294 | 115.9497 | 21.2612 | 0.003579 | 0.003166 | 234.8065 | 0.26666 | 8.4026 | 0.003683 | 0.006065 | 0.005974 | 4.7e-05 | 0.001027 | 0.30127 | 1.3591 | 2.548 | 101.19 | 20.79 | 17.611 | 107.9963 | 110.5262 | 98.0029 | 0.004278 | 0.95121 | 0.023252 | 24.4441 | 0.1136 | 0.15864 | 0.023773 | 1.6584 | 1.6584 | 0.003979
2019-03-19 15:53 | 1553025189.764 | 660 | 0.20479 | 0.003848 | 0.000102 | 95.402 | 94.2779 | 116.0795 | 21.2612 | 0.003529 | 0.003162 | 234.7448 | 0.26596 | 8.4041 | 0.003564 | 0.005542 | 0.005455 | 5.1e-05 | 0.00103 | 0.32947 | 1.3667 | 2.548 | 101.19 | 20.79 | 17.615 | 106.7828 | 107.6522 | 97.9019 | 0.004279 | 0.95147 | 0.023268 | 24.4545 | 0.10601 | 0.15822 | 0.023474 | 1.6386 | 1.6386 | 0.004017
2019-03-19 15:53 | 1553025190.7634 | 661 | 0.20479 | 0.003817 | 0.000102 | 95.4047 | 94.0008 | 115.7467 | 21.2612 | 0.003459 | 0.003116 | 234.6613 | 0.26646 | 8.4017 | 0.003634 | 0.00531 | 0.005175 | 4.8e-05 | 0.001028 | 0.34722 | 1.3548 | 2.5482 | 101.19 | 20.79 | 17.622 | 105.1003 | 109.3432 | 96.8001 | 0.00428 | 0.95431 | 0.023269 | 24.3831 | 0.13259 | 0.15852 | 0.023485 | 1.6395 | 1.6395 | 0.004022
2019-03-19 15:53 | 1553025191.7645 | 662 | 0.20479 | 0.003782 | 0.000101 | 95.3293 | 93.8258 | 115.4249 | 21.2612 | 0.003477 | 0.003179 | 234.5757 | 0.26619 | 8.4034 | 0.00357 | 0.005055 | 0.004952 | 5.2e-05 | 0.001028 | 0.28854 | 1.3915 | 2.5481 | 101.19 | 20.79 | 17.623 | 105.5244 | 107.794 | 98.3173 | 0.00428 | 0.95041 | 0.023189 | 24.399 | 0.1422 | 0.15835 | 0.023533 | 1.6372 | 1.6372 | 0.004024
2019-03-19 15:53 | 1553025192.7633 | 663 | 0.20477 | 0.003761 | 0.000101 | 95.2762 | 93.8303 | 115.5089 | 21.2612 | 0.003586 | 0.003115 | 234.5369 | 0.26601 | 8.4039 | 0.00357 | 0.005441 | 0.005302 | 4.6e-05 | 0.001032 | 0.30901 | 1.3358 | 2.5481 | 101.19 | 20.79 | 17.62 | 108.1886 | 107.7836 | 96.7718 | 0.00428 | 0.95438 | 0.023246 | 24.3571 | 0.13679 | 0.15825 | 0.023673 | 1.6511 | 1.6511 | 0.003987
2019-03-19 15:53 | 1553025193.7633 | 664 | 0.20485 | 0.003731 | 0.000101 | 95.1697 | 93.6138 | 115.7786 | 21.2612 | 0.003503 | 0.003135 | 234.4642 | 0.26673 | 8.4023 | 0.003477 | 0.005382 | 0.005276 | 5e-05 | 0.001038 | 0.31868 | 1.3637 | 2.5481 | 101.19 | 20.8 | 17.617 | 106.1744 | 105.5392 | 97.2608 | 0.004282 | 0.95312 | 0.023258 | 24.4016 | 0.14744 | 0.15868 | 0.023245 | 1.622 | 1.622 | 0.00407
2019-03-19 15:53 | 1553025194.7647 | 665 | 0.20476 | 0.003709 | 0.000101 | 95.1094 | 93.7614 | 115.9598 | 21.2612 | 0.003496 | 0.00311 | 234.3893 | 0.26664 | 8.4017 | 0.003522 | 0.005109 | 0.005018 | 5.1e-05 | 0.001038 | 0.33644 | 1.3616 | 2.5481 | 101.19 | 20.8 | 17.612 | 105.9956 | 106.6229 | 96.6496 | 0.00428 | 0.95469 | 0.023295 | 24.4006 | 0.12767 | 0.15863 | 0.023786 | 1.6625 | 1.6625 | 0.003969
2019-03-19 15:53 | 1553025195.7634 | 666 | 0.20481 | 0.003698 | 0.000102 | 94.9926 | 93.3487 | 116.1835 | 21.2612 | 0.003518 | 0.003094 | 234.3156 | 0.26654 | 8.4031 | 0.003504 | 0.004933 | 0.004889 | 5.2e-05 | 0.00104 | 0.30879 | 1.3939 | 2.5482 | 101.19 | 20.8 | 17.602 | 106.5307 | 106.1766 | 96.2595 | 0.004278 | 0.9557 | 0.02333 | 24.4112 | 0.15615 | 0.15856 | 0.023505 | 1.6452 | 1.6452 | 0.004009
2019-03-19 15:53 | 1553025196.7644 | 667 | 0.20461 | 0.003681 | 0.000102 | 94.9236 | 93.41 | 116.061 | 21.2612 | 0.003456 | 0.003177 | 234.2272 | 0.2666 | 8.4044 | 0.003535 | 0.004862 | 0.004788 | 4.9e-05 | 0.001046 | 0.32908 | 1.3978 | 2.5481 | 101.19 | 20.8 | 17.598 | 105.0319 | 106.9323 | 98.2813 | 0.004277 | 0.9505 | 0.023254 | 24.465 | 0.14377 | 0.1586 | 0.024734 | 1.7257 | 1.7257 | 0.003823
2019-03-19 15:53 | 1553025197.7645 | 668 | 0.2046 | 0.003675 | 0.000101 | 94.8486 | 93.3806 | 115.6589 | 21.2612 | 0.003446 | 0.003156 | 234.1467 | 0.26548 | 8.4029 | 0.00348 | 0.004808 | 0.004718 | 5e-05 | 0.001049 | 0.32709 | 1.3852 | 2.5481 | 101.19 | 20.8 | 17.594 | 104.7772 | 105.5998 | 97.7553 | 0.004276 | 0.95185 | 0.02323 | 24.4053 | 0.13952 | 0.15793 | 0.024817 | 1.7298 | 1.7298 | 0.003798
2019-03-19 15:53 | 1553025198.7649 | 669 | 0.20455 | 0.003675 | 0.0001 | 94.7123 | 93.2494 | 115.6425 | 21.2612 | 0.003471 | 0.003073 | 234.0715 | 0.26673 | 8.4051 | 0.003537 | 0.005558 | 0.00542 | 5.2e-05 | 0.001053 | 0.31725 | 1.3307 | 2.5481 | 101.19 | 20.8 | 17.587 | 105.3874 | 106.9804 | 95.7692 | 0.004274 | 0.95697 | 0.023291 | 24.3381 | 0.13923 | 0.15868 | 0.025123 | 1.7557 | 1.7557 | 0.00376
2019-03-19 15:53 | 1553025199.7645 | 670 | 0.20469 | 0.003678 | 9.9e-05 | 94.5914 | 93.1172 | 115.3636 | 21.2612 | 0.003431 | 0.00315 | 233.9898 | 0.26658 | 8.4009 | 0.00351 | 0.005087 | 0.004956 | 5.1e-05 | 0.001053 | 0.28744 | 1.3367 | 2.5479 | 101.19 | 20.8 | 17.584 | 104.4216 | 106.3277 | 97.6244 | 0.004274 | 0.95218 | 0.023205 | 24.3698 | 0.1405 | 0.15859 | 0.024264 | 1.6894 | 1.6894 | 0.003905
2019-03-19 15:53 | 1553025200.7647 | 671 | 0.20472 | 0.003685 | 9.9e-05 | 94.483 | 93.0415 | 115.7009 | 21.2612 | 0.003413 | 0.003094 | 233.9307 | 0.26657 | 8.4021 | 0.003547 | 0.005202 | 0.005159 | 5.1e-05 | 0.001053 | 0.27815 | 1.3115 | 2.5479 | 101.19 | 20.8 | 17.582 | 103.9932 | 107.2206 | 96.2706 | 0.004273 | 0.95567 | 0.023281 | 24.3608 | 0.13752 | 0.15858 | 0.024081 | 1.6822 | 1.6822 | 0.003922
2019-03-19 15:53 | 1553025201.7637 | 672 | 0.20471 | 0.003686 | 9.8e-05 | 94.3342 | 92.9268 | 116.0532 | 21.2612 | 0.003496 | 0.00317 | 233.8653 | 0.26636 | 8.4028 | 0.00355 | 0.004836 | 0.004716 | 5e-05 | 0.001045 | 0.30397 | 1.2947 | 2.548 | 101.19 | 20.8 | 17.582 | 105.9871 | 107.3049 | 98.1034 | 0.004273 | 0.95095 | 0.023259 | 24.4583 | 0.13445 | 0.15846 | 0.024134 | 1.6844 | 1.6844 | 0.003914
2019-03-19 15:53 | 1553025202.7641 | 673 | 0.20477 | 0.003684 | 9.8e-05 | 94.2145 | 92.9898 | 116.1896 | 21.2612 | 0.003425 | 0.003101 | 233.7772 | 0.26615 | 8.4019 | 0.003464 | 0.004991 | 0.00487 | 5.2e-05 | 0.00104 | 0.30541 | 1.3209 | 2.5482 | 101.19 | 20.8 | 17.585 | 104.2667 | 105.2286 | 96.4451 | 0.004274 | 0.95522 | 0.023325 | 24.418 | 0.11703 | 0.15833 | 0.023775 | 1.664 | 1.664 | 0.003958
2019-03-19 15:53 | 1553025203.7652 | 674 | 0.20476 | 0.003684 | 9.8e-05 | 94.0992 | 92.5237 | 116.1962 | 21.2612 | 0.003479 | 0.003105 | 233.722 | 0.26744 | 8.4036 | 0.003535 | 0.004974 | 0.004927 | 4.9e-05 | 0.001038 | 0.30486 | 1.323 | 2.548 | 101.19 | 20.8 | 17.591 | 105.577 | 106.9299 | 96.5299 | 0.004275 | 0.955 | 0.023323 | 24.4215 | 0.15103 | 0.1591 | 0.023816 | 1.6667 | 1.6667 | 0.003971
2019-03-19 15:53 | 1553025204.7638 | 675 | 0.20477 | 0.003677 | 9.7e-05 | 93.9678 | 92.4418 | 116.387 | 21.2612 | 0.003495 | 0.003084 | 233.6567 | 0.26589 | 8.403 | 0.003561 | 0.005133 | 0.005016 | 5.4e-05 | 0.001036 | 0.29878 | 1.314 | 2.5482 | 101.19 | 20.8 | 17.594 | 105.965 | 107.5639 | 96.0347 | 0.004276 | 0.95628 | 0.023357 | 24.4251 | 0.14644 | 0.15818 | 0.023814 | 1.6691 | 1.6691 | 0.003942
2019-03-19 15:53 | 1553025205.7639 | 676 | 0.2048 | 0.003664 | 9.6e-05 | 93.7601 | 92.3963 | 116.4844 | 21.2612 | 0.003504 | 0.003115 | 233.6068 | 0.26576 | 8.4013 | 0.00359 | 0.00545 | 0.005289 | 5.4e-05 | 0.001033 | 0.32549 | 1.3343 | 2.5482 | 101.19 | 20.8 | 17.604 | 106.1758 | 108.2831 | 96.7788 | 0.004279 | 0.95436 | 0.023344 | 24.46 | 0.13106 | 0.1581 | 0.023647 | 1.6565 | 1.6565 | 0.00397
2019-03-19 15:53 | 1553025206.7643 | 677 | 0.20478 | 0.003659 | 9.6e-05 | 93.6674 | 92.062 | 116.6816 | 21.2612 | 0.003422 | 0.003078 | 233.5839 | 0.26587 | 8.4015 | 0.003614 | 0.004892 | 0.004846 | 5.6e-05 | 0.001036 | 0.30055 | 1.3332 | 2.5479 | 101.19 | 20.8 | 17.607 | 104.188 | 108.8655 | 95.8749 | 0.004279 | 0.9567 | 0.023392 | 24.4507 | 0.15463 | 0.15817 | 0.023735 | 1.6661 | 1.6661 | 0.003949
2019-03-19 15:53 | 1553025207.7643 | 678 | 0.20476 | 0.003649 | 9.6e-05 | 93.5341 | 92.051 | 117.0632 | 21.2612 | 0.003373 | 0.003092 | 233.5255 | 0.26619 | 8.4037 | 0.00358 | 0.005021 | 0.004884 | 5.4e-05 | 0.001032 | 0.33572 | 1.3582 | 2.548 | 101.19 | 20.8 | 17.611 | 103.024 | 108.0295 | 96.2172 | 0.00428 | 0.95581 | 0.023419 | 24.502 | 0.14296 | 0.15835 | 0.023905 | 1.68 | 1.68 | 0.003921
2019-03-19 15:53 | 1553025208.7641 | 679 | 0.20478 | 0.003643 | 9.6e-05 | 93.2901 | 91.7226 | 117.0093 | 21.2612 | 0.003497 | 0.003116 | 233.4486 | 0.26694 | 8.4047 | 0.003495 | 0.005032 | 0.005004 | 5.6e-05 | 0.001018 | 0.30901 | 1.3388 | 2.5481 | 101.19 | 20.8 | 17.621 | 106.0247 | 105.9728 | 96.8071 | 0.004283 | 0.95429 | 0.023395 | 24.5159 | 0.15157 | 0.1588 | 0.023797 | 1.6707 | 1.6707 | 0.003954
2019-03-19 15:53 | 1553025209.7634 | 680 | 0.20475 | 0.003639 | 9.5e-05 | 93.2661 | 91.763 | 117.1872 | 21.2612 | 0.003441 | 0.00312 | 233.4 | 0.26546 | 8.3998 | 0.003548 | 0.005518 | 0.005433 | 5.6e-05 | 0.00101 | 0.30237 | 1.3874 | 2.5481 | 101.19 | 20.8 | 17.624 | 104.6521 | 107.2568 | 96.8901 | 0.004283 | 0.95407 | 0.02341 | 24.5373 | 0.14533 | 0.15792 | 0.023953 | 1.6828 | 1.6828 | 0.003904
2019-03-19 15:53 | 1553025210.7642 | 681 | 0.20476 | 0.003639 | 9.5e-05 | 93.1152 | 91.5174 | 116.9562 | 21.2612 | 0.003475 | 0.003116 | 233.3144 | 0.26617 | 8.4048 | 0.00358 | 0.004757 | 0.004609 | 5.4e-05 | 0.001008 | 0.31078 | 1.3791 | 2.5482 | 101.19 | 20.8 | 17.629 | 105.4827 | 108.038 | 96.793 | 0.004285 | 0.95432 | 0.02339 | 24.5099 | 0.15481 | 0.15834 | 0.023915 | 1.6787 | 1.6787 | 0.003924
2019-03-19 15:54 | 1553025211.7643 | 682 | 0.20477 | 0.003631 | 9.4e-05 | 93.0188 | 91.5342 | 116.741 | 21.2612 | 0.003496 | 0.003065 | 233.2278 | 0.26582 | 8.4028 | 0.003592 | 0.005225 | 0.005074 | 5.6e-05 | 0.001009 | 0.31061 | 1.3951 | 2.5482 | 101.19 | 20.8 | 17.63 | 105.9854 | 108.3328 | 95.5751 | 0.004285 | 0.95748 | 0.023407 | 24.447 | 0.14391 | 0.15814 | 0.023834 | 1.6742 | 1.6742 | 0.003929
2019-03-19 15:54 | 1553025212.7633 | 683 | 0.20475 | 0.003632 | 9.3e-05 | 92.9324 | 91.3342 | 116.4893 | 21.2612 | 0.003416 | 0.003102 | 233.175 | 0.26645 | 8.4051 | 0.003679 | 0.00521 | 0.00506 | 5.8e-05 | 0.001014 | 0.29535 | 1.3839 | 2.5481 | 101.19 | 20.8 | 17.629 | 104.052 | 110.4482 | 96.4584 | 0.004285 | 0.95519 | 0.023354 | 24.4499 | 0.15516 | 0.15851 | 0.023967 | 1.6799 | 1.6799 | 0.003925
2019-03-19 15:54 | 1553025213.7648 | 684 | 0.20477 | 0.003629 | 9.2e-05 | 92.7803 | 91.4904 | 116.5602 | 21.2612 | 0.003436 | 0.003115 | 233.0853 | 0.26677 | 8.4029 | 0.003673 | 0.005347 | 0.005311 | 5.5e-05 | 0.001012 | 0.31078 | 1.4056 | 2.5481 | 101.19 | 20.8 | 17.622 | 104.5309 | 110.2872 | 96.7836 | 0.004283 | 0.95435 | 0.023351 | 24.4681 | 0.12523 | 0.1587 | 0.023845 | 1.6711 | 1.6711 | 0.003951
2019-03-19 15:54 | 1553025214.7644 | 685 | 0.20477 | 0.003624 | 9.1e-05 | 92.6362 | 91.1506 | 116.2757 | 21.2612 | 0.003478 | 0.003122 | 233.0253 | 0.26714 | 8.4031 | 0.003674 | 0.004279 | 0.00417 | 5.4e-05 | 0.001009 | 0.32549 | 1.3831 | 2.5481 | 101.19 | 20.8 | 17.617 | 105.5511 | 110.3195 | 96.9372 | 0.004282 | 0.95395 | 0.023318 | 24.4433 | 0.14461 | 0.15892 | 0.023854 | 1.6694 | 1.6694 | 0.00396
2019-03-19 15:54 | 1553025215.7635 | 686 | 0.2048 | 0.003624 | 9.1e-05 | 92.412 | 90.9194 | 116.1089 | 21.2612 | 0.003447 | 0.003058 | 232.9487 | 0.26678 | 8.402 | 0.0036 | 0.005211 | 0.005121 | 5.4e-05 | 0.001004 | 0.31708 | 1.2983 | 2.5481 | 101.19 | 20.8 | 17.612 | 104.7936 | 108.5084 | 95.4004 | 0.00428 | 0.95793 | 0.023349 | 24.375 | 0.14565 | 0.15871 | 0.0237 | 1.6609 | 1.6609 | 0.003975
2019-03-19 15:54 | 1553025216.7645 | 687 | 0.20477 | 0.003626 | 9e-05 | 92.3034 | 90.7119 | 115.8887 | 21.2612 | 0.003465 | 0.003094 | 232.8408 | 0.26623 | 8.403 | 0.003634 | 0.004851 | 0.004768 | 5.6e-05 | 0.001002 | 0.29717 | 1.3258 | 2.5479 | 101.19 | 20.8 | 17.598 | 105.2324 | 109.3463 | 96.2727 | 0.004277 | 0.95567 | 0.0233 | 24.3806 | 0.15557 | 0.15838 | 0.023865 | 1.669 | 1.669 | 0.003948
2019-03-19 15:54 | 1553025217.7635 | 688 | 0.20478 | 0.003623 | 9e-05 | 92.1864 | 90.455 | 115.7145 | 21.2612 | 0.003397 | 0.003051 | 232.7676 | 0.26565 | 8.4034 | 0.003554 | 0.005199 | 0.005061 | 5.5e-05 | 0.00101 | 0.33732 | 1.3468 | 2.5481 | 101.19 | 20.8 | 17.592 | 103.6033 | 107.3947 | 95.2235 | 0.004276 | 0.95839 | 0.023315 | 24.3277 | 0.16958 | 0.15803 | 0.023785 | 1.6645 | 1.6645 | 0.00395
2019-03-19 15:54 | 1553025218.7639 | 689 | 0.2047 | 0.003619 | 9.1e-05 | 91.988 | 90.4293 | 115.5101 | 21.2612 | 0.003401 | 0.003066 | 232.7086 | 0.26657 | 8.404 | 0.003551 | 0.00584 | 0.005675 | 5.7e-05 | 0.001012 | 0.30757 | 1.3474 | 2.5481 | 101.19 | 20.81 | 17.578 | 103.6811 | 107.3308 | 95.586 | 0.004275 | 0.95745 | 0.023283 | 24.3181 | 0.15286 | 0.15858 | 0.024272 | 1.6962 | 1.6962 | 0.003889
2019-03-19 15:54 | 1553025219.7647 | 690 | 0.2047 | 0.00361 | 9e-05 | 91.8346 | 90.1139 | 115.6518 | 21.2612 | 0.003439 | 0.003102 | 232.6222 | 0.26694 | 8.4022 | 0.003547 | 0.005058 | 0.004983 | 5.5e-05 | 0.001012 | 0.31061 | 1.3367 | 2.5482 | 101.19 | 20.8 | 17.571 | 104.6119 | 107.2393 | 96.4545 | 0.00427 | 0.9552 | 0.02327 | 24.3617 | 0.16919 | 0.1588 | 0.024292 | 1.6967 | 1.6967 | 0.003894
2019-03-19 15:54 | 1553025220.7641 | 691 | 0.20466 | 0.003619 | 9.1e-05 | 91.7005 | 89.9266 | 115.9549 | 21.2612 | 0.003369 | 0.003043 | 232.5594 | 0.26625 | 8.4005 | 0.003518 | 0.005264 | 0.005099 | 5.3e-05 | 0.001011 | 0.3027 | 1.335 | 2.548 | 101.19 | 20.81 | 17.563 | 102.9094 | 106.5339 | 95.0404 | 0.004271 | 0.95887 | 0.023345 | 24.3469 | 0.17473 | 0.15839 | 0.024518 | 1.718 | 1.718 | 0.003835
2019-03-19 15:54 | 1553025221.763 | 692 | 0.20463 | 0.003619 | 9.1e-05 | 91.5144 | 89.769 | 116.1694 | 21.2612 | 0.003446 | 0.003098 | 232.5035 | 0.26624 | 8.4014 | 0.003525 | 0.005324 | 0.005239 | 5.4e-05 | 0.001015 | 0.3126 | 1.3021 | 2.548 | 101.19 | 20.81 | 17.55 | 104.781 | 106.7088 | 96.3567 | 0.004268 | 0.95545 | 0.023325 | 24.4129 | 0.17224 | 0.15839 | 0.024706 | 1.7297 | 1.7297 | 0.003809
2019-03-19 15:54 | 1553025222.7628 | 693 | 0.20463 | 0.003626 | 9.1e-05 | 91.3746 | 89.6372 | 116.1905 | 21.2612 | 0.003371 | 0.00309 | 232.4186 | 0.26592 | 8.4025 | 0.00355 | 0.004918 | 0.004852 | 5.2e-05 | 0.001019 | 0.31116 | 1.3021 | 2.548 | 101.19 | 20.81 | 17.545 | 102.9549 | 107.301 | 96.1624 | 0.004267 | 0.95595 | 0.023334 | 24.4087 | 0.17171 | 0.1582 | 0.024691 | 1.7292 | 1.7292 | 0.003806
2019-03-19 15:54 | 1553025223.7638 | 694 | 0.20461 | 0.003639 | 9.1e-05 | 91.232 | 89.5241 | 116.087 | 21.2612 | 0.003443 | 0.003084 | 232.3482 | 0.26576 | 8.401 | 0.003652 | 0.004618 | 0.004467 | 5.5e-05 | 0.00102 | 0.32925 | 1.3465 | 2.5481 | 101.19 | 20.81 | 17.539 | 104.7166 | 109.7896 | 96.0284 | 0.004265 | 0.9563 | 0.023327 | 24.3934 | 0.16903 | 0.1581 | 0.024828 | 1.7384 | 1.7384 | 0.003783
2019-03-19 15:54 | 1553025224.7639 | 695 | 0.20464 | 0.003643 | 9e-05 | 91.073 | 89.1748 | 116.0178 | 21.2612 | 0.00343 | 0.003053 | 232.2497 | 0.26644 | 8.4019 | 0.00367 | 0.005156 | 0.005036 | 5.4e-05 | 0.001022 | 0.31222 | 1.3659 | 2.5481 | 101.19 | 20.81 | 17.538 | 104.3901 | 110.2092 | 95.2884 | 0.004265 | 0.95822 | 0.023344 | 24.3617 | 0.1884 | 0.15851 | 0.024609 | 1.7243 | 1.7243 | 0.003824
2019-03-19 15:54 | 1553025225.7633 | 696 | 0.20471 | 0.003653 | 9e-05 | 90.8521 | 89.0618 | 115.9488 | 21.2612 | 0.003437 | 0.003092 | 232.1668 | 0.26561 | 8.403 | 0.003513 | 0.005646 | 0.00562 | 5.5e-05 | 0.001025 | 0.29391 | 1.3659 | 2.5482 | 101.19 | 20.81 | 17.537 | 104.5726 | 106.3985 | 96.2117 | 0.004265 | 0.95583 | 0.023308 | 24.3849 | 0.17802 | 0.15801 | 0.024178 | 1.6915 | 1.6915 | 0.003886
2019-03-19 15:54 | 1553025226.7641 | 697 | 0.20476 | 0.003657 | 9e-05 | 90.6591 | 89.0029 | 116.0842 | 21.2612 | 0.003379 | 0.003082 | 232.0913 | 0.26705 | 8.4017 | 0.003538 | 0.005765 | 0.005571 | 5.3e-05 | 0.001026 | 0.30646 | 1.3338 | 2.548 | 101.19 | 20.81 | 17.537 | 103.1482 | 107.0134 | 95.983 | 0.004265 | 0.95642 | 0.023329 | 24.3916 | 0.16492 | 0.15887 | 0.023909 | 1.6741 | 1.6741 | 0.003948
2019-03-19 15:54 | 1553025227.7634 | 698 | 0.20472 | 0.003656 | 9e-05 | 90.6177 | 88.7429 | 115.9086 | 21.2612 | 0.003428 | 0.003086 | 232.0104 | 0.26578 | 8.4044 | 0.003532 | 0.005234 | 0.005221 | 5.3e-05 | 0.001027 | 0.31974 | 1.3331 | 2.548 | 101.19 | 20.81 | 17.539 | 104.3564 | 106.8583 | 96.0769 | 0.004265 | 0.95617 | 0.023308 | 24.3763 | 0.187 | 0.15811 | 0.024147 | 1.6893 | 1.6893 | 0.003894
2019-03-19 15:54 | 1553025228.7641 | 699 | 0.2047 | 0.003656 | 9e-05 | 90.5901 | 88.5628 | 115.988 | 21.2612 | 0.003437 | 0.003053 | 231.9398 | 0.26586 | 8.4019 | 0.003537 | 0.003924 | 0.0038 | 5.2e-05 | 0.001026 | 0.3361 | 1.3243 | 2.548 | 101.19 | 20.81 | 17.54 | 104.5726 | 106.978 | 95.279 | 0.004266 | 0.95825 | 0.023341 | 24.3583 | 0.20245 | 0.15816 | 0.024255 | 1.6993 | 1.6993 | 0.003872
2019-03-19 15:54 | 1553025229.7632 | 700 | 0.20468 | 0.003653 | 9.1e-05 | 90.5757 | 88.3915 | 115.8171 | 21.2612 | 0.003462 | 0.003094 | 231.859 | 0.26653 | 8.403 | 0.00353 | 0.005387 | 0.005319 | 5.6e-05 | 0.001022 | 0.31741 | 1.343 | 2.5479 | 101.19 | 20.81 | 17.542 | 105.1624 | 106.8292 | 96.2636 | 0.004266 | 0.95569 | 0.023293 | 24.3728 | 0.21834 | 0.15856 | 0.02437 | 1.7038 | 1.7038 | 0.003871
2019-03-19 15:54 | 1553025230.7646 | 701 | 0.20465 | 0.003646 | 9e-05 | 90.5364 | 88.6223 | 115.7103 | 21.2612 | 0.0034 | 0.003067 | 231.786 | 0.26576 | 8.4024 | 0.003531 | 0.004663 | 0.004539 | 5.2e-05 | 0.001022 | 0.32803 | 1.3539 | 2.5482 | 101.19 | 20.81 | 17.545 | 103.6559 | 106.8324 | 95.615 | 0.004267 | 0.95737 | 0.023303 | 24.3402 | 0.19112 | 0.1581 | 0.024559 | 1.7178 | 1.7178 | 0.003829
2019-03-19 15:54 | 1553025231.7643 | 702 | 0.20461 | 0.003647 | 9.1e-05 | 90.6897 | 88.6018 | 115.4533 | 21.2612 | 0.003402 | 0.003057 | 231.7288 | 0.26534 | 8.404 | 0.003609 | 0.005639 | 0.005603 | 5.5e-05 | 0.001021 | 0.29463 | 1.3542 | 2.548 | 101.19 | 20.81 | 17.548 | 103.7204 | 108.7331 | 95.3784 | 0.004268 | 0.95799 | 0.023284 | 24.3053 | 0.20834 | 0.15785 | 0.024772 | 1.7312 | 1.7312 | 0.003793
2019-03-19 15:54 | 1553025232.7644 | 703 | 0.20461 | 0.003652 | 9.1e-05 | 90.8251 | 88.8443 | 115.2843 | 21.2612 | 0.00341 | 0.003086 | 231.6584 | 0.26544 | 8.4034 | 0.003603 | 0.005411 | 0.005278 | 5e-05 | 0.001026 | 0.30845 | 1.3405 | 2.5479 | 101.19 | 20.81 | 17.549 | 103.9207 | 108.5857 | 96.0849 | 0.004268 | 0.95615 | 0.023245 | 24.3108 | 0.19723 | 0.15791 | 0.024805 | 1.7306 | 1.7306 | 0.003796
2019-03-19 15:54 | 1553025233.7644 | 704 | 0.20461 | 0.003654 | 9.2e-05 | 91.0133 | 89.2176 | 114.9965 | 21.2612 | 0.003421 | 0.00309 | 231.5897 | 0.26513 | 8.4037 | 0.003541 | 0.005135 | 0.005045 | 5.2e-05 | 0.001031 | 0.31564 | 1.3263 | 2.5482 | 101.19 | 20.81 | 17.548 | 104.1715 | 107.0748 | 96.1764 | 0.004268 | 0.95592 | 0.023213 | 24.2834 | 0.17824 | 0.15773 | 0.024817 | 1.729 | 1.729 | 0.003795
2019-03-19 15:54 | 1553025234.764 | 705 | 0.20461 | 0.00366 | 9.2e-05 | 91.2107 | 89.3378 | 115.066 | 21.2612 | 0.003436 | 0.003113 | 231.5199 | 0.26572 | 8.404 | 0.003597 | 0.005107 | 0.005074 | 5.2e-05 | 0.001027 | 0.32493 | 1.3078 | 2.548 | 101.19 | 20.81 | 17.546 | 104.5396 | 108.455 | 96.7177 | 0.004267 | 0.95452 | 0.023203 | 24.3086 | 0.18558 | 0.15808 | 0.024755 | 1.724 | 1.724 | 0.003814
2019-03-19 15:54 | 1553025235.7634 | 706 | 0.20459 | 0.003675 | 9.3e-05 | 91.4382 | 89.8827 | 114.8747 | 21.2612 | 0.003446 | 0.003081 | 231.444 | 0.26597 | 8.4031 | 0.003532 | 0.005068 | 0.004984 | 5.3e-05 | 0.001026 | 0.29999 | 1.2889 | 2.548 | 101.19 | 20.81 | 17.544 | 104.7755 | 106.8576 | 95.9635 | 0.004267 | 0.95647 | 0.023207 | 24.2636 | 0.15347 | 0.15822 | 0.024854 | 1.7311 | 1.7311 | 0.003802
2019-03-19 15:54 | 1553025236.7645 | 707 | 0.20458 | 0.003683 | 9.4e-05 | 91.6024 | 89.8423 | 114.7932 | 21.2612 | 0.003478 | 0.003087 | 231.3689 | 0.26569 | 8.4011 | 0.003515 | 0.004842 | 0.004763 | 5.5e-05 | 0.001024 | 0.31636 | 1.3062 | 2.548 | 101.19 | 20.81 | 17.537 | 105.5496 | 106.449 | 96.1045 | 0.004265 | 0.9561 | 0.023195 | 24.2596 | 0.17353 | 0.15806 | 0.024941 | 1.7362 | 1.7362 | 0.003787
2019-03-19 15:54 | 1553025237.763 | 708 | 0.20455 | 0.003691 | 9.6e-05 | 91.6819 | 89.9259 | 114.9805 | 21.2612 | 0.00341 | 0.003113 | 231.2906 | 0.26645 | 8.401 | 0.003462 | 0.005016 | 0.004917 | 5.6e-05 | 0.001026 | 0.29518 | 1.345 | 2.548 | 101.19 | 20.81 | 17.533 | 103.9081 | 105.1633 | 96.7146 | 0.004264 | 0.95453 | 0.023194 | 24.2994 | 0.17298 | 0.15851 | 0.025095 | 1.7468 | 1.7468 | 0.003775
2019-03-19 15:54 | 1553025238.7637 | 709 | 0.20459 | 0.003704 | 9.7e-05 | 91.7218 | 89.8331 | 115.0667 | 21.2612 | 0.003422 | 0.003139 | 231.2095 | 0.26613 | 8.4001 | 0.003421 | 0.005085 | 0.005007 | 5.1e-05 | 0.001024 | 0.30182 | 1.363 | 2.5483 | 101.19 | 20.81 | 17.528 | 104.196 | 104.1637 | 97.3603 | 0.004263 | 0.95286 | 0.023183 | 24.3298 | 0.18611 | 0.15832 | 0.024861 | 1.7296 | 1.7296 | 0.003808
2019-03-19 15:54 | 1553025239.7641 | 710 | 0.20457 | 0.003713 | 9.8e-05 | 91.964 | 90.1036 | 115.0207 | 21.2612 | 0.003436 | 0.003083 | 231.1341 | 0.26633 | 8.4038 | 0.003486 | 0.004965 | 0.004885 | 5.4e-05 | 0.001023 | 0.34987 | 1.3523 | 2.5481 | 101.19 | 20.81 | 17.518 | 104.5302 | 105.74 | 95.9995 | 0.00426 | 0.95637 | 0.023221 | 24.2802 | 0.1828 | 0.15844 | 0.024972 | 1.7401 | 1.7401 | 0.003788
2019-03-19 15:54 | 1553025240.7638 | 711 | 0.20456 | 0.003718 | 9.8e-05 | 92.0696 | 90.1632 | 114.8343 | 21.2612 | 0.003447 | 0.003078 | 231.0702 | 0.26626 | 8.4039 | 0.003467 | 0.004739 | 0.004564 | 5.2e-05 | 0.001015 | 0.31957 | 1.3429 | 2.548 | 101.19 | 20.81 | 17.513 | 104.8031 | 105.2977 | 95.8703 | 0.004259 | 0.95671 | 0.023206 | 24.2563 | 0.18715 | 0.1584 | 0.024979 | 1.7395 | 1.7395 | 0.003788
2019-03-19 15:54 | 1553025241.7632 | 712 | 0.2046 | 0.00372 | 9.8e-05 | 92.3274 | 90.5403 | 114.7576 | 21.2612 | 0.00343 | 0.003106 | 230.9882 | 0.26651 | 8.404 | 0.003458 | 0.005191 | 0.005115 | 5.3e-05 | 0.001016 | 0.32012 | 1.3284 | 2.5481 | 101.19 | 20.81 | 17.503 | 104.4027 | 105.0821 | 96.5494 | 0.004257 | 0.95495 | 0.023177 | 24.2704 | 0.17483 | 0.15854 | 0.024768 | 1.7226 | 1.7226 | 0.003829
2019-03-19 15:54 | 1553025242.7629 | 713 | 0.20467 | 0.003725 | 9.9e-05 | 92.5015 | 90.6924 | 114.8033 | 21.2612 | 0.003394 | 0.003099 | 230.913 | 0.26642 | 8.4035 | 0.003444 | 0.004761 | 0.004576 | 5.1e-05 | 0.001013 | 0.31238 | 1.333 | 2.5482 | 101.19 | 20.81 | 17.498 | 103.5239 | 104.7219 | 96.3934 | 0.004255 | 0.95536 | 0.023187 | 24.2702 | 0.17667 | 0.15849 | 0.024352 | 1.6943 | 1.6943 | 0.003891
2019-03-19 15:54 | 1553025243.7639 | 714 | 0.20465 | 0.003721 | 9.9e-05 | 92.5927 | 90.9188 | 114.9418 | 21.2612 | 0.003421 | 0.003131 | 230.8357 | 0.2655 | 8.4039 | 0.003472 | 0.004656 | 0.004581 | 5.5e-05 | 0.001017 | 0.30718 | 1.3557 | 2.548 | 101.19 | 20.81 | 17.494 | 104.1715 | 105.4228 | 97.1659 | 0.004254 | 0.95336 | 0.023176 | 24.3102 | 0.16318 | 0.15795 | 0.024443 | 1.6999 | 1.6999 | 0.003865
2019-03-19 15:54 | 1553025244.7633 | 715 | 0.20471 | 0.003716 | 0.0001 | 92.762 | 91.1862 | 114.601 | 21.2612 | 0.00339 | 0.003078 | 230.7475 | 0.26543 | 8.4031 | 0.003454 | 0.004853 | 0.004721 | 5.1e-05 | 0.00102 | 0.30934 | 1.3225 | 2.5481 | 101.19 | 20.81 | 17.488 | 103.4224 | 104.9696 | 95.8742 | 0.004253 | 0.9567 | 0.023182 | 24.2317 | 0.15324 | 0.1579 | 0.024088 | 1.6756 | 1.6756 | 0.00392
2019-03-19 15:54 | 1553025245.7631 | 716 | 0.20469 | 0.003707 | 0.0001 | 93.0068 | 91.2595 | 114.7378 | 21.2612 | 0.003465 | 0.003157 | 230.6781 | 0.26612 | 8.4026 | 0.003456 | 0.005344 | 0.005225 | 5.2e-05 | 0.001023 | 0.2975 | 1.3302 | 2.5481 | 101.19 | 20.81 | 17.484 | 105.2506 | 105.0295 | 97.7835 | 0.004252 | 0.95178 | 0.023137 | 24.3088 | 0.16964 | 0.15831 | 0.024265 | 1.6846 | 1.6846 | 0.003909
2019-03-19 15:54 | 1553025246.7646 | 717 | 0.2047 | 0.003698 | 9.9e-05 | 93.1602 | 91.5188 | 114.6209 | 21.2612 | 0.003454 | 0.003151 | 230.6057 | 0.26647 | 8.4035 | 0.003429 | 0.005271 | 0.005187 | 5.4e-05 | 0.001026 | 0.29446 | 1.3309 | 2.5482 | 101.19 | 20.81 | 17.476 | 104.9839 | 104.3586 | 97.6449 | 0.00425 | 0.95213 | 0.023129 | 24.2919 | 0.159 | 0.15852 | 0.024177 | 1.678 | 1.678 | 0.00393
2019-03-19 15:54 | 1553025247.7632 | 718 | 0.20467 | 0.003696 | 9.9e-05 | 93.2751 | 91.7685 | 114.4562 | 21.2612 | 0.003461 | 0.003103 | 230.54 | 0.26558 | 8.403 | 0.003431 | 0.004827 | 0.004766 | 4.9e-05 | 0.001024 | 0.30558 | 1.2954 | 2.548 | 101.19 | 20.81 | 17.472 | 105.1444 | 104.4294 | 96.4734 | 0.004249 | 0.95515 | 0.023149 | 24.2361 | 0.14566 | 0.15799 | 0.024379 | 1.6934 | 1.6934 | 0.003881
2019-03-19 15:54 | 1553025248.7631 | 719 | 0.20465 | 0.00369 | 9.8e-05 | 93.448 | 91.8551 | 114.1611 | 21.2612 | 0.003451 | 0.00311 | 230.4762 | 0.26558 | 8.4022 | 0.003442 | 0.005546 | 0.005468 | 4.8e-05 | 0.001027 | 0.29031 | 1.3099 | 2.548 | 101.19 | 20.81 | 17.468 | 104.9076 | 104.6875 | 96.6513 | 0.004248 | 0.95469 | 0.023114 | 24.2106 | 0.15378 | 0.158 | 0.024521 | 1.7008 | 1.7008 | 0.003864
2019-03-19 15:54 | 1553025249.7646 | 720 | 0.20463 | 0.003689 | 9.7e-05 | 93.6244 | 91.9565 | 114.22 | 21.2612 | 0.003446 | 0.003138 | 230.4183 | 0.26606 | 8.402 | 0.003455 | 0.004116 | 0.004086 | 5.1e-05 | 0.001018 | 0.31542 | 1.3036 | 2.5481 | 101.19 | 20.81 | 17.468 | 104.7786 | 104.9878 | 97.3219 | 0.004248 | 0.95296 | 0.023099 | 24.2388 | 0.16079 | 0.15828 | 0.024619 | 1.7065 | 1.7065 | 0.003859
2019-03-19 15:54 | 1553025250.7643 | 721 | 0.20463 | 0.003685 | 9.7e-05 | 93.7709 | 92.0723 | 114.2933 | 21.2612 | 0.00352 | 0.003189 | 230.3611 | 0.2662 | 8.4036 | 0.003419 | 0.004291 | 0.004171 | 5.3e-05 | 0.001017 | 0.31365 | 1.3215 | 2.548 | 101.19 | 20.81 | 17.455 | 106.5773 | 104.133 | 98.5604 | 0.004245 | 0.94979 | 0.023068 | 24.2871 | 0.16351 | 0.15836 | 0.024604 | 1.7032 | 1.7032 | 0.003868
2019-03-19 15:54 | 1553025251.7633 | 722 | 0.20464 | 0.003683 | 9.7e-05 | 93.8159 | 92.2836 | 114.3733 | 21.2612 | 0.003442 | 0.003104 | 230.2874 | 0.2661 | 8.4052 | 0.003503 | 0.004351 | 0.004268 | 5e-05 | 0.001014 | 0.28385 | 1.318 | 2.5481 | 101.19 | 20.81 | 17.451 | 104.6826 | 106.1642 | 96.4968 | 0.004244 | 0.95509 | 0.02314 | 24.2281 | 0.1473 | 0.1583 | 0.024559 | 1.7054 | 1.7054 | 0.003861
2019-03-19 15:54 | 1553025252.7633 | 723 | 0.20463 | 0.003684 | 9.8e-05 | 93.9822 | 92.4959 | 114.4402 | 21.2612 | 0.003418 | 0.00313 | 230.2175 | 0.26536 | 8.4052 | 0.003441 | 0.005322 | 0.005186 | 5.1e-05 | 0.00101 | 0.32997 | 1.3058 | 2.5481 | 101.19 | 20.81 | 17.446 | 104.0923 | 104.6504 | 97.1292 | 0.004243 | 0.95346 | 0.023127 | 24.2559 | 0.14258 | 0.15786 | 0.024625 | 1.709 | 1.709 | 0.003843
2019-03-19 15:54 | 1553025253.7629 | 724 | 0.20461 | 0.00368 | 9.8e-05 | 94.017 | 92.4812 | 114.5064 | 21.2612 | 0.003397 | 0.003125 | 230.1404 | 0.26582 | 8.4033 | 0.003464 | 0.00475 | 0.004561 | 5.1e-05 | 0.001007 | 0.30397 | 1.3525 | 2.5482 | 101.19 | 20.81 | 17.435 | 103.6033 | 105.2198 | 97.0179 | 0.00424 | 0.95374 | 0.023137 | 24.2592 | 0.14732 | 0.15813 | 0.024736 | 1.7174 | 1.7174 | 0.00383
2019-03-19 15:54 | 1553025254.7625 | 725 | 0.2046 | 0.003687 | 9.9e-05 | 94.2021 | 92.6439 | 114.8543 | 21.2612 | 0.003423 | 0.003172 | 230.0917 | 0.26513 | 8.4044 | 0.003492 | 0.005209 | 0.005205 | 5e-05 | 0.001007 | 0.29695 | 1.3585 | 2.5481 | 101.19 | 20.81 | 17.429 | 104.2243 | 105.902 | 98.1496 | 0.004239 | 0.95084 | 0.023137 | 24.3332 | 0.1492 | 0.15773 | 0.024832 | 1.7241 | 1.7241 | 0.003806
2019-03-19 15:54 | 1553025255.7626 | 726 | 0.20459 | 0.003686 | 0.0001 | 94.2856 | 92.9011 | 114.9503 | 21.2612 | 0.003476 | 0.003151 | 230.0265 | 0.26576 | 8.4048 | 0.00349 | 0.004319 | 0.004221 | 4.9e-05 | 0.001006 | 0.30215 | 1.3307 | 2.5481 | 101.19 | 20.81 | 17.414 | 105.4975 | 105.8375 | 97.6471 | 0.004235 | 0.95213 | 0.023162 | 24.3269 | 0.13232 | 0.1581 | 0.024867 | 1.7283 | 1.7283 | 0.003806
2019-03-19 15:54 | 1553025256.7627 | 727 | 0.2047 | 0.003688 | 0.0001 | 94.4539 | 92.8733 | 114.9104 | 21.2612 | 0.003446 | 0.003078 | 229.9658 | 0.26541 | 8.4011 | 0.003436 | 0.0046 | 0.004496 | 5.1e-05 | 0.00101 | 0.32925 | 1.3285 | 2.5481 | 101.19 | 20.81 | 17.407 | 104.781 | 104.5316 | 95.886 | 0.004233 | 0.95667 | 0.023213 | 24.2648 | 0.15094 | 0.15789 | 0.024175 | 1.6839 | 1.6839 | 0.003901
2019-03-19 15:54 | 1553025257.7637 | 728 | 0.20477 | 0.003694 | 0.0001 | 94.5641 | 93.2836 | 114.778 | 21.2612 | 0.003415 | 0.003103 | 229.905 | 0.26696 | 8.4033 | 0.003426 | 0.005379 | 0.005246 | 4.9e-05 | 0.001009 | 0.3131 | 1.3472 | 2.548 | 101.19 | 20.81 | 17.401 | 104.0261 | 104.2982 | 96.4922 | 0.004232 | 0.9551 | 0.023181 | 24.2707 | 0.12194 | 0.15881 | 0.023774 | 1.6536 | 1.6536 | 0.003995
2019-03-19 15:54 | 1553025258.7628 | 729 | 0.20479 | 0.003691 | 0.000101 | 94.733 | 93.0977 | 114.766 | 21.2612 | 0.003429 | 0.003138 | 229.8286 | 0.26564 | 8.4023 | 0.0034 | 0.005127 | 0.004969 | 5.3e-05 | 0.001009 | 0.30541 | 1.319 | 2.5482 | 101.19 | 20.81 | 17.39 | 104.3792 | 103.6622 | 97.3359 | 0.004229 | 0.95293 | 0.023153 | 24.2971 | 0.15576 | 0.15803 | 0.023669 | 1.6443 | 1.6443 | 0.003998
2019-03-19 15:54 | 1553025259.7637 | 730 | 0.2048 | 0.003681 | 0.000101 | 94.8522 | 93.246 | 114.6977 | 21.2612 | 0.003461 | 0.003175 | 229.7549 | 0.26463 | 8.4047 | 0.003392 | 0.004901 | 0.004819 | 4.7e-05 | 0.001013 | 0.32676 | 1.3274 | 2.5482 | 101.19 | 20.81 | 17.386 | 105.142 | 103.4619 | 98.2138 | 0.004228 | 0.95067 | 0.023119 | 24.3187 | 0.15276 | 0.15743 | 0.023592 | 1.6365 | 1.6365 | 0.004002
2019-03-19 15:54 | 1553025260.7629 | 731 | 0.20484 | 0.003665 | 0.000102 | 94.9173 | 93.471 | 114.659 | 21.2612 | 0.003417 | 0.00313 | 229.6752 | 0.26593 | 8.4029 | 0.003431 | 0.004935 | 0.004793 | 4.7e-05 | 0.001015 | 0.3533 | 1.3258 | 2.5481 | 101.19 | 20.81 | 17.384 | 104.0692 | 104.4129 | 97.1244 | 0.004228 | 0.95347 | 0.023149 | 24.2789 | 0.13734 | 0.1582 | 0.023391 | 1.6246 | 1.6246 | 0.004051
2019-03-19 15:54 | 1553025261.7639 | 732 | 0.2048 | 0.003651 | 0.000102 | 95.1409 | 93.6779 | 114.5644 | 21.2612 | 0.003513 | 0.003121 | 229.609 | 0.26519 | 8.4031 | 0.003413 | 0.004335 | 0.004286 | 5e-05 | 0.00101 | 0.32388 | 1.3196 | 2.5481 | 101.19 | 20.81 | 17.384 | 106.4009 | 103.9735 | 96.909 | 0.004228 | 0.95403 | 0.023146 | 24.2618 | 0.13861 | 0.15776 | 0.023651 | 1.6425 | 1.6425 | 0.003996
2019-03-19 15:54 | 1553025262.763 | 733 | 0.20483 | 0.003632 | 0.000103 | 95.2312 | 93.8806 | 114.3604 | 21.2612 | 0.003399 | 0.003161 | 229.5507 | 0.26572 | 8.401 | 0.00344 | 0.004471 | 0.004351 | 5.1e-05 | 0.001004 | 0.30807 | 1.3338 | 2.5479 | 101.19 | 20.81 | 17.385 | 103.6465 | 104.6441 | 97.8964 | 0.004228 | 0.95149 | 0.023095 | 24.2725 | 0.12776 | 0.15808 | 0.023488 | 1.6275 | 1.6275 | 0.00404
2019-03-19 15:54 | 1553025263.7644 | 734 | 0.20483 | 0.003622 | 0.000103 | 95.3855 | 93.9289 | 114.3191 | 21.2612 | 0.00342 | 0.003112 | 229.4918 | 0.2656 | 8.4024 | 0.003473 | 0.005104 | 0.005017 | 5e-05 | 0.001007 | 0.31188 | 1.2746 | 2.5479 | 101.19 | 20.81 | 17.388 | 104.1424 | 105.4339 | 96.7052 | 0.004229 | 0.95455 | 0.023128 | 24.2291 | 0.13764 | 0.15801 | 0.023523 | 1.6322 | 1.6322 | 0.004027
2019-03-19 15:54 | 1553025264.764 | 735 | 0.20486 | 0.003608 | 0.000103 | 95.4998 | 94.2008 | 114.3794 | 21.2612 | 0.003459 | 0.003128 | 229.429 | 0.26556 | 8.405 | 0.003458 | 0.004916 | 0.004833 | 4.9e-05 | 0.00101 | 0.30215 | 1.2658 | 2.5479 | 101.19 | 20.81 | 17.391 | 105.0947 | 105.0618 | 97.0853 | 0.004229 | 0.95357 | 0.023122 | 24.248 | 0.1225 | 0.15798 | 0.023316 | 1.6175 | 1.6175 | 0.004063
2019-03-19 15:54 | 1553025265.7641 | 736 | 0.20488 | 0.003596 | 0.000103 | 95.657 | 94.3039 | 114.359 | 21.2612 | 0.003419 | 0.003145 | 229.3815 | 0.26553 | 8.4052 | 0.003485 | 0.005206 | 0.00515 | 5e-05 | 0.001011 | 0.33251 | 1.2778 | 2.5481 | 101.19 | 20.81 | 17.395 | 104.1298 | 105.7298 | 97.4966 | 0.00423 | 0.95251 | 0.023107 | 24.2593 | 0.12743 | 0.15796 | 0.023224 | 1.61 | 1.61 | 0.004082
2019-03-19 15:54 | 1553025266.7632 | 737 | 0.20488 | 0.00358 | 0.000103 | 95.7966 | 94.6009 | 114.5156 | 21.2612 | 0.003399 | 0.003155 | 229.3155 | 0.2647 | 8.4023 | 0.003391 | 0.005625 | 0.005518 | 4.8e-05 | 0.001013 | 0.32621 | 1.3301 | 2.5482 | 101.19 | 20.81 | 17.404 | 103.6331 | 103.443 | 97.734 | 0.004233 | 0.9519 | 0.023116 | 24.2837 | 0.11235 | 0.15747 | 0.023247 | 1.6122 | 1.6122 | 0.004063
2019-03-19 15:54 | 1553025267.7632 | 738 | 0.2049 | 0.003565 | 0.000103 | 95.8563 | 94.5813 | 114.8899 | 21.2612 | 0.003492 | 0.003158 | 229.2383 | 0.26573 | 8.4022 | 0.003467 | 0.00538 | 0.005231 | 5.1e-05 | 0.001016 | 0.31382 | 1.3353 | 2.548 | 101.19 | 20.81 | 17.408 | 105.8964 | 105.2892 | 97.8087 | 0.004233 | 0.95171 | 0.023151 | 24.3258 | 0.11977 | 0.15809 | 0.023137 | 1.607 | 1.607 | 0.004092
2019-03-19 15:54 | 1553025268.7633 | 739 | 0.2049 | 0.003548 | 0.000103 | 95.8323 | 94.6218 | 115.1111 | 21.2612 | 0.003447 | 0.003127 | 229.1751 | 0.26535 | 8.4027 | 0.0035 | 0.005117 | 0.004996 | 4.8e-05 | 0.001013 | 0.34363 | 1.3306 | 2.5481 | 101.19 | 20.81 | 17.413 | 104.8038 | 106.0829 | 97.0633 | 0.004235 | 0.95363 | 0.023197 | 24.3247 | 0.11371 | 0.15786 | 0.023156 | 1.6115 | 1.6115 | 0.004075
2019-03-19 15:54 | 1553025269.7633 | 740 | 0.20492 | 0.003538 | 0.000103 | 95.9463 | 94.744 | 115.2387 | 21.2612 | 0.003469 | 0.003135 | 229.1295 | 0.26674 | 8.4014 | 0.003469 | 0.005024 | 0.004979 | 4.5e-05 | 0.001013 | 0.31868 | 1.3378 | 2.5482 | 101.19 | 20.81 | 17.423 | 105.334 | 105.3316 | 97.2475 | 0.004237 | 0.95315 | 0.023204 | 24.3442 | 0.11279 | 0.15868 | 0.023077 | 1.6065 | 1.6065 | 0.004109
2019-03-19 15:54 | 1553025270.7627 | 741 | 0.20491 | 0.003529 | 0.000103 | 96.0016 | 94.7581 | 115.3099 | 21.2612 | 0.003408 | 0.003116 | 229.0784 | 0.26616 | 8.4032 | 0.003405 | 0.005143 | 0.005 | 4.5e-05 | 0.001017 | 0.30071 | 1.341 | 2.5481 | 101.19 | 20.81 | 17.428 | 103.8555 | 103.7957 | 96.7955 | 0.004238 | 0.95432 | 0.023225 | 24.3369 | 0.11661 | 0.15834 | 0.023127 | 1.6114 | 1.6114 | 0.004087
2019-03-19 15:54 | 1553025271.7634 | 742 | 0.20497 | 0.003523 | 0.000103 | 96.1273 | 94.7581 | 115.3509 | 21.2612 | 0.00344 | 0.003139 | 229.0199 | 0.26621 | 8.4023 | 0.003318 | 0.004945 | 0.004889 | 5.1e-05 | 0.001015 | 0.30254 | 1.3765 | 2.5482 | 101.19 | 20.81 | 17.432 | 104.6409 | 101.6793 | 97.3524 | 0.004239 | 0.95288 | 0.023212 | 24.3595 | 0.12832 | 0.15837 | 0.022797 | 1.5875 | 1.5875 | 0.00415
2019-03-19 15:55 | 1553025272.7629 | 743 | 0.20497 | 0.003503 | 0.000103 | 96.1999 | 95.041 | 115.6943 | 21.2612 | 0.003488 | 0.003117 | 228.9422 | 0.26564 | 8.4032 | 0.003371 | 0.00507 | 0.004936 | 4.7e-05 | 0.001006 | 0.35513 | 1.3592 | 2.548 | 101.19 | 20.81 | 17.433 | 105.7919 | 102.9581 | 96.815 | 0.00424 | 0.95427 | 0.023263 | 24.3781 | 0.10841 | 0.15803 | 0.022795 | 1.5909 | 1.5909 | 0.004132
2019-03-19 15:55 | 1553025273.7626 | 744 | 0.20497 | 0.003489 | 0.000103 | 96.293 | 95.0677 | 116.0598 | 21.2612 | 0.003456 | 0.003139 | 228.8654 | 0.26582 | 8.404 | 0.003406 | 0.004798 | 0.004704 | 4.4e-05 | 0.001 | 0.30469 | 1.3495 | 2.5481 | 101.19 | 20.81 | 17.432 | 105.0319 | 103.8194 | 97.3587 | 0.004239 | 0.95287 | 0.023283 | 24.4345 | 0.11454 | 0.15814 | 0.022849 | 1.596 | 1.596 | 0.004122
2019-03-19 15:55 | 1553025274.7636 | 745 | 0.20497 | 0.003475 | 0.000102 | 96.3929 | 95.0793 | 116.061 | 21.2612 | 0.003433 | 0.003141 | 228.7935 | 0.26641 | 8.4032 | 0.00336 | 0.004465 | 0.004386 | 4.9e-05 | 0.001003 | 0.31222 | 1.3426 | 2.5481 | 101.19 | 20.81 | 17.43 | 104.4626 | 102.6995 | 97.3901 | 0.004239 | 0.95279 | 0.023282 | 24.4356 | 0.12273 | 0.15849 | 0.022824 | 1.5943 | 1.5943 | 0.004135
2019-03-19 15:55 | 1553025275.764 | 746 | 0.20495 | 0.003457 | 0.0001 | 96.4958 | 95.3506 | 115.7955 | 21.2612 | 0.003462 | 0.00316 | 228.702 | 0.26565 | 8.4028 | 0.003397 | 0.004891 | 0.004849 | 4.5e-05 | 0.000999 | 0.27671 | 1.3939 | 2.5482 | 101.19 | 20.81 | 17.429 | 105.1696 | 103.6016 | 97.8705 | 0.004239 | 0.95155 | 0.02324 | 24.4235 | 0.10679 | 0.15803 | 0.022998 | 1.6037 | 1.6037 | 0.004099
2019-03-19 15:55 | 1553025276.7626 | 747 | 0.20492 | 0.003454 | 9.8e-05 | 96.5733 | 95.5119 | 115.892 | 21.2612 | 0.003462 | 0.003143 | 228.6255 | 0.26636 | 8.405 | 0.003428 | 0.005222 | 0.005064 | 4.3e-05 | 0.000998 | 0.31509 | 1.3301 | 2.5481 | 101.19 | 20.81 | 17.424 | 105.1616 | 104.3382 | 97.4473 | 0.004237 | 0.95264 | 0.023263 | 24.4197 | 0.098848 | 0.15846 | 0.023167 | 1.6172 | 1.6172 | 0.004076
2019-03-19 15:55 | 1553025277.7628 | 748 | 0.20492 | 0.003444 | 9.7e-05 | 96.615 | 95.4698 | 115.7239 | 21.2612 | 0.003471 | 0.003087 | 228.5585 | 0.26615 | 8.4022 | 0.003448 | 0.004986 | 0.004894 | 4.8e-05 | 0.000996 | 0.29103 | 1.3566 | 2.5482 | 101.19 | 20.81 | 17.422 | 105.3789 | 104.8203 | 96.0902 | 0.004237 | 0.95614 | 0.023289 | 24.3573 | 0.10666 | 0.15833 | 0.023184 | 1.6202 | 1.6202 | 0.004065
2019-03-19 15:55 | 1553025278.7628 | 749 | 0.20492 | 0.003443 | 9.6e-05 | 96.7275 | 95.6828 | 115.5995 | 21.2612 | 0.003408 | 0.003138 | 228.4817 | 0.26668 | 8.4034 | 0.003447 | 0.005056 | 0.004986 | 4.4e-05 | 0.000991 | 0.30093 | 1.3637 | 2.548 | 101.19 | 20.81 | 17.419 | 103.8509 | 104.8126 | 97.3359 | 0.004236 | 0.95293 | 0.023237 | 24.3852 | 0.097132 | 0.15865 | 0.023198 | 1.6177 | 1.6177 | 0.00408
2019-03-19 15:55 | 1553025279.7628 | 750 | 0.20493 | 0.003437 | 9.6e-05 | 96.7914 | 95.7468 | 115.1592 | 21.2612 | 0.003417 | 0.003163 | 228.3992 | 0.26534 | 8.4024 | 0.003417 | 0.004215 | 0.004046 | 4.4e-05 | 0.00099 | 0.27616 | 1.2878 | 2.548 | 101.19 | 20.81 | 17.411 | 104.0748 | 104.078 | 97.9278 | 0.004234 | 0.9514 | 0.023174 | 24.3582 | 0.097056 | 0.15785 | 0.02313 | 1.6086 | 1.6086 | 0.004082
2019-03-19 15:55 | 1553025280.7638 | 751 | 0.20486 | 0.003443 | 9.5e-05 | 96.8365 | 96.0541 | 115.1221 | 21.2612 | 0.003456 | 0.003123 | 228.3141 | 0.26676 | 8.4009 | 0.003464 | 0.004991 | 0.004847 | 4.6e-05 | 0.000994 | 0.30663 | 1.3326 | 2.548 | 101.19 | 20.81 | 17.407 | 105.0208 | 105.2128 | 96.9725 | 0.004233 | 0.95386 | 0.023201 | 24.3229 | 0.072562 | 0.1587 | 0.023533 | 1.6386 | 1.6386 | 0.004029
2019-03-19 15:55 | 1553025281.764 | 752 | 0.20489 | 0.003443 | 9.4e-05 | 96.9445 | 95.9982 | 115.169 | 21.2612 | 0.003433 | 0.003138 | 228.2414 | 0.26625 | 8.4025 | 0.003496 | 0.00473 | 0.004654 | 4.8e-05 | 0.000992 | 0.34202 | 1.307 | 2.548 | 101.19 | 20.81 | 17.399 | 104.4563 | 105.9963 | 97.3251 | 0.004231 | 0.95295 | 0.023194 | 24.3394 | 0.087735 | 0.15839 | 0.023344 | 1.625 | 1.625 | 0.004055
2019-03-19 15:55 | 1553025282.7643 | 753 | 0.20487 | 0.003451 | 9.3e-05 | 97.0249 | 96.0568 | 115.3709 | 21.2612 | 0.003424 | 0.003128 | 228.1804 | 0.2665 | 8.4045 | 0.003501 | 0.00493 | 0.004825 | 4.7e-05 | 0.000983 | 0.31421 | 1.3303 | 2.548 | 101.19 | 20.81 | 17.397 | 104.2369 | 106.1043 | 97.0853 | 0.004231 | 0.95357 | 0.023222 | 24.3529 | 0.089698 | 0.15854 | 0.023481 | 1.6366 | 1.6366 | 0.00403
2019-03-19 15:55 | 1553025283.7624 | 754 | 0.20485 | 0.003457 | 9.2e-05 | 97.1519 | 96.215 | 115.4167 | 21.2612 | 0.003419 | 0.003162 | 228.1173 | 0.26701 | 8.404 | 0.003475 | 0.005252 | 0.005094 | 4.7e-05 | 0.000982 | 0.28744 | 1.3271 | 2.5481 | 101.19 | 20.81 | 17.395 | 104.1267 | 105.489 | 97.9167 | 0.00423 | 0.95143 | 0.023201 | 24.385 | 0.086671 | 0.15884 | 0.023623 | 1.645 | 1.645 | 0.004017
2019-03-19 15:55 | 1553025284.7631 | 755 | 0.20487 | 0.003462 | 9.2e-05 | 97.2086 | 96.3989 | 115.6455 | 21.2612 | 0.003483 | 0.003121 | 228.0538 | 0.26603 | 8.4027 | 0.003499 | 0.004615 | 0.004551 | 4.6e-05 | 0.000983 | 0.30574 | 1.3334 | 2.5481 | 101.19 | 20.81 | 17.392 | 105.6818 | 106.0689 | 96.9145 | 0.00423 | 0.95401 | 0.023255 | 24.3762 | 0.07481 | 0.15826 | 0.023458 | 1.6374 | 1.6374 | 0.004021
2019-03-19 15:55 | 1553025285.7631 | 756 | 0.2049 | 0.003471 | 9.1e-05 | 97.2974 | 96.4335 | 116.011 | 21.2612 | 0.003508 | 0.003118 | 227.9972 | 0.26661 | 8.4035 | 0.003433 | 0.004831 | 0.00474 | 5.1e-05 | 0.000979 | 0.31061 | 1.3774 | 2.5482 | 101.19 | 20.81 | 17.391 | 106.2932 | 104.4759 | 96.8392 | 0.004229 | 0.9542 | 0.023294 | 24.4122 | 0.079771 | 0.15861 | 0.023272 | 1.6271 | 1.6271 | 0.004055
2019-03-19 15:55 | 1553025286.7626 | 757 | 0.20491 | 0.003474 | 9.1e-05 | 97.3763 | 96.5596 | 115.8279 | 21.2612 | 0.003459 | 0.003102 | 227.9436 | 0.26534 | 8.4038 | 0.003466 | 0.005301 | 0.005236 | 5e-05 | 0.000972 | 0.28871 | 1.3268 | 2.548 | 101.19 | 20.82 | 17.389 | 105.1027 | 105.2654 | 96.4577 | 0.004231 | 0.95519 | 0.023288 | 24.3803 | 0.075341 | 0.15785 | 0.023216 | 1.6227 | 1.6227 | 0.004047
2019-03-19 15:55 | 1553025287.7633 | 758 | 0.20492 | 0.003483 | 9.1e-05 | 97.4631 | 96.6613 | 115.6824 | 21.2612 | 0.003463 | 0.003095 | 227.8716 | 0.26523 | 8.4014 | 0.003542 | 0.004773 | 0.004734 | 4.7e-05 | 0.00098 | 0.30558 | 1.3485 | 2.5482 | 101.19 | 20.82 | 17.384 | 105.2026 | 107.1112 | 96.3017 | 0.00423 | 0.95559 | 0.023278 | 24.3599 | 0.073886 | 0.15779 | 0.023144 | 1.6171 | 1.6171 | 0.004059
2019-03-19 15:55 | 1553025288.7633 | 759 | 0.20498 | 0.003481 | 9.1e-05 | 97.554 | 96.8819 | 115.7324 | 21.2612 | 0.003424 | 0.003131 | 227.8024 | 0.26523 | 8.4049 | 0.003547 | 0.005175 | 0.005041 | 4.6e-05 | 0.000976 | 0.34363 | 1.3597 | 2.5481 | 101.19 | 20.82 | 17.382 | 104.2408 | 107.227 | 97.1596 | 0.00423 | 0.95338 | 0.023256 | 24.3934 | 0.061841 | 0.15779 | 0.022769 | 1.5894 | 1.5894 | 0.00413
2019-03-19 15:55 | 1553025289.7634 | 760 | 0.20501 | 0.003475 | 9.1e-05 | 97.6056 | 96.9257 | 116.0361 | 21.2612 | 0.003475 | 0.003157 | 227.7354 | 0.26609 | 8.403 | 0.003507 | 0.004798 | 0.004682 | 4.8e-05 | 0.000977 | 0.30071 | 1.3221 | 2.5481 | 101.19 | 20.82 | 17.376 | 105.4739 | 106.26 | 97.7835 | 0.004228 | 0.95178 | 0.023267 | 24.446 | 0.062526 | 0.1583 | 0.022617 | 1.5795 | 1.5795 | 0.004169
2019-03-19 15:55 | 1553025290.7628 | 761 | 0.20502 | 0.00347 | 9.1e-05 | 97.7113 | 96.7281 | 116.0765 | 21.2612 | 0.003454 | 0.003101 | 227.6879 | 0.26482 | 8.4018 | 0.003492 | 0.005418 | 0.00538 | 4.5e-05 | 0.000972 | 0.30165 | 1.3408 | 2.5481 | 101.19 | 20.82 | 17.375 | 104.9713 | 105.8869 | 96.4357 | 0.004228 | 0.95525 | 0.023314 | 24.4058 | 0.090456 | 0.15754 | 0.022569 | 1.5792 | 1.5792 | 0.00415
2019-03-19 15:55 | 1553025291.7636 | 762 | 0.20504 | 0.003455 | 9.1e-05 | 97.8178 | 97.1579 | 116.283 | 21.2612 | 0.003489 | 0.003127 | 227.6244 | 0.26575 | 8.4024 | 0.003502 | 0.004967 | 0.004797 | 4.8e-05 | 0.000972 | 0.32012 | 1.3353 | 2.548 | 101.19 | 20.82 | 17.373 | 105.8147 | 106.1428 | 97.0688 | 0.004228 | 0.95361 | 0.023314 | 24.4484 | 0.060545 | 0.1581 | 0.022468 | 1.5722 | 1.5722 | 0.004183
2019-03-19 15:55 | 1553025292.7636 | 763 | 0.20503 | 0.003443 | 9e-05 | 97.9009 | 97.2963 | 115.8615 | 21.2612 | 0.003404 | 0.003155 | 227.5525 | 0.2664 | 8.4009 | 0.003508 | 0.005362 | 0.005293 | 4.5e-05 | 0.000972 | 0.30215 | 1.3967 | 2.5481 | 101.19 | 20.82 | 17.368 | 103.7589 | 106.2915 | 97.7294 | 0.004226 | 0.95191 | 0.023251 | 24.4258 | 0.055414 | 0.15848 | 0.022576 | 1.5755 | 1.5755 | 0.004184
2019-03-19 15:55 | 1553025293.764 | 764 | 0.20504 | 0.003434 | 9e-05 | 98.0165 | 97.3226 | 115.4434 | 21.2612 | 0.00338 | 0.003166 | 227.5055 | 0.26595 | 8.4046 | 0.003571 | 0.005416 | 0.005319 | 4.6e-05 | 0.000979 | 0.35817 | 1.363 | 2.5481 | 101.19 | 20.82 | 17.364 | 103.1717 | 107.8035 | 98.0036 | 0.004225 | 0.95121 | 0.023201 | 24.3907 | 0.06354 | 0.15821 | 0.022516 | 1.568 | 1.568 | 0.004197
2019-03-19 15:55 | 1553025294.7637 | 765 | 0.20503 | 0.003423 | 8.8e-05 | 98.1473 | 97.5227 | 114.8982 | 21.2612 | 0.003445 | 0.003123 | 227.437 | 0.26589 | 8.4002 | 0.003569 | 0.005345 | 0.005234 | 4.5e-05 | 0.000979 | 0.32621 | 1.3379 | 2.548 | 101.19 | 20.82 | 17.356 | 104.759 | 107.7671 | 96.9739 | 0.004223 | 0.95386 | 0.023178 | 24.2993 | 0.057108 | 0.15818 | 0.022604 | 1.5726 | 1.5726 | 0.004184
2019-03-19 15:55 | 1553025295.7637 | 766 | 0.20507 | 0.00341 | 8.8e-05 | 98.2319 | 97.6353 | 114.9834 | 21.2612 | 0.003429 | 0.003179 | 227.3569 | 0.26556 | 8.4023 | 0.003585 | 0.005102 | 0.004997 | 4.5e-05 | 0.000986 | 0.34396 | 1.3576 | 2.5482 | 101.19 | 20.82 | 17.35 | 104.3768 | 108.1507 | 98.3173 | 0.004222 | 0.95041 | 0.023145 | 24.3523 | 0.054489 | 0.15798 | 0.02236 | 1.5534 | 1.5534 | 0.004231
2019-03-19 15:55 | 1553025296.7628 | 767 | 0.20503 | 0.003397 | 8.7e-05 | 98.3067 | 97.7069 | 115.2812 | 21.2612 | 0.003459 | 0.003125 | 227.2984 | 0.26617 | 8.4004 | 0.003443 | 0.005162 | 0.00506 | 4.6e-05 | 0.000988 | 0.30165 | 1.3138 | 2.5481 | 101.19 | 20.82 | 17.345 | 105.101 | 104.7078 | 97.0162 | 0.004221 | 0.95375 | 0.023215 | 24.3411 | 0.054735 | 0.15834 | 0.022583 | 1.5738 | 1.5738 | 0.004186
2019-03-19 15:55 | 1553025297.8309 | 768 | 0.20509 | 0.003388 | 8.6e-05 | 98.4111 | 97.723 | 115.4789 | 21.2612 | 0.003453 | 0.003072 | 227.2599 | 0.26593 | 8.4025 | 0.003513 | 0.005127 | 0.005095 | 4.6e-05 | 0.000991 | 0.33964 | 1.3622 | 2.5482 | 101.19 | 20.82 | 17.338 | 104.958 | 106.4041 | 95.7286 | 0.004219 | 0.95708 | 0.023276 | 24.3196 | 0.062759 | 0.1582 | 0.02226 | 1.5554 | 1.5554 | 0.004231
2019-03-19 15:55 | 1553025298.7753 | 769 | 0.20508 | 0.003376 | 8.6e-05 | 98.4807 | 97.8251 | 115.4434 | 21.2612 | 0.003433 | 0.003123 | 227.2168 | 0.26594 | 8.4036 | 0.003442 | 0.005785 | 0.005586 | 5e-05 | 0.000994 | 0.30718 | 1.3552 | 2.5481 | 101.19 | 20.82 | 17.336 | 104.457 | 104.6756 | 96.9631 | 0.004219 | 0.95388 | 0.023233 | 24.3565 | 0.059748 | 0.15821 | 0.02236 | 1.5595 | 1.5595 | 0.00422
2019-03-19 15:55 | 1553025299.7755 | 770 | 0.20506 | 0.003366 | 8.5e-05 | 98.5269 | 97.9042 | 115.3223 | 21.2612 | 0.003423 | 0.003133 | 227.1441 | 0.26649 | 8.4053 | 0.003553 | 0.004885 | 0.004822 | 4.7e-05 | 0.000996 | 0.32045 | 1.3191 | 2.5481 | 101.19 | 20.82 | 17.333 | 104.2352 | 107.3874 | 97.2091 | 0.004218 | 0.95325 | 0.023213 | 24.3518 | 0.056714 | 0.15854 | 0.022455 | 1.5649 | 1.5649 | 0.004215
2019-03-19 15:55 | 1553025300.7754 | 771 | 0.20505 | 0.003358 | 8.5e-05 | 98.677 | 98.2316 | 115.1157 | 21.2612 | 0.003413 | 0.003099 | 227.0718 | 0.26608 | 8.4031 | 0.003493 | 0.005279 | 0.005123 | 4.7e-05 | 0.000991 | 0.29966 | 1.3228 | 2.5482 | 101.19 | 20.82 | 17.333 | 103.982 | 105.9304 | 96.3763 | 0.004218 | 0.9554 | 0.023219 | 24.3026 | 0.040462 | 0.15829 | 0.022548 | 1.5717 | 1.5717 | 0.00419
2019-03-19 15:55 | 1553025301.7745 | 772 | 0.2051 | 0.003353 | 8.5e-05 | 98.7301 | 98.1977 | 115.2087 | 21.2612 | 0.003417 | 0.003094 | 227.0087 | 0.26571 | 8.4017 | 0.003532 | 0.005154 | 0.005129 | 4.7e-05 | 0.000989 | 0.31636 | 1.3283 | 2.548 | 101.19 | 20.82 | 17.334 | 104.0741 | 106.8717 | 96.2783 | 0.004218 | 0.95565 | 0.023231 | 24.3092 | 0.048364 | 0.15807 | 0.022245 | 1.5514 | 1.5514 | 0.004238
2019-03-19 15:55 | 1553025302.7749 | 773 | 0.20508 | 0.003347 | 8.4e-05 | 98.8285 | 98.1292 | 115.1716 | 21.2612 | 0.003444 | 0.00314 | 226.9465 | 0.26611 | 8.4043 | 0.003531 | 0.005775 | 0.005595 | 4.9e-05 | 0.00099 | 0.30165 | 1.3507 | 2.5482 | 101.19 | 20.82 | 17.336 | 104.7282 | 106.8316 | 97.3657 | 0.004219 | 0.95285 | 0.023193 | 24.341 | 0.063518 | 0.15831 | 0.022402 | 1.5599 | 1.5599 | 0.004222
2019-03-19 15:55 | 1553025303.7756 | 774 | 0.20508 | 0.003342 | 8.4e-05 | 98.8567 | 98.3916 | 115.1941 | 21.2612 | 0.003401 | 0.003143 | 226.8907 | 0.2654 | 8.4027 | 0.003535 | 0.005331 | 0.005268 | 4.9e-05 | 0.000996 | 0.30574 | 1.3404 | 2.548 | 101.19 | 20.82 | 17.338 | 103.6864 | 106.9474 | 97.4449 | 0.004219 | 0.95264 | 0.023193 | 24.346 | 0.04219 | 0.15789 | 0.022356 | 1.5567 | 1.5567 | 0.004219
2019-03-19 15:55 | 1553025304.7761 | 775 | 0.20515 | 0.003337 | 8.4e-05 | 98.9534 | 98.4193 | 115.3303 | 21.2612 | 0.003387 | 0.003078 | 226.845 | 0.26564 | 8.4047 | 0.003545 | 0.005657 | 0.005568 | 4.7e-05 | 0.001 | 0.3027 | 1.3427 | 2.5481 | 101.19 | 20.82 | 17.34 | 103.3618 | 107.1709 | 95.871 | 0.00422 | 0.95671 | 0.023256 | 24.3086 | 0.048407 | 0.15803 | 0.021979 | 1.5346 | 1.5346 | 0.004284
2019-03-19 15:55 | 1553025305.7761 | 776 | 0.20516 | 0.003331 | 8.4e-05 | 98.9963 | 98.5364 | 115.2641 | 21.2612 | 0.003422 | 0.003112 | 226.8108 | 0.26554 | 8.4018 | 0.003537 | 0.004998 | 0.004912 | 4.7e-05 | 0.001003 | 0.32444 | 1.337 | 2.5482 | 101.19 | 20.82 | 17.349 | 104.196 | 106.9852 | 96.703 | 0.004222 | 0.95456 | 0.023223 | 24.329 | 0.041647 | 0.15797 | 0.021902 | 1.527 | 1.527 | 0.004304
2019-03-19 15:55 | 1553025306.7748 | 777 | 0.20521 | 0.003317 | 8.4e-05 | 99.0335 | 98.5001 | 115.1254 | 21.2612 | 0.003386 | 0.003081 | 226.753 | 0.26558 | 8.4031 | 0.003492 | 0.005429 | 0.005329 | 4.6e-05 | 0.001001 | 0.32029 | 1.3595 | 2.548 | 101.19 | 20.82 | 17.355 | 103.3173 | 105.902 | 95.9439 | 0.004223 | 0.95652 | 0.023233 | 24.2894 | 0.048304 | 0.15799 | 0.02165 | 1.5101 | 1.5101 | 0.004352
2019-03-19 15:55 | 1553025307.7759 | 778 | 0.20521 | 0.003311 | 8.4e-05 | 99.1562 | 98.5833 | 115.1991 | 21.2612 | 0.003385 | 0.003136 | 226.6881 | 0.26724 | 8.402 | 0.003599 | 0.00483 | 0.004805 | 4.7e-05 | 0.000998 | 0.31974 | 1.3208 | 2.548 | 101.19 | 20.82 | 17.369 | 103.2959 | 108.5053 | 97.2818 | 0.004227 | 0.95306 | 0.023199 | 24.3412 | 0.051829 | 0.15898 | 0.021659 | 1.5085 | 1.5085 | 0.004384
2019-03-19 15:55 | 1553025308.7748 | 779 | 0.20525 | 0.003297 | 8.4e-05 | 99.1496 | 98.4214 | 115.603 | 21.2612 | 0.003445 | 0.003125 | 226.624 | 0.26579 | 8.4028 | 0.003504 | 0.00521 | 0.005123 | 4.9e-05 | 0.000998 | 0.3298 | 1.3223 | 2.548 | 101.19 | 20.82 | 17.377 | 104.7597 | 106.1996 | 97.021 | 0.004229 | 0.95374 | 0.023248 | 24.3752 | 0.065944 | 0.15812 | 0.021413 | 1.4944 | 1.4944 | 0.004402
2019-03-19 15:55 | 1553025309.7753 | 780 | 0.20525 | 0.00328 | 8.4e-05 | 99.1175 | 98.5833 | 115.9093 | 21.2612 | 0.003476 | 0.003095 | 226.591 | 0.26616 | 8.4047 | 0.003522 | 0.004979 | 0.004838 | 5.1e-05 | 0.001 | 0.27865 | 1.3199 | 2.548 | 101.19 | 20.82 | 17.383 | 105.5086 | 106.6135 | 96.2863 | 0.00423 | 0.95563 | 0.023301 | 24.3833 | 0.048336 | 0.15834 | 0.021466 | 1.5016 | 1.5016 | 0.004387
2019-03-19 15:55 | 1553025310.776 | 781 | 0.20525 | 0.00326 | 8.3e-05 | 99.1704 | 98.7977 | 115.903 | 21.2612 | 0.003412 | 0.00314 | 226.5172 | 0.26565 | 8.4014 | 0.003566 | 0.004648 | 0.004428 | 4.9e-05 | 0.001 | 0.33301 | 1.3538 | 2.548 | 101.19 | 20.82 | 17.396 | 103.9592 | 107.6867 | 97.3806 | 0.004233 | 0.95281 | 0.023266 | 24.4187 | 0.03367 | 0.15804 | 0.021454 | 1.4986 | 1.4986 | 0.004387
2019-03-19 15:55 | 1553025311.7764 | 782 | 0.20525 | 0.003244 | 8.3e-05 | 99.1559 | 98.7388 | 115.7295 | 21.2612 | 0.003386 | 0.003126 | 226.4424 | 0.26612 | 8.4016 | 0.003542 | 0.00469 | 0.004636 | 5.2e-05 | 0.001005 | 0.3235 | 1.3231 | 2.5481 | 101.19 | 20.82 | 17.402 | 103.3289 | 107.1119 | 97.0304 | 0.004235 | 0.95371 | 0.02326 | 24.3889 | 0.037705 | 0.15831 | 0.021475 | 1.4996 | 1.4996 | 0.004392
2019-03-19 15:55 | 1553025312.7752 | 783 | 0.20526 | 0.003235 | 8.3e-05 | 99.2448 | 98.8361 | 116.0283 | 21.2612 | 0.003442 | 0.003106 | 226.4009 | 0.2662 | 8.4038 | 0.003518 | 0.004686 | 0.004579 | 5.3e-05 | 0.001008 | 0.3246 | 1.3224 | 2.548 | 101.19 | 20.82 | 17.41 | 104.6875 | 106.5261 | 96.5673 | 0.004237 | 0.95491 | 0.023305 | 24.4051 | 0.036908 | 0.15836 | 0.02145 | 1.5007 | 1.5007 | 0.00439
2019-03-19 15:55 | 1553025313.7752 | 784 | 0.20528 | 0.003218 | 8.2e-05 | 99.2361 | 98.6426 | 115.9469 | 21.2612 | 0.003432 | 0.00313 | 226.3521 | 0.26692 | 8.4031 | 0.003538 | 0.005083 | 0.004966 | 5.3e-05 | 0.001015 | 0.31293 | 1.3461 | 2.548 | 101.19 | 20.82 | 17.414 | 104.4538 | 107.008 | 97.1299 | 0.004238 | 0.95346 | 0.023279 | 24.415 | 0.053653 | 0.15879 | 0.02132 | 1.4901 | 1.4901 | 0.004433
2019-03-19 15:55 | 1553025314.7747 | 785 | 0.20527 | 0.003209 | 8.1e-05 | 99.2298 | 98.8642 | 115.9845 | 21.2612 | 0.003395 | 0.003084 | 226.3093 | 0.26595 | 8.402 | 0.003502 | 0.00504 | 0.004996 | 5.3e-05 | 0.001012 | 0.3361 | 1.3186 | 2.5481 | 101.19 | 20.82 | 17.418 | 103.5341 | 106.1404 | 96.0277 | 0.004239 | 0.9563 | 0.023317 | 24.3826 | 0.033014 | 0.15821 | 0.021398 | 1.498 | 1.498 | 0.004394
2019-03-19 15:55 | 1553025315.7743 | 786 | 0.20526 | 0.003206 | 8.1e-05 | 99.2373 | 98.9621 | 115.9164 | 21.2612 | 0.003462 | 0.003054 | 226.2301 | 0.26597 | 8.4049 | 0.003474 | 0.004791 | 0.004629 | 4.9e-05 | 0.001013 | 0.31741 | 1.3254 | 2.5481 | 101.19 | 20.82 | 17.424 | 105.1585 | 105.455 | 95.3097 | 0.00424 | 0.95817 | 0.023333 | 24.3517 | 0.024833 | 0.15822 | 0.021483 | 1.505 | 1.505 | 0.004373
2019-03-19 15:55 | 1553025316.7747 | 787 | 0.20529 | 0.003197 | 8e-05 | 99.2664 | 98.8762 | 116.0704 | 21.2612 | 0.003396 | 0.003125 | 226.167 | 0.26642 | 8.4037 | 0.00344 | 0.005071 | 0.004987 | 5.3e-05 | 0.001007 | 0.32189 | 1.3737 | 2.5481 | 101.19 | 20.82 | 17.426 | 103.5663 | 104.6426 | 97.0273 | 0.00424 | 0.95372 | 0.023294 | 24.4246 | 0.03523 | 0.15849 | 0.021283 | 1.4885 | 1.4885 | 0.004429
2019-03-19 15:55 | 1553025317.7744 | 788 | 0.20528 | 0.003196 | 8e-05 | 99.3495 | 98.9789 | 115.964 | 21.2612 | 0.003419 | 0.003081 | 226.1184 | 0.26634 | 8.4031 | 0.003312 | 0.00504 | 0.004934 | 5.3e-05 | 0.00101 | 0.33179 | 1.3394 | 2.548 | 101.19 | 20.82 | 17.429 | 104.1371 | 101.5332 | 95.9517 | 0.004241 | 0.9565 | 0.023317 | 24.378 | 0.033427 | 0.15845 | 0.021382 | 1.4969 | 1.4969 | 0.004403
2019-03-19 15:55 | 1553025318.7748 | 789 | 0.20526 | 0.003191 | 8.1e-05 | 99.3948 | 98.9998 | 115.6913 | 21.2612 | 0.00339 | 0.0031 | 226.0747 | 0.2663 | 8.4016 | 0.003416 | 0.005172 | 0.005037 | 5.1e-05 | 0.001005 | 0.30685 | 1.3267 | 2.548 | 101.19 | 20.82 | 17.43 | 103.4161 | 104.045 | 96.4021 | 0.004241 | 0.95533 | 0.023276 | 24.3641 | 0.035618 | 0.15842 | 0.021495 | 1.5022 | 1.5022 | 0.004387
2019-03-19 15:55 | 1553025319.7745 | 790 | 0.20526 | 0.003183 | 8e-05 | 99.3906 | 99.1478 | 115.5347 | 21.2612 | 0.003428 | 0.003099 | 226.0335 | 0.26542 | 8.4034 | 0.003453 | 0.004629 | 0.004506 | 5.2e-05 | 0.001006 | 0.31669 | 1.3027 | 2.5482 | 101.19 | 20.82 | 17.431 | 104.3494 | 104.9398 | 96.392 | 0.004242 | 0.95536 | 0.02326 | 24.3473 | 0.021882 | 0.1579 | 0.021491 | 1.5009 | 1.5009 | 0.004376
2019-03-19 15:55 | 1553025320.7749 | 791 | 0.20534 | 0.003176 | 8e-05 | 99.4518 | 99.2228 | 115.7678 | 21.2612 | 0.003396 | 0.003084 | 226.0092 | 0.26591 | 8.4022 | 0.003431 | 0.005134 | 0.005052 | 5.5e-05 | 0.000997 | 0.34739 | 1.3241 | 2.5482 | 101.19 | 20.82 | 17.433 | 103.5694 | 104.4223 | 96.0308 | 0.004242 | 0.95629 | 0.023295 | 24.3599 | 0.020626 | 0.15819 | 0.021002 | 1.4689 | 1.4689 | 0.00448
2019-03-19 15:55 | 1553025321.7762 | 792 | 0.20539 | 0.003167 | 8.1e-05 | 99.5578 | 99.0639 | 115.9774 | 21.2612 | 0.003372 | 0.003068 | 225.9787 | 0.26695 | 8.4033 | 0.003396 | 0.004823 | 0.004732 | 5.5e-05 | 0.000998 | 0.30613 | 1.3018 | 2.5481 | 101.19 | 20.81 | 17.432 | 102.9879 | 103.5631 | 95.6331 | 0.004239 | 0.95732 | 0.023329 | 24.3688 | 0.044486 | 0.15881 | 0.020709 | 1.4505 | 1.4505 | 0.004555
2019-03-19 15:55 | 1553025322.776 | 793 | 0.20549 | 0.003161 | 8.2e-05 | 99.5686 | 99.1259 | 116.0021 | 21.2612 | 0.003406 | 0.003126 | 225.9348 | 0.2659 | 8.4008 | 0.003365 | 0.004873 | 0.004747 | 4.9e-05 | 0.000997 | 0.30381 | 1.3159 | 2.5481 | 101.19 | 20.81 | 17.432 | 103.8061 | 102.8244 | 97.0406 | 0.004239 | 0.95369 | 0.023287 | 24.4179 | 0.039861 | 0.15818 | 0.020173 | 1.4103 | 1.4103 | 0.004666
2019-03-19 15:55 | 1553025323.7761 | 794 | 0.20552 | 0.003148 | 8.2e-05 | 99.6706 | 99.056 | 116.3412 | 21.2612 | 0.003361 | 0.003089 | 225.8853 | 0.26615 | 8.4032 | 0.003412 | 0.005147 | 0.005009 | 5.3e-05 | 0.000995 | 0.30917 | 1.3256 | 2.548 | 101.19 | 20.82 | 17.433 | 102.72 | 103.9483 | 96.1429 | 0.004242 | 0.956 | 0.023349 | 24.4239 | 0.055329 | 0.15833 | 0.020011 | 1.4027 | 1.4027 | 0.004696
2019-03-19 15:55 | 1553025324.7751 | 795 | 0.20551 | 0.003125 | 8.3e-05 | 99.7024 | 99.3992 | 116.3621 | 21.2612 | 0.003419 | 0.003097 | 225.8271 | 0.26556 | 8.4029 | 0.00336 | 0.004529 | 0.004454 | 5.4e-05 | 0.000998 | 0.286 | 1.3296 | 2.5481 | 101.19 | 20.82 | 17.434 | 104.115 | 102.6941 | 96.3488 | 0.004242 | 0.95547 | 0.023345 | 24.4329 | 0.027249 | 0.15798 | 0.020079 | 1.4071 | 1.4071 | 0.004671
2019-03-19 15:55 | 1553025325.7759 | 796 | 0.20568 | 0.003104 | 8.3e-05 | 99.7177 | 99.3632 | 116.4065 | 21.2612 | 0.003445 | 0.003059 | 225.7985 | 0.26646 | 8.4034 | 0.003456 | 0.005001 | 0.004802 | 5.5e-05 | 0.000998 | 0.30254 | 1.3258 | 2.548 | 101.19 | 20.82 | 17.434 | 104.752 | 105.0201 | 95.4129 | 0.004242 | 0.9579 | 0.023379 | 24.4066 | 0.031858 | 0.15852 | 0.01905 | 1.3368 | 1.3368 | 0.004933
2019-03-19 15:55 | 1553025326.7749 | 797 | 0.20558 | 0.003081 | 8.3e-05 | 99.7724 | 99.5317 | 116.5138 | 21.2612 | 0.003367 | 0.003108 | 225.7422 | 0.26619 | 8.3997 | 0.003349 | 0.005062 | 0.004933 | 5.5e-05 | 0.000996 | 0.32228 | 1.3341 | 2.5481 | 101.19 | 20.82 | 17.433 | 102.8692 | 102.4348 | 96.6161 | 0.004242 | 0.95478 | 0.023352 | 24.4577 | 0.021601 | 0.15835 | 0.019691 | 1.3803 | 1.3803 | 0.004773
2019-03-19 15:55 | 1553025327.7763 | 798 | 0.20564 | 0.003052 | 8.5e-05 | 99.8813 | 99.3923 | 116.7553 | 21.2612 | 0.0034 | 0.003075 | 225.6877 | 0.26601 | 8.4015 | 0.00337 | 0.0045 | 0.004393 | 4.8e-05 | 0.000985 | 0.32084 | 1.3646 | 2.5482 | 101.19 | 20.82 | 17.433 | 103.6709 | 102.9448 | 95.7998 | 0.004242 | 0.95689 | 0.023402 | 24.456 | 0.043897 | 0.15825 | 0.019348 | 1.359 | 1.359 | 0.004844
2019-03-19 15:55 | 1553025328.7763 | 799 | 0.20571 | 0.003034 | 8.4e-05 | 99.9227 | 99.3954 | 116.9109 | 21.2612 | 0.003416 | 0.003062 | 225.6354 | 0.26487 | 8.404 | 0.003336 | 0.004521 | 0.004399 | 5e-05 | 0.000982 | 0.31581 | 1.323 | 2.5483 | 101.19 | 20.82 | 17.435 | 104.0632 | 102.1153 | 95.503 | 0.004243 | 0.95766 | 0.023427 | 24.4624 | 0.047328 | 0.15757 | 0.018978 | 1.3344 | 1.3344 | 0.004912
2019-03-19 15:55 | 1553025329.776 | 800 | 0.20564 | 0.003009 | 8.5e-05 | 100.0076 | 99.7334 | 116.8586 | 21.2612 | 0.003429 | 0.003069 | 225.5671 | 0.26627 | 8.4042 | 0.003291 | 0.005047 | 0.004989 | 5.2e-05 | 0.000982 | 0.33107 | 1.3633 | 2.5481 | 101.19 | 20.82 | 17.435 | 104.3705 | 101.0347 | 95.6753 | 0.004243 | 0.95722 | 0.023416 | 24.4626 | 0.024558 | 0.15841 | 0.019429 | 1.3656 | 1.3656 | 0.004826
2019-03-19 15:55 | 1553025330.7742 | 801 | 0.2058 | 0.002986 | 8.5e-05 | 100.0884 | 99.6742 | 116.6872 | 21.2612 | 0.003379 | 0.003061 | 225.5005 | 0.26625 | 8.4036 | 0.003361 | 0.004699 | 0.004616 | 4.8e-05 | 0.00098 | 0.33771 | 1.3763 | 2.548 | 101.19 | 20.82 | 17.435 | 103.1528 | 102.7341 | 95.4747 | 0.004243 | 0.95774 | 0.023405 | 24.4381 | 0.037091 | 0.15839 | 0.018482 | 1.2983 | 1.2983 | 0.005075
2019-03-19 15:56 | 1553025331.7752 | 802 | 0.20582 | 0.002959 | 8.4e-05 | 100.1106 | 99.8105 | 116.5959 | 21.2612 | 0.00337 | 0.003059 | 225.4516 | 0.26631 | 8.4055 | 0.003347 | 0.004817 | 0.004675 | 5.3e-05 | 0.00098 | 0.29966 | 1.3426 | 2.548 | 101.19 | 20.82 | 17.435 | 102.9455 | 102.3813 | 95.4332 | 0.004243 | 0.95784 | 0.023397 | 24.4271 | 0.026852 | 0.15843 | 0.018417 | 1.2933 | 1.2933 | 0.005096
2019-03-19 15:56 | 1553025332.7743 | 803 | 0.20603 | 0.002935 | 8.4e-05 | 100.1781 | 99.8358 | 116.5954 | 21.2612 | 0.003368 | 0.003026 | 225.3741 | 0.26595 | 8.4036 | 0.003312 | 0.004856 | 0.004809 | 5.1e-05 | 0.000985 | 0.35601 | 1.3313 | 2.5481 | 101.19 | 20.82 | 17.434 | 102.899 | 101.5293 | 94.6152 | 0.004242 | 0.95998 | 0.023423 | 24.3999 | 0.030612 | 0.15822 | 0.017194 | 1.2086 | 1.2086 | 0.005446
2019-03-19 15:56 | 1553025333.7746 | 804 | 0.20575 | 0.002914 | 8.3e-05 | 100.2276 | 99.9143 | 116.5614 | 21.2612 | 0.003333 | 0.003077 | 225.3042 | 0.26545 | 8.401 | 0.003435 | 0.004644 | 0.004551 | 5.5e-05 | 0.000985 | 0.32676 | 1.3451 | 2.5481 | 101.19 | 20.83 | 17.432 | 102.0351 | 104.5159 | 95.8498 | 0.004245 | 0.95676 | 0.023381 | 24.4373 | 0.028008 | 0.15792 | 0.018903 | 1.3266 | 1.3266 | 0.004952
2019-03-19 15:56 | 1553025334.7742 | 805 | 0.20585 | 0.002894 | 8.3e-05 | 100.2726 | 100.1383 | 116.9419 | 21.2612 | 0.003308 | 0.002977 | 225.2431 | 0.2656 | 8.4014 | 0.003355 | 0.005014 | 0.004877 | 5.4e-05 | 0.000986 | 0.3178 | 1.3238 | 2.548 | 101.19 | 20.83 | 17.431 | 101.4289 | 102.5849 | 93.45 | 0.004244 | 0.96303 | 0.023495 | 24.3974 | 0.011995 | 0.15801 | 0.018318 | 1.2918 | 1.2918 | 0.005088
2019-03-19 15:56 | 1553025335.7754 | 806 | 0.20589 | 0.002863 | 8.3e-05 | 100.3522 | 100.0828 | 117.0698 | 21.2612 | 0.003334 | 0.002995 | 225.1808 | 0.26584 | 8.4044 | 0.003343 | 0.005047 | 0.004934 | 5.3e-05 | 0.000978 | 0.31222 | 1.2966 | 2.5481 | 101.19 | 20.83 | 17.429 | 102.0642 | 102.2879 | 93.8764 | 0.004244 | 0.96191 | 0.023495 | 24.425 | 0.024043 | 0.15815 | 0.018081 | 1.275 | 1.275 | 0.00516
2019-03-19 15:56 | 1553025336.7747 | 807 | 0.20587 | 0.002841 | 8.3e-05 | 100.393 | 100.1516 | 116.9909 | 21.2612 | 0.003314 | 0.003102 | 225.1391 | 0.26593 | 8.4035 | 0.003372 | 0.004917 | 0.004811 | 5.2e-05 | 0.000984 | 0.31166 | 1.3175 | 2.5481 | 101.19 | 20.83 | 17.428 | 101.5811 | 102.999 | 96.4688 | 0.004244 | 0.95516 | 0.023404 | 24.5028 | 0.02153 | 0.1582 | 0.018238 | 1.2812 | 1.2812 | 0.005137
2019-03-19 15:56 | 1553025337.7759 | 808 | 0.20598 | 0.00282 | 8.2e-05 | 100.4716 | 100.1999 | 116.9665 | 21.2612 | 0.003291 | 0.003019 | 225.089 | 0.26601 | 8.4004 | 0.00337 | 0.005199 | 0.005081 | 5.4e-05 | 0.000987 | 0.31045 | 1.2675 | 2.5481 | 101.19 | 20.83 | 17.429 | 101.0284 | 102.9501 | 94.4454 | 0.004244 | 0.96042 | 0.023466 | 24.4331 | 0.024221 | 0.15825 | 0.017635 | 1.2421 | 1.2421 | 0.0053
2019-03-19 15:56 | 1553025338.7745 | 809 | 0.20587 | 0.002793 | 8.2e-05 | 100.4812 | 100.192 | 116.6612 | 21.2612 | 0.003392 | 0.003028 | 225.0227 | 0.26642 | 8.4021 | 0.003416 | 0.005064 | 0.004977 | 5.7e-05 | 0.000986 | 0.31116 | 1.3196 | 2.5481 | 101.19 | 20.83 | 17.432 | 103.4791 | 104.0433 | 94.6787 | 0.004245 | 0.95981 | 0.023428 | 24.4089 | 0.025779 | 0.15849 | 0.018307 | 1.2873 | 1.2873 | 0.005122
2019-03-19 15:56 | 1553025339.7745 | 810 | 0.2061 | 0.002772 | 8.1e-05 | 100.5322 | 100.3242 | 116.6421 | 21.2612 | 0.003274 | 0.003009 | 224.9809 | 0.26593 | 8.4026 | 0.00339 | 0.005055 | 0.004926 | 5.4e-05 | 0.00099 | 0.32604 | 1.3027 | 2.548 | 101.19 | 20.83 | 17.432 | 100.6147 | 103.4192 | 94.2083 | 0.004245 | 0.96104 | 0.023441 | 24.3913 | 0.018525 | 0.1582 | 0.01694 | 1.1918 | 1.1918 | 0.005522
2019-03-19 15:56 | 1553025340.775 | 811 | 0.20594 | 0.002753 | 8.1e-05 | 100.5091 | 100.3074 | 116.8658 | 21.2612 | 0.003347 | 0.003071 | 224.9469 | 0.26669 | 8.401 | 0.00341 | 0.004521 | 0.004378 | 5.7e-05 | 0.000988 | 0.29247 | 1.3063 | 2.5481 | 101.19 | 20.83 | 17.436 | 102.3845 | 103.9122 | 95.7004 | 0.004246 | 0.95715 | 0.023416 | 24.4642 | 0.017965 | 0.15865 | 0.017944 | 1.2612 | 1.2612 | 0.005233
2019-03-19 15:56 | 1553025341.7747 | 812 | 0.20606 | 0.002737 | 8.1e-05 | 100.4617 | 100.2523 | 116.8496 | 21.2612 | 0.003252 | 0.002998 | 224.8961 | 0.26587 | 8.404 | 0.003422 | 0.005044 | 0.004968 | 5.7e-05 | 0.000994 | 0.28423 | 1.3363 | 2.548 | 101.19 | 20.82 | 17.439 | 100.0962 | 104.1999 | 93.9437 | 0.004244 | 0.96173 | 0.02347 | 24.4042 | 0.018662 | 0.15817 | 0.017265 | 1.2162 | 1.2162 | 0.00541
2019-03-19 15:56 | 1553025342.7758 | 813 | 0.20608 | 0.002715 | 8.2e-05 | 100.5427 | 100.1112 | 116.7438 | 21.2612 | 0.003314 | 0.003048 | 224.8622 | 0.26622 | 8.4051 | 0.003423 | 0.005227 | 0.005119 | 5.6e-05 | 0.000993 | 0.32516 | 1.3686 | 2.5481 | 101.19 | 20.82 | 17.447 | 101.5905 | 104.2163 | 95.1608 | 0.004246 | 0.95855 | 0.023421 | 24.4336 | 0.038472 | 0.15837 | 0.017169 | 1.2068 | 1.2068 | 0.005459
2019-03-19 15:56 | 1553025343.7758 | 814 | 0.20602 | 0.002693 | 8.1e-05 | 100.5106 | 100.4592 | 116.7126 | 21.2612 | 0.003326 | 0.003016 | 224.8298 | 0.26666 | 8.4007 | 0.003327 | 0.004539 | 0.004438 | 5.9e-05 | 0.000999 | 0.33716 | 1.3668 | 2.5481 | 101.19 | 20.82 | 17.452 | 101.8788 | 101.8991 | 94.3959 | 0.004247 | 0.96055 | 0.023442 | 24.4049 | 0.004579 | 0.15864 | 0.017565 | 1.2359 | 1.2359 | 0.00534
2019-03-19 15:56 | 1553025344.7758 | 815 | 0.20608 | 0.002671 | 8.1e-05 | 100.6436 | 100.4849 | 116.5203 | 21.2612 | 0.003371 | 0.003015 | 224.7809 | 0.26545 | 8.4046 | 0.003387 | 0.005308 | 0.005192 | 5.6e-05 | 0.000999 | 0.29463 | 1.3345 | 2.5481 | 101.19 | 20.82 | 17.458 | 102.9557 | 103.3408 | 94.3585 | 0.004248 | 0.96064 | 0.023424 | 24.3836 | 0.014116 | 0.15791 | 0.017211 | 1.21 | 1.21 | 0.005429
2019-03-19 15:56 | 1553025345.7749 | 816 | 0.20598 | 0.002657 | 8.1e-05 | 100.6592 | 100.402 | 116.6278 | 21.2612 | 0.003331 | 0.003044 | 224.7332 | 0.26629 | 8.403 | 0.003402 | 0.004531 | 0.004481 | 5.8e-05 | 0.001003 | 0.32892 | 1.3721 | 2.5481 | 101.19 | 20.82 | 17.475 | 101.9935 | 103.7187 | 95.0592 | 0.004252 | 0.95882 | 0.023412 | 24.4181 | 0.022883 | 0.15841 | 0.017826 | 1.2527 | 1.2527 | 0.00526
2019-03-19 15:56 | 1553025346.7749 | 817 | 0.20618 | 0.002642 | 8e-05 | 100.7171 | 100.3722 | 116.4445 | 21.2612 | 0.003284 | 0.003028 | 224.6853 | 0.26625 | 8.4036 | 0.003402 | 0.005222 | 0.005079 | 6e-05 | 0.001006 | 0.30254 | 1.3519 | 2.5481 | 101.19 | 20.82 | 17.484 | 100.8668 | 103.7032 | 94.6739 | 0.004255 | 0.95982 | 0.023406 | 24.3861 | 0.030683 | 0.15839 | 0.016664 | 1.1706 | 1.1706 | 0.005629
2019-03-19 15:56 | 1553025347.7741 | 818 | 0.20597 | 0.002624 | 7.9e-05 | 100.727 | 100.4725 | 116.4574 | 21.2612 | 0.003301 | 0.003016 | 224.6497 | 0.26655 | 8.4007 | 0.003404 | 0.004601 | 0.00453 | 5.9e-05 | 0.001004 | 0.29911 | 1.3288 | 2.5481 | 101.19 | 20.82 | 17.501 | 101.2757 | 103.7635 | 94.3795 | 0.004259 | 0.96059 | 0.023417 | 24.3777 | 0.022625 | 0.15857 | 0.017906 | 1.2587 | 1.2587 | 0.005241
2019-03-19 15:56 | 1553025348.7751 | 819 | 0.20636 | 0.00261 | 7.8e-05 | 100.7708 | 100.6236 | 116.6182 | 21.2612 | 0.003344 | 0.003047 | 224.6007 | 0.2653 | 8.4032 | 0.003356 | 0.00484 | 0.004648 | 6.2e-05 | 0.001006 | 0.30165 | 1.2877 | 2.548 | 101.19 | 20.82 | 17.509 | 102.3155 | 102.5927 | 95.1265 | 0.004261 | 0.95864 | 0.023409 | 24.4193 | 0.013078 | 0.15783 | 0.015616 | 1.0971 | 1.0971 | 0.005985
2019-03-19 15:56 | 1553025349.7741 | 820 | 0.20616 | 0.002594 | 7.8e-05 | 100.811 | 100.6393 | 116.9473 | 21.2612 | 0.003286 | 0.002997 | 224.5485 | 0.26592 | 8.4011 | 0.003351 | 0.004911 | 0.004885 | 6.1e-05 | 0.001006 | 0.32803 | 1.3228 | 2.5481 | 101.19 | 20.82 | 17.517 | 100.9123 | 102.482 | 93.929 | 0.004263 | 0.96177 | 0.023481 | 24.4139 | 0.015247 | 0.1582 | 0.016833 | 1.1864 | 1.1864 | 0.005547
2019-03-19 15:56 | 1553025350.7741 | 821 | 0.20617 | 0.002585 | 7.7e-05 | 100.7594 | 100.5999 | 116.9513 | 21.2612 | 0.003285 | 0.003037 | 224.4897 | 0.26634 | 8.4024 | 0.003346 | 0.004866 | 0.004756 | 6.2e-05 | 0.001003 | 0.31636 | 1.3191 | 2.5481 | 101.19 | 20.82 | 17.53 | 100.892 | 102.361 | 94.8953 | 0.004266 | 0.95924 | 0.02345 | 24.4465 | 0.014167 | 0.15844 | 0.0168 | 1.1826 | 1.1826 | 0.005574
2019-03-19 15:56 | 1553025351.7746 | 822 | 0.20619 | 0.002561 | 7.7e-05 | 100.8155 | 100.6167 | 116.905 | 21.2612 | 0.003257 | 0.003032 | 224.4347 | 0.26578 | 8.4052 | 0.003462 | 0.00465 | 0.004585 | 6.3e-05 | 0.001005 | 0.32278 | 1.3073 | 2.5481 | 101.19 | 20.82 | 17.535 | 100.2098 | 105.1735 | 94.7638 | 0.004267 | 0.95959 | 0.02345 | 24.4373 | 0.017656 | 0.15811 | 0.016723 | 1.1771 | 1.1771 | 0.005588
2019-03-19 15:56 | 1553025352.7753 | 823 | 0.20622 | 0.002556 | 7.6e-05 | 100.8173 | 100.5342 | 117.2719 | 21.2612 | 0.003286 | 0.002962 | 224.3812 | 0.26606 | 8.4035 | 0.003454 | 0.00502 | 0.004932 | 6.3e-05 | 0.001006 | 0.31996 | 1.3292 | 2.5481 | 101.19 | 20.82 | 17.544 | 100.9154 | 104.9871 | 93.0917 | 0.004269 | 0.96397 | 0.02354 | 24.4199 | 0.025156 | 0.15828 | 0.016515 | 1.167 | 1.167 | 0.005642
2019-03-19 15:56 | 1553025353.7753 | 824 | 0.20629 | 0.002537 | 7.6e-05 | 100.8665 | 100.8753 | 117.413 | 21.2612 | 0.003311 | 0.003 | 224.3363 | 0.26635 | 8.404 | 0.00338 | 0.005364 | 0.005296 | 6.2e-05 | 0.001008 | 0.347 | 1.3429 | 2.5481 | 101.19 | 20.82 | 17.548 | 101.5003 | 103.186 | 93.9994 | 0.00427 | 0.96159 | 0.023525 | 24.4648 | -0.000777 | 0.15845 | 0.016129 | 1.139 | 1.139 | 0.005787
2019-03-19 15:56 | 1553025354.7745 | 825 | 0.20627 | 0.002523 | 7.5e-05 | 100.8662 | 100.6811 | 117.5341 | 21.2612 | 0.003296 | 0.002996 | 224.297 | 0.26576 | 8.402 | 0.003355 | 0.004498 | 0.004354 | 6.1e-05 | 0.001007 | 0.29087 | 1.3884 | 2.5481 | 101.19 | 20.82 | 17.551 | 101.154 | 102.5818 | 93.8945 | 0.004271 | 0.96186 | 0.023541 | 24.4739 | 0.016432 | 0.1581 | 0.016283 | 1.1506 | 1.1506 | 0.005716
2019-03-19 15:56 | 1553025355.7748 | 826 | 0.20635 | 0.002502 | 7.5e-05 | 100.9437 | 100.7602 | 117.6526 | 21.2612 | 0.00328 | 0.002992 | 224.2584 | 0.26496 | 8.4014 | 0.003371 | 0.005522 | 0.005384 | 6.1e-05 | 0.001001 | 0.3147 | 1.354 | 2.5482 | 101.19 | 20.82 | 17.556 | 100.7577 | 102.9605 | 93.8067 | 0.004272 | 0.96209 | 0.023555 | 24.4834 | 0.016269 | 0.15762 | 0.015807 | 1.1177 | 1.1177 | 0.005867
2019-03-19 15:56 | 1553025356.7745 | 827 | 0.20632 | 0.002488 | 7.5e-05 | 100.9908 | 100.552 | 117.8693 | 21.2612 | 0.003293 | 0.002987 | 224.223 | 0.26521 | 8.4029 | 0.003389 | 0.004443 | 0.004295 | 6.4e-05 | 0.001004 | 0.29286 | 1.343 | 2.5481 | 101.19 | 20.82 | 17.557 | 101.0708 | 103.3972 | 93.6731 | 0.004272 | 0.96244 | 0.023581 | 24.5014 | 0.038948 | 0.15777 | 0.016015 | 1.1336 | 1.1336 | 0.00579
2019-03-19 15:56 | 1553025357.7745 | 828 | 0.2063 | 0.002462 | 7.4e-05 | 100.9404 | 100.6866 | 118.023 | 21.2612 | 0.003259 | 0.002987 | 224.1665 | 0.26687 | 8.4057 | 0.003395 | 0.0054 | 0.005317 | 6.4e-05 | 0.001002 | 0.31349 | 1.3503 | 2.548 | 101.19 | 20.81 | 17.558 | 100.2476 | 103.5481 | 93.6769 | 0.00427 | 0.96243 | 0.023596 | 24.5175 | 0.022516 | 0.15876 | 0.016132 | 1.1427 | 1.1427 | 0.005779
2019-03-19 15:56 | 1553025358.7761 | 829 | 0.20633 | 0.002444 | 7.4e-05 | 101.0391 | 100.975 | 118.454 | 21.2612 | 0.003181 | 0.002991 | 224.1374 | 0.26503 | 8.4029 | 0.003361 | 0.00528 | 0.005241 | 6.3e-05 | 0.001003 | 0.29717 | 1.3315 | 2.548 | 101.19 | 20.81 | 17.56 | 98.3801 | 102.7317 | 93.777 | 0.00427 | 0.96217 | 0.023636 | 24.5656 | 0.005679 | 0.15767 | 0.015981 | 1.1339 | 1.1339 | 0.005784
2019-03-19 15:56 | 1553025359.7756 | 830 | 0.20637 | 0.002428 | 7.4e-05 | 101.0046 | 100.828 | 118.5868 | 21.2612 | 0.003226 | 0.002896 | 224.0801 | 0.2661 | 8.4017 | 0.003313 | 0.005099 | 0.004976 | 6.4e-05 | 0.001003 | 0.32244 | 1.3148 | 2.5481 | 101.19 | 20.81 | 17.562 | 99.467 | 101.5489 | 91.4916 | 0.004271 | 0.9682 | 0.023723 | 24.5027 | 0.015648 | 0.1583 | 0.015778 | 1.1236 | 1.1236 | 0.005861
2019-03-19 15:56 | 1553025360.7747 | 831 | 0.20642 | 0.002418 | 7.3e-05 | 101.0094 | 100.9732 | 118.6622 | 21.2612 | 0.003251 | 0.002968 | 224.0636 | 0.26494 | 8.4031 | 0.003427 | 0.004799 | 0.004732 | 6.4e-05 | 0.001012 | 0.30359 | 1.328 | 2.5481 | 101.19 | 20.81 | 17.562 | 100.0483 | 104.3099 | 93.22 | 0.004271 | 0.96363 | 0.023675 | 24.5685 | 0.0032 | 0.15761 | 0.015501 | 1.1016 | 1.1016 | 0.005952
2019-03-19 15:56 | 1553025361.7748 | 832 | 0.20651 | 0.002403 | 7.4e-05 | 101.0781 | 101.0513 | 119.0103 | 21.2612 | 0.003235 | 0.002944 | 224.0406 | 0.26575 | 8.4017 | 0.003355 | 0.004853 | 0.004747 | 6e-05 | 0.001011 | 0.31094 | 1.303 | 2.5481 | 101.19 | 20.81 | 17.562 | 99.6655 | 102.5808 | 92.6599 | 0.004271 | 0.96511 | 0.023728 | 24.5857 | 0.002369 | 0.1581 | 0.014963 | 1.0656 | 1.0656 | 0.006172
2019-03-19 15:56 | 1553025362.7749 | 833 | 0.20644 | 0.002388 | 7.4e-05 | 101.1147 | 101.0082 | 119.1999 | 21.2612 | 0.003253 | 0.002989 | 224.0045 | 0.26578 | 8.4045 | 0.003317 | 0.004861 | 0.004749 | 6.3e-05 | 0.00101 | 0.29999 | 1.3327 | 2.5481 | 101.19 | 20.81 | 17.563 | 100.1165 | 101.6612 | 93.7293 | 0.004271 | 0.96229 | 0.023712 | 24.6412 | 0.009429 | 0.15812 | 0.015416 | 1.0973 | 1.0973 | 0.005994
2019-03-19 15:56 | 1553025363.7739 | 834 | 0.20647 | 0.002372 | 7.5e-05 | 101.0853 | 100.8849 | 119.3769 | 21.2612 | 0.003239 | 0.002912 | 223.9686 | 0.26555 | 8.4067 | 0.003353 | 0.005013 | 0.004883 | 5.9e-05 | 0.001013 | 0.28766 | 1.364 | 2.5481 | 101.19 | 20.81 | 17.564 | 99.7697 | 102.5377 | 91.8849 | 0.004271 | 0.96715 | 0.02379 | 24.5975 | 0.017752 | 0.15798 | 0.015222 | 1.0869 | 1.0869 | 0.006046
2019-03-19 15:56 | 1553025364.7738 | 835 | 0.20647 | 0.002349 | 7.5e-05 | 101.1519 | 101.1534 | 119.4706 | 21.2612 | 0.00323 | 0.00301 | 223.9258 | 0.26591 | 8.403 | 0.003369 | 0.00518 | 0.005121 | 5.9e-05 | 0.001008 | 0.32101 | 1.3575 | 2.548 | 101.19 | 20.81 | 17.563 | 99.5492 | 102.9147 | 94.2334 | 0.004271 | 0.96097 | 0.023723 | 24.6861 | -0.00013 | 0.15819 | 0.015281 | 1.0881 | 1.0881 | 0.006048
2019-03-19 15:56 | 1553025365.7753 | 836 | 0.20655 | 0.002337 | 7.5e-05 | 101.203 | 101.0647 | 119.4985 | 21.2612 | 0.003248 | 0.002953 | 223.8992 | 0.26573 | 8.4025 | 0.003395 | 0.005449 | 0.005345 | 5.9e-05 | 0.001009 | 0.31957 | 1.3411 | 2.5481 | 101.19 | 20.81 | 17.559 | 99.9973 | 103.5544 | 92.8725 | 0.00427 | 0.96455 | 0.02377 | 24.6432 | 0.012228 | 0.15808 | 0.014822 | 1.0574 | 1.0574 | 0.006219
2019-03-19 15:56 | 1553025366.775 | 837 | 0.20657 | 0.002316 | 7.5e-05 | 101.1555 | 101.2377 | 119.6978 | 21.2612 | 0.003248 | 0.002981 | 223.8642 | 0.26584 | 8.4009 | 0.003416 | 0.004915 | 0.004765 | 6e-05 | 0.001006 | 0.28423 | 1.3377 | 2.5482 | 101.19 | 20.81 | 17.554 | 99.9965 | 104.0615 | 93.5493 | 0.004269 | 0.96276 | 0.023767 | 24.6866 | -0.007258 | 0.15815 | 0.014698 | 1.0484 | 1.0484 | 0.006275
2019-03-19 15:56 | 1553025367.7751 | 838 | 0.20664 | 0.002295 | 7.5e-05 | 101.2258 | 101.2681 | 119.5954 | 21.2612 | 0.003201 | 0.002917 | 223.8363 | 0.26606 | 8.406 | 0.003416 | 0.005148 | 0.005099 | 6e-05 | 0.001007 | 0.31492 | 1.3303 | 2.5482 | 101.19 | 20.81 | 17.547 | 98.8567 | 104.0537 | 91.9944 | 0.004267 | 0.96686 | 0.023808 | 24.6236 | -0.003743 | 0.15828 | 0.014301 | 1.0217 | 1.0217 | 0.006444
2019-03-19 15:56 | 1553025368.7755 | 839 | 0.20658 | 0.002278 | 7.6e-05 | 101.26 | 101.1096 | 119.5307 | 21.2612 | 0.003237 | 0.002967 | 223.7894 | 0.26584 | 8.4009 | 0.003425 | 0.004577 | 0.004415 | 6e-05 | 0.001009 | 0.32156 | 1.3256 | 2.5482 | 101.19 | 20.82 | 17.53 | 99.7156 | 104.2706 | 93.1973 | 0.004266 | 0.96369 | 0.023762 | 24.6575 | 0.013296 | 0.15815 | 0.014724 | 1.05 | 1.05 | 0.006266
2019-03-19 15:56 | 1553025369.7746 | 840 | 0.20665 | 0.002264 | 7.6e-05 | 101.2951 | 101.213 | 119.5365 | 21.2612 | 0.003185 | 0.002915 | 223.7663 | 0.26664 | 8.4036 | 0.003526 | 0.004415 | 0.004283 | 6e-05 | 0.001008 | 0.31094 | 1.3209 | 2.548 | 101.19 | 20.82 | 17.518 | 98.4758 | 106.7229 | 91.9483 | 0.004263 | 0.96699 | 0.023803 | 24.616 | 0.007252 | 0.15862 | 0.014307 | 1.0219 | 1.0219 | 0.006457
2019-03-19 15:56 | 1553025370.7752 | 841 | 0.2066 | 0.002239 | 7.6e-05 | 101.263 | 101.2685 | 119.6966 | 21.2612 | 0.003194 | 0.002856 | 223.7152 | 0.26448 | 8.4014 | 0.003485 | 0.005094 | 0.005004 | 6e-05 | 0.001009 | 0.27649 | 1.3148 | 2.5482 | 101.19 | 20.82 | 17.504 | 98.6771 | 105.7179 | 90.5255 | 0.004259 | 0.97077 | 0.023866 | 24.5845 | -0.000486 | 0.15734 | 0.014633 | 1.048 | 1.048 | 0.006246
2019-03-19 15:56 | 1553025371.7743 | 842 | 0.20676 | 0.002217 | 7.6e-05 | 101.2927 | 101.3798 | 119.4119 | 21.2612 | 0.003131 | 0.002952 | 223.6519 | 0.26644 | 8.4043 | 0.003484 | 0.005366 | 0.005238 | 6.2e-05 | 0.00101 | 0.29822 | 1.3042 | 2.5481 | 101.19 | 20.82 | 17.472 | 97.1589 | 105.7029 | 92.8342 | 0.004252 | 0.96465 | 0.023762 | 24.633 | -0.007688 | 0.15851 | 0.013703 | 0.97697 | 0.97697 | 0.006749
2019-03-19 15:56 | 1553025372.7753 | 843 | 0.2068 | 0.0022 | 7.5e-05 | 101.3077 | 101.3945 | 119.3621 | 21.2612 | 0.003187 | 0.002838 | 223.6401 | 0.26604 | 8.4021 | 0.003485 | 0.005561 | 0.005486 | 6.2e-05 | 0.001013 | 0.31902 | 1.3547 | 2.5481 | 101.19 | 20.82 | 17.456 | 98.5086 | 105.7361 | 90.0987 | 0.004248 | 0.97191 | 0.023846 | 24.5357 | -0.007663 | 0.15827 | 0.01347 | 0.96371 | 0.96371 | 0.006832
2019-03-19 15:56 | 1553025373.7758 | 844 | 0.20685 | 0.002178 | 7.6e-05 | 101.3683 | 101.3784 | 119.4257 | 21.2612 | 0.003166 | 0.002909 | 223.5879 | 0.26608 | 8.4054 | 0.003514 | 0.005803 | 0.005671 | 6e-05 | 0.001014 | 0.30342 | 1.3436 | 2.5479 | 101.19 | 20.82 | 17.421 | 98.0147 | 106.4206 | 91.7972 | 0.004239 | 0.96739 | 0.023797 | 24.5995 | -0.000892 | 0.15829 | 0.013228 | 0.9444 | 0.9444 | 0.006973
2019-03-19 15:56 | 1553025374.7755 | 845 | 0.20689 | 0.002157 | 7.6e-05 | 101.3932 | 101.3 | 119.6267 | 21.2612 | 0.003124 | 0.00289 | 223.5493 | 0.2663 | 8.4017 | 0.003547 | 0.005342 | 0.005323 | 5.7e-05 | 0.001018 | 0.33987 | 1.3617 | 2.5481 | 101.19 | 20.82 | 17.404 | 96.999 | 107.2228 | 91.3415 | 0.004235 | 0.9686 | 0.023832 | 24.6048 | 0.008228 | 0.15842 | 0.013002 | 0.92957 | 0.92957 | 0.00709
2019-03-19 15:56 | 1553025375.7755 | 846 | 0.20696 | 0.002132 | 7.6e-05 | 101.3959 | 101.511 | 119.2875 | 21.2612 | 0.003136 | 0.002937 | 223.5089 | 0.26581 | 8.4016 | 0.00362 | 0.005109 | 0.005009 | 5.7e-05 | 0.001021 | 0.33323 | 1.3514 | 2.5482 | 101.19 | 20.82 | 17.389 | 97.2734 | 108.9955 | 92.4699 | 0.004231 | 0.96561 | 0.023762 | 24.6079 | -0.010143 | 0.15813 | 0.012606 | 0.89844 | 0.89844 | 0.007322
2019-03-19 15:56 | 1553025376.7755 | 847 | 0.20691 | 0.002106 | 7.7e-05 | 101.4283 | 101.2476 | 119.2401 | 21.2612 | 0.00314 | 0.002836 | 223.4721 | 0.26602 | 8.4027 | 0.003564 | 0.005211 | 0.005104 | 5.7e-05 | 0.001016 | 0.3147 | 1.3742 | 2.548 | 101.19 | 20.82 | 17.357 | 97.3862 | 107.634 | 90.0471 | 0.004224 | 0.97205 | 0.023836 | 24.5214 | 0.015953 | 0.15826 | 0.012932 | 0.92467 | 0.92467 | 0.00712
2019-03-19 15:56 | 1553025377.7749 | 848 | 0.20689 | 0.002085 | 7.7e-05 | 101.4895 | 101.3363 | 119.2856 | 21.2612 | 0.003161 | 0.002909 | 223.4269 | 0.26592 | 8.4005 | 0.003539 | 0.004907 | 0.004829 | 5.7e-05 | 0.001019 | 0.33445 | 1.3579 | 2.5482 | 101.19 | 20.82 | 17.344 | 97.8768 | 107.0347 | 91.8114 | 0.004221 | 0.96735 | 0.023783 | 24.5856 | 0.013517 | 0.15819 | 0.013051 | 0.93111 | 0.93111 | 0.007068
2019-03-19 15:56 | 1553025378.7747 | 849 | 0.20698 | 0.002061 | 7.8e-05 | 101.4628 | 101.4281 | 119.4128 | 21.2612 | 0.003071 | 0.002917 | 223.4108 | 0.2662 | 8.4036 | 0.003576 | 0.005286 | 0.005162 | 5.8e-05 | 0.001014 | 0.29966 | 1.3606 | 2.5481 | 101.19 | 20.82 | 17.321 | 95.7011 | 107.9443 | 92.0052 | 0.004215 | 0.96684 | 0.023789 | 24.6052 | 0.003065 | 0.15836 | 0.012563 | 0.89638 | 0.89638 | 0.007349
2019-03-19 15:56 | 1553025379.7737 | 850 | 0.20694 | 0.002039 | 7.8e-05 | 101.4328 | 101.6199 | 119.4478 | 21.2612 | 0.003207 | 0.002882 | 223.3768 | 0.26573 | 8.4042 | 0.003592 | 0.005567 | 0.005505 | 5.5e-05 | 0.001013 | 0.34435 | 1.37 | 2.5481 | 101.19 | 20.82 | 17.311 | 98.9995 | 108.3185 | 91.143 | 0.004212 | 0.96913 | 0.023821 | 24.5797 | -0.016479 | 0.15808 | 0.012812 | 0.91543 | 0.91543 | 0.007184
2019-03-19 15:56 | 1553025380.7747 | 851 | 0.20697 | 0.002013 | 7.8e-05 | 101.4935 | 101.688 | 119.6105 | 21.2612 | 0.003168 | 0.002875 | 223.3523 | 0.26578 | 8.405 | 0.003689 | 0.00489 | 0.004774 | 5.6e-05 | 0.001012 | 0.31957 | 1.3778 | 2.5481 | 101.19 | 20.82 | 17.303 | 98.0524 | 110.6801 | 90.9897 | 0.004211 | 0.96953 | 0.023842 | 24.5913 | -0.017131 | 0.15812 | 0.012669 | 0.90595 | 0.90595 | 0.00726
2019-03-19 15:56 | 1553025381.7752 | 852 | 0.20699 | 0.001996 | 7.8e-05 | 101.501 | 101.5216 | 119.9304 | 21.2612 | 0.003163 | 0.00285 | 223.325 | 0.26548 | 8.4046 | 0.003645 | 0.005705 | 0.005638 | 5.6e-05 | 0.001011 | 0.33411 | 1.335 | 2.5481 | 101.19 | 20.82 | 17.291 | 97.9418 | 109.6066 | 90.3918 | 0.004208 | 0.97113 | 0.023894 | 24.6039 | -0.001817 | 0.15793 | 0.012611 | 0.90368 | 0.90368 | 0.00727
2019-03-19 15:56 | 1553025382.7738 | 853 | 0.20701 | 0.00198 | 7.9e-05 | 101.5823 | 101.6418 | 119.8619 | 21.2612 | 0.003188 | 0.002887 | 223.2887 | 0.2665 | 8.4007 | 0.003558 | 0.005786 | 0.005614 | 5.4e-05 | 0.001012 | 0.30862 | 1.3358 | 2.5481 | 101.19 | 20.82 | 17.287 | 98.5454 | 107.4954 | 91.2859 | 0.004207 | 0.96874 | 0.023857 | 24.6271 | -0.005239 | 0.15854 | 0.012509 | 0.89498 | 0.89498 | 0.007369
2019-03-19 15:56 | 1553025383.7743 | 854 | 0.20695 | 0.00196 | 7.9e-05 | 101.5859 | 101.7254 | 119.8642 | 21.2612 | 0.00318 | 0.002933 | 223.2738 | 0.26585 | 8.4033 | 0.003568 | 0.005002 | 0.004975 | 5.5e-05 | 0.001008 | 0.31813 | 1.3339 | 2.5481 | 101.19 | 20.81 | 17.281 | 98.3542 | 107.7443 | 92.3783 | 0.004203 | 0.96585 | 0.023822 | 24.6643 | -0.012272 | 0.15815 | 0.012857 | 0.91862 | 0.91862 | 0.007162
2019-03-19 15:56 | 1553025384.774 | 855 | 0.20706 | 0.00194 | 8e-05 | 101.5049 | 101.588 | 119.7806 | 21.2612 | 0.00317 | 0.002904 | 223.2518 | 0.26607 | 8.4029 | 0.003698 | 0.005179 | 0.005077 | 5.5e-05 | 0.001009 | 0.32421 | 1.3251 | 2.5481 | 101.19 | 20.81 | 17.28 | 98.1034 | 110.8907 | 91.6753 | 0.004202 | 0.96771 | 0.023837 | 24.6319 | -0.007325 | 0.15829 | 0.012252 | 0.87575 | 0.87575 | 0.007519
2019-03-19 15:56 | 1553025385.7764 | 856 | 0.20701 | 0.001923 | 8e-05 | 101.5829 | 101.6024 | 120.124 | 21.2612 | 0.003142 | 0.00286 | 223.226 | 0.26476 | 8.4046 | 0.003692 | 0.005821 | 0.005715 | 5.5e-05 | 0.001012 | 0.33627 | 1.3185 | 2.5482 | 101.19 | 20.81 | 17.279 | 97.4362 | 110.7614 | 90.6216 | 0.004202 | 0.97051 | 0.023905 | 24.6316 | -0.001719 | 0.15751 | 0.012572 | 0.90128 | 0.90128 | 0.00727
2019-03-19 15:56 | 1553025386.775 | 857 | 0.20701 | 0.001913 | 8.1e-05 | 101.591 | 101.5904 | 120.0778 | 21.2612 | 0.003134 | 0.00284 | 223.1978 | 0.2655 | 8.4055 | 0.003648 | 0.00502 | 0.004831 | 5.6e-05 | 0.001009 | 0.29015 | 1.3154 | 2.5481 | 101.19 | 20.81 | 17.274 | 97.2239 | 109.68 | 90.1541 | 0.004201 | 0.97176 | 0.023916 | 24.611 | 5e-05 | 0.15795 | 0.012549 | 0.89998 | 0.89998 | 0.007301
2019-03-19 15:56 | 1553025387.7745 | 858 | 0.20713 | 0.001895 | 8.1e-05 | 101.6234 | 101.8404 | 119.9811 | 21.2612 | 0.003107 | 0.00282 | 223.1735 | 0.26558 | 8.4011 | 0.003677 | 0.005651 | 0.005549 | 5.4e-05 | 0.001009 | 0.33804 | 1.3305 | 2.5481 | 101.19 | 20.81 | 17.274 | 96.5815 | 110.3772 | 89.6495 | 0.004201 | 0.97311 | 0.023923 | 24.584 | -0.019083 | 0.158 | 0.011889 | 0.85262 | 0.85262 | 0.007709
2019-03-19 15:56 | 1553025388.7752 | 859 | 0.2072 | 0.001881 | 8.1e-05 | 101.5604 | 101.7158 | 119.9325 | 21.2612 | 0.003108 | 0.002849 | 223.1315 | 0.26569 | 8.4033 | 0.003729 | 0.004851 | 0.004756 | 5.4e-05 | 0.001 | 0.31022 | 1.2984 | 2.548 | 101.19 | 20.81 | 17.273 | 96.5941 | 111.6658 | 90.3542 | 0.004201 | 0.97123 | 0.023895 | 24.6029 | -0.013675 | 0.15806 | 0.011468 | 0.82139 | 0.82139 | 0.008005
2019-03-19 15:56 | 1553025389.7752 | 860 | 0.20725 | 0.001864 | 8.2e-05 | 101.6369 | 101.5853 | 119.7151 | 21.2612 | 0.003094 | 0.002832 | 223.1036 | 0.26502 | 8.4055 | 0.003624 | 0.005863 | 0.005726 | 5.5e-05 | 0.000996 | 0.31868 | 1.3235 | 2.5481 | 101.19 | 20.81 | 17.276 | 96.2595 | 109.1068 | 89.9556 | 0.004201 | 0.97229 | 0.023886 | 24.5671 | 0.004544 | 0.15766 | 0.011208 | 0.80234 | 0.80234 | 0.008174
2019-03-19 15:56 | 1553025390.7752 | 861 | 0.20703 | 0.001848 | 8.3e-05 | 101.6105 | 101.6942 | 119.1612 | 21.2612 | 0.00311 | 0.002847 | 223.0622 | 0.26525 | 8.4027 | 0.003618 | 0.004796 | 0.004762 | 5.4e-05 | 0.000997 | 0.32139 | 1.327 | 2.5482 | 101.19 | 20.81 | 17.277 | 96.645 | 108.961 | 90.3183 | 0.004202 | 0.97132 | 0.023819 | 24.5224 | -0.007365 | 0.1578 | 0.012531 | 0.89487 | 0.89487 | 0.007336
2019-03-19 15:56 | 1553025391.7747 | 862 | 0.20722 | 0.00182 | 8.4e-05 | 101.6324 | 101.7141 | 118.9913 | 21.2612 | 0.003106 | 0.002811 | 223.0472 | 0.26578 | 8.4018 | 0.003586 | 0.00494 | 0.004751 | 5.3e-05 | 0.000992 | 0.31166 | 1.3064 | 2.5481 | 101.19 | 20.81 | 17.284 | 96.5494 | 108.176 | 89.4385 | 0.004203 | 0.97368 | 0.023831 | 24.4753 | -0.007184 | 0.15811 | 0.011432 | 0.81641 | 0.81641 | 0.008057
2019-03-19 15:57 | 1553025392.7744 | 863 | 0.20719 | 0.001806 | 8.5e-05 | 101.6837 | 101.6182 | 118.9863 | 21.2612 | 0.003093 | 0.002816 | 223.0002 | 0.26603 | 8.4037 | 0.003653 | 0.004701 | 0.004578 | 5.3e-05 | 0.000994 | 0.29479 | 1.3468 | 2.5481 | 101.19 | 20.81 | 17.287 | 96.243 | 109.8069 | 89.5712 | 0.004204 | 0.97332 | 0.023826 | 24.4792 | 0.005768 | 0.15826 | 0.011591 | 0.8276 | 0.8276 | 0.007955
2019-03-19 15:57 | 1553025393.7744 | 864 | 0.20722 | 0.001785 | 8.5e-05 | 101.6501 | 101.6781 | 118.987 | 21.2612 | 0.003071 | 0.002793 | 222.9665 | 0.26586 | 8.4015 | 0.003549 | 0.0049 | 0.004792 | 5.7e-05 | 0.000997 | 0.29695 | 1.3526 | 2.548 | 101.19 | 20.81 | 17.293 | 95.7231 | 107.2789 | 89.0118 | 0.004206 | 0.97483 | 0.023845 | 24.4604 | -0.002463 | 0.15816 | 0.011473 | 0.81975 | 0.81975 | 0.008026
2019-03-19 15:57 | 1553025394.7739 | 865 | 0.20725 | 0.001772 | 8.6e-05 | 101.6657 | 101.8363 | 118.988 | 21.2612 | 0.003043 | 0.002823 | 222.9211 | 0.26616 | 8.4042 | 0.003546 | 0.004875 | 0.00477 | 5.4e-05 | 0.001002 | 0.33588 | 1.3408 | 2.548 | 101.19 | 20.81 | 17.305 | 95.0317 | 107.2143 | 89.729 | 0.004208 | 0.9729 | 0.023821 | 24.4847 | -0.014998 | 0.15834 | 0.011266 | 0.80405 | 0.80405 | 0.008192
2019-03-19 15:57 | 1553025395.7746 | 866 | 0.20725 | 0.001762 | 8.6e-05 | 101.5859 | 101.6891 | 119.0359 | 21.2612 | 0.003056 | 0.002846 | 222.8734 | 0.26674 | 8.4028 | 0.003549 | 0.00537 | 0.005264 | 5.7e-05 | 0.001002 | 0.33467 | 1.3419 | 2.5481 | 101.19 | 20.81 | 17.311 | 95.3425 | 107.2915 | 90.2839 | 0.00421 | 0.97142 | 0.023808 | 24.5084 | -0.00908 | 0.15868 | 0.011322 | 0.80761 | 0.80761 | 0.008174
2019-03-19 15:57 | 1553025396.7746 | 867 | 0.2073 | 0.001746 | 8.7e-05 | 101.6435 | 101.7726 | 119.0638 | 21.2612 | 0.003039 | 0.002766 | 222.8463 | 0.26466 | 8.4035 | 0.003506 | 0.004878 | 0.004844 | 5.6e-05 | 0.000994 | 0.31885 | 1.3199 | 2.5482 | 101.19 | 20.81 | 17.323 | 94.9409 | 106.227 | 88.3534 | 0.004213 | 0.9766 | 0.023874 | 24.4461 | -0.011355 | 0.15745 | 0.011033 | 0.78898 | 0.78898 | 0.008302
2019-03-19 15:57 | 1553025397.775 | 868 | 0.20735 | 0.00173 | 8.8e-05 | 101.639 | 101.7723 | 119.1727 | 21.2612 | 0.003042 | 0.002836 | 222.8238 | 0.26551 | 8.4024 | 0.003456 | 0.005225 | 0.005104 | 5.3e-05 | 0.000991 | 0.31277 | 1.2985 | 2.548 | 101.19 | 20.81 | 17.328 | 95.0114 | 105.0239 | 90.0394 | 0.004214 | 0.97207 | 0.02383 | 24.5142 | -0.011721 | 0.15795 | 0.010729 | 0.76564 | 0.76564 | 0.008582
2019-03-19 15:57 | 1553025398.7741 | 869 | 0.2073 | 0.001712 | 9e-05 | 101.696 | 101.7147 | 119.0995 | 21.2612 | 0.003064 | 0.002809 | 222.8036 | 0.26541 | 8.4032 | 0.003463 | 0.005903 | 0.005717 | 5.3e-05 | 0.000986 | 0.30127 | 1.3415 | 2.5481 | 101.19 | 20.81 | 17.332 | 95.5327 | 105.1876 | 89.4055 | 0.004215 | 0.97377 | 0.023843 | 24.4853 | -0.001646 | 0.1579 | 0.011041 | 0.78842 | 0.78842 | 0.008331
2019-03-19 15:57 | 1553025399.7738 | 870 | 0.20731 | 0.001701 | 9.1e-05 | 101.6327 | 101.6832 | 119.1143 | 21.2612 | 0.003068 | 0.002807 | 222.7721 | 0.26557 | 8.4036 | 0.003488 | 0.005877 | 0.005676 | 5.3e-05 | 0.000993 | 0.32621 | 1.3591 | 2.5481 | 101.19 | 20.81 | 17.335 | 95.6376 | 105.7919 | 89.3408 | 0.004216 | 0.97394 | 0.023847 | 24.4846 | -0.004446 | 0.15799 | 0.011006 | 0.7859 | 0.7859 | 0.008363
2019-03-19 15:57 | 1553025400.7751 | 871 | 0.20723 | 0.00169 | 9.2e-05 | 101.6726 | 101.7466 | 119.2086 | 21.2612 | 0.003051 | 0.002803 | 222.753 | 0.26497 | 8.4048 | 0.003488 | 0.004887 | 0.004878 | 5.3e-05 | 0.000995 | 0.32045 | 1.3738 | 2.5481 | 101.19 | 20.81 | 17.335 | 95.225 | 105.8115 | 89.2565 | 0.004216 | 0.97417 | 0.023859 | 24.4915 | -0.006506 | 0.15763 | 0.011514 | 0.82271 | 0.82271 | 0.00797
2019-03-19 15:57 | 1553025401.7739 | 872 | 0.20742 | 0.001673 | 9.4e-05 | 101.6339 | 101.7877 | 119.0673 | 21.2612 | 0.003095 | 0.002808 | 222.7308 | 0.26649 | 8.4036 | 0.003415 | 0.004736 | 0.004572 | 5.5e-05 | 0.000989 | 0.32228 | 1.3586 | 2.5481 | 101.19 | 20.82 | 17.331 | 96.2988 | 104.031 | 89.3627 | 0.004217 | 0.97388 | 0.023841 | 24.4805 | -0.013525 | 0.15853 | 0.010369 | 0.73979 | 0.73979 | 0.008915
2019-03-19 15:57 | 1553025402.7753 | 873 | 0.20735 | 0.001659 | 9.6e-05 | 101.6855 | 101.6579 | 119.038 | 21.2612 | 0.002993 | 0.002698 | 222.6972 | 0.26591 | 8.4041 | 0.003514 | 0.004554 | 0.004418 | 5.4e-05 | 0.000986 | 0.29175 | 1.3183 | 2.5482 | 101.19 | 20.82 | 17.328 | 93.8255 | 106.4419 | 86.7269 | 0.004217 | 0.98102 | 0.023925 | 24.3884 | 0.00243 | 0.15819 | 0.010841 | 0.77632 | 0.77632 | 0.008477
2019-03-19 15:57 | 1553025403.774 | 874 | 0.20724 | 0.00164 | 9.8e-05 | 101.5808 | 101.7897 | 119.1093 | 21.2612 | 0.003068 | 0.00285 | 222.6584 | 0.26558 | 8.4012 | 0.003524 | 0.004947 | 0.004799 | 5.3e-05 | 0.000983 | 0.32549 | 1.3378 | 2.548 | 101.19 | 20.81 | 17.326 | 95.6338 | 106.6843 | 90.3761 | 0.004214 | 0.97117 | 0.023812 | 24.519 | -0.018381 | 0.15799 | 0.011535 | 0.82232 | 0.82232 | 0.007993
2019-03-19 15:57 | 1553025404.775 | 875 | 0.20742 | 0.001627 | 9.9e-05 | 101.603 | 101.7264 | 119.2114 | 21.2612 | 0.003026 | 0.002796 | 222.6391 | 0.26578 | 8.4008 | 0.003558 | 0.005189 | 0.005092 | 5.1e-05 | 0.000988 | 0.28368 | 1.3043 | 2.5481 | 101.19 | 20.81 | 17.321 | 94.6246 | 107.4874 | 89.0745 | 0.004212 | 0.97466 | 0.023865 | 24.4856 | -0.010857 | 0.15811 | 0.01046 | 0.74679 | 0.74679 | 0.008808
2019-03-19 15:57 | 1553025405.7751 | 876 | 0.20746 | 0.001618 | 0.000101 | 101.5901 | 101.6986 | 119.2924 | 21.2612 | 0.003091 | 0.0028 | 222.6213 | 0.26581 | 8.4036 | 0.003531 | 0.004892 | 0.004806 | 5.2e-05 | 0.00099 | 0.32709 | 1.3341 | 2.548 | 101.19 | 20.81 | 17.321 | 96.1945 | 106.8372 | 89.1727 | 0.004212 | 0.9744 | 0.02387 | 24.4972 | -0.009553 | 0.15813 | 0.010209 | 0.72879 | 0.72879 | 0.009026
2019-03-19 15:57 | 1553025406.7746 | 877 | 0.20752 | 0.001603 | 0.000102 | 101.5901 | 101.5329 | 119.2551 | 21.2612 | 0.003059 | 0.002804 | 222.5829 | 0.26601 | 8.404 | 0.003639 | 0.005192 | 0.005075 | 5.6e-05 | 0.000993 | 0.30071 | 1.3385 | 2.5481 | 101.19 | 20.81 | 17.319 | 95.4168 | 109.4703 | 89.2837 | 0.004212 | 0.9741 | 0.023863 | 24.4971 | 0.005038 | 0.15825 | 0.009873 | 0.70435 | 0.70435 | 0.009346
2019-03-19 15:57 | 1553025407.7734 | 878 | 0.20756 | 0.001591 | 0.000104 | 101.606 | 101.7336 | 119.4931 | 21.2612 | 0.003062 | 0.002821 | 222.5425 | 0.26484 | 8.4032 | 0.003622 | 0.0051 | 0.005078 | 5.6e-05 | 0.000992 | 0.33859 | 1.3461 | 2.5482 | 101.19 | 20.81 | 17.318 | 95.4998 | 109.0548 | 89.6962 | 0.004212 | 0.97299 | 0.023873 | 24.5355 | -0.011225 | 0.15755 | 0.00964 | 0.68775 | 0.68775 | 0.00953
2019-03-19 15:57 | 1553025408.7732 | 879 | 0.20749 | 0.00157 | 0.000105 | 101.5286 | 101.5089 | 119.7203 | 21.2612 | 0.003018 | 0.002806 | 222.5165 | 0.26634 | 8.4025 | 0.003574 | 0.005084 | 0.004932 | 5.2e-05 | 0.000992 | 0.3027 | 1.4016 | 2.5482 | 101.19 | 20.81 | 17.32 | 94.4423 | 107.879 | 89.3196 | 0.004212 | 0.974 | 0.023908 | 24.5461 | 0.001732 | 0.15845 | 0.010088 | 0.721 | 0.721 | 0.009142
2019-03-19 15:57 | 1553025409.7733 | 880 | 0.20743 | 0.00156 | 0.000106 | 101.4931 | 101.5822 | 119.7813 | 21.2612 | 0.003028 | 0.002754 | 222.4939 | 0.26597 | 8.404 | 0.003584 | 0.005011 | 0.004932 | 5.5e-05 | 0.000994 | 0.30486 | 1.3483 | 2.5482 | 101.19 | 20.82 | 17.331 | 94.6833 | 108.1247 | 88.0674 | 0.004217 | 0.97738 | 0.023955 | 24.5099 | -0.007844 | 0.15823 | 0.010452 | 0.74859 | 0.74859 | 0.008793
2019-03-19 15:57 | 1553025410.7737 | 881 | 0.20758 | 0.001547 | 0.000108 | 101.5088 | 101.5288 | 120.2905 | 21.2612 | 0.003083 | 0.002722 | 222.4561 | 0.26602 | 8.4015 | 0.003583 | 0.004788 | 0.00463 | 5.4e-05 | 0.000996 | 0.3454 | 1.3416 | 2.548 | 101.19 | 20.82 | 17.34 | 95.9956 | 108.098 | 87.3117 | 0.00422 | 0.97943 | 0.024031 | 24.5362 | -0.001763 | 0.15826 | 0.009568 | 0.68688 | 0.68688 | 0.009584
2019-03-19 15:57 | 1553025411.7744 | 882 | 0.20756 | 0.001533 | 0.000111 | 101.5688 | 101.5788 | 120.5808 | 21.2612 | 0.003034 | 0.002733 | 222.4241 | 0.26574 | 8.4046 | 0.003562 | 0.005605 | 0.005508 | 5.3e-05 | 0.001001 | 0.31598 | 1.3478 | 2.5481 | 101.19 | 20.82 | 17.351 | 94.8218 | 107.601 | 87.574 | 0.004222 | 0.97871 | 0.024052 | 24.5748 | -0.00088 | 0.15809 | 0.009694 | 0.69644 | 0.69644 | 0.009443
2019-03-19 15:57 | 1553025412.7744 | 883 | 0.20761 | 0.00151 | 0.000114 | 101.4028 | 101.5271 | 120.4308 | 21.2612 | 0.003006 | 0.002745 | 222.4136 | 0.26617 | 8.4046 | 0.00349 | 0.005082 | 0.005042 | 5.5e-05 | 0.001004 | 0.31133 | 1.3057 | 2.548 | 101.19 | 20.81 | 17.372 | 94.1431 | 105.8532 | 87.8558 | 0.004225 | 0.97795 | 0.024027 | 24.5691 | -0.010952 | 0.15835 | 0.009411 | 0.67502 | 0.67502 | 0.009759
2019-03-19 15:57 | 1553025413.7745 | 884 | 0.20751 | 0.001495 | 0.000117 | 101.4727 | 101.4887 | 120.0909 | 21.2612 | 0.002989 | 0.00267 | 222.3917 | 0.26565 | 8.4017 | 0.00358 | 0.005643 | 0.005499 | 5.4e-05 | 0.001003 | 0.30558 | 1.3298 | 2.548 | 101.19 | 20.81 | 17.382 | 93.7411 | 108.0349 | 86.0438 | 0.004227 | 0.98288 | 0.024054 | 24.4727 | -0.001406 | 0.15803 | 0.01006 | 0.72266 | 0.72266 | 0.009097
2019-03-19 15:57 | 1553025414.7739 | 885 | 0.20749 | 0.001479 | 0.000119 | 101.3602 | 101.512 | 119.6642 | 21.2612 | 0.003004 | 0.002776 | 222.3811 | 0.26557 | 8.4017 | 0.003535 | 0.005591 | 0.005549 | 5.2e-05 | 0.001001 | 0.33428 | 1.3255 | 2.5481 | 101.19 | 20.82 | 17.399 | 94.1065 | 106.9355 | 88.6057 | 0.004234 | 0.97592 | 0.023926 | 24.5162 | -0.013384 | 0.15799 | 0.010158 | 0.72571 | 0.72571 | 0.009056
2019-03-19 15:57 | 1553025415.7736 | 886 | 0.20755 | 0.001459 | 0.000122 | 101.3941 | 101.6281 | 119.1626 | 21.2612 | 0.002932 | 0.002726 | 222.375 | 0.2655 | 8.4033 | 0.00349 | 0.005253 | 0.005111 | 5.4e-05 | 0.001001 | 0.33268 | 1.3455 | 2.5482 | 101.19 | 20.82 | 17.406 | 92.3706 | 105.8501 | 87.4077 | 0.004236 | 0.97916 | 0.023915 | 24.4242 | -0.020616 | 0.15795 | 0.00987 | 0.7045 | 0.7045 | 0.009327
2019-03-19 15:57 | 1553025416.7737 | 887 | 0.2074 | 0.001443 | 0.000124 | 101.3482 | 101.4007 | 119.364 | 21.2612 | 0.003019 | 0.002752 | 222.3502 | 0.26493 | 8.4046 | 0.003556 | 0.005053 | 0.004961 | 5.2e-05 | 0.000998 | 0.30309 | 1.3598 | 2.5481 | 101.19 | 20.82 | 17.412 | 94.4468 | 107.4418 | 88.0323 | 0.004237 | 0.97747 | 0.023915 | 24.466 | -0.00463 | 0.15761 | 0.010748 | 0.76755 | 0.76755 | 0.008542
2019-03-19 15:57 | 1553025417.7751 | 888 | 0.20742 | 0.001432 | 0.000127 | 101.383 | 101.3949 | 119.2225 | 21.2612 | 0.002943 | 0.002691 | 222.3381 | 0.26528 | 8.4028 | 0.003606 | 0.00481 | 0.004689 | 5.3e-05 | 0.000993 | 0.31078 | 1.35 | 2.548 | 101.19 | 20.81 | 17.422 | 92.6324 | 108.6583 | 86.5614 | 0.004237 | 0.98147 | 0.023949 | 24.4016 | -0.001045 | 0.15781 | 0.010655 | 0.76181 | 0.76181 | 0.008618
2019-03-19 15:57 | 1553025418.7747 | 889 | 0.20747 | 0.001418 | 0.000129 | 101.3368 | 101.2383 | 119.2114 | 21.2612 | 0.00299 | 0.002719 | 222.3103 | 0.26465 | 8.402 | 0.003569 | 0.004676 | 0.004543 | 5.2e-05 | 0.000993 | 0.3131 | 1.3473 | 2.5481 | 101.19 | 20.81 | 17.427 | 93.7488 | 107.771 | 87.232 | 0.004238 | 0.97964 | 0.023926 | 24.4232 | 0.008695 | 0.15744 | 0.010377 | 0.74091 | 0.74091 | 0.00884
2019-03-19 15:57 | 1553025419.7742 | 890 | 0.20758 | 0.001405 | 0.000131 | 101.3341 | 101.2849 | 119.3377 | 21.2612 | 0.002958 | 0.00277 | 222.2939 | 0.26514 | 8.4031 | 0.003421 | 0.004848 | 0.004755 | 5e-05 | 0.000991 | 0.32659 | 1.3327 | 2.548 | 101.19 | 20.81 | 17.436 | 92.9822 | 104.1771 | 88.4634 | 0.00424 | 0.97631 | 0.023898 | 24.4779 | 0.004342 | 0.15773 | 0.00972 | 0.69261 | 0.69261 | 0.009474
2019-03-19 15:57 | 1553025420.7739 | 891 | 0.20781 | 0.001394 | 0.000132 | 101.2387 | 101.4411 | 119.1814 | 21.2612 | 0.002925 | 0.002714 | 222.281 | 0.26547 | 8.4032 | 0.003449 | 0.005215 | 0.005128 | 4.9e-05 | 0.000987 | 0.32244 | 1.3309 | 2.5479 | 101.19 | 20.81 | 17.439 | 92.1975 | 104.8525 | 87.1172 | 0.004241 | 0.97995 | 0.023927 | 24.4163 | -0.017867 | 0.15793 | 0.008362 | 0.59545 | 0.59545 | 0.011033
2019-03-19 15:57 | 1553025421.7753 | 892 | 0.20777 | 0.00138 | 0.000133 | 101.2258 | 101.1692 | 119.0398 | 21.2612 | 0.002953 | 0.002693 | 222.2609 | 0.26541 | 8.4029 | 0.003441 | 0.005258 | 0.005145 | 5.4e-05 | 0.000989 | 0.3147 | 1.334 | 2.5481 | 101.19 | 20.81 | 17.44 | 92.8655 | 104.6647 | 86.5996 | 0.004241 | 0.98136 | 0.02393 | 24.3842 | 0.005003 | 0.1579 | 0.008591 | 0.61201 | 0.61201 | 0.010733
2019-03-19 15:57 | 1553025422.7744 | 893 | 0.20772 | 0.001367 | 0.000135 | 101.1648 | 101.2298 | 119.1114 | 21.2612 | 0.002908 | 0.002704 | 222.2395 | 0.26571 | 8.4055 | 0.003405 | 0.005051 | 0.004997 | 5.2e-05 | 0.000995 | 0.31614 | 1.3243 | 2.5479 | 101.19 | 20.81 | 17.441 | 91.7753 | 103.7872 | 86.8626 | 0.004242 | 0.98065 | 0.023928 | 24.4005 | -0.00574 | 0.15807 | 0.008923 | 0.63582 | 0.63582 | 0.010342
2019-03-19 15:57 | 1553025423.7731 | 894 | 0.20784 | 0.001356 | 0.000137 | 101.2027 | 101.3099 | 119.2997 | 21.2612 | 0.002977 | 0.002717 | 222.2234 | 0.26509 | 8.4032 | 0.00331 | 0.005101 | 0.004937 | 5.1e-05 | 0.000997 | 0.30397 | 1.3288 | 2.5482 | 101.19 | 20.81 | 17.44 | 93.4507 | 101.4988 | 87.1694 | 0.004241 | 0.97981 | 0.023937 | 24.4302 | -0.009475 | 0.1577 | 0.008192 | 0.58316 | 0.58316 | 0.01125
2019-03-19 15:57 | 1553025424.7742 | 895 | 0.20774 | 0.001338 | 0.000139 | 101.0841 | 101.0346 | 119.5142 | 21.2612 | 0.002952 | 0.002684 | 222.2005 | 0.26509 | 8.4025 | 0.003504 | 0.004917 | 0.004803 | 5e-05 | 0.000999 | 0.34346 | 1.3367 | 2.5479 | 101.19 | 20.81 | 17.44 | 92.8342 | 106.1916 | 86.3756 | 0.004241 | 0.98198 | 0.023985 | 24.4252 | 0.004386 | 0.1577 | 0.008827 | 0.63012 | 0.63012 | 0.010411
2019-03-19 15:57 | 1553025425.7741 | 896 | 0.20781 | 0.001323 | 0.000142 | 101.1195 | 101.0185 | 120.0048 | 21.2612 | 0.002982 | 0.002757 | 222.1592 | 0.26426 | 8.4021 | 0.0034 | 0.005456 | 0.005298 | 4.7e-05 | 0.001 | 0.32477 | 1.3366 | 2.548 | 101.19 | 20.81 | 17.445 | 93.5706 | 103.6755 | 88.1363 | 0.004242 | 0.97719 | 0.023975 | 24.5351 | 0.008944 | 0.15721 | 0.008433 | 0.6012 | 0.6012 | 0.010878
2019-03-19 15:57 | 1553025426.7734 | 897 | 0.20778 | 0.001309 | 0.000144 | 100.9812 | 101.1126 | 119.8778 | 21.2612 | 0.002983 | 0.002729 | 222.1414 | 0.26554 | 8.4043 | 0.003461 | 0.005 | 0.004941 | 5e-05 | 0.001005 | 0.31797 | 1.3728 | 2.5482 | 101.19 | 20.81 | 17.45 | 93.5814 | 105.1546 | 87.4686 | 0.004244 | 0.979 | 0.023985 | 24.4994 | -0.011637 | 0.15797 | 0.008639 | 0.61618 | 0.61618 | 0.010665
2019-03-19 15:57 | 1553025427.7743 | 898 | 0.20785 | 0.001294 | 0.000147 | 100.9794 | 100.9572 | 120.0278 | 21.2612 | 0.002932 | 0.002736 | 222.1235 | 0.26508 | 8.4017 | 0.003373 | 0.004824 | 0.00467 | 4.8e-05 | 0.001007 | 0.32693 | 1.3017 | 2.5481 | 101.19 | 20.81 | 17.46 | 92.3588 | 103.0029 | 87.6279 | 0.004246 | 0.97857 | 0.023995 | 24.5202 | 0.001969 | 0.15769 | 0.008223 | 0.58618 | 0.58618 | 0.011191
2019-03-19 15:57 | 1553025428.7747 | 899 | 0.20783 | 0.001274 | 0.000149 | 100.9572 | 100.953 | 119.8151 | 21.2612 | 0.002924 | 0.002713 | 222.1135 | 0.26446 | 8.4023 | 0.003524 | 0.004828 | 0.004833 | 4.7e-05 | 0.00101 | 0.31205 | 1.3427 | 2.5483 | 101.19 | 20.81 | 17.464 | 92.1595 | 106.6654 | 87.079 | 0.004247 | 0.98006 | 0.023992 | 24.4798 | 0.000365 | 0.15733 | 0.008356 | 0.59565 | 0.59565 | 0.010988
2019-03-19 15:57 | 1553025429.774 | 900 | 0.20795 | 0.001255 | 0.000152 | 100.8809 | 100.9818 | 120.0278 | 21.2612 | 0.002914 | 0.002682 | 222.096 | 0.2653 | 8.4043 | 0.003572 | 0.005258 | 0.005131 | 4.7e-05 | 0.001011 | 0.30055 | 1.3766 | 2.5481 | 101.19 | 20.81 | 17.464 | 91.9223 | 107.8302 | 86.3364 | 0.004247 | 0.98208 | 0.024038 | 24.4763 | -0.008939 | 0.15782 | 0.007639 | 0.54465 | 0.54465 | 0.012055
2019-03-19 15:57 | 1553025430.7734 | 901 | 0.20786 | 0.001243 | 0.000154 | 100.8872 | 101.053 | 120.1496 | 21.2612 | 0.00295 | 0.002659 | 222.0866 | 0.26463 | 8.4017 | 0.003476 | 0.004848 | 0.004732 | 4.4e-05 | 0.001014 | 0.32549 | 1.3545 | 2.5481 | 101.19 | 20.81 | 17.47 | 92.792 | 105.5173 | 85.7963 | 0.004249 | 0.98356 | 0.024068 | 24.4703 | -0.014688 | 0.15743 | 0.008191 | 0.58524 | 0.58524 | 0.011191
2019-03-19 15:57 | 1553025431.7745 | 902 | 0.20799 | 0.001228 | 0.000157 | 100.8017 | 100.7993 | 120.1919 | 21.2612 | 0.002914 | 0.002683 | 222.0665 | 0.26444 | 8.4027 | 0.00347 | 0.004684 | 0.004563 | 4.9e-05 | 0.001012 | 0.30735 | 1.3508 | 2.5481 | 101.19 | 20.81 | 17.473 | 91.9223 | 105.3551 | 86.3576 | 0.004249 | 0.98203 | 0.024053 | 24.4937 | 0.000217 | 0.15731 | 0.007465 | 0.53213 | 0.53213 | 0.012298
2019-03-19 15:57 | 1553025432.7739 | 903 | 0.20792 | 0.001215 | 0.000158 | 100.7759 | 100.7256 | 120.3585 | 21.2612 | 0.002945 | 0.002669 | 222.0488 | 0.26483 | 8.4059 | 0.003465 | 0.005021 | 0.004934 | 4.5e-05 | 0.001012 | 0.34036 | 1.3288 | 2.548 | 101.19 | 20.81 | 17.474 | 92.6707 | 105.2349 | 86.0258 | 0.00425 | 0.98293 | 0.024081 | 24.4994 | 0.004463 | 0.15755 | 0.007898 | 0.56406 | 0.56406 | 0.01162
2019-03-19 15:57 | 1553025433.7746 | 904 | 0.20791 | 0.0012 | 0.00016 | 100.6997 | 100.7445 | 120.3764 | 21.2612 | 0.002864 | 0.002638 | 222.0283 | 0.26621 | 8.4053 | 0.003441 | 0.00487 | 0.004807 | 5e-05 | 0.001011 | 0.31758 | 1.336 | 2.5481 | 101.19 | 20.81 | 17.474 | 90.7262 | 104.6511 | 85.289 | 0.00425 | 0.98495 | 0.024108 | 24.476 | -0.003977 | 0.15837 | 0.007976 | 0.57023 | 0.57023 | 0.011553
2019-03-19 15:57 | 1553025434.7747 | 905 | 0.20804 | 0.00118 | 0.000163 | 100.7681 | 100.8215 | 120.5181 | 21.2612 | 0.002899 | 0.002651 | 221.9989 | 0.26515 | 8.3998 | 0.003448 | 0.004998 | 0.004895 | 4.9e-05 | 0.001012 | 0.32836 | 1.3679 | 2.5481 | 101.19 | 20.81 | 17.474 | 91.5565 | 104.8407 | 85.5942 | 0.00425 | 0.98411 | 0.024112 | 24.5009 | -0.004741 | 0.15774 | 0.00717 | 0.51162 | 0.51162 | 0.012826
2019-03-19 15:57 | 1553025435.7746 | 906 | 0.20808 | 0.001169 | 0.000164 | 100.6676 | 100.7294 | 120.5099 | 21.2612 | 0.002894 | 0.002653 | 221.9888 | 0.26489 | 8.4047 | 0.003469 | 0.004062 | 0.00396 | 4.6e-05 | 0.001012 | 0.32189 | 1.3163 | 2.5479 | 101.19 | 20.81 | 17.477 | 91.4321 | 105.3411 | 85.6505 | 0.00425 | 0.98396 | 0.024109 | 24.502 | -0.005492 | 0.15758 | 0.006988 | 0.49822 | 0.49822 | 0.013158
2019-03-19 15:57 | 1553025436.7738 | 907 | 0.20785 | 0.001155 | 0.000168 | 100.6045 | 100.6568 | 120.635 | 21.2612 | 0.00288 | 0.002687 | 221.9856 | 0.26611 | 8.4049 | 0.003448 | 0.004991 | 0.004844 | 5e-05 | 0.001004 | 0.32388 | 1.31 | 2.5481 | 101.19 | 20.81 | 17.479 | 91.1068 | 104.8344 | 86.4661 | 0.004251 | 0.98173 | 0.024094 | 24.5425 | -0.004643 | 0.15831 | 0.008385 | 0.59913 | 0.59913 | 0.010992
2019-03-19 15:57 | 1553025437.7747 | 908 | 0.20818 | 0.001136 | 0.00017 | 100.6346 | 100.6626 | 120.6808 | 21.2612 | 0.00287 | 0.002643 | 221.9532 | 0.26492 | 8.4018 | 0.00336 | 0.004931 | 0.004893 | 4.8e-05 | 0.001 | 0.3126 | 1.3046 | 2.5481 | 101.19 | 20.81 | 17.483 | 90.87 | 102.6932 | 85.3929 | 0.004252 | 0.98467 | 0.024135 | 24.5105 | -0.002493 | 0.1576 | 0.006381 | 0.45423 | 0.45423 | 0.014434
2019-03-19 15:57 | 1553025438.7748 | 909 | 0.20806 | 0.001122 | 0.000173 | 100.5778 | 100.604 | 120.9458 | 21.2612 | 0.002904 | 0.002696 | 221.9285 | 0.26576 | 8.4019 | 0.003433 | 0.005297 | 0.005099 | 4.8e-05 | 0.001002 | 0.31797 | 1.3236 | 2.5482 | 101.19 | 20.81 | 17.484 | 91.6784 | 104.4742 | 86.6817 | 0.004252 | 0.98114 | 0.024118 | 24.5815 | -0.00233 | 0.1581 | 0.007114 | 0.50706 | 0.50706 | 0.012971
2019-03-19 15:57 | 1553025439.774 | 910 | 0.2081 | 0.001105 | 0.000175 | 100.6036 | 100.5582 | 120.8118 | 21.2612 | 0.002913 | 0.002694 | 221.8969 | 0.26549 | 8.4015 | 0.003384 | 0.005025 | 0.00502 | 4.6e-05 | 0.000999 | 0.29645 | 1.3117 | 2.5481 | 101.19 | 20.8 | 17.488 | 91.91 | 103.2732 | 86.6261 | 0.00425 | 0.98129 | 0.024106 | 24.566 | 0.004046 | 0.15794 | 0.006899 | 0.49103 | 0.49103 | 0.013381
2019-03-19 15:57 | 1553025440.774 | 911 | 0.20788 | 0.001086 | 0.000178 | 100.5277 | 100.4773 | 121.089 | 21.2612 | 0.002897 | 0.002638 | 221.8659 | 0.26569 | 8.4054 | 0.003405 | 0.004765 | 0.004512 | 4.7e-05 | 0.001 | 0.32516 | 1.3322 | 2.548 | 101.19 | 20.8 | 17.49 | 91.5198 | 103.7862 | 85.2899 | 0.004251 | 0.98495 | 0.024179 | 24.5484 | 0.004485 | 0.15806 | 0.008263 | 0.59174 | 0.59174 | 0.011112
2019-03-19 15:57 | 1553025441.773 | 912 | 0.20832 | 0.001079 | 0.000181 | 100.5853 | 100.467 | 121.4927 | 21.2612 | 0.002957 | 0.002658 | 221.8565 | 0.2658 | 8.4049 | 0.003405 | 0.005051 | 0.00495 | 4.6e-05 | 0.000992 | 0.30359 | 1.3467 | 2.5481 | 101.19 | 20.8 | 17.493 | 92.9627 | 103.7935 | 85.7653 | 0.004252 | 0.98365 | 0.024203 | 24.6056 | 0.010525 | 0.15812 | 0.005585 | 0.3968 | 0.3968 | 0.016577
2019-03-19 15:57 | 1553025442.7745 | 913 | 0.20838 | 0.001058 | 0.000184 | 100.4704 | 100.3164 | 121.7507 | 21.2612 | 0.00286 | 0.00265 | 221.8449 | 0.26602 | 8.4041 | 0.003406 | 0.005449 | 0.005396 | 4.9e-05 | 0.000992 | 0.32444 | 1.3716 | 2.5482 | 101.19 | 20.8 | 17.499 | 90.6356 | 103.8076 | 85.5699 | 0.004253 | 0.98418 | 0.024235 | 24.625 | 0.013725 | 0.15826 | 0.005246 | 0.3723 | 0.3723 | 0.017683
2019-03-19 15:57 | 1553025443.7742 | 914 | 0.20827 | 0.001047 | 0.000186 | 100.3816 | 100.3681 | 122.0056 | 21.2612 | 0.00282 | 0.002653 | 221.8387 | 0.26547 | 8.4041 | 0.0034 | 0.004955 | 0.004836 | 4.8e-05 | 0.000993 | 0.31797 | 1.4057 | 2.5481 | 101.19 | 20.8 | 17.504 | 89.6596 | 103.67 | 85.6394 | 0.004254 | 0.98399 | 0.024258 | 24.6532 | 0.001203 | 0.15793 | 0.005914 | 0.4214 | 0.4214 | 0.01559
2019-03-19 15:57 | 1553025444.7732 | 915 | 0.20821 | 0.001038 | 0.000187 | 100.4206 | 100.3016 | 122.0521 | 21.2612 | 0.002908 | 0.002638 | 221.8453 | 0.26534 | 8.4035 | 0.003447 | 0.005093 | 0.005018 | 4.6e-05 | 0.000995 | 0.3079 | 1.3436 | 2.5481 | 101.19 | 20.8 | 17.517 | 91.77 | 104.8094 | 85.2743 | 0.004257 | 0.98499 | 0.024275 | 24.6453 | 0.010602 | 0.15785 | 0.006329 | 0.45206 | 0.45206 | 0.014526
2019-03-19 15:57 | 1553025445.7742 | 916 | 0.20813 | 0.001021 | 0.000188 | 100.3489 | 100.2992 | 122.1072 | 21.2612 | 0.002854 | 0.00264 | 221.8256 | 0.26569 | 8.4032 | 0.003446 | 0.00522 | 0.005128 | 4.5e-05 | 0.001 | 0.30381 | 1.3368 | 2.5482 | 101.19 | 20.8 | 17.525 | 90.4855 | 104.7895 | 85.3289 | 0.004259 | 0.98484 | 0.024279 | 24.6527 | 0.004425 | 0.15806 | 0.006793 | 0.48597 | 0.48597 | 0.01353
2019-03-19 15:57 | 1553025446.7737 | 917 | 0.20828 | 0.001008 | 0.000189 | 100.3714 | 100.292 | 122.1762 | 21.2612 | 0.002871 | 0.002598 | 221.7875 | 0.26493 | 8.4032 | 0.0035 | 0.005016 | 0.004947 | 4.7e-05 | 0.001005 | 0.2907 | 1.3655 | 2.5483 | 101.19 | 20.8 | 17.534 | 90.8895 | 106.1026 | 84.3245 | 0.004262 | 0.98761 | 0.02432 | 24.6251 | 0.007072 | 0.15761 | 0.005923 | 0.42301 | 0.42301 | 0.0155
2019-03-19 15:57 | 1553025447.7734 | 918 | 0.20821 | 0.000995 | 0.000191 | 100.3567 | 100.3167 | 122.3218 | 21.2612 | 0.002797 | 0.002617 | 221.7933 | 0.26514 | 8.4027 | 0.003513 | 0.004598 | 0.004473 | 4.5e-05 | 0.001009 | 0.32892 | 1.3613 | 2.5481 | 101.19 | 20.8 | 17.553 | 89.0978 | 106.4104 | 84.7632 | 0.004266 | 0.9864 | 0.02432 | 24.6549 | 0.003563 | 0.15773 | 0.00638 | 0.45637 | 0.45637 | 0.014378
2019-03-19 15:57 | 1553025448.7747 | 919 | 0.20819 | 0.000977 | 0.000192 | 100.2867 | 100.264 | 122.3025 | 21.2612 | 0.00286 | 0.002621 | 221.7797 | 0.26554 | 8.403 | 0.003461 | 0.005744 | 0.005631 | 4.6e-05 | 0.001008 | 0.30287 | 1.3869 | 2.5483 | 101.19 | 20.8 | 17.563 | 90.6199 | 105.1522 | 84.877 | 0.004269 | 0.98609 | 0.024314 | 24.6569 | 0.002033 | 0.15797 | 0.006528 | 0.46699 | 0.46699 | 0.014072
2019-03-19 15:57 | 1553025449.7745 | 920 | 0.20825 | 0.00097 | 0.000195 | 100.3417 | 100.2294 | 122.5015 | 21.2612 | 0.002844 | 0.002608 | 221.7682 | 0.26521 | 8.4026 | 0.003505 | 0.004788 | 0.004708 | 4.4e-05 | 0.001008 | 0.29839 | 1.3341 | 2.548 | 101.19 | 20.79 | 17.581 | 90.2487 | 106.2224 | 84.5689 | 0.00427 | 0.98694 | 0.024344 | 24.6663 | 0.010016 | 0.15777 | 0.006139 | 0.43885 | 0.43885 | 0.014956
2019-03-19 15:57 | 1553025450.7739 | 921 | 0.20819 | 0.000955 | 0.000198 | 100.2846 | 100.1609 | 122.3882 | 21.2612 | 0.002848 | 0.002583 | 221.7358 | 0.26547 | 8.4044 | 0.003556 | 0.005057 | 0.004968 | 4.4e-05 | 0.001005 | 0.30309 | 1.3249 | 2.548 | 101.19 | 20.79 | 17.59 | 90.3347 | 107.4544 | 83.9532 | 0.004272 | 0.98864 | 0.024354 | 24.6337 | 0.011046 | 0.15793 | 0.006543 | 0.46857 | 0.46857 | 0.014021
2019-03-19 15:58 | 1553025451.7745 | 922 | 0.20828 | 0.000938 | 0.0002 | 100.2696 | 100.1992 | 122.3194 | 21.2612 | 0.002812 | 0.002624 | 221.7211 | 0.26534 | 8.4021 | 0.0035 | 0.005243 | 0.005179 | 4.5e-05 | 0.001008 | 0.31238 | 1.3641 | 2.548 | 101.19 | 20.79 | 17.599 | 89.4727 | 106.0822 | 84.9364 | 0.004275 | 0.98592 | 0.024313 | 24.6606 | 0.006283 | 0.15785 | 0.006024 | 0.42958 | 0.42958 | 0.015286
2019-03-19 15:58 | 1553025452.774 | 923 | 0.20831 | 0.000926 | 0.000202 | 100.2501 | 100.1047 | 122.2074 | 21.2612 | 0.002879 | 0.00261 | 221.7078 | 0.26549 | 8.4028 | 0.003469 | 0.00476 | 0.00463 | 4.3e-05 | 0.001009 | 0.34916 | 1.3209 | 2.548 | 101.19 | 20.79 | 17.615 | 91.0904 | 105.3481 | 84.6179 | 0.004279 | 0.9868 | 0.024313 | 24.6384 | 0.012985 | 0.15794 | 0.005849 | 0.41665 | 0.41665 | 0.01577
2019-03-19 15:58 | 1553025453.774 | 924 | 0.20833 | 0.000912 | 0.000204 | 100.1679 | 99.9283 | 122.0183 | 21.2612 | 0.002816 | 0.002598 | 221.6801 | 0.26666 | 8.406 | 0.003438 | 0.004604 | 0.004514 | 4.2e-05 | 0.001005 | 0.29574 | 1.2916 | 2.548 | 101.19 | 20.79 | 17.621 | 89.5751 | 104.5876 | 84.3238 | 0.00428 | 0.98761 | 0.024304 | 24.6092 | 0.021419 | 0.15863 | 0.005762 | 0.40998 | 0.40998 | 0.016096
2019-03-19 15:58 | 1553025454.7737 | 925 | 0.20828 | 0.000907 | 0.000206 | 100.1625 | 100.1718 | 122.0476 | 21.2612 | 0.002837 | 0.002584 | 221.6495 | 0.26488 | 8.4022 | 0.003416 | 0.005568 | 0.005465 | 4.3e-05 | 0.000997 | 0.31133 | 1.3221 | 2.5481 | 101.19 | 20.79 | 17.627 | 90.076 | 104.0506 | 83.979 | 0.004281 | 0.98857 | 0.024319 | 24.6003 | -0.000833 | 0.15758 | 0.006021 | 0.42912 | 0.42912 | 0.015276
2019-03-19 15:58 | 1553025455.7727 | 926 | 0.20833 | 0.000889 | 0.000207 | 100.2099 | 100.0088 | 121.7903 | 21.2612 | 0.002837 | 0.002658 | 221.6398 | 0.26538 | 8.4027 | 0.00342 | 0.004905 | 0.004759 | 4.1e-05 | 0.000993 | 0.31116 | 1.3304 | 2.548 | 101.19 | 20.79 | 17.632 | 90.0736 | 104.1526 | 85.758 | 0.004283 | 0.98367 | 0.024233 | 24.6355 | 0.017969 | 0.15788 | 0.005724 | 0.40581 | 0.40581 | 0.016184
2019-03-19 15:58 | 1553025456.7743 | 927 | 0.20832 | 0.000875 | 0.00021 | 100.1502 | 99.9684 | 121.6469 | 21.2612 | 0.002815 | 0.002559 | 221.6417 | 0.26623 | 8.4034 | 0.003423 | 0.004841 | 0.004698 | 4.5e-05 | 0.000988 | 0.30685 | 1.3511 | 2.5481 | 101.19 | 20.79 | 17.631 | 89.5399 | 104.2258 | 83.3814 | 0.004282 | 0.99022 | 0.024299 | 24.5393 | 0.016252 | 0.15838 | 0.005844 | 0.41557 | 0.41557 | 0.015854
2019-03-19 15:58 | 1553025457.7739 | 928 | 0.20836 | 0.000868 | 0.000213 | 100.0971 | 100.0324 | 121.1937 | 21.2612 | 0.00288 | 0.002521 | 221.6009 | 0.26575 | 8.4024 | 0.003408 | 0.005012 | 0.004921 | 4.5e-05 | 0.000987 | 0.30807 | 1.3766 | 2.5482 | 101.19 | 20.79 | 17.624 | 91.0998 | 103.8524 | 82.4758 | 0.004281 | 0.99274 | 0.024285 | 24.4624 | 0.005778 | 0.1581 | 0.005606 | 0.39768 | 0.39768 | 0.016538
2019-03-19 15:58 | 1553025458.773 | 929 | 0.20838 | 0.000861 | 0.000217 | 100.1703 | 99.8667 | 121.0761 | 21.2612 | 0.002782 | 0.002629 | 221.5917 | 0.2652 | 8.4044 | 0.003415 | 0.005042 | 0.004808 | 4.3e-05 | 0.000988 | 0.29214 | 1.3412 | 2.5481 | 101.19 | 20.79 | 17.62 | 88.7509 | 104.0363 | 85.0674 | 0.00428 | 0.98556 | 0.024185 | 24.5395 | 0.027154 | 0.15777 | 0.005453 | 0.3846 | 0.3846 | 0.017065
2019-03-19 15:58 | 1553025459.7739 | 930 | 0.20838 | 0.000845 | 0.000222 | 100.0731 | 99.8009 | 120.7928 | 21.2612 | 0.002774 | 0.00264 | 221.6036 | 0.26684 | 8.4018 | 0.003432 | 0.005172 | 0.005123 | 4.4e-05 | 0.000983 | 0.32372 | 1.3593 | 2.5481 | 101.19 | 20.79 | 17.614 | 88.5525 | 104.4507 | 85.3226 | 0.004278 | 0.98486 | 0.024148 | 24.5195 | 0.024359 | 0.15874 | 0.005498 | 0.38698 | 0.38698 | 0.017065
2019-03-19 15:58 | 1553025460.7742 | 931 | 0.20841 | 0.000832 | 0.000228 | 100.116 | 100.0574 | 120.6491 | 21.2612 | 0.002822 | 0.002577 | 221.6215 | 0.26564 | 8.4034 | 0.003422 | 0.004869 | 0.004772 | 4.1e-05 | 0.000968 | 0.31332 | 1.3241 | 2.5481 | 101.19 | 20.78 | 17.601 | 89.715 | 104.196 | 83.8229 | 0.004273 | 0.989 | 0.024185 | 24.4536 | 0.005232 | 0.15803 | 0.005346 | 0.37619 | 0.37619 | 0.017475
2019-03-19 15:58 | 1553025461.7745 | 932 | 0.20847 | 0.00082 | 0.000234 | 100.0623 | 99.8382 | 120.444 | 21.2612 | 0.002794 | 0.002558 | 221.5995 | 0.26522 | 8.4027 | 0.003412 | 0.004456 | 0.004355 | 4.4e-05 | 0.000963 | 0.31924 | 1.349 | 2.5481 | 101.19 | 20.78 | 17.595 | 89.0446 | 103.9624 | 83.3564 | 0.004271 | 0.99029 | 0.02418 | 24.4168 | 0.020048 | 0.15778 | 0.004987 | 0.3496 | 0.3496 | 0.018775
2019-03-19 15:58 | 1553025462.7732 | 933 | 0.20847 | 0.000803 | 0.000238 | 100.0154 | 99.8697 | 120.4015 | 21.2612 | 0.002784 | 0.002603 | 221.5831 | 0.26638 | 8.4006 | 0.003398 | 0.005139 | 0.005052 | 4.8e-05 | 0.000959 | 0.2844 | 1.3349 | 2.5481 | 101.19 | 20.78 | 17.581 | 88.7947 | 103.6079 | 84.4479 | 0.004268 | 0.98727 | 0.024139 | 24.4498 | 0.01304 | 0.15847 | 0.004951 | 0.3461 | 0.3461 | 0.019047
2019-03-19 15:58 | 1553025463.7732 | 934 | 0.20846 | 0.000793 | 0.000241 | 100.0971 | 99.9129 | 120.4557 | 21.2612 | 0.002757 | 0.002564 | 221.5918 | 0.26579 | 8.4 | 0.003442 | 0.004778 | 0.004716 | 4.7e-05 | 0.000958 | 0.30309 | 1.3528 | 2.5482 | 101.19 | 20.78 | 17.573 | 88.1464 | 104.6858 | 83.492 | 0.004266 | 0.98991 | 0.024176 | 24.4226 | 0.016473 | 0.15812 | 0.005076 | 0.35555 | 0.35555 | 0.0185
2019-03-19 15:58 | 1553025464.7743 | 935 | 0.2085 | 0.000773 | 0.000245 | 100.0707 | 99.9372 | 120.4745 | 21.2612 | 0.00283 | 0.002603 | 221.5843 | 0.2653 | 8.4023 | 0.003473 | 0.004995 | 0.004846 | 4.8e-05 | 0.000961 | 0.30901 | 1.3173 | 2.548 | 101.19 | 20.78 | 17.569 | 89.8988 | 105.4434 | 84.4486 | 0.004265 | 0.98727 | 0.024146 | 24.4573 | 0.011936 | 0.15782 | 0.004804 | 0.3351 | 0.3351 | 0.019593
2019-03-19 15:58 | 1553025465.7733 | 936 | 0.20845 | 0.000756 | 0.000249 | 100.0884 | 99.9413 | 120.4965 | 21.2612 | 0.00279 | 0.002623 | 221.5656 | 0.26573 | 8.4011 | 0.003498 | 0.004213 | 0.004165 | 4.5e-05 | 0.000957 | 0.33284 | 1.3073 | 2.548 | 101.19 | 20.78 | 17.562 | 88.9391 | 106.0397 | 84.927 | 0.004263 | 0.98595 | 0.024132 | 24.4759 | 0.01315 | 0.15808 | 0.005148 | 0.35965 | 0.35965 | 0.018285
2019-03-19 15:58 | 1553025466.7744 | 937 | 0.20848 | 0.000743 | 0.000254 | 100.0488 | 99.8533 | 120.9155 | 21.2612 | 0.002772 | 0.002587 | 221.5687 | 0.26565 | 8.4023 | 0.003494 | 0.004859 | 0.004754 | 4.7e-05 | 0.000952 | 0.31725 | 1.303 | 2.548 | 101.19 | 20.78 | 17.56 | 88.5135 | 105.9517 | 84.0585 | 0.004263 | 0.98835 | 0.024203 | 24.4886 | 0.017491 | 0.15804 | 0.004992 | 0.34904 | 0.34904 | 0.018836
2019-03-19 15:58 | 1553025467.7747 | 938 | 0.20855 | 0.000731 | 0.00026 | 100.0554 | 99.9992 | 121.2441 | 21.2612 | 0.002737 | 0.00256 | 221.5726 | 0.26563 | 8.4013 | 0.003477 | 0.004953 | 0.004788 | 4.7e-05 | 0.000944 | 0.32764 | 1.3203 | 2.548 | 101.19 | 20.78 | 17.56 | 87.6645 | 105.5377 | 83.3961 | 0.004263 | 0.99018 | 0.024259 | 24.4992 | 0.005021 | 0.15802 | 0.00453 | 0.31563 | 0.31563 | 0.020828
2019-03-19 15:58 | 1553025468.7735 | 939 | 0.20856 | 0.000711 | 0.000267 | 100.0848 | 99.8352 | 121.3815 | 21.2612 | 0.002789 | 0.002555 | 221.569 | 0.26628 | 8.4054 | 0.003474 | 0.004319 | 0.004209 | 4.9e-05 | 0.00095 | 0.3183 | 1.3552 | 2.5481 | 101.19 | 20.78 | 17.561 | 88.911 | 105.4676 | 83.2814 | 0.004263 | 0.9905 | 0.024276 | 24.5091 | 0.022336 | 0.15841 | 0.004516 | 0.31447 | 0.31447 | 0.020955
2019-03-19 15:58 | 1553025469.7735 | 940 | 0.20856 | 0.0007 | 0.000273 | 100.0539 | 100.0588 | 121.5831 | 21.2612 | 0.002717 | 0.002586 | 221.5466 | 0.26673 | 8.4021 | 0.003505 | 0.004661 | 0.004591 | 4.8e-05 | 0.000955 | 0.33964 | 1.3568 | 2.548 | 101.19 | 20.78 | 17.563 | 87.1781 | 106.2151 | 84.0234 | 0.004263 | 0.98844 | 0.024271 | 24.5549 | -0.000442 | 0.15868 | 0.004545 | 0.31611 | 0.31611 | 0.020882
2019-03-19 15:58 | 1553025470.7726 | 941 | 0.20857 | 0.000683 | 0.000279 | 100.0262 | 100.0557 | 121.4601 | 21.2612 | 0.002774 | 0.002621 | 221.531 | 0.26565 | 8.4036 | 0.003481 | 0.004846 | 0.004778 | 4.7e-05 | 0.000953 | 0.29983 | 1.3742 | 2.548 | 101.19 | 20.78 | 17.567 | 88.5636 | 105.6274 | 84.8637 | 0.004264 | 0.98612 | 0.02423 | 24.5714 | -0.002635 | 0.15803 | 0.004482 | 0.31065 | 0.31065 | 0.021163
2019-03-19 15:58 | 1553025471.774 | 942 | 0.20857 | 0.000674 | 0.000283 | 100.0989 | 99.9184 | 121.489 | 21.2612 | 0.002778 | 0.002517 | 221.5308 | 0.26553 | 8.4012 | 0.003425 | 0.00481 | 0.00473 | 4.8e-05 | 0.000945 | 0.29911 | 1.3386 | 2.5481 | 101.19 | 20.78 | 17.569 | 88.6572 | 104.2769 | 82.3681 | 0.004265 | 0.99304 | 0.024318 | 24.4885 | 0.016143 | 0.15796 | 0.004513 | 0.31375 | 0.31375 | 0.020945
2019-03-19 15:58 | 1553025472.7727 | 943 | 0.20855 | 0.000659 | 0.000287 | 100.1274 | 100.0656 | 121.5845 | 21.2612 | 0.002773 | 0.002532 | 221.5271 | 0.26541 | 8.4016 | 0.003369 | 0.005343 | 0.005203 | 5.1e-05 | 0.000941 | 0.29264 | 1.3112 | 2.548 | 101.19 | 20.78 | 17.577 | 88.5323 | 102.9259 | 82.7385 | 0.004267 | 0.99201 | 0.024315 | 24.5109 | 0.005516 | 0.15789 | 0.004651 | 0.32356 | 0.32356 | 0.0203
2019-03-19 15:58 | 1553025473.7727 | 944 | 0.20858 | 0.000648 | 0.000291 | 100.1358 | 100.1222 | 121.68 | 21.2612 | 0.002747 | 0.002579 | 221.5063 | 0.26565 | 8.4007 | 0.003369 | 0.005663 | 0.005457 | 5.1e-05 | 0.000936 | 0.32173 | 1.3242 | 2.5482 | 101.19 | 20.78 | 17.584 | 87.9146 | 102.9172 | 83.8525 | 0.004268 | 0.98892 | 0.024287 | 24.5589 | 0.001217 | 0.15804 | 0.00446 | 0.30898 | 0.30898 | 0.021278
2019-03-19 15:58 | 1553025474.7737 | 945 | 0.20857 | 0.000637 | 0.000293 | 100.1565 | 100.0574 | 121.5831 | 21.2612 | 0.002813 | 0.002549 | 221.4967 | 0.26515 | 8.4055 | 0.003413 | 0.005098 | 0.005068 | 4.5e-05 | 0.000941 | 0.28926 | 1.3002 | 2.5479 | 101.19 | 20.78 | 17.591 | 89.4978 | 103.9932 | 83.1451 | 0.00427 | 0.99088 | 0.024301 | 24.5247 | 0.008851 | 0.15774 | 0.004529 | 0.31409 | 0.31409 | 0.020892
2019-03-19 15:58 | 1553025475.7732 | 946 | 0.20854 | 0.000627 | 0.000295 | 100.1757 | 100.067 | 121.7009 | 21.2612 | 0.002777 | 0.002509 | 221.4842 | 0.26559 | 8.4053 | 0.003411 | 0.005014 | 0.00489 | 5e-05 | 0.00095 | 0.32836 | 1.3191 | 2.5479 | 101.19 | 20.78 | 17.608 | 88.6228 | 103.9445 | 82.1693 | 0.004274 | 0.9936 | 0.024346 | 24.503 | 0.009709 | 0.158 | 0.004685 | 0.32599 | 0.32599 | 0.020162
2019-03-19 15:58 | 1553025476.7739 | 947 | 0.20858 | 0.00062 | 0.000296 | 100.1913 | 100.1537 | 121.2465 | 21.2612 | 0.002839 | 0.002495 | 221.446 | 0.26482 | 8.4008 | 0.003378 | 0.004915 | 0.004808 | 5e-05 | 0.000949 | 0.31188 | 1.3009 | 2.548 | 101.19 | 20.78 | 17.617 | 90.1143 | 103.1388 | 81.8479 | 0.004276 | 0.9945 | 0.024312 | 24.4461 | 0.003361 | 0.15754 | 0.004464 | 0.30929 | 0.30929 | 0.02119
2019-03-19 15:58 | 1553025477.7743 | 948 | 0.20858 | 0.00061 | 0.000296 | 100.1703 | 100.1317 | 121.3157 | 21.2612 | 0.002752 | 0.002568 | 221.4352 | 0.26524 | 8.4009 | 0.003428 | 0.005345 | 0.005228 | 4.8e-05 | 0.000941 | 0.31078 | 1.3483 | 2.5481 | 101.19 | 20.77 | 17.626 | 88.0231 | 104.3351 | 83.599 | 0.004276 | 0.98962 | 0.024259 | 24.5134 | 0.003443 | 0.15779 | 0.004481 | 0.30988 | 0.30988 | 0.021183
2019-03-19 15:58 | 1553025478.774 | 949 | 0.20858 | 0.000607 | 0.000296 | 100.2291 | 100.166 | 121.3404 | 21.2612 | 0.002753 | 0.002519 | 221.4162 | 0.26542 | 8.4044 | 0.003443 | 0.004685 | 0.004676 | 4.8e-05 | 0.000939 | 0.3126 | 1.3515 | 2.5482 | 101.19 | 20.77 | 17.647 | 88.0402 | 104.7071 | 82.408 | 0.004281 | 0.99293 | 0.024302 | 24.4749 | 0.005635 | 0.1579 | 0.004518 | 0.31309 | 0.31309 | 0.02098
2019-03-19 15:58 | 1553025479.7731 | 950 | 0.2086 | 0.000596 | 0.000295 | 100.2696 | 100.342 | 121.1721 | 21.2612 | 0.002722 | 0.002542 | 221.3892 | 0.26538 | 8.4019 | 0.0034 | 0.004491 | 0.004353 | 4.7e-05 | 0.000948 | 0.31902 | 1.3228 | 2.548 | 101.19 | 20.77 | 17.66 | 87.3093 | 103.678 | 82.9804 | 0.004284 | 0.99134 | 0.024266 | 24.4776 | -0.006457 | 0.15788 | 0.004403 | 0.3043 | 0.3043 | 0.021583
2019-03-19 15:58 | 1553025480.7734 | 951 | 0.20857 | 0.000596 | 0.000295 | 100.3221 | 100.1842 | 121.3566 | 21.2612 | 0.002757 | 0.002567 | 221.383 | 0.26487 | 8.405 | 0.003456 | 0.004629 | 0.004533 | 4.9e-05 | 0.00095 | 0.30558 | 1.3618 | 2.5481 | 101.19 | 20.77 | 17.695 | 88.1432 | 105.0239 | 83.5623 | 0.004293 | 0.98972 | 0.024264 | 24.5162 | 0.012312 | 0.15757 | 0.004546 | 0.31473 | 0.31473 | 0.020827
2019-03-19 15:58 | 1553025481.7736 | 952 | 0.20859 | 0.00059 | 0.000293 | 100.3071 | 100.3985 | 121.4953 | 21.2612 | 0.002694 | 0.002535 | 221.364 | 0.26556 | 8.4022 | 0.003473 | 0.004446 | 0.004333 | 5.3e-05 | 0.000956 | 0.31509 | 1.3683 | 2.5479 | 101.19 | 20.77 | 17.716 | 86.6292 | 105.4371 | 82.7924 | 0.004298 | 0.99186 | 0.024304 | 24.5037 | -0.008147 | 0.15798 | 0.004476 | 0.31026 | 0.31026 | 0.021183
2019-03-19 15:58 | 1553025482.7732 | 953 | 0.20863 | 0.000583 | 0.000291 | 100.3306 | 100.342 | 121.5756 | 21.2612 | 0.002699 | 0.002557 | 221.3438 | 0.2657 | 8.4061 | 0.003447 | 0.004912 | 0.004803 | 4.8e-05 | 0.000958 | 0.29407 | 1.3369 | 2.5481 | 101.19 | 20.77 | 17.739 | 86.7394 | 104.8116 | 83.3206 | 0.004303 | 0.99039 | 0.024294 | 24.53 | -0.001024 | 0.15807 | 0.004206 | 0.29043 | 0.29043 | 0.02264
2019-03-19 15:58 | 1553025483.7743 | 954 | 0.2086 | 0.000577 | 0.000289 | 100.3858 | 100.5472 | 121.6317 | 21.2612 | 0.0027 | 0.002552 | 221.3241 | 0.26545 | 8.4022 | 0.003473 | 0.004671 | 0.004564 | 4.6e-05 | 0.000963 | 0.31542 | 1.3059 | 2.548 | 101.19 | 20.76 | 17.787 | 86.7808 | 105.4465 | 83.209 | 0.004312 | 0.9907 | 0.024304 | 24.5318 | -0.014371 | 0.15792 | 0.004411 | 0.3057 | 0.3057 | 0.02149
2019-03-19 15:58 | 1553025484.7747 | 955 | 0.20865 | 0.00058 | 0.000287 | 100.4812 | 100.4232 | 121.6922 | 21.2612 | 0.002755 | 0.002549 | 221.3119 | 0.26528 | 8.3993 | 0.003464 | 0.005317 | 0.005141 | 4.7e-05 | 0.000964 | 0.30342 | 1.338 | 2.5481 | 101.19 | 20.76 | 17.811 | 88.0987 | 105.2089 | 83.1347 | 0.004318 | 0.99091 | 0.024312 | 24.5354 | 0.005164 | 0.15781 | 0.004126 | 0.28503 | 0.28503 | 0.023033
2019-03-19 15:58 | 1553025485.7744 | 956 | 0.20864 | 0.00057 | 0.000285 | 100.4548 | 100.4982 | 121.7265 | 21.2612 | 0.002671 | 0.002484 | 221.3123 | 0.26478 | 8.4035 | 0.003458 | 0.005197 | 0.005055 | 4.9e-05 | 0.000971 | 0.3126 | 1.3672 | 2.5481 | 101.19 | 20.76 | 17.835 | 86.071 | 105.0768 | 81.571 | 0.004324 | 0.99528 | 0.024369 | 24.4849 | -0.003865 | 0.15752 | 0.004181 | 0.28985 | 0.28985 | 0.022607
2019-03-19 15:58 | 1553025486.7724 | 957 | 0.20866 | 0.000568 | 0.000284 | 100.4866 | 100.4797 | 122.2189 | 21.2612 | 0.002756 | 0.002534 | 221.3126 | 0.26461 | 8.4039 | 0.003495 | 0.00458 | 0.004537 | 5.1e-05 | 0.000977 | 0.30342 | 1.3223 | 2.548 | 101.19 | 20.76 | 17.884 | 88.1307 | 105.9641 | 82.7712 | 0.004336 | 0.99192 | 0.024377 | 24.5759 | 0.000613 | 0.15741 | 0.004109 | 0.2848 | 0.2848 | 0.022993
2019-03-19 15:58 | 1553025487.7738 | 958 | 0.20862 | 0.000562 | 0.000282 | 100.5037 | 100.5441 | 122.2621 | 21.2612 | 0.002703 | 0.002509 | 221.298 | 0.26571 | 8.4044 | 0.003472 | 0.005167 | 0.005027 | 4.9e-05 | 0.000976 | 0.28766 | 1.2889 | 2.5481 | 101.19 | 20.76 | 17.908 | 86.833 | 105.4228 | 82.1676 | 0.004342 | 0.9936 | 0.024402 | 24.5594 | -0.003595 | 0.15807 | 0.004324 | 0.30101 | 0.30101 | 0.021846
2019-03-19 15:58 | 1553025488.7728 | 959 | 0.20862 | 0.000561 | 0.000281 | 100.5904 | 100.6047 | 122.0575 | 21.2612 | 0.00266 | 0.002488 | 221.2804 | 0.2653 | 8.4013 | 0.003487 | 0.005683 | 0.005628 | 5e-05 | 0.000978 | 0.32997 | 1.3651 | 2.5481 | 101.19 | 20.76 | 17.958 | 85.7987 | 105.7831 | 81.6623 | 0.004354 | 0.99502 | 0.024399 | 24.5213 | -0.00127 | 0.15783 | 0.004322 | 0.30091 | 0.30091 | 0.02182
2019-03-19 15:58 | 1553025489.7736 | 960 | 0.20862 | 0.00056 | 0.00028 | 100.5871 | 100.4801 | 122.1356 | 21.2612 | 0.002698 | 0.002513 | 221.2613 | 0.26543 | 8.405 | 0.003505 | 0.005186 | 0.005113 | 4.9e-05 | 0.000976 | 0.2855 | 1.3405 | 2.548 | 101.19 | 20.76 | 17.985 | 86.73 | 106.2049 | 82.2722 | 0.00436 | 0.99331 | 0.024386 | 24.5503 | 0.009527 | 0.1579 | 0.00435 | 0.3028 | 0.3028 | 0.021693
2019-03-19 15:58 | 1553025490.7736 | 961 | 0.20865 | 0.000549 | 0.000279 | 100.6093 | 100.627 | 122.26 | 21.2612 | 0.002709 | 0.002535 | 221.2407 | 0.26513 | 8.4031 | 0.003443 | 0.005326 | 0.005208 | 4.6e-05 | 0.000979 | 0.31924 | 1.3121 | 2.5482 | 101.19 | 20.76 | 18.037 | 86.9892 | 104.7023 | 82.8126 | 0.004373 | 0.99181 | 0.02438 | 24.5814 | -0.001568 | 0.15773 | 0.004189 | 0.29096 | 0.29096 | 0.022551
2019-03-19 15:58 | 1553025491.7736 | 962 | 0.20862 | 0.000547 | 0.000278 | 100.691 | 100.7212 | 122.3656 | 21.2612 | 0.002711 | 0.002458 | 221.2113 | 0.2648 | 8.3997 | 0.003542 | 0.005298 | 0.005192 | 4.8e-05 | 0.000981 | 0.323 | 1.3181 | 2.548 | 101.19 | 20.76 | 18.059 | 87.0392 | 107.1039 | 80.9557 | 0.004378 | 0.99701 | 0.024454 | 24.5278 | -0.002682 | 0.15753 | 0.00435 | 0.30373 | 0.30373 | 0.021576
2019-03-19 15:58 | 1553025492.7742 | 963 | 0.20864 | 0.000551 | 0.000277 | 100.7267 | 100.6393 | 122.3957 | 21.2612 | 0.002677 | 0.002527 | 221.2009 | 0.26493 | 8.4012 | 0.003525 | 0.005047 | 0.004911 | 4.7e-05 | 0.000982 | 0.33644 | 1.3575 | 2.5481 | 101.19 | 20.76 | 18.08 | 86.2233 | 106.7088 | 82.6114 | 0.004383 | 0.99237 | 0.0244 | 24.5881 | 0.007761 | 0.15761 | 0.00425 | 0.29578 | 0.29578 | 0.022167
2019-03-19 15:58 | 1553025493.7728 | 964 | 0.20865 | 0.000554 | 0.000274 | 100.8092 | 100.852 | 122.1147 | 21.2612 | 0.002728 | 0.002561 | 221.1916 | 0.26379 | 8.4025 | 0.003503 | 0.005448 | 0.00536 | 4.6e-05 | 0.000982 | 0.31006 | 1.3768 | 2.5482 | 101.19 | 20.77 | 18.114 | 87.4342 | 106.1554 | 83.4204 | 0.004394 | 0.99011 | 0.024345 | 24.5878 | -0.003796 | 0.15693 | 0.004176 | 0.28988 | 0.28988 | 0.02252
2019-03-19 15:58 | 1553025494.7732 | 965 | 0.20863 | 0.000552 | 0.00027 | 100.8338 | 100.7479 | 122.1255 | 21.2612 | 0.002687 | 0.002454 | 221.182 | 0.26538 | 8.4032 | 0.003505 | 0.005095 | 0.004898 | 4.6e-05 | 0.00098 | 0.3027 | 1.3864 | 2.548 | 101.19 | 20.76 | 18.13 | 86.4473 | 106.2238 | 80.8497 | 0.004395 | 0.9973 | 0.024434 | 24.5001 | 0.007626 | 0.15787 | 0.004291 | 0.29963 | 0.29963 | 0.021918
2019-03-19 15:58 | 1553025495.773 | 966 | 0.20865 | 0.000549 | 0.000266 | 100.8377 | 100.8373 | 122.2617 | 21.2612 | 0.002693 | 0.00249 | 221.1664 | 0.26603 | 8.4048 | 0.00354 | 0.005239 | 0.005141 | 4.5e-05 | 0.000987 | 0.27793 | 1.3784 | 2.5481 | 101.19 | 20.77 | 18.146 | 86.6153 | 107.0646 | 81.7287 | 0.004402 | 0.99483 | 0.024417 | 24.5441 | 3.9e-05 | 0.15826 | 0.004208 | 0.2936 | 0.2936 | 0.022424
2019-03-19 15:58 | 1553025496.7744 | 967 | 0.20866 | 0.000546 | 0.000261 | 100.9119 | 100.9712 | 121.8983 | 21.2612 | 0.002676 | 0.002543 | 221.1501 | 0.26596 | 8.4022 | 0.003493 | 0.0049 | 0.004837 | 4.9e-05 | 0.000981 | 0.29806 | 1.3231 | 2.548 | 101.19 | 20.77 | 18.171 | 86.1913 | 105.9223 | 82.9936 | 0.004408 | 0.9913 | 0.024338 | 24.5513 | -0.005259 | 0.15822 | 0.004167 | 0.28992 | 0.28992 | 0.022703
2019-03-19 15:58 | 1553025497.773 | 968 | 0.20868 | 0.000546 | 0.000257 | 100.9533 | 101.0171 | 121.9593 | 21.2612 | 0.002659 | 0.002514 | 221.1593 | 0.26577 | 8.4027 | 0.003545 | 0.00508 | 0.00503 | 4.7e-05 | 0.000987 | 0.30414 | 1.3115 | 2.5478 | 101.19 | 20.77 | 18.18 | 85.7917 | 107.1709 | 82.2986 | 0.00441 | 0.99324 | 0.024368 | 24.5335 | -0.005653 | 0.15811 | 0.004012 | 0.27911 | 0.27911 | 0.023565
2019-03-19 15:58 | 1553025498.7741 | 969 | 0.20868 | 0.00055 | 0.000254 | 101.0448 | 100.9554 | 121.9971 | 21.2612 | 0.002667 | 0.0025 | 221.1507 | 0.26447 | 8.4041 | 0.003557 | 0.005194 | 0.005055 | 4.7e-05 | 0.000993 | 0.31454 | 1.3218 | 2.548 | 101.19 | 20.77 | 18.189 | 85.9729 | 107.4631 | 81.9664 | 0.004413 | 0.99417 | 0.024383 | 24.5258 | 0.007913 | 0.15733 | 0.004045 | 0.28194 | 0.28194 | 0.023215
2019-03-19 15:58 | 1553025499.7731 | 970 | 0.20869 | 0.000542 | 0.000251 | 100.9605 | 101.0746 | 122.0443 | 21.2612 | 0.002681 | 0.002522 | 221.1517 | 0.26504 | 8.404 | 0.003547 | 0.005011 | 0.004929 | 4.6e-05 | 0.000996 | 0.27688 | 1.316 | 2.5481 | 101.19 | 20.77 | 18.189 | 86.3075 | 107.234 | 82.4859 | 0.004413 | 0.99272 | 0.02437 | 24.5485 | -0.010108 | 0.15767 | 0.004005 | 0.27897 | 0.27897 | 0.023512
2019-03-19 15:58 | 1553025500.7735 | 971 | 0.20871 | 0.000538 | 0.000249 | 101.1351 | 101.0393 | 121.9438 | 21.2612 | 0.0027 | 0.002517 | 221.1262 | 0.26541 | 8.3991 | 0.003473 | 0.004895 | 0.004749 | 5e-05 | 0.000998 | 0.3011 | 1.3359 | 2.548 | 101.19 | 20.77 | 18.186 | 86.7791 | 105.4339 | 82.3587 | 0.004412 | 0.99307 | 0.024364 | 24.534 | 0.008475 | 0.15789 | 0.003885 | 0.27025 | 0.27025 | 0.024305
2019-03-19 15:58 | 1553025501.7733 | 972 | 0.20867 | 0.000537 | 0.000247 | 101.1234 | 101.2188 | 121.887 | 21.2612 | 0.002678 | 0.002531 | 221.1231 | 0.26576 | 8.4017 | 0.003471 | 0.004666 | 0.00456 | 4.8e-05 | 0.000992 | 0.30182 | 1.3188 | 2.5482 | 101.19 | 20.77 | 18.168 | 86.2365 | 105.3968 | 82.7065 | 0.004407 | 0.9921 | 0.024346 | 24.5403 | -0.008433 | 0.1581 | 0.00412 | 0.28748 | 0.28748 | 0.022878
2019-03-19 15:58 | 1553025502.7722 | 973 | 0.20869 | 0.000536 | 0.000246 | 101.2285 | 101.2466 | 121.6148 | 21.2612 | 0.00266 | 0.002507 | 221.1167 | 0.26489 | 8.4047 | 0.003473 | 0.005023 | 0.004961 | 4.6e-05 | 0.000989 | 0.3095 | 1.3315 | 2.548 | 101.19 | 20.77 | 18.155 | 85.8057 | 105.4259 | 82.131 | 0.004404 | 0.99371 | 0.024339 | 24.493 | -0.001599 | 0.15758 | 0.004002 | 0.2788 | 0.2788 | 0.023513
2019-03-19 15:58 | 1553025503.7733 | 974 | 0.20869 | 0.000536 | 0.000245 | 101.3095 | 101.576 | 121.4622 | 21.2612 | 0.002668 | 0.002495 | 221.1252 | 0.26508 | 8.4034 | 0.003436 | 0.004612 | 0.004524 | 4.6e-05 | 0.000987 | 0.33826 | 1.356 | 2.5482 | 101.19 | 20.77 | 18.141 | 86.0017 | 104.5386 | 81.8485 | 0.004401 | 0.9945 | 0.024333 | 24.4679 | -0.023502 | 0.15769 | 0.004019 | 0.27998 | 0.27998 | 0.02343
2019-03-19 15:58 | 1553025504.7737 | 975 | 0.20865 | 0.00054 | 0.000244 | 101.2855 | 101.4709 | 121.6211 | 21.2612 | 0.002687 | 0.002488 | 221.0962 | 0.26424 | 8.4019 | 0.003359 | 0.005089 | 0.004922 | 5e-05 | 0.00099 | 0.28783 | 1.3545 | 2.5481 | 101.19 | 20.77 | 18.106 | 86.4661 | 102.6658 | 81.6662 | 0.004392 | 0.99501 | 0.024355 | 24.4776 | -0.016359 | 0.15719 | 0.004214 | 0.29462 | 0.29462 | 0.022196
2019-03-19 15:58 | 1553025505.7725 | 976 | 0.20871 | 0.00054 | 0.000243 | 101.3497 | 101.5007 | 121.5446 | 21.2612 | 0.002649 | 0.002509 | 221.0606 | 0.26527 | 8.4051 | 0.00337 | 0.005138 | 0.005021 | 5.3e-05 | 0.000988 | 0.28368 | 1.3829 | 2.5482 | 101.19 | 20.77 | 18.087 | 85.5372 | 102.9406 | 82.1863 | 0.004388 | 0.99355 | 0.02433 | 24.4878 | -0.013315 | 0.15781 | 0.003841 | 0.26701 | 0.26701 | 0.024587
2019-03-19 15:58 | 1553025506.7734 | 977 | 0.20863 | 0.000535 | 0.000242 | 101.422 | 101.5897 | 121.6636 | 21.2612 | 0.002671 | 0.00245 | 221.046 | 0.26474 | 8.4039 | 0.003477 | 0.00484 | 0.004733 | 5.1e-05 | 0.000991 | 0.31045 | 1.3336 | 2.5481 | 101.19 | 20.77 | 18.043 | 86.0727 | 105.5345 | 80.7694 | 0.004377 | 0.99753 | 0.024391 | 24.4509 | -0.014778 | 0.15749 | 0.004344 | 0.30483 | 0.30483 | 0.021493
2019-03-19 15:58 | 1553025507.7729 | 978 | 0.20869 | 0.000536 | 0.00024 | 101.4998 | 101.6545 | 121.6096 | 21.2612 | 0.00267 | 0.002482 | 221.0075 | 0.26505 | 8.4026 | 0.003441 | 0.004647 | 0.004535 | 5.2e-05 | 0.000986 | 0.32068 | 1.2917 | 2.548 | 101.2 | 20.77 | 18.022 | 86.0523 | 104.6654 | 81.5243 | 0.004372 | 0.99541 | 0.024359 | 24.4716 | -0.013623 | 0.15768 | 0.003988 | 0.27831 | 0.27831 | 0.023569
2019-03-19 15:58 | 1553025508.7739 | 979 | 0.20868 | 0.000539 | 0.000238 | 101.5826 | 101.713 | 121.781 | 21.2612 | 0.002677 | 0.002534 | 220.9858 | 0.26526 | 8.4034 | 0.00338 | 0.00515 | 0.005016 | 5.3e-05 | 0.000991 | 0.29479 | 1.3054 | 2.5481 | 101.2 | 20.77 | 17.979 | 86.207 | 103.1913 | 82.7753 | 0.004361 | 0.99191 | 0.024333 | 24.5319 | -0.011479 | 0.1578 | 0.004058 | 0.28332 | 0.28332 | 0.02317
2019-03-19 15:58 | 1553025509.7727 | 980 | 0.20869 | 0.000544 | 0.000236 | 101.6087 | 101.7606 | 121.8246 | 21.2612 | 0.002662 | 0.002455 | 220.988 | 0.26576 | 8.4024 | 0.003456 | 0.004727 | 0.004674 | 5.5e-05 | 0.000989 | 0.33339 | 1.3021 | 2.548 | 101.2 | 20.77 | 17.958 | 85.8456 | 105.0295 | 80.8785 | 0.004356 | 0.99722 | 0.024403 | 24.4709 | -0.013366 | 0.1581 | 0.004017 | 0.28122 | 0.28122 | 0.023387
2019-03-19 15:58 | 1553025510.7736 | 981 | 0.20868 | 0.000539 | 0.000234 | 101.5991 | 101.7082 | 122.0617 | 21.2612 | 0.002694 | 0.00248 | 220.9933 | 0.26493 | 8.4035 | 0.003393 | 0.004845 | 0.004671 | 5.2e-05 | 0.000988 | 0.31885 | 1.3323 | 2.5481 | 101.2 | 20.77 | 17.939 | 86.616 | 103.5081 | 81.4806 | 0.004351 | 0.99553 | 0.024406 | 24.5155 | -0.009604 | 0.1576 | 0.004037 | 0.28291 | 0.28291 | 0.023174
2019-03-19 15:58 | 1553025511.7737 | 982 | 0.20868 | 0.000542 | 0.00023 | 101.7056 | 101.9398 | 121.819 | 21.2612 | 0.002633 | 0.002451 | 220.9823 | 0.26547 | 8.4024 | 0.003457 | 0.004989 | 0.00487 | 5.6e-05 | 0.000991 | 0.31885 | 1.3243 | 2.5481 | 101.2 | 20.77 | 17.903 | 85.1588 | 105.0477 | 80.7951 | 0.004343 | 0.99746 | 0.024405 | 24.4674 | -0.020567 | 0.15793 | 0.004055 | 0.28444 | 0.28444 | 0.023097
2019-03-19 15:59 | 1553025512.7731 | 983 | 0.20871 | 0.000541 | 0.000227 | 101.6678 | 101.7863 | 121.5979 | 21.2612 | 0.002626 | 0.002497 | 220.9818 | 0.26559 | 8.4013 | 0.003389 | 0.005004 | 0.004911 | 5.3e-05 | 0.000989 | 0.30613 | 1.3367 | 2.5479 | 101.2 | 20.77 | 17.887 | 85.0035 | 103.4076 | 81.8924 | 0.004339 | 0.99438 | 0.024345 | 24.4831 | -0.01042 | 0.158 | 0.00391 | 0.27324 | 0.27324 | 0.024055
2019-03-19 15:59 | 1553025513.7738 | 984 | 0.2087 | 0.000543 | 0.000224 | 101.7494 | 101.9791 | 121.9934 | 21.2612 | 0.002714 | 0.002542 | 220.9545 | 0.26485 | 8.4027 | 0.003469 | 0.004989 | 0.004888 | 5.8e-05 | 0.000994 | 0.30198 | 1.3216 | 2.548 | 101.2 | 20.77 | 17.873 | 87.1172 | 105.3308 | 82.9741 | 0.004335 | 0.99135 | 0.024348 | 24.5602 | -0.02017 | 0.15756 | 0.003971 | 0.27791 | 0.27791 | 0.023584
2019-03-19 15:59 | 1553025514.7739 | 985 | 0.2087 | 0.000536 | 0.000224 | 101.862 | 101.8922 | 122.0591 | 21.2612 | 0.002641 | 0.002443 | 220.9522 | 0.26458 | 8.4031 | 0.003474 | 0.005384 | 0.005224 | 6e-05 | 0.000995 | 0.29446 | 1.3509 | 2.5481 | 101.2 | 20.77 | 17.843 | 85.349 | 105.4637 | 80.5908 | 0.004328 | 0.99803 | 0.024436 | 24.4845 | -0.002649 | 0.1574 | 0.003982 | 0.27979 | 0.27979 | 0.023402
2019-03-19 15:59 | 1553025515.7733 | 986 | 0.20872 | 0.000545 | 0.000222 | 101.859 | 102.0158 | 121.9692 | 21.2612 | 0.002716 | 0.002458 | 220.9861 | 0.26517 | 8.4018 | 0.003468 | 0.00443 | 0.004287 | 5.7e-05 | 0.000998 | 0.31238 | 1.3055 | 2.5481 | 101.2 | 20.77 | 17.83 | 87.1687 | 105.3267 | 80.9463 | 0.004325 | 0.99703 | 0.024415 | 24.4877 | -0.01376 | 0.15775 | 0.003841 | 0.26926 | 0.26926 | 0.024373
2019-03-19 15:59 | 1553025516.7744 | 987 | 0.2087 | 0.000541 | 0.00022 | 101.8968 | 101.9795 | 121.9476 | 21.2612 | 0.002694 | 0.002439 | 220.9907 | 0.26543 | 8.4011 | 0.003468 | 0.005281 | 0.005102 | 5.9e-05 | 0.001005 | 0.31598 | 1.3223 | 2.5482 | 101.2 | 20.77 | 17.807 | 86.6292 | 105.3221 | 80.5067 | 0.004319 | 0.99827 | 0.024428 | 24.4704 | -0.007256 | 0.1579 | 0.00398 | 0.27983 | 0.27983 | 0.023474
2019-03-19 15:59 | 1553025517.7725 | 988 | 0.20869 | 0.00054 | 0.000217 | 101.9097 | 102.1062 | 121.9032 | 21.2612 | 0.002654 | 0.002448 | 220.9872 | 0.26537 | 8.4007 | 0.003394 | 0.005001 | 0.004985 | 5.9e-05 | 0.001004 | 0.29695 | 1.3455 | 2.5483 | 101.2 | 20.77 | 17.796 | 85.6574 | 103.5301 | 80.721 | 0.004317 | 0.99767 | 0.024416 | 24.4733 | -0.017231 | 0.15787 | 0.004032 | 0.28362 | 0.28362 | 0.023155
2019-03-19 15:59 | 1553025518.7732 | 989 | 0.20872 | 0.00054 | 0.000213 | 101.9622 | 102.2158 | 121.9596 | 21.2612 | 0.002651 | 0.002441 | 220.9913 | 0.2652 | 8.4036 | 0.003398 | 0.00466 | 0.004554 | 5.4e-05 | 0.001005 | 0.3115 | 1.4102 | 2.5482 | 101.2 | 20.77 | 17.785 | 85.6019 | 103.6236 | 80.5526 | 0.004314 | 0.99814 | 0.024428 | 24.4731 | -0.022219 | 0.15777 | 0.003826 | 0.26886 | 0.26886 | 0.02441
2019-03-19 15:59 | 1553025519.7736 | 990 | 0.2087 | 0.000538 | 0.000209 | 101.9088 | 102.0863 | 121.9899 | 21.2611 | 0.00262 | 0.002489 | 220.9645 | 0.26531 | 8.4014 | 0.003339 | 0.005482 | 0.005382 | 5.6e-05 | 0.001012 | 0.29878 | 1.38 | 2.548 | 101.2 | 20.77 | 17.785 | 84.8449 | 102.1803 | 81.6935 | 0.004314 | 0.99493 | 0.024391 | 24.5156 | -0.01557 | 0.15783 | 0.003991 | 0.28082 | 0.28082 | 0.023381
2019-03-19 15:59 | 1553025520.7736 | 991 | 0.20872 | 0.000533 | 0.000204 | 102.069 | 102.0976 | 121.453 | 21.2611 | 0.002634 | 0.002496 | 220.9454 | 0.26553 | 8.403 | 0.003436 | 0.004765 | 0.004647 | 5.7e-05 | 0.001015 | 0.31188 | 1.365 | 2.5481 | 101.19 | 20.77 | 17.756 | 85.1845 | 104.5277 | 81.8548 | 0.004307 | 0.99448 | 0.024332 | 24.4672 | -0.002506 | 0.15796 | 0.003897 | 0.2736 | 0.2736 | 0.024017
2019-03-19 15:59 | 1553025521.7723 | 992 | 0.20873 | 0.00054 | 0.0002 | 102.0141 | 102.1956 | 121.5852 | 21.2612 | 0.002594 | 0.002451 | 220.9413 | 0.26575 | 8.4026 | 0.003401 | 0.005284 | 0.005223 | 5.7e-05 | 0.001017 | 0.28799 | 1.3706 | 2.5482 | 101.19 | 20.77 | 17.747 | 84.2293 | 103.6896 | 80.7937 | 0.004305 | 0.99746 | 0.024382 | 24.4439 | -0.015897 | 0.15809 | 0.003785 | 0.26617 | 0.26617 | 0.024709
2019-03-19 15:59 | 1553025522.7733 | 993 | 0.20872 | 0.00054 | 0.000197 | 101.9853 | 102.2394 | 122.2039 | 21.2611 | 0.002673 | 0.002517 | 220.9229 | 0.26549 | 8.4036 | 0.00332 | 0.004986 | 0.004957 | 5.7e-05 | 0.001019 | 0.3183 | 1.3467 | 2.5481 | 101.19 | 20.77 | 17.738 | 86.1147 | 101.7366 | 82.3673 | 0.004303 | 0.99305 | 0.02439 | 24.5604 | -0.02226 | 0.15794 | 0.003848 | 0.27109 | 0.27109 | 0.024236
2019-03-19 15:59 | 1553025523.7727 | 994 | 0.20868 | 0.000538 | 0.000196 | 101.9985 | 102.2953 | 122.2875 | 21.2612 | 0.002636 | 0.002486 | 220.9072 | 0.26454 | 8.406 | 0.003239 | 0.004797 | 0.004645 | 5.8e-05 | 0.001019 | 0.28982 | 1.3553 | 2.5481 | 101.2 | 20.77 | 17.724 | 85.2437 | 99.7806 | 81.6171 | 0.004299 | 0.99515 | 0.024424 | 24.5429 | -0.025984 | 0.15738 | 0.004104 | 0.2904 | 0.2904 | 0.022544
2019-03-19 15:59 | 1553025524.7725 | 995 | 0.20867 | 0.000538 | 0.000194 | 102.0426 | 102.2309 | 122.1492 | 21.261 | 0.002679 | 0.002441 | 220.8968 | 0.26607 | 8.4027 | 0.003345 | 0.004789 | 0.004721 | 5.8e-05 | 0.00102 | 0.31598 | 1.3964 | 2.5482 | 101.2 | 20.77 | 17.717 | 86.2777 | 102.3331 | 80.5493 | 0.004298 | 0.99815 | 0.024447 | 24.4921 | -0.016485 | 0.15829 | 0.004167 | 0.29539 | 0.29539 | 0.022291
2019-03-19 15:59 | 1553025525.7725 | 996 | 0.20867 | 0.000544 | 0.000192 | 102.0402 | 102.2131 | 122.1889 | 21.261 | 0.002629 | 0.00251 | 220.8563 | 0.26528 | 8.4018 | 0.003343 | 0.004701 | 0.004627 | 6e-05 | 0.001026 | 0.34954 | 1.3482 | 2.5479 | 101.2 | 20.77 | 17.711 | 85.0626 | 102.2925 | 82.1962 | 0.004296 | 0.99352 | 0.024394 | 24.553 | -0.015137 | 0.15781 | 0.004166 | 0.29484 | 0.29484 | 0.022266
2019-03-19 15:59 | 1553025526.7735 | 997 | 0.20867 | 0.000538 | 0.000189 | 101.9817 | 102.1271 | 122.1244 | 21.2611 | 0.002655 | 0.002437 | 220.8351 | 0.26508 | 8.4033 | 0.003291 | 0.004961 | 0.004868 | 5.8e-05 | 0.001023 | 0.32444 | 1.3665 | 2.5481 | 101.2 | 20.77 | 17.702 | 85.6845 | 101.0227 | 80.4441 | 0.004294 | 0.99845 | 0.024448 | 24.4859 | -0.012742 | 0.15769 | 0.004151 | 0.29454 | 0.29454 | 0.022272
2019-03-19 15:59 | 1553025527.7731 | 998 | 0.20869 | 0.000541 | 0.000187 | 102.0135 | 102.2733 | 122.2887 | 21.261 | 0.002688 | 0.002478 | 220.8512 | 0.26538 | 8.4027 | 0.003351 | 0.005019 | 0.004966 | 6.2e-05 | 0.001026 | 0.29767 | 1.372 | 2.5481 | 101.2 | 20.77 | 17.698 | 86.4738 | 102.4881 | 81.4289 | 0.004293 | 0.99567 | 0.02443 | 24.5365 | -0.022752 | 0.15788 | 0.004078 | 0.28908 | 0.28908 | 0.022719
2019-03-19 15:59 | 1553025528.7727 | 999 | 0.20868 | 0.000542 | 0.000184 | 101.9622 | 102.0857 | 121.9429 | 21.261 | 0.002709 | 0.002434 | 220.845 | 0.26601 | 8.4018 | 0.003368 | 0.004886 | 0.00482 | 6.1e-05 | 0.001027 | 0.32211 | 1.348 | 2.548 | 101.2 | 20.77 | 17.691 | 86.9873 | 102.8973 | 80.3724 | 0.004291 | 0.99865 | 0.024432 | 24.4652 | -0.010823 | 0.15825 | 0.004104 | 0.29125 | 0.29125 | 0.022603
2019-03-19 15:59 | 1553025529.7731 | 1000 | 0.20871 | 0.000549 | 0.000181 | 101.9886 | 101.9932 | 121.9429 | 21.2611 | 0.002666 | 0.002447 | 220.8377 | 0.26538 | 8.4027 | 0.003418 | 0.005132 | 0.004973 | 6.3e-05 | 0.001027 | 0.32532 | 1.3211 | 2.548 | 101.2 | 20.78 | 17.689 | 85.946 | 104.1125 | 80.6866 | 0.004293 | 0.99776 | 0.024421 | 24.4761 | -0.000401 | 0.15788 | 0.00396 | 0.28066 | 0.28066 | 0.023401
2019-03-19 15:59 | 1553025530.7731 | 1001 | 0.20871 | 0.000547 | 0.000178 | 101.919 | 102.0997 | 122.3549 | 21.261 | 0.002623 | 0.002475 | 220.833 | 0.26531 | 8.4016 | 0.003446 | 0.004375 | 0.004331 | 5.8e-05 | 0.001022 | 0.31758 | 1.3374 | 2.5481 | 101.2 | 20.77 | 17.688 | 84.9159 | 104.7753 | 81.3564 | 0.004291 | 0.99588 | 0.024439 | 24.5406 | -0.015845 | 0.15783 | 0.003952 | 0.28047 | 0.28047 | 0.02341
2019-03-19 15:59 | 1553025531.7742 | 1002 | 0.20875 | 0.000555 | 0.000175 | 101.8614 | 102.0826 | 122.5581 | 21.2609 | 0.00267 | 0.002426 | 220.8385 | 0.26547 | 8.4032 | 0.003392 | 0.004999 | 0.004875 | 6e-05 | 0.001022 | 0.32244 | 1.3135 | 2.548 | 101.2 | 20.77 | 17.686 | 86.0606 | 103.4662 | 80.1827 | 0.00429 | 0.99919 | 0.0245 | 24.5203 | -0.019403 | 0.15793 | 0.003701 | 0.26279 | 0.26279 | 0.025001
2019-03-19 15:59 | 1553025532.7719 | 1003 | 0.20869 | 0.000552 | 0.000174 | 101.8053 | 102.0911 | 122.2321 | 21.2609 | 0.002645 | 0.002474 | 220.8258 | 0.26575 | 8.4048 | 0.003377 | 0.004586 | 0.0045 | 6.2e-05 | 0.001025 | 0.31404 | 1.353 | 2.548 | 101.2 | 20.77 | 17.685 | 85.4458 | 103.0975 | 81.3313 | 0.00429 | 0.99595 | 0.024428 | 24.5274 | -0.025082 | 0.1581 | 0.004041 | 0.2871 | 0.2871 | 0.022907
2019-03-19 15:59 | 1553025533.7733 | 1004 | 0.20869 | 0.00055 | 0.000172 | 101.7047 | 102.0233 | 121.8877 | 21.2608 | 0.002658 | 0.002456 | 220.8195 | 0.26536 | 8.4026 | 0.003406 | 0.005563 | 0.005421 | 5.7e-05 | 0.001024 | 0.33467 | 1.3797 | 2.5481 | 101.2 | 20.78 | 17.685 | 85.7593 | 103.811 | 80.9156 | 0.004292 | 0.99712 | 0.024408 | 24.4785 | -0.027976 | 0.15786 | 0.004042 | 0.28705 | 0.28705 | 0.022877
2019-03-19 15:59 | 1553025534.7719 | 1005 | 0.2087 | 0.000552 | 0.000173 | 101.6198 | 101.8987 | 121.8375 | 21.2609 | 0.002641 | 0.002457 | 220.8169 | 0.26459 | 8.4026 | 0.003371 | 0.004332 | 0.004254 | 5.5e-05 | 0.001017 | 0.31045 | 1.3595 | 2.5481 | 101.2 | 20.78 | 17.685 | 85.3457 | 102.968 | 80.9314 | 0.004292 | 0.99707 | 0.024402 | 24.474 | -0.024512 | 0.15741 | 0.004005 | 0.28423 | 0.28423 | 0.023039
2019-03-19 15:59 | 1553025535.773 | 1006 | 0.2087 | 0.000552 | 0.000173 | 101.6243 | 101.8007 | 122.0753 | 21.2608 | 0.002638 | 0.002472 | 220.8125 | 0.26559 | 8.4039 | 0.003354 | 0.004722 | 0.004652 | 5.6e-05 | 0.001017 | 0.32045 | 1.3429 | 2.548 | 101.2 | 20.78 | 17.687 | 85.2753 | 102.553 | 81.2922 | 0.004293 | 0.99606 | 0.024414 | 24.5103 | -0.015514 | 0.158 | 0.004025 | 0.28584 | 0.28584 | 0.022994
2019-03-19 15:59 | 1553025536.7734 | 1007 | 0.20871 | 0.000556 | 0.000173 | 101.5169 | 101.5897 | 122.0532 | 21.2608 | 0.002642 | 0.002478 | 220.811 | 0.26545 | 8.4037 | 0.003268 | 0.00452 | 0.004343 | 5.7e-05 | 0.001021 | 0.28838 | 1.3556 | 2.5482 | 101.2 | 20.78 | 17.695 | 85.3682 | 100.4692 | 81.442 | 0.004295 | 0.99564 | 0.024406 | 24.5133 | -0.006419 | 0.15792 | 0.003962 | 0.28108 | 0.28108 | 0.023372
2019-03-19 15:59 | 1553025537.7721 | 1008 | 0.20869 | 0.000549 | 0.000174 | 101.4625 | 101.7106 | 121.6629 | 21.2609 | 0.002684 | 0.002447 | 220.775 | 0.26509 | 8.4043 | 0.003338 | 0.005442 | 0.005366 | 5.7e-05 | 0.001018 | 0.31006 | 1.3841 | 2.548 | 101.2 | 20.78 | 17.7 | 86.3962 | 102.1706 | 80.6934 | 0.004296 | 0.99774 | 0.024393 | 24.4482 | -0.021844 | 0.1577 | 0.004076 | 0.28934 | 0.28934 | 0.022673
2019-03-19 15:59 | 1553025538.7731 | 1009 | 0.20872 | 0.00055 | 0.000172 | 101.3839 | 101.4349 | 121.7544 | 21.2605 | 0.002651 | 0.002472 | 220.7617 | 0.26497 | 8.4031 | 0.003333 | 0.004164 | 0.00406 | 6e-05 | 0.001019 | 0.3246 | 1.3707 | 2.5483 | 101.2 | 20.78 | 17.705 | 85.5866 | 102.0398 | 81.2965 | 0.004297 | 0.99605 | 0.024381 | 24.4783 | -0.004499 | 0.15763 | 0.003899 | 0.27626 | 0.27626 | 0.023736
2019-03-19 15:59 | 1553025539.7732 | 1010 | 0.20875 | 0.00055 | 0.000172 | 101.311 | 101.4688 | 121.8171 | 21.2604 | 0.002566 | 0.0025 | 220.73 | 0.26555 | 8.4024 | 0.003321 | 0.004775 | 0.004626 | 6e-05 | 0.001014 | 0.29789 | 1.331 | 2.5482 | 101.2 | 20.78 | 17.717 | 83.5475 | 101.7632 | 81.9704 | 0.0043 | 0.99416 | 0.024365 | 24.5078 | -0.013925 | 0.15798 | 0.003691 | 0.26085 | 0.26085 | 0.025194
2019-03-19 15:59 | 1553025540.7737 | 1011 | 0.20871 | 0.000552 | 0.000169 | 101.1324 | 101.2681 | 121.73 | 21.2606 | 0.002623 | 0.002423 | 220.7042 | 0.26495 | 8.4043 | 0.003304 | 0.004149 | 0.004061 | 5.8e-05 | 0.00101 | 0.31598 | 1.3281 | 2.5481 | 101.2 | 20.78 | 17.722 | 84.9239 | 101.3496 | 80.1066 | 0.004301 | 0.9994 | 0.02442 | 24.4347 | -0.011994 | 0.15762 | 0.003947 | 0.28041 | 0.28041 | 0.023383
2019-03-19 15:59 | 1553025541.7724 | 1012 | 0.20872 | 0.000552 | 0.000167 | 100.9785 | 101.0003 | 121.3826 | 21.2604 | 0.002624 | 0.002423 | 220.685 | 0.26536 | 8.4053 | 0.003272 | 0.004736 | 0.004552 | 5.7e-05 | 0.001007 | 0.28335 | 1.3182 | 2.548 | 101.2 | 20.78 | 17.729 | 84.9379 | 100.5715 | 80.119 | 0.004303 | 0.99937 | 0.024385 | 24.4002 | -0.001934 | 0.15786 | 0.003888 | 0.27576 | 0.27576 | 0.023814
2019-03-19 15:59 | 1553025542.7735 | 1013 | 0.20875 | 0.000542 | 0.000165 | 100.8563 | 100.9736 | 121.466 | 21.2604 | 0.002629 | 0.002454 | 220.6749 | 0.26463 | 8.4019 | 0.003345 | 0.00446 | 0.004389 | 6.3e-05 | 0.001005 | 0.33804 | 1.339 | 2.548 | 101.2 | 20.78 | 17.73 | 85.058 | 102.3382 | 80.8465 | 0.004303 | 0.99731 | 0.024368 | 24.4337 | -0.010393 | 0.15743 | 0.003716 | 0.26308 | 0.26308 | 0.024894
2019-03-19 15:59 | 1553025543.7724 | 1014 | 0.20874 | 0.000547 | 0.000163 | 100.8089 | 101.0859 | 121.4143 | 21.2604 | 0.002621 | 0.00245 | 220.6654 | 0.26525 | 8.4024 | 0.003335 | 0.005314 | 0.005197 | 6.2e-05 | 0.000993 | 0.29734 | 1.3264 | 2.548 | 101.2 | 20.78 | 17.733 | 84.8809 | 102.0956 | 80.7669 | 0.004304 | 0.99754 | 0.024366 | 24.4258 | -0.024545 | 0.1578 | 0.003789 | 0.26854 | 0.26854 | 0.024445
2019-03-19 15:59 | 1553025544.7726 | 1015 | 0.20878 | 0.000549 | 0.000161 | 100.612 | 100.6701 | 121.4329 | 21.2603 | 0.002651 | 0.002441 | 220.6532 | 0.26608 | 8.4037 | 0.003415 | 0.004363 | 0.004315 | 5.9e-05 | 0.00098 | 0.29822 | 1.3223 | 2.5481 | 101.2 | 20.78 | 17.735 | 85.5825 | 104.0378 | 80.5471 | 0.004305 | 0.99816 | 0.024375 | 24.4201 | -0.005163 | 0.15829 | 0.003552 | 0.25136 | 0.25136 | 0.026197
2019-03-19 15:59 | 1553025545.7727 | 1016 | 0.20872 | 0.000548 | 0.000159 | 100.4434 | 100.4345 | 121.4007 | 21.2603 | 0.002629 | 0.002501 | 220.6423 | 0.26566 | 8.4039 | 0.003322 | 0.005875 | 0.005702 | 6.2e-05 | 0.000977 | 0.3282 | 1.3431 | 2.5481 | 101.2 | 20.78 | 17.737 | 85.0533 | 101.7868 | 81.9796 | 0.004305 | 0.99413 | 0.024323 | 24.4662 | 0.000791 | 0.15804 | 0.0039 | 0.2765 | 0.2765 | 0.023778
2019-03-19 15:59 | 1553025546.7726 | 1017 | 0.20871 | 0.00055 | 0.000157 | 100.2591 | 100.1492 | 121.7591 | 21.26 | 0.002617 | 0.002436 | 220.6154 | 0.26604 | 8.4008 | 0.00332 | 0.00487 | 0.004884 | 6.4e-05 | 0.000977 | 0.28926 | 1.3223 | 2.5482 | 101.2 | 20.78 | 17.739 | 84.7805 | 101.7345 | 80.4306 | 0.004306 | 0.99849 | 0.024412 | 24.4488 | 0.009811 | 0.15827 | 0.003944 | 0.28079 | 0.28079 | 0.023448
2019-03-19 15:59 | 1553025547.7727 | 1018 | 0.2087 | 0.000549 | 0.000155 | 100.2192 | 100.2468 | 122.0539 | 21.2599 | 0.002607 | 0.002462 | 220.6175 | 0.26538 | 8.4004 | 0.003354 | 0.005146 | 0.005013 | 6.1e-05 | 0.000977 | 0.30845 | 1.3295 | 2.5479 | 101.2 | 20.79 | 17.738 | 84.5301 | 102.5529 | 81.0496 | 0.004308 | 0.99674 | 0.02442 | 24.4998 | -0.002464 | 0.15787 | 0.004008 | 0.28577 | 0.28577 | 0.022982
2019-03-19 15:59 | 1553025548.7718 | 1019 | 0.2087 | 0.00055 | 0.000153 | 99.9518 | 99.8173 | 122.2105 | 21.26 | 0.002648 | 0.00245 | 220.6026 | 0.26571 | 8.4035 | 0.003347 | 0.004842 | 0.004769 | 5.8e-05 | 0.000978 | 0.286 | 1.3524 | 2.5481 | 101.2 | 20.79 | 17.736 | 85.5079 | 102.3762 | 80.7651 | 0.004308 | 0.99754 | 0.024445 | 24.5057 | 0.012042 | 0.15807 | 0.004035 | 0.28823 | 0.28823 | 0.022814
2019-03-19 15:59 | 1553025549.7718 | 1020 | 0.2087 | 0.000547 | 0.00015 | 99.7928 | 99.6153 | 122.0089 | 21.2599 | 0.002664 | 0.00242 | 220.6206 | 0.26569 | 8.4045 | 0.003358 | 0.004526 | 0.004426 | 6.2e-05 | 0.000975 | 0.29695 | 1.3337 | 2.548 | 101.2 | 20.79 | 17.728 | 85.9076 | 102.6558 | 80.0414 | 0.004306 | 0.99959 | 0.02445 | 24.4604 | 0.015923 | 0.15806 | 0.004026 | 0.28776 | 0.28776 | 0.02285
2019-03-19 15:59 | 1553025550.7732 | 1021 | 0.20869 | 0.000545 | 0.000146 | 99.5317 | 99.4238 | 121.6803 | 21.2597 | 0.002676 | 0.002415 | 220.639 | 0.26561 | 8.4035 | 0.003356 | 0.004523 | 0.004429 | 6.3e-05 | 0.000977 | 0.30055 | 1.3512 | 2.5481 | 101.2 | 20.79 | 17.723 | 86.1852 | 102.5897 | 79.9133 | 0.004304 | 0.99995 | 0.024422 | 24.423 | 0.009698 | 0.15801 | 0.004103 | 0.29335 | 0.29335 | 0.022408
2019-03-19 15:59 | 1553025551.7719 | 1022 | 0.20875 | 0.000549 | 0.000143 | 99.3156 | 99.1557 | 121.8173 | 21.2595 | 0.002653 | 0.002461 | 220.6306 | 0.2656 | 8.4016 | 0.003385 | 0.004335 | 0.004192 | 6.6e-05 | 0.00098 | 0.31852 | 1.3477 | 2.548 | 101.2 | 20.79 | 17.715 | 85.6315 | 103.302 | 81.0196 | 0.004302 | 0.99683 | 0.024397 | 24.475 | 0.014417 | 0.158 | 0.003775 | 0.26911 | 0.26911 | 0.024425
2019-03-19 15:59 | 1553025552.7729 | 1023 | 0.20877 | 0.000549 | 0.000141 | 99.1169 | 98.9954 | 121.7617 | 21.2597 | 0.002661 | 0.002484 | 220.6277 | 0.26586 | 8.4056 | 0.003411 | 0.004776 | 0.004704 | 6.6e-05 | 0.000984 | 0.29927 | 1.3781 | 2.5481 | 101.2 | 20.79 | 17.711 | 85.8323 | 103.9413 | 81.5813 | 0.004301 | 0.99525 | 0.024372 | 24.4888 | 0.010977 | 0.15816 | 0.003656 | 0.26017 | 0.26017 | 0.025289
2019-03-19 15:59 | 1553025553.7731 | 1024 | 0.20874 | 0.00055 | 0.000141 | 98.9111 | 98.8015 | 122.0661 | 21.2595 | 0.00264 | 0.002493 | 220.6163 | 0.26538 | 8.4032 | 0.003316 | 0.005008 | 0.004827 | 6e-05 | 0.00099 | 0.33196 | 1.3741 | 2.5481 | 101.2 | 20.79 | 17.706 | 85.3147 | 101.6322 | 81.8034 | 0.0043 | 0.99462 | 0.024395 | 24.5271 | 0.009912 | 0.15788 | 0.003821 | 0.27257 | 0.27257 | 0.024095
2019-03-19 15:59 | 1553025554.773 | 1025 | 0.20874 | 0.00055 | 0.00014 | 98.7514 | 98.5094 | 121.9417 | 21.2599 | 0.002626 | 0.002452 | 220.6234 | 0.26535 | 8.4023 | 0.003285 | 0.005022 | 0.004922 | 6.2e-05 | 0.001001 | 0.35834 | 1.3534 | 2.5481 | 101.2 | 20.79 | 17.7 | 84.9943 | 100.8866 | 80.7955 | 0.004299 | 0.99746 | 0.024418 | 24.4798 | 0.02195 | 0.15786 | 0.003784 | 0.27017 | 0.27017 | 0.024306
2019-03-19 15:59 | 1553025555.7734 | 1026 | 0.20873 | 0.000548 | 0.000139 | 98.5173 | 98.2936 | 121.6601 | 21.2592 | 0.0026 | 0.002467 | 220.6164 | 0.26538 | 8.4019 | 0.003298 | 0.004605 | 0.004484 | 6.2e-05 | 0.001003 | 0.3298 | 1.3504 | 2.5481 | 101.2 | 20.79 | 17.698 | 84.3531 | 101.1883 | 81.1666 | 0.004298 | 0.99641 | 0.024377 | 24.4643 | 0.020336 | 0.15788 | 0.003865 | 0.27576 | 0.27576 | 0.023816
2019-03-19 15:59 | 1553025556.7721 | 1027 | 0.20872 | 0.000545 | 0.000138 | 98.2043 | 97.8717 | 121.2589 | 21.2592 | 0.002639 | 0.00247 | 220.6142 | 0.26565 | 8.4015 | 0.003356 | 0.004844 | 0.004691 | 6.4e-05 | 0.001011 | 0.3147 | 1.3792 | 2.548 | 101.2 | 20.79 | 17.695 | 85.2974 | 102.5956 | 81.24 | 0.004298 | 0.99621 | 0.024334 | 24.4265 | 0.030351 | 0.15803 | 0.003939 | 0.28074 | 0.28074 | 0.023418
2019-03-19 15:59 | 1553025557.7723 | 1028 | 0.20871 | 0.000546 | 0.000137 | 97.9958 | 97.7042 | 120.9744 | 21.2595 | 0.002566 | 0.002462 | 220.5997 | 0.26508 | 8.4035 | 0.00335 | 0.005207 | 0.005049 | 6.3e-05 | 0.001017 | 0.30934 | 1.3653 | 2.5483 | 101.2 | 20.79 | 17.695 | 83.5442 | 102.4488 | 81.0477 | 0.004298 | 0.99675 | 0.024312 | 24.3912 | 0.026653 | 0.1577 | 0.004 | 0.28502 | 0.28502 | 0.023017
2019-03-19 15:59 | 1553025558.7732 | 1029 | 0.20873 | 0.000548 | 0.000136 | 97.8835 | 97.5552 | 121.0409 | 21.2589 | 0.002639 | 0.002392 | 220.5932 | 0.26537 | 8.4008 | 0.00342 | 0.004492 | 0.004401 | 6.2e-05 | 0.001014 | 0.29695 | 1.3763 | 2.5481 | 101.2 | 20.79 | 17.696 | 85.2992 | 104.1488 | 79.3745 | 0.004298 | 1.0015 | 0.024376 | 24.3402 | 0.030052 | 0.15787 | 0.003879 | 0.27692 | 0.27692 | 0.023715
2019-03-19 15:59 | 1553025559.7733 | 1030 | 0.20875 | 0.000548 | 0.000135 | 97.6407 | 97.3442 | 121.161 | 21.2591 | 0.00266 | 0.002444 | 220.5847 | 0.26498 | 8.4051 | 0.003308 | 0.004595 | 0.004484 | 6e-05 | 0.001016 | 0.30773 | 1.3178 | 2.5481 | 101.2 | 20.79 | 17.699 | 85.8115 | 101.4439 | 80.6183 | 0.004299 | 0.99796 | 0.024345 | 24.3952 | 0.027204 | 0.15764 | 0.003732 | 0.26588 | 0.26588 | 0.024663
2019-03-19 15:59 | 1553025560.7723 | 1031 | 0.20875 | 0.000551 | 0.000134 | 97.3961 | 97.0236 | 121.5974 | 21.2592 | 0.002563 | 0.002458 | 220.5607 | 0.26628 | 8.4021 | 0.003342 | 0.005141 | 0.004985 | 5.9e-05 | 0.001015 | 0.30757 | 1.335 | 2.548 | 101.2 | 20.79 | 17.701 | 83.4667 | 102.2546 | 80.9443 | 0.004299 | 0.99704 | 0.024378 | 24.4503 | 0.034275 | 0.15841 | 0.003787 | 0.27029 | 0.27029 | 0.024381
2019-03-19 15:59 | 1553025561.7728 | 1032 | 0.2087 | 0.000553 | 0.000132 | 97.2668 | 96.9195 | 121.703 | 21.2587 | 0.002599 | 0.00244 | 220.5465 | 0.26616 | 8.4035 | 0.00337 | 0.005214 | 0.005129 | 6.4e-05 | 0.00101 | 0.30646 | 1.3394 | 2.548 | 101.2 | 20.79 | 17.703 | 84.3471 | 102.9375 | 80.5283 | 0.004299 | 0.99821 | 0.024403 | 24.4466 | 0.031992 | 0.15834 | 0.004051 | 0.29018 | 0.29018 | 0.022699
2019-03-19 15:59 | 1553025562.7725 | 1033 | 0.20874 | 0.000558 | 0.00013 | 97.0609 | 96.6157 | 121.7685 | 21.2587 | 0.002628 | 0.002455 | 220.5254 | 0.26588 | 8.4009 | 0.003298 | 0.004382 | 0.004214 | 5.9e-05 | 0.001006 | 0.31094 | 1.339 | 2.5481 | 101.2 | 20.79 | 17.707 | 85.0323 | 101.1931 | 80.8861 | 0.0043 | 0.9972 | 0.024397 | 24.4655 | 0.04112 | 0.15817 | 0.003787 | 0.27078 | 0.27078 | 0.0243
2019-03-19 15:59 | 1553025563.7715 | 1034 | 0.20872 | 0.000552 | 0.000127 | 96.9025 | 96.4078 | 121.6441 | 21.2592 | 0.00261 | 0.00247 | 220.5421 | 0.26559 | 8.4027 | 0.003279 | 0.005361 | 0.005226 | 6.2e-05 | 0.001011 | 0.31614 | 1.3238 | 2.5482 | 101.2 | 20.79 | 17.71 | 84.6166 | 100.7462 | 81.2462 | 0.004301 | 0.99619 | 0.024372 | 24.4654 | 0.045773 | 0.158 | 0.003953 | 0.28288 | 0.28288 | 0.023236
2019-03-19 15:59 | 1553025564.772 | 1035 | 0.20874 | 0.000556 | 0.000124 | 96.681 | 96.041 | 121.8767 | 21.2582 | 0.002613 | 0.002435 | 220.5337 | 0.26598 | 8.4004 | 0.003334 | 0.005107 | 0.005098 | 5.8e-05 | 0.001009 | 0.31813 | 1.348 | 2.548 | 101.2 | 20.79 | 17.714 | 84.6782 | 102.0619 | 80.3913 | 0.004302 | 0.9986 | 0.024425 | 24.4593 | 0.059403 | 0.15823 | 0.003854 | 0.27638 | 0.27638 | 0.023816
2019-03-19 15:59 | 1553025565.773 | 1036 | 0.20876 | 0.000556 | 0.000123 | 96.425 | 95.7513 | 121.8697 | 21.2579 | 0.002624 | 0.002491 | 220.5275 | 0.26576 | 8.403 | 0.00338 | 0.004571 | 0.004493 | 5.6e-05 | 0.001014 | 0.27151 | 1.3564 | 2.5481 | 101.2 | 20.8 | 17.717 | 84.9362 | 103.1669 | 81.7379 | 0.004306 | 0.99481 | 0.024378 | 24.5051 | 0.062715 | 0.1581 | 0.003719 | 0.26601 | 0.26601 | 0.024724
2019-03-19 15:59 | 1553025566.7727 | 1037 | 0.20876 | 0.00055 | 0.000122 | 96.1975 | 95.6944 | 122.0786 | 21.2582 | 0.002565 | 0.002458 | 220.5311 | 0.26588 | 8.4004 | 0.003333 | 0.0048 | 0.004707 | 6e-05 | 0.001015 | 0.3131 | 1.352 | 2.5481 | 101.2 | 20.79 | 17.72 | 83.5304 | 102.0438 | 80.9506 | 0.004304 | 0.99702 | 0.024426 | 24.4989 | 0.046901 | 0.15817 | 0.003695 | 0.26483 | 0.26483 | 0.024846
2019-03-19 15:59 | 1553025567.7718 | 1038 | 0.20878 | 0.000549 | 0.000122 | 95.9938 | 95.3889 | 122.279 | 21.2581 | 0.002683 | 0.0025 | 220.5313 | 0.26558 | 8.4002 | 0.003295 | 0.004848 | 0.004686 | 6e-05 | 0.001019 | 0.31868 | 1.337 | 2.5481 | 101.2 | 20.79 | 17.722 | 86.349 | 101.1226 | 81.9588 | 0.004304 | 0.99419 | 0.024411 | 24.5538 | 0.056535 | 0.15799 | 0.003588 | 0.25678 | 0.25678 | 0.025596
2019-03-19 15:59 | 1553025568.7722 | 1039 | 0.20878 | 0.000548 | 0.00012 | 95.7321 | 95.0615 | 122.4267 | 21.2579 | 0.002659 | 0.00247 | 220.5243 | 0.2652 | 8.4023 | 0.003312 | 0.004936 | 0.004849 | 5.8e-05 | 0.001023 | 0.29574 | 1.331 | 2.5481 | 101.2 | 20.79 | 17.721 | 85.7837 | 101.5372 | 81.227 | 0.004304 | 0.99624 | 0.024451 | 24.5433 | 0.062873 | 0.15777 | 0.003594 | 0.25774 | 0.25774 | 0.025464
2019-03-19 15:59 | 1553025569.7719 | 1040 | 0.20872 | 0.000549 | 0.00012 | 95.4992 | 94.9163 | 122.2152 | 21.2585 | 0.002616 | 0.002459 | 220.5165 | 0.26577 | 8.4019 | 0.003258 | 0.004594 | 0.004416 | 5.9e-05 | 0.001025 | 0.32139 | 1.324 | 2.5482 | 101.2 | 20.79 | 17.716 | 84.7596 | 100.2324 | 80.985 | 0.004303 | 0.99692 | 0.024438 | 24.5138 | 0.054762 | 0.15811 | 0.003942 | 0.28327 | 0.28327 | 0.023219
2019-03-19 15:59 | 1553025570.7719 | 1041 | 0.20874 | 0.00055 | 0.000119 | 95.3095 | 94.595 | 121.9389 | 21.2582 | 0.00264 | 0.002437 | 220.5086 | 0.26528 | 8.4018 | 0.003318 | 0.005081 | 0.004989 | 5.7e-05 | 0.001022 | 0.30198 | 1.31 | 2.5481 | 101.2 | 20.8 | 17.712 | 85.3315 | 101.6725 | 80.4444 | 0.004304 | 0.99845 | 0.024429 | 24.4673 | 0.067305 | 0.15781 | 0.003842 | 0.27585 | 0.27585 | 0.023799
2019-03-19 15:59 | 1553025571.772 | 1042 | 0.20876 | 0.000549 | 0.000117 | 95.0502 | 94.2532 | 122.0375 | 21.2581 | 0.002679 | 0.002439 | 220.5034 | 0.26507 | 8.4026 | 0.003223 | 0.004804 | 0.004788 | 5.6e-05 | 0.001022 | 0.30558 | 1.2961 | 2.5481 | 101.2 | 20.8 | 17.707 | 86.2609 | 99.3902 | 80.4895 | 0.004303 | 0.99832 | 0.024438 | 24.4788 | 0.075317 | 0.15769 | 0.003742 | 0.26872 | 0.26872 | 0.024412
2019-03-19 16:00 | 1553025572.7733 | 1043 | 0.20872 | 0.00055 | 0.000117 | 94.8312 | 94.2875 | 122.084 | 21.2575 | 0.002584 | 0.002393 | 220.4983 | 0.26595 | 8.3997 | 0.003319 | 0.004625 | 0.004518 | 5.7e-05 | 0.001023 | 0.3178 | 1.3718 | 2.5481 | 101.2 | 20.8 | 17.698 | 83.9799 | 101.6914 | 79.3755 | 0.004301 | 1.0015 | 0.024481 | 24.4449 | 0.051428 | 0.15822 | 0.003985 | 0.28712 | 0.28712 | 0.022923
2019-03-19 16:00 | 1553025573.7721 | 1044 | 0.20875 | 0.000552 | 0.000116 | 94.5851 | 93.8512 | 122.3854 | 21.2575 | 0.002622 | 0.00243 | 220.4573 | 0.26544 | 8.4025 | 0.003364 | 0.005572 | 0.005491 | 5.6e-05 | 0.001026 | 0.30215 | 1.3572 | 2.5481 | 101.2 | 20.8 | 17.695 | 84.8914 | 102.7823 | 80.2652 | 0.0043 | 0.99895 | 0.02448 | 24.5059 | 0.069675 | 0.15791 | 0.003804 | 0.27383 | 0.27383 | 0.023989
2019-03-19 16:00 | 1553025574.7744 | 1045 | 0.20874 | 0.000552 | 0.000116 | 94.4086 | 93.4275 | 122.3675 | 21.2571 | 0.002586 | 0.002436 | 220.4353 | 0.26517 | 8.4033 | 0.003357 | 0.004884 | 0.004813 | 5.7e-05 | 0.001027 | 0.29574 | 1.3978 | 2.5481 | 101.2 | 20.8 | 17.692 | 84.0264 | 102.6233 | 80.4107 | 0.004299 | 0.99854 | 0.024473 | 24.5091 | 0.093441 | 0.15775 | 0.003814 | 0.27446 | 0.27446 | 0.02391
2019-03-19 16:00 | 1553025575.7731 | 1046 | 0.20877 | 0.000551 | 0.000115 | 94.2352 | 93.348 | 122.6011 | 21.2568 | 0.002545 | 0.002421 | 220.4291 | 0.26632 | 8.4008 | 0.003312 | 0.004663 | 0.004472 | 5.7e-05 | 0.001029 | 0.30414 | 1.3278 | 2.548 | 101.2 | 20.8 | 17.692 | 83.0331 | 101.5344 | 80.0485 | 0.004299 | 0.99957 | 0.024509 | 24.5199 | 0.084602 | 0.15843 | 0.003635 | 0.26169 | 0.26169 | 0.025186
2019-03-19 16:00 | 1553025576.7726 | 1047 | 0.20877 | 0.000556 | 0.000114 | 93.9119 | 93.0336 | 122.5123 | 21.2562 | 0.002554 | 0.002446 | 220.4191 | 0.26562 | 8.4035 | 0.003213 | 0.004704 | 0.004518 | 5.7e-05 | 0.001023 | 0.3126 | 1.2967 | 2.548 | 101.2 | 20.8 | 17.695 | 83.2657 | 99.1396 | 80.6607 | 0.0043 | 0.99784 | 0.024479 | 24.5323 | 0.084048 | 0.15801 | 0.003659 | 0.26323 | 0.26323 | 0.024972
2019-03-19 16:00 | 1553025577.7733 | 1048 | 0.20873 | 0.000551 | 0.000112 | 93.7064 | 92.7004 | 122.3715 | 21.2567 | 0.00259 | 0.002432 | 220.4088 | 0.26549 | 8.4033 | 0.003238 | 0.004789 | 0.004766 | 5.5e-05 | 0.001024 | 0.30541 | 1.3244 | 2.5483 | 101.2 | 20.8 | 17.702 | 84.1165 | 99.7433 | 80.3238 | 0.004302 | 0.99879 | 0.024477 | 24.5065 | 0.096544 | 0.15794 | 0.003914 | 0.28212 | 0.28212 | 0.023289
2019-03-19 16:00 | 1553025578.7734 | 1049 | 0.2088 | 0.000552 | 0.000112 | 93.4498 | 92.5692 | 122.0563 | 21.2561 | 0.002572 | 0.002419 | 220.4141 | 0.26609 | 8.4047 | 0.003268 | 0.00512 | 0.004994 | 5.4e-05 | 0.001028 | 0.31797 | 1.3676 | 2.5481 | 101.2 | 20.8 | 17.707 | 83.6883 | 100.474 | 80.0197 | 0.004303 | 0.99965 | 0.024456 | 24.4644 | 0.084684 | 0.1583 | 0.003465 | 0.24878 | 0.24878 | 0.02647
2019-03-19 16:00 | 1553025579.7727 | 1050 | 0.20881 | 0.00055 | 0.000112 | 93.2238 | 92.3048 | 122.0384 | 21.2559 | 0.002589 | 0.002471 | 220.4014 | 0.26612 | 8.402 | 0.003309 | 0.004816 | 0.0048 | 5.7e-05 | 0.001028 | 0.29999 | 1.3214 | 2.5481 | 101.2 | 20.8 | 17.714 | 84.0854 | 101.4488 | 81.2687 | 0.004305 | 0.99613 | 0.024411 | 24.5058 | 0.088613 | 0.15832 | 0.003437 | 0.24627 | 0.24627 | 0.026743
2019-03-19 16:00 | 1553025580.7724 | 1051 | 0.20878 | 0.000555 | 0.000112 | 92.9003 | 91.9099 | 122.0999 | 21.2562 | 0.002604 | 0.002411 | 220.4007 | 0.26511 | 8.4033 | 0.003199 | 0.005524 | 0.005323 | 5.8e-05 | 0.001031 | 0.31581 | 1.333 | 2.5481 | 101.2 | 20.8 | 17.731 | 84.4694 | 98.808 | 79.8179 | 0.004309 | 1.0002 | 0.024467 | 24.4618 | 0.095867 | 0.15772 | 0.003577 | 0.25713 | 0.25713 | 0.025516
2019-03-19 16:00 | 1553025581.7725 | 1052 | 0.20877 | 0.000554 | 0.000113 | 92.7418 | 91.5822 | 121.9967 | 21.2556 | 0.002579 | 0.002453 | 220.3913 | 0.26556 | 8.4023 | 0.003237 | 0.005146 | 0.005126 | 5.5e-05 | 0.00103 | 0.3131 | 1.3328 | 2.5481 | 101.2 | 20.8 | 17.742 | 83.8518 | 99.7069 | 80.8233 | 0.004312 | 0.99738 | 0.024422 | 24.4862 | 0.11255 | 0.15798 | 0.00365 | 0.26208 | 0.26208 | 0.025077
2019-03-19 16:00 | 1553025582.7726 | 1053 | 0.20878 | 0.000549 | 0.000113 | 92.4471 | 91.3431 | 121.9063 | 21.256 | 0.002566 | 0.002444 | 220.3813 | 0.26581 | 8.4006 | 0.003207 | 0.004505 | 0.004506 | 5.7e-05 | 0.001029 | 0.30525 | 1.3297 | 2.5481 | 101.2 | 20.8 | 17.767 | 83.5439 | 98.9888 | 80.6028 | 0.004318 | 0.998 | 0.024421 | 24.4695 | 0.10746 | 0.15813 | 0.003591 | 0.25768 | 0.25768 | 0.025528
2019-03-19 16:00 | 1553025583.7725 | 1054 | 0.20876 | 0.000552 | 0.000112 | 92.1053 | 90.9825 | 121.8084 | 21.2561 | 0.002565 | 0.002403 | 220.3765 | 0.26596 | 8.4024 | 0.003285 | 0.004363 | 0.004214 | 5.4e-05 | 0.001025 | 0.28982 | 1.3498 | 2.5481 | 101.2 | 20.8 | 17.781 | 83.5176 | 100.8807 | 79.6292 | 0.004321 | 1.0008 | 0.024444 | 24.426 | 0.10971 | 0.15822 | 0.0037 | 0.26601 | 0.26601 | 0.024743
2019-03-19 16:00 | 1553025584.7726 | 1055 | 0.20877 | 0.000553 | 0.000112 | 91.8739 | 90.7167 | 121.6396 | 21.2557 | 0.002605 | 0.002409 | 220.3533 | 0.26589 | 8.4022 | 0.003191 | 0.004574 | 0.004475 | 5.4e-05 | 0.001022 | 0.3194 | 1.3481 | 2.548 | 101.2 | 20.8 | 17.794 | 84.4759 | 98.5989 | 79.7808 | 0.004324 | 1.0003 | 0.024422 | 24.4143 | 0.11338 | 0.15818 | 0.00364 | 0.26136 | 0.26136 | 0.025177
2019-03-19 16:00 | 1553025585.7718 | 1056 | 0.20877 | 0.000554 | 0.000111 | 91.6141 | 90.431 | 121.3589 | 21.2556 | 0.002607 | 0.002436 | 220.3559 | 0.26629 | 8.4033 | 0.003222 | 0.005557 | 0.005363 | 5.7e-05 | 0.001018 | 0.3183 | 1.3557 | 2.5481 | 101.2 | 20.8 | 17.821 | 84.5196 | 99.3641 | 80.4263 | 0.004331 | 0.9985 | 0.024372 | 24.4084 | 0.11625 | 0.15841 | 0.003698 | 0.26506 | 0.26506 | 0.024862
2019-03-19 16:00 | 1553025586.7727 | 1057 | 0.20877 | 0.000554 | 0.000111 | 91.3026 | 90.0663 | 121.2197 | 21.2553 | 0.002594 | 0.002464 | 220.3523 | 0.2658 | 8.4022 | 0.003295 | 0.00468 | 0.004641 | 5.6e-05 | 0.001013 | 0.28512 | 1.366 | 2.548 | 101.2 | 20.8 | 17.833 | 84.2252 | 101.1223 | 81.0989 | 0.004334 | 0.9966 | 0.024335 | 24.4176 | 0.12194 | 0.15812 | 0.003689 | 0.26407 | 0.26407 | 0.02491
2019-03-19 16:00 | 1553025587.7728 | 1058 | 0.20876 | 0.000551 | 0.00011 | 90.9904 | 89.7406 | 121.4303 | 21.2546 | 0.002531 | 0.002449 | 220.3401 | 0.26543 | 8.4026 | 0.003187 | 0.004721 | 0.00458 | 5.2e-05 | 0.001009 | 0.30215 | 1.3471 | 2.5481 | 101.2 | 20.8 | 17.844 | 82.6932 | 98.514 | 80.7387 | 0.004336 | 0.99762 | 0.024368 | 24.4264 | 0.12372 | 0.1579 | 0.003737 | 0.26797 | 0.26797 | 0.024513
2019-03-19 16:00 | 1553025588.7729 | 1059 | 0.20875 | 0.000556 | 0.00011 | 90.7975 | 89.2351 | 121.4502 | 21.2546 | 0.002614 | 0.002413 | 220.3256 | 0.26633 | 8.4025 | 0.003215 | 0.004567 | 0.00442 | 5.2e-05 | 0.001014 | 0.29551 | 1.3608 | 2.5481 | 101.2 | 20.8 | 17.863 | 84.6974 | 99.1845 | 79.8757 | 0.004341 | 1.0001 | 0.0244 | 24.3986 | 0.15526 | 0.15844 | 0.003795 | 0.27263 | 0.27263 | 0.024176
2019-03-19 16:00 | 1553025589.7719 | 1060 | 0.20877 | 0.000557 | 0.000109 | 90.5751 | 89.2385 | 121.8455 | 21.2539 | 0.002581 | 0.002453 | 220.2869 | 0.26613 | 8.4057 | 0.003167 | 0.00482 | 0.004637 | 5.6e-05 | 0.001013 | 0.32388 | 1.3596 | 2.548 | 101.2 | 20.8 | 17.872 | 83.903 | 98.0295 | 80.8192 | 0.004343 | 0.99739 | 0.024407 | 24.4709 | 0.13298 | 0.15832 | 0.003669 | 0.26346 | 0.26346 | 0.024999
2019-03-19 16:00 | 1553025590.773 | 1061 | 0.20877 | 0.000552 | 0.000108 | 90.5658 | 89.2286 | 121.7983 | 21.2548 | 0.002627 | 0.002419 | 220.2515 | 0.26573 | 8.4046 | 0.003142 | 0.004965 | 0.004878 | 5.7e-05 | 0.001013 | 0.29214 | 1.3528 | 2.5479 | 101.2 | 20.8 | 17.884 | 85.0049 | 97.4214 | 80.0162 | 0.004346 | 0.99966 | 0.02443 | 24.4384 | 0.13306 | 0.15808 | 0.003673 | 0.26411 | 0.26411 | 0.024899
2019-03-19 16:00 | 1553025591.772 | 1062 | 0.20877 | 0.000556 | 0.000107 | 90.5262 | 89.2423 | 121.7371 | 21.2534 | 0.002592 | 0.002428 | 220.2483 | 0.26511 | 8.4017 | 0.003257 | 0.005181 | 0.005122 | 5.1e-05 | 0.001014 | 0.28208 | 1.3923 | 2.548 | 101.2 | 20.8 | 17.889 | 84.1601 | 100.1881 | 80.2268 | 0.004347 | 0.99906 | 0.024417 | 24.4396 | 0.12777 | 0.15772 | 0.003691 | 0.26534 | 0.26534 | 0.024727
2019-03-19 16:00 | 1553025592.7721 | 1063 | 0.20878 | 0.000557 | 0.000105 | 90.6132 | 89.2498 | 121.8281 | 21.2541 | 0.002619 | 0.002413 | 220.2513 | 0.2659 | 8.4041 | 0.003368 | 0.004639 | 0.004484 | 5.1e-05 | 0.001015 | 0.30807 | 1.3929 | 2.5481 | 101.2 | 20.8 | 17.893 | 84.8202 | 102.8937 | 79.8759 | 0.004348 | 1.0001 | 0.024438 | 24.4366 | 0.13561 | 0.15818 | 0.00361 | 0.25971 | 0.25971 | 0.025337
2019-03-19 16:00 | 1553025593.7731 | 1064 | 0.2088 | 0.000559 | 0.000105 | 90.7858 | 89.4556 | 121.7716 | 21.2537 | 0.002608 | 0.002469 | 220.2575 | 0.26597 | 8.401 | 0.003365 | 0.004524 | 0.004374 | 5.3e-05 | 0.001018 | 0.31669 | 1.3731 | 2.548 | 101.2 | 20.8 | 17.892 | 84.5435 | 102.7999 | 81.2159 | 0.004348 | 0.99627 | 0.024386 | 24.4772 | 0.13202 | 0.15823 | 0.003507 | 0.2516 | 0.2516 | 0.026162
2019-03-19 16:00 | 1553025594.7721 | 1065 | 0.20877 | 0.000557 | 0.000102 | 91.0049 | 89.77 | 121.8171 | 21.2534 | 0.002563 | 0.002376 | 220.2619 | 0.26556 | 8.4022 | 0.003336 | 0.005044 | 0.004911 | 5.4e-05 | 0.001019 | 0.26792 | 1.3362 | 2.5481 | 101.2 | 20.8 | 17.889 | 83.458 | 102.1178 | 78.9645 | 0.004347 | 1.0026 | 0.024468 | 24.4039 | 0.1222 | 0.15798 | 0.003694 | 0.26642 | 0.26642 | 0.024669
2019-03-19 16:00 | 1553025595.7712 | 1066 | 0.20879 | 0.000561 | 0.000101 | 91.1303 | 90.0122 | 121.968 | 21.253 | 0.002676 | 0.002453 | 220.235 | 0.26568 | 8.403 | 0.003312 | 0.004528 | 0.004395 | 5.5e-05 | 0.00102 | 0.31581 | 1.3365 | 2.5482 | 101.2 | 20.8 | 17.882 | 86.1973 | 101.5292 | 80.833 | 0.004346 | 0.99735 | 0.024419 | 24.4837 | 0.11042 | 0.15805 | 0.003525 | 0.25353 | 0.25353 | 0.025934
2019-03-19 16:00 | 1553025596.7726 | 1067 | 0.20878 | 0.000556 | 0.0001 | 91.4367 | 90.1183 | 122.1481 | 21.2534 | 0.00254 | 0.002451 | 220.2321 | 0.26566 | 8.4019 | 0.003377 | 0.004944 | 0.004906 | 5.4e-05 | 0.001022 | 0.29806 | 1.3219 | 2.5481 | 101.2 | 20.8 | 17.878 | 82.9151 | 103.1129 | 80.7868 | 0.004345 | 0.99748 | 0.024438 | 24.5002 | 0.1299 | 0.15804 | 0.003619 | 0.26068 | 0.26068 | 0.02522
2019-03-19 16:00 | 1553025597.7724 | 1068 | 0.20878 | 0.000556 | 0.0001 | 91.7131 | 90.531 | 122.0476 | 21.2528 | 0.002634 | 0.002431 | 220.2253 | 0.26632 | 8.4036 | 0.003333 | 0.004326 | 0.00419 | 5.4e-05 | 0.001023 | 0.30359 | 1.2968 | 2.548 | 101.2 | 20.8 | 17.873 | 85.1702 | 102.0259 | 80.2964 | 0.004343 | 0.99887 | 0.024445 | 24.4731 | 0.11603 | 0.15843 | 0.003613 | 0.26032 | 0.26032 | 0.025318
2019-03-19 16:00 | 1553025598.7907 | 1069 | 0.20875 | 0.000554 | 9.9e-05 | 91.8769 | 90.7362 | 122.0154 | 21.2525 | 0.002538 | 0.002408 | 220.2001 | 0.26569 | 8.402 | 0.003282 | 0.004456 | 0.004379 | 5.4e-05 | 0.001024 | 0.30989 | 1.2978 | 2.5481 | 101.2 | 20.8 | 17.865 | 82.858 | 100.8128 | 79.7458 | 0.004342 | 1.0004 | 0.024461 | 24.4508 | 0.11175 | 0.15806 | 0.003823 | 0.27602 | 0.27602 | 0.023822
2019-03-19 16:00 | 1553025599.7805 | 1070 | 0.20875 | 0.000555 | 9.9e-05 | 92.0675 | 90.7722 | 121.8723 | 21.2527 | 0.002643 | 0.002421 | 220.2108 | 0.26495 | 8.4005 | 0.003273 | 0.004498 | 0.004399 | 5.2e-05 | 0.001026 | 0.27649 | 1.3168 | 2.5481 | 101.2 | 20.8 | 17.86 | 85.386 | 100.59 | 80.0445 | 0.00434 | 0.99958 | 0.024436 | 24.4468 | 0.12674 | 0.15762 | 0.003777 | 0.2724 | 0.2724 | 0.024071
2019-03-19 16:00 | 1553025600.7805 | 1071 | 0.20878 | 0.000561 | 9.9e-05 | 92.3703 | 91.2534 | 121.9922 | 21.2525 | 0.002637 | 0.00243 | 220.2436 | 0.26552 | 8.3999 | 0.003233 | 0.005062 | 0.004984 | 5.3e-05 | 0.001027 | 0.32997 | 1.3304 | 2.5481 | 101.2 | 20.8 | 17.854 | 85.2377 | 99.6277 | 80.2704 | 0.004339 | 0.99894 | 0.024441 | 24.4667 | 0.10882 | 0.15796 | 0.003629 | 0.26152 | 0.26152 | 0.025126
2019-03-19 16:00 | 1553025601.7812 | 1072 | 0.20877 | 0.000557 | 9.9e-05 | 92.4961 | 91.5116 | 121.9023 | 21.2519 | 0.002584 | 0.00243 | 220.235 | 0.26655 | 8.4021 | 0.003304 | 0.004944 | 0.004845 | 5.5e-05 | 0.001032 | 0.32803 | 1.3468 | 2.548 | 101.2 | 20.8 | 17.854 | 83.9629 | 101.3309 | 80.2729 | 0.004339 | 0.99893 | 0.024432 | 24.4577 | 0.09571 | 0.15857 | 0.003665 | 0.26405 | 0.26405 | 0.024982
2019-03-19 16:00 | 1553025602.7813 | 1073 | 0.20879 | 0.000565 | 9.8e-05 | 92.7088 | 91.7767 | 121.722 | 21.2514 | 0.002549 | 0.002441 | 220.1993 | 0.26575 | 8.4012 | 0.003246 | 0.004965 | 0.004846 | 5.3e-05 | 0.001033 | 0.31492 | 1.3199 | 2.548 | 101.2 | 20.8 | 17.853 | 83.119 | 99.9195 | 80.5275 | 0.004339 | 0.99821 | 0.024405 | 24.4484 | 0.090389 | 0.1581 | 0.003572 | 0.25701 | 0.25701 | 0.025589
2019-03-19 16:00 | 1553025603.7807 | 1074 | 0.20877 | 0.000564 | 9.9e-05 | 92.9621 | 92.0157 | 121.5622 | 21.2505 | 0.002555 | 0.00246 | 220.1766 | 0.26615 | 8.4021 | 0.003199 | 0.004797 | 0.004626 | 5.5e-05 | 0.001033 | 0.30596 | 1.339 | 2.5481 | 101.2 | 20.8 | 17.855 | 83.2782 | 98.7954 | 80.9991 | 0.004339 | 0.99688 | 0.024372 | 24.4487 | 0.091524 | 0.15833 | 0.00368 | 0.26456 | 0.26456 | 0.024897
2019-03-19 16:00 | 1553025604.7805 | 1075 | 0.20877 | 0.000564 | 9.8e-05 | 93.1716 | 92.3339 | 121.5882 | 21.2509 | 0.002584 | 0.002484 | 220.1532 | 0.2659 | 8.4024 | 0.003255 | 0.004655 | 0.004624 | 5.2e-05 | 0.001035 | 0.29358 | 1.3552 | 2.5481 | 101.2 | 20.8 | 17.858 | 83.9728 | 100.1487 | 81.5609 | 0.00434 | 0.9953 | 0.024356 | 24.4707 | 0.080781 | 0.15818 | 0.003653 | 0.26246 | 0.26246 | 0.025073
2019-03-19 16:00 | 1553025605.7815 | 1076 | 0.20874 | 0.000567 | 9.7e-05 | 93.3982 | 92.6086 | 121.9093 | 21.2504 | 0.002638 | 0.002463 | 220.1464 | 0.26613 | 8.403 | 0.003237 | 0.005074 | 0.004991 | 5.4e-05 | 0.001029 | 0.28279 | 1.3875 | 2.5481 | 101.2 | 20.8 | 17.862 | 85.2816 | 99.7099 | 81.0647 | 0.004341 | 0.9967 | 0.024405 | 24.4858 | 0.075937 | 0.15832 | 0.003833 | 0.2763 | 0.2763 | 0.023837
2019-03-19 16:00 | 1553025606.7805 | 1077 | 0.20876 | 0.000568 | 9.7e-05 | 93.6584 | 92.5829 | 121.8279 | 21.2495 | 0.002609 | 0.002435 | 220.1183 | 0.26565 | 8.4035 | 0.003166 | 0.004645 | 0.004482 | 5.3e-05 | 0.001027 | 0.31349 | 1.3731 | 2.548 | 101.2 | 20.8 | 17.872 | 84.5645 | 97.99 | 80.386 | 0.004343 | 0.99861 | 0.02442 | 24.4542 | 0.1033 | 0.15804 | 0.003757 | 0.27083 | 0.27083 | 0.024275
2019-03-19 16:00 | 1553025607.7809 | 1078 | 0.20876 | 0.000559 | 9.5e-05 | 93.924 | 92.9333 | 121.7908 | 21.2492 | 0.002613 | 0.002427 | 220.119 | 0.26452 | 8.4048 | 0.003239 | 0.004662 | 0.004533 | 5.2e-05 | 0.001018 | 0.31669 | 1.3485 | 2.5481 | 101.2 | 20.8 | 17.877 | 84.6562 | 99.7485 | 80.1946 | 0.004344 | 0.99915 | 0.024423 | 24.4438 | 0.094845 | 0.15736 | 0.003735 | 0.26938 | 0.26938 | 0.024301
2019-03-19 16:00 | 1553025608.7806 | 1079 | 0.20879 | 0.000566 | 9.5e-05 | 94.1325 | 93.1501 | 121.7605 | 21.2504 | 0.002609 | 0.002392 | 220.1217 | 0.26506 | 8.4022 | 0.003218 | 0.004794 | 0.004699 | 5.1e-05 | 0.001015 | 0.33323 | 1.3367 | 2.5481 | 101.2 | 20.8 | 17.888 | 84.5833 | 99.2415 | 79.3454 | 0.004347 | 1.0016 | 0.024449 | 24.4114 | 0.093843 | 0.15769 | 0.003551 | 0.25612 | 0.25612 | 0.025612
2019-03-19 16:00 | 1553025609.7807 | 1080 | 0.20871 | 0.000567 | 9.5e-05 | 94.3978 | 93.5559 | 121.8936 | 21.2494 | 0.002589 | 0.00236 | 220.1234 | 0.26579 | 8.4039 | 0.003158 | 0.005454 | 0.005289 | 5.2e-05 | 0.001014 | 0.32261 | 1.3226 | 2.5481 | 101.2 | 20.8 | 17.888 | 84.097 | 97.8096 | 78.574 | 0.004347 | 1.0038 | 0.02449 | 24.398 | 0.08013 | 0.15812 | 0.004036 | 0.29241 | 0.29241 | 0.022495
2019-03-19 16:00 | 1553025610.7812 | 1081 | 0.20875 | 0.000569 | 9.6e-05 | 94.5983 | 93.8035 | 122.1192 | 21.2484 | 0.002588 | 0.002423 | 220.0874 | 0.26588 | 8.4048 | 0.003168 | 0.004706 | 0.004769 | 5e-05 | 0.001018 | 0.29966 | 1.3052 | 2.548 | 101.2 | 20.8 | 17.895 | 84.067 | 98.0413 | 80.1003 | 0.004349 | 0.99942 | 0.024459 | 24.4735 | 0.075462 | 0.15817 | 0.003786 | 0.2735 | 0.2735 | 0.024058
2019-03-19 16:00 | 1553025611.7811 | 1082 | 0.20875 | 0.000562 | 9.5e-05 | 94.8927 | 93.9683 | 122.106 | 21.2479 | 0.002559 | 0.002428 | 220.0911 | 0.2667 | 8.4025 | 0.003172 | 0.004412 | 0.004324 | 5.4e-05 | 0.00102 | 0.30574 | 1.3378 | 2.5481 | 101.2 | 20.8 | 17.897 | 83.3692 | 98.1397 | 80.2246 | 0.004349 | 0.99907 | 0.024454 | 24.4765 | 0.087561 | 0.15866 | 0.003804 | 0.27478 | 0.27478 | 0.02402
2019-03-19 16:00 | 1553025612.7799 | 1083 | 0.20867 | 0.000569 | 9.5e-05 | 95.0793 | 94.3841 | 122.1077 | 21.2491 | 0.002557 | 0.002442 | 220.1009 | 0.26577 | 8.4023 | 0.003134 | 0.004434 | 0.004376 | 5e-05 | 0.001027 | 0.29142 | 1.2996 | 2.5481 | 101.2 | 20.8 | 17.898 | 83.3089 | 97.2222 | 80.5635 | 0.00435 | 0.99811 | 0.024442 | 24.4884 | 0.065648 | 0.15811 | 0.004285 | 0.31011 | 0.31011 | 0.02121
2019-03-19 16:00 | 1553025613.7798 | 1084 | 0.20878 | 0.000563 | 9.6e-05 | 95.194 | 94.5214 | 122.2218 | 21.2478 | 0.00262 | 0.002463 | 220.1065 | 0.26606 | 8.4044 | 0.003175 | 0.00448 | 0.004385 | 5.2e-05 | 0.001031 | 0.29446 | 1.3282 | 2.5481 | 101.2 | 20.8 | 17.897 | 84.8343 | 98.2179 | 81.0604 | 0.004349 | 0.99671 | 0.024436 | 24.517 | 0.063421 | 0.15828 | 0.00364 | 0.26251 | 0.26251 | 0.025082
2019-03-19 16:00 | 1553025614.7801 | 1085 | 0.20872 | 0.000565 | 9.6e-05 | 95.3053 | 94.7111 | 122.4168 | 21.2481 | 0.002573 | 0.002397 | 220.1038 | 0.26604 | 8.4028 | 0.003182 | 0.004467 | 0.004375 | 5.1e-05 | 0.001031 | 0.31277 | 1.3226 | 2.5482 | 101.2 | 20.8 | 17.894 | 83.6955 | 98.3876 | 79.4825 | 0.004349 | 1.0012 | 0.024511 | 24.4819 | 0.055939 | 0.15827 | 0.003996 | 0.28957 | 0.28957 | 0.022737
2019-03-19 16:00 | 1553025615.781 | 1086 | 0.2088 | 0.000562 | 9.7e-05 | 95.5172 | 94.8139 | 121.9002 | 21.2472 | 0.002554 | 0.00244 | 220.1258 | 0.26589 | 8.4014 | 0.003202 | 0.004672 | 0.00459 | 5e-05 | 0.00103 | 0.31598 | 1.3414 | 2.5481 | 101.2 | 20.8 | 17.887 | 83.25 | 98.8736 | 80.5001 | 0.004347 | 0.99829 | 0.024424 | 24.4654 | 0.066103 | 0.15818 | 0.003472 | 0.24989 | 0.24989 | 0.026333
2019-03-19 16:00 | 1553025616.7801 | 1087 | 0.20875 | 0.000566 | 9.8e-05 | 95.6342 | 95.118 | 121.842 | 21.2478 | 0.002591 | 0.002409 | 220.127 | 0.26577 | 8.4026 | 0.003121 | 0.004513 | 0.004438 | 4.9e-05 | 0.001036 | 0.33627 | 1.3373 | 2.5481 | 101.2 | 20.8 | 17.883 | 84.1313 | 96.9084 | 79.7575 | 0.004346 | 1.0004 | 0.024443 | 24.4338 | 0.048411 | 0.15811 | 0.003762 | 0.27145 | 0.27145 | 0.02423
2019-03-19 16:00 | 1553025617.7802 | 1088 | 0.20871 | 0.000565 | 0.0001 | 95.7411 | 95.3348 | 121.8955 | 21.2485 | 0.002568 | 0.00242 | 220.0992 | 0.26582 | 8.4039 | 0.003196 | 0.004834 | 0.004712 | 5e-05 | 0.001036 | 0.28191 | 1.3433 | 2.5482 | 101.2 | 20.8 | 17.881 | 83.5756 | 98.7149 | 80.0296 | 0.004345 | 0.99962 | 0.024439 | 24.4486 | 0.038035 | 0.15814 | 0.004026 | 0.2907 | 0.2907 | 0.02263
2019-03-19 16:00 | 1553025618.7815 | 1089 | 0.20873 | 0.000568 | 0.000101 | 95.8503 | 95.1249 | 122.2746 | 21.2473 | 0.002606 | 0.002419 | 220.0808 | 0.26585 | 8.4018 | 0.003192 | 0.004599 | 0.004473 | 5.1e-05 | 0.001028 | 0.28783 | 1.3668 | 2.5481 | 101.2 | 20.8 | 17.878 | 84.4921 | 98.6181 | 79.9957 | 0.004345 | 0.99972 | 0.024478 | 24.4854 | 0.067954 | 0.15815 | 0.003921 | 0.28337 | 0.28337 | 0.023217
2019-03-19 16:00 | 1553025619.7812 | 1090 | 0.20871 | 0.000574 | 0.000103 | 95.8998 | 95.4807 | 122.0786 | 21.2466 | 0.002593 | 0.002396 | 220.0691 | 0.26578 | 8.4026 | 0.003167 | 0.004648 | 0.004495 | 5.1e-05 | 0.001032 | 0.31813 | 1.3467 | 2.5481 | 101.2 | 20.8 | 17.878 | 84.1893 | 98.0169 | 79.4406 | 0.004345 | 1.0013 | 0.024478 | 24.4466 | 0.039176 | 0.15812 | 0.003993 | 0.28858 | 0.28858 | 0.022793
2019-03-19 16:00 | 1553025620.7803 | 1091 | 0.20875 | 0.000571 | 0.000105 | 95.9923 | 95.4249 | 121.8746 | 21.2463 | 0.002563 | 0.002419 | 220.0828 | 0.2657 | 8.4002 | 0.003111 | 0.005533 | 0.005457 | 5.2e-05 | 0.001035 | 0.2995 | 1.3572 | 2.548 | 101.2 | 20.8 | 17.877 | 83.4511 | 96.6585 | 80.0063 | 0.004344 | 0.99969 | 0.024438 | 24.4457 | 0.053024 | 0.15807 | 0.003796 | 0.27343 | 0.27343 | 0.024049
2019-03-19 16:00 | 1553025621.7807 | 1092 | 0.20873 | 0.000572 | 0.000106 | 96.0664 | 95.5033 | 121.7948 | 21.2466 | 0.002555 | 0.00239 | 220.0559 | 0.26606 | 8.4038 | 0.003112 | 0.004452 | 0.00434 | 5.2e-05 | 0.001038 | 0.31166 | 1.3331 | 2.548 | 101.2 | 20.8 | 17.877 | 83.2814 | 96.6981 | 79.3049 | 0.004344 | 1.0017 | 0.024454 | 24.4134 | 0.052579 | 0.15828 | 0.003918 | 0.28251 | 0.28251 | 0.023307
2019-03-19 16:00 | 1553025622.7817 | 1093 | 0.20869 | 0.000574 | 0.000108 | 96.0112 | 95.5092 | 121.6969 | 21.2465 | 0.002585 | 0.002459 | 220.0394 | 0.26645 | 8.4058 | 0.003133 | 0.004365 | 0.004298 | 5.2e-05 | 0.001033 | 0.29192 | 1.3327 | 2.548 | 101.2 | 20.8 | 17.877 | 83.9859 | 97.2042 | 80.9681 | 0.004344 | 0.99697 | 0.024387 | 24.4611 | 0.046891 | 0.15851 | 0.004121 | 0.29659 | 0.29659 | 0.022232
2019-03-19 16:00 | 1553025623.7804 | 1094 | 0.20866 | 0.000574 | 0.00011 | 96.1495 | 95.6095 | 121.6979 | 21.2458 | 0.002633 | 0.002447 | 220.0108 | 0.26602 | 8.4017 | 0.003141 | 0.003998 | 0.003834 | 4.7e-05 | 0.001029 | 0.31492 | 1.3077 | 2.5481 | 101.2 | 20.8 | 17.879 | 85.1532 | 97.3858 | 80.6703 | 0.004345 | 0.99781 | 0.024397 | 24.4509 | 0.050378 | 0.15826 | 0.004307 | 0.31031 | 0.31031 | 0.021216
2019-03-19 16:00 | 1553025624.7805 | 1095 | 0.20869 | 0.000577 | 0.000112 | 96.2224 | 95.8003 | 121.7018 | 21.2459 | 0.002583 | 0.002431 | 220.0123 | 0.26654 | 8.4019 | 0.003178 | 0.004842 | 0.004789 | 4.9e-05 | 0.001026 | 0.28335 | 1.3131 | 2.5481 | 101.2 | 20.8 | 17.88 | 83.9345 | 98.2701 | 80.2921 | 0.004345 | 0.99888 | 0.024411 | 24.4383 | 0.03933 | 0.15856 | 0.004101 | 0.2951 | 0.2951 | 0.022353
2019-03-19 16:00 | 1553025625.7808 | 1096 | 0.20867 | 0.00058 | 0.000113 | 96.2597 | 95.6314 | 121.7868 | 21.2462 | 0.002581 | 0.002478 | 219.9697 | 0.26539 | 8.4028 | 0.00318 | 0.005088 | 0.004968 | 5.2e-05 | 0.001027 | 0.34506 | 1.31 | 2.5481 | 101.2 | 20.8 | 17.882 | 83.9035 | 98.3326 | 81.4192 | 0.004346 | 0.9957 | 0.024381 | 24.4857 | 0.058567 | 0.15788 | 0.004233 | 0.30437 | 0.30437 | 0.021578
2019-03-19 16:00 | 1553025626.7796 | 1097 | 0.20867 | 0.000586 | 0.000114 | 96.3092 | 95.7722 | 121.7556 | 21.2468 | 0.002611 | 0.002443 | 219.9497 | 0.26668 | 8.4042 | 0.003197 | 0.004563 | 0.004431 | 4.8e-05 | 0.001027 | 0.29911 | 1.3347 | 2.548 | 101.2 | 20.8 | 17.888 | 84.6095 | 98.7408 | 80.58 | 0.004347 | 0.99806 | 0.024406 | 24.4536 | 0.050011 | 0.15865 | 0.004252 | 0.30607 | 0.30607 | 0.021563
2019-03-19 16:00 | 1553025627.7797 | 1098 | 0.20865 | 0.000587 | 0.000115 | 96.3476 | 95.8732 | 121.5878 | 21.2465 | 0.00259 | 0.002431 | 219.9412 | 0.2655 | 8.3996 | 0.003204 | 0.003995 | 0.003874 | 5e-05 | 0.001026 | 0.30397 | 1.2941 | 2.548 | 101.2 | 20.8 | 17.894 | 84.1106 | 98.8996 | 80.299 | 0.004349 | 0.99886 | 0.024399 | 24.4271 | 0.044147 | 0.15795 | 0.004372 | 0.31476 | 0.31476 | 0.020876
2019-03-19 16:00 | 1553025628.7797 | 1099 | 0.20866 | 0.00059 | 0.000118 | 96.3824 | 96.0191 | 121.5192 | 21.2458 | 0.002609 | 0.002474 | 219.9471 | 0.26589 | 8.4034 | 0.003146 | 0.004422 | 0.00432 | 4.9e-05 | 0.001009 | 0.29463 | 1.3356 | 2.548 | 101.2 | 20.8 | 17.904 | 84.5764 | 97.5073 | 81.3283 | 0.004351 | 0.99596 | 0.024357 | 24.4557 | 0.033777 | 0.15818 | 0.004265 | 0.3062 | 0.3062 | 0.02149
2019-03-19 16:00 | 1553025629.7808 | 1100 | 0.20864 | 0.0006 | 0.00012 | 96.476 | 95.9486 | 121.845 | 21.2467 | 0.002613 | 0.002447 | 219.9295 | 0.26604 | 8.4041 | 0.003165 | 0.004985 | 0.004848 | 4.8e-05 | 0.001006 | 0.30934 | 1.3536 | 2.548 | 101.2 | 20.8 | 17.91 | 84.6646 | 97.9793 | 80.6712 | 0.004352 | 0.99781 | 0.024412 | 24.4658 | 0.049036 | 0.15827 | 0.004402 | 0.31679 | 0.31679 | 0.020783
2019-03-19 16:00 | 1553025630.7798 | 1101 | 0.20858 | 0.000606 | 0.000123 | 96.4433 | 96.0674 | 121.8476 | 21.2457 | 0.002597 | 0.002389 | 219.8958 | 0.26636 | 8.4012 | 0.003189 | 0.004551 | 0.004546 | 5.2e-05 | 0.001006 | 0.29856 | 1.3915 | 2.5481 | 101.2 | 20.8 | 17.915 | 84.2743 | 98.5428 | 79.2737 | 0.004354 | 1.0018 | 0.024461 | 24.4177 | 0.034931 | 0.15846 | 0.004714 | 0.34031 | 0.34031 | 0.019371
2019-03-19 16:00 | 1553025631.7809 | 1102 | 0.20857 | 0.000608 | 0.000126 | 96.5745 | 96.1078 | 122.1408 | 21.2445 | 0.002541 | 0.002419 | 197.3277 | 0.26551 | 8.4033 | 0.003161 | 0.004871 | 0.004726 | 4.7e-05 | 0.000999 | 0.28655 | 1.3468 | 2.5481 | 101.2 | 20.8 | 17.925 | 82.9308 | 97.8714 | 79.9983 | 0.004356 | 0.99971 | 0.024465 | 24.4721 | 0.043324 | 0.15795 | 0.004804 | 0.3468 | 0.3468 | 0.018947
2019-03-19 16:01 | 1553025632.7812 | 1103 | 0.20858 | 0.000618 | 0.000127 | 96.6435 | 95.9886 | 122.4037 | 21.245 | 0.002554 | 0.002395 | 160.0945 | 0.26545 | 8.4024 | 0.003202 | 0.004333 | 0.004247 | 5e-05 | 0.001005 | 0.31222 | 1.3683 | 2.5482 | 101.2 | 20.8 | 17.931 | 83.2493 | 98.862 | 79.4328 | 0.004358 | 1.0013 | 0.024511 | 24.4788 | 0.060814 | 0.15791 | 0.004712 | 0.3406 | 0.3406 | 0.019287
2019-03-19 16:01 | 1553025633.7806 | 1104 | 0.20851 | 0.000627 | 0.000128 | 96.6168 | 96.1075 | 122.7809 | 21.2452 | 0.002602 | 0.002399 | 122.8659 | 0.26612 | 8.4033 | 0.003214 | 0.004186 | 0.00409 | 4.8e-05 | 0.001005 | 0.28423 | 1.3477 | 2.5482 | 101.2 | 20.8 | 17.946 | 84.3924 | 99.1494 | 79.5141 | 0.004361 | 1.0011 | 0.024546 | 24.5193 | 0.047274 | 0.15831 | 0.005138 | 0.37255 | 0.37255 | 0.017678
2019-03-19 16:01 | 1553025634.78 | 1105 | 0.20862 | 0.000633 | 0.000128 | 96.6399 | 96.2387 | 123.183 | 21.2449 | 0.002558 | 0.002414 | 85.6169 | 0.26604 | 8.4041 | 0.003261 | 0.004391 | 0.004295 | 5e-05 | 0.000998 | 0.32278 | 1.3247 | 2.5481 | 101.2 | 20.8 | 17.956 | 83.3288 | 100.2882 | 79.8841 | 0.004364 | 1 | 0.024573 | 24.5723 | 0.037213 | 0.15827 | 0.004455 | 0.32235 | 0.32235 | 0.020425
2019-03-19 16:01 | 1553025635.78 | 1106 | 0.20858 | 0.00064 | 0.000123 | 96.747 | 96.2335 | 123.6407 | 21.2439 | 0.002629 | 0.00241 | 48.3861 | 0.26642 | 8.401 | 0.003313 | 0.004023 | 0.003979 | 4.8e-05 | 0.000997 | 0.30165 | 1.3294 | 2.5481 | 101.2 | 20.8 | 17.965 | 85.0357 | 101.5414 | 79.7833 | 0.004366 | 1.0003 | 0.024622 | 24.6144 | 0.047601 | 0.15849 | 0.004725 | 0.34337 | 0.34337 | 0.019202
2019-03-19 16:01 | 1553025636.7811 | 1107 | 0.20856 | 0.00065 | 0.000117 | 96.768 | 96.2804 | 123.9683 | 21.2446 | 0.002612 | 0.002391 | 11.1357 | 0.2481 | 8.4036 | 0.003243 | 0.004791 | 0.004633 | 4.9e-05 | 0.000996 | 0.29358 | 1.3277 | 2.5481 | 101.2 | 20.8 | 17.985 | 84.6494 | 99.8444 | 79.3171 | 0.004371 | 1.0016 | 0.024671 | 24.6307 | 0.045183 | 0.1476 | 0.004786 | 0.34898 | 0.34898 | 0.017594
2019-03-19 16:01 | 1553025637.7801 | 1108 | 0.20875 | 0.000652 | 0.000108 | 96.8163 | 96.4572 | 124.3838 | 21.2437 | 0.002565 | 0.002324 | -26.1026 | 0.23696 | 8.4032 | 0.003287 | 0.004791 | 0.004553 | 4.8e-05 | 0.000997 | 0.29103 | 1.3866 | 2.548 | 101.2 | 20.8 | 17.995 | 83.5055 | 100.9077 | 77.7196 | 0.004373 | 1.0062 | 0.024769 | 24.616 | 0.033245 | 0.14097 | 0.003683 | 0.26853 | 0.26853 | 0.021839
2019-03-19 16:01 | 1553025638.7802 | 1109 | 0.20893 | 0.000655 | 9.9e-05 | 96.8614 | 96.4904 | 124.7178 | 21.2432 | 0.002606 | 0.002367 | -63.3601 | 0.23853 | 8.4025 | 0.003234 | 0.004259 | 0.003965 | 5.2e-05 | 0.000995 | 0.33085 | 1.3369 | 2.5481 | 101.2 | 20.8 | 18.004 | 84.4811 | 99.6196 | 78.7451 | 0.004375 | 1.0033 | 0.024766 | 24.685 | 0.034323 | 0.1419 | 0.002603 | 0.18835 | 0.18835 | 0.031341
2019-03-19 16:01 | 1553025639.7813 | 1110 | 0.20894 | 0.000653 | 9.1e-05 | 96.9322 | 96.3873 | 124.6432 | 21.2433 | 0.002537 | 0.002268 | -100.5983 | 0.2371 | 8.402 | 0.003267 | 0.004416 | 0.004194 | 5e-05 | 0.000994 | 0.29142 | 1.3765 | 2.548 | 101.2 | 20.8 | 18.021 | 82.8334 | 100.4388 | 76.3793 | 0.004379 | 1.0101 | 0.024842 | 24.5945 | 0.050422 | 0.14105 | 0.002532 | 0.18413 | 0.18413 | 0.031868
2019-03-19 16:01 | 1553025640.7804 | 1111 | 0.20907 | 0.000637 | 8.3e-05 | 96.9832 | 96.4589 | 125.1607 | 21.2438 | 0.00256 | 0.002373 | -137.8412 | 0.23717 | 8.4013 | 0.003197 | 0.004747 | 0.00453 | 4.9e-05 | 0.000997 | 0.29192 | 1.3448 | 2.5481 | 101.2 | 20.8 | 18.029 | 83.3785 | 98.7268 | 78.9 | 0.004381 | 1.0028 | 0.024804 | 24.7343 | 0.048488 | 0.14109 | 0.001773 | 0.12757 | 0.12757 | 0.046008
2019-03-19 16:01 | 1553025641.7814 | 1112 | 0.20911 | 0.000611 | 7.5e-05 | 96.9106 | 96.4294 | 125.6213 | 21.2433 | 0.002506 | 0.002301 | -152.5126 | 0.23745 | 8.4014 | 0.003291 | 0.004542 | 0.004363 | 4.9e-05 | 0.000994 | 0.31957 | 1.3221 | 2.5481 | 101.2 | 20.8 | 18.042 | 82.0963 | 101.0101 | 77.1594 | 0.004385 | 1.0078 | 0.024911 | 24.7184 | 0.04452 | 0.14126 | 0.001627 | 0.11759 | 0.11759 | 0.049976
2019-03-19 16:01 | 1553025642.7794 | 1113 | 0.20916 | 0.000576 | 6.8e-05 | 97.0642 | 96.5651 | 126.0811 | 21.2422 | 0.002504 | 0.002305 | -152.5254 | 0.2362 | 8.4035 | 0.003253 | 0.004362 | 0.00419 | 4.9e-05 | 0.000991 | 0.30807 | 1.3397 | 2.5481 | 101.2 | 20.8 | 18.048 | 82.0494 | 100.0839 | 77.2552 | 0.004386 | 1.0075 | 0.024954 | 24.767 | 0.046118 | 0.14052 | 0.001339 | 0.096566 | 0.096566 | 0.060533
2019-03-19 16:01 | 1553025643.7795 | 1114 | 0.20919 | 0.000539 | 6e-05 | 97.1002 | 96.5743 | 125.9445 | 21.2426 | 0.002483 | 0.002288 | -152.5288 | 0.23644 | 8.4047 | 0.003305 | 0.004153 | 0.003898 | 4.7e-05 | 0.000992 | 0.32405 | 1.3193 | 2.5481 | 101.2 | 20.8 | 18.053 | 81.5265 | 101.3548 | 76.8448 | 0.004387 | 1.0087 | 0.024955 | 24.7391 | 0.04858 | 0.14066 | 0.001209 | 0.087274 | 0.087274 | 0.067046
2019-03-19 16:01 | 1553025644.7805 | 1115 | 0.20927 | 0.0005 | 5.3e-05 | 97.0819 | 96.8264 | 126.013 | 21.2422 | 0.002555 | 0.002251 | -152.5152 | 0.23648 | 8.4019 | 0.003219 | 0.004925 | 0.004603 | 4.8e-05 | 0.000992 | 0.30917 | 1.317 | 2.5482 | 101.19 | 20.8 | 18.063 | 83.2646 | 99.271 | 75.9689 | 0.00439 | 1.0112 | 0.024993 | 24.7148 | 0.023577 | 0.14068 | 0.000773 | 0.054996 | 0.054996 | 0.10641
2019-03-19 16:01 | 1553025645.781 | 1116 | 0.20932 | 0.000451 | 4.6e-05 | 97.1555 | 96.6825 | 126.3914 | 21.2419 | 0.002412 | 0.002217 | -152.515 | 0.23586 | 8.4068 | 0.003319 | 0.004345 | 0.004126 | 5e-05 | 0.000992 | 0.30342 | 1.3109 | 2.548 | 101.19 | 20.8 | 18.066 | 79.8297 | 101.6871 | 75.1416 | 0.004391 | 1.0136 | 0.02506 | 24.7225 | 0.043647 | 0.14031 | 0.00059 | 0.041701 | 0.041701 | 0.13997
2019-03-19 16:01 | 1553025646.7807 | 1117 | 0.20936 | 0.000415 | 3.9e-05 | 97.2602 | 96.9329 | 126.4407 | 21.2418 | 0.002502 | 0.002249 | -152.5078 | 0.23646 | 8.4034 | 0.003276 | 0.004442 | 0.0043 | 5.4e-05 | 0.000997 | 0.29247 | 1.3381 | 2.5481 | 101.19 | 20.8 | 18.067 | 81.9826 | 100.6359 | 75.9233 | 0.004391 | 1.0114 | 0.025037 | 24.7551 | 0.030152 | 0.14067 | 0.000389 | 0.026926 | 0.026926 | 0.21733
2019-03-19 16:01 | 1553025647.7797 | 1118 | 0.20934 | 0.000374 | 3.4e-05 | 97.3934 | 97.0017 | 126.8115 | 21.2422 | 0.002374 | 0.002264 | -152.5257 | 0.23683 | 8.4027 | 0.003342 | 0.004415 | 0.00416 | 4.9e-05 | 0.000993 | 0.31061 | 1.333 | 2.5481 | 101.19 | 20.8 | 18.066 | 78.9164 | 102.2544 | 76.266 | 0.004391 | 1.0104 | 0.025061 | 24.8035 | 0.036048 | 0.14089 | 0.000573 | 0.041219 | 0.041219 | 0.14219
2019-03-19 16:01 | 1553025648.7798 | 1119 | 0.20942 | 0.000335 | 2.9e-05 | 97.3829 | 97.0137 | 127.1695 | 21.2407 | 0.002494 | 0.002229 | -152.5277 | 0.23653 | 8.402 | 0.003265 | 0.004725 | 0.004491 | 5.2e-05 | 0.000989 | 0.30287 | 1.3115 | 2.548 | 101.19 | 20.8 | 18.065 | 81.8092 | 100.3765 | 75.4242 | 0.004391 | 1.0128 | 0.025127 | 24.8086 | 0.033978 | 0.14071 | 0.000159 | 0.010232 | 0.010232 | 0.57212
2019-03-19 16:01 | 1553025649.7801 | 1120 | 0.20943 | 0.000297 | 2.4e-05 | 97.5927 | 97.1545 | 127.5044 | 21.2416 | 0.002458 | 0.002219 | -152.515 | 0.23675 | 8.4031 | 0.003314 | 0.004478 | 0.004323 | 4.5e-05 | 0.000986 | 0.31133 | 1.3413 | 2.5481 | 101.19 | 20.8 | 18.061 | 80.9273 | 101.5636 | 75.1865 | 0.00439 | 1.0135 | 0.025168 | 24.8328 | 0.040255 | 0.14084 | 0.000144 | 0.009424 | 0.009424 | 0.62174
2019-03-19 16:01 | 1553025650.781 | 1121 | 0.20945 | 0.000269 | 2e-05 | 97.59 | 97.1576 | 127.3674 | 21.2411 | 0.002412 | 0.002158 | -152.5112 | 0.23554 | 8.4036 | 0.003308 | 0.004365 | 0.004076 | 5.1e-05 | 0.000979 | 0.29535 | 1.3441 | 2.5481 | 101.19 | 20.8 | 18.059 | 79.8352 | 101.4312 | 73.7229 | 0.004389 | 1.0178 | 0.025208 | 24.7672 | 0.039724 | 0.14012 | 4.7e-05 | 0.002269 | 0.002269 | 2.5694
2019-03-19 16:01 | 1553025651.7801 | 1122 | 0.20946 | 0.000237 | 1.7e-05 | 97.6884 | 97.3939 | 127.4307 | 21.2401 | 0.002323 | 0.002183 | -152.5113 | 0.23668 | 8.401 | 0.00327 | 0.004429 | 0.004233 | 5.2e-05 | 0.000975 | 0.32493 | 1.3679 | 2.548 | 101.2 | 20.8 | 18.053 | 77.6918 | 100.4874 | 74.3199 | 0.004387 | 1.016 | 0.025193 | 24.7947 | 0.027012 | 0.1408 | 4.5e-05 | 0.002343 | 0.002343 | 2.4994
2019-03-19 16:01 | 1553025652.7813 | 1123 | 0.20947 | 0.000213 | 1.4e-05 | 97.7314 | 97.4699 | 127.4251 | 21.2405 | 0.002386 | 0.002209 | -152.4891 | 0.23637 | 8.4032 | 0.003342 | 0.004847 | 0.004648 | 5.2e-05 | 0.000975 | 0.28567 | 1.3157 | 2.5479 | 101.2 | 20.8 | 18.05 | 79.193 | 102.2386 | 74.9399 | 0.004386 | 1.0142 | 0.02517 | 24.8162 | 0.02396 | 0.14062 | -2e-05 | -0.002419 | -0.002419 | -2.4184
2019-03-19 16:01 | 1553025653.781 | 1124 | 0.20949 | 0.000183 | 1.2e-05 | 97.8226 | 97.4665 | 127.5533 | 21.2409 | 0.002359 | 0.002206 | -152.4883 | 0.23743 | 8.3993 | 0.003325 | 0.004976 | 0.004788 | 5.4e-05 | 0.000984 | 0.32709 | 1.321 | 2.5481 | 101.2 | 20.79 | 18.046 | 78.548 | 101.8274 | 74.8819 | 0.004383 | 1.0144 | 0.025184 | 24.8267 | 0.03262 | 0.14125 | -9.4e-05 | -0.007921 | -0.007921 | -0.74177
2019-03-19 16:01 | 1553025654.7801 | 1125 | 0.20946 | 0.000163 | 1e-05 | 97.8355 | 97.4706 | 127.4401 | 21.2404 | 0.00239 | 0.002125 | -152.4875 | 0.23688 | 8.4031 | 0.003381 | 0.005068 | 0.004877 | 5.2e-05 | 0.000989 | 0.31492 | 1.3167 | 2.5482 | 101.2 | 20.79 | 18.039 | 79.3006 | 103.1813 | 72.9358 | 0.004381 | 1.0201 | 0.025244 | 24.7462 | 0.033423 | 0.14092 | 0.000114 | 0.008033 | 0.008033 | 0.72977
2019-03-19 16:01 | 1553025655.7811 | 1126 | 0.20947 | 0.000147 | 8e-06 | 97.8151 | 97.4278 | 127.3242 | 21.2409 | 0.002405 | 0.002192 | -152.4876 | 0.23731 | 8.4059 | 0.00335 | 0.004199 | 0.003956 | 5.2e-05 | 0.000985 | 0.34036 | 1.3226 | 2.548 | 101.19 | 20.79 | 18.035 | 79.6659 | 102.4394 | 74.5521 | 0.004381 | 1.0154 | 0.025174 | 24.7926 | 0.035487 | 0.14118 | 9.5e-05 | 0.006717 | 0.006717 | 0.87432
2019-03-19 16:01 | 1553025656.7816 | 1127 | 0.2095 | 0.00013 | 7e-06 | 97.9597 | 97.648 | 127.5523 | 21.2405 | 0.002301 | 0.002193 | -152.4849 | 0.23643 | 8.4022 | 0.003325 | 0.005208 | 0.005026 | 5.3e-05 | 0.000991 | 0.30502 | 1.3126 | 2.5482 | 101.19 | 20.79 | 18.03 | 77.1692 | 101.8232 | 74.5697 | 0.004379 | 1.0153 | 0.025195 | 24.8154 | 0.028509 | 0.14065 | -5.1e-05 | -0.004282 | -0.004282 | -1.3664
2019-03-19 16:01 | 1553025657.7792 | 1128 | 0.20948 | 0.000113 | 6e-06 | 98.0555 | 97.7333 | 127.2754 | 21.2397 | 0.002357 | 0.002173 | -152.4799 | 0.23697 | 8.4004 | 0.003324 | 0.004289 | 0.00401 | 5.4e-05 | 0.000991 | 0.3079 | 1.3149 | 2.5481 | 101.19 | 20.79 | 18.023 | 78.5069 | 101.8011 | 74.0896 | 0.004378 | 1.0167 | 0.025186 | 24.7713 | 0.029438 | 0.14098 | 6.3e-05 | 0.004377 | 0.004377 | 1.3397
2019-03-19 16:01 | 1553025658.7803 | 1129 | 0.2095 | 0.00011 | 5e-06 | 98.1431 | 97.9573 | 126.9524 | 21.2402 | 0.00234 | 0.00217 | -152.4713 | 0.23639 | 8.4037 | 0.003368 | 0.004682 | 0.004605 | 5.7e-05 | 0.000996 | 0.33157 | 1.3225 | 2.5482 | 101.19 | 20.79 | 18.02 | 78.099 | 102.8683 | 74.0176 | 0.004377 | 1.0169 | 0.025156 | 24.7373 | 0.016952 | 0.14063 | -8e-06 | -0.000941 | -0.000941 | -6.2174
2019-03-19 16:01 | 1553025659.7803 | 1130 | 0.20948 | 9.5e-05 | 4e-06 | 98.1272 | 97.8436 | 126.8442 | 21.2388 | 0.002305 | 0.002177 | -152.4606 | 0.23697 | 8.4019 | 0.003328 | 0.004762 | 0.00449 | 5.7e-05 | 0.000998 | 0.30021 | 1.3495 | 2.5482 | 101.19 | 20.79 | 18.017 | 77.2456 | 101.8937 | 74.1791 | 0.004376 | 1.0165 | 0.02514 | 24.7325 | 0.025891 | 0.14098 | 8.3e-05 | 0.006039 | 0.006039 | 0.97108
2019-03-19 16:01 | 1553025660.7794 | 1131 | 0.20949 | 8.7e-05 | 3e-06 | 98.3115 | 98.2022 | 126.7993 | 21.2391 | 0.002355 | 0.002122 | -152.4503 | 0.23629 | 8.4012 | 0.003275 | 0.0056 | 0.005355 | 5.7e-05 | 0.000995 | 0.31636 | 1.3312 | 2.5481 | 101.19 | 20.8 | 18.016 | 78.4595 | 100.6253 | 72.8558 | 0.004379 | 1.0203 | 0.025183 | 24.681 | 0.00995 | 0.14057 | 7e-05 | 0.005139 | 0.005139 | 1.1378
2019-03-19 16:01 | 1553025661.7795 | 1132 | 0.20949 | 7.7e-05 | 3e-06 | 98.3727 | 98.1398 | 126.6003 | 21.24 | 0.002307 | 0.00213 | -152.4669 | 0.2365 | 8.4052 | 0.003382 | 0.004674 | 0.004469 | 5.6e-05 | 0.000997 | 0.27743 | 1.3425 | 2.5481 | 101.19 | 20.8 | 18.014 | 77.3058 | 103.2131 | 73.0421 | 0.004378 | 1.0198 | 0.025157 | 24.6683 | 0.021198 | 0.14069 | 5.6e-05 | 0.004053 | 0.004053 | 1.444
2019-03-19 16:01 | 1553025662.7808 | 1133 | 0.20951 | 7.3e-05 | 3e-06 | 98.493 | 98.2361 | 126.7141 | 21.2384 | 0.002348 | 0.002088 | -152.4917 | 0.23687 | 8.4002 | 0.003315 | 0.004807 | 0.004629 | 5.5e-05 | 0.000999 | 0.31708 | 1.3489 | 2.5482 | 101.19 | 20.8 | 18.012 | 78.2921 | 101.5857 | 72.042 | 0.004378 | 1.0228 | 0.025204 | 24.6437 | 0.023366 | 0.14091 | -5.2e-05 | -0.004167 | -0.004167 | -1.4069
2019-03-19 16:01 | 1553025663.7796 | 1134 | 0.20949 | 6.5e-05 | 2e-06 | 98.5764 | 98.3957 | 126.6346 | 21.2397 | 0.002279 | 0.002117 | -152.4999 | 0.23629 | 8.4036 | 0.003349 | 0.004928 | 0.004785 | 5.5e-05 | 0.000999 | 0.31349 | 1.3637 | 2.5482 | 101.19 | 20.8 | 18.011 | 76.6266 | 102.4099 | 72.7347 | 0.004377 | 1.0207 | 0.025171 | 24.6607 | 0.016419 | 0.14057 | 8e-05 | 0.005968 | 0.005968 | 0.97992
2019-03-19 16:01 | 1553025664.7806 | 1135 | 0.20952 | 6.6e-05 | 2e-06 | 98.6563 | 98.5005 | 126.586 | 21.2387 | 0.002331 | 0.002043 | -152.5199 | 0.2368 | 8.4022 | 0.003339 | 0.004681 | 0.004422 | 5.6e-05 | 0.001003 | 0.32676 | 1.3606 | 2.5482 | 101.19 | 20.8 | 18.011 | 77.8878 | 102.1685 | 70.9492 | 0.004377 | 1.026 | 0.025232 | 24.5922 | 0.014138 | 0.14087 | -8.9e-05 | -0.006894 | -0.006894 | -0.85001
2019-03-19 16:01 | 1553025665.7796 | 1136 | 0.2095 | 5.7e-05 | 1e-06 | 98.7817 | 98.707 | 126.4374 | 21.2397 | 0.002324 | 0.002073 | -152.5299 | 0.23664 | 8.4029 | 0.00334 | 0.005291 | 0.005109 | 5.7e-05 | 0.001006 | 0.3011 | 1.3522 | 2.5482 | 101.19 | 20.79 | 18.012 | 77.7039 | 102.1845 | 71.6854 | 0.004375 | 1.0238 | 0.02519 | 24.604 | 0.006769 | 0.14078 | 3.7e-05 | 0.002719 | 0.002719 | 2.1541
2019-03-19 16:01 | 1553025666.7801 | 1137 | 0.20951 | 5.8e-05 | 2e-06 | 98.8937 | 98.7765 | 126.4543 | 21.239 | 0.00218 | 0.002008 | -152.5259 | 0.23667 | 8.4029 | 0.003285 | 0.004378 | 0.004206 | 5.5e-05 | 0.001001 | 0.31293 | 1.3472 | 2.5481 | 101.19 | 20.8 | 18.013 | 74.2563 | 100.8567 | 70.1268 | 0.004378 | 1.0285 | 0.025249 | 24.55 | 0.010602 | 0.1408 | -2e-06 | -0.000299 | -0.000299 | -19.5791
2019-03-19 16:01 | 1553025667.7809 | 1138 | 0.20949 | 4.9e-05 | 1e-06 | 98.972 | 98.6953 | 126.4052 | 21.2384 | 0.002246 | 0.002108 | -152.528 | 0.23722 | 8.4033 | 0.003273 | 0.004194 | 0.00396 | 5.4e-05 | 0.000999 | 0.30879 | 1.287 | 2.548 | 101.19 | 20.8 | 18.016 | 75.8266 | 100.5791 | 72.5089 | 0.004379 | 1.0214 | 0.025157 | 24.6303 | 0.025036 | 0.14112 | 0.000109 | 0.008188 | 0.008188 | 0.71701
2019-03-19 16:01 | 1553025668.7802 | 1139 | 0.20951 | 5e-05 | 1e-06 | 99.0887 | 99.0957 | 126.5656 | 21.2384 | 0.002222 | 0.002085 | -152.5272 | 0.23582 | 8.4022 | 0.003306 | 0.00437 | 0.004141 | 5.5e-05 | 0.001002 | 0.32068 | 1.3052 | 2.5479 | 101.19 | 20.8 | 18.02 | 75.2708 | 101.3663 | 71.9631 | 0.00438 | 1.023 | 0.025192 | 24.6264 | -0.000631 | 0.14029 | 4e-06 | 0.000232 | 0.000232 | 25.1813
2019-03-19 16:01 | 1553025669.7798 | 1140 | 0.2095 | 4.2e-05 | 1e-06 | 99.2505 | 99.108 | 126.6125 | 21.238 | 0.002281 | 0.002099 | -152.5233 | 0.23684 | 8.4041 | 0.003308 | 0.004957 | 0.004702 | 5.7e-05 | 0.000999 | 0.3063 | 1.3125 | 2.5479 | 101.19 | 20.8 | 18.025 | 76.6663 | 101.4282 | 72.2948 | 0.004381 | 1.022 | 0.025185 | 24.6428 | 0.012846 | 0.1409 | 7.8e-05 | 0.00584 | 0.00584 | 1.0037
2019-03-19 16:01 | 1553025670.78 | 1141 | 0.2095 | 4.5e-05 | 1e-06 | 99.3471 | 99.1998 | 126.531 | 21.238 | 0.002238 | 0.002103 | -152.5564 | 0.23639 | 8.4 | 0.003291 | 0.00415 | 0.003972 | 5.6e-05 | 0.000999 | 0.32637 | 1.3627 | 2.5481 | 101.19 | 20.79 | 18.035 | 75.6382 | 101.0151 | 72.4118 | 0.004381 | 1.0217 | 0.025173 | 24.6391 | 0.01327 | 0.14063 | 5.1e-05 | 0.003851 | 0.003851 | 1.519
2019-03-19 16:01 | 1553025671.78 | 1142 | 0.20951 | 3.9e-05 | 1e-06 | 99.4504 | 99.3331 | 126.6038 | 21.2381 | 0.002255 | 0.002126 | -152.5604 | 0.2365 | 8.4031 | 0.003277 | 0.004497 | 0.004337 | 5.6e-05 | 0.001005 | 0.32316 | 1.3162 | 2.548 | 101.19 | 20.79 | 18.041 | 76.0576 | 100.6645 | 72.943 | 0.004382 | 1.0201 | 0.025161 | 24.6651 | 0.010555 | 0.1407 | -1.9e-05 | -0.001494 | -0.001494 | -3.917
2019-03-19 16:01 | 1553025672.7803 | 1143 | 0.20947 | 4.4e-05 | 1e-06 | 99.5014 | 99.3718 | 126.5949 | 21.2376 | 0.002218 | 0.002101 | -152.5535 | 0.23651 | 8.4037 | 0.003194 | 0.004262 | 0.004037 | 5.2e-05 | 0.001 | 0.3194 | 1.3595 | 2.5481 | 101.19 | 20.8 | 18.055 | 75.1617 | 98.6571 | 72.342 | 0.004388 | 1.0219 | 0.025182 | 24.6428 | 0.011658 | 0.1407 | 0.000258 | 0.019558 | 0.019558 | 0.29927
2019-03-19 16:01 | 1553025673.7794 | 1144 | 0.20952 | 3.9e-05 | 1e-06 | 99.5959 | 99.5598 | 126.7346 | 21.2374 | 0.002233 | 0.00206 | -152.5443 | 0.23621 | 8.4053 | 0.003216 | 0.00424 | 0.004071 | 5.8e-05 | 0.001002 | 0.29999 | 1.3534 | 2.548 | 101.19 | 20.8 | 18.062 | 75.5308 | 99.1817 | 71.3728 | 0.00439 | 1.0247 | 0.025231 | 24.6218 | 0.003244 | 0.14052 | -3.3e-05 | -0.002538 | -0.002538 | -2.3031
2019-03-19 16:01 | 1553025674.7802 | 1145 | 0.20949 | 4e-05 | 1e-06 | 99.7111 | 99.856 | 126.7498 | 21.2365 | 0.002214 | 0.002082 | -152.5451 | 0.23702 | 8.4031 | 0.003184 | 0.004889 | 0.004671 | 5.8e-05 | 0.000995 | 0.27417 | 1.3437 | 2.548 | 101.19 | 20.8 | 18.068 | 75.0663 | 98.4234 | 71.8835 | 0.004391 | 1.0232 | 0.025214 | 24.6415 | -0.012989 | 0.141 | 0.000125 | 0.009459 | 0.009459 | 0.62009
2019-03-19 16:01 | 1553025675.7802 | 1146 | 0.20952 | 3.9e-05 | 1e-06 | 99.8252 | 100.0831 | 126.6069 | 21.2359 | 0.002187 | 0.002005 | -152.5491 | 0.23657 | 8.4024 | 0.003242 | 0.004526 | 0.00439 | 5.8e-05 | 0.000995 | 0.3147 | 1.3473 | 2.5481 | 101.19 | 20.8 | 18.081 | 74.4124 | 99.8282 | 70.0318 | 0.004394 | 1.0287 | 0.025267 | 24.5614 | -0.023082 | 0.14074 | -5.4e-05 | -0.004153 | -0.004153 | -1.4096
2019-03-19 16:01 | 1553025676.7806 | 1147 | 0.20953 | 3.5e-05 | 0 | 99.9377 | 100.0475 | 126.4285 | 21.2362 | 0.002166 | 0.002035 | -152.5508 | 0.23588 | 8.4025 | 0.003242 | 0.004928 | 0.0048 | 5.5e-05 | 0.000993 | 0.2995 | 1.3249 | 2.5481 | 101.19 | 20.8 | 18.086 | 73.917 | 99.8299 | 70.7636 | 0.004396 | 1.0266 | 0.025223 | 24.5703 | -0.00982 | 0.14033 | -0.000102 | -0.007768 | -0.007768 | -0.75151
2019-03-19 16:01 | 1553025677.7803 | 1148 | 0.20953 | 3.7e-05 | 0 | 100.071 | 99.9519 | 126.6836 | 21.2363 | 0.002199 | 0.002062 | -152.5446 | 0.23628 | 8.4027 | 0.003253 | 0.004702 | 0.004462 | 5.6e-05 | 0.001002 | 0.29839 | 1.3479 | 2.5481 | 101.19 | 20.8 | 18.093 | 74.7035 | 100.0858 | 71.4061 | 0.004397 | 1.0246 | 0.025225 | 24.618 | 0.010645 | 0.14056 | -0.000104 | -0.007922 | -0.007922 | -0.73814
2019-03-19 16:01 | 1553025678.7804 | 1149 | 0.20953 | 3.1e-05 | 1e-06 | 100.1907 | 100.2338 | 126.4259 | 21.2357 | 0.002187 | 0.002077 | -152.5546 | 0.23735 | 8.4042 | 0.003291 | 0.00448 | 0.00429 | 5.5e-05 | 0.001007 | 0.32925 | 1.3649 | 2.548 | 101.19 | 20.8 | 18.093 | 74.4098 | 101.0136 | 71.7712 | 0.004397 | 1.0236 | 0.025186 | 24.606 | -0.003848 | 0.1412 | -9e-05 | -0.006921 | -0.006921 | -0.84867
2019-03-19 16:01 | 1553025679.7794 | 1150 | 0.20954 | 3.5e-05 | 1e-06 | 100.2987 | 100.2427 | 126.4597 | 21.2352 | 0.002194 | 0.002037 | -152.5922 | 0.23671 | 8.4028 | 0.003266 | 0.004685 | 0.004488 | 5.4e-05 | 0.001012 | 0.29574 | 1.3497 | 2.5481 | 101.19 | 20.8 | 18.092 | 74.5887 | 100.3868 | 70.8155 | 0.004397 | 1.0264 | 0.025224 | 24.5752 | 0.004998 | 0.14082 | -0.000145 | -0.011117 | -0.011117 | -0.52693
2019-03-19 16:01 | 1553025680.7798 | 1151 | 0.20951 | 2.7e-05 | 0 | 100.3125 | 100.3794 | 126.532 | 21.2358 | 0.002197 | 0.002001 | -152.5635 | 0.23646 | 8.4054 | 0.00324 | 0.00547 | 0.005281 | 5.7e-05 | 0.001013 | 0.28783 | 1.3365 | 2.5482 | 101.19 | 20.8 | 18.084 | 74.6585 | 99.7599 | 69.9461 | 0.004395 | 1.029 | 0.025263 | 24.5511 | -0.005957 | 0.14067 | 2.5e-05 | 0.001861 | 0.001861 | 3.1444
2019-03-19 16:01 | 1553025681.7805 | 1152 | 0.2095 | 3.3e-05 | 1e-06 | 100.4491 | 100.4177 | 126.6571 | 21.235 | 0.002223 | 0.002091 | -152.566 | 0.23775 | 8.4043 | 0.003328 | 0.004792 | 0.004589 | 5.5e-05 | 0.001012 | 0.31421 | 1.3902 | 2.5481 | 101.19 | 20.8 | 18.078 | 75.2884 | 101.8899 | 72.1197 | 0.004394 | 1.0225 | 0.025196 | 24.6409 | 0.002793 | 0.14144 | 9.2e-05 | 0.006963 | 0.006963 | 0.84501
2019-03-19 16:01 | 1553025682.7796 | 1153 | 0.20955 | 3e-05 | 1e-06 | 100.4311 | 100.4273 | 126.7766 | 21.2357 | 0.002171 | 0.001964 | -152.5613 | 0.23789 | 8.4037 | 0.003283 | 0.004827 | 0.004549 | 5.6e-05 | 0.001009 | 0.28799 | 1.3666 | 2.5481 | 101.19 | 20.8 | 18.065 | 74.0326 | 100.8143 | 69.0483 | 0.004391 | 1.0317 | 0.025321 | 24.5426 | 0.000336 | 0.14152 | -0.000224 | -0.017182 | -0.017182 | -0.34264
2019-03-19 16:01 | 1553025683.7796 | 1154 | 0.20956 | 2.9e-05 | 1e-06 | 100.5961 | 100.6698 | 126.8437 | 21.2351 | 0.002171 | 0.00201 | -152.588 | 0.23786 | 8.4027 | 0.00335 | 0.004437 | 0.004342 | 5.8e-05 | 0.001012 | 0.31277 | 1.3892 | 2.548 | 101.19 | 20.8 | 18.058 | 74.0258 | 102.4415 | 70.1623 | 0.004389 | 1.0284 | 0.025286 | 24.5891 | -0.006547 | 0.1415 | -0.000265 | -0.020281 | -0.020281 | -0.29025
2019-03-19 16:01 | 1553025684.7811 | 1155 | 0.20953 | 3.1e-05 | 0 | 100.6189 | 100.5585 | 126.627 | 21.2341 | 0.00209 | 0.002016 | -152.5859 | 0.2379 | 8.4026 | 0.003297 | 0.004625 | 0.00443 | 5.6e-05 | 0.001016 | 0.28224 | 1.3214 | 2.548 | 101.19 | 20.8 | 18.051 | 72.0729 | 101.1463 | 70.3086 | 0.004387 | 1.0279 | 0.025259 | 24.5733 | 0.005376 | 0.14153 | -0.000101 | -0.007748 | -0.007748 | -0.7599
2019-03-19 16:01 | 1553025685.7808 | 1156 | 0.20957 | 3.9e-05 | 0 | 100.6316 | 100.7616 | 126.8805 | 21.2361 | 0.002121 | 0.00206 | -152.5851 | 0.23698 | 8.4033 | 0.00328 | 0.005158 | 0.004938 | 5.5e-05 | 0.001019 | 0.31868 | 1.2934 | 2.548 | 101.19 | 20.8 | 18.035 | 72.8286 | 100.73 | 71.3544 | 0.004383 | 1.0248 | 0.025246 | 24.6353 | -0.011551 | 0.14098 | -0.000375 | -0.0286 | -0.0286 | -0.20506
2019-03-19 16:01 | 1553025686.7801 | 1157 | 0.20954 | 3.1e-05 | 1e-06 | 100.772 | 100.7544 | 127.1927 | 21.2342 | 0.00216 | 0.001985 | -152.5856 | 0.23779 | 8.4039 | 0.003297 | 0.004783 | 0.004605 | 5.2e-05 | 0.001024 | 0.31686 | 1.3344 | 2.5482 | 101.19 | 20.8 | 18.027 | 73.7748 | 101.1566 | 69.5677 | 0.004381 | 1.0301 | 0.025343 | 24.6015 | 0.001563 | 0.14146 | -0.000189 | -0.014463 | -0.014463 | -0.40689
2019-03-19 16:01 | 1553025687.7809 | 1158 | 0.20957 | 2.5e-05 | 0 | 100.7573 | 100.7328 | 127.1934 | 21.2341 | 0.00215 | 0.002032 | -152.5771 | 0.2366 | 8.4051 | 0.003323 | 0.00421 | 0.003976 | 5.4e-05 | 0.001021 | 0.30989 | 1.337 | 2.5481 | 101.19 | 20.8 | 18.011 | 73.5261 | 101.7831 | 70.6929 | 0.004377 | 1.0268 | 0.025301 | 24.642 | 0.002173 | 0.14075 | -0.000352 | -0.02692 | -0.02692 | -0.2175
2019-03-19 16:01 | 1553025688.7789 | 1159 | 0.20952 | 2.9e-05 | 1e-06 | 100.8638 | 100.95 | 127.1024 | 21.2349 | 0.002159 | 0.001975 | -152.5625 | 0.23794 | 8.4035 | 0.003244 | 0.004015 | 0.003812 | 5.1e-05 | 0.001027 | 0.31708 | 1.3647 | 2.548 | 101.19 | 20.81 | 18.004 | 73.7449 | 99.8639 | 69.3105 | 0.004378 | 1.0309 | 0.025343 | 24.5836 | -0.007634 | 0.14155 | -6.7e-05 | -0.005164 | -0.005164 | -1.1402
2019-03-19 16:01 | 1553025689.7789 | 1160 | 0.20949 | 3e-05 | 0 | 100.8728 | 100.9465 | 127.0165 | 21.2341 | 0.002076 | 0.002047 | -152.5385 | 0.23679 | 8.4037 | 0.003275 | 0.004774 | 0.004623 | 5.4e-05 | 0.001025 | 0.2995 | 1.3741 | 2.5482 | 101.19 | 20.81 | 17.998 | 71.7516 | 100.6078 | 71.043 | 0.004377 | 1.0257 | 0.025271 | 24.6373 | -0.006532 | 0.14087 | 0.000134 | 0.010206 | 0.010206 | 0.5742
2019-03-19 16:01 | 1553025690.7789 | 1161 | 0.20948 | 2.7e-05 | 0 | 100.8785 | 100.9301 | 127.3716 | 21.2352 | 0.002144 | 0.002027 | -152.5352 | 0.23775 | 8.4045 | 0.003267 | 0.004391 | 0.004151 | 5.3e-05 | 0.001027 | 0.29789 | 1.3764 | 2.5482 | 101.19 | 20.81 | 17.994 | 73.3757 | 100.4277 | 70.5747 | 0.004376 | 1.0271 | 0.025324 | 24.655 | -0.00457 | 0.14144 | 0.000185 | 0.014133 | 0.014133 | 0.41633
2019-03-19 16:02 | 1553025691.7791 | 1162 | 0.20952 | 2.5e-05 | 0 | 100.8836 | 100.9298 | 127.4549 | 21.2341 | 0.002133 | 0.002 | -152.524 | 0.23684 | 8.403 | 0.003238 | 0.0048 | 0.00456 | 5e-05 | 0.001026 | 0.29192 | 1.331 | 2.5482 | 101.19 | 20.81 | 17.994 | 73.1277 | 99.7205 | 69.9279 | 0.004376 | 1.0291 | 0.025356 | 24.6398 | -0.004087 | 0.14089 | -5.4e-05 | -0.004169 | -0.004169 | -1.406
2019-03-19 16:02 | 1553025692.7801 | 1163 | 0.20948 | 3e-05 | -1e-06 | 100.9893 | 101.1637 | 126.9998 | 21.2345 | 0.002099 | 0.001952 | -152.5305 | 0.23766 | 8.4024 | 0.003276 | 0.004504 | 0.004383 | 5.5e-05 | 0.001024 | 0.30182 | 1.3563 | 2.5481 | 101.19 | 20.81 | 17.995 | 72.3018 | 100.6302 | 68.7626 | 0.004376 | 1.0326 | 0.025353 | 24.554 | -0.015433 | 0.14138 | 0.000206 | 0.015802 | 0.015802 | 0.37219
2019-03-19 16:02 | 1553025693.7792 | 1164 | 0.20954 | 2.7e-05 | 0 | 101.0238 | 101.089 | 127.1319 | 21.2336 | 0.002115 | 0.001973 | -152.5252 | 0.23686 | 8.4036 | 0.003304 | 0.004857 | 0.004512 | 4.7e-05 | 0.001024 | 0.30879 | 1.3132 | 2.548 | 101.19 | 20.81 | 17.999 | 72.6907 | 101.3085 | 69.2655 | 0.004377 | 1.031 | 0.025348 | 24.5848 | -0.005773 | 0.14091 | -0.000163 | -0.012458 | -0.012458 | -0.47054
2019-03-19 16:02 | 1553025694.7805 | 1165 | 0.20951 | 2.4e-05 | 0 | 101.1147 | 101.1126 | 127.5469 | 21.2348 | 0.002075 | 0.002004 | -152.5389 | 0.23679 | 8.4034 | 0.003327 | 0.0043 | 0.004164 | 5e-05 | 0.001027 | 0.29286 | 1.3018 | 2.548 | 101.19 | 20.81 | 18 | 71.736 | 101.8631 | 70.02 | 0.004377 | 1.0288 | 0.025361 | 24.652 | 0.000184 | 0.14087 | 5.1e-05 | 0.003894 | 0.003894 | 1.5048
2019-03-19 16:02 | 1553025695.7793 | 1166 | 0.2095 | 2.3e-05 | 0 | 101.0466 | 101.2678 | 127.6218 | 21.2338 | 0.00213 | 0.00195 | -152.5339 | 0.23751 | 8.4019 | 0.003357 | 0.004662 | 0.004476 | 5.2e-05 | 0.001029 | 0.3027 | 1.3317 | 2.5481 | 101.19 | 20.81 | 18.005 | 73.0337 | 102.6076 | 68.7174 | 0.004379 | 1.0327 | 0.025417 | 24.6124 | -0.019558 | 0.1413 | 5.6e-05 | 0.004267 | 0.004267 | 1.3775
2019-03-19 16:02 | 1553025696.7793 | 1167 | 0.20951 | 2.6e-05 | 0 | 101.1609 | 101.1188 | 127.3744 | 21.234 | 0.002079 | 0.001997 | -152.5362 | 0.23697 | 8.4032 | 0.003299 | 0.004252 | 0.004075 | 5.5e-05 | 0.001034 | 0.30685 | 1.3247 | 2.548 | 101.19 | 20.82 | 18.007 | 71.8303 | 101.1997 | 69.8514 | 0.004382 | 1.0293 | 0.02535 | 24.6293 | 0.003726 | 0.14098 | 0 | -9e-06 | -9e-06 | -642.8038
2019-03-19 16:02 | 1553025697.7797 | 1168 | 0.20951 | 3e-05 | 0 | 101.1675 | 101.324 | 127.1369 | 21.2332 | 0.002083 | 0.001995 | -152.5421 | 0.23654 | 8.4047 | 0.003294 | 0.00448 | 0.004312 | 4.9e-05 | 0.001035 | 0.29679 | 1.3312 | 2.5481 | 101.19 | 20.82 | 18.008 | 71.9224 | 101.064 | 69.8071 | 0.004382 | 1.0294 | 0.025328 | 24.6047 | -0.013819 | 0.14072 | 8e-06 | 0.000669 | 0.000669 | 8.7551
2019-03-19 16:02 | 1553025698.7793 | 1169 | 0.20952 | 3.4e-05 | 0 | 101.2042 | 101.3062 | 127.316 | 21.2335 | 0.002097 | 0.00198 | -152.5325 | 0.23755 | 8.4029 | 0.003298 | 0.00446 | 0.004201 | 4.9e-05 | 0.001038 | 0.31542 | 1.3329 | 2.5481 | 101.19 | 20.82 | 18.01 | 72.2547 | 101.1622 | 69.4416 | 0.004383 | 1.0305 | 0.02536 | 24.6089 | -0.00901 | 0.14132 | -7.4e-05 | -0.005606 | -0.005606 | -1.0487
2019-03-19 16:02 | 1553025699.7805 | 1170 | 0.20951 | 2.3e-05 | -1e-06 | 101.266 | 101.45 | 127.4568 | 21.2334 | 0.002116 | 0.00199 | -152.5454 | 0.23716 | 8.4027 | 0.003337 | 0.004989 | 0.004796 | 5e-05 | 0.001034 | 0.3027 | 1.3281 | 2.548 | 101.19 | 20.82 | 18.01 | 72.6967 | 102.1045 | 69.6754 | 0.004383 | 1.0298 | 0.025365 | 24.6309 | -0.01624 | 0.14108 | 8e-06 | 0.000644 | 0.000644 | 9.1093
2019-03-19 16:02 | 1553025700.7794 | 1171 | 0.20951 | 2.9e-05 | 0 | 101.2591 | 101.3103 | 127.3937 | 21.2327 | 0.002128 | 0.001972 | -152.5511 | 0.23653 | 8.4032 | 0.003293 | 0.005146 | 0.004886 | 5e-05 | 0.001027 | 0.31116 | 1.3113 | 2.548 | 101.19 | 20.82 | 18.013 | 73.0006 | 101.039 | 69.2483 | 0.004383 | 1.0311 | 0.025375 | 24.6095 | -0.00452 | 0.14071 | 3.1e-05 | 0.002407 | 0.002407 | 2.4319
2019-03-19 16:02 | 1553025701.7799 | 1172 | 0.20952 | 3e-05 | 1e-06 | 101.2744 | 101.3452 | 127.2348 | 21.2319 | 0.002132 | 0.002004 | -152.5425 | 0.23775 | 8.4044 | 0.00323 | 0.004604 | 0.004518 | 5e-05 | 0.001026 | 0.32549 | 1.3637 | 2.5481 | 101.19 | 20.82 | 18.015 | 73.0895 | 99.5324 | 70.0102 | 0.004384 | 1.0288 | 0.025331 | 24.6215 | -0.006252 | 0.14144 | -4.5e-05 | -0.003519 | -0.003519 | -1.6721
2019-03-19 16:02 | 1553025702.7809 | 1173 | 0.20948 | 2.8e-05 | 0 | 101.3398 | 101.4462 | 127.5643 | 21.2331 | 0.002068 | 0.001947 | -152.5487 | 0.23671 | 8.4018 | 0.003213 | 0.004104 | 0.003826 | 5e-05 | 0.001025 | 0.27992 | 1.3497 | 2.5481 | 101.19 | 20.82 | 18.018 | 71.555 | 99.1124 | 68.6434 | 0.004385 | 1.0329 | 0.025414 | 24.6042 | -0.009389 | 0.14082 | 0.0002 | 0.01534 | 0.01534 | 0.38188
2019-03-19 16:02 | 1553025703.7797 | 1174 | 0.20951 | 3.2e-05 | 0 | 101.3074 | 101.5408 | 127.3683 | 21.2334 | 0.002094 | 0.001965 | -152.5442 | 0.23676 | 8.4056 | 0.003281 | 0.004405 | 0.004253 | 4.8e-05 | 0.001021 | 0.30989 | 1.365 | 2.548 | 101.19 | 20.82 | 18.018 | 72.1813 | 100.7673 | 69.0867 | 0.004385 | 1.0316 | 0.025378 | 24.6012 | -0.020581 | 0.14085 | 4.1e-05 | 0.003143 | 0.003143 | 1.864
2019-03-19 16:02 | 1553025704.7797 | 1175 | 0.20951 | 2.9e-05 | -1e-06 | 101.4538 | 101.5897 | 127.1488 | 21.2326 | 0.002077 | 0.00198 | -152.5211 | 0.2365 | 8.4034 | 0.003146 | 0.004388 | 0.004168 | 5.3e-05 | 0.001023 | 0.31166 | 1.3631 | 2.5481 | 101.19 | 20.82 | 18.017 | 71.7644 | 97.4844 | 69.4393 | 0.004384 | 1.0305 | 0.025343 | 24.5927 | -0.011973 | 0.1407 | 4.9e-05 | 0.003755 | 0.003755 | 1.5588
2019-03-19 16:02 | 1553025705.7799 | 1176 | 0.2095 | 2.5e-05 | 0 | 101.4154 | 101.5716 | 127.081 | 21.2317 | 0.002109 | 0.001999 | -152.516 | 0.23635 | 8.4025 | 0.00319 | 0.004765 | 0.004543 | 5e-05 | 0.00102 | 0.32156 | 1.3696 | 2.5482 | 101.19 | 20.83 | 18.016 | 72.5446 | 98.5472 | 69.9022 | 0.004387 | 1.0291 | 0.025319 | 24.6027 | -0.013762 | 0.1406 | 9.2e-05 | 0.007045 | 0.007045 | 0.8303
2019-03-19 16:02 | 1553025706.7802 | 1177 | 0.2095 | 2e-05 | 0 | 101.4235 | 101.6069 | 127.4096 | 21.2324 | 0.002069 | 0.001986 | -152.5287 | 0.23714 | 8.4046 | 0.003248 | 0.004422 | 0.004233 | 5.3e-05 | 0.001022 | 0.29446 | 1.3553 | 2.5481 | 101.19 | 20.83 | 18.015 | 71.5825 | 99.9611 | 69.584 | 0.004387 | 1.0301 | 0.025364 | 24.6231 | -0.016153 | 0.14108 | 9.9e-05 | 0.007568 | 0.007568 | 0.77549
2019-03-19 16:02 | 1553025707.7789 | 1178 | 0.20947 | 2.2e-05 | -1e-06 | 101.4466 | 101.6034 | 127.6272 | 21.2318 | 0.002133 | 0.001957 | -152.5368 | 0.23721 | 8.4032 | 0.003213 | 0.004525 | 0.004364 | 5e-05 | 0.001024 | 0.30879 | 1.3301 | 2.5481 | 101.19 | 20.83 | 18.011 | 73.1098 | 99.122 | 68.8738 | 0.004386 | 1.0322 | 0.025412 | 24.6186 | -0.013814 | 0.14111 | 0.00028 | 0.021499 | 0.021499 | 0.27305
2019-03-19 16:02 | 1553025708.7793 | 1179 | 0.20954 | 2.3e-05 | 0 | 101.4613 | 101.4586 | 127.6706 | 21.2322 | 0.002086 | 0.001931 | -152.5453 | 0.23679 | 8.4037 | 0.003186 | 0.005207 | 0.004992 | 5.3e-05 | 0.001022 | 0.32084 | 1.3584 | 2.5481 | 101.19 | 20.83 | 18.009 | 71.9891 | 98.4709 | 68.2588 | 0.004385 | 1.0341 | 0.025439 | 24.6006 | 0.000245 | 0.14087 | -0.000164 | -0.012645 | -0.012645 | -0.46344
2019-03-19 16:02 | 1553025709.7804 | 1180 | 0.20952 | 1.9e-05 | 0 | 101.4998 | 101.5709 | 127.4397 | 21.2306 | 0.002107 | 0.001932 | -152.5432 | 0.23692 | 8.4022 | 0.003129 | 0.004568 | 0.00433 | 5.2e-05 | 0.001011 | 0.31382 | 1.3515 | 2.5481 | 101.19 | 20.83 | 18.007 | 72.4991 | 97.0913 | 68.2876 | 0.004385 | 1.034 | 0.025415 | 24.5794 | -0.006267 | 0.14094 | -3.5e-05 | -0.002669 | -0.002669 | -2.1969
2019-03-19 16:02 | 1553025710.7803 | 1181 | 0.20951 | 3e-05 | 0 | 101.5103 | 101.711 | 127.2162 | 21.231 | 0.002035 | 0.001934 | -152.5332 | 0.23662 | 8.4022 | 0.003192 | 0.004852 | 0.004723 | 4.9e-05 | 0.001006 | 0.31116 | 1.36 | 2.5482 | 101.19 | 20.83 | 18.002 | 70.7554 | 98.5936 | 68.3355 | 0.004383 | 1.0339 | 0.025391 | 24.5595 | -0.017669 | 0.14077 | -1e-06 | -8e-05 | -8e-05 | -72.8601
2019-03-19 16:02 | 1553025711.7791 | 1182 | 0.20955 | 2.3e-05 | 0 | 101.5475 | 101.8377 | 127.1646 | 21.2308 | 0.002018 | 0.001929 | -152.5645 | 0.23774 | 8.4032 | 0.003244 | 0.004053 | 0.003853 | 5.4e-05 | 0.001007 | 0.27151 | 1.3488 | 2.548 | 101.19 | 20.83 | 17.998 | 70.3468 | 99.8641 | 68.2051 | 0.004382 | 1.0342 | 0.025391 | 24.5499 | -0.025527 | 0.14143 | -0.000235 | -0.017975 | -0.017975 | -0.32731
2019-03-19 16:02 | 1553025712.7806 | 1183 | 0.20956 | 1.8e-05 | 0 | 101.4989 | 101.6719 | 126.9113 | 21.2314 | 0.002006 | 0.001925 | -152.5817 | 0.23789 | 8.4008 | 0.003218 | 0.004152 | 0.00397 | 5.4e-05 | 0.001013 | 0.30397 | 1.3386 | 2.5481 | 101.19 | 20.83 | 17.986 | 70.0666 | 99.2409 | 68.1261 | 0.004379 | 1.0345 | 0.025368 | 24.5226 | -0.015239 | 0.14152 | -0.000284 | -0.021726 | -0.021726 | -0.27098
2019-03-19 16:02 | 1553025713.7802 | 1184 | 0.20956 | 2.2e-05 | 0 | 101.5541 | 101.7528 | 127.0552 | 21.2311 | 0.002084 | 0.001904 | -152.5505 | 0.23718 | 8.4033 | 0.003236 | 0.004271 | 0.004008 | 5.1e-05 | 0.001019 | 0.30596 | 1.3457 | 2.5481 | 101.19 | 20.83 | 17.979 | 71.9289 | 99.676 | 67.6111 | 0.004378 | 1.036 | 0.025402 | 24.5179 | -0.017483 | 0.1411 | -0.000268 | -0.020502 | -0.020502 | -0.28631
2019-03-19 16:02 | 1553025714.7793 | 1185 | 0.20951 | 2.5e-05 | 0 | 101.5829 | 101.887 | 127.3462 | 21.2301 | 0.002065 | 0.001951 | -152.5834 | 0.23715 | 8.4019 | 0.00324 | 0.004919 | 0.004743 | 5.1e-05 | 0.001023 | 0.31526 | 1.37 | 2.5481 | 101.19 | 20.83 | 17.965 | 71.4834 | 99.7645 | 68.7443 | 0.004374 | 1.0326 | 0.025389 | 24.5868 | -0.02674 | 0.14108 | 3.5e-05 | 0.002632 | 0.002632 | 2.23
2019-03-19 16:02 | 1553025715.7794 | 1186 | 0.2095 | 2.1e-05 | 0 | 101.5541 | 101.8535 | 127.2674 | 21.2303 | 0.002065 | 0.001966 | -152.5787 | 0.23745 | 8.4023 | 0.003224 | 0.004306 | 0.004088 | 4.9e-05 | 0.001029 | 0.3183 | 1.317 | 2.5479 | 101.19 | 20.83 | 17.956 | 71.4728 | 99.3701 | 69.1 | 0.004372 | 1.0315 | 0.025368 | 24.592 | -0.02633 | 0.14126 | 0.000113 | 0.008677 | 0.008677 | 0.67723
2019-03-19 16:02 | 1553025716.7808 | 1187 | 0.20951 | 2.8e-05 | 0 | 101.6165 | 101.8298 | 127.2256 | 21.2297 | 0.002052 | 0.001922 | -152.5705 | 0.23773 | 8.404 | 0.003183 | 0.004148 | 0.003955 | 5e-05 | 0.001032 | 0.30143 | 1.361 | 2.548 | 101.19 | 20.83 | 17.948 | 71.1652 | 98.3962 | 68.0334 | 0.00437 | 1.0348 | 0.025403 | 24.5496 | -0.018758 | 0.14142 | 6e-06 | 0.000438 | 0.000438 | 13.444
2019-03-19 16:02 | 1553025717.7806 | 1188 | 0.2095 | 2.4e-05 | 0 | 101.5631 | 101.612 | 127.243 | 21.2306 | 0.001998 | 0.00198 | -152.5719 | 0.23671 | 8.4009 | 0.003224 | 0.005069 | 0.004896 | 5.2e-05 | 0.001036 | 0.29175 | 1.3727 | 2.5482 | 101.19 | 20.84 | 17.932 | 69.8661 | 99.3773 | 69.4325 | 0.004369 | 1.0305 | 0.025353 | 24.6015 | -0.004307 | 0.14082 | 7.4e-05 | 0.005639 | 0.005639 | 1.0389
2019-03-19 16:02 | 1553025718.7794 | 1189 | 0.20951 | 2.2e-05 | 0 | 101.6042 | 101.5671 | 127.1209 | 21.2303 | 0.001993 | 0.001938 | -152.5824 | 0.23673 | 8.4029 | 0.003276 | 0.005033 | 0.004824 | 5e-05 | 0.001041 | 0.33284 | 1.3542 | 2.5481 | 101.19 | 20.84 | 17.923 | 69.7581 | 100.6411 | 68.4347 | 0.004367 | 1.0336 | 0.025378 | 24.5539 | 0.003263 | 0.14083 | 2.2e-05 | 0.001696 | 0.001696 | 3.4537
2019-03-19 16:02 | 1553025719.7796 | 1190 | 0.20954 | 2.6e-05 | 0 | 101.5679 | 101.6915 | 127.377 | 21.2304 | 0.00208 | 0.001914 | -152.5881 | 0.23804 | 8.4007 | 0.003221 | 0.004316 | 0.004101 | 5e-05 | 0.00104 | 0.34755 | 1.3402 | 2.5481 | 101.19 | 20.84 | 17.913 | 71.8494 | 99.3042 | 67.8529 | 0.004364 | 1.0353 | 0.025425 | 24.5577 | -0.010876 | 0.14161 | -0.000133 | -0.01021 | -0.01021 | -0.577
2019-03-19 16:02 | 1553025720.78 | 1191 | 0.20954 | 2.2e-05 | 0 | 101.5811 | 101.8202 | 127.512 | 21.2294 | 0.002044 | 0.001887 | -152.6067 | 0.23673 | 8.4036 | 0.003184 | 0.004923 | 0.004746 | 5e-05 | 0.001037 | 0.31078 | 1.3301 | 2.5482 | 101.19 | 20.84 | 17.893 | 70.9659 | 98.4179 | 67.2038 | 0.004359 | 1.0373 | 0.025463 | 24.5473 | -0.021033 | 0.14083 | -0.000133 | -0.010258 | -0.010258 | -0.57113
2019-03-19 16:02 | 1553025721.7787 | 1192 | 0.20956 | 2.3e-05 | 0 | 101.573 | 101.5993 | 127.4786 | 21.2298 | 0.002078 | 0.001961 | -152.5845 | 0.2375 | 8.4036 | 0.003269 | 0.004707 | 0.004506 | 5e-05 | 0.001043 | 0.32444 | 1.3087 | 2.5481 | 101.19 | 20.84 | 17.883 | 71.7818 | 100.4633 | 68.9912 | 0.004357 | 1.0319 | 0.025393 | 24.6084 | -0.00232 | 0.14129 | -0.000248 | -0.019013 | -0.019013 | -0.30914
2019-03-19 16:02 | 1553025722.7787 | 1193 | 0.20956 | 2.5e-05 | 0 | 101.597 | 101.6915 | 127.6164 | 21.2296 | 0.00202 | 0.001932 | -152.5736 | 0.2378 | 8.4026 | 0.003275 | 0.004312 | 0.0041 | 5.2e-05 | 0.001047 | 0.3027 | 1.2906 | 2.5481 | 101.19 | 20.84 | 17.866 | 70.3876 | 100.6051 | 68.2734 | 0.004353 | 1.034 | 0.025433 | 24.5959 | -0.008313 | 0.14147 | -0.000284 | -0.021783 | -0.021783 | -0.27016
2019-03-19 16:02 | 1553025723.7798 | 1194 | 0.20955 | 2.2e-05 | 0 | 101.5865 | 101.7928 | 127.6089 | 21.2303 | 0.002041 | 0.001993 | -152.5912 | 0.23738 | 8.4017 | 0.003234 | 0.004766 | 0.004581 | 4.6e-05 | 0.001044 | 0.32117 | 1.3223 | 2.5481 | 101.19 | 20.84 | 17.858 | 70.9139 | 99.6085 | 69.7387 | 0.004351 | 1.0296 | 0.025378 | 24.6479 | -0.01815 | 0.14122 | -0.00022 | -0.016858 | -0.016858 | -0.34849
2019-03-19 16:02 | 1553025724.7798 | 1195 | 0.20957 | 2.3e-05 | 0 | 101.5799 | 101.8315 | 127.6784 | 21.2299 | 0.00203 | 0.001948 | -152.5665 | 0.23639 | 8.4032 | 0.003227 | 0.004529 | 0.004295 | 5e-05 | 0.001045 | 0.3027 | 1.3426 | 2.5481 | 101.19 | 20.84 | 17.851 | 70.6374 | 99.4512 | 68.6691 | 0.004349 | 1.0328 | 0.025425 | 24.6161 | -0.022132 | 0.14063 | -0.000319 | -0.024476 | -0.024476 | -0.23901
2019-03-19 16:02 | 1553025725.7799 | 1196 | 0.20955 | 3e-05 | 0 | 101.6354 | 101.7363 | 127.8199 | 21.2297 | 0.002038 | 0.001947 | -152.5915 | 0.23869 | 8.4008 | 0.003294 | 0.004919 | 0.004706 | 4.8e-05 | 0.001045 | 0.33555 | 1.318 | 2.5482 | 101.19 | 20.84 | 17.84 | 70.84 | 101.0745 | 68.655 | 0.004347 | 1.0329 | 0.025439 | 24.6293 | -0.008877 | 0.14199 | -0.000242 | -0.018575 | -0.018575 | -0.318
2019-03-19 16:02 | 1553025726.7792 | 1197 | 0.20958 | 2.5e-05 | 0 | 101.5172 | 101.8767 | 127.6544 | 21.2303 | 0.002083 | 0.00191 | -152.5853 | 0.23719 | 8.4029 | 0.003187 | 0.004228 | 0.004161 | 5.1e-05 | 0.001044 | 0.29374 | 1.3255 | 2.5481 | 101.19 | 20.84 | 17.834 | 71.9156 | 98.4902 | 67.7451 | 0.004345 | 1.0356 | 0.025457 | 24.5805 | -0.031626 | 0.14111 | -0.000377 | -0.028964 | -0.028964 | -0.20267
2019-03-19 16:02 | 1553025727.78 | 1198 | 0.20955 | 2.5e-05 | 0 | 101.6003 | 101.8161 | 127.4718 | 21.2301 | 0.002098 | 0.002005 | -152.5791 | 0.23593 | 8.4035 | 0.003241 | 0.004579 | 0.004396 | 5.1e-05 | 0.001036 | 0.29574 | 1.3033 | 2.5481 | 101.19 | 20.83 | 17.829 | 72.2663 | 99.7865 | 70.0324 | 0.004341 | 1.0287 | 0.025353 | 24.6452 | -0.018981 | 0.14036 | -0.000229 | -0.017579 | -0.017579 | -0.33216
2019-03-19 16:02 | 1553025728.779 | 1199 | 0.20956 | 2.6e-05 | 0 | 101.5463 | 101.8791 | 127.5131 | 21.231 | 0.00203 | 0.001912 | -152.5824 | 0.23707 | 8.4049 | 0.003289 | 0.004134 | 0.0039 | 5e-05 | 0.001032 | 0.32748 | 1.3388 | 2.548 | 101.19 | 20.83 | 17.819 | 70.6392 | 100.9471 | 67.8146 | 0.004339 | 1.0354 | 0.02544 | 24.5694 | -0.029272 | 0.14103 | -0.000278 | -0.021351 | -0.021351 | -0.27479
2019-03-19 16:02 | 1553025729.7791 | 1200 | 0.20953 | 2.5e-05 | 0 | 101.5544 | 101.686 | 127.5779 | 21.23 | 0.002041 | 0.001915 | -152.5904 | 0.23678 | 8.405 | 0.003294 | 0.004412 | 0.004183 | 5.1e-05 | 0.001034 | 0.29927 | 1.3619 | 2.5482 | 101.19 | 20.83 | 17.815 | 70.9119 | 101.0591 | 67.872 | 0.004338 | 1.0353 | 0.025444 | 24.5777 | -0.011583 | 0.14086 | -0.00012 | -0.009177 | -0.009177 | -0.63849
2019-03-19 16:02 | 1553025730.7808 | 1201 | 0.20955 | 2.3e-05 | 0 | 101.5598 | 101.7216 | 127.9011 | 21.2294 | 0.00203 | 0.001932 | -152.5928 | 0.23719 | 8.4027 | 0.003194 | 0.003972 | 0.003781 | 5.2e-05 | 0.001037 | 0.29518 | 1.3993 | 2.548 | 101.19 | 20.83 | 17.809 | 70.6417 | 98.6439 | 68.2833 | 0.004336 | 1.034 | 0.025461 | 24.6237 | -0.01424 | 0.14111 | -0.00022 | -0.016882 | -0.016882 | -0.34771
2019-03-19 16:02 | 1553025731.7792 | 1202 | 0.20955 | 2.1e-05 | 0 | 101.5922 | 101.7127 | 127.8777 | 21.2289 | 0.002001 | 0.001911 | -152.5823 | 0.237 | 8.403 | 0.003276 | 0.004462 | 0.004232 | 4.9e-05 | 0.001044 | 0.3027 | 1.3906 | 2.5479 | 101.19 | 20.83 | 17.806 | 69.9284 | 100.631 | 67.7649 | 0.004336 | 1.0356 | 0.025478 | 24.6027 | -0.010603 | 0.14099 | -0.000203 | -0.015588 | -0.015588 | -0.37626
2019-03-19 16:02 | 1553025732.7792 | 1203 | 0.20951 | 2.4e-05 | 0 | 101.5781 | 101.7024 | 127.8542 | 21.2302 | 0.002016 | 0.001905 | -152.5733 | 0.23725 | 8.4029 | 0.003323 | 0.004703 | 0.004507 | 5.1e-05 | 0.001044 | 0.32084 | 1.3655 | 2.5483 | 101.19 | 20.83 | 17.805 | 70.3007 | 101.7729 | 67.6328 | 0.004335 | 1.036 | 0.025481 | 24.5957 | -0.010941 | 0.14114 | 5.2e-05 | 0.003937 | 0.003937 | 1.4915
2019-03-19 16:02 | 1553025733.7793 | 1204 | 0.20949 | 2.7e-05 | 0 | 101.5466 | 101.8096 | 127.7842 | 21.2302 | 0.002007 | 0.001921 | -152.5882 | 0.23707 | 8.4013 | 0.003264 | 0.004935 | 0.004719 | 4.9e-05 | 0.001042 | 0.28495 | 1.3463 | 2.5481 | 101.19 | 20.83 | 17.804 | 70.0912 | 100.3537 | 68.0283 | 0.004335 | 1.0348 | 0.025459 | 24.6032 | -0.02314 | 0.14103 | 0.000158 | 0.012159 | 0.012159 | 0.48254
2019-03-19 16:02 | 1553025734.7794 | 1205 | 0.20949 | 2.3e-05 | 0 | 101.5166 | 101.7189 | 127.842 | 21.2293 | 0.001979 | 0.001906 | -152.5867 | 0.2363 | 8.4043 | 0.003233 | 0.004567 | 0.004304 | 5.1e-05 | 0.00105 | 0.31349 | 1.3037 | 2.548 | 101.19 | 20.83 | 17.804 | 69.4157 | 99.5872 | 67.6507 | 0.004335 | 1.0359 | 0.025479 | 24.5952 | -0.017806 | 0.14057 | 0.000184 | 0.014146 | 0.014146 | 0.41339
2019-03-19 16:02 | 1553025735.7794 | 1206 | 0.20953 | 2.3e-05 | 0 | 101.5415 | 101.6164 | 128.0924 | 21.2301 | 0.002007 | 0.001854 | -152.579 | 0.23641 | 8.4018 | 0.003195 | 0.00371 | 0.003577 | 5e-05 | 0.001046 | 0.28263 | 1.3429 | 2.5481 | 101.19 | 20.83 | 17.803 | 70.0872 | 98.6836 | 66.4181 | 0.004335 | 1.0397 | 0.02555 | 24.5747 | -0.006602 | 0.14064 | -6.8e-05 | -0.005297 | -0.005297 | -1.1046
2019-03-19 16:02 | 1553025736.7789 | 1207 | 0.20954 | 2.4e-05 | 0 | 101.4844 | 101.5623 | 128.1788 | 21.2294 | 0.002007 | 0.001894 | -152.5892 | 0.23671 | 8.406 | 0.003242 | 0.004222 | 0.003945 | 4.8e-05 | 0.001054 | 0.29695 | 1.3078 | 2.5481 | 101.19 | 20.83 | 17.803 | 70.0779 | 99.8093 | 67.3762 | 0.004335 | 1.0368 | 0.025523 | 24.6176 | -0.006862 | 0.14082 | -0.000158 | -0.012168 | -0.012168 | -0.48144
2019-03-19 16:02 | 1553025737.7795 | 1208 | 0.20953 | 2e-05 | 0 | 101.5895 | 101.8117 | 128.23 | 21.2298 | 0.001983 | 0.001847 | -152.5892 | 0.23666 | 8.4018 | 0.003194 | 0.00471 | 0.004491 | 4.7e-05 | 0.001048 | 0.2912 | 1.291 | 2.548 | 101.19 | 20.84 | 17.803 | 69.5004 | 98.641 | 66.2398 | 0.004338 | 1.0402 | 0.02557 | 24.5814 | -0.019541 | 0.14079 | -5.7e-05 | -0.004395 | -0.004395 | -1.3325
2019-03-19 16:02 | 1553025738.7786 | 1209 | 0.20952 | 2.2e-05 | 0 | 101.4998 | 101.5873 | 128.487 | 21.2292 | 0.001937 | 0.001918 | -152.5964 | 0.23614 | 8.4035 | 0.00329 | 0.004089 | 0.00392 | 4.4e-05 | 0.001046 | 0.29789 | 1.3199 | 2.5481 | 101.19 | 20.84 | 17.803 | 68.4102 | 100.9635 | 67.9494 | 0.004338 | 1.035 | 0.025532 | 24.6679 | -0.007715 | 0.14048 | -1.3e-05 | -0.001026 | -0.001026 | -5.6952
2019-03-19 16:02 | 1553025739.7796 | 1210 | 0.20954 | 2.2e-05 | 0 | 101.5289 | 101.6517 | 128.6335 | 21.2297 | 0.001975 | 0.001952 | -152.5677 | 0.23656 | 8.4023 | 0.003235 | 0.004879 | 0.004703 | 4.9e-05 | 0.001049 | 0.31636 | 1.3342 | 2.548 | 101.19 | 20.83 | 17.8 | 69.3117 | 99.6425 | 68.7728 | 0.004334 | 1.0325 | 0.025516 | 24.7118 | -0.010817 | 0.14073 | -0.000123 | -0.009469 | -0.009469 | -0.61827
2019-03-19 16:02 | 1553025740.779 | 1211 | 0.20954 | 2.4e-05 | 0 | 101.4889 | 101.7123 | 128.3044 | 21.2285 | 0.002031 | 0.001895 | -152.5797 | 0.23642 | 8.4023 | 0.003206 | 0.005027 | 0.004758 | 4.7e-05 | 0.001052 | 0.30143 | 1.3 | 2.5482 | 101.19 | 20.83 | 17.796 | 70.6641 | 98.9368 | 67.3854 | 0.004333 | 1.0367 | 0.025535 | 24.63 | -0.019667 | 0.14065 | -0.000154 | -0.011866 | -0.011866 | -0.49306
2019-03-19 16:02 | 1553025741.7797 | 1212 | 0.20955 | 2.8e-05 | 0 | 101.5595 | 101.6877 | 128.0389 | 21.229 | 0.002049 | 0.001891 | -152.6017 | 0.23722 | 8.4045 | 0.003249 | 0.004895 | 0.004686 | 4.8e-05 | 0.001052 | 0.28224 | 1.3252 | 2.5481 | 101.19 | 20.83 | 17.793 | 71.1059 | 99.9714 | 67.2885 | 0.004332 | 1.037 | 0.025512 | 24.601 | -0.011285 | 0.14112 | -0.000187 | -0.014451 | -0.014451 | -0.40625
2019-03-19 16:02 | 1553025742.7798 | 1213 | 0.20955 | 2.4e-05 | 0 | 101.4901 | 101.7456 | 128.1276 | 21.2291 | 0.002083 | 0.001843 | -152.6023 | 0.23719 | 8.403 | 0.003268 | 0.004702 | 0.004494 | 5.1e-05 | 0.001048 | 0.31669 | 1.3537 | 2.5481 | 101.19 | 20.83 | 17.791 | 71.9176 | 100.45 | 66.1501 | 0.004332 | 1.0405 | 0.025564 | 24.5684 | -0.022482 | 0.14111 | -0.000183 | -0.014127 | -0.014127 | -0.41553
2019-03-19 16:02 | 1553025743.7799 | 1214 | 0.20954 | 2.2e-05 | 0 | 101.3965 | 101.4863 | 128.2654 | 21.229 | 0.001979 | 0.001889 | -152.5868 | 0.23723 | 8.4025 | 0.003247 | 0.003696 | 0.003422 | 4.4e-05 | 0.001044 | 0.32908 | 1.3754 | 2.5481 | 101.19 | 20.83 | 17.789 | 69.4163 | 99.9406 | 67.2423 | 0.004331 | 1.0372 | 0.025536 | 24.6211 | -0.007916 | 0.14113 | -0.000121 | -0.009348 | -0.009348 | -0.62805
2019-03-19 16:02 | 1553025744.7801 | 1215 | 0.20953 | 2.7e-05 | 0 | 101.4394 | 101.5908 | 128.2478 | 21.2287 | 0.001995 | 0.001877 | -152.5856 | 0.23602 | 8.4052 | 0.003253 | 0.00456 | 0.004424 | 4.5e-05 | 0.001034 | 0.3298 | 1.3507 | 2.5482 | 101.19 | 20.83 | 17.788 | 69.802 | 100.0878 | 66.9635 | 0.004331 | 1.038 | 0.025545 | 24.6093 | -0.013334 | 0.14041 | -0.000105 | -0.008131 | -0.008131 | -0.71836
2019-03-19 16:02 | 1553025745.7799 | 1216 | 0.20956 | 2.9e-05 | 0 | 101.4328 | 101.598 | 128.0072 | 21.2287 | 0.002076 | 0.001887 | -152.5975 | 0.23644 | 8.4035 | 0.003226 | 0.004244 | 0.003996 | 5e-05 | 0.001031 | 0.28512 | 1.4035 | 2.5481 | 101.19 | 20.83 | 17.789 | 71.7543 | 99.4164 | 67.1983 | 0.004331 | 1.0373 | 0.025512 | 24.5947 | -0.014549 | 0.14066 | -0.000274 | -0.021086 | -0.021086 | -0.2775
2019-03-19 16:02 | 1553025746.7794 | 1217 | 0.20955 | 2.6e-05 | 1e-06 | 101.4478 | 101.5456 | 127.9063 | 21.2293 | 0.002042 | 0.001924 | -152.6018 | 0.23791 | 8.4048 | 0.003267 | 0.00491 | 0.004608 | 4.8e-05 | 0.001027 | 0.31222 | 1.3586 | 2.5481 | 101.19 | 20.83 | 17.79 | 70.9331 | 100.4094 | 68.093 | 0.004332 | 1.0346 | 0.025469 | 24.6173 | -0.008611 | 0.14153 | -0.000215 | -0.01659 | -0.01659 | -0.3549
2019-03-19 16:02 | 1553025747.7801 | 1218 | 0.20948 | 2.1e-05 | 0 | 101.4391 | 101.6599 | 127.6857 | 21.2283 | 0.001982 | 0.001901 | -152.6093 | 0.23708 | 8.4039 | 0.003311 | 0.004509 | 0.004308 | 5e-05 | 0.001029 | 0.29087 | 1.3369 | 2.5481 | 101.19 | 20.83 | 17.791 | 69.4911 | 101.4792 | 67.5451 | 0.004332 | 1.0362 | 0.025467 | 24.5763 | -0.019449 | 0.14104 | 0.000193 | 0.014833 | 0.014833 | 0.39555
2019-03-19 16:02 | 1553025748.7791 | 1219 | 0.20955 | 3.2e-05 | 0 | 101.4904 | 101.5695 | 127.6265 | 21.2289 | 0.00201 | 0.00189 | -152.6005 | 0.23691 | 8.4006 | 0.003275 | 0.004862 | 0.004676 | 4.8e-05 | 0.001033 | 0.31421 | 1.3333 | 2.5481 | 101.19 | 20.83 | 17.794 | 70.1646 | 100.6194 | 67.2704 | 0.004333 | 1.0371 | 0.025472 | 24.5607 | -0.006966 | 0.14094 | -0.000202 | -0.015548 | -0.015548 | -0.37709
2019-03-19 16:02 | 1553025749.7792 | 1220 | 0.20954 | 2.5e-05 | 0 | 101.5205 | 101.7459 | 127.4474 | 21.2287 | 0.001991 | 0.001879 | -152.6212 | 0.23696 | 8.4024 | 0.003233 | 0.004532 | 0.004229 | 5.3e-05 | 0.001036 | 0.28783 | 1.3091 | 2.5481 | 101.19 | 20.83 | 17.795 | 69.6968 | 99.5943 | 67.0189 | 0.004333 | 1.0379 | 0.025463 | 24.5344 | -0.019841 | 0.14097 | -0.000166 | -0.012738 | -0.012738 | -0.46038
2019-03-19 16:02 | 1553025750.7796 | 1221 | 0.2095 | 2.3e-05 | 0 | 101.4742 | 101.726 | 127.1427 | 21.2273 | 0.001984 | 0.00188 | -152.5981 | 0.23705 | 8.4021 | 0.003187 | 0.004466 | 0.004318 | 4.7e-05 | 0.001044 | 0.32045 | 1.326 | 2.548 | 101.19 | 20.82 | 17.796 | 69.5259 | 98.488 | 67.0317 | 0.004331 | 1.0378 | 0.025432 | 24.5055 | -0.022168 | 0.14102 | 7.9e-05 | 0.006037 | 0.006037 | 0.97177
2019-03-19 16:03 | 1553025751.7792 | 1222 | 0.20954 | 2.7e-05 | 0 | 101.4796 | 101.8637 | 127.5493 | 21.2271 | 0.001973 | 0.001852 | -152.5906 | 0.23787 | 8.4004 | 0.003297 | 0.004889 | 0.004628 | 4.6e-05 | 0.001047 | 0.29911 | 1.3172 | 2.5481 | 101.19 | 20.82 | 17.799 | 69.2623 | 101.1482 | 66.3613 | 0.004331 | 1.0399 | 0.025498 | 24.5205 | -0.033789 | 0.14151 | -0.00014 | -0.010784 | -0.010784 | -0.54588
2019-03-19 16:03 | 1553025752.7793 | 1223 | 0.20953 | 2.6e-05 | 0 | 101.4409 | 101.5956 | 127.6211 | 21.2274 | 0.001952 | 0.001895 | -152.5982 | 0.23722 | 8.4013 | 0.003271 | 0.004821 | 0.004738 | 4.5e-05 | 0.001047 | 0.29214 | 1.3816 | 2.5481 | 101.19 | 20.82 | 17.801 | 68.7713 | 100.512 | 67.3812 | 0.004332 | 1.0367 | 0.025467 | 24.5642 | -0.013623 | 0.14112 | -0.000107 | -0.008221 | -0.008221 | -0.71413
2019-03-19 16:03 | 1553025753.7787 | 1224 | 0.20952 | 2.1e-05 | 0 | 101.4586 | 101.5048 | 127.8608 | 21.2273 | 0.00206 | 0.001924 | -152.5895 | 0.23774 | 8.4036 | 0.003279 | 0.004386 | 0.004102 | 4.6e-05 | 0.001051 | 0.30055 | 1.3741 | 2.548 | 101.19 | 20.82 | 17.803 | 71.3671 | 100.7135 | 68.0863 | 0.004332 | 1.0346 | 0.025464 | 24.6127 | -0.004068 | 0.14143 | -3.6e-05 | -0.002772 | -0.002772 | -2.1227
2019-03-19 16:03 | 1553025754.779 | 1225 | 0.20951 | 2.9e-05 | 0 | 101.4682 | 101.5027 | 127.7333 | 21.2274 | 0.002009 | 0.001876 | -152.6042 | 0.23708 | 8.4015 | 0.003208 | 0.004385 | 0.004245 | 4.7e-05 | 0.001056 | 0.28423 | 1.3586 | 2.5482 | 101.19 | 20.82 | 17.803 | 70.1226 | 98.9993 | 66.9296 | 0.004332 | 1.0381 | 0.025495 | 24.5587 | -0.00304 | 0.14104 | 2.1e-05 | 0.001627 | 0.001627 | 3.6058
2019-03-19 16:03 | 1553025755.7785 | 1226 | 0.2095 | 2e-05 | 0 | 101.495 | 101.8596 | 127.7605 | 21.2268 | 0.001957 | 0.001855 | -152.5944 | 0.23755 | 8.4004 | 0.003205 | 0.004131 | 0.00385 | 4.5e-05 | 0.001051 | 0.27887 | 1.3106 | 2.5481 | 101.19 | 20.82 | 17.8 | 68.8829 | 98.9204 | 66.4382 | 0.004331 | 1.0396 | 0.025516 | 24.5436 | -0.03208 | 0.14132 | 0.000114 | 0.008769 | 0.008769 | 0.6704
2019-03-19 16:03 | 1553025756.7795 | 1227 | 0.2095 | 2.6e-05 | 1e-06 | 101.4938 | 101.6349 | 127.6772 | 21.2259 | 0.001955 | 0.001853 | -152.5979 | 0.23762 | 8.3998 | 0.00329 | 0.004882 | 0.00463 | 4.6e-05 | 0.001051 | 0.29463 | 1.3616 | 2.5481 | 101.19 | 20.82 | 17.793 | 68.8431 | 100.9617 | 66.3702 | 0.00433 | 1.0398 | 0.02551 | 24.5331 | -0.012434 | 0.14136 | 7.1e-05 | 0.005454 | 0.005454 | 1.0783
2019-03-19 16:03 | 1553025757.7795 | 1228 | 0.20949 | 2.7e-05 | 0 | 101.4505 | 101.4418 | 127.8943 | 21.2273 | 0.001972 | 0.001826 | -152.6008 | 0.23713 | 8.403 | 0.003278 | 0.00469 | 0.004472 | 4.9e-05 | 0.001057 | 0.30486 | 1.3314 | 2.5482 | 101.19 | 20.82 | 17.788 | 69.2382 | 100.6718 | 65.7339 | 0.004329 | 1.0418 | 0.025556 | 24.5309 | 0.000772 | 0.14107 | 0.00016 | 0.012347 | 0.012347 | 0.47531
2019-03-19 16:03 | 1553025758.779 | 1229 | 0.20947 | 2.7e-05 | 0 | 101.4553 | 101.8267 | 127.919 | 21.2273 | 0.001954 | 0.001869 | -152.6063 | 0.23732 | 8.4037 | 0.003288 | 0.005203 | 0.005078 | 4.7e-05 | 0.001063 | 0.33069 | 1.3724 | 2.5481 | 101.19 | 20.82 | 17.777 | 68.8001 | 100.9254 | 66.7648 | 0.004326 | 1.0386 | 0.02552 | 24.5706 | -0.032684 | 0.14118 | 0.000265 | 0.020394 | 0.020394 | 0.28798
2019-03-19 16:03 | 1553025759.7787 | 1230 | 0.20948 | 2e-05 | 0 | 101.4016 | 101.6473 | 128.018 | 21.2266 | 0.001927 | 0.001893 | -152.6069 | 0.23719 | 8.404 | 0.003297 | 0.004532 | 0.004342 | 4.4e-05 | 0.001052 | 0.29999 | 1.3593 | 2.5481 | 101.19 | 20.82 | 17.77 | 68.1503 | 101.1376 | 67.3328 | 0.004324 | 1.0369 | 0.025508 | 24.6006 | -0.021643 | 0.14111 | 0.000228 | 0.017533 | 0.017533 | 0.33481
2019-03-19 16:03 | 1553025760.7791 | 1231 | 0.20949 | 2.4e-05 | 0 | 101.4727 | 101.7719 | 128.0521 | 21.2262 | 0.001981 | 0.001822 | -152.6101 | 0.2381 | 8.4047 | 0.003224 | 0.00507 | 0.004876 | 4.5e-05 | 0.001048 | 0.30021 | 1.3143 | 2.548 | 101.19 | 20.82 | 17.763 | 69.4525 | 99.3801 | 65.6439 | 0.004322 | 1.0421 | 0.025575 | 24.5428 | -0.026334 | 0.14164 | 0.000127 | 0.009775 | 0.009775 | 0.60282
2019-03-19 16:03 | 1553025761.7802 | 1232 | 0.20946 | 2.5e-05 | 0 | 101.431 | 101.7682 | 128.1579 | 21.2266 | 0.001959 | 0.001867 | -152.5984 | 0.23798 | 8.4039 | 0.003181 | 0.00467 | 0.004455 | 4.7e-05 | 0.001039 | 0.27959 | 1.3222 | 2.548 | 101.19 | 20.82 | 17.748 | 68.9241 | 98.3407 | 66.7221 | 0.004319 | 1.0388 | 0.025545 | 24.592 | -0.029681 | 0.14157 | 0.000325 | 0.025043 | 0.025043 | 0.23517
2019-03-19 16:03 | 1553025762.7789 | 1233 | 0.20953 | 2.6e-05 | 0 | 101.4878 | 101.6332 | 127.8997 | 21.2263 | 0.001948 | 0.001859 | -152.5736 | 0.23675 | 8.4037 | 0.003262 | 0.004738 | 0.004531 | 4.6e-05 | 0.001035 | 0.28102 | 1.2969 | 2.5478 | 101.19 | 20.82 | 17.742 | 68.6726 | 100.2872 | 66.5229 | 0.004317 | 1.0394 | 0.025527 | 24.56 | -0.012812 | 0.14084 | -6.9e-05 | -0.00531 | -0.00531 | -1.1034
2019-03-19 16:03 | 1553025763.7788 | 1234 | 0.2095 | 2.2e-05 | 0 | 101.4637 | 101.7897 | 127.6004 | 21.2266 | 0.00199 | 0.001894 | -152.5729 | 0.23727 | 8.4028 | 0.00319 | 0.004838 | 0.004569 | 4.6e-05 | 0.001042 | 0.30773 | 1.3083 | 2.5481 | 101.19 | 20.82 | 17.736 | 69.6705 | 98.5626 | 67.3555 | 0.004316 | 1.0368 | 0.025466 | 24.5613 | -0.028693 | 0.14115 | 0.000125 | 0.009617 | 0.009617 | 0.6106
2019-03-19 16:03 | 1553025764.7789 | 1235 | 0.20953 | 2.7e-05 | 0 | 101.5328 | 101.7017 | 127.4875 | 21.2253 | 0.001961 | 0.001894 | -152.5701 | 0.23694 | 8.4013 | 0.003212 | 0.004545 | 0.004345 | 4.8e-05 | 0.001048 | 0.29502 | 1.3066 | 2.548 | 101.19 | 20.82 | 17.728 | 68.9868 | 99.0797 | 67.3753 | 0.004314 | 1.0368 | 0.025454 | 24.5511 | -0.014872 | 0.14095 | -9.6e-05 | -0.007377 | -0.007377 | -0.79483
2019-03-19 16:03 | 1553025765.7793 | 1236 | 0.20952 | 2.7e-05 | 0 | 101.531 | 101.8185 | 127.5265 | 21.2255 | 0.00204 | 0.001912 | -152.5598 | 0.23673 | 8.4029 | 0.003204 | 0.004597 | 0.004458 | 5e-05 | 0.001048 | 0.28008 | 1.2996 | 2.5479 | 101.19 | 20.82 | 17.725 | 70.8811 | 98.8863 | 67.7929 | 0.004313 | 1.0355 | 0.025442 | 24.5699 | -0.025297 | 0.14083 | -5.8e-05 | -0.004484 | -0.004484 | -1.3065
2019-03-19 16:03 | 1553025766.779 | 1237 | 0.20951 | 2.6e-05 | 0 | 101.51 | 101.8339 | 127.8694 | 21.2261 | 0.002001 | 0.001906 | -152.5595 | 0.23705 | 8.4039 | 0.003202 | 0.004554 | 0.004339 | 4.9e-05 | 0.001044 | 0.29214 | 1.3217 | 2.5481 | 101.19 | 20.81 | 17.726 | 69.9483 | 98.8437 | 67.6563 | 0.004311 | 1.0359 | 0.025481 | 24.598 | -0.028501 | 0.14102 | 4.5e-05 | 0.003463 | 0.003463 | 1.6939
2019-03-19 16:03 | 1553025767.7794 | 1238 | 0.20951 | 2.5e-05 | 0 | 101.4941 | 101.726 | 127.8976 | 21.2252 | 0.001943 | 0.001859 | -152.5568 | 0.23685 | 8.4025 | 0.003299 | 0.00466 | 0.004452 | 5e-05 | 0.001042 | 0.27887 | 1.335 | 2.5481 | 101.19 | 20.81 | 17.726 | 68.553 | 101.1943 | 66.5319 | 0.004311 | 1.0393 | 0.025526 | 24.5601 | -0.020422 | 0.1409 | 5.8e-05 | 0.004499 | 0.004499 | 1.3027
2019-03-19 16:03 | 1553025768.7792 | 1239 | 0.20953 | 2.6e-05 | 0 | 101.5124 | 101.8202 | 127.7619 | 21.2258 | 0.001978 | 0.001891 | -152.5635 | 0.23737 | 8.4012 | 0.003248 | 0.004256 | 0.004025 | 4.8e-05 | 0.001043 | 0.28263 | 1.3568 | 2.5482 | 101.19 | 20.81 | 17.727 | 69.3903 | 99.9414 | 67.2986 | 0.004311 | 1.037 | 0.025484 | 24.5748 | -0.027086 | 0.14121 | -9.8e-05 | -0.007558 | -0.007558 | -0.77724
2019-03-19 16:03 | 1553025769.7792 | 1240 | 0.20955 | 2.5e-05 | 1e-06 | 101.4953 | 101.6897 | 127.65 | 21.2242 | 0.001965 | 0.001894 | -152.5494 | 0.23776 | 8.4007 | 0.003148 | 0.004592 | 0.004416 | 5e-05 | 0.001049 | 0.31349 | 1.3618 | 2.5481 | 101.19 | 20.81 | 17.731 | 69.0771 | 97.524 | 67.3572 | 0.004312 | 1.0368 | 0.025471 | 24.5661 | -0.017123 | 0.14144 | -0.000177 | -0.013689 | -0.013689 | -0.42984
2019-03-19 16:03 | 1553025770.7793 | 1241 | 0.20955 | 2.5e-05 | 0 | 101.5874 | 101.6288 | 127.8558 | 21.2247 | 0.001944 | 0.001884 | -152.5471 | 0.23671 | 8.4028 | 0.003215 | 0.00433 | 0.004096 | 5e-05 | 0.001045 | 0.32587 | 1.3734 | 2.5482 | 101.19 | 20.81 | 17.732 | 68.5737 | 99.1473 | 67.1305 | 0.004312 | 1.0375 | 0.0255 | 24.5777 | -0.003644 | 0.14082 | -0.000197 | -0.015112 | -0.015112 | -0.38764
2019-03-19 16:03 | 1553025771.7796 | 1242 | 0.20958 | 2.7e-05 | 0 | 101.5403 | 101.7733 | 127.6159 | 21.2247 | 0.001958 | 0.001898 | -152.5605 | 0.23705 | 8.4043 | 0.003237 | 0.004182 | 0.003976 | 4.9e-05 | 0.001047 | 0.27472 | 1.3793 | 2.5482 | 101.19 | 20.81 | 17.739 | 68.9074 | 99.6767 | 67.4661 | 0.004314 | 1.0365 | 0.025463 | 24.5668 | -0.020505 | 0.14102 | -0.000363 | -0.027883 | -0.027883 | -0.21039
2019-03-19 16:03 | 1553025772.7783 | 1243 | 0.2095 | 2e-05 | 0 | 101.5337 | 101.8329 | 127.7204 | 21.225 | 0.001935 | 0.001849 | -152.5789 | 0.23824 | 8.4046 | 0.003186 | 0.004649 | 0.004411 | 4.7e-05 | 0.001049 | 0.31636 | 1.382 | 2.5481 | 101.19 | 20.81 | 17.741 | 68.3586 | 98.453 | 66.2797 | 0.004314 | 1.0401 | 0.025518 | 24.534 | -0.026322 | 0.14173 | 0.000106 | 0.008147 | 0.008147 | 0.72372
2019-03-19 16:03 | 1553025773.7794 | 1244 | 0.20959 | 2.1e-05 | 0 | 101.5169 | 101.9254 | 127.7108 | 21.2246 | 0.001922 | 0.00187 | -152.5873 | 0.23812 | 8.4017 | 0.003332 | 0.004927 | 0.004755 | 5.2e-05 | 0.001049 | 0.31205 | 1.4034 | 2.5482 | 101.19 | 20.81 | 17.743 | 68.0291 | 101.992 | 66.7778 | 0.004315 | 1.0386 | 0.025498 | 24.551 | -0.035922 | 0.14166 | -0.000445 | -0.034235 | -0.034235 | -0.17213
2019-03-19 16:03 | 1553025774.7795 | 1245 | 0.20955 | 2.6e-05 | 0 | 101.5526 | 101.6973 | 127.6645 | 21.224 | 0.001939 | 0.001832 | -152.5862 | 0.23755 | 8.4046 | 0.003236 | 0.004703 | 0.004457 | 5e-05 | 0.001051 | 0.28119 | 1.3756 | 2.5482 | 101.19 | 20.81 | 17.745 | 68.4372 | 99.6557 | 65.8767 | 0.004315 | 1.0413 | 0.025528 | 24.514 | -0.012736 | 0.14132 | -0.000229 | -0.017622 | -0.017622 | -0.3336
2019-03-19 16:03 | 1553025775.7799 | 1246 | 0.20956 | 2.6e-05 | 0 | 101.5553 | 101.6795 | 127.5669 | 21.2241 | 0.001993 | 0.001853 | -152.5997 | 0.23825 | 8.4021 | 0.003226 | 0.004761 | 0.004542 | 5.1e-05 | 0.00105 | 0.3079 | 1.3588 | 2.5482 | 101.19 | 20.81 | 17.745 | 69.7438 | 99.4119 | 66.3818 | 0.004315 | 1.0398 | 0.025499 | 24.5229 | -0.010932 | 0.14174 | -0.000291 | -0.022392 | -0.022392 | -0.26332
2019-03-19 16:03 | 1553025776.7796 | 1247 | 0.20954 | 2.8e-05 | 0 | 101.5451 | 101.5836 | 127.504 | 21.2245 | 0.001966 | 0.001872 | -152.5934 | 0.23789 | 8.4038 | 0.003274 | 0.004592 | 0.004397 | 5.1e-05 | 0.001052 | 0.28816 | 1.331 | 2.5481 | 101.19 | 20.81 | 17.745 | 69.0905 | 100.589 | 66.8409 | 0.004315 | 1.0384 | 0.025475 | 24.5334 | -0.003391 | 0.14152 | -0.000137 | -0.010538 | -0.010538 | -0.55866
2019-03-19 16:03 | 1553025777.7787 | 1248 | 0.20955 | 2.8e-05 | 0 | 101.5412 | 101.7815 | 127.6612 | 21.2241 | 0.001978 | 0.001889 | -152.5945 | 0.23719 | 8.4038 | 0.003313 | 0.004954 | 0.004704 | 4.8e-05 | 0.001049 | 0.30574 | 1.3106 | 2.548 | 101.19 | 20.81 | 17.743 | 69.3803 | 101.5174 | 67.2477 | 0.004315 | 1.0372 | 0.025476 | 24.5632 | -0.021148 | 0.1411 | -0.000188 | -0.014421 | -0.014421 | -0.40705
2019-03-19 16:03 | 1553025778.7796 | 1249 | 0.20954 | 2.9e-05 | 0 | 101.5556 | 101.5993 | 127.8622 | 21.2242 | 0.001988 | 0.001836 | -152.5829 | 0.23727 | 8.3996 | 0.003254 | 0.004131 | 0.003933 | 5.1e-05 | 0.001041 | 0.31078 | 1.3269 | 2.548 | 101.19 | 20.81 | 17.739 | 69.6236 | 100.0943 | 65.9659 | 0.004314 | 1.0411 | 0.025544 | 24.5362 | -0.003852 | 0.14115 | -0.000127 | -0.009784 | -0.009784 | -0.60016
2019-03-19 16:03 | 1553025779.7791 | 1250 | 0.20955 | 2.4e-05 | 1e-06 | 101.5592 | 101.748 | 127.7244 | 21.2235 | 0.001968 | 0.001827 | -152.5957 | 0.23759 | 8.4043 | 0.003254 | 0.004804 | 0.004599 | 5.2e-05 | 0.001042 | 0.30309 | 1.3095 | 2.5481 | 101.19 | 20.81 | 17.733 | 69.1449 | 100.1 | 65.7536 | 0.004313 | 1.0417 | 0.025538 | 24.5153 | -0.016612 | 0.14134 | -0.000196 | -0.015137 | -0.015137 | -0.38845
2019-03-19 16:03 | 1553025780.7788 | 1251 | 0.20955 | 2.5e-05 | 0 | 101.6273 | 101.8148 | 127.677 | 21.2239 | 0.001978 | 0.001838 | -152.6095 | 0.23769 | 8.4026 | 0.003301 | 0.005139 | 0.004959 | 5.1e-05 | 0.001042 | 0.30646 | 1.3022 | 2.548 | 101.19 | 20.81 | 17.73 | 69.3707 | 101.2329 | 66.0265 | 0.004312 | 1.0409 | 0.025523 | 24.5206 | -0.016483 | 0.1414 | -0.000193 | -0.014911 | -0.014911 | -0.39449
2019-03-19 16:03 | 1553025781.7799 | 1252 | 0.20952 | 2.8e-05 | 0 | 101.5385 | 101.8168 | 127.7169 | 21.2235 | 0.001902 | 0.001849 | -152.6049 | 0.23813 | 8.4009 | 0.003323 | 0.004722 | 0.004547 | 5.1e-05 | 0.001038 | 0.29894 | 1.317 | 2.5482 | 101.19 | 20.81 | 17.723 | 67.5585 | 101.7601 | 66.2789 | 0.00431 | 1.0401 | 0.025518 | 24.5336 | -0.024485 | 0.14167 | -2.8e-05 | -0.002217 | -0.002217 | -2.6586
2019-03-19 16:03 | 1553025782.7778 | 1253 | 0.20952 | 2.3e-05 | 0 | 101.5787 | 101.8034 | 128.0863 | 21.2231 | 0.001959 | 0.001868 | -152.6167 | 0.2382 | 8.4057 | 0.003277 | 0.004445 | 0.004346 | 5.3e-05 | 0.00103 | 0.26737 | 1.3503 | 2.5481 | 101.19 | 20.81 | 17.719 | 68.9284 | 100.6623 | 66.7294 | 0.004309 | 1.0387 | 0.025538 | 24.5854 | -0.01977 | 0.14171 | -4e-06 | -0.000294 | -0.000294 | -20.0378
2019-03-19 16:03 | 1553025783.7783 | 1254 | 0.20955 | 2.8e-05 | 1e-06 | 101.5592 | 101.651 | 127.7068 | 21.2237 | 0.002 | 0.001841 | -152.6299 | 0.23726 | 8.4042 | 0.003291 | 0.004542 | 0.004303 | 5.1e-05 | 0.001027 | 0.28871 | 1.3463 | 2.5481 | 101.19 | 20.81 | 17.716 | 69.9031 | 100.9941 | 66.0778 | 0.004308 | 1.0407 | 0.025524 | 24.5254 | -0.008087 | 0.14115 | -0.000176 | -0.01363 | -0.01363 | -0.43078
2019-03-19 16:03 | 1553025784.779 | 1255 | 0.20953 | 2.8e-05 | 1e-06 | 101.5454 | 101.849 | 127.5913 | 21.2226 | 0.001941 | 0.00187 | -152.6286 | 0.23796 | 8.4017 | 0.003341 | 0.005075 | 0.004814 | 5.1e-05 | 0.001021 | 0.30558 | 1.4197 | 2.5481 | 101.19 | 20.8 | 17.707 | 68.4959 | 102.2108 | 66.7814 | 0.004304 | 1.0386 | 0.025486 | 24.5397 | -0.026705 | 0.14156 | -7e-05 | -0.005443 | -0.005443 | -1.0819
2019-03-19 16:03 | 1553025785.7791 | 1256 | 0.20954 | 3.3e-05 | 0 | 101.5517 | 101.7233 | 127.7863 | 21.2225 | 0.001937 | 0.001845 | -152.6347 | 0.23784 | 8.4033 | 0.003396 | 0.005069 | 0.00489 | 5.2e-05 | 0.001025 | 0.27616 | 1.355 | 2.5481 | 101.19 | 20.8 | 17.703 | 68.3983 | 103.5221 | 66.1835 | 0.004303 | 1.0404 | 0.025528 | 24.5368 | -0.015104 | 0.14149 | -0.000127 | -0.009832 | -0.009832 | -0.59865
2019-03-19 16:03 | 1553025786.7794 | 1257 | 0.20953 | 2.9e-05 | 0 | 101.6075 | 101.8397 | 127.6495 | 21.2232 | 0.001947 | 0.001883 | -152.6187 | 0.23725 | 8.4 | 0.003261 | 0.005064 | 0.004855 | 5.2e-05 | 0.001027 | 0.33069 | 1.3751 | 2.5481 | 101.19 | 20.8 | 17.699 | 68.6383 | 100.266 | 67.1025 | 0.004302 | 1.0376 | 0.02548 | 24.5569 | -0.020423 | 0.14114 | -0.000104 | -0.008004 | -0.008004 | -0.73356
2019-03-19 16:03 | 1553025787.7796 | 1258 | 0.20953 | 2.5e-05 | 1e-06 | 101.5445 | 101.7569 | 127.5458 | 21.2231 | 0.002013 | 0.00182 | -152.6091 | 0.2368 | 8.4042 | 0.003284 | 0.004759 | 0.004578 | 5.4e-05 | 0.001031 | 0.31061 | 1.3599 | 2.5481 | 101.19 | 20.8 | 17.69 | 70.234 | 100.8153 | 65.5849 | 0.004299 | 1.0422 | 0.025527 | 24.4921 | -0.01869 | 0.14088 | -6e-05 | -0.004642 | -0.004642 | -1.2626
2019-03-19 16:03 | 1553025788.7793 | 1259 | 0.20953 | 3.1e-05 | 0 | 101.5946 | 101.7808 | 127.1852 | 21.2232 | 0.00198 | 0.001911 | -152.6275 | 0.23747 | 8.4038 | 0.003315 | 0.004763 | 0.004528 | 5.1e-05 | 0.001033 | 0.34506 | 1.3469 | 2.5481 | 101.19 | 20.8 | 17.687 | 69.4402 | 101.577 | 67.7712 | 0.004299 | 1.0356 | 0.025409 | 24.5362 | -0.016383 | 0.14127 | -6.9e-05 | -0.005268 | -0.005268 | -1.1156
2019-03-19 16:03 | 1553025789.7786 | 1260 | 0.20953 | 2.8e-05 | 0 | 101.5433 | 101.7182 | 127.4894 | 21.2237 | 0.001913 | 0.001839 | -152.6226 | 0.23742 | 8.4043 | 0.00326 | 0.004493 | 0.004294 | 5.3e-05 | 0.001035 | 0.28224 | 1.3365 | 2.5483 | 101.19 | 20.8 | 17.687 | 67.8255 | 100.2322 | 66.0419 | 0.004299 | 1.0408 | 0.025504 | 24.5032 | -0.015393 | 0.14124 | -9.4e-05 | -0.007236 | -0.007236 | -0.81195
2019-03-19 16:03 | 1553025790.7787 | 1261 | 0.20953 | 2.7e-05 | 0 | 101.588 | 101.7582 | 127.5265 | 21.224 | 0.002005 | 0.001875 | -152.607 | 0.23698 | 8.405 | 0.003248 | 0.005253 | 0.005035 | 5.2e-05 | 0.001035 | 0.29031 | 1.3395 | 2.548 | 101.19 | 20.8 | 17.687 | 70.0275 | 99.9578 | 66.9107 | 0.004299 | 1.0382 | 0.025475 | 24.5381 | -0.014978 | 0.14098 | -0.000107 | -0.008204 | -0.008204 | -0.71483
2019-03-19 16:03 | 1553025791.7783 | 1262 | 0.20952 | 2.7e-05 | 0 | 101.5976 | 101.7158 | 127.5106 | 21.2231 | 0.001904 | 0.001845 | -152.6056 | 0.23737 | 8.4037 | 0.003256 | 0.00479 | 0.004607 | 5.1e-05 | 0.00104 | 0.30342 | 1.3352 | 2.5481 | 101.19 | 20.8 | 17.688 | 67.5987 | 100.138 | 66.1724 | 0.004299 | 1.0404 | 0.025501 | 24.5099 | -0.010399 | 0.14121 | -1e-06 | -0.000131 | -0.000131 | -44.7455
2019-03-19 16:03 | 1553025792.7783 | 1263 | 0.20951 | 2.4e-05 | 0 | 101.6006 | 101.712 | 127.0634 | 21.2234 | 0.001963 | 0.001934 | -152.6094 | 0.2379 | 8.4035 | 0.003299 | 0.00531 | 0.005159 | 5.3e-05 | 0.00104 | 0.29839 | 1.3334 | 2.5481 | 101.19 | 20.8 | 17.691 | 69.0162 | 101.194 | 68.3307 | 0.0043 | 1.0339 | 0.025376 | 24.5446 | -0.009803 | 0.14153 | 2.7e-05 | 0.002087 | 0.002087 | 2.8208
2019-03-19 16:03 | 1553025793.7784 | 1264 | 0.20952 | 2.4e-05 | 0 | 101.6231 | 101.9223 | 127.5122 | 21.2221 | 0.001878 | 0.001877 | -152.5994 | 0.23774 | 8.4016 | 0.003283 | 0.004388 | 0.004128 | 5.2e-05 | 0.001029 | 0.27273 | 1.3171 | 2.548 | 101.19 | 20.8 | 17.696 | 66.9798 | 100.8002 | 66.9596 | 0.004301 | 1.038 | 0.025472 | 24.5385 | -0.026296 | 0.14143 | 1.3e-05 | 0.000949 | 0.000949 | 6.1968
2019-03-19 16:03 | 1553025794.7786 | 1265 | 0.20949 | 2.7e-05 | 0 | 101.6204 | 101.8397 | 127.442 | 21.2241 | 0.001976 | 0.001805 | -152.5867 | 0.23649 | 8.4025 | 0.003249 | 0.004791 | 0.004672 | 5.2e-05 | 0.001031 | 0.29822 | 1.3472 | 2.5483 | 101.19 | 20.8 | 17.699 | 69.3404 | 99.983 | 65.2243 | 0.004302 | 1.0434 | 0.02553 | 24.4691 | -0.019287 | 0.14069 | 0.000134 | 0.0103 | 0.0103 | 0.56822
2019-03-19 16:03 | 1553025795.7785 | 1266 | 0.20949 | 2.7e-05 | 0 | 101.5634 | 101.6319 | 127.115 | 21.2241 | 0.00201 | 0.001882 | -152.5636 | 0.23736 | 8.4022 | 0.003184 | 0.004283 | 0.004083 | 5.1e-05 | 0.001035 | 0.32316 | 1.3206 | 2.5482 | 101.19 | 20.79 | 17.702 | 70.1607 | 98.3987 | 67.0832 | 0.0043 | 1.0377 | 0.025427 | 24.5047 | -0.006029 | 0.1412 | 0.00014 | 0.010769 | 0.010769 | 0.54547
2019-03-19 16:03 | 1553025796.7786 | 1267 | 0.20953 | 3.1e-05 | 0 | 101.6072 | 101.8072 | 126.9681 | 21.2236 | 0.001945 | 0.001896 | -152.5938 | 0.23739 | 8.4037 | 0.003277 | 0.004385 | 0.00418 | 5.1e-05 | 0.001037 | 0.29175 | 1.2993 | 2.5482 | 101.19 | 20.79 | 17.713 | 68.5998 | 100.6558 | 67.408 | 0.004302 | 1.0367 | 0.025401 | 24.5022 | -0.017591 | 0.14122 | -0.000118 | -0.009092 | -0.009092 | -0.64614
2019-03-19 16:03 | 1553025797.7797 | 1268 | 0.20949 | 2.7e-05 | 0 | 101.6114 | 101.9692 | 127.0169 | 21.2233 | 0.002002 | 0.001883 | -152.5848 | 0.23687 | 8.4025 | 0.003324 | 0.004447 | 0.004245 | 5e-05 | 0.001033 | 0.29479 | 1.3101 | 2.548 | 101.19 | 20.79 | 17.718 | 69.9542 | 101.786 | 67.087 | 0.004304 | 1.0376 | 0.025418 | 24.4954 | -0.031443 | 0.14091 | 0.00019 | 0.014568 | 0.014568 | 0.40238
2019-03-19 16:03 | 1553025798.7778 | 1269 | 0.20952 | 2.7e-05 | 0 | 101.5178 | 101.7175 | 127.285 | 21.2234 | 0.001943 | 0.001862 | -152.5613 | 0.23757 | 8.4047 | 0.003329 | 0.004864 | 0.004675 | 5.1e-05 | 0.001035 | 0.30541 | 1.3449 | 2.5481 | 101.19 | 20.79 | 17.722 | 68.5326 | 101.9108 | 66.598 | 0.004305 | 1.0391 | 0.025463 | 24.5036 | -0.01758 | 0.14133 | -3.1e-05 | -0.002396 | -0.002396 | -2.4535
2019-03-19 16:03 | 1553025799.7778 | 1270 | 0.20954 | 2e-05 | 0 | 101.6231 | 101.7189 | 127.2087 | 21.2237 | 0.001933 | 0.001836 | -152.5792 | 0.23716 | 8.4047 | 0.003297 | 0.004481 | 0.004286 | 5e-05 | 0.001029 | 0.31437 | 1.3632 | 2.5481 | 101.19 | 20.79 | 17.728 | 68.2965 | 101.1393 | 65.9717 | 0.004306 | 1.0411 | 0.025479 | 24.4737 | -0.008424 | 0.14109 | -0.000127 | -0.009793 | -0.009793 | -0.59933
2019-03-19 16:03 | 1553025800.7778 | 1271 | 0.20955 | 2.3e-05 | 0 | 101.5721 | 101.7387 | 127.3629 | 21.2226 | 0.001961 | 0.001815 | -152.5839 | 0.2379 | 8.401 | 0.003345 | 0.004207 | 0.004009 | 5.4e-05 | 0.001032 | 0.3246 | 1.3473 | 2.5481 | 101.19 | 20.79 | 17.73 | 68.9796 | 102.3082 | 65.4718 | 0.004306 | 1.0426 | 0.025513 | 24.4704 | -0.014662 | 0.14153 | -0.000214 | -0.016436 | -0.016436 | -0.35822
2019-03-19 16:03 | 1553025801.7789 | 1272 | 0.20955 | 2.9e-05 | 1e-06 | 101.5739 | 101.8267 | 127.3671 | 21.2228 | 0.001918 | 0.001833 | -152.5742 | 0.23796 | 8.4027 | 0.003303 | 0.005126 | 0.004862 | 5.5e-05 | 0.001039 | 0.29103 | 1.3848 | 2.5481 | 101.19 | 20.79 | 17.729 | 67.9478 | 101.2805 | 65.8944 | 0.004306 | 1.0413 | 0.025497 | 24.4861 | -0.022239 | 0.14156 | -0.000175 | -0.013548 | -0.013548 | -0.43468
2019-03-19 16:03 | 1553025802.7789 | 1273 | 0.20952 | 2.3e-05 | 0 | 101.5997 | 101.8678 | 127.7249 | 21.2227 | 0.001904 | 0.001831 | -152.5792 | 0.23771 | 8.4008 | 0.0033 | 0.005554 | 0.005403 | 5.2e-05 | 0.001045 | 0.33483 | 1.3635 | 2.5482 | 101.19 | 20.79 | 17.725 | 67.6077 | 101.2115 | 65.8412 | 0.004305 | 1.0415 | 0.025535 | 24.5185 | -0.023576 | 0.14141 | 1.3e-05 | 0.000982 | 0.000982 | 5.9925
2019-03-19 16:03 | 1553025803.7791 | 1274 | 0.20953 | 2.4e-05 | 0 | 101.6036 | 101.8137 | 127.7824 | 21.2226 | 0.00198 | 0.00185 | -152.5914 | 0.23811 | 8.4023 | 0.003257 | 0.004827 | 0.004618 | 5.4e-05 | 0.001039 | 0.31581 | 1.3679 | 2.5481 | 101.19 | 20.79 | 17.721 | 69.4303 | 100.162 | 66.3024 | 0.004304 | 1.04 | 0.025523 | 24.5407 | -0.01848 | 0.14165 | -5.7e-05 | -0.004406 | -0.004406 | -1.3375
2019-03-19 16:03 | 1553025804.7784 | 1275 | 0.20957 | 2.4e-05 | 0 | 101.5931 | 101.76 | 127.7868 | 21.2231 | 0.00194 | 0.001837 | -152.6148 | 0.23779 | 8.4028 | 0.003283 | 0.005083 | 0.004934 | 5.5e-05 | 0.001042 | 0.29446 | 1.3552 | 2.5479 | 101.19 | 20.79 | 17.71 | 68.459 | 100.7855 | 65.9938 | 0.004302 | 1.041 | 0.025536 | 24.53 | -0.014679 | 0.14146 | -0.000335 | -0.025803 | -0.025803 | -0.22806
2019-03-19 16:03 | 1553025805.7792 | 1276 | 0.20953 | 2.5e-05 | 0 | 101.5622 | 101.7189 | 127.7624 | 21.2235 | 0.00199 | 0.0018 | -152.6197 | 0.23771 | 8.4021 | 0.003257 | 0.004263 | 0.004111 | 5.6e-05 | 0.001047 | 0.31404 | 1.3273 | 2.548 | 101.19 | 20.79 | 17.706 | 69.6761 | 100.1684 | 65.1106 | 0.004301 | 1.0437 | 0.025566 | 24.4957 | -0.013788 | 0.14141 | -8.4e-05 | -0.006526 | -0.006526 | -0.9014
2019-03-19 16:03 | 1553025806.7792 | 1277 | 0.20952 | 2.5e-05 | 1e-06 | 101.5754 | 101.8175 | 127.7612 | 21.2223 | 0.001939 | 0.001819 | -152.609 | 0.23769 | 8.4027 | 0.003296 | 0.00487 | 0.004563 | 5.2e-05 | 0.001042 | 0.30431 | 1.3016 | 2.5481 | 101.19 | 20.79 | 17.7 | 68.4449 | 101.1058 | 65.5653 | 0.004299 | 1.0423 | 0.025549 | 24.5121 | -0.021294 | 0.1414 | -4e-06 | -0.000373 | -0.000373 | -15.7873
2019-03-19 16:03 | 1553025807.7796 | 1278 | 0.20953 | 2.6e-05 | 1e-06 | 101.6594 | 101.8894 | 128.0309 | 21.2232 | 0.001958 | 0.001849 | -152.6289 | 0.23797 | 8.4014 | 0.00334 | 0.00518 | 0.00494 | 5.7e-05 | 0.001043 | 0.30917 | 1.3382 | 2.548 | 101.19 | 20.79 | 17.698 | 68.8989 | 102.1705 | 66.2738 | 0.004299 | 1.0401 | 0.025549 | 24.5636 | -0.020214 | 0.14157 | -0.000102 | -0.00789 | -0.00789 | -0.74644
2019-03-19 16:03 | 1553025808.7786 | 1279 | 0.20952 | 2.5e-05 | 0 | 101.6048 | 101.7326 | 127.7655 | 21.2226 | 0.001972 | 0.001889 | -152.631 | 0.23694 | 8.4042 | 0.003319 | 0.004946 | 0.004765 | 5.6e-05 | 0.001038 | 0.2896 | 1.3169 | 2.548 | 101.19 | 20.79 | 17.696 | 69.2368 | 101.6674 | 67.2447 | 0.004298 | 1.0372 | 0.025486 | 24.5732 | -0.01124 | 0.14095 | -1.3e-05 | -0.000993 | -0.000993 | -5.9071
2019-03-19 16:03 | 1553025809.7783 | 1280 | 0.20952 | 2.7e-05 | 0 | 101.6105 | 101.7993 | 127.3692 | 21.223 | 0.001943 | 0.001864 | -152.6179 | 0.23762 | 8.4031 | 0.003279 | 0.004279 | 0.004097 | 5.4e-05 | 0.001034 | 0.33931 | 1.3304 | 2.548 | 101.19 | 20.79 | 17.697 | 68.5401 | 100.6952 | 66.6351 | 0.004298 | 1.039 | 0.02547 | 24.513 | -0.016608 | 0.14136 | -3e-06 | -0.000224 | -0.000224 | -26.2325
2019-03-19 16:03 | 1553025810.7794 | 1281 | 0.2095 | 2.8e-05 | 1e-06 | 101.6411 | 101.6986 | 126.9836 | 21.2215 | 0.001936 | 0.001853 | -152.6235 | 0.23716 | 8.4015 | 0.003336 | 0.004396 | 0.004271 | 5.7e-05 | 0.001034 | 0.33948 | 1.3384 | 2.5481 | 101.19 | 20.79 | 17.7 | 68.3662 | 102.0779 | 66.3705 | 0.004299 | 1.0398 | 0.025441 | 24.4664 | -0.005062 | 0.14109 | 0.00011 | 0.008398 | 0.008398 | 0.69891
2019-03-19 16:04 | 1553025811.7784 | 1282 | 0.20952 | 2.6e-05 | 0 | 101.6009 | 101.8617 | 127.2157 | 21.2222 | 0.002004 | 0.001802 | -152.6342 | 0.23814 | 8.407 | 0.003294 | 0.004583 | 0.004347 | 5.8e-05 | 0.001026 | 0.29391 | 1.3076 | 2.548 | 101.19 | 20.79 | 17.704 | 70.0046 | 101.0616 | 65.1584 | 0.0043 | 1.0436 | 0.02551 | 24.445 | -0.022929 | 0.14167 | -1.4e-05 | -0.001055 | -0.001055 | -5.5861
2019-03-19 16:04 | 1553025812.7795 | 1283 | 0.20948 | 2.4e-05 | 0 | 101.6258 | 101.8308 | 127.2256 | 21.2222 | 0.001887 | 0.001919 | -152.6169 | 0.2374 | 8.4027 | 0.003285 | 0.004626 | 0.004437 | 5.7e-05 | 0.00103 | 0.29446 | 1.3118 | 2.5481 | 101.19 | 20.79 | 17.714 | 67.2 | 100.8459 | 67.9566 | 0.004303 | 1.035 | 0.025406 | 24.5468 | -0.018029 | 0.14123 | 0.00021 | 0.016082 | 0.016082 | 0.36531
2019-03-19 16:04 | 1553025813.7789 | 1284 | 0.2095 | 2.3e-05 | 0 | 101.5514 | 101.7086 | 127.3035 | 21.2206 | 0.001987 | 0.00187 | -152.6115 | 0.23723 | 8.4037 | 0.003326 | 0.005074 | 0.00485 | 5.8e-05 | 0.001031 | 0.28744 | 1.3437 | 2.5482 | 101.19 | 20.79 | 17.719 | 69.6069 | 101.8279 | 66.784 | 0.004304 | 1.0386 | 0.025458 | 24.5121 | -0.013836 | 0.14113 | 9.3e-05 | 0.007119 | 0.007119 | 0.82462
2019-03-19 16:04 | 1553025814.7786 | 1285 | 0.20952 | 2.1e-05 | 1e-06 | 101.5637 | 101.7993 | 127.4944 | 21.2212 | 0.001884 | 0.001829 | -152.5974 | 0.23716 | 8.4015 | 0.003255 | 0.004108 | 0.003857 | 5.7e-05 | 0.00103 | 0.28744 | 1.3847 | 2.5482 | 101.19 | 20.79 | 17.731 | 67.1249 | 100.1214 | 65.8077 | 0.004307 | 1.0416 | 0.025513 | 24.4952 | -0.02073 | 0.14108 | -3.5e-05 | -0.002745 | -0.002745 | -2.1382
2019-03-19 16:04 | 1553025815.7776 | 1286 | 0.20954 | 2.6e-05 | 0 | 101.5667 | 101.749 | 127.7042 | 21.2207 | 0.001951 | 0.001864 | -152.5969 | 0.2367 | 8.4026 | 0.003286 | 0.004593 | 0.004371 | 5.9e-05 | 0.001028 | 0.29336 | 1.363 | 2.5482 | 101.19 | 20.79 | 17.736 | 68.7374 | 100.8734 | 66.6507 | 0.004308 | 1.039 | 0.025503 | 24.5458 | -0.016041 | 0.14081 | -0.000145 | -0.011201 | -0.011201 | -0.52299
2019-03-19 16:04 | 1553025816.7787 | 1287 | 0.20959 | 2.7e-05 | 0 | 101.5688 | 101.6863 | 127.9225 | 21.2215 | 0.001905 | 0.001853 | -152.5843 | 0.23653 | 8.4024 | 0.003301 | 0.004597 | 0.004316 | 5.8e-05 | 0.001028 | 0.30525 | 1.3616 | 2.5481 | 101.19 | 20.79 | 17.741 | 67.6322 | 101.2322 | 66.3847 | 0.004309 | 1.0398 | 0.025534 | 24.5572 | -0.010345 | 0.14071 | -0.000436 | -0.033602 | -0.033602 | -0.1742
2019-03-19 16:04 | 1553025817.7787 | 1288 | 0.20951 | 2.8e-05 | 0 | 101.5736 | 101.7997 | 127.4584 | 21.2211 | 0.001875 | 0.001841 | -152.5786 | 0.23692 | 8.4036 | 0.003254 | 0.004588 | 0.004394 | 5.9e-05 | 0.001034 | 0.31813 | 1.3261 | 2.5481 | 101.18 | 20.79 | 17.747 | 66.8961 | 100.1 | 66.084 | 0.004311 | 1.0407 | 0.025499 | 24.5017 | -0.019888 | 0.14094 | 2.2e-05 | 0.001686 | 0.001686 | 3.478
2019-03-19 16:04 | 1553025818.7778 | 1289 | 0.20952 | 2.8e-05 | 0 | 101.4769 | 101.7106 | 127.4061 | 21.2216 | 0.00198 | 0.001861 | -152.5904 | 0.23659 | 8.4037 | 0.003339 | 0.00421 | 0.004044 | 5.8e-05 | 0.001037 | 0.28584 | 1.3503 | 2.5482 | 101.18 | 20.79 | 17.749 | 69.42 | 102.1557 | 66.5657 | 0.004312 | 1.0392 | 0.025476 | 24.5141 | -0.020574 | 0.14075 | -3.3e-05 | -0.002545 | -0.002545 | -2.3003
2019-03-19 16:04 | 1553025819.7792 | 1290 | 0.20953 | 2.8e-05 | 0 | 101.5412 | 101.6589 | 127.5913 | 21.2223 | 0.001911 | 0.001866 | -152.5967 | 0.23732 | 8.4041 | 0.003282 | 0.004768 | 0.004556 | 6.1e-05 | 0.001035 | 0.29087 | 1.4044 | 2.5481 | 101.18 | 20.79 | 17.75 | 67.7773 | 100.7697 | 66.6852 | 0.004312 | 1.0389 | 0.02549 | 24.5362 | -0.010366 | 0.14118 | -0.00011 | -0.008508 | -0.008508 | -0.69033
2019-03-19 16:04 | 1553025820.778 | 1291 | 0.20955 | 2.6e-05 | 0 | 101.4853 | 101.6185 | 127.9823 | 21.221 | 0.001964 | 0.001856 | -152.5953 | 0.23623 | 8.4004 | 0.003291 | 0.004348 | 0.004132 | 5.7e-05 | 0.00103 | 0.32875 | 1.341 | 2.548 | 101.18 | 20.79 | 17.75 | 69.0483 | 100.9804 | 66.4381 | 0.004312 | 1.0396 | 0.025538 | 24.5648 | -0.011728 | 0.14053 | -0.000182 | -0.013997 | -0.013997 | -0.41767
2019-03-19 16:04 | 1553025821.7789 | 1292 | 0.20954 | 3.1e-05 | 1e-06 | 101.4878 | 101.7644 | 127.826 | 21.2204 | 0.001973 | 0.00181 | -152.6012 | 0.23724 | 8.4071 | 0.003328 | 0.004134 | 0.003976 | 5.5e-05 | 0.001024 | 0.2995 | 1.3223 | 2.5481 | 101.18 | 20.79 | 17.75 | 69.259 | 101.8939 | 65.3473 | 0.004312 | 1.043 | 0.025564 | 24.5104 | -0.024349 | 0.14113 | -0.000119 | -0.009231 | -0.009231 | -0.63605
2019-03-19 16:04 | 1553025822.7791 | 1293 | 0.20951 | 2.9e-05 | 0 | 101.4913 | 101.6558 | 127.8669 | 21.2205 | 0.001967 | 0.001803 | -152.6236 | 0.23704 | 8.4024 | 0.003282 | 0.004905 | 0.00466 | 5.7e-05 | 0.001026 | 0.29502 | 1.3385 | 2.5482 | 101.18 | 20.79 | 17.748 | 69.1148 | 100.7689 | 65.1647 | 0.004311 | 1.0435 | 0.025575 | 24.5077 | -0.014484 | 0.14101 | 6e-05 | 0.004635 | 0.004635 | 1.2655
2019-03-19 16:04 | 1553025823.7795 | 1294 | 0.20949 | 3.2e-05 | 0 | 101.4712 | 101.7192 | 127.6791 | 21.2197 | 0.00194 | 0.001888 | -152.6405 | 0.23802 | 8.4017 | 0.003313 | 0.004372 | 0.004211 | 5.6e-05 | 0.001033 | 0.28672 | 1.3544 | 2.5483 | 101.18 | 20.79 | 17.748 | 68.4534 | 101.5139 | 67.2107 | 0.004311 | 1.0373 | 0.025479 | 24.5636 | -0.02183 | 0.14159 | 0.000132 | 0.010139 | 0.010139 | 0.58098
2019-03-19 16:04 | 1553025824.7792 | 1295 | 0.20951 | 3e-05 | 0 | 101.4379 | 101.5808 | 127.6075 | 21.22 | 0.001934 | 0.001855 | -152.6308 | 0.23665 | 8.4035 | 0.003312 | 0.00408 | 0.003868 | 5.6e-05 | 0.00104 | 0.29607 | 1.3288 | 2.548 | 101.18 | 20.79 | 17.751 | 68.3203 | 101.5047 | 66.4261 | 0.004312 | 1.0397 | 0.025501 | 24.5284 | -0.012591 | 0.14078 | 5.3e-05 | 0.004098 | 0.004098 | 1.4291
2019-03-19 16:04 | 1553025825.7782 | 1296 | 0.20952 | 3e-05 | 0 | 101.4463 | 101.4733 | 127.2484 | 21.22 | 0.001947 | 0.001794 | -152.6309 | 0.23802 | 8.4027 | 0.003297 | 0.00463 | 0.004385 | 5.7e-05 | 0.001046 | 0.30685 | 1.3447 | 2.548 | 101.18 | 20.79 | 17.755 | 68.6431 | 101.1301 | 64.9479 | 0.004313 | 1.0442 | 0.025521 | 24.4405 | -0.002376 | 0.14159 | -2.1e-05 | -0.001625 | -0.001625 | -3.6247
2019-03-19 16:04 | 1553025826.7782 | 1297 | 0.20952 | 2.8e-05 | 0 | 101.4781 | 101.474 | 126.9402 | 21.219 | 0.001932 | 0.001866 | -152.6484 | 0.23788 | 8.402 | 0.003214 | 0.005081 | 0.004902 | 5.7e-05 | 0.001046 | 0.33229 | 1.3363 | 2.548 | 101.18 | 20.79 | 17.76 | 68.2617 | 99.1323 | 66.6858 | 0.004314 | 1.0389 | 0.025425 | 24.4735 | 0.000368 | 0.14151 | -1.1e-05 | -0.000805 | -0.000805 | -7.3172
2019-03-19 16:04 | 1553025827.7787 | 1298 | 0.20952 | 2.6e-05 | 0 | 101.4415 | 101.5531 | 127.1019 | 21.2191 | 0.001965 | 0.001836 | -152.6482 | 0.23695 | 8.403 | 0.003195 | 0.004192 | 0.004019 | 5.5e-05 | 0.001042 | 0.27201 | 1.3504 | 2.5482 | 101.18 | 20.79 | 17.773 | 69.0728 | 98.669 | 65.9578 | 0.004317 | 1.0411 | 0.025468 | 24.4629 | -0.009831 | 0.14096 | 1.7e-05 | 0.001306 | 0.001306 | 4.4904
2019-03-19 16:04 | 1553025828.7784 | 1299 | 0.20955 | 2.7e-05 | 0 | 101.3974 | 101.5277 | 127.1845 | 21.2193 | 0.001906 | 0.001835 | -152.6505 | 0.23782 | 8.4025 | 0.003162 | 0.004447 | 0.004148 | 5.3e-05 | 0.001042 | 0.33428 | 1.3414 | 2.5481 | 101.18 | 20.79 | 17.781 | 67.6464 | 97.8641 | 65.9518 | 0.004319 | 1.0411 | 0.025477 | 24.4706 | -0.011489 | 0.14148 | -0.000184 | -0.014182 | -0.014182 | -0.41501
2019-03-19 16:04 | 1553025829.7787 | 1300 | 0.20952 | 3e-05 | 0 | 101.422 | 101.4781 | 127.3662 | 21.2195 | 0.001945 | 0.001823 | -152.6512 | 0.23791 | 8.4029 | 0.003151 | 0.004736 | 0.004555 | 5.7e-05 | 0.001047 | 0.29374 | 1.3696 | 2.5481 | 101.18 | 20.79 | 17.79 | 68.5936 | 97.6041 | 65.6624 | 0.004321 | 1.042 | 0.025506 | 24.4776 | -0.004942 | 0.14153 | -5e-05 | -0.003902 | -0.003902 | -1.5091
2019-03-19 16:04 | 1553025830.7778 | 1301 | 0.20951 | 3.4e-05 | 0 | 101.3932 | 101.5493 | 127.3502 | 21.2198 | 0.001962 | 0.001793 | -152.6504 | 0.23701 | 8.4024 | 0.003187 | 0.004541 | 0.004391 | 5.3e-05 | 0.001053 | 0.29286 | 1.3305 | 2.5481 | 101.18 | 20.79 | 17.807 | 68.9845 | 98.4728 | 64.9275 | 0.004326 | 1.0443 | 0.025532 | 24.4495 | -0.01376 | 0.141 | 3.8e-05 | 0.002912 | 0.002912 | 2.0145
2019-03-19 16:04 | 1553025831.7775 | 1302 | 0.20952 | 2.9e-05 | 0 | 101.3152 | 101.3825 | 127.1045 | 21.2193 | 0.001965 | 0.001834 | -152.6579 | 0.23762 | 8.403 | 0.003184 | 0.004741 | 0.004473 | 5.8e-05 | 0.001061 | 0.29391 | 1.3444 | 2.5481 | 101.18 | 20.79 | 17.817 | 69.062 | 98.3964 | 65.922 | 0.004328 | 1.0412 | 0.02547 | 24.4618 | -0.005943 | 0.14136 | -3.9e-05 | -0.003009 | -0.003009 | -1.9545
2019-03-19 16:04 | 1553025832.7789 | 1303 | 0.20949 | 3.1e-05 | 0 | 101.3635 | 101.4942 | 126.5876 | 21.2193 | 0.001927 | 0.001759 | -152.6357 | 0.23797 | 8.404 | 0.003169 | 0.004809 | 0.00464 | 5.7e-05 | 0.001067 | 0.31166 | 1.3128 | 2.5481 | 101.18 | 20.79 | 17.835 | 68.1442 | 98.0468 | 64.1212 | 0.004332 | 1.0468 | 0.025486 | 24.3471 | -0.011523 | 0.14157 | 0.000163 | 0.012519 | 0.012519 | 0.47042
2019-03-19 16:04 | 1553025833.7777 | 1304 | 0.20954 | 2.4e-05 | 0 | 101.3191 | 101.5161 | 126.5494 | 21.2198 | 0.001912 | 0.001825 | -152.6188 | 0.23721 | 8.4009 | 0.003234 | 0.004732 | 0.004474 | 5.4e-05 | 0.001069 | 0.30807 | 1.3357 | 2.5481 | 101.18 | 20.79 | 17.843 | 67.7975 | 99.6052 | 65.7103 | 0.004334 | 1.0419 | 0.025422 | 24.4007 | -0.017375 | 0.14111 | -0.000118 | -0.009073 | -0.009073 | -0.64702
2019-03-19 16:04 | 1553025834.7778 | 1305 | 0.20949 | 3.1e-05 | 0 | 101.2933 | 101.389 | 126.5623 | 21.2191 | 0.001954 | 0.001811 | -152.6375 | 0.23772 | 8.4014 | 0.003239 | 0.004873 | 0.004642 | 5.9e-05 | 0.001067 | 0.28926 | 1.3191 | 2.5481 | 101.18 | 20.79 | 17.85 | 68.7914 | 99.7349 | 65.3546 | 0.004336 | 1.043 | 0.025437 | 24.3892 | -0.008451 | 0.14142 | 0.000132 | 0.010126 | 0.010126 | 0.58099
2019-03-19 16:04 | 1553025835.7777 | 1306 | 0.20953 | 2.8e-05 | 0 | 101.2846 | 101.4462 | 127.0646 | 21.2192 | 0.001966 | 0.001896 | -152.6556 | 0.23719 | 8.4028 | 0.003235 | 0.004545 | 0.004382 | 5.8e-05 | 0.001063 | 0.30901 | 1.3272 | 2.5481 | 101.18 | 20.79 | 17.862 | 69.0792 | 99.6267 | 67.4045 | 0.004339 | 1.0367 | 0.02541 | 24.5114 | -0.014264 | 0.1411 | -5.7e-05 | -0.004382 | -0.004382 | -1.3394
2019-03-19 16:04 | 1553025836.7788 | 1307 | 0.20951 | 2.6e-05 | 0 | 101.2957 | 101.2798 | 127.5709 | 21.2192 | 0.001937 | 0.001887 | -152.6454 | 0.23804 | 8.402 | 0.003259 | 0.004559 | 0.004375 | 5.8e-05 | 0.001054 | 0.30182 | 1.3343 | 2.5481 | 101.18 | 20.79 | 17.865 | 68.3905 | 100.2215 | 67.19 | 0.00434 | 1.0373 | 0.025469 | 24.5525 | 0.001405 | 0.14161 | 2.6e-05 | 0.001951 | 0.001951 | 3.0191
2019-03-19 16:04 | 1553025837.7778 | 1308 | 0.20952 | 2.6e-05 | 0 | 101.2255 | 101.2962 | 127.6418 | 21.2181 | 0.00191 | 0.001823 | -152.6431 | 0.23785 | 8.4034 | 0.003239 | 0.005305 | 0.005034 | 6e-05 | 0.001049 | 0.31454 | 1.3284 | 2.548 | 101.18 | 20.79 | 17.867 | 67.7521 | 99.7221 | 65.6541 | 0.00434 | 1.042 | 0.025534 | 24.5038 | -0.00625 | 0.14149 | -3.3e-05 | -0.002588 | -0.002588 | -2.2748
2019-03-19 16:04 | 1553025838.7789 | 1309 | 0.2095 | 2.7e-05 | 0 | 101.2243 | 101.3027 | 127.9619 | 21.2196 | 0.001926 | 0.001915 | -152.638 | 0.23785 | 8.4041 | 0.003212 | 0.003889 | 0.003759 | 5.7e-05 | 0.001049 | 0.30359 | 1.3442 | 2.5483 | 101.18 | 20.79 | 17.867 | 68.1312 | 99.0787 | 67.8619 | 0.00434 | 1.0353 | 0.025483 | 24.6143 | -0.006931 | 0.14149 | 0.000116 | 0.008949 | 0.008949 | 0.65772
2019-03-19 16:04 | 1553025839.778 | 1310 | 0.20951 | 3e-05 | 0 | 101.281 | 101.413 | 127.8563 | 21.2182 | 0.001895 | 0.001855 | -152.635 | 0.2378 | 8.4015 | 0.003216 | 0.004263 | 0.004061 | 5.7e-05 | 0.001042 | 0.28528 | 1.3382 | 2.548 | 101.18 | 20.79 | 17.867 | 67.3793 | 99.1801 | 66.4174 | 0.00434 | 1.0397 | 0.025527 | 24.552 | -0.011653 | 0.14147 | 4e-05 | 0.003131 | 0.003131 | 1.8796
2019-03-19 16:04 | 1553025840.7784 | 1311 | 0.20952 | 2.8e-05 | 0 | 101.2027 | 101.2202 | 127.8197 | 21.218 | 0.001971 | 0.001829 | -152.628 | 0.23754 | 8.4029 | 0.003198 | 0.00485 | 0.004659 | 5.7e-05 | 0.001043 | 0.30055 | 1.3676 | 2.548 | 101.18 | 20.79 | 17.862 | 69.2128 | 98.74 | 65.7983 | 0.004339 | 1.0416 | 0.025546 | 24.5261 | -0.00155 | 0.14131 | -3.3e-05 | -0.00259 | -0.00259 | -2.2698
2019-03-19 16:04 | 1553025841.7791 | 1312 | 0.20952 | 3e-05 | 1e-06 | 101.1579 | 101.3671 | 127.8352 | 21.218 | 0.001884 | 0.001916 | -152.6238 | 0.2374 | 8.4004 | 0.003196 | 0.004481 | 0.00426 | 5.7e-05 | 0.001045 | 0.31653 | 1.3223 | 2.5481 | 101.18 | 20.79 | 17.859 | 67.1107 | 98.6931 | 67.876 | 0.004338 | 1.0352 | 0.02547 | 24.6026 | -0.018477 | 0.14123 | -1e-06 | -9.3e-05 | -9.3e-05 | -63.0431
2019-03-19 16:04 | 1553025842.779 | 1313 | 0.20949 | 3e-05 | 0 | 101.1978 | 101.4466 | 128.0389 | 21.2174 | 0.001946 | 0.001833 | -152.6266 | 0.23762 | 8.4042 | 0.003305 | 0.004492 | 0.004298 | 5.5e-05 | 0.001047 | 0.32516 | 1.3164 | 2.5481 | 101.18 | 20.79 | 17.858 | 68.5993 | 101.3209 | 65.8804 | 0.004338 | 1.0413 | 0.025565 | 24.5501 | -0.021957 | 0.14136 | 0.00015 | 0.011559 | 0.011559 | 0.50874
2019-03-19 16:04 | 1553025843.7792 | 1314 | 0.20952 | 3.5e-05 | 1e-06 | 101.1879 | 101.2613 | 127.9964 | 21.2181 | 0.001922 | 0.001853 | -152.6303 | 0.23729 | 8.4027 | 0.003229 | 0.0047 | 0.004534 | 5.8e-05 | 0.001051 | 0.31492 | 1.311 | 2.5481 | 101.18 | 20.79 | 17.856 | 68.0235 | 99.4805 | 66.3682 | 0.004338 | 1.0398 | 0.025542 | 24.5637 | -0.00648 | 0.14117 | 8e-06 | 0.00056 | 0.00056 | 10.4934
2019-03-19 16:04 | 1553025844.7782 | 1315 | 0.2095 | 3.2e-05 | 1e-06 | 101.0952 | 101.3339 | 128.0478 | 21.2177 | 0.001991 | 0.001878 | -152.6237 | 0.23796 | 8.4029 | 0.003335 | 0.004771 | 0.004527 | 5.9e-05 | 0.001051 | 0.30663 | 1.3081 | 2.5481 | 101.18 | 20.79 | 17.856 | 69.6937 | 102.056 | 66.9661 | 0.004338 | 1.038 | 0.025525 | 24.5902 | -0.021092 | 0.14156 | 7.1e-05 | 0.005449 | 0.005449 | 1.0808
2019-03-19 16:04 | 1553025845.7782 | 1316 | 0.2095 | 3.7e-05 | 1e-06 | 101.0781 | 101.2644 | 127.9281 | 21.2173 | 0.001922 | 0.001849 | -152.623 | 0.23756 | 8.4 | 0.003327 | 0.004324 | 0.004153 | 5.6e-05 | 0.001055 | 0.29247 | 1.3015 | 2.5482 | 101.18 | 20.79 | 17.857 | 68.0244 | 101.853 | 66.2667 | 0.004338 | 1.0402 | 0.025539 | 24.5534 | -0.016467 | 0.14132 | 0.000101 | 0.007733 | 0.007733 | 0.76029
2019-03-19 16:04 | 1553025846.7774 | 1317 | 0.2095 | 3.8e-05 | 0 | 101.0772 | 101.3236 | 127.4681 | 21.2169 | 0.001936 | 0.00184 | -152.6221 | 0.23785 | 8.4012 | 0.003272 | 0.00507 | 0.004782 | 5.8e-05 | 0.001053 | 0.28224 | 1.2822 | 2.5481 | 101.18 | 20.79 | 17.859 | 68.3546 | 100.5291 | 66.0587 | 0.004338 | 1.0408 | 0.025501 | 24.5017 | -0.021779 | 0.1415 | 6.2e-05 | 0.004723 | 0.004723 | 1.2464
2019-03-19 16:04 | 1553025847.7787 | 1318 | 0.20949 | 3e-05 | 0 | 101.0487 | 101.1835 | 127.2545 | 21.2174 | 0.001912 | 0.001798 | -152.6177 | 0.23719 | 8.403 | 0.003321 | 0.005467 | 0.005307 | 5.7e-05 | 0.001056 | 0.29734 | 1.3135 | 2.5482 | 101.18 | 20.79 | 17.862 | 67.8007 | 101.7159 | 65.0453 | 0.004339 | 1.0439 | 0.025518 | 24.4446 | -0.011928 | 0.14111 | 0.000182 | 0.013998 | 0.013998 | 0.41934
2019-03-19 16:04 | 1553025848.7785 | 1319 | 0.20947 | 3e-05 | 0 | 101.0886 | 101.1914 | 126.9385 | 21.2168 | 0.00192 | 0.001853 | -152.6255 | 0.23824 | 8.4029 | 0.003343 | 0.004662 | 0.004531 | 5.5e-05 | 0.001046 | 0.30381 | 1.3096 | 2.5481 | 101.18 | 20.79 | 17.868 | 67.9788 | 102.2343 | 66.3739 | 0.00434 | 1.0398 | 0.025436 | 24.4622 | -0.009092 | 0.14173 | 0.000256 | 0.019635 | 0.019635 | 0.30028
2019-03-19 16:04 | 1553025849.7784 | 1320 | 0.20952 | 2.9e-05 | 0 | 101.0736 | 101.0781 | 126.7726 | 21.2168 | 0.001899 | 0.001845 | -152.6186 | 0.23695 | 8.404 | 0.003359 | 0.004639 | 0.004406 | 5.5e-05 | 0.001043 | 0.30901 | 1.3184 | 2.548 | 101.18 | 20.79 | 17.872 | 67.4647 | 102.6192 | 66.1685 | 0.004341 | 1.0405 | 0.025427 | 24.4388 | -0.000393 | 0.14096 | -1.6e-05 | -0.001232 | -0.001232 | -4.7583
2019-03-19 16:04 | 1553025850.7779 | 1321 | 0.20952 | 2.6e-05 | 1e-06 | 101.0673 | 101.0483 | 126.5756 | 21.2166 | 0.001947 | 0.001838 | -152.6298 | 0.23662 | 8.4024 | 0.003295 | 0.004654 | 0.004389 | 5.6e-05 | 0.001043 | 0.32621 | 1.2843 | 2.5482 | 101.18 | 20.79 | 17.879 | 68.6187 | 101.0719 | 66.0131 | 0.004343 | 1.0409 | 0.025413 | 24.4142 | 0.001686 | 0.14077 | 1.3e-05 | 0.000975 | 0.000975 | 6.0049
2019-03-19 16:04 | 1553025851.7776 | 1322 | 0.20954 | 3.4e-05 | 0 | 100.9995 | 101.1263 | 126.8193 | 21.2172 | 0.001931 | 0.001841 | -152.6231 | 0.23785 | 8.4004 | 0.003397 | 0.004469 | 0.004365 | 5.6e-05 | 0.001044 | 0.34092 | 1.3259 | 2.548 | 101.18 | 20.79 | 17.881 | 68.2373 | 103.5571 | 66.0778 | 0.004344 | 1.0407 | 0.025436 | 24.44 | -0.011228 | 0.14149 | -0.000129 | -0.009906 | -0.009906 | -0.5942
2019-03-19 16:04 | 1553025852.7777 | 1323 | 0.2095 | 3.4e-05 | 0 | 100.959 | 101.1075 | 126.7754 | 21.2155 | 0.001953 | 0.001825 | -152.6169 | 0.23745 | 8.4023 | 0.003337 | 0.004812 | 0.004636 | 5.7e-05 | 0.001042 | 0.33069 | 1.3846 | 2.5481 | 101.18 | 20.79 | 17.884 | 68.7691 | 102.1028 | 65.6874 | 0.004344 | 1.0419 | 0.025446 | 24.4217 | -0.01315 | 0.14126 | 9.8e-05 | 0.007546 | 0.007546 | 0.77872
2019-03-19 16:04 | 1553025853.7777 | 1324 | 0.20954 | 3.3e-05 | 1e-06 | 100.9518 | 101.0794 | 126.6893 | 21.2162 | 0.001962 | 0.001867 | -152.6209 | 0.23775 | 8.4018 | 0.003336 | 0.004472 | 0.004208 | 5.4e-05 | 0.001038 | 0.32781 | 1.3557 | 2.5481 | 101.18 | 20.79 | 17.885 | 69.0006 | 102.0846 | 66.7019 | 0.004345 | 1.0388 | 0.025399 | 24.4499 | -0.011303 | 0.14144 | -0.000121 | -0.009352 | -0.009352 | -0.62918
2019-03-19 16:04 | 1553025854.7781 | 1325 | 0.2095 | 3.4e-05 | 1e-06 | 100.914 | 100.8602 | 126.7519 | 21.2162 | 0.001931 | 0.001875 | -152.6339 | 0.23735 | 8.4062 | 0.003384 | 0.004661 | 0.004478 | 6.2e-05 | 0.00103 | 0.29302 | 1.3005 | 2.5482 | 101.18 | 20.79 | 17.883 | 68.2347 | 103.2267 | 66.8921 | 0.004344 | 1.0382 | 0.025398 | 24.4628 | 0.004763 | 0.1412 | 9.3e-05 | 0.007113 | 0.007113 | 0.82582
2019-03-19 16:04 | 1553025855.7788 | 1326 | 0.20953 | 3.4e-05 | 1e-06 | 100.9812 | 101.0085 | 126.7625 | 21.2158 | 0.001961 | 0.001843 | -152.6273 | 0.23729 | 8.4003 | 0.003357 | 0.004659 | 0.004467 | 5.8e-05 | 0.001026 | 0.29645 | 1.284 | 2.548 | 101.18 | 20.79 | 17.872 | 68.9744 | 102.5727 | 66.124 | 0.004341 | 1.0406 | 0.025428 | 24.4362 | -0.002422 | 0.14116 | -9.9e-05 | -0.00762 | -0.00762 | -0.77068
2019-03-19 16:04 | 1553025856.7779 | 1327 | 0.20952 | 2.6e-05 | 0 | 100.8368 | 100.9996 | 127.0714 | 21.2159 | 0.001921 | 0.00185 | -152.6335 | 0.23782 | 8.4027 | 0.003373 | 0.00512 | 0.004916 | 5.6e-05 | 0.001026 | 0.34307 | 1.3233 | 2.548 | 101.18 | 20.79 | 17.866 | 67.9963 | 102.9607 | 66.291 | 0.00434 | 1.0401 | 0.025453 | 24.472 | -0.014429 | 0.14148 | -3.3e-05 | -0.002519 | -0.002519 | -2.3364
2019-03-19 16:04 | 1553025857.7779 | 1328 | 0.20949 | 2.9e-05 | 0 | 100.8338 | 101.065 | 127.0665 | 21.216 | 0.001893 | 0.00187 | -152.6419 | 0.2377 | 8.4037 | 0.003319 | 0.004701 | 0.004497 | 5.9e-05 | 0.001026 | 0.3402 | 1.3252 | 2.548 | 101.18 | 20.79 | 17.858 | 67.3413 | 101.6649 | 66.7781 | 0.004338 | 1.0386 | 0.025434 | 24.489 | -0.020486 | 0.14141 | 0.000161 | 0.012368 | 0.012368 | 0.47564
2019-03-19 16:04 | 1553025858.7784 | 1329 | 0.20952 | 2.8e-05 | 0 | 100.9134 | 101.0034 | 127.1829 | 21.2153 | 0.001963 | 0.001826 | -152.6335 | 0.23711 | 8.4025 | 0.003353 | 0.00488 | 0.004665 | 5.8e-05 | 0.001031 | 0.32156 | 1.3175 | 2.5482 | 101.18 | 20.79 | 17.834 | 69.0053 | 102.4914 | 65.7241 | 0.004332 | 1.0418 | 0.025485 | 24.4622 | -0.007977 | 0.14106 | 1.4e-05 | 0.001091 | 0.001091 | 5.378
2019-03-19 16:04 | 1553025859.7771 | 1330 | 0.2095 | 2.8e-05 | 0 | 100.7759 | 100.8051 | 127.339 | 21.2155 | 0.001948 | 0.001833 | -152.6458 | 0.2375 | 8.4024 | 0.003341 | 0.004269 | 0.004035 | 6e-05 | 0.001035 | 0.34971 | 1.3396 | 2.5479 | 101.18 | 20.79 | 17.821 | 68.6652 | 102.1812 | 65.8957 | 0.004329 | 1.0413 | 0.025494 | 24.4834 | -0.00259 | 0.14129 | 8.7e-05 | 0.006706 | 0.006706 | 0.87651
2019-03-19 16:04 | 1553025860.7785 | 1331 | 0.2095 | 2.7e-05 | 0 | 100.8548 | 100.9572 | 127.4488 | 21.2146 | 0.001909 | 0.00178 | -152.6465 | 0.23691 | 8.4029 | 0.003293 | 0.004349 | 0.004092 | 5.7e-05 | 0.001032 | 0.33572 | 1.3548 | 2.5482 | 101.18 | 20.79 | 17.806 | 67.7083 | 101.0293 | 64.6008 | 0.004325 | 1.0453 | 0.025554 | 24.4472 | -0.00907 | 0.14094 | 9.5e-05 | 0.007348 | 0.007348 | 0.79788
2019-03-19 16:04 | 1553025861.7771 | 1332 | 0.20951 | 3e-05 | 0 | 100.8101 | 100.6599 | 127.247 | 21.2147 | 0.001911 | 0.001851 | -152.6486 | 0.23766 | 8.4042 | 0.003347 | 0.004864 | 0.004718 | 5.8e-05 | 0.001034 | 0.30431 | 1.337 | 2.548 | 101.18 | 20.79 | 17.775 | 67.762 | 102.3327 | 66.3226 | 0.004318 | 1.04 | 0.025469 | 24.49 | 0.013341 | 0.14138 | 6.9e-05 | 0.005275 | 0.005275 | 1.115
2019-03-19 16:04 | 1553025862.7782 | 1333 | 0.20951 | 2.9e-05 | 0 | 100.7657 | 100.8126 | 127.4091 | 21.2146 | 0.001928 | 0.00188 | -152.6632 | 0.2371 | 8.403 | 0.003354 | 0.004522 | 0.004297 | 5.6e-05 | 0.001038 | 0.31708 | 1.3299 | 2.548 | 101.18 | 20.79 | 17.76 | 68.1699 | 102.5062 | 67.0271 | 0.004314 | 1.0378 | 0.025459 | 24.531 | -0.004164 | 0.14105 | 5.1e-05 | 0.003891 | 0.003891 | 1.508
2019-03-19 16:04 | 1553025863.7773 | 1334 | 0.2095 | 2.9e-05 | 0 | 100.7636 | 100.7643 | 127.6073 | 21.2152 | 0.001921 | 0.001843 | -152.6723 | 0.23751 | 8.4018 | 0.003282 | 0.004373 | 0.004176 | 5.7e-05 | 0.001031 | 0.31365 | 1.3203 | 2.548 | 101.18 | 20.79 | 17.733 | 68.0043 | 100.7575 | 66.128 | 0.004308 | 1.0406 | 0.025513 | 24.5176 | -6.5e-05 | 0.14129 | 8.7e-05 | 0.006693 | 0.006693 | 0.87817
2019-03-19 16:04 | 1553025864.7782 | 1335 | 0.2095 | 2.8e-05 | 0 | 100.7027 | 100.7527 | 127.6779 | 21.2146 | 0.001963 | 0.001829 | -152.6468 | 0.23801 | 8.4041 | 0.003334 | 0.004521 | 0.004324 | 5.9e-05 | 0.001026 | 0.30879 | 1.3248 | 2.548 | 101.18 | 20.79 | 17.721 | 69.0046 | 102.0312 | 65.7958 | 0.004305 | 1.0416 | 0.025532 | 24.5124 | -0.00444 | 0.14159 | 0.000102 | 0.00786 | 0.00786 | 0.74943
2019-03-19 16:04 | 1553025865.7778 | 1336 | 0.20951 | 2.9e-05 | 0 | 100.6919 | 100.6972 | 127.2758 | 21.2152 | 0.001942 | 0.001832 | -152.6424 | 0.23774 | 8.4013 | 0.003345 | 0.005034 | 0.00485 | 5.5e-05 | 0.00103 | 0.30773 | 1.3404 | 2.5481 | 101.18 | 20.79 | 17.711 | 68.5207 | 102.2995 | 65.8519 | 0.004302 | 1.0414 | 0.02549 | 24.4758 | -0.000473 | 0.14143 | 3.4e-05 | 0.002594 | 0.002594 | 2.2679
2019-03-19 16:04 | 1553025866.7774 | 1337 | 0.20951 | 3e-05 | 0 | 100.6637 | 100.6106 | 127.0019 | 21.2144 | 0.001929 | 0.001857 | -152.6255 | 0.23798 | 8.4021 | 0.003378 | 0.004121 | 0.003978 | 5.7e-05 | 0.001034 | 0.29502 | 1.3813 | 2.5481 | 101.18 | 20.79 | 17.695 | 68.186 | 103.0859 | 66.4605 | 0.004298 | 1.0396 | 0.025439 | 24.4714 | 0.004721 | 0.14158 | 2.3e-05 | 0.0018 | 0.0018 | 3.2725
2019-03-19 16:04 | 1553025867.7778 | 1338 | 0.20952 | 2.7e-05 | 0 | 100.609 | 100.6849 | 127.3477 | 21.214 | 0.001887 | 0.001791 | -152.6228 | 0.2368 | 8.4041 | 0.003341 | 0.004267 | 0.004039 | 5.5e-05 | 0.001038 | 0.31332 | 1.3906 | 2.548 | 101.18 | 20.79 | 17.689 | 67.1934 | 102.1885 | 64.882 | 0.004297 | 1.0444 | 0.025533 | 24.4477 | -0.006739 | 0.14088 | -1.5e-05 | -0.001169 | -0.001169 | -5.0119
2019-03-19 16:04 | 1553025868.7774 | 1339 | 0.20953 | 3.1e-05 | 0 | 100.6361 | 100.6099 | 127.3164 | 21.2141 | 0.001927 | 0.001852 | -152.6272 | 0.23699 | 8.4024 | 0.003334 | 0.004576 | 0.004397 | 5.4e-05 | 0.001038 | 0.31957 | 1.3766 | 2.5481 | 101.18 | 20.79 | 17.679 | 68.1422 | 102.0336 | 66.3407 | 0.004295 | 1.0399 | 0.025475 | 24.4973 | 0.002328 | 0.14099 | -8.2e-05 | -0.006288 | -0.006288 | -0.93275
2019-03-19 16:04 | 1553025869.7776 | 1340 | 0.20952 | 2.9e-05 | 0 | 100.6322 | 100.49 | 127.1934 | 21.2142 | 0.001902 | 0.001804 | -152.6167 | 0.23757 | 8.4047 | 0.003399 | 0.004718 | 0.004528 | 5.8e-05 | 0.001042 | 0.30901 | 1.3523 | 2.5481 | 101.18 | 20.79 | 17.675 | 67.5376 | 103.5878 | 65.1812 | 0.004294 | 1.0435 | 0.025507 | 24.4437 | 0.012645 | 0.14133 | 2e-05 | 0.001517 | 0.001517 | 3.8745
2019-03-19 16:04 | 1553025870.7776 | 1341 | 0.20951 | 2.9e-05 | 0 | 100.5274 | 100.5938 | 127.1953 | 21.2135 | 0.001953 | 0.001854 | -152.599 | 0.23775 | 8.4019 | 0.00343 | 0.004427 | 0.004222 | 5.7e-05 | 0.001044 | 0.32997 | 1.3618 | 2.5481 | 101.18 | 20.79 | 17.672 | 68.7718 | 104.356 | 66.3965 | 0.004293 | 1.0398 | 0.025461 | 24.4877 | -0.005902 | 0.14144 | 5.8e-05 | 0.004449 | 0.004449 | 1.3225
2019-03-19 16:04 | 1553025871.7777 | 1342 | 0.20951 | 2.7e-05 | 0 | 100.6069 | 100.5643 | 127.3193 | 21.2141 | 0.00191 | 0.001806 | -152.6053 | 0.23713 | 8.4012 | 0.003356 | 0.004147 | 0.003944 | 5.6e-05 | 0.001041 | 0.31686 | 1.363 | 2.548 | 101.18 | 20.79 | 17.67 | 67.742 | 102.5623 | 65.235 | 0.004292 | 1.0433 | 0.025517 | 24.4577 | 0.003791 | 0.14107 | 7.8e-05 | 0.006 | 0.006 | 0.97803
2019-03-19 16:05 | 1553025872.7787 | 1343 | 0.20952 | 2.9e-05 | 0 | 100.4812 | 100.5465 | 127.4249 | 21.2138 | 0.001911 | 0.001876 | -152.5931 | 0.23669 | 8.403 | 0.003336 | 0.00459 | 0.004363 | 5.6e-05 | 0.00104 | 0.33019 | 1.34 | 2.548 | 101.18 | 20.78 | 17.67 | 67.7534 | 102.0742 | 66.914 | 0.00429 | 1.0382 | 0.025465 | 24.5285 | -0.005811 | 0.14081 | -1.5e-05 | -0.001142 | -0.001142 | -5.1312
2019-03-19 16:05 | 1553025873.7778 | 1344 | 0.20949 | 3.1e-05 | 0 | 100.4314 | 100.4194 | 127.3094 | 21.2132 | 0.001895 | 0.001875 | -152.5887 | 0.23874 | 8.401 | 0.003441 | 0.004515 | 0.004331 | 5.6e-05 | 0.001035 | 0.2756 | 1.3544 | 2.5481 | 101.18 | 20.78 | 17.672 | 67.3713 | 104.6051 | 66.9048 | 0.00429 | 1.0382 | 0.025454 | 24.517 | 0.001064 | 0.14202 | 0.000141 | 0.010818 | 0.010818 | 0.54613
2019-03-19 16:05 | 1553025874.7792 | 1345 | 0.20954 | 2.8e-05 | 0 | 100.3987 | 100.4575 | 127.0122 | 21.2146 | 0.001893 | 0.001914 | -152.6134 | 0.2366 | 8.4043 | 0.003412 | 0.004568 | 0.004392 | 5.4e-05 | 0.001037 | 0.27688 | 1.3656 | 2.5482 | 101.18 | 20.79 | 17.674 | 67.3197 | 103.9097 | 67.8463 | 0.004293 | 1.0353 | 0.025389 | 24.5223 | -0.005235 | 0.14075 | -0.000144 | -0.011042 | -0.011042 | -0.53026
2019-03-19 16:05 | 1553025875.7781 | 1346 | 0.2095 | 3.1e-05 | 1e-06 | 100.3873 | 100.4417 | 127.2474 | 21.2137 | 0.001918 | 0.001839 | -152.6124 | 0.23658 | 8.4021 | 0.003348 | 0.004897 | 0.004696 | 5.8e-05 | 0.001036 | 0.31797 | 1.3527 | 2.5481 | 101.18 | 20.79 | 17.676 | 67.9383 | 102.3664 | 66.0229 | 0.004294 | 1.0409 | 0.02548 | 24.4792 | -0.004849 | 0.14074 | 0.000134 | 0.010267 | 0.010267 | 0.57029
2019-03-19 16:05 | 1553025876.7769 | 1347 | 0.20949 | 2.4e-05 | 1e-06 | 100.3291 | 100.2321 | 127.622 | 21.2141 | 0.001925 | 0.001827 | -152.6225 | 0.23684 | 8.403 | 0.003396 | 0.004559 | 0.004393 | 5.8e-05 | 0.001037 | 0.32261 | 1.3319 | 2.5481 | 101.18 | 20.78 | 17.681 | 68.0947 | 103.5241 | 65.7368 | 0.004292 | 1.0418 | 0.025529 | 24.5049 | 0.008648 | 0.1409 | 0.00016 | 0.012271 | 0.012271 | 0.47764
2019-03-19 16:05 | 1553025877.7769 | 1348 | 0.20955 | 2.3e-05 | 1e-06 | 100.3246 | 100.2691 | 127.2183 | 21.214 | 0.001911 | 0.001825 | -152.6229 | 0.23664 | 8.4047 | 0.003393 | 0.004313 | 0.004104 | 5.8e-05 | 0.001036 | 0.34307 | 1.2933 | 2.548 | 101.18 | 20.78 | 17.684 | 67.7623 | 103.4462 | 65.6877 | 0.004293 | 1.0419 | 0.02549 | 24.4643 | 0.004946 | 0.14078 | -0.000172 | -0.013299 | -0.013299 | -0.44037
2019-03-19 16:05 | 1553025878.7784 | 1349 | 0.20952 | 2.4e-05 | 1e-06 | 100.3029 | 100.1903 | 127.1672 | 21.2132 | 0.00187 | 0.001828 | -152.6175 | 0.23695 | 8.4047 | 0.003359 | 0.004703 | 0.004471 | 5.6e-05 | 0.001035 | 0.31581 | 1.3163 | 2.5481 | 101.18 | 20.78 | 17.688 | 66.7785 | 102.6234 | 65.7571 | 0.004294 | 1.0417 | 0.025482 | 24.4619 | 0.010049 | 0.14096 | 4e-06 | 0.00028 | 0.00028 | 20.9567
2019-03-19 16:05 | 1553025879.7781 | 1350 | 0.20952 | 2.1e-05 | 0 | 100.2399 | 100.3592 | 127.3585 | 21.2132 | 0.001943 | 0.001796 | -152.6057 | 0.23659 | 8.4029 | 0.003484 | 0.00449 | 0.004271 | 5.4e-05 | 0.001037 | 0.29551 | 1.3382 | 2.5482 | 101.18 | 20.78 | 17.688 | 68.5382 | 105.6463 | 65.007 | 0.004294 | 1.044 | 0.02553 | 24.4532 | -0.010634 | 0.14075 | 3.2e-05 | 0.002449 | 0.002449 | 2.3905
2019-03-19 16:05 | 1553025880.7784 | 1351 | 0.20952 | 2.4e-05 | 0 | 100.2801 | 100.1074 | 127.469 | 21.2141 | 0.001954 | 0.00184 | -152.6213 | 0.23712 | 8.404 | 0.003519 | 0.004355 | 0.00422 | 5.9e-05 | 0.001041 | 0.30558 | 1.3149 | 2.5479 | 101.18 | 20.78 | 17.694 | 68.7928 | 106.5052 | 66.0584 | 0.004296 | 1.0408 | 0.025501 | 24.5018 | 0.015418 | 0.14106 | -5e-06 | -0.000409 | -0.000409 | -14.3491
2019-03-19 16:05 | 1553025881.7781 | 1352 | 0.20953 | 2.2e-05 | 1e-06 | 100.2507 | 100.1177 | 127.3319 | 21.2141 | 0.001941 | 0.001803 | -152.6259 | 0.23722 | 8.4023 | 0.003533 | 0.005053 | 0.004799 | 5.7e-05 | 0.001043 | 0.32516 | 1.333 | 2.548 | 101.18 | 20.78 | 17.698 | 68.4846 | 106.8563 | 65.1561 | 0.004296 | 1.0436 | 0.025522 | 24.4561 | 0.011876 | 0.14112 | -4.5e-05 | -0.003535 | -0.003535 | -1.6605
2019-03-19 16:05 | 1553025882.7776 | 1353 | 0.20951 | 2.5e-05 | 0 | 100.2438 | 100.1317 | 127.1446 | 21.2129 | 0.001879 | 0.00181 | -152.6171 | 0.23725 | 8.405 | 0.00353 | 0.004874 | 0.004716 | 5.5e-05 | 0.001041 | 0.30525 | 1.3182 | 2.5479 | 101.18 | 20.78 | 17.701 | 66.9865 | 106.7637 | 65.3243 | 0.004297 | 1.043 | 0.025496 | 24.4441 | 0.010005 | 0.14114 | 3.6e-05 | 0.002706 | 0.002706 | 2.1698
2019-03-19 16:05 | 1553025883.7782 | 1354 | 0.20949 | 2.4e-05 | 1e-06 | 100.1388 | 100.2191 | 127.1477 | 21.2131 | 0.001912 | 0.001821 | -152.6063 | 0.23618 | 8.4002 | 0.00346 | 0.005004 | 0.004812 | 5.5e-05 | 0.001038 | 0.32532 | 1.3278 | 2.5482 | 101.18 | 20.78 | 17.709 | 67.7956 | 105.0854 | 65.6016 | 0.004299 | 1.0422 | 0.025486 | 24.4544 | -0.00717 | 0.1405 | 0.000153 | 0.011698 | 0.011698 | 0.49966
2019-03-19 16:05 | 1553025884.7774 | 1355 | 0.20947 | 2.7e-05 | 0 | 100.1358 | 99.9598 | 127.393 | 21.2124 | 0.001884 | 0.001865 | -152.5992 | 0.23703 | 8.4038 | 0.003536 | 0.004906 | 0.004691 | 5.9e-05 | 0.001034 | 0.30829 | 1.3416 | 2.5482 | 101.18 | 20.78 | 17.712 | 67.1226 | 106.9153 | 66.6467 | 0.0043 | 1.039 | 0.025472 | 24.5157 | 0.015732 | 0.14101 | 0.000299 | 0.022935 | 0.022935 | 0.25576
2019-03-19 16:05 | 1553025885.7783 | 1356 | 0.20952 | 2.6e-05 | 1e-06 | 100.173 | 100.2359 | 127.3695 | 21.2123 | 0.00194 | 0.001815 | -152.6077 | 0.23703 | 8.4025 | 0.00351 | 0.004481 | 0.004291 | 6e-05 | 0.001029 | 0.29319 | 1.3275 | 2.5481 | 101.18 | 20.78 | 17.72 | 68.4523 | 106.2941 | 65.4572 | 0.004302 | 1.0426 | 0.025514 | 24.4705 | -0.005612 | 0.14101 | -3.1e-05 | -0.00241 | -0.00241 | -2.4341
2019-03-19 16:05 | 1553025886.7786 | 1357 | 0.20948 | 2.3e-05 | 1e-06 | 100.1622 | 100.0427 | 127.0869 | 21.2134 | 0.001976 | 0.001825 | -152.6234 | 0.23632 | 8.4019 | 0.003505 | 0.005185 | 0.004937 | 5.9e-05 | 0.001028 | 0.31614 | 1.3935 | 2.5481 | 101.18 | 20.78 | 17.723 | 69.3154 | 106.1637 | 65.6916 | 0.004303 | 1.0419 | 0.025477 | 24.4518 | 0.010677 | 0.14058 | 0.000216 | 0.016554 | 0.016554 | 0.3533
2019-03-19 16:05 | 1553025887.7785 | 1358 | 0.20953 | 3e-05 | 1e-06 | 100.113 | 100.1461 | 127.3425 | 21.212 | 0.001943 | 0.001814 | -152.623 | 0.23597 | 8.3984 | 0.003462 | 0.005014 | 0.004809 | 5.6e-05 | 0.001026 | 0.33987 | 1.3564 | 2.5481 | 101.18 | 20.78 | 17.727 | 68.5308 | 105.1127 | 65.4389 | 0.004304 | 1.0427 | 0.025512 | 24.4673 | -0.002962 | 0.14038 | -7.7e-05 | -0.005972 | -0.005972 | -0.97779
2019-03-19 16:05 | 1553025888.7775 | 1359 | 0.20952 | 2.4e-05 | 1e-06 | 100.1031 | 100.0718 | 127.3826 | 21.2128 | 0.001895 | 0.001839 | -152.6297 | 0.23617 | 8.404 | 0.003548 | 0.004549 | 0.004261 | 5.9e-05 | 0.001031 | 0.32405 | 1.3377 | 2.548 | 101.18 | 20.78 | 17.731 | 67.3851 | 107.2164 | 66.0329 | 0.004304 | 1.0409 | 0.025494 | 24.4926 | 0.002794 | 0.1405 | 1e-05 | 0.000753 | 0.000753 | 7.7615
2019-03-19 16:05 | 1553025889.779 | 1360 | 0.20954 | 2.9e-05 | 1e-06 | 100.0692 | 100.054 | 127.4333 | 21.2122 | 0.001934 | 0.001864 | -152.651 | 0.23729 | 8.4016 | 0.00352 | 0.005462 | 0.005313 | 5.9e-05 | 0.001036 | 0.32908 | 1.3616 | 2.548 | 101.18 | 20.78 | 17.734 | 68.3111 | 106.5339 | 66.6316 | 0.004305 | 1.039 | 0.025476 | 24.5191 | 0.001355 | 0.14117 | -0.00015 | -0.011612 | -0.011612 | -0.50572
2019-03-19 16:05 | 1553025890.7767 | 1361 | 0.20948 | 2.7e-05 | 1e-06 | 100.0503 | 99.9393 | 127.357 | 21.2125 | 0.001913 | 0.0018 | -152.6495 | 0.23667 | 8.4037 | 0.003534 | 0.005741 | 0.005577 | 5.9e-05 | 0.001033 | 0.32189 | 1.4039 | 2.548 | 101.18 | 20.79 | 17.738 | 67.8061 | 106.8688 | 65.0923 | 0.004309 | 1.0438 | 0.025526 | 24.4562 | 0.009928 | 0.1408 | 0.000212 | 0.016281 | 0.016281 | 0.35976
2019-03-19 16:05 | 1553025891.7767 | 1362 | 0.20953 | 2.8e-05 | 0 | 100.0983 | 100.2557 | 127.6441 | 21.2127 | 0.001896 | 0.00184 | -152.6281 | 0.23697 | 8.406 | 0.003564 | 0.005 | 0.004712 | 5.9e-05 | 0.001031 | 0.32875 | 1.3519 | 2.5481 | 101.18 | 20.79 | 17.738 | 67.4027 | 107.5959 | 66.0553 | 0.004309 | 1.0408 | 0.025519 | 24.5185 | -0.014059 | 0.14097 | -7e-05 | -0.005415 | -0.005415 | -1.083
2019-03-19 16:05 | 1553025892.7777 | 1363 | 0.20951 | 2.8e-05 | 0 | 100.0749 | 100.0543 | 127.768 | 21.2123 | 0.001972 | 0.001826 | -152.6336 | 0.23642 | 8.4002 | 0.003464 | 0.005111 | 0.004973 | 5.9e-05 | 0.001035 | 0.35386 | 1.3232 | 2.548 | 101.18 | 20.79 | 17.737 | 69.2191 | 105.1664 | 65.725 | 0.004309 | 1.0418 | 0.025544 | 24.5185 | 0.001834 | 0.14065 | 4.2e-05 | 0.003228 | 0.003228 | 1.8126
2019-03-19 16:05 | 1553025893.7777 | 1364 | 0.20947 | 2.9e-05 | 0 | 100.0677 | 100.0472 | 127.8033 | 21.2129 | 0.00193 | 0.001837 | -152.6339 | 0.2367 | 8.4014 | 0.003458 | 0.004651 | 0.00438 | 6e-05 | 0.001039 | 0.31238 | 1.3135 | 2.5479 | 101.18 | 20.79 | 17.737 | 68.2151 | 105.0348 | 65.9721 | 0.004309 | 1.0411 | 0.025538 | 24.5308 | 0.001833 | 0.14081 | 0.000304 | 0.023406 | 0.023406 | 0.25027
2019-03-19 16:05 | 1553025894.7773 | 1365 | 0.20951 | 3.1e-05 | 1e-06 | 99.9878 | 99.8947 | 127.8931 | 21.2126 | 0.001945 | 0.001832 | -152.6404 | 0.23626 | 8.4043 | 0.003511 | 0.004284 | 0.004108 | 6.3e-05 | 0.001047 | 0.30934 | 1.299 | 2.548 | 101.18 | 20.79 | 17.739 | 68.5797 | 106.3108 | 65.8493 | 0.004309 | 1.0414 | 0.025552 | 24.535 | 0.008331 | 0.14055 | 5.9e-05 | 0.00449 | 0.00449 | 1.3022
2019-03-19 16:05 | 1553025895.7773 | 1366 | 0.2095 | 2.6e-05 | 0 | 100.0611 | 99.8906 | 128.0065 | 21.2133 | 0.001879 | 0.001837 | -152.6251 | 0.23733 | 8.4041 | 0.003544 | 0.005065 | 0.004797 | 6e-05 | 0.001048 | 0.32261 | 1.325 | 2.548 | 101.18 | 20.79 | 17.741 | 67.0001 | 107.1028 | 65.9912 | 0.00431 | 1.041 | 0.025558 | 24.551 | 0.015247 | 0.14119 | 8.8e-05 | 0.006771 | 0.006771 | 0.86737
2019-03-19 16:05 | 1553025896.7779 | 1367 | 0.2095 | 2.9e-05 | 0 | 99.9899 | 100.1167 | 128.0373 | 21.2129 | 0.001937 | 0.00181 | -152.6162 | 0.23681 | 8.4024 | 0.0035 | 0.004693 | 0.004429 | 6.4e-05 | 0.001055 | 0.35242 | 1.3733 | 2.548 | 101.18 | 20.79 | 17.747 | 68.375 | 106.0388 | 65.3416 | 0.004311 | 1.043 | 0.025585 | 24.5304 | -0.011331 | 0.14088 | 9.2e-05 | 0.007066 | 0.007066 | 0.82935
2019-03-19 16:05 | 1553025897.777 | 1368 | 0.20949 | 2.9e-05 | 0 | 100.0368 | 100.1434 | 127.957 | 21.2136 | 0.001985 | 0.001768 | -152.6319 | 0.23641 | 8.4045 | 0.003486 | 0.00492 | 0.004799 | 6.2e-05 | 0.001046 | 0.31094 | 1.3509 | 2.5481 | 101.18 | 20.79 | 17.749 | 69.5331 | 105.7018 | 64.318 | 0.004312 | 1.0462 | 0.025616 | 24.4856 | -0.00953 | 0.14064 | 0.000157 | 0.01213 | 0.01213 | 0.48233
2019-03-19 16:05 | 1553025898.7781 | 1369 | 0.2095 | 2.7e-05 | 1e-06 | 99.9872 | 99.9314 | 128.1363 | 21.2132 | 0.00194 | 0.001827 | -152.6431 | 0.23673 | 8.4023 | 0.003435 | 0.005159 | 0.004949 | 6.3e-05 | 0.001051 | 0.3361 | 1.3185 | 2.5481 | 101.18 | 20.79 | 17.753 | 68.4503 | 104.4699 | 65.753 | 0.004312 | 1.0417 | 0.025579 | 24.5548 | 0.004997 | 0.14083 | 0.000113 | 0.008675 | 0.008675 | 0.67538
2019-03-19 16:05 | 1553025899.8392 | 1370 | 0.20954 | 2.8e-05 | 1e-06 | 99.9734 | 99.9872 | 128.3086 | 21.2121 | 0.001986 | 0.001834 | -152.6542 | 0.23713 | 8.402 | 0.00347 | 0.004281 | 0.004055 | 6.2e-05 | 0.001057 | 0.31454 | 1.3597 | 2.5481 | 101.18 | 20.79 | 17.755 | 69.5691 | 105.3235 | 65.9182 | 0.004313 | 1.0412 | 0.02559 | 24.5773 | -0.001234 | 0.14107 | -0.0001 | -0.007817 | -0.007817 | -0.75069
2019-03-19 16:05 | 1553025900.7828 | 1371 | 0.20952 | 3.2e-05 | 1e-06 | 99.9353 | 100.0016 | 128.2776 | 21.2117 | 0.001893 | 0.001822 | -152.6681 | 0.23695 | 8.4017 | 0.003394 | 0.00472 | 0.004469 | 6.3e-05 | 0.00106 | 0.31852 | 1.3851 | 2.5482 | 101.18 | 20.79 | 17.756 | 67.334 | 103.4674 | 65.6274 | 0.004313 | 1.0421 | 0.025598 | 24.5638 | -0.005931 | 0.14096 | -1.9e-05 | -0.001526 | -0.001526 | -3.8432
2019-03-19 16:05 | 1553025901.7829 | 1372 | 0.20949 | 3e-05 | 1e-06 | 100.0103 | 100.0478 | 128.1023 | 21.2125 | 0.00192 | 0.00187 | -152.6778 | 0.23731 | 8.4012 | 0.003383 | 0.005135 | 0.004867 | 6e-05 | 0.001048 | 0.32565 | 1.4185 | 2.548 | 101.18 | 20.79 | 17.754 | 67.9697 | 103.2168 | 66.7653 | 0.004313 | 1.0386 | 0.025538 | 24.5882 | -0.003353 | 0.14117 | 0.00014 | 0.010794 | 0.010794 | 0.54409
2019-03-19 16:05 | 1553025902.7829 | 1373 | 0.20953 | 2.7e-05 | 1e-06 | 99.9884 | 99.9293 | 127.9915 | 21.212 | 0.001866 | 0.001848 | -152.6834 | 0.23695 | 8.3997 | 0.003385 | 0.004752 | 0.004557 | 6.2e-05 | 0.001042 | 0.31382 | 1.3992 | 2.5481 | 101.18 | 20.79 | 17.754 | 66.6884 | 103.2629 | 66.2514 | 0.004313 | 1.0402 | 0.025546 | 24.559 | 0.005288 | 0.14096 | -3.8e-05 | -0.00298 | -0.00298 | -1.9679
2019-03-19 16:05 | 1553025903.7833 | 1374 | 0.20954 | 3.1e-05 | 0 | 100.0001 | 99.9687 | 127.9838 | 21.2134 | 0.00197 | 0.001786 | -152.6952 | 0.23695 | 8.4027 | 0.003429 | 0.004419 | 0.004231 | 6.3e-05 | 0.001048 | 0.30215 | 1.376 | 2.5482 | 101.18 | 20.78 | 17.755 | 69.1757 | 104.3212 | 64.7558 | 0.00431 | 1.0448 | 0.025602 | 24.5041 | 0.00281 | 0.14096 | -0.000119 | -0.009196 | -0.009196 | -0.6377
2019-03-19 16:05 | 1553025904.7831 | 1375 | 0.20951 | 2.8e-05 | 1e-06 | 99.9995 | 99.9321 | 127.7917 | 21.2117 | 0.001895 | 0.001885 | -152.6912 | 0.23721 | 8.4016 | 0.003538 | 0.004676 | 0.004468 | 6e-05 | 0.001051 | 0.31188 | 1.3952 | 2.5483 | 101.18 | 20.78 | 17.757 | 67.3708 | 106.961 | 67.1359 | 0.004311 | 1.0375 | 0.025493 | 24.5718 | 0.006036 | 0.14111 | 4.5e-05 | 0.00339 | 0.00339 | 1.7315
2019-03-19 16:05 | 1553025905.7834 | 1376 | 0.20949 | 2.8e-05 | 0 | 99.9968 | 100.1081 | 127.8129 | 21.2121 | 0.001905 | 0.001799 | -152.691 | 0.236 | 8.4034 | 0.003488 | 0.005139 | 0.004919 | 5.7e-05 | 0.001054 | 0.32333 | 1.3368 | 2.5482 | 101.18 | 20.78 | 17.758 | 67.6232 | 105.7501 | 65.0669 | 0.004311 | 1.0438 | 0.025573 | 24.499 | -0.009948 | 0.1404 | 0.000187 | 0.014383 | 0.014383 | 0.40607
2019-03-19 16:05 | 1553025906.7811 | 1377 | 0.2095 | 2.7e-05 | 1e-06 | 99.9794 | 100.0239 | 127.8685 | 21.2121 | 0.001958 | 0.001805 | -152.7043 | 0.23644 | 8.4017 | 0.00341 | 0.00511 | 0.004968 | 6e-05 | 0.001055 | 0.33948 | 1.3584 | 2.5481 | 101.18 | 20.78 | 17.757 | 68.8889 | 103.8656 | 65.2025 | 0.004311 | 1.0434 | 0.025574 | 24.5092 | -0.003974 | 0.14066 | 9.8e-05 | 0.007522 | 0.007522 | 0.77791
2019-03-19 16:05 | 1553025907.7826 | 1378 | 0.2095 | 3.3e-05 | 1e-06 | 99.9854 | 99.8766 | 127.9647 | 21.2124 | 0.001891 | 0.001769 | -152.6814 | 0.23672 | 8.4044 | 0.003451 | 0.004913 | 0.004633 | 6e-05 | 0.001049 | 0.3131 | 1.3215 | 2.5482 | 101.18 | 20.78 | 17.755 | 67.2895 | 104.844 | 64.3489 | 0.00431 | 1.0461 | 0.025615 | 24.4875 | 0.009742 | 0.14083 | 0.000137 | 0.010558 | 0.010558 | 0.55488
2019-03-19 16:05 | 1553025908.7813 | 1379 | 0.2095 | 2.9e-05 | 0 | 100.077 | 100.0656 | 127.7373 | 21.2107 | 0.001925 | 0.001858 | -152.672 | 0.23707 | 8.4045 | 0.003433 | 0.005007 | 0.004755 | 6.1e-05 | 0.001055 | 0.30735 | 1.3226 | 2.548 | 101.18 | 20.78 | 17.754 | 68.0906 | 104.4225 | 66.481 | 0.00431 | 1.0395 | 0.025512 | 24.5429 | 0.001011 | 0.14103 | 0.000112 | 0.008645 | 0.008645 | 0.67866
2019-03-19 16:05 | 1553025909.7813 | 1380 | 0.2095 | 2.5e-05 | 0 | 100.0154 | 100.0098 | 127.4817 | 21.2121 | 0.001928 | 0.00182 | -152.6443 | 0.23659 | 8.4022 | 0.003462 | 0.004342 | 0.004126 | 5.7e-05 | 0.001056 | 0.30807 | 1.3212 | 2.548 | 101.18 | 20.78 | 17.753 | 68.1761 | 105.112 | 65.5686 | 0.00431 | 1.0423 | 0.025521 | 24.4853 | 0.000503 | 0.14075 | 0.000111 | 0.008524 | 0.008524 | 0.68692
2019-03-19 16:05 | 1553025910.7823 | 1381 | 0.20952 | 2.4e-05 | 1e-06 | 100.0202 | 99.9783 | 127.6725 | 21.2118 | 0.001888 | 0.00182 | -152.6381 | 0.23697 | 8.404 | 0.003374 | 0.0052 | 0.004988 | 6e-05 | 0.00106 | 0.3366 | 1.3518 | 2.5481 | 101.18 | 20.78 | 17.754 | 67.2141 | 102.9805 | 65.5753 | 0.00431 | 1.0423 | 0.02554 | 24.5039 | 0.003751 | 0.14097 | 2e-06 | 7.4e-05 | 7.4e-05 | 79.3104
2019-03-19 16:05 | 1553025911.7828 | 1382 | 0.20953 | 3.2e-05 | 0 | 99.9839 | 99.9704 | 128.0706 | 21.2117 | 0.001906 | 0.001806 | -152.6289 | 0.23703 | 8.4019 | 0.003347 | 0.004508 | 0.004474 | 6.2e-05 | 0.001063 | 0.31598 | 1.3207 | 2.548 | 101.18 | 20.78 | 17.757 | 67.6429 | 102.3265 | 65.2271 | 0.004311 | 1.0433 | 0.025593 | 24.5295 | 0.001207 | 0.14101 | -7e-05 | -0.005402 | -0.005402 | -1.086
2019-03-19 16:05 | 1553025912.7815 | 1383 | 0.20952 | 3.2e-05 | 1e-06 | 100.0653 | 99.9167 | 128.218 | 21.2114 | 0.001912 | 0.001832 | -152.636 | 0.23709 | 8.4051 | 0.003411 | 0.004468 | 0.004214 | 5.9e-05 | 0.001065 | 0.3194 | 1.3492 | 2.5481 | 101.18 | 20.78 | 17.759 | 67.773 | 103.8782 | 65.8489 | 0.004311 | 1.0414 | 0.025584 | 24.5661 | 0.013292 | 0.14105 | -3.5e-05 | -0.002724 | -0.002724 | -2.1536
2019-03-19 16:05 | 1553025913.7819 | 1384 | 0.20952 | 2.8e-05 | 0 | 100.1034 | 100.204 | 128.403 | 21.2114 | 0.001941 | 0.001768 | -152.6199 | 0.2364 | 8.4031 | 0.003364 | 0.004221 | 0.004114 | 6e-05 | 0.001059 | 0.33395 | 1.3531 | 2.5481 | 101.18 | 20.78 | 17.767 | 68.4701 | 102.7498 | 64.3173 | 0.004313 | 1.0462 | 0.02566 | 24.5282 | -0.008988 | 0.14064 | 1.8e-05 | 0.001376 | 0.001376 | 4.2513
2019-03-19 16:05 | 1553025914.7815 | 1385 | 0.20953 | 3.1e-05 | 1e-06 | 100.0629 | 99.9879 | 128.1398 | 21.211 | 0.00187 | 0.001817 | -152.6195 | 0.23614 | 8.404 | 0.003392 | 0.004911 | 0.00467 | 5.8e-05 | 0.001061 | 0.33229 | 1.3381 | 2.548 | 101.18 | 20.78 | 17.773 | 66.7833 | 103.4127 | 65.5004 | 0.004315 | 1.0425 | 0.025589 | 24.546 | 0.006703 | 0.14048 | -6.2e-05 | -0.004824 | -0.004824 | -1.2114
2019-03-19 16:05 | 1553025915.7817 | 1386 | 0.20952 | 2.8e-05 | 1e-06 | 100.0824 | 100.0194 | 127.8124 | 21.212 | 0.001865 | 0.001771 | -152.6309 | 0.23694 | 8.4 | 0.003355 | 0.004286 | 0.004096 | 5.5e-05 | 0.001062 | 0.32764 | 1.3589 | 2.5481 | 101.18 | 20.78 | 17.778 | 66.6586 | 102.5211 | 64.3892 | 0.004316 | 1.0459 | 0.025599 | 24.4744 | 0.005628 | 0.14095 | -3.5e-05 | -0.002728 | -0.002728 | -2.1497
2019-03-19 16:05 | 1553025916.7817 | 1387 | 0.20955 | 2.7e-05 | 1e-06 | 100.041 | 100.1218 | 127.6136 | 21.2109 | 0.001896 | 0.001773 | -152.6253 | 0.2377 | 8.4041 | 0.003405 | 0.004877 | 0.004684 | 5.9e-05 | 0.001068 | 0.30901 | 1.3782 | 2.548 | 101.18 | 20.78 | 17.787 | 67.4051 | 103.737 | 64.4348 | 0.004318 | 1.0458 | 0.025577 | 24.457 | -0.007227 | 0.14141 | -0.000157 | -0.012195 | -0.012195 | -0.48238
2019-03-19 16:05 | 1553025917.7827 | 1388 | 0.20952 | 2.7e-05 | 0 | 99.9977 | 99.754 | 127.5983 | 21.2125 | 0.001872 | 0.00183 | -152.6435 | 0.23742 | 8.4028 | 0.003438 | 0.004538 | 0.004299 | 5.9e-05 | 0.001076 | 0.29767 | 1.3506 | 2.548 | 101.18 | 20.78 | 17.792 | 66.8314 | 104.5476 | 65.8093 | 0.004319 | 1.0416 | 0.025524 | 24.5052 | 0.02183 | 0.14124 | 1.3e-05 | 0.000932 | 0.000932 | 6.3052
2019-03-19 16:05 | 1553025918.7828 | 1389 | 0.20953 | 2.1e-05 | 0 | 100.1112 | 100.0386 | 127.8084 | 21.2114 | 0.001917 | 0.001819 | -152.6463 | 0.23751 | 8.4039 | 0.003378 | 0.004783 | 0.00456 | 5.6e-05 | 0.001078 | 0.34164 | 1.3302 | 2.5481 | 101.18 | 20.78 | 17.796 | 67.911 | 103.0869 | 65.5453 | 0.00432 | 1.0424 | 0.025555 | 24.5159 | 0.006488 | 0.1413 | -7.7e-05 | -0.005974 | -0.005974 | -0.98387
2019-03-19 16:05 | 1553025919.7819 | 1390 | 0.20956 | 2.2e-05 | 0 | 100.089 | 100.1465 | 127.8493 | 21.2109 | 0.00191 | 0.001792 | -152.6524 | 0.2368 | 8.4038 | 0.003358 | 0.004365 | 0.004141 | 6e-05 | 0.001086 | 0.34794 | 1.3262 | 2.5482 | 101.18 | 20.78 | 17.801 | 67.7429 | 102.6023 | 64.89 | 0.004321 | 1.0444 | 0.025583 | 24.496 | -0.005138 | 0.14088 | -0.000222 | -0.017191 | -0.017191 | -0.34091
2019-03-19 16:05 | 1553025920.7823 | 1391 | 0.20951 | 2.2e-05 | 1e-06 | 100.1262 | 100.1047 | 127.6507 | 21.2107 | 0.001913 | 0.001794 | -152.6407 | 0.23737 | 8.4018 | 0.003288 | 0.004817 | 0.004654 | 5.7e-05 | 0.001089 | 0.31813 | 1.3335 | 2.5481 | 101.18 | 20.78 | 17.802 | 67.8098 | 100.9116 | 64.9406 | 0.004322 | 1.0442 | 0.025562 | 24.4788 | 0.00192 | 0.14121 | 7.2e-05 | 0.005474 | 0.005474 | 1.0732
2019-03-19 16:05 | 1553025921.781 | 1392 | 0.20956 | 2.4e-05 | 1e-06 | 100.2231 | 100.1807 | 127.2871 | 21.2102 | 0.002002 | 0.001876 | -152.6611 | 0.23735 | 8.4021 | 0.003297 | 0.00481 | 0.004534 | 5.7e-05 | 0.001086 | 0.32836 | 1.3281 | 2.5481 | 101.18 | 20.78 | 17.802 | 69.9391 | 101.1121 | 66.9081 | 0.004322 | 1.0382 | 0.025451 | 24.515 | 0.003784 | 0.1412 | -0.00023 | -0.017682 | -0.017682 | -0.3322
2019-03-19 16:05 | 1553025922.782 | 1393 | 0.20952 | 2.8e-05 | 1e-06 | 100.1799 | 100.2194 | 126.9813 | 21.2106 | 0.001936 | 0.001819 | -152.6635 | 0.23841 | 8.4024 | 0.003255 | 0.004916 | 0.004735 | 5.5e-05 | 0.001083 | 0.30973 | 1.3422 | 2.548 | 101.18 | 20.78 | 17.8 | 68.3602 | 100.1195 | 65.5366 | 0.004321 | 1.0424 | 0.025472 | 24.4361 | -0.003528 | 0.14183 | 1.1e-05 | 0.000816 | 0.000816 | 7.2274
2019-03-19 16:05 | 1553025923.781 | 1394 | 0.20953 | 2.3e-05 | 1e-06 | 100.233 | 100.1294 | 126.6552 | 21.2107 | 0.001957 | 0.001755 | -152.6659 | 0.23721 | 8.4031 | 0.003247 | 0.005013 | 0.004717 | 5.6e-05 | 0.001075 | 0.31006 | 1.3536 | 2.5481 | 101.18 | 20.78 | 17.798 | 68.8743 | 99.9046 | 64.0074 | 0.004321 | 1.0471 | 0.025497 | 24.3495 | 0.009256 | 0.14111 | -3.3e-05 | -0.002616 | -0.002616 | -2.2439
2019-03-19 16:05 | 1553025924.7822 | 1395 | 0.20954 | 2.6e-05 | 1e-06 | 100.2441 | 100.2256 | 127.1035 | 21.211 | 0.001908 | 0.001802 | -152.6709 | 0.23697 | 8.4004 | 0.003336 | 0.004246 | 0.004094 | 5.2e-05 | 0.001077 | 0.30469 | 1.3444 | 2.5481 | 101.18 | 20.79 | 17.791 | 67.6923 | 102.0574 | 65.1329 | 0.004322 | 1.0436 | 0.0255 | 24.4333 | 0.001654 | 0.14097 | -0.000112 | -0.008724 | -0.008724 | -0.67218
2019-03-19 16:05 | 1553025925.7812 | 1396 | 0.20954 | 2.6e-05 | 1e-06 | 100.1979 | 100.1417 | 127.4237 | 21.2112 | 0.001882 | 0.00181 | -152.6821 | 0.23757 | 8.4047 | 0.003363 | 0.00432 | 0.004186 | 5.8e-05 | 0.001086 | 0.33732 | 1.3425 | 2.5482 | 101.18 | 20.79 | 17.786 | 67.0582 | 102.7181 | 65.3328 | 0.004321 | 1.043 | 0.025524 | 24.4713 | 0.00502 | 0.14133 | -0.00011 | -0.008483 | -0.008483 | -0.69304
2019-03-19 16:05 | 1553025926.7823 | 1397 | 0.20955 | 2.5e-05 | 1e-06 | 100.2375 | 100.2054 | 127.7859 | 21.2108 | 0.001851 | 0.001777 | -152.6773 | 0.23712 | 8.4027 | 0.003282 | 0.004881 | 0.004693 | 5.8e-05 | 0.001088 | 0.3027 | 1.3304 | 2.5482 | 101.18 | 20.79 | 17.771 | 66.3063 | 100.7728 | 64.5417 | 0.004317 | 1.0455 | 0.02559 | 24.4774 | 0.002868 | 0.14106 | -0.000162 | -0.012524 | -0.012524 | -0.46857
2019-03-19 16:05 | 1553025927.7816 | 1398 | 0.20953 | 2.5e-05 | 0 | 100.2168 | 100.093 | 127.8183 | 21.2114 | 0.001961 | 0.001786 | -152.6654 | 0.2378 | 8.4026 | 0.003218 | 0.004407 | 0.004212 | 5.7e-05 | 0.001096 | 0.29518 | 1.3449 | 2.548 | 101.18 | 20.79 | 17.763 | 68.9584 | 99.2105 | 64.7594 | 0.004315 | 1.0448 | 0.025585 | 24.4883 | 0.011053 | 0.14147 | -7.6e-05 | -0.005898 | -0.005898 | -0.99784
2019-03-19 16:05 | 1553025928.7824 | 1399 | 0.20953 | 2.4e-05 | 0 | 100.2126 | 100.3225 | 127.6859 | 21.2116 | 0.001931 | 0.001848 | -152.6607 | 0.23714 | 8.4014 | 0.003172 | 0.00466 | 0.004502 | 5.7e-05 | 0.001097 | 0.30397 | 1.3481 | 2.5482 | 101.18 | 20.79 | 17.756 | 68.235 | 98.1018 | 66.2556 | 0.004313 | 1.0402 | 0.025516 | 24.5298 | -0.009804 | 0.14108 | -3.2e-05 | -0.002459 | -0.002459 | -2.3864
2019-03-19 16:05 | 1553025929.7814 | 1400 | 0.20956 | 2.5e-05 | 1e-06 | 100.2108 | 100.1674 | 127.6481 | 21.2119 | 0.001892 | 0.001792 | -152.6669 | 0.23802 | 8.404 | 0.003227 | 0.004536 | 0.004393 | 5.7e-05 | 0.001099 | 0.32565 | 1.2952 | 2.5482 | 101.18 | 20.79 | 17.745 | 67.3085 | 99.4227 | 64.898 | 0.004311 | 1.0444 | 0.025563 | 24.4771 | 0.003878 | 0.14159 | -0.000236 | -0.018279 | -0.018279 | -0.32224
2019-03-19 16:05 | 1553025930.7824 | 1401 | 0.20953 | 2.2e-05 | 1e-06 | 100.2021 | 100.2807 | 127.4563 | 21.2121 | 0.001926 | 0.001858 | -152.675 | 0.23821 | 8.4022 | 0.00321 | 0.004761 | 0.004547 | 5.8e-05 | 0.001095 | 0.32709 | 1.3049 | 2.5481 | 101.18 | 20.79 | 17.741 | 68.1239 | 99.0299 | 66.4831 | 0.00431 | 1.0395 | 0.025484 | 24.5159 | -0.007016 | 0.14171 | -4.3e-05 | -0.003386 | -0.003386 | -1.7411
2019-03-19 16:06 | 1553025931.7825 | 1402 | 0.2095 | 2.8e-05 | 1e-06 | 100.1898 | 100.1163 | 127.4528 | 21.2113 | 0.001903 | 0.001846 | -152.6617 | 0.23681 | 8.4019 | 0.003209 | 0.004196 | 0.003989 | 5.7e-05 | 0.001097 | 0.31614 | 1.308 | 2.5479 | 101.18 | 20.79 | 17.738 | 67.5616 | 98.9907 | 66.192 | 0.004309 | 1.0404 | 0.025495 | 24.5051 | 0.006561 | 0.14088 | 0.000137 | 0.010511 | 0.010511 | 0.55758
2019-03-19 16:06 | 1553025932.7816 | 1403 | 0.20954 | 2.6e-05 | 1e-06 | 100.1793 | 100.1407 | 127.6549 | 21.2124 | 0.001914 | 0.001766 | -152.6555 | 0.23758 | 8.4028 | 0.003275 | 0.004815 | 0.004626 | 6.2e-05 | 0.0011 | 0.2959 | 1.2966 | 2.5481 | 101.18 | 20.79 | 17.734 | 67.8418 | 100.5971 | 64.2669 | 0.004308 | 1.0463 | 0.025588 | 24.4548 | 0.003451 | 0.14134 | -0.000122 | -0.009515 | -0.009515 | -0.61796
2019-03-19 16:06 | 1553025933.7826 | 1404 | 0.20956 | 2.7e-05 | 1e-06 | 100.2762 | 100.3352 | 127.9394 | 21.2115 | 0.001964 | 0.001756 | -152.6581 | 0.23791 | 8.4028 | 0.00333 | 0.00479 | 0.004647 | 5.6e-05 | 0.001104 | 0.35723 | 1.302 | 2.548 | 101.18 | 20.79 | 17.732 | 69.046 | 101.9212 | 64.0319 | 0.004307 | 1.047 | 0.025625 | 24.4736 | -0.005257 | 0.14153 | -0.000264 | -0.02048 | -0.02048 | -0.28749
2019-03-19 16:06 | 1553025934.7816 | 1405 | 0.20954 | 3e-05 | 1e-06 | 100.2633 | 100.3335 | 127.7368 | 21.212 | 0.001878 | 0.001774 | -152.6616 | 0.23746 | 8.4016 | 0.003333 | 0.004865 | 0.004667 | 5.6e-05 | 0.00111 | 0.32405 | 1.38 | 2.5482 | 101.18 | 20.8 | 17.729 | 66.9762 | 102.0056 | 64.4577 | 0.004309 | 1.0457 | 0.025589 | 24.4696 | -0.006256 | 0.14127 | -9.8e-05 | -0.007633 | -0.007633 | -0.76986
2019-03-19 16:06 | 1553025935.7821 | 1406 | 0.2095 | 2.9e-05 | 0 | 100.221 | 100.3164 | 127.7472 | 21.2118 | 0.001909 | 0.001774 | -152.6398 | 0.23698 | 8.4036 | 0.003315 | 0.005301 | 0.005041 | 5.8e-05 | 0.001113 | 0.32875 | 1.3677 | 2.5482 | 101.18 | 20.8 | 17.726 | 67.7085 | 101.5548 | 64.4684 | 0.004309 | 1.0457 | 0.025589 | 24.471 | -0.008505 | 0.14098 | 0.000101 | 0.007814 | 0.007814 | 0.75052
2019-03-19 16:06 | 1553025936.7828 | 1407 | 0.20959 | 2.9e-05 | 1e-06 | 100.2516 | 100.3164 | 127.4495 | 21.2117 | 0.001879 | 0.001803 | -152.6474 | 0.23788 | 8.4033 | 0.003378 | 0.004933 | 0.004743 | 6e-05 | 0.001114 | 0.31868 | 1.3327 | 2.5481 | 101.18 | 20.8 | 17.724 | 66.994 | 103.0745 | 65.1631 | 0.004308 | 1.0435 | 0.025533 | 24.4676 | -0.005773 | 0.14151 | -0.000444 | -0.034242 | -0.034242 | -0.17192
2019-03-19 16:06 | 1553025937.782 | 1408 | 0.20953 | 2.6e-05 | 1e-06 | 100.2672 | 100.2715 | 127.4894 | 21.212 | 0.001891 | 0.001819 | -152.6688 | 0.2368 | 8.4033 | 0.003306 | 0.005028 | 0.004809 | 5.3e-05 | 0.001114 | 0.29231 | 1.3688 | 2.5482 | 101.18 | 20.8 | 17.717 | 67.286 | 101.3368 | 65.5537 | 0.004306 | 1.0423 | 0.025522 | 24.4855 | -0.00038 | 0.14087 | -6.9e-05 | -0.005394 | -0.005394 | -1.0865
2019-03-19 16:06 | 1553025938.7811 | 1409 | 0.20953 | 2.5e-05 | 1e-06 | 100.293 | 100.1804 | 127.4347 | 21.2109 | 0.001991 | 0.001834 | -152.6635 | 0.23667 | 8.4025 | 0.003289 | 0.004417 | 0.004244 | 5.6e-05 | 0.001112 | 0.33572 | 1.3627 | 2.5482 | 101.18 | 20.8 | 17.712 | 69.6723 | 100.9306 | 65.8961 | 0.004305 | 1.0413 | 0.025504 | 24.4927 | 0.010053 | 0.1408 | -6.5e-05 | -0.005058 | -0.005058 | -1.1581
2019-03-19 16:06 | 1553025939.7813 | 1410 | 0.20953 | 3e-05 | 1e-06 | 100.2948 | 100.3461 | 127.5828 | 21.2113 | 0.00193 | 0.00179 | -152.6597 | 0.23738 | 8.4017 | 0.003325 | 0.004833 | 0.004579 | 5.4e-05 | 0.001105 | 0.30541 | 1.313 | 2.5481 | 101.18 | 20.8 | 17.7 | 68.2097 | 101.8087 | 64.849 | 0.004302 | 1.0445 | 0.025558 | 24.469 | -0.004575 | 0.14122 | -8.7e-05 | -0.006764 | -0.006764 | -0.86853
2019-03-19 16:06 | 1553025940.7821 | 1411 | 0.20949 | 2.8e-05 | 1e-06 | 100.293 | 100.0341 | 128.0645 | 21.2115 | 0.00191 | 0.00175 | -152.6629 | 0.23729 | 8.4044 | 0.003308 | 0.00456 | 0.004324 | 5.5e-05 | 0.001104 | 0.33395 | 1.3071 | 2.548 | 101.18 | 20.8 | 17.692 | 67.7418 | 101.3982 | 63.892 | 0.0043 | 1.0475 | 0.025643 | 24.4804 | 0.023119 | 0.14117 | 0.00015 | 0.011533 | 0.011533 | 0.5092
2019-03-19 16:06 | 1553025941.7821 | 1412 | 0.20958 | 2.8e-05 | 1e-06 | 100.2441 | 100.3948 | 128.1593 | 21.211 | 0.001883 | 0.001757 | -152.6819 | 0.23753 | 8.4057 | 0.003362 | 0.004272 | 0.004086 | 5.2e-05 | 0.001109 | 0.33085 | 1.3292 | 2.548 | 101.18 | 20.8 | 17.682 | 67.0945 | 102.7024 | 64.0532 | 0.004298 | 1.047 | 0.025646 | 24.4954 | -0.013433 | 0.1413 | -0.000386 | -0.029911 | -0.029911 | -0.19652
2019-03-19 16:06 | 1553025942.781 | 1413 | 0.20954 | 2.8e-05 | 1e-06 | 100.3095 | 100.3657 | 127.9903 | 21.2114 | 0.001884 | 0.001803 | -152.6967 | 0.23669 | 8.4038 | 0.003286 | 0.004794 | 0.004651 | 5.5e-05 | 0.001109 | 0.32117 | 1.3118 | 2.5482 | 101.18 | 20.8 | 17.662 | 67.1021 | 100.8692 | 65.1667 | 0.004293 | 1.0435 | 0.025587 | 24.5196 | -0.005003 | 0.14081 | -0.000106 | -0.008191 | -0.008191 | -0.71513
2019-03-19 16:06 | 1553025943.7816 | 1414 | 0.20954 | 2.7e-05 | 1e-06 | 100.278 | 100.3564 | 127.8446 | 21.2111 | 0.001859 | 0.001807 | -152.6989 | 0.23693 | 8.4011 | 0.003337 | 0.004453 | 0.004235 | 5.2e-05 | 0.001102 | 0.32029 | 1.3664 | 2.5481 | 101.18 | 20.8 | 17.652 | 66.4987 | 102.1023 | 65.253 | 0.004291 | 1.0433 | 0.025569 | 24.5087 | -0.006989 | 0.14095 | -0.000129 | -0.009983 | -0.009983 | -0.58733
2019-03-19 16:06 | 1553025944.7823 | 1415 | 0.20954 | 2.8e-05 | 1e-06 | 100.2786 | 100.3074 | 127.7915 | 21.2111 | 0.001947 | 0.001746 | -152.6957 | 0.23762 | 8.4048 | 0.003299 | 0.004961 | 0.004745 | 5.4e-05 | 0.001094 | 0.31188 | 1.3868 | 2.548 | 101.18 | 20.81 | 17.642 | 68.622 | 101.1742 | 63.7854 | 0.004291 | 1.0478 | 0.02562 | 24.4505 | -0.00257 | 0.14136 | -0.000102 | -0.007888 | -0.007888 | -0.7455
2019-03-19 16:06 | 1553025945.7826 | 1416 | 0.20954 | 2.6e-05 | 1e-06 | 100.2201 | 100.3379 | 127.9032 | 21.2117 | 0.001876 | 0.001851 | -152.7012 | 0.23831 | 8.4008 | 0.003213 | 0.005148 | 0.004872 | 5.2e-05 | 0.001088 | 0.32405 | 1.3481 | 2.5482 | 101.18 | 20.81 | 17.626 | 66.912 | 99.0962 | 66.3238 | 0.004287 | 1.04 | 0.025535 | 24.5531 | -0.010509 | 0.14177 | -0.0001 | -0.007776 | -0.007776 | -0.75846
2019-03-19 16:06 | 1553025946.7823 | 1417 | 0.20953 | 2.5e-05 | 1e-06 | 100.3273 | 100.4794 | 128.095 | 21.2105 | 0.001866 | 0.001776 | -152.6851 | 0.23745 | 8.4017 | 0.003345 | 0.004845 | 0.004596 | 5.4e-05 | 0.001076 | 0.3246 | 1.3576 | 2.5482 | 101.18 | 20.81 | 17.618 | 66.6674 | 102.2862 | 64.5179 | 0.004285 | 1.0455 | 0.025622 | 24.5061 | -0.013552 | 0.14126 | -2.9e-05 | -0.002318 | -0.002318 | -2.5353
2019-03-19 16:06 | 1553025947.7814 | 1418 | 0.20955 | 2.2e-05 | 1e-06 | 100.2795 | 100.267 | 128.2065 | 21.2119 | 0.001954 | 0.001816 | -152.6696 | 0.23755 | 8.406 | 0.00331 | 0.004854 | 0.004604 | 5.6e-05 | 0.001077 | 0.2724 | 1.3325 | 2.548 | 101.18 | 20.81 | 17.609 | 68.8 | 101.4402 | 65.4787 | 0.004283 | 1.0426 | 0.025597 | 24.5516 | 0.001115 | 0.14132 | -0.000184 | -0.014248 | -0.014248 | -0.41261
2019-03-19 16:06 | 1553025948.7814 | 1419 | 0.20954 | 2.7e-05 | 1e-06 | 100.2771 | 100.5123 | 128.2835 | 21.2114 | 0.001893 | 0.001774 | -152.679 | 0.23787 | 8.4048 | 0.003311 | 0.00464 | 0.004496 | 5.2e-05 | 0.001076 | 0.31797 | 1.3315 | 2.548 | 101.18 | 20.81 | 17.608 | 67.3361 | 101.4634 | 64.4603 | 0.004283 | 1.0457 | 0.025643 | 24.522 | -0.020948 | 0.14151 | -0.000111 | -0.008621 | -0.008621 | -0.68285
2019-03-19 16:06 | 1553025949.7828 | 1420 | 0.2095 | 2.5e-05 | 1e-06 | 100.3717 | 100.5218 | 128.0671 | 21.2111 | 0.001911 | 0.001738 | -152.6823 | 0.23598 | 8.4033 | 0.003266 | 0.005299 | 0.00505 | 5.1e-05 | 0.001076 | 0.33052 | 1.3361 | 2.5479 | 101.18 | 20.81 | 17.611 | 67.7563 | 100.3822 | 63.6098 | 0.004283 | 1.0484 | 0.025654 | 24.4704 | -0.013373 | 0.14038 | 0.000124 | 0.009513 | 0.009513 | 0.61386
2019-03-19 16:06 | 1553025950.7815 | 1421 | 0.20953 | 2.7e-05 | 1e-06 | 100.2855 | 100.2766 | 128.0429 | 21.2105 | 0.00188 | 0.001784 | -152.6924 | 0.23699 | 8.4027 | 0.003333 | 0.005188 | 0.005008 | 5e-05 | 0.00108 | 0.32139 | 1.3862 | 2.5482 | 101.18 | 20.81 | 17.616 | 67.0161 | 101.9964 | 64.7053 | 0.004284 | 1.045 | 0.02561 | 24.5079 | 0.000795 | 0.14098 | -7e-05 | -0.005472 | -0.005472 | -1.0719
2019-03-19 16:06 | 1553025951.7816 | 1422 | 0.20954 | 2.4e-05 | 1e-06 | 100.3351 | 100.3917 | 128.1614 | 21.2113 | 0.001881 | 0.001759 | -152.6772 | 0.23609 | 8.4034 | 0.003319 | 0.004977 | 0.004806 | 5e-05 | 0.00108 | 0.31636 | 1.3763 | 2.5482 | 101.18 | 20.81 | 17.621 | 67.0378 | 101.6564 | 64.1026 | 0.004286 | 1.0468 | 0.025645 | 24.4974 | -0.005048 | 0.14045 | -0.000101 | -0.007883 | -0.007883 | -0.74114
2019-03-19 16:06 | 1553025952.7822 | 1423 | 0.20956 | 2.8e-05 | 1e-06 | 100.2753 | 100.4016 | 128.3185 | 21.2115 | 0.001925 | 0.001791 | -152.6583 | 0.237 | 8.4006 | 0.003342 | 0.004302 | 0.004091 | 5.2e-05 | 0.001079 | 0.32117 | 1.3578 | 2.5481 | 101.18 | 20.81 | 17.634 | 68.0967 | 102.2165 | 64.867 | 0.004289 | 1.0445 | 0.025631 | 24.5401 | -0.011259 | 0.14099 | -0.000261 | -0.02026 | -0.02026 | -0.2895
2019-03-19 16:06 | 1553025953.7821 | 1424 | 0.20953 | 2.3e-05 | 1e-06 | 100.2711 | 100.4462 | 128.3401 | 21.2111 | 0.001893 | 0.001802 | -152.6646 | 0.23779 | 8.4013 | 0.003384 | 0.004844 | 0.004621 | 5.2e-05 | 0.001073 | 0.31868 | 1.344 | 2.5482 | 101.18 | 20.81 | 17.641 | 67.3327 | 103.2186 | 65.1306 | 0.004291 | 1.0436 | 0.025623 | 24.5517 | -0.015599 | 0.14146 | -4.1e-05 | -0.003254 | -0.003254 | -1.8084
2019-03-19 16:06 | 1553025954.7811 | 1425 | 0.20955 | 2.8e-05 | 1e-06 | 100.3137 | 100.2944 | 128.3757 | 21.2109 | 0.001862 | 0.001815 | -152.6532 | 0.23749 | 8.4027 | 0.00328 | 0.00516 | 0.004988 | 5e-05 | 0.001069 | 0.30381 | 1.3397 | 2.548 | 101.18 | 20.81 | 17.647 | 66.5719 | 100.7243 | 65.452 | 0.004292 | 1.0427 | 0.025615 | 24.5668 | 0.001722 | 0.14128 | -0.000167 | -0.012982 | -0.012982 | -0.45274
2019-03-19 16:06 | 1553025955.7808 | 1426 | 0.20958 | 2.6e-05 | 1e-06 | 100.2933 | 100.3328 | 128.4605 | 21.2102 | 0.001861 | 0.001753 | -152.6728 | 0.23776 | 8.4013 | 0.003325 | 0.004725 | 0.004455 | 5e-05 | 0.001071 | 0.32029 | 1.3243 | 2.5481 | 101.18 | 20.81 | 17.657 | 66.5515 | 101.8044 | 63.9477 | 0.004294 | 1.0473 | 0.02568 | 24.5203 | -0.003518 | 0.14144 | -0.000336 | -0.026137 | -0.026137 | -0.22512
2019-03-19 16:06 | 1553025956.781 | 1427 | 0.20955 | 2.7e-05 | 0 | 100.2807 | 100.2441 | 128.1314 | 21.2109 | 0.001894 | 0.00178 | -152.6856 | 0.23712 | 8.4032 | 0.003313 | 0.004424 | 0.004235 | 5.1e-05 | 0.001074 | 0.32029 | 1.3125 | 2.548 | 101.18 | 20.81 | 17.663 | 67.3425 | 101.506 | 64.602 | 0.004296 | 1.0453 | 0.025623 | 24.5126 | 0.003269 | 0.14106 | -0.000178 | -0.013788 | -0.013788 | -0.4256
2019-03-19 16:06 | 1553025957.7819 | 1428 | 0.20947 | 2.4e-05 | 0 | 100.2651 | 100.3759 | 127.8758 | 21.2114 | 0.001903 | 0.001797 | -152.689 | 0.23671 | 8.404 | 0.003276 | 0.004655 | 0.004405 | 5.4e-05 | 0.001082 | 0.34899 | 1.36 | 2.5481 | 101.18 | 20.82 | 17.67 | 67.5561 | 100.6265 | 65.0247 | 0.0043 | 1.044 | 0.025581 | 24.5035 | -0.00988 | 0.14082 | 0.000326 | 0.025161 | 0.025161 | 0.23283
2019-03-19 16:06 | 1553025958.7823 | 1429 | 0.20951 | 2.8e-05 | 1e-06 | 100.2279 | 100.4112 | 128.1366 | 21.2097 | 0.00185 | 0.0018 | -152.684 | 0.23634 | 8.4029 | 0.003262 | 0.005001 | 0.004727 | 5.1e-05 | 0.001088 | 0.32803 | 1.336 | 2.5481 | 101.18 | 20.82 | 17.672 | 66.2936 | 100.2748 | 65.0979 | 0.004301 | 1.0437 | 0.025604 | 24.5311 | -0.016344 | 0.1406 | 7.3e-05 | 0.00559 | 0.00559 | 1.0464
2019-03-19 16:06 | 1553025959.7812 | 1430 | 0.20952 | 2.6e-05 | 0 | 100.3035 | 100.2729 | 128.3295 | 21.2104 | 0.001866 | 0.001746 | -152.6814 | 0.23719 | 8.4011 | 0.003329 | 0.00433 | 0.004179 | 5.2e-05 | 0.001092 | 0.32709 | 1.3456 | 2.5481 | 101.18 | 20.82 | 17.672 | 66.6705 | 101.9061 | 63.8019 | 0.004301 | 1.0478 | 0.025673 | 24.5025 | 0.002737 | 0.14111 | -1.4e-05 | -0.00108 | -0.00108 | -5.4362
2019-03-19 16:06 | 1553025960.7811 | 1431 | 0.20954 | 2.5e-05 | 0 | 100.2939 | 100.5184 | 127.9608 | 21.2105 | 0.001899 | 0.001817 | -152.6587 | 0.23729 | 8.4034 | 0.003288 | 0.004207 | 0.00405 | 5.6e-05 | 0.0011 | 0.32068 | 1.3168 | 2.5481 | 101.18 | 20.82 | 17.671 | 67.4677 | 100.9074 | 65.5055 | 0.004301 | 1.0425 | 0.025571 | 24.529 | -0.019998 | 0.14117 | -0.000115 | -0.008909 | -0.008909 | -0.65918
2019-03-19 16:06 | 1553025961.7811 | 1432 | 0.20949 | 2.2e-05 | 1e-06 | 100.3251 | 100.2759 | 127.6943 | 21.209 | 0.001887 | 0.001811 | -152.6563 | 0.23693 | 8.3996 | 0.003332 | 0.004031 | 0.003832 | 5.4e-05 | 0.001098 | 0.30596 | 1.3108 | 2.5481 | 101.18 | 20.82 | 17.668 | 67.1815 | 101.9768 | 65.3431 | 0.0043 | 1.043 | 0.025551 | 24.4976 | 0.004388 | 0.14095 | 0.000182 | 0.013981 | 0.013981 | 0.41939
2019-03-19 16:06 | 1553025962.7815 | 1433 | 0.2095 | 2.8e-05 | 0 | 100.3561 | 100.2547 | 127.7612 | 21.21 | 0.001922 | 0.001777 | -152.6562 | 0.23641 | 8.4039 | 0.00326 | 0.004932 | 0.004721 | 5.7e-05 | 0.001093 | 0.31116 | 1.2972 | 2.548 | 101.18 | 20.82 | 17.655 | 68.0198 | 100.2297 | 64.5237 | 0.004297 | 1.0455 | 0.025588 | 24.4743 | 0.009038 | 0.14064 | 0.000126 | 0.00969 | 0.00969 | 0.60382
2019-03-19 16:06 | 1553025963.7822 | 1434 | 0.20952 | 2.6e-05 | 1e-06 | 100.2477 | 100.2002 | 127.6037 | 21.2101 | 0.001901 | 0.001743 | -152.6642 | 0.23669 | 8.4051 | 0.003234 | 0.004658 | 0.00442 | 6e-05 | 0.001095 | 0.29231 | 1.2956 | 2.5481 | 101.18 | 20.82 | 17.648 | 67.508 | 99.6045 | 63.7166 | 0.004295 | 1.048 | 0.025603 | 24.43 | 0.004237 | 0.14081 | -9e-06 | -0.000747 | -0.000747 | -7.837
2019-03-19 16:06 | 1553025964.7812 | 1435 | 0.20947 | 2.5e-05 | 0 | 100.2858 | 100.3324 | 127.554 | 21.21 | 0.001835 | 0.001775 | -152.6626 | 0.23801 | 8.4049 | 0.003164 | 0.00485 | 0.004652 | 5.7e-05 | 0.001096 | 0.3131 | 1.3095 | 2.548 | 101.18 | 20.82 | 17.639 | 65.9224 | 97.9068 | 64.4948 | 0.004293 | 1.0456 | 0.025569 | 24.4534 | -0.004156 | 0.14159 | 0.000341 | 0.026277 | 0.026277 | 0.22416
2019-03-19 16:06 | 1553025965.7813 | 1436 | 0.20952 | 2.7e-05 | 1e-06 | 100.3083 | 100.2105 | 127.535 | 21.21 | 0.001905 | 0.001802 | -152.6538 | 0.23742 | 8.4031 | 0.003321 | 0.004699 | 0.004552 | 5.7e-05 | 0.001103 | 0.32139 | 1.3299 | 2.5481 | 101.18 | 20.82 | 17.622 | 67.6101 | 101.714 | 65.1403 | 0.004289 | 1.0436 | 0.025542 | 24.475 | 0.008727 | 0.14124 | -2.6e-05 | -0.002066 | -0.002066 | -2.8447
2019-03-19 16:06 | 1553025966.7818 | 1437 | 0.2095 | 2.8e-05 | 1e-06 | 100.3236 | 100.1725 | 128.0354 | 21.2101 | 0.00184 | 0.001821 | -152.6324 | 0.23762 | 8.4023 | 0.003257 | 0.004664 | 0.004449 | 6e-05 | 0.001108 | 0.31797 | 1.3553 | 2.5481 | 101.18 | 20.82 | 17.613 | 66.0451 | 100.1638 | 65.6049 | 0.004286 | 1.0422 | 0.025575 | 24.5398 | 0.013486 | 0.14136 | 0.000135 | 0.010379 | 0.010379 | 0.56661
2019-03-19 16:06 | 1553025967.7824 | 1438 | 0.20955 | 3.1e-05 | 0 | 100.2483 | 100.1653 | 128.0558 | 21.2098 | 0.001894 | 0.001798 | -152.6356 | 0.23776 | 8.4028 | 0.003238 | 0.004559 | 0.004354 | 6.1e-05 | 0.001109 | 0.33085 | 1.3899 | 2.5481 | 101.18 | 20.82 | 17.606 | 67.3501 | 99.7055 | 65.0451 | 0.004285 | 1.0439 | 0.025598 | 24.5214 | 0.007409 | 0.14144 | -0.000186 | -0.014414 | -0.014414 | -0.40823
2019-03-19 16:06 | 1553025968.7805 | 1439 | 0.20951 | 2.7e-05 | 1e-06 | 100.2969 | 100.2705 | 127.7065 | 21.2104 | 0.001923 | 0.001806 | -152.6308 | 0.23721 | 8.4034 | 0.003118 | 0.005236 | 0.005054 | 5.7e-05 | 0.001117 | 0.31382 | 1.3466 | 2.5481 | 101.18 | 20.82 | 17.589 | 68.0442 | 96.7944 | 65.2384 | 0.004281 | 1.0433 | 0.025556 | 24.495 | 0.002362 | 0.14112 | 6.1e-05 | 0.004673 | 0.004673 | 1.2562
2019-03-19 16:06 | 1553025969.7804 | 1440 | 0.2095 | 2.5e-05 | 0 | 100.2396 | 100.3225 | 127.6143 | 21.2104 | 0.001875 | 0.001737 | -152.6371 | 0.23766 | 8.4016 | 0.003192 | 0.004531 | 0.004286 | 6e-05 | 0.001117 | 0.34468 | 1.2957 | 2.548 | 101.18 | 20.82 | 17.589 | 66.8989 | 98.5735 | 63.5778 | 0.004281 | 1.0485 | 0.02561 | 24.426 | -0.007394 | 0.14138 | 0.000104 | 0.008013 | 0.008013 | 0.73396
2019-03-19 16:06 | 1553025970.7816 | 1441 | 0.20952 | 2.9e-05 | 0 | 100.2816 | 100.1684 | 127.8479 | 21.2095 | 0.001889 | 0.00177 | -152.6424 | 0.23713 | 8.4012 | 0.003212 | 0.0049 | 0.004691 | 6e-05 | 0.001122 | 0.32477 | 1.2923 | 2.5481 | 101.18 | 20.82 | 17.575 | 67.2304 | 99.0563 | 64.3617 | 0.004277 | 1.046 | 0.025603 | 24.4768 | 0.010106 | 0.14107 | 2.6e-05 | 0.001964 | 0.001964 | 2.9874
2019-03-19 16:06 | 1553025971.7807 | 1442 | 0.2095 | 2.5e-05 | 1e-06 | 100.3717 | 100.2174 | 127.6925 | 21.2093 | 0.001859 | 0.001781 | -152.6423 | 0.23673 | 8.4029 | 0.003184 | 0.003849 | 0.003676 | 5.9e-05 | 0.001122 | 0.32532 | 1.3294 | 2.5481 | 101.18 | 20.82 | 17.568 | 66.5141 | 98.385 | 64.6409 | 0.004275 | 1.0452 | 0.025577 | 24.472 | 0.01376 | 0.14083 | 9.8e-05 | 0.007521 | 0.007521 | 0.77895
2019-03-19 16:06 | 1553025972.7817 | 1443 | 0.20946 | 2.7e-05 | 0 | 100.266 | 100.3047 | 127.4857 | 21.2094 | 0.001946 | 0.001739 | -152.652 | 0.23735 | 8.4036 | 0.003218 | 0.004782 | 0.004548 | 6e-05 | 0.001125 | 0.30973 | 1.3307 | 2.5481 | 101.18 | 20.82 | 17.556 | 68.5939 | 99.2053 | 63.6276 | 0.004273 | 1.0483 | 0.025595 | 24.4155 | -0.00345 | 0.1412 | 0.000403 | 0.031114 | 0.031114 | 0.18879
2019-03-19 16:06 | 1553025973.7807 | 1444 | 0.20949 | 2.6e-05 | 1e-06 | 100.2615 | 100.4205 | 127.6066 | 21.2084 | 0.001902 | 0.001782 | -152.6587 | 0.23779 | 8.4032 | 0.003229 | 0.004714 | 0.004557 | 5.8e-05 | 0.001127 | 0.34451 | 1.3403 | 2.5481 | 101.18 | 20.82 | 17.551 | 67.5421 | 99.4647 | 64.6573 | 0.004271 | 1.0451 | 0.025568 | 24.4644 | -0.014168 | 0.14146 | 0.000192 | 0.014745 | 0.014745 | 0.3991
2019-03-19 16:06 | 1553025974.7808 | 1445 | 0.20955 | 2.2e-05 | 1e-06 | 100.23 | 100.3568 | 127.834 | 21.2088 | 0.001863 | 0.001827 | -152.6861 | 0.23802 | 8.4034 | 0.003235 | 0.004935 | 0.004725 | 5.7e-05 | 0.00113 | 0.31421 | 1.3919 | 2.5481 | 101.18 | 20.82 | 17.543 | 66.6032 | 99.6102 | 65.7261 | 0.004269 | 1.0418 | 0.02555 | 24.5249 | -0.011304 | 0.14159 | -0.000134 | -0.010433 | -0.010433 | -0.56458
2019-03-19 16:06 | 1553025975.7822 | 1446 | 0.20953 | 2.7e-05 | 0 | 100.2903 | 100.1513 | 127.7448 | 21.2094 | 0.001875 | 0.001795 | -152.6742 | 0.23756 | 8.4016 | 0.003247 | 0.005166 | 0.004985 | 5.8e-05 | 0.001126 | 0.34291 | 1.3582 | 2.5481 | 101.18 | 20.82 | 17.54 | 66.8948 | 99.9246 | 64.972 | 0.004269 | 1.0441 | 0.02557 | 24.489 | 0.012411 | 0.14132 | -7.9e-05 | -0.006131 | -0.006131 | -0.95892
2019-03-19 16:06 | 1553025976.7809 | 1447 | 0.20954 | 2.9e-05 | 1e-06 | 100.239 | 100.3181 | 127.7699 | 21.2086 | 0.001887 | 0.001796 | -152.6938 | 0.23671 | 8.401 | 0.003332 | 0.005044 | 0.004816 | 5.5e-05 | 0.001112 | 0.3027 | 1.3535 | 2.5481 | 101.18 | 20.82 | 17.537 | 67.1829 | 101.9624 | 64.9985 | 0.004268 | 1.0441 | 0.025571 | 24.4924 | -0.00705 | 0.14082 | -9.5e-05 | -0.00738 | -0.00738 | -0.79371
2019-03-19 16:06 | 1553025977.781 | 1448 | 0.20954 | 2.2e-05 | 1e-06 | 100.2648 | 100.2194 | 127.4859 | 21.2088 | 0.001846 | 0.001761 | -152.6765 | 0.23712 | 8.4046 | 0.003285 | 0.00462 | 0.00436 | 5.6e-05 | 0.001099 | 0.32068 | 1.3724 | 2.5481 | 101.18 | 20.82 | 17.529 | 66.1875 | 100.8257 | 64.1416 | 0.004266 | 1.0467 | 0.025575 | 24.4341 | 0.004051 | 0.14106 | -8.6e-05 | -0.006716 | -0.006716 | -0.87377
2019-03-19 16:06 | 1553025978.781 | 1449 | 0.20957 | 2.5e-05 | 1e-06 | 100.3011 | 100.3167 | 127.3845 | 21.2097 | 0.001818 | 0.001765 | -152.6736 | 0.23744 | 8.4039 | 0.003234 | 0.005113 | 0.004814 | 5.5e-05 | 0.001093 | 0.3183 | 1.3854 | 2.5481 | 101.18 | 20.82 | 17.524 | 65.5125 | 99.5954 | 64.258 | 0.004265 | 1.0463 | 0.025561 | 24.4286 | -0.001387 | 0.14125 | -0.000305 | -0.023546 | -0.023546 | -0.24955
2019-03-19 16:06 | 1553025979.782 | 1450 | 0.2096 | 2e-05 | 0 | 100.2717 | 100.2509 | 127.1756 | 21.21 | 0.001893 | 0.001776 | -152.6748 | 0.23747 | 8.4019 | 0.003302 | 0.00456 | 0.004427 | 5.8e-05 | 0.001091 | 0.32405 | 1.3536 | 2.548 | 101.18 | 20.82 | 17.512 | 67.3254 | 101.239 | 64.5105 | 0.004262 | 1.0456 | 0.02553 | 24.4177 | 0.001855 | 0.14127 | -0.000467 | -0.036057 | -0.036057 | -0.16299
2019-03-19 16:06 | 1553025980.7811 | 1451 | 0.20954 | 2.1e-05 | 1e-06 | 100.2987 | 100.403 | 127.2702 | 21.2093 | 0.001867 | 0.001788 | -152.6852 | 0.23688 | 8.4043 | 0.003274 | 0.004121 | 0.003952 | 5.7e-05 | 0.001093 | 0.31404 | 1.3554 | 2.5481 | 101.18 | 20.82 | 17.505 | 66.6904 | 100.564 | 64.7901 | 0.00426 | 1.0447 | 0.025529 | 24.4369 | -0.009293 | 0.14092 | -0.000114 | -0.008846 | -0.008846 | -0.66272
2019-03-19 16:06 | 1553025981.7821 | 1452 | 0.2095 | 2.6e-05 | 0 | 100.3215 | 100.5287 | 127.8368 | 21.2094 | 0.001895 | 0.001758 | -152.7016 | 0.23665 | 8.401 | 0.003223 | 0.004713 | 0.004393 | 5.8e-05 | 0.00109 | 0.29231 | 1.4053 | 2.5483 | 101.18 | 20.82 | 17.498 | 67.377 | 99.3239 | 64.089 | 0.004258 | 1.0469 | 0.025613 | 24.4658 | -0.01845 | 0.14078 | 0.000114 | 0.008791 | 0.008791 | 0.66621
2019-03-19 16:06 | 1553025982.7812 | 1453 | 0.20958 | 2.7e-05 | 0 | 100.3603 | 100.4287 | 128.0509 | 21.2092 | 0.001909 | 0.001772 | -152.6932 | 0.23684 | 8.4021 | 0.003264 | 0.004289 | 0.004132 | 5.6e-05 | 0.001093 | 0.33915 | 1.3722 | 2.5482 | 101.18 | 20.82 | 17.489 | 67.7 | 100.3136 | 64.4029 | 0.004256 | 1.0459 | 0.025622 | 24.4977 | -0.006096 | 0.14089 | -0.000379 | -0.029371 | -0.029371 | -0.19956
2019-03-19 16:06 | 1553025983.7813 | 1454 | 0.20949 | 2.6e-05 | 1e-06 | 100.3179 | 100.4962 | 127.9877 | 21.2084 | 0.001887 | 0.001766 | -152.6844 | 0.23716 | 8.4008 | 0.003282 | 0.004742 | 0.004539 | 5.4e-05 | 0.00109 | 0.31116 | 1.3692 | 2.5481 | 101.18 | 20.82 | 17.486 | 67.1771 | 100.7586 | 64.2685 | 0.004256 | 1.0463 | 0.025621 | 24.4868 | -0.015876 | 0.14108 | 0.000164 | 0.012671 | 0.012671 | 0.46321
2019-03-19 16:06 | 1553025984.7815 | 1455 | 0.20958 | 2.7e-05 | 1e-06 | 100.3173 | 100.3164 | 127.5436 | 21.2096 | 0.001966 | 0.001752 | -152.6616 | 0.23793 | 8.4038 | 0.003247 | 0.00455 | 0.004299 | 5.2e-05 | 0.001093 | 0.28816 | 1.3992 | 2.5481 | 101.18 | 20.82 | 17.483 | 69.0893 | 99.9108 | 63.9433 | 0.004255 | 1.0473 | 0.025589 | 24.4325 | 8.9e-05 | 0.14155 | -0.000384 | -0.029736 | -0.029736 | -0.19802
2019-03-19 16:06 | 1553025985.7813 | 1456 | 0.20949 | 2.4e-05 | 0 | 100.296 | 100.3003 | 127.2913 | 21.209 | 0.001882 | 0.00181 | -152.6615 | 0.23686 | 8.405 | 0.003294 | 0.005054 | 0.004869 | 5.3e-05 | 0.001091 | 0.30198 | 1.3288 | 2.5481 | 101.18 | 20.83 | 17.485 | 67.0505 | 101.0456 | 65.3378 | 0.004258 | 1.043 | 0.025511 | 24.4587 | -0.000376 | 0.14091 | 0.000213 | 0.01637 | 0.01637 | 0.35809
2019-03-19 16:06 | 1553025986.7805 | 1457 | 0.20951 | 3e-05 | 1e-06 | 100.3915 | 100.5236 | 127.0585 | 21.2091 | 0.00192 | 0.001736 | -152.6488 | 0.2374 | 8.4023 | 0.003322 | 0.004469 | 0.004396 | 5.2e-05 | 0.001099 | 0.32637 | 1.3389 | 2.5482 | 101.18 | 20.83 | 17.489 | 67.9665 | 101.7155 | 63.5578 | 0.004259 | 1.0485 | 0.025555 | 24.372 | -0.011761 | 0.14123 | 4e-05 | 0.003055 | 0.003055 | 1.9229
2019-03-19 16:06 | 1553025987.7815 | 1458 | 0.20952 | 2.3e-05 | 0 | 100.3822 | 100.4859 | 127.4012 | 21.2079 | 0.0019 | 0.001762 | -152.6764 | 0.23742 | 8.4014 | 0.00328 | 0.005336 | 0.005038 | 5.3e-05 | 0.001106 | 0.2907 | 1.3143 | 2.5481 | 101.18 | 20.83 | 17.501 | 67.4957 | 100.7017 | 64.1845 | 0.004262 | 1.0466 | 0.025565 | 24.4275 | -0.009237 | 0.14124 | -1e-06 | -7e-05 | -7e-05 | -83.3854
2019-03-19 16:06 | 1553025988.7805 | 1459 | 0.20955 | 3e-05 | 1e-06 | 100.3429 | 100.3602 | 127.6366 | 21.2082 | 0.001848 | 0.001801 | -152.678 | 0.23755 | 8.4039 | 0.003375 | 0.004629 | 0.004477 | 5.4e-05 | 0.00111 | 0.2896 | 1.2989 | 2.5481 | 101.18 | 20.83 | 17.508 | 66.2376 | 103.0021 | 65.1208 | 0.004263 | 1.0437 | 0.025553 | 24.484 | -0.001545 | 0.14132 | -0.000149 | -0.011555 | -0.011555 | -0.50878
2019-03-19 16:06 | 1553025989.7819 | 1460 | 0.20952 | 2.5e-05 | 1e-06 | 100.4371 | 100.4674 | 127.804 | 21.2088 | 0.001867 | 0.001793 | -152.6715 | 0.23714 | 8.4087 | 0.003324 | 0.004808 | 0.004607 | 5.6e-05 | 0.001112 | 0.30845 | 1.3447 | 2.548 | 101.18 | 20.83 | 17.518 | 66.6999 | 101.7772 | 64.9231 | 0.004266 | 1.0443 | 0.025578 | 24.4929 | -0.002698 | 0.14107 | 2e-06 | 0.000113 | 0.000113 | 51.8676
2019-03-19 16:06 | 1553025990.7817 | 1461 | 0.20951 | 2.2e-05 | 1e-06 | 100.3492 | 100.2859 | 127.6244 | 21.2088 | 0.001923 | 0.001755 | -152.6616 | 0.23762 | 8.4039 | 0.003246 | 0.004964 | 0.004711 | 5.2e-05 | 0.001114 | 0.29966 | 1.3648 | 2.5483 | 101.18 | 20.83 | 17.54 | 68.0474 | 99.8856 | 64.0159 | 0.004271 | 1.0471 | 0.025594 | 24.4428 | 0.005643 | 0.14136 | 5.1e-05 | 0.003877 | 0.003877 | 1.517
2019-03-19 16:06 | 1553025991.7817 | 1462 | 0.20949 | 2.1e-05 | 1e-06 | 100.3849 | 100.4071 | 127.3204 | 21.2094 | 0.001882 | 0.001811 | -152.6534 | 0.23797 | 8.403 | 0.003259 | 0.004284 | 0.004112 | 5.6e-05 | 0.001117 | 0.32156 | 1.3668 | 2.5481 | 101.18 | 20.83 | 17.553 | 67.0634 | 100.2032 | 65.3427 | 0.004274 | 1.043 | 0.025513 | 24.4617 | -0.001982 | 0.14157 | 0.000209 | 0.016032 | 0.016032 | 0.36735
2019-03-19 16:07 | 1553025992.7811 | 1463 | 0.20944 | 2.6e-05 | 1e-06 | 100.3903 | 100.4033 | 127.2847 | 21.2085 | 0.001838 | 0.00181 | -152.661 | 0.23683 | 8.4022 | 0.003295 | 0.00472 | 0.004505 | 5.5e-05 | 0.001118 | 0.28982 | 1.3785 | 2.5481 | 101.18 | 20.83 | 17.581 | 65.9897 | 101.0703 | 65.3176 | 0.004281 | 1.0431 | 0.025511 | 24.4574 | -0.001165 | 0.14089 | 0.000524 | 0.040276 | 0.040276 | 0.14552
2019-03-19 16:07 | 1553025993.7808 | 1464 | 0.20949 | 2.2e-05 | 1e-06 | 100.3912 | 100.4492 | 127.3815 | 21.2077 | 0.001898 | 0.001767 | -152.6557 | 0.23809 | 8.3988 | 0.003276 | 0.005053 | 0.00481 | 5.7e-05 | 0.001114 | 0.30807 | 1.3651 | 2.5482 | 101.18 | 20.83 | 17.594 | 67.4486 | 100.6064 | 64.3022 | 0.004284 | 1.0462 | 0.025559 | 24.4299 | -0.005172 | 0.14164 | 0.000183 | 0.014081 | 0.014081 | 0.41845
2019-03-19 16:07 | 1553025994.7822 | 1465 | 0.20955 | 2.2e-05 | 0 | 100.3624 | 100.3965 | 127.7842 | 21.2083 | 0.001928 | 0.001749 | -152.6703 | 0.23705 | 8.4024 | 0.003327 | 0.004501 | 0.00431 | 5.9e-05 | 0.001116 | 0.30237 | 1.3505 | 2.5481 | 101.18 | 20.83 | 17.607 | 68.1606 | 101.8505 | 63.853 | 0.004288 | 1.0476 | 0.025616 | 24.4522 | -0.003041 | 0.14102 | -0.000144 | -0.011147 | -0.011147 | -0.52627
2019-03-19 16:07 | 1553025995.7819 | 1466 | 0.20957 | 2.4e-05 | 1e-06 | 100.3891 | 100.3979 | 128.0781 | 21.2085 | 0.001881 | 0.001807 | -152.6755 | 0.2377 | 8.4027 | 0.003358 | 0.004536 | 0.004294 | 5.5e-05 | 0.001117 | 0.28727 | 1.3556 | 2.5481 | 101.18 | 20.83 | 17.629 | 67.0476 | 102.604 | 65.2639 | 0.004293 | 1.0432 | 0.025592 | 24.5315 | -0.000783 | 0.14141 | -0.000257 | -0.019933 | -0.019933 | -0.29512
2019-03-19 16:07 | 1553025996.782 | 1467 | 0.20942 | 2.1e-05 | 1e-06 | 100.4029 | 100.4849 | 128.1321 | 21.2083 | 0.001911 | 0.001816 | -152.6862 | 0.23743 | 8.4022 | 0.003288 | 0.004788 | 0.004578 | 6e-05 | 0.001119 | 0.31094 | 1.3559 | 2.5482 | 101.18 | 20.83 | 17.638 | 67.7582 | 100.9068 | 65.4815 | 0.004295 | 1.0426 | 0.025589 | 24.5446 | -0.0073 | 0.14125 | 0.000646 | 0.049869 | 0.049869 | 0.11782
2019-03-19 16:07 | 1553025997.781 | 1468 | 0.20958 | 2.6e-05 | 1e-06 | 100.3768 | 100.414 | 128.0391 | 21.2079 | 0.001914 | 0.001774 | -152.6702 | 0.23836 | 8.4035 | 0.003357 | 0.004644 | 0.004441 | 5.6e-05 | 0.001119 | 0.31924 | 1.3823 | 2.5483 | 101.18 | 20.83 | 17.654 | 67.8241 | 102.5829 | 64.4561 | 0.004299 | 1.0457 | 0.025619 | 24.4985 | -0.003314 | 0.1418 | -0.000362 | -0.028005 | -0.028005 | -0.21064
2019-03-19 16:07 | 1553025998.7811 | 1469 | 0.20943 | 2.3e-05 | 1e-06 | 100.378 | 100.6106 | 127.991 | 21.2081 | 0.001912 | 0.001787 | -152.6803 | 0.23766 | 8.4045 | 0.003307 | 0.004977 | 0.004748 | 5.7e-05 | 0.001121 | 0.29031 | 1.3364 | 2.5482 | 101.18 | 20.83 | 17.661 | 67.7704 | 101.3517 | 64.7638 | 0.004301 | 1.0448 | 0.025602 | 24.505 | -0.020701 | 0.14139 | 0.000589 | 0.045501 | 0.045501 | 0.12926
2019-03-19 16:07 | 1553025999.7815 | 1470 | 0.20953 | 2.4e-05 | 0 | 100.4407 | 100.2078 | 127.926 | 21.2081 | 0.001826 | 0.001763 | -152.685 | 0.23739 | 8.3996 | 0.003352 | 0.004828 | 0.004694 | 5.8e-05 | 0.001124 | 0.30237 | 1.3404 | 2.5481 | 101.18 | 20.83 | 17.666 | 65.7233 | 102.4505 | 64.191 | 0.004302 | 1.0466 | 0.025618 | 24.478 | 0.020765 | 0.14122 | -5.9e-05 | -0.004615 | -0.004615 | -1.273
2019-03-19 16:07 | 1553026000.7805 | 1471 | 0.20954 | 3.1e-05 | 0 | 100.3825 | 100.2283 | 127.8035 | 21.2084 | 0.001889 | 0.001715 | -152.6863 | 0.23759 | 8.4028 | 0.0033 | 0.004635 | 0.004394 | 5.9e-05 | 0.001129 | 0.29911 | 1.2762 | 2.548 | 101.18 | 20.83 | 17.675 | 67.2192 | 101.1989 | 63.0532 | 0.004304 | 1.0501 | 0.025649 | 24.425 | 0.013745 | 0.14134 | -0.000105 | -0.008126 | -0.008126 | -0.7236
2019-03-19 16:07 | 1553026001.7803 | 1472 | 0.20941 | 2.6e-05 | 0 | 100.4326 | 100.3283 | 127.791 | 21.208 | 0.001879 | 0.001773 | -152.6901 | 0.23808 | 8.4016 | 0.003345 | 0.004026 | 0.00396 | 5.5e-05 | 0.001131 | 0.29966 | 1.3369 | 2.5479 | 101.18 | 20.83 | 17.68 | 66.9957 | 102.2932 | 64.432 | 0.004305 | 1.0458 | 0.025595 | 24.4739 | 0.009289 | 0.14163 | 0.000665 | 0.051372 | 0.051372 | 0.11469
2019-03-19 16:07 | 1553026002.7803 | 1473 | 0.20962 | 2.3e-05 | 0 | 100.4335 | 100.3338 | 127.9971 | 21.2086 | 0.001912 | 0.0018 | -152.6916 | 0.23753 | 8.4029 | 0.003322 | 0.004653 | 0.0044 | 5.5e-05 | 0.001126 | 0.30469 | 1.3409 | 2.548 | 101.18 | 20.83 | 17.687 | 67.781 | 101.7385 | 65.0975 | 0.004307 | 1.0437 | 0.02559 | 24.5177 | 0.008881 | 0.1413 | -0.000559 | -0.043208 | -0.043208 | -0.13605
2019-03-19 16:07 | 1553026003.7817 | 1474 | 0.20943 | 2e-05 | 1e-06 | 100.4446 | 100.2951 | 128.0692 | 21.2074 | 0.001849 | 0.001784 | -152.6716 | 0.23831 | 8.4045 | 0.00337 | 0.005079 | 0.004947 | 5.8e-05 | 0.001124 | 0.30469 | 1.3235 | 2.548 | 101.18 | 20.83 | 17.689 | 66.2547 | 102.8849 | 64.7132 | 0.004308 | 1.0449 | 0.025612 | 24.5107 | 0.01332 | 0.14177 | 0.000577 | 0.044561 | 0.044561 | 0.13235
2019-03-19 16:07 | 1553026004.7813 | 1475 | 0.20954 | 2.3e-05 | 1e-06 | 100.3954 | 100.2848 | 127.8777 | 21.2079 | 0.001923 | 0.001754 | -152.6622 | 0.23705 | 8.4046 | 0.003308 | 0.004606 | 0.004379 | 6e-05 | 0.001125 | 0.31238 | 1.315 | 2.5481 | 101.18 | 20.83 | 17.693 | 68.0379 | 101.3939 | 63.9838 | 0.004309 | 1.0472 | 0.025621 | 24.4659 | 0.009853 | 0.14102 | -0.000113 | -0.008774 | -0.008774 | -0.66864
2019-03-19 16:07 | 1553026005.7805 | 1476 | 0.2096 | 2.5e-05 | 0 | 100.4899 | 100.6092 | 127.6371 | 21.2088 | 0.001904 | 0.00172 | -152.6737 | 0.23714 | 8.4041 | 0.003361 | 0.004079 | 0.003827 | 5.6e-05 | 0.001128 | 0.32156 | 1.373 | 2.5481 | 101.18 | 20.83 | 17.699 | 67.5964 | 102.6719 | 63.1733 | 0.00431 | 1.0497 | 0.025627 | 24.4135 | -0.01061 | 0.14108 | -0.000492 | -0.038031 | -0.038031 | -0.15432
2019-03-19 16:07 | 1553026006.7814 | 1477 | 0.20944 | 2.3e-05 | 0 | 100.4146 | 100.504 | 127.6237 | 21.2073 | 0.001922 | 0.001747 | -152.6769 | 0.23774 | 8.404 | 0.00337 | 0.004761 | 0.004504 | 5.7e-05 | 0.001123 | 0.3027 | 1.3454 | 2.5481 | 101.18 | 20.83 | 17.701 | 68.0154 | 102.8983 | 63.8162 | 0.00431 | 1.0477 | 0.025602 | 24.4355 | -0.007965 | 0.14143 | 0.000501 | 0.038743 | 0.038743 | 0.15186
2019-03-19 16:07 | 1553026007.7819 | 1478 | 0.20963 | 2.4e-05 | 0 | 100.4605 | 100.4506 | 127.6903 | 21.208 | 0.001905 | 0.001761 | -152.6642 | 0.23738 | 8.4046 | 0.00329 | 0.004394 | 0.004174 | 5.8e-05 | 0.001127 | 0.33555 | 1.3211 | 2.548 | 101.18 | 20.83 | 17.704 | 67.6222 | 100.941 | 64.1525 | 0.004311 | 1.0467 | 0.025595 | 24.4541 | 0.00088 | 0.14122 | -0.000649 | -0.050185 | -0.050185 | -0.11706
2019-03-19 16:07 | 1553026008.7817 | 1479 | 0.20959 | 2.1e-05 | 0 | 100.4608 | 100.4883 | 127.8014 | 21.2086 | 0.001848 | 0.001774 | -152.6539 | 0.2369 | 8.4011 | 0.003283 | 0.0046 | 0.004372 | 5.6e-05 | 0.001122 | 0.31686 | 1.3212 | 2.5482 | 101.18 | 20.83 | 17.707 | 66.252 | 100.7731 | 64.4569 | 0.004312 | 1.0457 | 0.025595 | 24.4758 | -0.002447 | 0.14093 | -0.000381 | -0.02946 | -0.02946 | -0.19901
2019-03-19 16:07 | 1553026009.7817 | 1480 | 0.20953 | 2.2e-05 | 0 | 100.4347 | 100.389 | 128.2689 | 21.2085 | 0.001797 | 0.001799 | -152.6538 | 0.2377 | 8.4029 | 0.003313 | 0.004371 | 0.004237 | 6.1e-05 | 0.001126 | 0.30198 | 1.3195 | 2.5481 | 101.18 | 20.84 | 17.708 | 65.007 | 101.4966 | 65.0609 | 0.004315 | 1.0439 | 0.025619 | 24.5424 | 0.004073 | 0.14141 | -2.6e-05 | -0.001996 | -0.001996 | -2.9478
2019-03-19 16:07 | 1553026010.7817 | 1481 | 0.20954 | 2.3e-05 | 0 | 100.4068 | 100.7301 | 128.5311 | 21.2079 | 0.001916 | 0.001767 | -152.6466 | 0.23722 | 8.4009 | 0.003273 | 0.004121 | 0.003951 | 5.7e-05 | 0.001121 | 0.31421 | 1.301 | 2.5482 | 101.18 | 20.84 | 17.71 | 67.8772 | 100.543 | 64.3 | 0.004315 | 1.0462 | 0.025674 | 24.5398 | -0.028755 | 0.14112 | -8e-05 | -0.006205 | -0.006205 | -0.94613
2019-03-19 16:07 | 1553026011.7821 | 1482 | 0.20951 | 1.8e-05 | 1e-06 | 100.4581 | 100.476 | 128.6098 | 21.2082 | 0.00182 | 0.001749 | -152.6496 | 0.2376 | 8.4031 | 0.003229 | 0.004281 | 0.004059 | 5.8e-05 | 0.001119 | 0.32676 | 1.3129 | 2.5481 | 101.18 | 20.84 | 17.71 | 65.5711 | 99.4754 | 63.8693 | 0.004315 | 1.0476 | 0.025698 | 24.5317 | -0.00159 | 0.14135 | 6.8e-05 | 0.005231 | 0.005231 | 1.1241
2019-03-19 16:07 | 1553026012.7818 | 1483 | 0.20951 | 2.3e-05 | 1e-06 | 100.4098 | 100.4259 | 128.2349 | 21.2081 | 0.001878 | 0.001742 | -152.6495 | 0.23854 | 8.4038 | 0.003269 | 0.005119 | 0.004868 | 5.8e-05 | 0.001125 | 0.31133 | 1.3135 | 2.5479 | 101.18 | 20.84 | 17.71 | 66.9537 | 100.4366 | 63.6991 | 0.004315 | 1.0481 | 0.025667 | 24.4897 | -0.00144 | 0.14191 | 8.8e-05 | 0.006776 | 0.006776 | 0.87118
2019-03-19 16:07 | 1553026013.7809 | 1484 | 0.20962 | 2.4e-05 | 0 | 100.4644 | 100.5513 | 128.0685 | 21.2078 | 0.001873 | 0.00173 | -152.6598 | 0.23752 | 8.402 | 0.003243 | 0.004794 | 0.004635 | 5.8e-05 | 0.001119 | 0.31725 | 1.3213 | 2.5481 | 101.18 | 20.84 | 17.706 | 66.8404 | 99.8242 | 63.4059 | 0.004314 | 1.049 | 0.025662 | 24.4632 | -0.007734 | 0.1413 | -0.000563 | -0.043594 | -0.043594 | -0.13484
2019-03-19 16:07 | 1553026014.7812 | 1485 | 0.20954 | 2.6e-05 | 0 | 100.429 | 100.4496 | 128.3642 | 21.2087 | 0.001862 | 0.001782 | -152.6602 | 0.23784 | 8.4011 | 0.003251 | 0.004256 | 0.004021 | 5.5e-05 | 0.001118 | 0.32444 | 1.3111 | 2.5481 | 101.18 | 20.84 | 17.705 | 66.5782 | 99.997 | 64.6653 | 0.004314 | 1.0451 | 0.025643 | 24.5372 | -0.001834 | 0.14149 | -0.00012 | -0.009266 | -0.009266 | -0.63523
2019-03-19 16:07 | 1553026015.781 | 1486 | 0.20953 | 2.1e-05 | 0 | 100.3987 | 100.5427 | 128.4147 | 21.2084 | 0.00192 | 0.001764 | -152.6611 | 0.23677 | 8.4024 | 0.003347 | 0.004397 | 0.004196 | 5.7e-05 | 0.001117 | 0.28998 | 1.2983 | 2.5482 | 101.18 | 20.84 | 17.7 | 67.979 | 102.3202 | 64.2317 | 0.004313 | 1.0464 | 0.025665 | 24.5262 | -0.012825 | 0.14086 | -6.7e-05 | -0.005199 | -0.005199 | -1.127
2019-03-19 16:07 | 1553026016.78 | 1487 | 0.2095 | 2.7e-05 | 0 | 100.4959 | 100.5273 | 128.0729 | 21.2089 | 0.001885 | 0.001768 | -152.6603 | 0.23757 | 8.4009 | 0.003361 | 0.0041 | 0.003921 | 5.7e-05 | 0.001102 | 0.32477 | 1.3345 | 2.5481 | 101.18 | 20.84 | 17.697 | 67.1237 | 102.6822 | 64.3228 | 0.004312 | 1.0461 | 0.025627 | 24.4969 | -0.002796 | 0.14133 | 0.000163 | 0.012647 | 0.012647 | 0.4649
2019-03-19 16:07 | 1553026017.7801 | 1488 | 0.20941 | 2.1e-05 | 0 | 100.4995 | 100.5948 | 127.9321 | 21.2088 | 0.001862 | 0.001789 | -152.6846 | 0.23694 | 8.4008 | 0.003293 | 0.004942 | 0.004665 | 5.6e-05 | 0.001106 | 0.33069 | 1.3361 | 2.5482 | 101.18 | 20.84 | 17.694 | 66.5862 | 101.0251 | 64.824 | 0.004311 | 1.0446 | 0.025594 | 24.5016 | -0.008477 | 0.14095 | 0.000693 | 0.053529 | 0.053529 | 0.10954
2019-03-19 16:07 | 1553026018.7815 | 1489 | 0.20978 | 2.3e-05 | 0 | 100.4512 | 100.4804 | 127.7272 | 21.2072 | 0.001841 | 0.001784 | -152.6828 | 0.239 | 8.4021 | 0.003226 | 0.004984 | 0.004778 | 5.7e-05 | 0.001109 | 0.32637 | 1.32 | 2.5481 | 101.18 | 20.84 | 17.691 | 66.0774 | 99.407 | 64.7007 | 0.004311 | 1.045 | 0.025578 | 24.4775 | -0.002601 | 0.14218 | -0.001544 | -0.1192 | -0.1192 | -0.049619
2019-03-19 16:07 | 1553026019.7802 | 1490 | 0.20935 | 2.2e-05 | 1e-06 | 100.4899 | 100.6284 | 127.4521 | 21.2087 | 0.00189 | 0.001741 | -152.6821 | 0.23753 | 8.4033 | 0.003229 | 0.004532 | 0.004381 | 5.7e-05 | 0.001113 | 0.30381 | 1.3453 | 2.5482 | 101.18 | 20.84 | 17.69 | 67.2469 | 99.4822 | 63.6645 | 0.00431 | 1.0482 | 0.02559 | 24.4136 | -0.012315 | 0.14131 | 0.001022 | 0.07892 | 0.07892 | 0.074486
2019-03-19 16:07 | 1553026020.7812 | 1491 | 0.20961 | 2.7e-05 | 0 | 100.435 | 100.5325 | 127.1465 | 21.209 | 0.001876 | 0.001755 | -152.7033 | 0.23651 | 8.4025 | 0.003247 | 0.004548 | 0.004388 | 5.6e-05 | 0.001112 | 0.31509 | 1.3804 | 2.5481 | 101.18 | 20.84 | 17.69 | 66.9242 | 99.9187 | 63.9981 | 0.00431 | 1.0472 | 0.025547 | 24.3964 | -0.008677 | 0.1407 | -0.000504 | -0.038875 | -0.038875 | -0.15057
2019-03-19 16:07 | 1553026021.7803 | 1492 | 0.20945 | 2.3e-05 | 0 | 100.4926 | 100.4393 | 127.1514 | 21.207 | 0.001831 | 0.001789 | -152.7065 | 0.23717 | 8.4032 | 0.003245 | 0.004713 | 0.004497 | 5.8e-05 | 0.001106 | 0.31188 | 1.3677 | 2.5481 | 101.18 | 20.84 | 17.691 | 65.8252 | 99.8643 | 64.8307 | 0.004311 | 1.0446 | 0.025516 | 24.427 | 0.004745 | 0.14109 | 0.000441 | 0.033967 | 0.033967 | 0.1728
2019-03-19 16:07 | 1553026022.7817 | 1493 | 0.20972 | 2.6e-05 | 1e-06 | 100.5067 | 100.5575 | 126.9913 | 21.2073 | 0.001922 | 0.001704 | -152.7054 | 0.23706 | 8.4041 | 0.003178 | 0.004068 | 0.003908 | 5.8e-05 | 0.001105 | 0.31166 | 1.3914 | 2.5482 | 101.18 | 20.84 | 17.692 | 68.0185 | 98.2317 | 62.7782 | 0.004311 | 1.051 | 0.025577 | 24.3373 | -0.004516 | 0.14103 | -0.001184 | -0.091452 | -0.091452 | -0.064151
2019-03-19 16:07 | 1553026023.7804 | 1494 | 0.20922 | 1.9e-05 | 0 | 100.5622 | 100.5623 | 126.9465 | 21.2074 | 0.001889 | 0.001787 | -152.7145 | 0.23726 | 8.4026 | 0.003226 | 0.004653 | 0.004298 | 5.7e-05 | 0.001106 | 0.30917 | 1.3463 | 2.5482 | 101.18 | 20.85 | 17.695 | 67.2339 | 99.3949 | 64.768 | 0.004314 | 1.0448 | 0.025498 | 24.405 | -3e-06 | 0.14115 | 0.001862 | 0.14332 | 0.14332 | 0.040968
2019-03-19 16:07 | 1553026024.7815 | 1495 | 0.20993 | 2.2e-05 | 0 | 100.5088 | 100.5931 | 126.5264 | 21.2073 | 0.001901 | 0.001782 | -152.7149 | 0.23775 | 8.4031 | 0.003227 | 0.003874 | 0.003725 | 5.6e-05 | 0.001111 | 0.30525 | 1.3444 | 2.548 | 101.18 | 20.84 | 17.696 | 67.5045 | 99.4269 | 64.6452 | 0.004312 | 1.0451 | 0.02546 | 24.3602 | -0.007496 | 0.14144 | -0.002444 | -0.18794 | -0.18794 | -0.031308
2019-03-19 16:07 | 1553026025.7808 | 1496 | 0.20992 | 2.2e-05 | 0 | 100.426 | 100.3835 | 126.4806 | 21.2076 | 0.001886 | 0.001802 | -152.7078 | 0.23757 | 8.4027 | 0.003241 | 0.005417 | 0.005167 | 5.5e-05 | 0.001108 | 0.29734 | 1.3617 | 2.5481 | 101.18 | 20.84 | 17.699 | 67.1467 | 99.7722 | 65.1365 | 0.004313 | 1.0436 | 0.025437 | 24.3735 | 0.003786 | 0.14133 | -0.002398 | -0.18423 | -0.18423 | -0.031913
2019-03-19 16:07 | 1553026026.7806 | 1497 | 0.20953 | 2.7e-05 | 1e-06 | 100.4698 | 100.501 | 127.0533 | 21.2067 | 0.001849 | 0.001733 | -152.6929 | 0.23719 | 8.4045 | 0.003205 | 0.004047 | 0.003903 | 5.6e-05 | 0.001111 | 0.2672 | 1.338 | 2.548 | 101.18 | 20.85 | 17.708 | 66.2549 | 98.9062 | 63.482 | 0.004318 | 1.0488 | 0.025557 | 24.3688 | -0.002773 | 0.1411 | -2.6e-05 | -0.002043 | -0.002043 | -2.8728
2019-03-19 16:07 | 1553026027.782 | 1498 | 0.20948 | 2.3e-05 | 0 | 100.5289 | 100.4174 | 127.2711 | 21.2066 | 0.001918 | 0.001724 | -152.6778 | 0.23766 | 8.4043 | 0.003269 | 0.004228 | 0.004 | 5.5e-05 | 0.00111 | 0.31741 | 1.3334 | 2.548 | 101.18 | 20.84 | 17.713 | 67.9297 | 100.4414 | 63.254 | 0.004316 | 1.0495 | 0.025587 | 24.3814 | 0.009929 | 0.14138 | 0.000286 | 0.022053 | 0.022053 | 0.26669
2019-03-19 16:07 | 1553026028.7806 | 1499 | 0.20919 | 2.4e-05 | 0 | 100.3858 | 100.5297 | 127.2563 | 21.2056 | 0.001847 | 0.001804 | -152.6906 | 0.23783 | 8.4026 | 0.003241 | 0.005144 | 0.005001 | 5.6e-05 | 0.001108 | 0.32875 | 1.3658 | 2.5482 | 101.18 | 20.85 | 17.724 | 66.2132 | 99.7546 | 65.1837 | 0.004321 | 1.0435 | 0.025513 | 24.4498 | -0.012817 | 0.14148 | 0.002008 | 0.15458 | 0.15458 | 0.038076
2019-03-19 16:07 | 1553026029.7821 | 1500 | 0.20915 | 2.6e-05 | 1e-06 | 100.4725 | 100.5766 | 126.9646 | 21.207 | 0.001838 | 0.001751 | -152.6786 | 0.2371 | 8.4037 | 0.003285 | 0.004728 | 0.004467 | 5.6e-05 | 0.001102 | 0.30773 | 1.3632 | 2.5481 | 101.18 | 20.85 | 17.729 | 65.9897 | 100.8254 | 63.9009 | 0.004323 | 1.0475 | 0.025532 | 24.3754 | -0.009267 | 0.14105 | 0.002239 | 0.17245 | 0.17245 | 0.034026
2019-03-19 16:07 | 1553026030.7818 | 1501 | 0.20939 | 2.1e-05 | 1e-06 | 100.4626 | 100.601 | 127.1031 | 21.2062 | 0.001866 | 0.001809 | -152.6849 | 0.23747 | 8.4017 | 0.003265 | 0.004455 | 0.004262 | 5.4e-05 | 0.001104 | 0.27959 | 1.3965 | 2.5482 | 101.18 | 20.85 | 17.737 | 66.6695 | 100.3554 | 65.3091 | 0.004325 | 1.0431 | 0.025493 | 24.4396 | -0.012311 | 0.14127 | 0.000833 | 0.064022 | 0.064022 | 0.091793
2019-03-19 16:07 | 1553026031.7801 | 1502 | 0.2095 | 3.1e-05 | 1e-06 | 100.3831 | 100.3962 | 127.0139 | 21.2063 | 0.001863 | 0.001772 | -152.6595 | 0.23757 | 8.4051 | 0.0033 | 0.0043 | 0.004074 | 5.5e-05 | 0.001111 | 0.30414 | 1.4157 | 2.548 | 101.18 | 20.85 | 17.76 | 66.5914 | 101.1898 | 64.4194 | 0.00433 | 1.0458 | 0.025517 | 24.3989 | -0.001166 | 0.14133 | 0.000144 | 0.011037 | 0.011037 | 0.53272
2019-03-19 16:07 | 1553026032.7813 | 1503 | 0.20947 | 2.4e-05 | 1e-06 | 100.4182 | 100.4821 | 127.3239 | 21.2057 | 0.001917 | 0.001772 | -152.6509 | 0.23781 | 8.403 | 0.003273 | 0.004635 | 0.004343 | 5.6e-05 | 0.001116 | 0.29551 | 1.3417 | 2.5481 | 101.18 | 20.85 | 17.772 | 67.893 | 100.5332 | 64.4182 | 0.004333 | 1.0458 | 0.025549 | 24.4286 | -0.005693 | 0.14147 | 0.000301 | 0.02318 | 0.02318 | 0.2539
2019-03-19 16:07 | 1553026033.781 | 1504 | 0.20948 | 2.8e-05 | 1e-06 | 100.4578 | 100.4992 | 127.6903 | 21.2061 | 0.001905 | 0.001757 | -152.6446 | 0.23827 | 8.4018 | 0.003246 | 0.005014 | 0.004885 | 5.8e-05 | 0.001106 | 0.30486 | 1.4046 | 2.5479 | 101.18 | 20.85 | 17.798 | 67.6101 | 99.8805 | 64.0617 | 0.004339 | 1.047 | 0.025599 | 24.4508 | -0.003689 | 0.14175 | 0.000256 | 0.01973 | 0.01973 | 0.29887
2019-03-19 16:07 | 1553026034.78 | 1505 | 0.20998 | 2.4e-05 | 0 | 100.39 | 100.5787 | 127.8455 | 21.2051 | 0.001793 | 0.001759 | -152.6357 | 0.23852 | 8.4031 | 0.003278 | 0.004569 | 0.004349 | 6e-05 | 0.001102 | 0.30021 | 1.3826 | 2.5482 | 101.18 | 20.85 | 17.81 | 64.9132 | 100.6588 | 64.0934 | 0.004342 | 1.0469 | 0.025613 | 24.4668 | -0.0168 | 0.14189 | -0.002781 | -0.21508 | -0.21508 | -0.027444
2019-03-19 16:07 | 1553026035.781 | 1506 | 0.21015 | 2.5e-05 | 0 | 100.4485 | 100.4623 | 127.7054 | 21.2062 | 0.001897 | 0.001761 | -152.6257 | 0.23735 | 8.4029 | 0.003288 | 0.00427 | 0.004079 | 5.5e-05 | 0.001106 | 0.32333 | 1.3645 | 2.5481 | 101.18 | 20.85 | 17.821 | 67.42 | 100.8946 | 64.1398 | 0.004345 | 1.0467 | 0.025597 | 24.4551 | -0.001225 | 0.1412 | -0.003759 | -0.29063 | -0.29063 | -0.020211
2019-03-19 16:07 | 1553026036.7801 | 1507 | 0.20958 | 2.2e-05 | 0 | 100.4068 | 100.5496 | 127.2862 | 21.2059 | 0.001857 | 0.001774 | -152.6506 | 0.23745 | 8.4025 | 0.003303 | 0.004448 | 0.004182 | 5.8e-05 | 0.001111 | 0.3194 | 1.3635 | 2.5481 | 101.18 | 20.85 | 17.843 | 66.4482 | 101.2751 | 64.4542 | 0.00435 | 1.0457 | 0.025543 | 24.4263 | -0.012713 | 0.14126 | -0.000366 | -0.028226 | -0.028226 | -0.2082
2019-03-19 16:07 | 1553026037.7802 | 1508 | 0.20956 | 2.4e-05 | 1e-06 | 100.4821 | 100.702 | 127.3653 | 21.2046 | 0.001869 | 0.001768 | -152.6434 | 0.23804 | 8.4033 | 0.003298 | 0.004724 | 0.004555 | 5.9e-05 | 0.001108 | 0.30198 | 1.3384 | 2.5481 | 101.18 | 20.85 | 17.852 | 66.7514 | 101.149 | 64.3194 | 0.004353 | 1.0462 | 0.025557 | 24.429 | -0.019553 | 0.14161 | -0.000239 | -0.01845 | -0.01845 | -0.31929
2019-03-19 16:07 | 1553026038.7812 | 1509 | 0.20949 | 2.3e-05 | 0 | 100.4356 | 100.5917 | 127.676 | 21.2048 | 0.001913 | 0.001719 | -152.642 | 0.23803 | 8.4015 | 0.003315 | 0.004743 | 0.004567 | 5.7e-05 | 0.0011 | 0.27992 | 1.31 | 2.5481 | 101.18 | 20.84 | 17.861 | 67.8016 | 101.5511 | 63.1388 | 0.004352 | 1.0498 | 0.025633 | 24.416 | -0.013893 | 0.14161 | 0.000218 | 0.016833 | 0.016833 | 0.34995
2019-03-19 16:07 | 1553026039.7813 | 1510 | 0.20885 | 2.5e-05 | 1e-06 | 100.4173 | 100.3533 | 128.0025 | 21.2051 | 0.001843 | 0.001788 | -152.6536 | 0.23779 | 8.4034 | 0.003281 | 0.005035 | 0.004761 | 5.7e-05 | 0.001087 | 0.29192 | 1.3088 | 2.5482 | 101.18 | 20.84 | 17.878 | 66.122 | 100.7428 | 64.7853 | 0.004356 | 1.0447 | 0.025603 | 24.5069 | 0.005697 | 0.14146 | 0.004053 | 0.31306 | 0.31306 | 0.018798
2019-03-19 16:07 | 1553026040.7817 | 1511 | 0.20948 | 1.9e-05 | 1e-06 | 100.3912 | 100.3496 | 128.1335 | 21.2068 | 0.001899 | 0.001808 | -152.6398 | 0.23761 | 8.4038 | 0.00342 | 0.00519 | 0.005013 | 5.8e-05 | 0.001081 | 0.29717 | 1.3103 | 2.5481 | 101.18 | 20.84 | 17.888 | 67.4588 | 104.0898 | 65.2856 | 0.004359 | 1.0432 | 0.025597 | 24.5376 | 0.003707 | 0.14136 | 0.000255 | 0.019654 | 0.019654 | 0.2992
2019-03-19 16:07 | 1553026041.7804 | 1512 | 0.20999 | 2.3e-05 | 0 | 100.4785 | 100.4225 | 128.0391 | 21.2061 | 0.001858 | 0.001758 | -152.6593 | 0.23684 | 8.4026 | 0.003397 | 0.004294 | 0.004145 | 5.9e-05 | 0.001083 | 0.30989 | 1.3337 | 2.5479 | 101.18 | 20.84 | 17.903 | 66.4759 | 103.5327 | 64.0851 | 0.004362 | 1.0469 | 0.025633 | 24.485 | 0.004985 | 0.1409 | -0.002783 | -0.21542 | -0.21542 | -0.027209
2019-03-19 16:07 | 1553026042.7814 | 1513 | 0.20904 | 2.2e-05 | 0 | 100.4674 | 100.4629 | 127.8577 | 21.2056 | 0.001904 | 0.001773 | -152.6607 | 0.23762 | 8.4041 | 0.003369 | 0.004607 | 0.004415 | 5.9e-05 | 0.00109 | 0.32693 | 1.3378 | 2.5481 | 101.18 | 20.84 | 17.905 | 67.5887 | 102.8564 | 64.4303 | 0.004363 | 1.0458 | 0.025602 | 24.4802 | 0.000397 | 0.14136 | 0.002908 | 0.22467 | 0.22467 | 0.026174
2019-03-19 16:07 | 1553026043.7814 | 1514 | 0.20949 | 1.7e-05 | 0 | 100.4413 | 100.5616 | 127.5512 | 21.2045 | 0.001858 | 0.001786 | -152.6611 | 0.23704 | 8.402 | 0.003433 | 0.004712 | 0.004477 | 5.4e-05 | 0.001098 | 0.26488 | 1.3476 | 2.5483 | 101.18 | 20.84 | 17.906 | 66.475 | 104.3991 | 64.7392 | 0.004363 | 1.0449 | 0.025559 | 24.462 | -0.010705 | 0.14101 | 0.000206 | 0.015924 | 0.015924 | 0.36839
2019-03-19 16:07 | 1553026044.7819 | 1515 | 0.20991 | 2.2e-05 | 0 | 100.4227 | 100.7119 | 127.4979 | 21.2052 | 0.001836 | 0.001776 | -152.6738 | 0.23801 | 8.4027 | 0.003345 | 0.004976 | 0.004858 | 5.6e-05 | 0.001103 | 0.32925 | 1.3262 | 2.5482 | 101.18 | 20.84 | 17.906 | 65.9413 | 102.2803 | 64.4971 | 0.004363 | 1.0456 | 0.025563 | 24.4481 | -0.025726 | 0.14159 | -0.002309 | -0.17821 | -0.17821 | -0.033051
2019-03-19 16:07 | 1553026045.7806 | 1516 | 0.20925 | 1.8e-05 | 1e-06 | 100.4941 | 100.5421 | 127.6537 | 21.205 | 0.001853 | 0.001796 | -152.691 | 0.23708 | 8.4028 | 0.003364 | 0.003835 | 0.003604 | 5.8e-05 | 0.001104 | 0.32853 | 1.3344 | 2.5482 | 101.18 | 20.84 | 17.907 | 66.3555 | 102.7278 | 64.9967 | 0.004363 | 1.0441 | 0.02556 | 24.4812 | -0.004266 | 0.14104 | 0.001649 | 0.12715 | 0.12715 | 0.046144
2019-03-19 16:07 | 1553026046.7807 | 1517 | 0.20956 | 2.2e-05 | 1e-06 | 100.4758 | 100.5434 | 127.888 | 21.2055 | 0.001923 | 0.001756 | -152.6764 | 0.2381 | 8.4035 | 0.003242 | 0.004864 | 0.004635 | 5.7e-05 | 0.001113 | 0.31061 | 1.3192 | 2.5481 | 101.18 | 20.84 | 17.909 | 68.0527 | 99.7821 | 64.0354 | 0.004364 | 1.047 | 0.02562 | 24.4688 | -0.006018 | 0.14164 | -0.000225 | -0.01742 | -0.01742 | -0.33825
2019-03-19 16:07 | 1553026047.78 | 1518 | 0.20972 | 2.3e-05 | 1e-06 | 100.5805 | 100.6064 | 127.9342 | 21.2046 | 0.001911 | 0.001783 | -152.6929 | 0.23766 | 8.4034 | 0.003311 | 0.004604 | 0.004373 | 5.6e-05 | 0.001116 | 0.30502 | 1.2921 | 2.5482 | 101.18 | 20.84 | 17.911 | 67.7602 | 101.4568 | 64.6684 | 0.004364 | 1.0451 | 0.0256 | 24.4961 | -0.002303 | 0.14138 | -0.00116 | -0.089678 | -0.089678 | -0.065584
2019-03-19 16:07 | 1553026048.7801 | 1519 | 0.20936 | 1.9e-05 | 0 | 100.5061 | 100.5732 | 127.7103 | 21.2047 | 0.001823 | 0.001759 | -152.6827 | 0.23743 | 8.4026 | 0.003288 | 0.00395 | 0.00379 | 5.6e-05 | 0.001113 | 0.31492 | 1.3094 | 2.5478 | 101.18 | 20.84 | 17.914 | 65.6324 | 100.8962 | 64.0875 | 0.004365 | 1.0469 | 0.0256 | 24.4537 | -0.00597 | 0.14125 | 0.001012 | 0.078176 | 0.078176 | 0.075162
2019-03-19 16:07 | 1553026049.7797 | 1520 | 0.20957 | 2.1e-05 | 0 | 100.5136 | 100.4071 | 127.5744 | 21.2053 | 0.001833 | 0.001735 | -152.6686 | 0.23794 | 8.4041 | 0.00336 | 0.00456 | 0.004301 | 6e-05 | 0.001112 | 0.30397 | 1.3034 | 2.5478 | 101.18 | 20.84 | 17.919 | 65.8836 | 102.6472 | 63.53 | 0.004366 | 1.0486 | 0.025607 | 24.4204 | 0.009482 | 0.14155 | -0.000292 | -0.022588 | -0.022588 | -0.26068
2019-03-19 16:07 | 1553026050.7798 | 1521 | 0.20974 | 2.2e-05 | 0 | 100.5055 | 100.4153 | 127.4929 | 21.2051 | 0.001863 | 0.001742 | -152.652 | 0.23758 | 8.4021 | 0.003368 | 0.004253 | 0.004032 | 5.5e-05 | 0.001111 | 0.34202 | 1.3311 | 2.548 | 101.18 | 20.84 | 17.921 | 66.5896 | 102.8325 | 63.6819 | 0.004367 | 1.0481 | 0.025594 | 24.4181 | 0.008029 | 0.14134 | -0.001303 | -0.10068 | -0.10068 | -0.058399
2019-03-19 16:08 | 1553026051.78 | 1522 | 0.2093 | 2.3e-05 | 0 | 100.5238 | 100.5123 | 127.7098 | 21.2049 | 0.001836 | 0.001772 | -152.6473 | 0.23777 | 8.4039 | 0.003366 | 0.004944 | 0.004778 | 5.9e-05 | 0.001115 | 0.31598 | 1.3221 | 2.5479 | 101.18 | 20.84 | 17.923 | 65.9479 | 102.7889 | 64.4133 | 0.004367 | 1.0459 | 0.025588 | 24.4654 | 0.001029 | 0.14145 | 0.001331 | 0.10281 | 0.10281 | 0.057235
2019-03-19 16:08 | 1553026052.78 | 1523 | 0.20957 | 1.7e-05 | 0 | 100.579 | 100.5866 | 127.9105 | 21.205 | 0.001893 | 0.001785 | -152.6644 | 0.23775 | 8.4061 | 0.003275 | 0.004685 | 0.004451 | 5.9e-05 | 0.001114 | 0.29479 | 1.3275 | 2.5483 | 101.18 | 20.84 | 17.924 | 67.3198 | 100.5881 | 64.7288 | 0.004367 | 1.0449 | 0.025596 | 24.4961 | -0.000671 | 0.14144 | -0.000255 | -0.019709 | -0.019709 | -0.29854
2019-03-19 16:08 | 1553026053.78 | 1524 | 0.20974 | 2.2e-05 | 0 | 100.6313 | 100.5941 | 128.2506 | 21.206 | 0.001893 | 0.0018 | -152.667 | 0.2376 | 8.4016 | 0.003408 | 0.004519 | 0.004269 | 5.7e-05 | 0.001117 | 0.31741 | 1.3525 | 2.5481 | 101.18 | 20.84 | 17.924 | 67.3239 | 103.8102 | 65.0811 | 0.004367 | 1.0438 | 0.025616 | 24.5414 | 0.003302 | 0.14135 | -0.001311 | -0.10137 | -0.10137 | -0.058008
2019-03-19 16:08 | 1553026054.7801 | 1525 | 0.20957 | 2e-05 | 0 | 100.5658 | 100.7208 | 128.51 | 21.2054 | 0.001829 | 0.001769 | -152.6636 | 0.23729 | 8.4019 | 0.003427 | 0.005 | 0.004784 | 5.5e-05 | 0.001114 | 0.32875 | 1.3333 | 2.5481 | 101.18 | 20.84 | 17.925 | 65.7846 | 104.2646 | 64.3459 | 0.004368 | 1.0461 | 0.02567 | 24.5395 | -0.013775 | 0.14116 | -0.000296 | -0.022996 | -0.022996 | -0.25536
2019-03-19 16:08 | 1553026055.7804 | 1526 | 0.20954 | 2e-05 | 1e-06 | 100.6271 | 100.6232 | 128.3809 | 21.2047 | 0.001853 | 0.001804 | -152.6528 | 0.23758 | 8.4016 | 0.003354 | 0.004397 | 0.004245 | 5.4e-05 | 0.001115 | 0.32947 | 1.3502 | 2.5483 | 101.18 | 20.84 | 17.926 | 66.3636 | 102.4856 | 65.179 | 0.004368 | 1.0435 | 0.025626 | 24.5574 | 0.00034 | 0.14134 | -0.000114 | -0.008845 | -0.008845 | -0.6647
2019-03-19 16:08 | 1553026056.7802 | 1527 | 0.20953 | 1.8e-05 | 1e-06 | 100.5631 | 100.6534 | 128.153 | 21.2054 | 0.001844 | 0.001773 | -152.626 | 0.23784 | 8.4024 | 0.003364 | 0.004509 | 0.004358 | 5.2e-05 | 0.001116 | 0.32947 | 1.3184 | 2.5482 | 101.18 | 20.84 | 17.927 | 66.1473 | 102.7362 | 64.4232 | 0.004368 | 1.0458 | 0.025631 | 24.5082 | -0.008022 | 0.14149 | -5e-06 | -0.000444 | -0.000444 | -13.2709
2019-03-19 16:08 | 1553026057.7816 | 1528 | 0.20955 | 2.1e-05 | 0 | 100.6586 | 100.6612 | 127.9075 | 21.205 | 0.001852 | 0.001705 | -152.6298 | 0.23747 | 8.4032 | 0.003369 | 0.004641 | 0.004373 | 5.5e-05 | 0.001122 | 0.28799 | 1.383 | 2.548 | 101.18 | 20.84 | 17.933 | 66.3358 | 102.8583 | 62.8025 | 0.00437 | 1.0509 | 0.025669 | 24.4259 | -0.000238 | 0.14127 | -0.000171 | -0.013287 | -0.013287 | -0.4423
2019-03-19 16:08 | 1553026058.7813 | 1529 | 0.20954 | 2.2e-05 | 0 | 100.6181 | 100.6407 | 128.1844 | 21.2042 | 0.001855 | 0.001776 | -152.6574 | 0.23836 | 8.4041 | 0.003423 | 0.004429 | 0.004235 | 5.3e-05 | 0.00112 | 0.33251 | 1.3787 | 2.5481 | 101.18 | 20.84 | 17.937 | 66.3944 | 104.1686 | 64.5065 | 0.004371 | 1.0456 | 0.025631 | 24.5142 | -0.002012 | 0.1418 | -7.8e-05 | -0.00607 | -0.00607 | -0.97174
2019-03-19 16:08 | 1553026059.7803 | 1530 | 0.20955 | 2.3e-05 | 0 | 100.5952 | 100.5468 | 128.2492 | 21.205 | 0.001839 | 0.001787 | -152.675 | 0.2372 | 8.4031 | 0.003292 | 0.005104 | 0.004931 | 5.1e-05 | 0.001117 | 0.30596 | 1.3305 | 2.5482 | 101.18 | 20.84 | 17.937 | 66.0103 | 101.0057 | 64.7695 | 0.004371 | 1.0448 | 0.025628 | 24.53 | 0.004304 | 0.14111 | -0.000179 | -0.013861 | -0.013861 | -0.4235
2019-03-19 16:08 | 1553026060.7804 | 1531 | 0.20955 | 2.4e-05 | 0 | 100.6238 | 100.8767 | 128.072 | 21.2046 | 0.001846 | 0.001744 | -152.6986 | 0.23704 | 8.4045 | 0.00333 | 0.004894 | 0.004697 | 5.4e-05 | 0.001113 | 0.32228 | 1.3194 | 2.548 | 101.18 | 20.84 | 17.944 | 66.1911 | 101.9224 | 63.7277 | 0.004372 | 1.048 | 0.02565 | 24.4752 | -0.022454 | 0.14101 | -0.000132 | -0.010201 | -0.010201 | -0.57505
2019-03-19 16:08 | 1553026061.7818 | 1532 | 0.2095 | 2.3e-05 | 0 | 100.622 | 100.7304 | 127.6563 | 21.2055 | 0.001842 | 0.001734 | -152.7071 | 0.23684 | 8.401 | 0.003291 | 0.004624 | 0.004385 | 5.4e-05 | 0.001109 | 0.28816 | 1.3371 | 2.548 | 101.18 | 20.84 | 17.947 | 66.1008 | 100.9819 | 63.4965 | 0.004373 | 1.0487 | 0.025617 | 24.427 | -0.009637 | 0.14089 | 0.000157 | 0.012118 | 0.012118 | 0.48368
2019-03-19 16:08 | 1553026062.7798 | 1533 | 0.20948 | 2.3e-05 | 0 | 100.6652 | 100.6797 | 127.5782 | 21.2047 | 0.001879 | 0.001759 | -152.6767 | 0.23699 | 8.4024 | 0.003349 | 0.005023 | 0.004762 | 5.2e-05 | 0.001108 | 0.30646 | 1.359 | 2.5482 | 101.18 | 20.84 | 17.949 | 66.9918 | 102.3725 | 64.09 | 0.004374 | 1.0469 | 0.025587 | 24.4411 | -0.001294 | 0.14098 | 0.000249 | 0.019245 | 0.019245 | 0.30475
2019-03-19 16:08 | 1553026063.7796 | 1534 | 0.20957 | 1.5e-05 | 0 | 100.7372 | 100.7205 | 127.4108 | 21.205 | 0.001879 | 0.0018 | -152.6835 | 0.23758 | 8.4021 | 0.003359 | 0.004947 | 0.004799 | 5.1e-05 | 0.001103 | 0.3131 | 1.3309 | 2.5481 | 101.18 | 20.84 | 17.952 | 66.9913 | 102.6297 | 65.0818 | 0.004374 | 1.0438 | 0.025532 | 24.4609 | 0.001483 | 0.14134 | -0.000239 | -0.018434 | -0.018434 | -0.31895
2019-03-19 16:08 | 1553026064.7806 | 1535 | 0.20957 | 2e-05 | 0 | 100.6973 | 100.7859 | 127.5279 | 21.2046 | 0.001861 | 0.001735 | -152.6716 | 0.23746 | 8.4027 | 0.003354 | 0.003669 | 0.00347 | 5.1e-05 | 0.001092 | 0.29264 | 1.3236 | 2.548 | 101.18 | 20.84 | 17.952 | 66.5508 | 102.5005 | 63.5143 | 0.004374 | 1.0487 | 0.025603 | 24.4154 | -0.007869 | 0.14127 | -0.00026 | -0.020055 | -0.020055 | -0.29302
2019-03-19 16:08 | 1553026065.7796 | 1536 | 0.20965 | 2.6e-05 | 0 | 100.6751 | 100.8476 | 127.85 | 21.2048 | 0.001926 | 0.001792 | -152.6723 | 0.23697 | 8.4015 | 0.003367 | 0.004978 | 0.004761 | 5.1e-05 | 0.001085 | 0.31813 | 1.3209 | 2.548 | 101.18 | 20.83 | 17.953 | 68.1176 | 102.8063 | 64.8863 | 0.004372 | 1.0444 | 0.025584 | 24.496 | -0.015312 | 0.14097 | -0.000778 | -0.060107 | -0.060107 | -0.097566
2019-03-19 16:08 | 1553026066.7797 | 1537 | 0.20947 | 1.7e-05 | 0 | 100.7267 | 100.7441 | 127.9159 | 21.2048 | 0.001808 | 0.001764 | -152.7023 | 0.23768 | 8.4039 | 0.003337 | 0.004357 | 0.00418 | 4.8e-05 | 0.001076 | 0.29502 | 1.3347 | 2.5481 | 101.18 | 20.83 | 17.951 | 65.2695 | 102.0969 | 64.2205 | 0.004371 | 1.0465 | 0.025615 | 24.4782 | -0.001548 | 0.1414 | 0.000313 | 0.024216 | 0.024216 | 0.2429
2019-03-19 16:08 | 1553026067.7798 | 1538 | 0.20958 | 2.1e-05 | 1e-06 | 100.7639 | 100.7876 | 127.7298 | 21.2053 | 0.001823 | 0.001656 | -152.6959 | 0.23661 | 8.4034 | 0.00335 | 0.004978 | 0.004744 | 4.8e-05 | 0.001073 | 0.31614 | 1.3701 | 2.5481 | 101.18 | 20.83 | 17.95 | 65.6493 | 102.4116 | 61.6245 | 0.004371 | 1.0546 | 0.025696 | 24.366 | -0.002105 | 0.14076 | -0.000314 | -0.02438 | -0.02438 | -0.24018
2019-03-19 16:08 | 1553026068.7812 | 1539 | 0.20954 | 1.7e-05 | 0 | 100.7813 | 100.8047 | 127.6035 | 21.2045 | 0.00186 | 0.001757 | -152.6757 | 0.23798 | 8.4024 | 0.003387 | 0.004901 | 0.004696 | 5.3e-05 | 0.001077 | 0.28351 | 1.3502 | 2.5481 | 101.18 | 20.83 | 17.945 | 66.5149 | 103.2916 | 64.0602 | 0.00437 | 1.047 | 0.02559 | 24.4424 | -0.00208 | 0.14158 | -8.9e-05 | -0.006866 | -0.006866 | -0.85778
2019-03-19 16:08 | 1553026069.7809 | 1540 | 0.2096 | 1.9e-05 | 0 | 100.7858 | 100.8565 | 127.5702 | 21.205 | 0.001836 | 0.001801 | -152.6667 | 0.23813 | 8.4004 | 0.003288 | 0.004649 | 0.004514 | 5.1e-05 | 0.001087 | 0.34235 | 1.3499 | 2.5481 | 101.18 | 20.83 | 17.943 | 65.9387 | 100.891 | 65.1076 | 0.004369 | 1.0437 | 0.025547 | 24.4772 | -0.006268 | 0.14167 | -0.000448 | -0.034602 | -0.034602 | -0.17031
2019-03-19 16:08 | 1553026070.7809 | 1541 | 0.20946 | 1.8e-05 | 0 | 100.8389 | 100.8842 | 127.7396 | 21.2045 | 0.001892 | 0.001801 | -152.6496 | 0.2382 | 8.4031 | 0.003359 | 0.0048 | 0.004606 | 5.1e-05 | 0.001098 | 0.30502 | 1.3177 | 2.5482 | 101.18 | 20.83 | 17.941 | 67.2857 | 102.6159 | 65.1165 | 0.004369 | 1.0437 | 0.025564 | 24.4937 | -0.004015 | 0.14171 | 0.000411 | 0.03168 | 0.03168 | 0.18608
2019-03-19 16:08 | 1553026071.7809 | 1542 | 0.20961 | 2e-05 | 1e-06 | 100.7996 | 100.84 | 128.022 | 21.2052 | 0.001874 | 0.0018 | -152.6464 | 0.23788 | 8.4048 | 0.003365 | 0.005052 | 0.004849 | 5.2e-05 | 0.001101 | 0.31813 | 1.3274 | 2.5481 | 101.18 | 20.84 | 17.938 | 66.8595 | 102.7759 | 65.0827 | 0.004371 | 1.0438 | 0.025593 | 24.5196 | -0.003585 | 0.14151 | -0.000531 | -0.041058 | -0.041058 | -0.14338
2019-03-19 16:08 | 1553026072.7814 | 1543 | 0.20946 | 1.7e-05 | 0 | 100.805 | 101.0866 | 128.0269 | 21.205 | 0.001836 | 0.001744 | -152.6393 | 0.23764 | 8.4028 | 0.003338 | 0.005125 | 0.004959 | 5.5e-05 | 0.001106 | 0.31045 | 1.3236 | 2.5481 | 101.18 | 20.84 | 17.937 | 65.9575 | 102.1049 | 63.7444 | 0.004371 | 1.0479 | 0.025645 | 24.4715 | -0.024952 | 0.14137 | 0.000422 | 0.032612 | 0.032612 | 0.18034
2019-03-19 16:08 | 1553026073.7811 | 1544 | 0.20953 | 2.2e-05 | 0 | 100.853 | 100.9085 | 127.9645 | 21.2054 | 0.001845 | 0.001773 | -152.6477 | 0.23826 | 8.4017 | 0.00339 | 0.004789 | 0.004564 | 5.2e-05 | 0.001112 | 0.30453 | 1.3436 | 2.548 | 101.18 | 20.84 | 17.935 | 66.1755 | 103.3734 | 64.4326 | 0.00437 | 1.0458 | 0.025612 | 24.4905 | -0.00492 | 0.14174 | -6e-05 | -0.004618 | -0.004618 | -1.2769
2019-03-19 16:08 | 1553026074.7811 | 1545 | 0.20964 | 1.9e-05 | 0 | 100.865 | 100.9308 | 127.6134 | 21.204 | 0.001931 | 0.001722 | -152.6639 | 0.23778 | 8.4013 | 0.003363 | 0.003937 | 0.003769 | 5e-05 | 0.001112 | 0.3183 | 1.3613 | 2.5481 | 101.18 | 20.84 | 17.935 | 68.2276 | 102.7183 | 63.2115 | 0.00437 | 1.0496 | 0.025623 | 24.4126 | -0.005828 | 0.14146 | -0.000663 | -0.051314 | -0.051314 | -0.11468
2019-03-19 16:08 | 1553026075.7801 | 1546 | 0.20954 | 2e-05 | 0 | 100.8533 | 100.9945 | 127.4507 | 21.2052 | 0.001825 | 0.001746 | -152.6666 | 0.23758 | 8.406 | 0.003381 | 0.004317 | 0.004086 | 5.2e-05 | 0.001111 | 0.32189 | 1.3567 | 2.5481 | 101.18 | 20.84 | 17.935 | 65.6884 | 103.142 | 63.7961 | 0.00437 | 1.0478 | 0.025585 | 24.4182 | -0.01251 | 0.14134 | -6.6e-05 | -0.005109 | -0.005109 | -1.1508
2019-03-19 16:08 | 1553026076.7812 | 1547 | 0.20947 | 2.5e-05 | 0 | 100.8044 | 100.9681 | 127.4779 | 21.2047 | 0.001875 | 0.001767 | -152.6506 | 0.23821 | 8.4029 | 0.003394 | 0.005035 | 0.004759 | 5.1e-05 | 0.001115 | 0.32621 | 1.3295 | 2.5481 | 101.18 | 20.84 | 17.94 | 66.884 | 103.471 | 64.2913 | 0.004371 | 1.0462 | 0.025569 | 24.4388 | -0.014513 | 0.14171 | 0.000346 | 0.02669 | 0.02669 | 0.22088
2019-03-19 16:08 | 1553026077.7803 | 1548 | 0.20955 | 2.7e-05 | 0 | 100.8509 | 100.8904 | 127.887 | 21.2048 | 0.001879 | 0.00176 | -152.6572 | 0.2377 | 8.4046 | 0.0034 | 0.005062 | 0.004871 | 5.2e-05 | 0.001117 | 0.31277 | 1.3329 | 2.5482 | 101.18 | 20.84 | 17.942 | 66.9755 | 103.6032 | 64.1315 | 0.004372 | 1.0467 | 0.025616 | 24.4722 | -0.003497 | 0.14141 | -0.000141 | -0.010936 | -0.010936 | -0.53792
2019-03-19 16:08 | 1553026078.7803 | 1549 | 0.20962 | 1.6e-05 | 0 | 100.8857 | 101.0948 | 127.7319 | 21.204 | 0.001863 | 0.001785 | -152.6592 | 0.23848 | 8.4007 | 0.003433 | 0.00479 | 0.004658 | 5.3e-05 | 0.001111 | 0.32117 | 1.3215 | 2.548 | 101.18 | 20.84 | 17.945 | 66.5916 | 104.3985 | 64.724 | 0.004373 | 1.0449 | 0.025578 | 24.4788 | -0.018519 | 0.14187 | -0.000542 | -0.041857 | -0.041857 | -0.141
2019-03-19 16:08 | 1553026079.7793 | 1550 | 0.20955 | 2.1e-05 | 0 | 100.9176 | 100.9123 | 127.6744 | 21.2042 | 0.001876 | 0.001777 | -152.6466 | 0.23766 | 8.4047 | 0.003471 | 0.004474 | 0.004205 | 5.3e-05 | 0.001106 | 0.30287 | 1.3549 | 2.5482 | 101.18 | 20.84 | 17.95 | 66.9196 | 105.3308 | 64.5212 | 0.004374 | 1.0455 | 0.02558 | 24.4659 | 0.000467 | 0.14139 | -0.000144 | -0.011106 | -0.011106 | -0.52958
2019-03-19 16:08 | 1553026080.7805 | 1551 | 0.20954 | 2e-05 | 0 | 100.9638 | 100.8746 | 127.7432 | 21.2045 | 0.001927 | 0.001736 | -152.6541 | 0.23701 | 8.4024 | 0.003442 | 0.004728 | 0.00452 | 5.4e-05 | 0.001099 | 0.30397 | 1.3773 | 2.5481 | 101.18 | 20.84 | 17.95 | 68.1343 | 104.6193 | 63.5508 | 0.004374 | 1.0485 | 0.025624 | 24.4373 | 0.007902 | 0.141 | -0.000101 | -0.007848 | -0.007848 | -0.74733
2019-03-19 16:08 | 1553026081.7805 | 1552 | 0.20954 | 2.1e-05 | 0 | 100.9971 | 101.0983 | 128.0152 | 21.2044 | 0.001868 | 0.001753 | -152.6383 | 0.23839 | 8.4032 | 0.003465 | 0.005012 | 0.004787 | 5.3e-05 | 0.001097 | 0.30182 | 1.3289 | 2.5481 | 101.18 | 20.84 | 17.947 | 66.7137 | 105.1972 | 63.9578 | 0.004373 | 1.0473 | 0.025635 | 24.4781 | -0.008956 | 0.14182 | -7.9e-05 | -0.006129 | -0.006129 | -0.96258
2019-03-19 16:08 | 1553026082.7797 | 1553 | 0.20956 | 2.3e-05 | 0 | 101.0019 | 101.0037 | 128.115 | 21.2039 | 0.001839 | 0.0018 | -152.6661 | 0.23811 | 8.4033 | 0.003499 | 0.005129 | 0.004987 | 5.5e-05 | 0.001097 | 0.30862 | 1.3528 | 2.5482 | 101.18 | 20.84 | 17.945 | 66.0179 | 106.0018 | 65.0877 | 0.004373 | 1.0438 | 0.025602 | 24.5287 | -0.000164 | 0.14165 | -0.000192 | -0.01483 | -0.01483 | -0.39735
2019-03-19 16:08 | 1553026083.781 | 1554 | 0.20957 | 2.1e-05 | 0 | 101.0001 | 101.1346 | 128.1438 | 21.2046 | 0.001863 | 0.001786 | -152.6571 | 0.23761 | 8.4031 | 0.003496 | 0.004957 | 0.004766 | 5.3e-05 | 0.001102 | 0.3282 | 1.3923 | 2.548 | 101.18 | 20.84 | 17.942 | 66.5922 | 105.9316 | 64.7566 | 0.004372 | 1.0448 | 0.025618 | 24.5194 | -0.011902 | 0.14136 | -0.000263 | -0.020334 | -0.020334 | -0.28919
2019-03-19 16:08 | 1553026084.7807 | 1555 | 0.20955 | 2.1e-05 | 1e-06 | 101.0616 | 101.2222 | 128.2039 | 21.2044 | 0.001768 | 0.001812 | -152.6548 | 0.23778 | 8.4019 | 0.003616 | 0.004982 | 0.004776 | 5e-05 | 0.001101 | 0.32659 | 1.3386 | 2.548 | 101.18 | 20.84 | 17.933 | 64.3238 | 108.8523 | 65.3654 | 0.00437 | 1.0429 | 0.025601 | 24.5472 | -0.014206 | 0.14146 | -0.000167 | -0.012981 | -0.012981 | -0.45331
2019-03-19 16:08 | 1553026085.7807 | 1556 | 0.20952 | 2e-05 | 0 | 101.0835 | 101.2603 | 128.1995 | 21.2043 | 0.001924 | 0.001774 | -152.6435 | 0.23837 | 8.4001 | 0.003475 | 0.004688 | 0.004433 | 5.5e-05 | 0.0011 | 0.30326 | 1.3912 | 2.5481 | 101.18 | 20.84 | 17.928 | 68.0608 | 105.4198 | 64.4589 | 0.004368 | 1.0457 | 0.025635 | 24.5139 | -0.015626 | 0.1418 | 4.4e-05 | 0.003365 | 0.003365 | 1.7531
2019-03-19 16:08 | 1553026086.7797 | 1557 | 0.20953 | 2.2e-05 | 0 | 101.1162 | 101.2767 | 128.2422 | 21.2037 | 0.001914 | 0.001733 | -152.6534 | 0.23778 | 8.403 | 0.00348 | 0.004655 | 0.004455 | 5.2e-05 | 0.001106 | 0.31996 | 1.3211 | 2.5481 | 101.18 | 20.84 | 17.916 | 67.827 | 105.556 | 63.4588 | 0.004366 | 1.0488 | 0.025677 | 24.4817 | -0.014184 | 0.14146 | -1e-05 | -0.000764 | -0.000764 | -7.7007
2019-03-19 16:08 | 1553026087.7812 | 1558 | 0.20954 | 2.3e-05 | 0 | 101.1402 | 101.1876 | 128.1767 | 21.2043 | 0.001911 | 0.001783 | -152.6352 | 0.23744 | 8.4024 | 0.00352 | 0.005145 | 0.004968 | 5.3e-05 | 0.001107 | 0.30021 | 1.3711 | 2.5483 | 101.18 | 20.84 | 17.909 | 67.7536 | 106.5097 | 64.668 | 0.004364 | 1.0451 | 0.025625 | 24.5193 | -0.004193 | 0.14125 | -8.8e-05 | -0.006788 | -0.006788 | -0.86566
2019-03-19 16:08 | 1553026088.7809 | 1559 | 0.20954 | 2e-05 | 0 | 101.1609 | 101.275 | 128.1711 | 21.2039 | 0.001828 | 0.001743 | -152.6454 | 0.23812 | 8.4048 | 0.003546 | 0.004857 | 0.004767 | 5.3e-05 | 0.001109 | 0.32117 | 1.3821 | 2.5481 | 101.18 | 20.83 | 17.903 | 65.7519 | 107.1424 | 63.7058 | 0.00436 | 1.0481 | 0.025661 | 24.4839 | -0.010078 | 0.14166 | -8.5e-05 | -0.006639 | -0.006639 | -0.88762
2019-03-19 16:08 | 1553026089.7809 | 1560 | 0.20961 | 2e-05 | 0 | 101.2306 | 101.3544 | 128.1593 | 21.2036 | 0.001885 | 0.00176 | -152.6645 | 0.23765 | 8.4039 | 0.003566 | 0.005061 | 0.004834 | 5.3e-05 | 0.001111 | 0.32139 | 1.3614 | 2.5481 | 101.18 | 20.83 | 17.89 | 67.1342 | 107.6311 | 64.1122 | 0.004357 | 1.0468 | 0.025644 | 24.4975 | -0.010939 | 0.14138 | -0.000487 | -0.037702 | -0.037702 | -0.15599
2019-03-19 16:08 | 1553026090.7799 | 1561 | 0.20954 | 2.1e-05 | 0 | 101.2195 | 101.4497 | 128.1844 | 21.2039 | 0.001887 | 0.001757 | -152.6672 | 0.23702 | 8.4042 | 0.003508 | 0.004907 | 0.00466 | 5.2e-05 | 0.001115 | 0.32101 | 1.345 | 2.5481 | 101.18 | 20.84 | 17.884 | 67.1828 | 106.2259 | 64.0478 | 0.004358 | 1.047 | 0.025649 | 24.4976 | -0.020319 | 0.141 | -0.000108 | -0.008325 | -0.008325 | -0.70459
2019-03-19 16:08 | 1553026091.7812 | 1562 | 0.20953 | 2.3e-05 | 0 | 101.2189 | 101.4089 | 127.8162 | 21.2035 | 0.001849 | 0.001766 | -152.6522 | 0.23779 | 8.4046 | 0.003589 | 0.004678 | 0.004454 | 5.1e-05 | 0.001113 | 0.31404 | 1.3305 | 2.5482 | 101.18 | 20.84 | 17.873 | 66.261 | 108.1857 | 64.2589 | 0.004355 | 1.0463 | 0.025604 | 24.47 | -0.016778 | 0.14146 | 0 | 2.5e-05 | 2.5e-05 | 233.97
2019-03-19 16:08 | 1553026092.7802 | 1563 | 0.20949 | 2.2e-05 | 1e-06 | 101.2117 | 101.3836 | 127.5162 | 21.2024 | 0.00184 | 0.001791 | -152.6489 | 0.23866 | 8.4026 | 0.003599 | 0.005123 | 0.004913 | 5.2e-05 | 0.001109 | 0.28943 | 1.3237 | 2.5481 | 101.18 | 20.84 | 17.869 | 66.0408 | 108.4328 | 64.8693 | 0.004354 | 1.0445 | 0.025551 | 24.4634 | -0.015179 | 0.14198 | 0.000239 | 0.01843 | 0.01843 | 0.32047
2019-03-19 16:08 | 1553026093.7805 | 1564 | 0.20959 | 1.7e-05 | 0 | 101.2714 | 101.4606 | 127.4401 | 21.2024 | 0.001855 | 0.001708 | -152.6437 | 0.2374 | 8.4026 | 0.003495 | 0.005079 | 0.004927 | 5e-05 | 0.001111 | 0.27345 | 1.342 | 2.5481 | 101.18 | 20.83 | 17.866 | 66.4114 | 105.9137 | 62.8625 | 0.004351 | 1.0507 | 0.025619 | 24.3834 | -0.016698 | 0.14123 | -0.00036 | -0.027851 | -0.027851 | -0.21095
2019-03-19 16:08 | 1553026094.7801 | 1565 | 0.20948 | 2.6e-05 | 1e-06 | 101.2321 | 101.427 | 127.5256 | 21.2031 | 0.001871 | 0.001752 | -152.6437 | 0.23731 | 8.4021 | 0.003548 | 0.00454 | 0.004364 | 5.5e-05 | 0.001109 | 0.27041 | 1.3254 | 2.548 | 101.18 | 20.83 | 17.861 | 66.782 | 107.2023 | 63.9253 | 0.004349 | 1.0474 | 0.025588 | 24.4301 | -0.017212 | 0.14118 | 0.00027 | 0.020794 | 0.020794 | 0.28244
2019-03-19 16:08 | 1553026095.7794 | 1566 | 0.20959 | 2.3e-05 | 0 | 101.3158 | 101.3298 | 127.5258 | 21.2023 | 0.00188 | 0.001789 | -152.6542 | 0.23741 | 8.4059 | 0.003545 | 0.005127 | 0.004987 | 5.2e-05 | 0.001101 | 0.33859 | 1.3503 | 2.548 | 101.18 | 20.83 | 17.86 | 66.994 | 107.1275 | 64.816 | 0.004349 | 1.0446 | 0.025554 | 24.4624 | -0.001235 | 0.14123 | -0.000398 | -0.030714 | -0.030714 | -0.19129
2019-03-19 16:08 | 1553026096.7794 | 1567 | 0.20953 | 2.4e-05 | 0 | 101.3626 | 101.4901 | 127.3948 | 21.2026 | 0.001884 | 0.00176 | -152.6556 | 0.23744 | 8.4037 | 0.00344 | 0.004871 | 0.004632 | 5.4e-05 | 0.001097 | 0.27168 | 1.3949 | 2.5481 | 101.18 | 20.83 | 17.86 | 67.0981 | 104.5782 | 64.1129 | 0.004349 | 1.0468 | 0.025567 | 24.4243 | -0.011241 | 0.14125 | -2.1e-05 | -0.001606 | -0.001606 | -3.6581
2019-03-19 16:08 | 1553026097.7793 | 1568 | 0.20957 | 2.5e-05 | 0 | 101.374 | 101.7182 | 127.0559 | 21.2028 | 0.001856 | 0.001758 | -152.6656 | 0.23761 | 8.4042 | 0.003458 | 0.00519 | 0.004921 | 5.5e-05 | 0.001101 | 0.32388 | 1.3448 | 2.5482 | 101.18 | 20.83 | 17.862 | 66.4406 | 105.0045 | 64.0655 | 0.00435 | 1.0469 | 0.025535 | 24.3901 | -0.030314 | 0.14136 | -0.000287 | -0.022128 | -0.022128 | -0.26575
2019-03-19 16:08 | 1553026098.7797 | 1569 |  | 2e-05 |  | 101.3608 | 101.4606 | 127.2432 | 21.2028 | 0.001827 | 0.001791 | -152.6678 | 0.23797 | 8.4041 | 0.00345 | 0.00417 | 0.004046 | 5.3e-05 | 0.001088 | 0.28528 | 1.3405 | 2.5481 | 101.18 | 20.83 | 17.863 | 65.7304 | 104.8244 | 64.8766 | 0.00435 | 1.0444 | 0.025523 | 24.4374 | -0.008803 | 0.14157
2019-03-19 16:08 | 1553026099.7797 | 1570 |  | 2.1e-05 |  | 101.3761 | 101.5243 | 127.1564 | 21.2025 | 0.001868 | 0.001757 | -152.6731 | 0.23798 | 8.403 | 0.003388 | 0.005002 | 0.004857 | 5.7e-05 | 0.001085 | 0.31581 | 1.3506 | 2.548 | 101.18 | 20.83 | 17.861 | 66.7212 | 103.3264 | 64.0543 | 0.004349 | 1.047 | 0.025546 | 24.3994 | -0.013067 | 0.14158
2019-03-19 16:08 | 1553026100.7797 | 1571 |  |  |  | 101.3662 | 101.4794 | 126.9721 | 21.2039 | 0.001908 | 0.001763 | -152.6689 | 0.2381 | 8.4025 | 0.003392 | 0.005337 | 0.005083 | 5.4e-05 | 0.001087 | 0.29358 | 1.3364 | 2.548 | 101.18 | 20.83 | 17.86 | 67.6813 | 103.4104 | 64.1837 | 0.004349 | 1.0466 | 0.025522 | 24.3864 | -0.009987 | 0.14164
2019-03-19 16:08 | 1553026101.7797 | 1572 |  |  |  | 101.4169 | 101.5058 | 126.8885 | 21.2027 | 0.00191 | 0.001827 | -152.6857 | 0.23689 | 8.404 | 0.003358 | 0.004685 | 0.004522 | 5.8e-05 | 0.001092 | 0.30596 | 1.3515 | 2.5481 | 101.18 | 20.83 | 17.859 | 67.7287 | 102.6038 | 65.7407 | 0.004349 | 1.0418 | 0.025455 | 24.4345 | -0.007837 | 0.14092
2019-03-19 16:08 | 1553026102.7797 | 1573 |  |  |  | 101.4286 | 101.5682 | 126.8322 | 21.2025 | 0.001912 | 0.001812 | -152.683 | 0.23734 | 8.4051 | 0.003355 | 0.004676 | 0.004447 | 5.7e-05 | 0.001094 | 0.32997 | 1.3692 | 2.5479 | 101.18 | 20.83 | 17.855 | 67.7682 | 102.5212 | 65.3673 | 0.004348 | 1.0429 | 0.025463 | 24.4156 | -0.012296 | 0.14119
2019-03-19 16:08 | 1553026103.7797 | 1574 |  |  |  | 101.488 | 101.4832 |  | 21.2023 | 0.001918 | 0.001759 | -152.6881 | 0.23679 | 8.4016 | 0.003325 | 0.004855 | 0.004712 | 5.7e-05 | 0.001091 | 0.28982 | 1.3206 | 2.548 | 101.18 | 20.83 | 17.852 | 67.9077 | 101.7914 | 64.0938 | 0.004347 | 1.0469 |  |  | 0.000426 | 0.14087
2019-03-19 16:08 | 1553026104.7797 | 1575 |  |  |  | 101.4841 | 101.7134 |  | 21.2029 | 0.001866 |  | -152.6684 | 0.23712 | 8.4016 | 0.003328 | 0.005609 | 0.005417 | 5.8e-05 | 0.001094 | 0.30453 | 1.3456 | 2.548 | 101.18 | 20.83 | 17.849 | 66.6678 | 101.8665 |  | 0.004347 |  |  |  | -0.02018 | 0.14106
2019-03-19 16:08 | 1553026105.7797 | 1576 |  |  |  | 101.5004 | 101.5695 |  | 21.2029 | 0.001842 |  |  | 0.23876 | 8.4012 | 0.003418 | 0.004935 | 0.004754 | 5.7e-05 | 0.001094 | 0.29087 | 1.3455 | 2.548 | 101.18 | 20.83 | 17.848 | 66.1028 | 104.0563 |  | 0.004346 |  |  |  | -0.006093 | 0.14204
2019-03-19 16:08 | 1553026106.7797 | 1577 |  |  |  | 101.4965 | 101.587 |  | 21.2026 |  |  |  | 0.23737 | 8.4002 | 0.003387 | 0.004929 | 0.004713 | 5.7e-05 | 0.001096 | 0.28943 | 1.3179 | 2.5482 | 101.18 | 20.83 | 17.848 |  | 103.2835 |  | 0.004346 |  |  |  | -0.007975 | 0.14121
2019-03-19 16:08 | 1553026107.7797 | 1578 |  |  |  | 101.4754 | 101.6356 |  | 21.203 |  |  |  | 0.23771 | 8.4039 | 0.003444 | 0.004805 | 0.004581 | 5.9e-05 | 0.001095 | 0.24729 | 1.3348 | 2.548 | 101.18 | 20.83 | 17.853 |  | 104.6775 |  | 0.004347 |  |  |  | -0.014108 | 0.14141
2019-03-19 16:08 | 1553026108.7797 | 1579 |  |  |  | 101.5658 | 101.6336 |  | 21.2026 |  |  |  | 0.23725 | 8.4036 | 0.003422 | 0.004815 | 0.004657 | 5.6e-05 | 0.001101 | 0.34291 | 1.3119 | 2.5481 | 101.18 | 20.83 | 17.856 |  | 104.1526 |  | 0.004348 |  |  |  | -0.005968 | 0.14114
2019-03-19 16:08 | 1553026109.7797 | 1580 |  |  |  | 101.5457 | 101.7171 |  | 21.2025 |  |  |  | 0.23753 | 8.4027 | 0.003404 | 0.004929 | 0.004757 | 5.7e-05 | 0.001099 | 0.2995 | 1.3006 | 2.548 | 101.18 | 20.83 | 17.868 |  | 103.7108 |  | 0.004351 |  |  |  | -0.015091 | 0.1413
2019-03-19 16:08 | 1553026110.7797

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-19 15:42 | 0 | 0 | 0 | 0 | 0 | 48.9 | -0.019146
2019-03-19 15:42 | 1 | 0 | 0 | 0 | 0 | 48.9772 | -0.013465
2019-03-19 15:42 | 2 | 0 | 0 | 0 | 0 | 48.8279 | -0.014587
2019-03-19 15:42 | 3 | 0 | 0 | 0 | 0 | 49.0436 | -0.005807
2019-03-19 15:42 | 4 | 0 | 0 | 0 | 0 | 48.984 | 0.007215
2019-03-19 15:42 | 5 | 0 | 0 | 0 | 0 | 48.7444 | 0.028006
2019-03-19 15:42 | 6 | 0 | 0 | 0 | 0 | 48.5842 | 0.23422
2019-03-19 15:42 | 7 | 0 | 0 | 0 | 0 | 48.511 | 0.57021
2019-03-19 15:42 | 8 | 0 | 0 | 0 | 0 | 48.3911 | 1.102
2019-03-19 15:42 | 9 | 0 | 0 | 0 | 0 | 48.2746 | 1.8229
2019-03-19 15:42 | 10 | 0 | 0 | 0 | 0 | 48.1693 | 2.4576
2019-03-19 15:42 | 11 | 0 | 0 | 0 | 0 | 48.0528 | 2.9336
2019-03-19 15:42 | 12 | 0 | 0 | 0 | 0 | 47.9441 | 3.1474
2019-03-19 15:42 | 13 | 0 | 0 | 0 | 0 | 47.8227 | 3.3725
2019-03-19 15:42 | 14 | 0.12312 | 0 | 0 | 0 | 47.6945 | 3.4355
2019-03-19 15:42 | 15 | 0.12329 | 0 | 0 | 0 | 47.5756 | 3.5679
2019-03-19 15:42 | 16 | 0.13664 | 0 | 0 | 0 | 47.4449 | 3.7705
2019-03-19 15:42 | 17 | 0.134 | 0 | 0 | 0 | 47.3059 | 4.0232
2019-03-19 15:42 | 18 | 0.13206 | 0 | 0 | 0 | 47.1767 | 4.1587
2019-03-19 15:42 | 19 | 0.14915 | 0 | 0 | 0 | 47.0379 | 4.2566
2019-03-19 15:42 | 20 | 0.13563 | 0 | 0 | 0 | 46.8851 | 4.4071
2019-03-19 15:42 | 21 | 0.13991 | 0 | 0 | 0 | 46.7625 | 4.4306
2019-03-19 15:42 | 22 | 0.16063 | 0 | 0 | 0 | 46.6014 | 4.4455
2019-03-19 15:42 | 23 | 0.15359 | 0 | 0 | 0 | 46.4477 | 4.4764
2019-03-19 15:43 | 24 | 0.1515 | 0 | 0 | 0 | 46.2939 | 4.5279
2019-03-19 15:43 | 25 | 0.14444 | 0 | 0 | 0 | 46.1458 | 4.5475
2019-03-19 15:43 | 26 | 0.14073 | 0 | 0 | 0 | 46.0042 | 4.5158
2019-03-19 15:43 | 27 | 0.13873 | 0 | 0 | 0 | 45.8633 | 4.4997
2019-03-19 15:43 | 28 | 0.1531 | 0 | 0 | 0 | 45.723 | 4.4434
2019-03-19 15:43 | 29 | 0.15334 | 0 | 0 | 0 | 45.5606 | 4.3884
2019-03-19 15:43 | 30 | 0.14074 | 0 | 0 | 0 | 45.4201 | 4.3135
2019-03-19 15:43 | 31 | 0.12463 | 0 | 0 | 0 | 45.2804 | 4.2309
2019-03-19 15:43 | 32 | 0.11226 | 0 | 0 | 0 | 45.1696 | 4.0506
2019-03-19 15:43 | 33 | 0.11443 | 0 | 0 | 0 | 45.0522 | 3.9311
2019-03-19 15:43 | 34 | 0.10747 | 0 | 0 | 0 | 44.9418 | 3.7887
2019-03-19 15:43 | 35 | 0.12487 | 0.020431 | 2.358 | 29.176 | 44.8314 | 3.6432
2019-03-19 15:43 | 36 | 0.14116 | 0.017843 | 2.0798 | 25.0827 | 44.6926 | 3.5408
2019-03-19 15:43 | 37 | 0.12691 | 0.019554 | 2.2956 | 27.0654 | 44.5569 | 3.435
2019-03-19 15:43 | 38 | 0.10893 | 0.022046 | 2.6438 | 31.3397 | 44.4403 | 3.4138
2019-03-19 15:43 | 39 | 0.10337 | 0.022382 | 2.7425 | 32.5048 | 44.336 | 3.3602
2019-03-19 15:43 | 40 | 0.10252 | 0.021552 | 2.7241 | 32.2149 | 44.232 | 3.3026
2019-03-19 15:43 | 41 | 0.10235 | 0.020543 | 2.6863 | 31.3126 | 44.1305 | 3.2049
2019-03-19 15:43 | 42 | 0.10324 | 0.019413 | 2.632 | 29.8178 | 44.0269 | 3.0783
2019-03-19 15:43 | 43 | 0.10353 | 0.018355 | 2.5851 | 28.1237 | 43.9239 | 2.9116
2019-03-19 15:43 | 44 | 0.11366 | 0.016073 | 2.3288 | 24.7695 | 43.8181 | 2.8154
2019-03-19 15:43 | 45 | 0.10071 | 0.017337 | 2.5831 | 26.3825 | 43.7024 | 2.657
2019-03-19 15:43 | 46 | 0.07711 | 0.021574 | 3.3041 | 32.5009 | 43.6188 | 2.5061
2019-03-19 15:43 | 47 | 0.10041 | 0.015963 | 2.4821 | 24.3888 | 43.5376 | 2.4488
2019-03-19 15:43 | 48 | 0.093097 | 0.016543 | 2.6104 | 24.9731 | 43.4249 | 2.3249
2019-03-19 15:43 | 49 | 0.066116 | 0.022466 | 3.5967 | 33.5539 | 43.3563 | 2.2184
2019-03-19 15:43 | 50 | 0.070383 | 0.02054 | 3.2873 | 30.8082 | 43.2854 | 2.1684
2019-03-19 15:43 | 51 | 0.074861 | 0.018496 | 2.9988 | 27.5205 | 43.2147 | 2.0602
2019-03-19 15:43 | 52 | 0.076696 | 0.017537 | 2.8279 | 24.5645 | 43.1368 | 1.884
2019-03-19 15:43 | 53 | 0.057558 | 0.022277 | 3.5535 | 29.4022 | 43.0661 | 1.6923
2019-03-19 15:43 | 54 | 0.059776 | 0.020824 | 3.308 | 25.7367 | 43.0169 | 1.5384
2019-03-19 15:43 | 55 | 0.063459 | 0.018508 | 2.9168 | 22.3125 | 42.9461 | 1.4159
2019-03-19 15:43 | 56 | 0.050202 | 0.023063 | 3.5439 | 26.3121 | 42.8942 | 1.3209
2019-03-19 15:43 | 57 | 0.034134 | 0.03199 | 4.9988 | 36.376 | 42.8462 | 1.2417
2019-03-19 15:43 | 58 | 0.038671 | 0.027191 | 4.2274 | 30.3293 | 42.8204 | 1.1729
2019-03-19 15:43 | 59 | 0.055529 | 0.01812 | 2.8047 | 20.3432 | 42.7653 | 1.1296
2019-03-19 15:43 | 60 | 0.065196 | 0.015088 | 2.275 | 16.436 | 42.7112 | 1.0716
2019-03-19 15:43 | 61 | 0.062804 | 0.014976 | 2.2478 | 16.6397 | 42.6387 | 1.045
2019-03-19 15:43 | 62 | 0.042856 | 0.02089 | 3.1405 | 23.5153 | 42.5897 | 1.0078
2019-03-19 15:43 | 63 | 0.051507 | 0.016965 | 2.4913 | 18.9801 | 42.5468 | 0.97761
2019-03-19 15:43 | 64 | 0.039938 | 0.021333 | 3.057 | 23.8743 | 42.4905 | 0.95349
2019-03-19 15:43 | 65 | 0.041789 | 0.019819 | 2.7787 | 22.5718 | 42.4636 | 0.94326
2019-03-19 15:43 | 66 | 0.061918 | 0.013024 | 1.7975 | 15.2147 | 42.4028 | 0.94207
2019-03-19 15:43 | 67 | 0.05238 | 0.015379 | 2.0308 | 17.5722 | 42.3468 | 0.92044
2019-03-19 15:43 | 68 | 0.039155 | 0.020021 | 2.6113 | 23.7071 | 42.2992 | 0.92825
2019-03-19 15:43 | 69 | 0.046216 | 0.016975 | 2.1332 | 19.8829 | 42.2635 | 0.9189
2019-03-19 15:43 | 70 | 0.054993 | 0.014257 | 1.7349 | 16.8732 | 42.2064 | 0.9279
2019-03-19 15:43 | 71 | 0.040709 | 0.018684 | 2.2714 | 22.7505 | 42.1589 | 0.92616
2019-03-19 15:43 | 72 | 0.039882 | 0.019616 | 2.2538 | 23.3611 | 42.1224 | 0.93168
2019-03-19 15:43 | 73 | 0.034769 | 0.021898 | 2.5235 | 26.9093 | 42.0796 | 0.93561
2019-03-19 15:43 | 74 | 0.038351 | 0.019813 | 2.2426 | 24.9251 | 42.0508 | 0.9559
2019-03-19 15:43 | 75 | 0.038684 | 0.020219 | 2.1727 | 24.4224 | 42.0036 | 0.94476
2019-03-19 15:43 | 76 | 0.034077 | 0.02301 | 2.4279 | 28.2086 | 41.9742 | 0.96127
2019-03-19 15:43 | 77 | 0.053282 | 0.014727 | 1.5356 | 18.2855 | 41.9305 | 0.97428
2019-03-19 15:43 | 78 | 0.041176 | 0.019571 | 1.9577 | 23.3822 | 41.8751 | 0.96278
2019-03-19 15:43 | 79 | 0.021101 | 0.03829 | 3.7993 | 46.0796 | 41.8504 | 0.97235
2019-03-19 15:43 | 80 | 0.023477 | 0.034477 | 3.3792 | 40.6848 | 41.8272 | 0.95514
2019-03-19 15:43 | 81 | 0.034755 | 0.024 | 2.2734 | 28.139 | 41.8004 | 0.97797
2019-03-19 15:43 | 82 | 0.049191 | 0.016925 | 1.5906 | 19.5596 | 41.7569 | 0.96215
2019-03-19 15:43 | 83 | 0.05307 | 0.016131 | 1.4661 | 18.2404 | 41.7053 | 0.96802
2019-03-19 15:44 | 84 | 0.031912 | 0.026824 | 2.4278 | 30.3051 | 41.6567 | 0.96709
2019-03-19 15:44 | 85 | 0.034742 | 0.024648 | 2.2235 | 27.6722 | 41.636 | 0.9614
2019-03-19 15:44 | 86 | 0.042129 | 0.02033 | 1.8261 | 22.5893 | 41.5858 | 0.95167
2019-03-19 15:44 | 87 | 0.032795 | 0.026164 | 2.3391 | 28.5223 | 41.5548 | 0.93538
2019-03-19 15:44 | 88 | 0.056564 | 0.01557 | 1.3509 | 16.7462 | 41.5128 | 0.94724
2019-03-19 15:44 | 89 | 0.066009 | 0.013006 | 1.1528 | 14.388 | 41.445 | 0.94974
2019-03-19 15:44 | 90 | 0.055182 | 0.015523 | 1.3733 | 17.0564 | 41.3863 | 0.94121
2019-03-19 15:44 | 91 | 0.045565 | 0.018371 | 1.6599 | 20.5991 | 41.3351 | 0.9386
2019-03-19 15:44 | 92 | 0.033019 | 0.024598 | 2.2729 | 27.985 | 41.2951 | 0.92403
2019-03-19 15:44 | 93 | 0.047466 | 0.017142 | 1.5708 | 20.0313 | 41.2624 | 0.9508
2019-03-19 15:44 | 94 | 0.057644 | 0.014084 | 1.2875 | 16.4629 | 41.2012 | 0.94898
2019-03-19 15:44 | 95 | 0.04141 | 0.019046 | 1.7826 | 22.0011 | 41.1535 | 0.91107
2019-03-19 15:44 | 96 | 0.042641 | 0.018488 | 1.7133 | 22.4329 | 41.1153 | 0.95656
2019-03-19 15:44 | 97 | 0.025194 | 0.031257 | 2.8807 | 33.4577 | 41.0718 | 0.84293
2019-03-19 15:44 | 98 | 0.029027 | 0.027113 | 2.4812 | 29.3817 | 41.0594 | 0.85285
2019-03-19 15:44 | 99 | 0.053796 | 0.014636 | 1.3333 | 15.4933 | 41.0086 | 0.83347
2019-03-19 15:44 | 100 | 0.04871 | 0.016191 | 1.4644 | 16.9688 | 40.9586 | 0.82655
2019-03-19 15:44 | 101 | 0.046249 | 0.016538 | 1.527 | 19.4611 | 40.9115 | 0.90005
2019-03-19 15:44 | 102 | 0.043153 | 0.017197 | 1.6282 | 19.6946 | 40.8669 | 0.84989
2019-03-19 15:44 | 103 | 0.018737 | 0.038491 | 3.7228 | 46.0868 | 40.8299 | 0.86355
2019-03-19 15:44 | 104 | 0.018584 | 0.037516 | 3.7239 | 48.2222 | 40.8236 | 0.89616
2019-03-19 15:44 | 105 | 0.030139 | 0.022388 | 2.2722 | 25.9125 | 40.7895 | 0.78098
2019-03-19 15:44 | 106 | 0.023414 | 0.028865 | 2.9046 | 36.2749 | 40.7668 | 0.84936
2019-03-19 15:44 | 107 | 0.040794 | 0.016561 | 1.6566 | 19.9164 | 40.7378 | 0.81248
2019-03-19 15:44 | 108 | 0.034674 | 0.020235 | 1.9333 | 23.1804 | 40.6911 | 0.80375
2019-03-19 15:44 | 109 | 0.0076185 | 0.091989 | 8.6877 | 107.099 | 40.6734 | 0.81593
2019-03-19 15:44 | 110 | 0.015698 | 0.044702 | 4.2054 | 51.4534 | 40.6671 | 0.8077
2019-03-19 15:44 | 111 | 0.04289 | 0.016389 | 1.5238 | 18.4238 | 40.637 | 0.7902
2019-03-19 15:44 | 112 | 0.038044 | 0.018464 | 1.7041 | 21.1582 | 40.5885 | 0.80494
2019-03-19 15:44 | 113 | 0.045769 | 0.015314 | 1.399 | 17.0744 | 40.5591 | 0.78149
2019-03-19 15:44 | 114 | 0.041931 | 0.016178 | 1.521 | 18.7173 | 40.5002 | 0.78483
2019-03-19 15:44 | 115 | 0.018239 | 0.037207 | 3.4618 | 42.4053 | 40.4796 | 0.77345
2019-03-19 15:44 | 116 | 0.025796 | 0.026287 | 2.423 | 30.1691 | 40.4564 | 0.77824
2019-03-19 15:44 | 117 | 0.030267 | 0.021603 | 2.0429 | 25.3739 | 40.4277 | 0.76799
2019-03-19 15:44 | 118 | 0.044285 | 0.014779 | 1.3926 | 17.0311 | 40.3942 | 0.75422
2019-03-19 15:44 | 119 | 0.035338 | 0.017896 | 1.7311 | 21.6137 | 40.3451 | 0.76378
2019-03-19 15:44 | 120 | 0.013468 | 0.046885 | 4.497 | 54.3855 | 40.3271 | 0.73246
2019-03-19 15:44 | 121 | 0.0040654 | 0.15546 | 14.7658 | 180.7677 | 40.3149 | 0.7349
2019-03-19 15:44 | 122 | 0.0097671 | 0.064811 | 6.1183 | 75.8172 | 40.3143 | 0.74051
2019-03-19 15:44 | 123 | 0.028169 | 0.022503 | 2.1034 | 26.0791 | 40.2916 | 0.73462
2019-03-19 15:44 | 124 | 0.047246 | 0.013931 | 1.2462 | 15.4233 | 40.2578 | 0.72869
2019-03-19 15:44 | 125 | 0.05222 | 0.012168 | 1.121 | 13.9528 | 40.2012 | 0.72861
2019-03-19 15:44 | 126 | 0.040033 | 0.015832 | 1.4475 | 18.2625 | 40.1581 | 0.73111
2019-03-19 15:44 | 127 | 0.033209 | 0.019135 | 1.7373 | 21.8733 | 40.1204 | 0.72638
2019-03-19 15:44 | 128 | 0.02021 | 0.031384 | 2.8312 | 35.413 | 40.0919 | 0.71571
2019-03-19 15:44 | 129 | 0.050673 | 0.012059 | 1.121 | 13.9264 | 40.0703 | 0.70569
2019-03-19 15:44 | 130 | 0.042796 | 0.013716 | 1.319 | 16.8161 | 39.9992 | 0.71966
2019-03-19 15:44 | 131 | 0.02731 | 0.02152 | 2.0683 | 25.7009 | 39.9867 | 0.70189
2019-03-19 15:44 | 132 | 0.048654 | 0.01159 | 1.1512 | 14.109 | 39.9367 | 0.68646
2019-03-19 15:44 | 133 | 0.028822 | 0.019566 | 1.9345 | 23.9804 | 39.8986 | 0.69115
2019-03-19 15:44 | 134 | 0.016209 | 0.034753 | 3.4201 | 40.9302 | 39.8781 | 0.66344
2019-03-19 15:44 | 135 | 0.011054 | 0.04879 | 4.9669 | 59.0274 | 39.8636 | 0.65251
2019-03-19 15:44 | 136 | 0.022606 | 0.024944 | 2.4188 | 29.4394 | 39.8517 | 0.66551
2019-03-19 15:44 | 137 | 0.027722 | 0.020345 | 1.9609 | 24.0002 | 39.8199 | 0.66534
2019-03-19 15:44 | 138 | 0.019425 | 0.029142 | 2.7859 | 35.3916 | 39.7998 | 0.68747
2019-03-19 15:44 | 139 | 0.0093423 | 0.060375 | 5.7162 | 69.2921 | 39.7809 | 0.64735
2019-03-19 15:44 | 140 | 0.034261 | 0.016515 | 1.5582 | 19.6177 | 39.7726 | 0.67212
2019-03-19 15:44 | 141 | 0.050641 | 0.010738 | 1.0506 | 13.2401 | 39.7146 | 0.6705
2019-03-19 15:44 | 142 | 0.027686 | 0.020439 | 1.9111 | 23.9646 | 39.6805 | 0.66348
2019-03-19 15:44 | 143 | 0.039559 | 0.013701 | 1.3227 | 16.3444 | 39.652 | 0.64658
2019-03-19 15:45 | 144 | 0.038995 | 0.013891 | 1.3391 | 16.4318 | 39.6037 | 0.64076
2019-03-19 15:45 | 145 | 0.034944 | 0.016175 | 1.4901 | 18.7394 | 39.5749 | 0.65482
2019-03-19 15:45 | 146 | 0.039565 | 0.013679 | 1.3047 | 16.0046 | 39.5319 | 0.63323
2019-03-19 15:45 | 147 | 0.040987 | 0.013233 | 1.2612 | 15.7525 | 39.4967 | 0.64565
2019-03-19 15:45 | 148 | 0.030738 | 0.017635 | 1.6724 | 21.4417 | 39.4532 | 0.65907
2019-03-19 15:45 | 149 | 0.011315 | 0.045768 | 4.5179 | 56.4143 | 39.4371 | 0.63831
2019-03-19 15:45 | 150 | 0.020065 | 0.025757 | 2.5297 | 31.3892 | 39.424 | 0.62983
2019-03-19 15:45 | 151 | 0.013735 | 0.039347 | 3.691 | 44.4865 | 39.3991 | 0.61103
2019-03-19 15:45 | 152 | 0.036005 | 0.015038 | 1.4045 | 17.1701 | 39.3899 | 0.61821
2019-03-19 15:45 | 153 | 0.050199 | 0.010332 | 1.0007 | 12.462 | 39.329 | 0.62558
2019-03-19 15:45 | 154 | 0.042481 | 0.01219 | 1.178 | 14.2465 | 39.2952 | 0.60521
2019-03-19 15:45 | 155 | 0.040425 | 0.012838 | 1.2315 | 15.2354 | 39.2437 | 0.61589
2019-03-19 15:45 | 156 | 0.023883 | 0.021732 | 2.0754 | 25.3068 | 39.2173 | 0.60439
2019-03-19 15:45 | 157 | 0.026639 | 0.019466 | 1.8451 | 22.5939 | 39.1912 | 0.60189
2019-03-19 15:45 | 158 | 0.035366 | 0.01401 | 1.3849 | 16.8195 | 39.1621 | 0.59484
2019-03-19 15:45 | 159 | 0.040249 | 0.012322 | 1.2152 | 14.5529 | 39.1215 | 0.58575
2019-03-19 15:45 | 160 | 0.038324 | 0.012934 | 1.2689 | 15.3368 | 39.0834 | 0.58776
2019-03-19 15:45 | 161 | 0.039419 | 0.012579 | 1.2255 | 14.9913 | 39.0449 | 0.59094
2019-03-19 15:45 | 162 | 0.018309 | 0.027182 | 2.632 | 32.1801 | 39.0099 | 0.5892
2019-03-19 15:45 | 163 | 0.0080083 | 0.062016 | 6.0004 | 74.0919 | 39.0058 | 0.59335
2019-03-19 15:45 | 164 | 0.010124 | 0.049036 | 4.6967 | 58.3911 | 38.9902 | 0.59113
2019-03-19 15:45 | 165 | 0.018892 | 0.026267 | 2.5022 | 31.4134 | 38.9837 | 0.59348
2019-03-19 15:45 | 166 | 0.020034 | 0.024848 | 2.3651 | 29.5036 | 38.9545 | 0.59108
2019-03-19 15:45 | 167 | 0.0040236 | 0.12948 | 11.7549 | 145.2302 | 38.9473 | 0.58435
2019-03-19 15:45 | 168 | 0.016433 | 0.031636 | 2.854 | 35.7907 | 38.9394 | 0.58815
2019-03-19 15:45 | 169 | 0.041078 | 0.012688 | 1.141 | 14.2805 | 38.911 | 0.58662
2019-03-19 15:45 | 170 | 0.049864 | 0.010431 | 0.93055 | 11.8015 | 38.8613 | 0.58846
2019-03-19 15:45 | 171 | 0.045188 | 0.011511 | 1.0286 | 12.9038 | 38.8154 | 0.5831
2019-03-19 15:45 | 172 | 0.033637 | 0.015429 | 1.3787 | 17.0298 | 38.7729 | 0.57282
2019-03-19 15:45 | 173 | 0.027741 | 0.018746 | 1.6644 | 20.6761 | 38.7462 | 0.57358
2019-03-19 15:45 | 174 | 0.046808 | 0.01059 | 0.9795 | 12.0149 | 38.7116 | 0.56239
2019-03-19 15:45 | 175 | 0.045016 | 0.010484 | 1.019 | 12.3371 | 38.6573 | 0.55536
2019-03-19 15:45 | 176 | 0.036603 | 0.01288 | 1.2427 | 15.0549 | 38.6233 | 0.55105
2019-03-19 15:45 | 177 | 0.045251 | 0.010943 | 1.0031 | 12.4031 | 38.5803 | 0.56125
2019-03-19 15:45 | 178 | 0.046848 | 0.010584 | 0.96627 | 11.8101 | 38.5347 | 0.55328
2019-03-19 15:45 | 179 | 0.029519 | 0.016786 | 1.5236 | 18.5628 | 38.4915 | 0.54796
2019-03-19 15:45 | 180 | 0.016682 | 0.028307 | 2.6934 | 33.0488 | 38.4746 | 0.55131
2019-03-19 15:45 | 181 | 0.019676 | 0.025214 | 2.283 | 27.5826 | 38.4538 | 0.5427
2019-03-19 15:45 | 182 | 0.026147 | 0.018982 | 1.7101 | 20.7613 | 38.4339 | 0.54286
2019-03-19 15:45 | 183 | 0.032065 | 0.015484 | 1.3951 | 17.046 | 38.4017 | 0.54658
2019-03-19 15:45 | 184 | 0.031459 | 0.015794 | 1.4159 | 17.3919 | 38.3714 | 0.54714
2019-03-19 15:45 | 185 | 0.029353 | 0.016142 | 1.5144 | 18.8165 | 38.3392 | 0.55233
2019-03-19 15:45 | 186 | 0.03403 | 0.014628 | 1.3125 | 16.0629 | 38.3112 | 0.54662
2019-03-19 15:45 | 187 | 0.035411 | 0.013372 | 1.2503 | 15.2603 | 38.272 | 0.54038
2019-03-19 15:45 | 188 | 0.029142 | 0.015462 | 1.5193 | 18.2689 | 38.2423 | 0.5324
2019-03-19 15:45 | 189 | 0.027072 | 0.016643 | 1.6284 | 19.633 | 38.2133 | 0.5315
2019-03-19 15:45 | 190 | 0.0077913 | 0.057791 | 5.64 | 67.9062 | 38.1915 | 0.52907
2019-03-19 15:45 | 191 | 0.018359 | 0.025791 | 2.3749 | 28.5844 | 38.1904 | 0.52479
2019-03-19 15:45 | 192 | 0.041243 | 0.011502 | 1.0528 | 12.855 | 38.1521 | 0.53018
2019-03-19 15:45 | 193 | 0.018494 | 0.026864 | 2.3459 | 28.5478 | 38.1185 | 0.52797
2019-03-19 15:45 | 194 | 0.010616 | 0.046827 | 4.0472 | 50.5814 | 38.1117 | 0.53695
2019-03-19 15:45 | 195 | 0.035397 | 0.01404 | 1.2019 | 15.5777 | 38.0891 | 0.5514
2019-03-19 15:45 | 196 | 0.047862 | 0.010877 | 0.89273 | 11.6021 | 38.0438 | 0.5553
2019-03-19 15:45 | 197 | 0.03419 | 0.01594 | 1.247 | 16.069 | 38.0002 | 0.5494
2019-03-19 15:45 | 198 | 0.022796 | 0.023905 | 1.8782 | 24.2832 | 37.9749 | 0.55356
2019-03-19 15:45 | 199 | 0.038894 | 0.014001 | 1.0982 | 13.8335 | 37.9478 | 0.53804
2019-03-19 15:45 | 200 | 0.050331 | 0.011301 | 0.85243 | 10.6903 | 37.8983 | 0.53805
2019-03-19 15:45 | 201 | 0.031423 | 0.017342 | 1.365 | 16.9617 | 37.8543 | 0.53299
2019-03-19 15:45 | 202 | 0.031581 | 0.017273 | 1.3583 | 16.9067 | 37.8312 | 0.53394
2019-03-19 15:46 | 203 | 0.044806 | 0.012163 | 0.95235 | 12.2119 | 37.7879 | 0.54717
2019-03-19 15:46 | 204 | 0.03958 | 0.013199 | 1.0899 | 13.618 | 37.7457 | 0.539
2019-03-19 15:46 | 205 | 0.042516 | 0.012863 | 1.0116 | 12.6774 | 37.7078 | 0.539
2019-03-19 15:46 | 206 | 0.022434 | 0.024373 | 1.925 | 23.8327 | 37.6657 | 0.53465
2019-03-19 15:46 | 207 | 0.022499 | 0.024248 | 1.9151 | 24.0012 | 37.658 | 0.54
2019-03-19 15:46 | 208 | 0.039692 | 0.013743 | 1.0864 | 13.5569 | 37.6167 | 0.53811
2019-03-19 15:46 | 209 | 0.024336 | 0.02343 | 1.7719 | 22.2385 | 37.5857 | 0.54119
2019-03-19 15:46 | 210 | 0.030893 | 0.018464 | 1.3927 | 17.5955 | 37.5634 | 0.54357
2019-03-19 15:46 | 211 | 0.036054 | 0.015827 | 1.1948 | 14.8648 | 37.5243 | 0.53593
2019-03-19 15:46 | 212 | 0.021881 | 0.027133 | 1.9596 | 24.4936 | 37.4956 | 0.53594
2019-03-19 15:46 | 213 | 0.030383 | 0.019545 | 1.414 | 17.366 | 37.4759 | 0.52763
2019-03-19 15:46 | 214 | 0.020619 | 0.028767 | 2.0812 | 24.9856 | 37.4388 | 0.51519
2019-03-19 15:46 | 215 | 0.0088588 | 0.069685 | 4.8393 | 57.8997 | 37.4345 | 0.51292
2019-03-19 15:46 | 216 | 0.033555 | 0.018385 | 1.2756 | 15.292 | 37.4123 | 0.51313
2019-03-19 15:46 | 217 | 0.046835 | 0.012673 | 0.90808 | 10.8851 | 37.3699 | 0.5098
2019-03-19 15:46 | 218 | 0.036757 | 0.016162 | 1.1571 | 13.8719 | 37.3249 | 0.50988
2019-03-19 15:46 | 219 | 0.024112 | 0.025612 | 1.7523 | 20.9916 | 37.2973 | 0.50616
2019-03-19 15:46 | 220 | 0.029286 | 0.020274 | 1.4375 | 17.2593 | 37.2724 | 0.50545
2019-03-19 15:46 | 221 | 0.030989 | 0.019121 | 1.3545 | 16.3834 | 37.2391 | 0.5077
2019-03-19 15:46 | 222 | 0.031982 | 0.018519 | 1.3026 | 15.6812 | 37.211 | 0.50152
2019-03-19 15:46 | 223 | 0.021015 | 0.028232 | 1.9804 | 23.6451 | 37.1777 | 0.49691
2019-03-19 15:46 | 224 | 0.028813 | 0.021443 | 1.4398 | 17.1628 | 37.1649 | 0.49451
2019-03-19 15:46 | 225 | 0.024515 | 0.025209 | 1.6821 | 20.342 | 37.1227 | 0.49868
2019-03-19 15:46 | 226 | 0.016719 | 0.038376 | 2.4748 | 29.9585 | 37.1165 | 0.50086
2019-03-19 15:46 | 227 | 0.028728 | 0.022333 | 1.4233 | 17.4069 | 37.0847 | 0.50007
2019-03-19 15:46 | 228 | 0.026311 | 0.024383 | 1.5568 | 18.9203 | 37.0617 | 0.49782
2019-03-19 15:46 | 229 | 0.038691 | 0.016571 | 1.0513 | 12.959 | 37.0295 | 0.50139
2019-03-19 15:46 | 230 | 0.023367 | 0.028484 | 1.7457 | 20.7093 | 36.9905 | 0.48391
2019-03-19 15:46 | 231 | 0.024959 | 0.024815 | 1.6363 | 19.3924 | 36.9788 | 0.48401
2019-03-19 15:46 | 232 | 0.045511 | 0.013623 | 0.89994 | 10.5588 | 36.9361 | 0.48054
2019-03-19 15:46 | 233 | 0.037125 | 0.01672 | 1.0944 | 12.8966 | 36.8942 | 0.47879
2019-03-19 15:46 | 234 | 0.039654 | 0.015656 | 1.0229 | 12.0656 | 36.8598 | 0.47844
2019-03-19 15:46 | 235 | 0.049395 | 0.013068 | 0.81687 | 9.7349 | 36.8135 | 0.48086
2019-03-19 15:46 | 236 | 0.036552 | 0.016966 | 1.0972 | 13.1256 | 36.7669 | 0.47976
2019-03-19 15:46 | 237 | 0.0089003 | 0.069708 | 4.4744 | 53.5171 | 36.7437 | 0.47632
2019-03-19 15:46 | 238 | 0.019576 | 0.032914 | 2.0306 | 24.557 | 36.7396 | 0.48073
2019-03-19 15:46 | 239 | 0.03694 | 0.016777 | 1.0678 | 13.0306 | 36.7024 | 0.48134
2019-03-19 15:46 | 240 | 0.027543 | 0.023305 | 1.4269 | 17.5446 | 36.6718 | 0.48323
2019-03-19 15:46 | 241 | 0.035223 | 0.018273 | 1.1156 | 13.8947 | 36.6443 | 0.48941
2019-03-19 15:46 | 242 | 0.025162 | 0.025518 | 1.5639 | 19.4218 | 36.6051 | 0.48869
2019-03-19 15:46 | 243 | 0.027398 | 0.023423 | 1.4301 | 17.9918 | 36.5913 | 0.49294
2019-03-19 15:46 | 244 | 0.031155 | 0.020635 | 1.2622 | 15.8178 | 36.5501 | 0.4928
2019-03-19 15:46 | 245 | 0.015177 | 0.043894 | 2.5817 | 32.2201 | 36.533 | 0.48902
2019-03-19 15:46 | 246 | 0.030257 | 0.02119 | 1.2925 | 16.0595 | 36.5122 | 0.48591
2019-03-19 15:46 | 247 | 0.052476 | 0.012697 | 0.74467 | 9.2275 | 36.4703 | 0.48423
2019-03-19 15:46 | 248 | 0.059211 | 0.011236 | 0.66213 | 8.2354 | 36.4117 | 0.48762
2019-03-19 15:46 | 249 | 0.035906 | 0.017856 | 1.085 | 13.6516 | 36.3592 | 0.49018
2019-03-19 15:46 | 250 | 0.040899 | 0.015718 | 0.95687 | 11.8786 | 36.3339 | 0.48583
2019-03-19 15:46 | 251 | 0.03801 | 0.016883 | 1.0258 | 12.8286 | 36.279 | 0.48762
2019-03-19 15:46 | 252 | 0.024261 | 0.02638 | 1.5968 | 20.3715 | 36.2603 | 0.49423
2019-03-19 15:46 | 253 | 0.02677 | 0.023961 | 1.4461 | 18.2862 | 36.2265 | 0.48953
2019-03-19 15:46 | 254 | 0.031418 | 0.021195 | 1.2359 | 15.9625 | 36.2057 | 0.50152
2019-03-19 15:46 | 255 | 0.040373 | 0.01649 | 0.95601 | 12.4351 | 36.163 | 0.50205
2019-03-19 15:46 | 256 | 0.031575 | 0.021072 | 1.224 | 16.0017 | 36.1294 | 0.50525
2019-03-19 15:46 | 257 | 0.022183 | 0.029972 | 1.746 | 22.0417 | 36.1003 | 0.48895
2019-03-19 15:46 | 258 | 0.020054 | 0.032109 | 1.9305 | 24.2575 | 36.0826 | 0.48646
2019-03-19 15:46 | 259 | 0.039169 | 0.017031 | 0.98925 | 12.3437 | 36.0549 | 0.48349
2019-03-19 15:46 | 260 | 0.046969 | 0.014177 | 0.8219 | 10.4551 | 36.0072 | 0.49106
2019-03-19 15:46 | 261 | 0.029387 | 0.021852 | 1.3214 | 16.7561 | 35.9674 | 0.4924
2019-03-19 15:46 | 262 | 0.022998 | 0.027951 | 1.6707 | 21.7036 | 35.9457 | 0.49914
2019-03-19 15:47 | 263 | 0.039242 | 0.016376 | 0.97982 | 12.7523 | 35.916 | 0.50042
2019-03-19 15:47 | 264 | 0.034009 | 0.018885 | 1.1255 | 14.8587 | 35.8719 | 0.50533
2019-03-19 15:47 | 265 | 0.03139 | 0.019679 | 1.2239 | 16.2028 | 35.8476 | 0.5086
2019-03-19 15:47 | 266 | 0.037342 | 0.016534 | 1.0253 | 13.538 | 35.8071 | 0.50554
2019-03-19 15:47 | 267 | 0.040962 | 0.015079 | 0.93415 | 12.3774 | 35.7737 | 0.507
2019-03-19 15:47 | 268 | 0.03027 | 0.019598 | 1.2626 | 16.6172 | 35.7291 | 0.50301
2019-03-19 15:47 | 269 | 0.011441 | 0.05194 | 3.3725 | 43.7158 | 35.7152 | 0.50017
2019-03-19 15:47 | 270 | 0.0089362 | 0.066501 | 4.3138 | 55.8862 | 35.7022 | 0.49941
2019-03-19 15:47 | 271 | 0.0097844 | 0.060612 | 3.9242 | 51.113 | 35.6958 | 0.50011
2019-03-19 15:47 | 272 | 0.016937 | 0.033629 | 2.2746 | 28.6033 | 35.6805 | 0.48447
2019-03-19 15:47 | 273 | 0.035067 | 0.016927 | 1.1012 | 13.6338 | 35.6595 | 0.47809
2019-03-19 15:47 | 274 | 0.031626 | 0.018789 | 1.2188 | 15.367 | 35.6156 | 0.486
2019-03-19 15:47 | 275 | 0.025865 | 0.022942 | 1.4896 | 18.5 | 35.5972 | 0.47851
2019-03-19 15:47 | 276 | 0.03215 | 0.018456 | 1.1972 | 14.9708 | 35.5614 | 0.48131
2019-03-19 15:47 | 277 | 0.022743 | 0.027196 | 1.693 | 20.7354 | 35.5361 | 0.47159
2019-03-19 15:47 | 278 | 0.027568 | 0.022441 | 1.397 | 17.0915 | 35.5129 | 0.47118
2019-03-19 15:47 | 279 | 0.026801 | 0.023931 | 1.4332 | 17.7646 | 35.482 | 0.4761
2019-03-19 15:47 | 280 | 0.023518 | 0.026291 | 1.6288 | 20.035 | 35.4593 | 0.47119
2019-03-19 15:47 | 281 | 0.044607 | 0.01385 | 0.86137 | 10.5328 | 35.4295 | 0.46984
2019-03-19 15:47 | 282 | 0.049232 | 0.012569 | 0.77788 | 9.6417 | 35.3746 | 0.47468
2019-03-19 15:47 | 283 | 0.01826 | 0.03393 | 2.0917 | 26.2332 | 35.3397 | 0.47902
2019-03-19 15:47 | 284 | 0.014695 | 0.042146 | 2.5881 | 32.2642 | 35.332 | 0.47412
2019-03-19 15:47 | 285 | 0.017832 | 0.033375 | 2.1272 | 26.5256 | 35.3079 | 0.473
2019-03-19 15:47 | 286 | 0.019573 | 0.030362 | 1.9371 | 24.4317 | 35.2963 | 0.47819
2019-03-19 15:47 | 287 | 0.028703 | 0.021596 | 1.3196 | 16.4703 | 35.2676 | 0.47275
2019-03-19 15:47 | 288 | 0.012524 | 0.04935 | 3.0125 | 37.5312 | 35.2444 | 0.47006
2019-03-19 15:47 | 289 | 0.016258 | 0.038055 | 2.323 | 28.9024 | 35.2377 | 0.46991
2019-03-19 15:47 | 290 | 0.0389 | 0.015927 | 0.97033 | 12.0343 | 35.2075 | 0.46813
2019-03-19 15:47 | 291 | 0.029768 | 0.021564 | 1.2613 | 15.579 | 35.1672 | 0.46375
2019-03-19 15:47 | 292 | 0.025496 | 0.024219 | 1.4667 | 18.3842 | 35.1473 | 0.46872
2019-03-19 15:47 | 293 | 0.034875 | 0.017702 | 1.072 | 13.3855 | 35.113 | 0.46682
2019-03-19 15:47 | 294 | 0.033087 | 0.019408 | 1.1205 | 14.1321 | 35.0797 | 0.46759
2019-03-19 15:47 | 295 | 0.042236 | 0.01462 | 0.87834 | 11.1114 | 35.0449 | 0.46931
2019-03-19 15:47 | 296 | 0.033696 | 0.018313 | 1.0991 | 13.972 | 34.9995 | 0.4708
2019-03-19 15:47 | 297 | 0.019861 | 0.031029 | 1.8604 | 23.7768 | 34.9785 | 0.47222
2019-03-19 15:47 | 298 | 0.032522 | 0.018234 | 1.1289 | 14.1749 | 34.9535 | 0.461
2019-03-19 15:47 | 299 | 0.039312 | 0.015697 | 0.93829 | 11.6545 | 34.9149 | 0.45816
2019-03-19 15:47 | 300 | 0.022901 | 0.026957 | 1.6097 | 20.1974 | 34.8804 | 0.46254
2019-03-19 15:47 | 301 | 0.021193 | 0.028004 | 1.7339 | 21.3355 | 34.8656 | 0.45216
2019-03-19 15:47 | 302 | 0.034977 | 0.017638 | 1.0426 | 12.9835 | 34.8342 | 0.45412
2019-03-19 15:47 | 303 | 0.031222 | 0.01979 | 1.171 | 14.7546 | 34.7998 | 0.46066
2019-03-19 15:47 | 304 | 0.022743 | 0.027089 | 1.6045 | 19.5878 | 34.7735 | 0.44549
2019-03-19 15:47 | 305 | 0.011814 | 0.052216 | 3.0738 | 38.4981 | 34.7549 | 0.4548
2019-03-19 15:47 | 306 | 0.0039741 | 0.15554 | 9.194 | 113.7012 | 34.7485 | 0.45186
2019-03-19 15:47 | 307 | 0.019913 | 0.030995 | 1.819 | 22.367 | 34.741 | 0.44539
2019-03-19 15:47 | 308 | 0.032496 | 0.018965 | 1.1119 | 14.0209 | 34.7094 | 0.45563
2019-03-19 15:47 | 309 | 0.023887 | 0.026901 | 1.5187 | 18.9234 | 34.6817 | 0.45203
2019-03-19 15:47 | 310 | 0.0060037 | 0.10305 | 6.0292 | 75.3241 | 34.6643 | 0.45223
2019-03-19 15:47 | 311 | 0.0011018 | 0.58306 | 32.8825 | 411.2273 | 34.6665 | 0.4531
2019-03-19 15:47 | 312 | 0.002361 | 0.27194 | 15.1938 | 193.1571 | 34.6599 | 0.45604
2019-03-19 15:47 | 313 | 0.0097258 | 0.065953 | 3.7042 | 46.4984 | 34.6597 | 0.45223
2019-03-19 15:47 | 314 | 0.027412 | 0.024237 | 1.3072 | 16.3837 | 34.6379 | 0.44912
2019-03-19 15:47 | 315 | 0.036282 | 0.018353 | 0.98781 | 12.3016 | 34.6072 | 0.44632
2019-03-19 15:47 | 316 | 0.033385 | 0.019919 | 1.0676 | 13.1963 | 34.569 | 0.44055
2019-03-19 15:47 | 317 | 0.0166 | 0.041514 | 2.1573 | 26.5776 | 34.544 | 0.44119
2019-03-19 15:47 | 318 | 0.01234 | 0.054018 | 2.9013 | 36.1592 | 34.5326 | 0.44622
2019-03-19 15:47 | 319 | 0.020197 | 0.033007 | 1.7598 | 22.2104 | 34.5155 | 0.4486
2019-03-19 15:47 | 320 | 0.038038 | 0.017493 | 0.93562 | 11.6221 | 34.4895 | 0.44208
2019-03-19 15:47 | 321 | 0.049811 | 0.013357 | 0.70988 | 8.9249 | 34.4415 | 0.44455
2019-03-19 15:47 | 322 | 0.036491 | 0.018248 | 0.96781 | 11.8133 | 34.3968 | 0.43108
2019-03-19 15:48 | 323 | 0.014344 | 0.048202 | 2.4578 | 29.942 | 34.3711 | 0.42948
2019-03-19 15:48 | 324 | 0.0061478 | 0.10851 | 5.7368 | 69.4666 | 34.3644 | 0.42707
2019-03-19 15:48 | 325 | 0.012424 | 0.05567 | 2.8296 | 34.5819 | 34.3545 | 0.42966
2019-03-19 15:48 | 326 | 0.025681 | 0.026944 | 1.3651 | 17.043 | 34.3375 | 0.43768
2019-03-19 15:48 | 327 | 0.031895 | 0.022411 | 1.0941 | 13.3449 | 34.3048 | 0.42563
2019-03-19 15:48 | 328 | 0.035583 | 0.020082 | 0.97619 | 12.2644 | 34.2749 | 0.43641
2019-03-19 15:48 | 329 | 0.026479 | 0.027027 | 1.3053 | 16.3932 | 34.2371 | 0.43408
2019-03-19 15:48 | 330 | 0.010905 | 0.065513 | 3.1571 | 38.9171 | 34.2236 | 0.42439
2019-03-19 15:48 | 331 | 0.0083173 | 0.088579 | 4.122 | 52.7825 | 34.212 | 0.439
2019-03-19 15:48 | 332 | 0.01139 | 0.062676 | 3.0171 | 37.2462 | 34.2048 | 0.42423
2019-03-19 15:48 | 333 | 0.027104 | 0.026353 | 1.2645 | 15.4933 | 34.1861 | 0.41993
2019-03-19 15:48 | 334 | 0.034049 | 0.020979 | 1.0011 | 12.331 | 34.1527 | 0.41986
2019-03-19 15:48 | 335 | 0.032136 | 0.022193 | 1.052 | 12.8057 | 34.1206 | 0.41152
2019-03-19 15:48 | 336 | 0.024558 | 0.029121 | 1.3805 | 17.0489 | 34.09 | 0.41868
2019-03-19 15:48 | 337 | 0.021624 | 0.032975 | 1.5511 | 19.4469 | 34.0703 | 0.42052
2019-03-19 15:48 | 338 | 0.028537 | 0.025021 | 1.1769 | 14.737 | 34.0444 | 0.42055
2019-03-19 15:48 | 339 | 0.023192 | 0.03183 | 1.4409 | 18.5819 | 34.0153 | 0.43095
2019-03-19 15:48 | 340 | 0.041032 | 0.018024 | 0.81314 | 10.2719 | 33.993 | 0.42148
2019-03-19 15:48 | 341 | 0.041187 | 0.017952 | 0.80716 | 9.5968 | 33.9372 | 0.39526
2019-03-19 15:48 | 342 | 0.040074 | 0.019013 | 0.82631 | 9.948 | 33.9115 | 0.39865
2019-03-19 15:48 | 343 | 0.039234 | 0.019422 | 0.84311 | 10.055 | 33.8576 | 0.39449
2019-03-19 15:48 | 344 | 0.019089 | 0.038642 | 1.7238 | 21.1311 | 33.8372 | 0.40337
2019-03-19 15:48 | 345 | 0.023625 | 0.031122 | 1.3867 | 17.0786 | 33.8138 | 0.40348
2019-03-19 15:48 | 346 | 0.025176 | 0.030209 | 1.2995 | 15.565 | 33.7902 | 0.39186
2019-03-19 15:48 | 347 | 0.022023 | 0.034515 | 1.4899 | 18.933 | 33.7648 | 0.41696
2019-03-19 15:48 | 348 | 0.010657 | 0.073618 | 3.0499 | 39.7059 | 33.7474 | 0.42315
2019-03-19 15:48 | 349 | 0.033298 | 0.023566 | 0.9707 | 12.6045 | 33.7358 | 0.41971
2019-03-19 15:48 | 350 | 0.032359 | 0.024217 | 1.0033 | 12.9704 | 33.6862 | 0.41971
2019-03-19 15:48 | 351 | 0.016811 | 0.048082 | 1.9334 | 24.5319 | 33.6749 | 0.4124
2019-03-19 15:48 | 352 | 0.019397 | 0.041654 | 1.6634 | 21.0105 | 33.6486 | 0.40754
2019-03-19 15:48 | 353 | 0.018994 | 0.04393 | 1.7041 | 21.8705 | 33.6363 | 0.41542
2019-03-19 15:48 | 354 | 0.022945 | 0.036377 | 1.4177 | 17.0926 | 33.6096 | 0.39219
2019-03-19 15:48 | 355 | 0.022913 | 0.036439 | 1.4053 | 17.9664 | 33.5908 | 0.41166
2019-03-19 15:48 | 356 | 0.037809 | 0.022711 | 0.85454 | 10.6562 | 33.5605 | 0.4029
2019-03-19 15:48 | 357 | 0.038462 | 0.02234 | 0.83767 | 10.5147 | 33.5187 | 0.40442
2019-03-19 15:48 | 358 | 0.032209 | 0.026694 | 1.0032 | 12.6042 | 33.4859 | 0.40597
2019-03-19 15:48 | 359 | 0.019106 | 0.045039 | 1.6927 | 20.5121 | 33.456 | 0.39191
2019-03-19 15:48 | 360 | 0.011496 | 0.07483 | 2.7895 | 31.9121 | 33.446 | 0.36685
2019-03-19 15:48 | 361 | 0.019436 | 0.045413 | 1.6491 | 19.7555 | 33.4287 | 0.38397
2019-03-19 15:48 | 362 | 0.03558 | 0.024823 | 0.90246 | 10.9172 | 33.4051 | 0.38844
2019-03-19 15:48 | 363 | 0.034487 | 0.026315 | 0.93041 | 12.1192 | 33.3616 | 0.41795
2019-03-19 15:48 | 364 | 0.037026 | 0.024512 | 0.85856 | 10.4179 | 33.3362 | 0.38573
2019-03-19 15:48 | 365 | 0.027244 | 0.033279 | 1.1684 | 13.8026 | 33.2913 | 0.37604
2019-03-19 15:48 | 366 | -0.0076316 | -0.11887 | -4.1587 | -53.0654 | 33.287 | 0.40497
2019-03-19 15:48 | 367 | 0.0056355 | 0.16081 | 5.6193 | 69.0915 | 33.2948 | 0.38937
2019-03-19 15:48 | 368 | 0.030112 | 0.030893 | 1.0469 | 13.3025 | 33.2723 | 0.40056
2019-03-19 15:48 | 369 | 0.023478 | 0.039631 | 1.3445 | 17.0365 | 33.2421 | 0.39998
2019-03-19 15:48 | 370 | 0.012388 | 0.077099 | 2.5473 | 32.0108 | 33.2268 | 0.39655
2019-03-19 15:48 | 371 | 0.02654 | 0.03605 | 1.1953 | 14.7673 | 33.2119 | 0.39193
2019-03-19 15:48 | 372 | 0.013997 | 0.068092 | 2.2417 | 28.4668 | 33.1793 | 0.39845
2019-03-19 15:48 | 373 | 0.0146 | 0.06683 | 2.1439 | 27.6159 | 33.1812 | 0.4032
2019-03-19 15:48 | 374 | 0.018269 | 0.053323 | 1.7106 | 20.9405 | 33.1494 | 0.38256
2019-03-19 15:48 | 375 | 0.0075517 | 0.13247 | 4.1287 | 37.2704 | 33.1474 | 0.28145
2019-03-19 15:48 | 376 | 0.024528 | 0.041835 | 1.2659 | 18.1701 | 33.1283 | 0.44567
2019-03-19 15:48 | 377 | 0.022426 | 0.045775 | 1.3834 | 17.5564 | 33.1027 | 0.39372
2019-03-19 15:48 | 378 | 0.0090087 | 0.11392 | 3.4389 | 44.0571 | 33.086 | 0.3969
2019-03-19 15:48 | 379 | 0.021584 | 0.048661 | 1.4268 | 16.8008 | 33.0784 | 0.36262
2019-03-19 15:48 | 380 | 0.038218 | 0.027526 | 0.80809 | 9.9484 | 33.0421 | 0.3802
2019-03-19 15:48 | 381 | 0.020199 | 0.052155 | 1.5199 | 18.1664 | 33.0104 | 0.36694
2019-03-19 15:48 | 382 | 0.0086527 | 0.12439 | 3.5662 | 46.0069 | 33.0006 | 0.39808
2019-03-19 15:48 | 383 | 0.011593 | 0.090827 | 2.6405 | 29.489 | 32.9887 | 0.34187
2019-03-19 15:49 | 384 | 0.034464 | 0.031229 | 0.88007 | 11.0776 | 32.972 | 0.38178
2019-03-19 15:49 | 385 | 0.051782 | 0.021261 | 0.58033 | 6.7002 | 32.9212 | 0.34695
2019-03-19 15:49 | 386 | 0.048562 | 0.022631 | 0.61694 | 7.3779 | 32.8744 | 0.35829
2019-03-19 15:49 | 387 | 0.029817 | 0.037592 | 1.003 | 12.248 | 32.8289 | 0.3652
2019-03-19 15:49 | 388 | 0.0017123 | 0.66945 | 17.358 | 205.2755 | 32.8167 | 0.35148
2019-03-19 15:49 | 389 | 0.0025332 | 0.45194 | 11.7329 | 139.1369 | 32.8179 | 0.35246
2019-03-19 15:49 | 390 | 0.020415 | 0.056172 | 1.4528 | 17.2661 | 32.8068 | 0.35249
2019-03-19 15:49 | 391 | 0.027381 | 0.041769 | 1.0735 | 13.3262 | 32.7794 | 0.36488
2019-03-19 15:49 | 392 | 0.028463 | 0.041188 | 1.0316 | 12.3956 | 32.754 | 0.35282
2019-03-19 15:49 | 393 | 0.027109 | 0.043204 | 1.0807 | 12.7475 | 32.7242 | 0.34557
2019-03-19 15:49 | 394 | -0.0025891 | -0.46075 | -11.3384 | -140.2689 | 32.7062 | 0.36317
2019-03-19 15:49 | 395 | -0.0077497 | -0.15402 | -3.7805 | -47.8939 | 32.7229 | 0.37116
2019-03-19 15:49 | 396 | 0.021609 | 0.056457 | 1.3517 | 15.9789 | 32.7129 | 0.34529
2019-03-19 15:49 | 397 | 0.027706 | 0.04399 | 1.0518 | 13.5608 | 32.6847 | 0.37572
2019-03-19 15:49 | 398 | 0.027823 | 0.04474 | 1.0492 | 12.7352 | 32.6597 | 0.35433
2019-03-19 15:49 | 399 | 0.032196 | 0.038672 | 0.90219 | 10.5693 | 32.6291 | 0.34029
2019-03-19 15:49 | 400 | 0.0073232 | 0.17347 | 3.95 | 48.0628 | 32.6023 | 0.35197
2019-03-19 15:49 | 401 | -0.0091601 | -0.13844 | -3.1565 | -37.503 | 32.6121 | 0.34353
2019-03-19 15:49 | 402 | 0.012005 | 0.10558 | 2.4041 | 29.8445 | 32.6114 | 0.35828
2019-03-19 15:49 | 403 | 0.01241 | 0.10217 | 2.3204 | 27.9241 | 32.5919 | 0.34654
2019-03-19 15:49 | 404 | 0.024151 | 0.053365 | 1.1802 | 14.2172 | 32.5846 | 0.34336
2019-03-19 15:49 | 405 | 0.026192 | 0.049177 | 1.0862 | 13.2144 | 32.5467 | 0.34612
2019-03-19 15:49 | 406 | -0.0038593 | -0.33506 | -7.3713 | -88.6886 | 32.5395 | 0.34228
2019-03-19 15:49 | 407 | 0.0043042 | 0.3003 | 6.5891 | 81.6323 | 32.5452 | 0.35136
2019-03-19 15:49 | 408 | 0.027763 | 0.047417 | 1.0188 | 12.7075 | 32.5268 | 0.35279
2019-03-19 15:49 | 409 | 0.019673 | 0.06563 | 1.4343 | 17.8237 | 32.4968 | 0.35065
2019-03-19 15:49 | 410 | 0.022675 | 0.056916 | 1.2455 | 15.2978 | 32.4853 | 0.34687
2019-03-19 15:49 | 411 | 0.035999 | 0.036488 | 0.7798 | 9.79 | 32.4493 | 0.35243
2019-03-19 15:49 | 412 | 0.023935 | 0.054939 | 1.1631 | 14.3354 | 32.419 | 0.34312
2019-03-19 15:49 | 413 | 0.024171 | 0.054365 | 1.1494 | 13.8101 | 32.399 | 0.33381
2019-03-19 15:49 | 414 | 0.026922 | 0.048965 | 1.0381 | 12.8504 | 32.37 | 0.34596
2019-03-19 15:49 | 415 | 0.02198 | 0.059893 | 1.2612 | 14.6745 | 32.3469 | 0.32255
2019-03-19 15:49 | 416 | 0.022336 | 0.060188 | 1.2414 | 15.3602 | 32.3247 | 0.34309
2019-03-19 15:49 | 417 | 0.028687 | 0.047615 | 0.95959 | 11.1945 | 32.3008 | 0.32114
2019-03-19 15:49 | 418 | 0.027565 | 0.049513 | 0.99646 | 11.6803 | 32.2692 | 0.32197
2019-03-19 15:49 | 419 | 0.021344 | 0.063936 | 1.2832 | 15.5038 | 32.2472 | 0.33092
2019-03-19 15:49 | 420 | 0.012351 | 0.11067 | 2.2151 | 26.258 | 32.2269 | 0.32432
2019-03-19 15:49 | 421 | 0.019558 | 0.071117 | 1.3969 | 16.8466 | 32.2185 | 0.32948
2019-03-19 15:49 | 422 | 0.027124 | 0.051346 | 1.0016 | 12.5904 | 32.1878 | 0.3415
2019-03-19 15:49 | 423 | 0.016687 | 0.083448 | 1.6256 | 20.5525 | 32.1685 | 0.34295
2019-03-19 15:49 | 424 | 0.016191 | 0.085653 | 1.6639 | 19.9337 | 32.1524 | 0.32275
2019-03-19 15:49 | 425 | 0.016875 | 0.082433 | 1.5902 | 19.5766 | 32.1355 | 0.33036
2019-03-19 15:49 | 426 | 0.019402 | 0.072797 | 1.3883 | 16.9741 | 32.1183 | 0.32934
2019-03-19 15:49 | 427 | 0.012002 | 0.11528 | 2.2144 | 28.3969 | 32.0986 | 0.34081
2019-03-19 15:49 | 428 | 0.021366 | 0.064956 | 1.246 | 15.2283 | 32.0903 | 0.32537
2019-03-19 15:49 | 429 | 0.032519 | 0.042598 | 0.81252 | 10.0661 | 32.0555 | 0.32734
2019-03-19 15:49 | 430 | 0.026358 | 0.052508 | 0.99157 | 12.0423 | 32.0295 | 0.31742
2019-03-19 15:49 | 431 | 0.02212 | 0.062719 | 1.1827 | 14.4452 | 32.0028 | 0.31953
2019-03-19 15:49 | 432 | 0.017506 | 0.081784 | 1.4929 | 18.7539 | 31.9849 | 0.3283
2019-03-19 15:49 | 433 | 0.027804 | 0.051603 | 0.93524 | 11.8061 | 31.964 | 0.32825
2019-03-19 15:49 | 434 | 0.038211 | 0.037493 | 0.68247 | 8.6166 | 31.9291 | 0.32925
2019-03-19 15:49 | 435 | 0.040049 | 0.035217 | 0.65002 | 8.2857 | 31.8899 | 0.33183
2019-03-19 15:49 | 436 | 0.035325 | 0.04048 | 0.73682 | 9.2439 | 31.851 | 0.32654
2019-03-19 15:49 | 437 | 0.032199 | 0.044429 | 0.80989 | 9.8954 | 31.8194 | 0.31862
2019-03-19 15:49 | 438 | 0.018648 | 0.076701 | 1.4002 | 16.4028 | 31.7889 | 0.30588
2019-03-19 15:49 | 439 | 0.012687 | 0.11274 | 2.0551 | 26.5537 | 31.7796 | 0.3369
2019-03-19 15:49 | 440 | 0.035627 | 0.040892 | 0.73529 | 8.6655 | 31.7567 | 0.30872
2019-03-19 15:49 | 441 | 0.032225 | 0.045183 | 0.80779 | 9.8794 | 31.7143 | 0.31836
2019-03-19 15:49 | 442 | 0.023771 | 0.060151 | 1.0902 | 13.7918 | 31.694 | 0.32784
2019-03-19 15:50 | 443 | 0.027832 | 0.05235 | 0.93592 | 12.1662 | 31.6645 | 0.33861
2019-03-19 15:50 | 444 | 0.010445 | 0.14198 | 2.4974 | 29.5762 | 31.643 | 0.30891
2019-03-19 15:50 | 445 | 0.0072417 | 0.20479 | 3.597 | 45.049 | 31.6403 | 0.32623
2019-03-19 15:50 | 446 | 0.012248 | 0.12319 | 2.1264 | 27.2191 | 31.626 | 0.33338
2019-03-19 15:50 | 447 | 0.018113 | 0.083153 | 1.427 | 17.8926 | 31.6152 | 0.3241
2019-03-19 15:50 | 448 | 0.027511 | 0.054836 | 0.93968 | 12.5311 | 31.5892 | 0.34475
2019-03-19 15:50 | 449 | 0.021694 | 0.070631 | 1.2001 | 14.9153 | 31.5639 | 0.32357
2019-03-19 15:50 | 450 | 0.018177 | 0.084664 | 1.4323 | 18.0894 | 31.5453 | 0.32881
2019-03-19 15:50 | 451 | 0.031174 | 0.050038 | 0.83348 | 10.711 | 31.5241 | 0.3339
2019-03-19 15:50 | 452 | 0.018285 | 0.085309 | 1.4148 | 18.2261 | 31.4892 | 0.33327
2019-03-19 15:50 | 453 | -0.0015252 | -1.0381 | -16.9807 | -217.7446 | 31.489 | 0.33211
2019-03-19 15:50 | 454 | 0.013269 | 0.11921 | 1.9471 | 25.3081 | 31.4838 | 0.3358
2019-03-19 15:50 | 455 | 0.022887 | 0.069036 | 1.1325 | 14.1122 | 31.4629 | 0.32299
2019-03-19 15:50 | 456 | 0.028746 | 0.055803 | 0.88737 | 11.1249 | 31.4394 | 0.31979
2019-03-19 15:50 | 457 | 0.022955 | 0.069982 | 1.1211 | 14.0969 | 31.4083 | 0.32359
2019-03-19 15:50 | 458 | 0.023167 | 0.069358 | 1.1159 | 13.5268 | 31.3927 | 0.31337
2019-03-19 15:50 | 459 | 0.013914 | 0.11486 | 1.8345 | 22.5424 | 31.3645 | 0.31365
2019-03-19 15:50 | 460 | -0.0028609 | -0.55935 | -8.8423 | -109.9364 | 31.3664 | 0.31452
2019-03-19 15:50 | 461 | -0.001578 | -1.0316 | -16.0198 | -197.9151 | 31.3652 | 0.31231
2019-03-19 15:50 | 462 | 0.016703 | 0.095941 | 1.5166 | 19.0413 | 31.3645 | 0.31804
2019-03-19 15:50 | 463 | 0.043413 | 0.0369 | 0.58245 | 6.9498 | 31.33 | 0.30172
2019-03-19 15:50 | 464 | 0.041484 | 0.038599 | 0.61108 | 7.3099 | 31.2853 | 0.30325
2019-03-19 15:50 | 465 | 0.017486 | 0.090368 | 1.448 | 17.8959 | 31.2529 | 0.31292
2019-03-19 15:50 | 466 | 0.014332 | 0.11201 | 1.7628 | 21.1014 | 31.2453 | 0.30242
2019-03-19 15:50 | 467 | 0.031376 | 0.051089 | 0.80043 | 9.7802 | 31.2194 | 0.30686
2019-03-19 15:50 | 468 | 0.018444 | 0.088256 | 1.3726 | 17.3784 | 31.1891 | 0.32052
2019-03-19 15:50 | 469 | 0.014516 | 0.11209 | 1.7278 | 22.2 | 31.1803 | 0.32225
2019-03-19 15:50 | 470 | 0.036124 | 0.045099 | 0.69514 | 8.8793 | 31.1539 | 0.32076
2019-03-19 15:50 | 471 | 0.040237 | 0.039821 | 0.62763 | 7.7977 | 31.1121 | 0.31375
2019-03-19 15:50 | 472 | 0.035268 | 0.046023 | 0.71045 | 9.0376 | 31.0766 | 0.31873
2019-03-19 15:50 | 473 | 0.019343 | 0.082825 | 1.2977 | 16.0611 | 31.0443 | 0.31067
2019-03-19 15:50 | 474 | 0.014136 | 0.11336 | 1.7734 | 22.3743 | 31.0349 | 0.31628
2019-03-19 15:50 | 475 | 0.029752 | 0.053776 | 0.83369 | 10.2596 | 31.0111 | 0.30525
2019-03-19 15:50 | 476 | 0.020032 | 0.079932 | 1.2448 | 15.9589 | 30.9809 | 0.3197
2019-03-19 15:50 | 477 | 0.0199 | 0.080369 | 1.246 | 15.3088 | 30.9692 | 0.30465
2019-03-19 15:50 | 478 | 0.022809 | 0.071303 | 1.0925 | 14.1537 | 30.94 | 0.32284
2019-03-19 15:50 | 479 | 0.030327 | 0.053627 | 0.82537 | 10.1725 | 30.9224 | 0.30851
2019-03-19 15:50 | 480 | 0.036619 | 0.044428 | 0.68173 | 8.6253 | 30.8802 | 0.31585
2019-03-19 15:50 | 481 | 0.022672 | 0.071746 | 1.1092 | 14.4044 | 30.8547 | 0.32657
2019-03-19 15:50 | 482 | -0.0038221 | -0.43213 | -6.5643 | -82.9791 | 30.838 | 0.31715
2019-03-19 15:50 | 483 | -0.0097051 | -0.17035 | -2.5878 | -34.0537 | 30.8568 | 0.33049
2019-03-19 15:50 | 484 | 0.0080612 | 0.20436 | 3.1137 | 39.5326 | 30.851 | 0.31868
2019-03-19 15:50 | 485 | 0.013936 | 0.11816 | 1.8057 | 22.2339 | 30.8427 | 0.30986
2019-03-19 15:50 | 486 | 0.029465 | 0.055945 | 0.84984 | 10.9766 | 30.8217 | 0.32343
2019-03-19 15:50 | 487 | 0.021995 | 0.074034 | 1.148 | 12.7718 | 30.7894 | 0.28091
2019-03-19 15:50 | 488 | 0.011813 | 0.13965 | 2.1437 | 26.4127 | 30.7786 | 0.31201
2019-03-19 15:50 | 489 | 0.017493 | 0.094251 | 1.4446 | 17.9074 | 30.7621 | 0.31325
2019-03-19 15:50 | 490 | 0.018661 | 0.088492 | 1.3624 | 17.1428 | 30.7445 | 0.31991
2019-03-19 15:50 | 491 | 0.01018 | 0.16476 | 2.4782 | 32.8914 | 30.7263 | 0.33483
2019-03-19 15:50 | 492 | 0.033812 | 0.048869 | 0.7557 | 9.7756 | 30.7161 | 0.33053
2019-03-19 15:50 | 493 | 0.064744 | 0.025929 | 0.39291 | 4.7013 | 30.6573 | 0.30438
2019-03-19 15:50 | 494 | 0.050418 | 0.033265 | 0.50779 | 6.8811 | 30.5984 | 0.34693
2019-03-19 15:50 | 495 | 0.019848 | 0.084446 | 1.3005 | 17.2063 | 30.5612 | 0.3415
2019-03-19 15:50 | 496 | 0.013663 | 0.12089 | 1.886 | 24.7496 | 30.5528 | 0.33815
2019-03-19 15:50 | 497 | 0.017102 | 0.097901 | 1.5033 | 19.398 | 30.5305 | 0.33174
2019-03-19 15:50 | 498 | 0.029134 | 0.056668 | 0.87888 | 11.1114 | 30.5165 | 0.32372
2019-03-19 15:50 | 499 | 0.023628 | 0.071105 | 1.1014 | 14.3094 | 30.4766 | 0.33811
2019-03-19 15:50 | 500 | 0.0069145 | 0.23854 | 3.7322 | 47.5551 | 30.4715 | 0.32882
2019-03-19 15:50 | 501 | 0.031559 | 0.052194 | 0.81306 | 10.3155 | 30.4537 | 0.32554
2019-03-19 15:50 | 502 | 0.024032 | 0.068543 | 1.0662 | 13.5285 | 30.416 | 0.32512
2019-03-19 15:50 | 503 | -0.0059337 | -0.27713 | -4.3422 | -53.9082 | 30.4111 | 0.31987
2019-03-19 15:51 | 504 | -0.0019309 | -0.8652 | -13.227 | -165.9999 | 30.4193 | 0.32053
2019-03-19 15:51 | 505 | 0.023501 | 0.070102 | 1.0856 | 13.7543 | 30.4093 | 0.32323
2019-03-19 15:51 | 506 | 0.022571 | 0.073043 | 1.1312 | 14.4662 | 30.3784 | 0.32651
2019-03-19 15:51 | 507 | 0.018589 | 0.088654 | 1.3649 | 17.7649 | 30.3656 | 0.33023
2019-03-19 15:51 | 508 | 0.016554 | 0.099315 | 1.5358 | 19.7419 | 30.3404 | 0.32682
2019-03-19 15:51 | 509 | 0.029215 | 0.057153 | 0.86476 | 11.1477 | 30.3286 | 0.32569
2019-03-19 15:51 | 510 | 0.049039 | 0.03411 | 0.51841 | 6.6737 | 30.2813 | 0.32727
2019-03-19 15:51 | 511 | 0.016014 | 0.10603 | 1.6005 | 20.2381 | 30.2428 | 0.32409
2019-03-19 15:51 | 512 | 0.011895 | 0.14274 | 2.1452 | 27.0406 | 30.2426 | 0.32165
2019-03-19 15:51 | 513 | 0.03956 | 0.042858 | 0.64218 | 8.4073 | 30.2112 | 0.33259
2019-03-19 15:51 | 514 | 0.034283 | 0.049432 | 0.73632 | 9.517 | 30.171 | 0.32627
2019-03-19 15:51 | 515 | 0.022726 | 0.074617 | 1.1164 | 14.8213 | 30.145 | 0.33683
2019-03-19 15:51 | 516 | 0.016396 | 0.10352 | 1.5489 | 20.176 | 30.1242 | 0.33081
2019-03-19 15:51 | 517 | 0.019416 | 0.087342 | 1.3049 | 17.2837 | 30.1095 | 0.33559
2019-03-19 15:51 | 518 | 0.028929 | 0.0586 | 0.87676 | 11.8703 | 30.0836 | 0.3434
2019-03-19 15:51 | 519 | 0.034766 | 0.048912 | 0.73505 | 9.9948 | 30.0534 | 0.34748
2019-03-19 15:51 | 520 | 0.002735 | 0.63132 | 9.3144 | 121.2984 | 30.0225 | 0.33175
2019-03-19 15:51 | 521 | 0.012537 | 0.13731 | 2.0109 | 26.9461 | 30.038 | 0.33782
2019-03-19 15:51 | 522 | 0.039824 | 0.04332 | 0.63536 | 8.65 | 29.9933 | 0.34448
2019-03-19 15:51 | 523 | 0.01587 | 0.10874 | 1.5948 | 20.6875 | 29.97 | 0.32832
2019-03-19 15:51 | 524 | 0.015963 | 0.10655 | 1.5988 | 21.471 | 29.9564 | 0.34274
2019-03-19 15:51 | 525 | 0.035068 | 0.047729 | 0.72003 | 9.6527 | 29.9331 | 0.3385
2019-03-19 15:51 | 526 | 0.039676 | 0.042251 | 0.64119 | 8.6242 | 29.8898 | 0.34218
2019-03-19 15:51 | 527 | 0.026609 | 0.061996 | 0.94739 | 13.2577 | 29.8584 | 0.35277
2019-03-19 15:51 | 528 | 0.023734 | 0.06952 | 1.0687 | 13.9248 | 29.8349 | 0.33049
2019-03-19 15:51 | 529 | 0.0082738 | 0.19906 | 3.0782 | 40.1451 | 29.8132 | 0.33215
2019-03-19 15:51 | 530 | 0.011966 | 0.13562 | 2.1156 | 28.6279 | 29.8134 | 0.34257
2019-03-19 15:51 | 531 | 0.033649 | 0.048856 | 0.75216 | 10.4867 | 29.7852 | 0.35287
2019-03-19 15:51 | 532 | 0.019409 | 0.083584 | 1.298 | 18.3375 | 29.754 | 0.3559
2019-03-19 15:51 | 533 | 0.025381 | 0.063955 | 1.0094 | 14.3024 | 29.7425 | 0.36301
2019-03-19 15:51 | 534 | 0.027245 | 0.060607 | 0.94412 | 12.853 | 29.7045 | 0.35018
2019-03-19 15:51 | 535 | 0.0048126 | 0.34358 | 5.3365 | 74.4147 | 29.6932 | 0.35813
2019-03-19 15:51 | 536 | 0.015559 | 0.10787 | 1.6709 | 22.3728 | 29.6871 | 0.34809
2019-03-19 15:51 | 537 | 0.030087 | 0.055794 | 0.86173 | 11.5363 | 29.6608 | 0.34709
2019-03-19 15:51 | 538 | 0.022951 | 0.0731 | 1.1372 | 15.3756 | 29.632 | 0.35289
2019-03-19 15:51 | 539 | 0.019207 | 0.088408 | 1.3599 | 17.7478 | 29.6142 | 0.34088
2019-03-19 15:51 | 540 | 0.034428 | 0.049371 | 0.75944 | 9.9892 | 29.5894 | 0.3439
2019-03-19 15:51 | 541 | 0.027673 | 0.061305 | 0.94301 | 12.3702 | 29.5501 | 0.34232
2019-03-19 15:51 | 542 | 0.029765 | 0.05847 | 0.87602 | 11.8344 | 29.5327 | 0.35224
2019-03-19 15:51 | 543 | 0.020761 | 0.083567 | 1.252 | 16.7603 | 29.4929 | 0.34797
2019-03-19 15:51 | 544 | 0.02335 | 0.075206 | 1.1147 | 15.0933 | 29.4882 | 0.35243
2019-03-19 15:51 | 545 | 0.037954 | 0.046829 | 0.69075 | 9.1412 | 29.4435 | 0.34694
2019-03-19 15:51 | 546 | 0.026547 | 0.06629 | 0.98821 | 13.3826 | 29.4184 | 0.35527
2019-03-19 15:51 | 547 | 0.0217 | 0.082075 | 1.2107 | 16.9871 | 29.3894 | 0.36862
2019-03-19 15:51 | 548 | 0.021232 | 0.08521 | 1.2368 | 17.2003 | 29.3736 | 0.3652
2019-03-19 15:51 | 549 | 0.024069 | 0.073964 | 1.1034 | 15.3364 | 29.3459 | 0.36914
2019-03-19 15:51 | 550 | 0.026671 | 0.067914 | 1.0012 | 13.8396 | 29.3253 | 0.36912
2019-03-19 15:51 | 551 | 0.036659 | 0.0494 | 0.73032 | 10.1905 | 29.2912 | 0.37358
2019-03-19 15:51 | 552 | 0.029605 | 0.061175 | 0.90567 | 12.4311 | 29.2561 | 0.36803
2019-03-19 15:51 | 553 | 0.022559 | 0.079238 | 1.192 | 16.2745 | 29.2323 | 0.36714
2019-03-19 15:51 | 554 | 0.02291 | 0.078088 | 1.1813 | 16.1488 | 29.209 | 0.36996
2019-03-19 15:51 | 555 | 0.025288 | 0.070602 | 1.0666 | 14.5888 | 29.1855 | 0.36893
2019-03-19 15:51 | 556 | 0.036006 | 0.048935 | 0.74923 | 10.1628 | 29.1567 | 0.36592
2019-03-19 15:51 | 557 | 0.031349 | 0.056851 | 0.8624 | 11.6525 | 29.1171 | 0.36529
2019-03-19 15:51 | 558 | 0.024021 | 0.074328 | 1.1306 | 15.3152 | 29.0941 | 0.36789
2019-03-19 15:51 | 559 | 0.052443 | 0.034434 | 0.51975 | 6.9468 | 29.0612 | 0.36431
2019-03-19 15:51 | 560 | 0.047177 | 0.038392 | 0.5771 | 7.5802 | 28.9975 | 0.35761
2019-03-19 15:51 | 561 | 0.014061 | 0.12841 | 1.9223 | 27.4949 | 28.9734 | 0.3866
2019-03-19 15:51 | 562 | 0.026487 | 0.069091 | 1.0237 | 14.4525 | 28.9584 | 0.38281
2019-03-19 15:52 | 563 | 0.050358 | 0.036405 | 0.54385 | 7.6088 | 28.9172 | 0.38317
2019-03-19 15:52 | 564 | 0.040272 | 0.046073 | 0.68038 | 9.3791 | 28.8658 | 0.37771
2019-03-19 15:52 | 565 | 0.034931 | 0.053843 | 0.78613 | 11.0638 | 28.8365 | 0.38646
2019-03-19 15:52 | 566 | 0.028062 | 0.067881 | 0.98669 | 13.4404 | 28.7958 | 0.37717
2019-03-19 15:52 | 567 | 0.035371 | 0.053859 | 0.78817 | 10.8219 | 28.7769 | 0.38278
2019-03-19 15:52 | 568 | 0.043 | 0.044817 | 0.64775 | 9.1103 | 28.7255 | 0.39174
2019-03-19 15:52 | 569 | 0.018295 | 0.10516 | 1.5219 | 22.0967 | 28.6982 | 0.40427
2019-03-19 15:52 | 570 | 0.019411 | 0.10031 | 1.4495 | 20.2989 | 28.6825 | 0.39403
2019-03-19 15:52 | 571 | 0.046674 | 0.041746 | 0.60001 | 8.7038 | 28.6528 | 0.40624
2019-03-19 15:52 | 572 | 0.049009 | 0.040247 | 0.57683 | 8.2128 | 28.595 | 0.4025
2019-03-19 15:52 | 573 | 0.039476 | 0.050576 | 0.71718 | 10.0037 | 28.5582 | 0.39491
2019-03-19 15:52 | 574 | 0.04389 | 0.046131 | 0.65157 | 9.1441 | 28.5136 | 0.40133
2019-03-19 15:52 | 575 | 0.026439 | 0.078349 | 1.0826 | 15.0441 | 28.4753 | 0.39776
2019-03-19 15:52 | 576 | 0.018389 | 0.11527 | 1.5732 | 22.0112 | 28.4586 | 0.40477
2019-03-19 15:52 | 577 | 0.020498 | 0.10573 | 1.4151 | 19.4265 | 28.4358 | 0.3982
2019-03-19 15:52 | 578 | 0.014322 | 0.15632 | 2.0383 | 28.3291 | 28.4185 | 0.40574
2019-03-19 15:52 | 579 | 0.035063 | 0.064603 | 0.83246 | 11.2147 | 28.4005 | 0.39322
2019-03-19 15:52 | 580 | 0.059332 | 0.03898 | 0.49446 | 6.8683 | 28.3478 | 0.40751
2019-03-19 15:52 | 581 | 0.048744 | 0.048778 | 0.60325 | 8.79 | 28.2903 | 0.42846
2019-03-19 15:52 | 582 | 0.046763 | 0.051922 | 0.63194 | 9.3086 | 28.2491 | 0.4353
2019-03-19 15:52 | 583 | 0.046629 | 0.053425 | 0.63724 | 9.6945 | 28.1965 | 0.45205
2019-03-19 15:52 | 584 | 0.031877 | 0.079375 | 0.93374 | 14.639 | 28.1583 | 0.46664
2019-03-19 15:52 | 585 | 0.054828 | 0.047059 | 0.54275 | 11.1287 | 28.1233 | 0.61016
2019-03-19 15:52 | 586 | 0.084507 | 0.030393 | 0.35667 | 8.5743 | 28.0465 | 0.72458
2019-03-19 15:52 | 587 | 0.089482 | 0.028919 | 0.3462 | 9.0658 | 27.96 | 0.81123
2019-03-19 15:52 | 588 | 0.11049 | 0.02357 | 0.29464 | 8.3453 | 27.8657 | 0.92204
2019-03-19 15:52 | 589 | 0.08692 | 0.029905 | 0.39755 | 12.0809 | 27.7502 | 1.0501
2019-03-19 15:52 | 590 | 0.054492 | 0.047905 | 0.68827 | 21.1773 | 27.6942 | 1.154
2019-03-19 15:52 | 591 | 0.060782 | 0.043184 | 0.6774 | 20.4519 | 27.6316 | 1.2431
2019-03-19 15:52 | 592 | 0.070154 | 0.037605 | 0.6495 | 19.429 | 27.5704 | 1.363
2019-03-19 15:52 | 593 | 0.08672 | 0.031121 | 0.58142 | 17.3063 | 27.4891 | 1.5008
2019-03-19 15:52 | 594 | 0.10766 | 0.025554 | 0.51735 | 15.661 | 27.3959 | 1.6861
2019-03-19 15:52 | 595 | 0.12298 | 0.023115 | 0.52538 | 15.6716 | 27.2758 | 1.9273
2019-03-19 15:52 | 596 | 0.11563 | 0.025928 | 0.64202 | 17.7076 | 27.1562 | 2.0476
2019-03-19 15:52 | 597 | 0.097842 | 0.032265 | 0.86084 | 22.2327 | 27.0473 | 2.1753
2019-03-19 15:52 | 598 | 0.09744 | 0.034248 | 0.97235 | 22.7628 | 26.9557 | 2.218
2019-03-19 15:52 | 599 | 0.12001 | 0.029497 | 0.8947 | 20.4811 | 26.846 | 2.458
2019-03-19 15:52 | 600 | 0.14233 | 0.026297 | 0.83722 | 18.5188 | 26.7148 | 2.6358
2019-03-19 15:52 | 601 | 0.17076 | 0.022913 | 0.78299 | 16.867 | 26.5605 | 2.8802
2019-03-19 15:52 | 602 | 0.17381 | 0.023398 | 0.84906 | 17.9928 | 26.3797 | 3.1272
2019-03-19 15:52 | 603 | 0.17563 | 0.023762 | 0.91865 | 19.6084 | 26.2141 | 3.4439
2019-03-19 15:52 | 604 | 0.17686 | 0.024117 | 0.99925 | 20.6596 | 26.0295 | 3.6538
2019-03-19 15:52 | 605 | 0.15093 | 0.02873 | 1.2596 | 25.564 | 25.8666 | 3.8585
2019-03-19 15:52 | 606 | 0.14669 | 0.02977 | 1.3884 | 26.513 | 25.7218 | 3.8891
2019-03-19 15:52 | 607 | 0.17903 | 0.024324 | 1.205 | 21.8773 | 25.5638 | 3.9166
2019-03-19 15:52 | 608 | 0.19441 | 0.022254 | 1.1526 | 20.6031 | 25.3675 | 4.0054
2019-03-19 15:52 | 609 | 0.19177 | 0.022437 | 1.2169 | 21.568 | 25.1803 | 4.1362
2019-03-19 15:52 | 610 | 0.18129 | 0.023718 | 1.3524 | 23.4586 | 24.9867 | 4.2529
2019-03-19 15:52 | 611 | 0.16543 | 0.026047 | 1.5244 | 25.997 | 24.819 | 4.3006
2019-03-19 15:52 | 612 | 0.16135 | 0.026767 | 1.6086 | 26.2747 | 24.6529 | 4.2395
2019-03-19 15:52 | 613 | 0.16225 | 0.026858 | 1.6456 | 26.0186 | 24.4947 | 4.2216
2019-03-19 15:52 | 614 | 0.16076 | 0.027192 | 1.7042 | 26.004 | 24.3285 | 4.1803
2019-03-19 15:52 | 615 | 0.16824 | 0.026188 | 1.6725 | 24.8204 | 24.171 | 4.1758
2019-03-19 15:52 | 616 | 0.17732 | 0.02497 | 1.6181 | 23.3424 | 23.9921 | 4.1391
2019-03-19 15:52 | 617 | 0.16286 | 0.027182 | 1.7868 | 25.1987 | 23.8224 | 4.1037
2019-03-19 15:52 | 618 | 0.14646 | 0.030435 | 2.017 | 28.0041 | 23.6675 | 4.1014
2019-03-19 15:52 | 619 | 0.13395 | 0.033549 | 2.225 | 30.4618 | 23.5282 | 4.0802
2019-03-19 15:52 | 620 | 0.13555 | 0.033541 | 2.2224 | 29.5658 | 23.396 | 4.0075
2019-03-19 15:52 | 621 | 0.13237 | 0.034532 | 2.2932 | 29.6455 | 23.258 | 3.9242
2019-03-19 15:52 | 622 | 0.12424 | 0.037193 | 2.4561 | 31.2808 | 23.1326 | 3.8862
2019-03-19 15:52 | 623 | 0.11723 | 0.039825 | 2.6212 | 32.95 | 23.0092 | 3.8627
2019-03-19 15:53 | 624 | 0.11896 | 0.039154 | 2.5844 | 31.1899 | 22.8962 | 3.7103
2019-03-19 15:53 | 625 | 0.11051 | 0.042086 | 2.7804 | 31.8081 | 22.7738 | 3.515
2019-03-19 15:53 | 626 | 0.094644 | 0.048609 | 3.2354 | 36.0208 | 22.6771 | 3.4092
2019-03-19 15:53 | 627 | 0.088069 | 0.051439 | 3.4555 | 37.008 | 22.5823 | 3.2593
2019-03-19 15:53 | 628 | 0.083339 | 0.052999 | 3.6006 | 36.8205 | 22.4998 | 3.0686
2019-03-19 15:53 | 629 | 0.092136 | 0.046469 | 3.2149 | 32.4744 | 22.4123 | 2.9921
2019-03-19 15:53 | 630 | 0.095604 | 0.043339 | 3.0419 | 28.4129 | 22.3169 | 2.7164
2019-03-19 15:53 | 631 | 0.084486 | 0.047314 | 3.3606 | 29.6798 | 22.2248 | 2.5075
2019-03-19 15:53 | 632 | 0.07794 | 0.049395 | 3.5094 | 29.7165 | 22.1473 | 2.3161
2019-03-19 15:53 | 633 | 0.071399 | 0.051857 | 3.6382 | 31.0832 | 22.0689 | 2.2193
2019-03-19 15:53 | 634 | 0.058862 | 0.060769 | 4.2254 | 36.2613 | 22.0051 | 2.1344
2019-03-19 15:53 | 635 | 0.073466 | 0.047017 | 3.2156 | 28.2523 | 21.9443 | 2.0756
2019-03-19 15:53 | 636 | 0.094424 | 0.035299 | 2.4085 | 20.9201 | 21.8566 | 1.9754
2019-03-19 15:53 | 637 | 0.085552 | 0.037562 | 2.5742 | 22.4174 | 21.7626 | 1.9179
2019-03-19 15:53 | 638 | 0.0815 | 0.038266 | 2.5999 | 22.4283 | 21.6847 | 1.8279
2019-03-19 15:53 | 639 | 0.095865 | 0.031933 | 2.1332 | 18.5616 | 21.5956 | 1.7794
2019-03-19 15:53 | 640 | 0.085088 | 0.034939 | 2.3126 | 21.0565 | 21.4989 | 1.7917
2019-03-19 15:53 | 641 | 0.070855 | 0.040703 | 2.6799 | 25.3101 | 21.4265 | 1.7933
2019-03-19 15:53 | 642 | 0.073556 | 0.038341 | 2.4927 | 23.9336 | 21.3529 | 1.7605
2019-03-19 15:53 | 643 | 0.08198 | 0.033848 | 2.1632 | 21.5378 | 21.2776 | 1.7657
2019-03-19 15:53 | 644 | 0.085917 | 0.031747 | 1.9988 | 20.4061 | 21.1903 | 1.7532
2019-03-19 15:53 | 645 | 0.07552 | 0.035498 | 2.2076 | 22.8988 | 21.1097 | 1.7293
2019-03-19 15:53 | 646 | 0.061252 | 0.043166 | 2.6598 | 28.0023 | 21.0403 | 1.7152
2019-03-19 15:53 | 647 | 0.059064 | 0.044397 | 2.7 | 29.2638 | 20.9841 | 1.7284
2019-03-19 15:53 | 648 | 0.06126 | 0.042491 | 2.5465 | 28.1098 | 20.9211 | 1.722
2019-03-19 15:53 | 649 | 0.048695 | 0.052395 | 3.1367 | 35.038 | 20.8647 | 1.7062
2019-03-19 15:53 | 650 | 0.053468 | 0.047369 | 2.8095 | 31.9006 | 20.8197 | 1.7057
2019-03-19 15:53 | 651 | 0.057186 | 0.043983 | 2.5925 | 30.1613 | 20.7572 | 1.7248
2019-03-19 15:53 | 652 | 0.072955 | 0.03393 | 2.0096 | 23.72 | 20.7028 | 1.7305
2019-03-19 15:53 | 653 | 0.07056 | 0.034715 | 2.0475 | 24.1945 | 20.6158 | 1.7072
2019-03-19 15:53 | 654 | 0.056516 | 0.042561 | 2.5296 | 30.1924 | 20.5644 | 1.7064
2019-03-19 15:53 | 655 | 0.068695 | 0.034672 | 2.0591 | 24.6787 | 20.497 | 1.6953
2019-03-19 15:53 | 656 | 0.065589 | 0.035987 | 2.1371 | 25.6799 | 20.43 | 1.6843
2019-03-19 15:53 | 657 | 0.06915 | 0.03349 | 2.0131 | 24.1278 | 20.3647 | 1.6684
2019-03-19 15:53 | 658 | 0.066754 | 0.034378 | 2.0719 | 24.8637 | 20.2932 | 1.6598
2019-03-19 15:53 | 659 | 0.058682 | 0.039063 | 2.3344 | 28.2607 | 20.2321 | 1.6584
2019-03-19 15:53 | 660 | 0.072486 | 0.031646 | 1.8761 | 22.6065 | 20.1704 | 1.6386
2019-03-19 15:53 | 661 | 0.090271 | 0.025412 | 1.4944 | 18.162 | 20.0869 | 1.6395
2019-03-19 15:53 | 662 | 0.059567 | 0.038002 | 2.2362 | 27.486 | 20.0013 | 1.6372
2019-03-19 15:53 | 663 | 0.051703 | 0.043889 | 2.5682 | 31.9341 | 19.9625 | 1.6511
2019-03-19 15:53 | 664 | 0.076686 | 0.029606 | 1.7186 | 21.1512 | 19.8897 | 1.622
2019-03-19 15:53 | 665 | 0.073263 | 0.031039 | 1.7912 | 22.6916 | 19.8149 | 1.6625
2019-03-19 15:53 | 666 | 0.081627 | 0.028177 | 1.6053 | 20.1558 | 19.7411 | 1.6452
2019-03-19 15:53 | 667 | 0.086051 | 0.026641 | 1.5108 | 20.0542 | 19.6528 | 1.7257
2019-03-19 15:53 | 668 | 0.076661 | 0.029581 | 1.6914 | 22.5645 | 19.5723 | 1.7298
2019-03-19 15:53 | 669 | 0.079941 | 0.02816 | 1.6262 | 21.9627 | 19.4971 | 1.7557
2019-03-19 15:53 | 670 | 0.070425 | 0.031528 | 1.8407 | 23.9888 | 19.4154 | 1.6894
2019-03-19 15:53 | 671 | 0.058473 | 0.038097 | 2.2284 | 28.7691 | 19.3563 | 1.6822
2019-03-19 15:53 | 672 | 0.079975 | 0.027547 | 1.6282 | 21.0613 | 19.2909 | 1.6844
2019-03-19 15:53 | 673 | 0.072687 | 0.030395 | 1.7955 | 22.8928 | 19.2028 | 1.664
2019-03-19 15:53 | 674 | 0.058806 | 0.037566 | 2.2191 | 28.3426 | 19.1476 | 1.6667
2019-03-19 15:53 | 675 | 0.060712 | 0.036068 | 2.1485 | 27.4923 | 19.0823 | 1.6691
2019-03-19 15:53 | 676 | 0.032155 | 0.067362 | 4.0401 | 51.5168 | 19.0324 | 1.6565
2019-03-19 15:53 | 677 | 0.036852 | 0.058896 | 3.5275 | 45.2098 | 19.0095 | 1.6661
2019-03-19 15:53 | 678 | 0.072953 | 0.029785 | 1.7791 | 23.0286 | 18.9511 | 1.68
2019-03-19 15:53 | 679 | 0.061226 | 0.035454 | 2.1142 | 27.287 | 18.8742 | 1.6707
2019-03-19 15:53 | 680 | 0.064664 | 0.033241 | 2.0009 | 26.0235 | 18.8255 | 1.6828
2019-03-19 15:53 | 681 | 0.091984 | 0.023348 | 1.4054 | 18.2498 | 18.74 | 1.6787
2019-03-19 15:53 | 682 | 0.066715 | 0.031876 | 1.9349 | 25.0954 | 18.6534 | 1.6742
2019-03-19 15:54 | 683 | 0.070884 | 0.029615 | 1.8175 | 23.6988 | 18.6006 | 1.6799
2019-03-19 15:54 | 684 | 0.076581 | 0.027113 | 1.6806 | 21.8213 | 18.5109 | 1.6711
2019-03-19 15:54 | 685 | 0.063205 | 0.032448 | 2.0306 | 26.4124 | 18.4509 | 1.6694
2019-03-19 15:54 | 686 | 0.096497 | 0.021282 | 1.3318 | 17.2121 | 18.3743 | 1.6609
2019-03-19 15:54 | 687 | 0.094392 | 0.021472 | 1.3594 | 17.6811 | 18.2664 | 1.669
2019-03-19 15:54 | 688 | 0.060912 | 0.033296 | 2.1062 | 27.3261 | 18.1931 | 1.6645
2019-03-19 15:54 | 689 | 0.073424 | 0.02789 | 1.743 | 23.1019 | 18.1342 | 1.6962
2019-03-19 15:54 | 690 | 0.077448 | 0.026136 | 1.6474 | 21.9075 | 18.0478 | 1.6967
2019-03-19 15:54 | 691 | 0.054993 | 0.037337 | 2.3334 | 31.2405 | 17.985 | 1.718
2019-03-19 15:54 | 692 | 0.071054 | 0.028872 | 1.8044 | 24.3428 | 17.9291 | 1.7297
2019-03-19 15:54 | 693 | 0.077744 | 0.026398 | 1.6529 | 22.2429 | 17.8442 | 1.7292
2019-03-19 15:54 | 694 | 0.084532 | 0.024271 | 1.5252 | 20.5652 | 17.7738 | 1.7384
2019-03-19 15:54 | 695 | 0.093618 | 0.02169 | 1.3797 | 18.4185 | 17.6753 | 1.7243
2019-03-19 15:54 | 696 | 0.077463 | 0.026174 | 1.6694 | 21.8363 | 17.5924 | 1.6915
2019-03-19 15:54 | 697 | 0.078468 | 0.025862 | 1.6513 | 21.3351 | 17.5169 | 1.6741
2019-03-19 15:54 | 698 | 0.07536 | 0.026904 | 1.7174 | 22.4169 | 17.436 | 1.6893
2019-03-19 15:54 | 699 | 0.075467 | 0.026904 | 1.7174 | 22.5167 | 17.3654 | 1.6993
2019-03-19 15:54 | 700 | 0.079051 | 0.025916 | 1.6348 | 21.5535 | 17.2846 | 1.7038
2019-03-19 15:54 | 701 | 0.06339 | 0.031977 | 2.0357 | 27.0979 | 17.2116 | 1.7178
2019-03-19 15:54 | 702 | 0.062635 | 0.032696 | 2.0591 | 27.6403 | 17.1544 | 1.7312
2019-03-19 15:54 | 703 | 0.070519 | 0.028992 | 1.8283 | 24.541 | 17.084 | 1.7306
2019-03-19 15:54 | 704 | 0.068576 | 0.030099 | 1.8786 | 25.2133 | 17.0153 | 1.729
2019-03-19 15:54 | 705 | 0.073044 | 0.028246 | 1.7658 | 23.6024 | 16.9455 | 1.724
2019-03-19 15:54 | 706 | 0.075723 | 0.027548 | 1.7106 | 22.8607 | 16.8696 | 1.7311
2019-03-19 15:54 | 707 | 0.07641 | 0.027579 | 1.6981 | 22.7222 | 16.7945 | 1.7362
2019-03-19 15:54 | 708 | 0.08045 | 0.02675 | 1.6162 | 21.7134 | 16.7162 | 1.7468
2019-03-19 15:54 | 709 | 0.079437 | 0.027361 | 1.6418 | 21.773 | 16.6351 | 1.7296
2019-03-19 15:54 | 710 | 0.0677 | 0.032488 | 1.9343 | 25.7031 | 16.5597 | 1.7401
2019-03-19 15:54 | 711 | 0.072532 | 0.030305 | 1.8067 | 23.982 | 16.4958 | 1.7395
2019-03-19 15:54 | 712 | 0.079914 | 0.027471 | 1.6386 | 21.5555 | 16.4138 | 1.7226
2019-03-19 15:54 | 713 | 0.074802 | 0.02966 | 1.7537 | 22.6506 | 16.3386 | 1.6943
2019-03-19 15:54 | 714 | 0.084471 | 0.026253 | 1.5506 | 20.1242 | 16.2613 | 1.6999
2019-03-19 15:54 | 715 | 0.079441 | 0.028204 | 1.647 | 21.0925 | 16.1731 | 1.6756
2019-03-19 15:54 | 716 | 0.069904 | 0.03199 | 1.8635 | 24.0988 | 16.1037 | 1.6846
2019-03-19 15:54 | 717 | 0.069468 | 0.031858 | 1.87 | 24.1542 | 16.0313 | 1.678
2019-03-19 15:54 | 718 | 0.064688 | 0.034242 | 2.0088 | 26.1783 | 15.9656 | 1.6934
2019-03-19 15:54 | 719 | 0.060771 | 0.036026 | 2.1316 | 27.9871 | 15.9018 | 1.7008
2019-03-19 15:54 | 720 | 0.05567 | 0.038901 | 2.3248 | 30.6537 | 15.8438 | 1.7065
2019-03-19 15:54 | 721 | 0.065686 | 0.032924 | 1.9655 | 25.9293 | 15.7867 | 1.7032
2019-03-19 15:54 | 722 | 0.072632 | 0.029869 | 1.7821 | 23.4804 | 15.713 | 1.7054
2019-03-19 15:54 | 723 | 0.075518 | 0.029007 | 1.7135 | 22.6298 | 15.6431 | 1.709
2019-03-19 15:54 | 724 | 0.062084 | 0.035299 | 2.083 | 27.6629 | 15.566 | 1.7174
2019-03-19 15:54 | 725 | 0.05498 | 0.040267 | 2.3566 | 31.3588 | 15.5173 | 1.7241
2019-03-19 15:54 | 726 | 0.064349 | 0.034789 | 2.0151 | 26.8583 | 15.4521 | 1.7283
2019-03-19 15:54 | 727 | 0.059046 | 0.037997 | 2.2021 | 28.5184 | 15.3914 | 1.6839
2019-03-19 15:54 | 728 | 0.068833 | 0.03255 | 1.8895 | 24.0239 | 15.3306 | 1.6536
2019-03-19 15:54 | 729 | 0.075876 | 0.029787 | 1.7106 | 21.6714 | 15.2542 | 1.6443
2019-03-19 15:54 | 730 | 0.077584 | 0.029089 | 1.666 | 21.094 | 15.1805 | 1.6365
2019-03-19 15:54 | 731 | 0.074086 | 0.030804 | 1.7393 | 21.929 | 15.1008 | 1.6246
2019-03-19 15:54 | 732 | 0.061067 | 0.037366 | 2.1018 | 26.896 | 15.0346 | 1.6425
2019-03-19 15:54 | 733 | 0.057663 | 0.039872 | 2.2094 | 28.225 | 14.9763 | 1.6275
2019-03-19 15:54 | 734 | 0.062158 | 0.037041 | 2.0469 | 26.259 | 14.9174 | 1.6322
2019-03-19 15:54 | 735 | 0.053886 | 0.042717 | 2.3514 | 30.0169 | 14.8546 | 1.6175
2019-03-19 15:54 | 736 | 0.054572 | 0.042152 | 2.3126 | 29.502 | 14.8071 | 1.61
2019-03-19 15:54 | 737 | 0.074354 | 0.03095 | 1.6904 | 21.6833 | 14.7411 | 1.6122
2019-03-19 15:54 | 738 | 0.072597 | 0.031747 | 1.7267 | 22.1365 | 14.6639 | 1.607
2019-03-19 15:54 | 739 | 0.05275 | 0.043778 | 2.3697 | 30.5497 | 14.6007 | 1.6115
2019-03-19 15:54 | 740 | 0.046231 | 0.049966 | 2.6971 | 34.7488 | 14.5551 | 1.6065
2019-03-19 15:54 | 741 | 0.05366 | 0.043087 | 2.3198 | 30.0306 | 14.504 | 1.6114
2019-03-19 15:54 | 742 | 0.068805 | 0.033584 | 1.8051 | 23.0727 | 14.4455 | 1.5875
2019-03-19 15:54 | 743 | 0.079254 | 0.029221 | 1.5617 | 20.0734 | 14.3678 | 1.5909
2019-03-19 15:55 | 744 | 0.072692 | 0.031886 | 1.6973 | 21.9559 | 14.291 | 1.596
2019-03-19 15:55 | 745 | 0.082539 | 0.027808 | 1.4888 | 19.3161 | 14.219 | 1.5943
2019-03-19 15:55 | 746 | 0.086398 | 0.025998 | 1.4123 | 18.5616 | 14.1276 | 1.6037
2019-03-19 15:55 | 747 | 0.069703 | 0.031612 | 1.7508 | 23.2011 | 14.0511 | 1.6172
2019-03-19 15:55 | 748 | 0.070631 | 0.030913 | 1.7247 | 22.9392 | 13.9841 | 1.6202
2019-03-19 15:55 | 749 | 0.080208 | 0.026881 | 1.515 | 20.1686 | 13.9073 | 1.6177
2019-03-19 15:55 | 750 | 0.085334 | 0.025198 | 1.4176 | 18.8511 | 13.8248 | 1.6086
2019-03-19 15:55 | 751 | 0.080122 | 0.026588 | 1.5142 | 20.4511 | 13.7397 | 1.6386
2019-03-19 15:55 | 752 | 0.065592 | 0.032126 | 1.8491 | 24.7746 | 13.667 | 1.625
2019-03-19 15:55 | 753 | 0.061064 | 0.034183 | 1.9933 | 26.8011 | 13.606 | 1.6366
2019-03-19 15:55 | 754 | 0.06406 | 0.032205 | 1.9016 | 25.6792 | 13.5429 | 1.645
2019-03-19 15:55 | 755 | 0.060333 | 0.034273 | 2.0267 | 27.1387 | 13.4794 | 1.6374
2019-03-19 15:55 | 756 | 0.052992 | 0.038662 | 2.3173 | 30.7045 | 13.4228 | 1.6271
2019-03-19 15:55 | 757 | 0.062734 | 0.03265 | 1.9587 | 25.867 | 13.3692 | 1.6227
2019-03-19 15:55 | 758 | 0.07232 | 0.02831 | 1.7027 | 22.3599 | 13.2972 | 1.6171
2019-03-19 15:55 | 759 | 0.069545 | 0.029412 | 1.768 | 22.8539 | 13.228 | 1.5894
2019-03-19 15:55 | 760 | 0.055718 | 0.036728 | 2.204 | 28.348 | 13.161 | 1.5795
2019-03-19 15:55 | 761 | 0.053174 | 0.038562 | 2.3107 | 29.6988 | 13.1135 | 1.5792
2019-03-19 15:55 | 762 | 0.071106 | 0.028838 | 1.7205 | 22.1107 | 13.05 | 1.5722
2019-03-19 15:55 | 763 | 0.058348 | 0.034663 | 2.0838 | 27.0025 | 12.9781 | 1.5755
2019-03-19 15:55 | 764 | 0.054759 | 0.036856 | 2.2098 | 28.6348 | 12.9311 | 1.568
2019-03-19 15:55 | 765 | 0.077905 | 0.025305 | 1.5468 | 20.1868 | 12.8625 | 1.5726
2019-03-19 15:55 | 766 | 0.0719 | 0.027379 | 1.6672 | 21.6057 | 12.7824 | 1.5534
2019-03-19 15:55 | 767 | 0.04626 | 0.042198 | 2.5892 | 34.0203 | 12.724 | 1.5738
2019-03-19 15:55 | 768 | 0.036682 | 0.052742 | 3.2651 | 42.4013 | 12.6855 | 1.5554
2019-03-19 15:55 | 769 | 0.058319 | 0.033113 | 2.0427 | 26.7406 | 12.6424 | 1.5595
2019-03-19 15:55 | 770 | 0.075741 | 0.025179 | 1.5668 | 20.6607 | 12.5697 | 1.5649
2019-03-19 15:55 | 771 | 0.06778 | 0.028143 | 1.7471 | 23.1887 | 12.4974 | 1.5717
2019-03-19 15:55 | 772 | 0.062388 | 0.030591 | 1.8963 | 24.8675 | 12.4343 | 1.5514
2019-03-19 15:55 | 773 | 0.059739 | 0.03152 | 1.9736 | 26.1115 | 12.3721 | 1.5599
2019-03-19 15:55 | 774 | 0.051184 | 0.036788 | 2.3 | 30.413 | 12.3163 | 1.5567
2019-03-19 15:55 | 775 | 0.037146 | 0.050829 | 3.1731 | 41.3121 | 12.2706 | 1.5346
2019-03-19 15:55 | 776 | 0.044454 | 0.042413 | 2.6429 | 34.3496 | 12.2364 | 1.527
2019-03-19 15:55 | 777 | 0.063361 | 0.029769 | 1.8473 | 23.833 | 12.1786 | 1.5101
2019-03-19 15:55 | 778 | 0.067695 | 0.027823 | 1.7234 | 22.2834 | 12.1137 | 1.5085
2019-03-19 15:55 | 779 | 0.04508 | 0.041869 | 2.5824 | 33.1507 | 12.0496 | 1.4944
2019-03-19 15:55 | 780 | 0.050744 | 0.037281 | 2.2876 | 29.5925 | 12.0166 | 1.5016
2019-03-19 15:55 | 781 | 0.080526 | 0.023178 | 1.4306 | 18.6096 | 11.9428 | 1.4986
2019-03-19 15:55 | 782 | 0.057615 | 0.032386 | 1.9891 | 26.0284 | 11.868 | 1.4996
2019-03-19 15:55 | 783 | 0.042835 | 0.043645 | 2.6732 | 35.0348 | 11.8265 | 1.5007
2019-03-19 15:55 | 784 | 0.043379 | 0.042532 | 2.6229 | 34.3504 | 11.7777 | 1.4901
2019-03-19 15:55 | 785 | 0.061855 | 0.029511 | 1.8372 | 24.2183 | 11.7349 | 1.498
2019-03-19 15:55 | 786 | 0.075373 | 0.024235 | 1.5074 | 19.9677 | 11.6557 | 1.505
2019-03-19 15:55 | 787 | 0.054933 | 0.032787 | 2.059 | 27.0973 | 11.5926 | 1.4885
2019-03-19 15:55 | 788 | 0.045167 | 0.039916 | 2.5059 | 33.1421 | 11.544 | 1.4969
2019-03-19 15:55 | 789 | 0.043444 | 0.041945 | 2.5966 | 34.5781 | 11.5003 | 1.5022
2019-03-19 15:55 | 790 | 0.032011 | 0.056184 | 3.5128 | 46.8879 | 11.4591 | 1.5009
2019-03-19 15:55 | 791 | 0.024863 | 0.072445 | 4.5195 | 59.0806 | 11.4348 | 1.4689
2019-03-19 15:55 | 792 | 0.037253 | 0.049026 | 3.0122 | 38.9352 | 11.4043 | 1.4505
2019-03-19 15:55 | 793 | 0.047121 | 0.039167 | 2.3726 | 29.9296 | 11.3604 | 1.4103
2019-03-19 15:55 | 794 | 0.056754 | 0.032606 | 1.967 | 24.7152 | 11.3109 | 1.4027
2019-03-19 15:55 | 795 | 0.04182 | 0.044782 | 2.6495 | 33.6477 | 11.2527 | 1.4071
2019-03-19 15:55 | 796 | 0.040178 | 0 | 0 | 0 | 11.2241 | 1.3368
2019-03-19 15:55 | 797 | 0.057892 | 0 | 0 | 0 | 11.1677 | 1.3803
2019-03-19 15:55 | 798 | 0.051908 | 0 | 0 | 0 | 11.1133 | 1.359
2019-03-19 15:55 | 799 | 0.060221 | 0 | 0 | 0 | 11.061 | 1.3344
2019-03-19 15:55 | 800 | 0.07026 | 0 | 0 | 0 | 10.9927 | 1.3656
2019-03-19 15:55 | 801 | 0.055242 | 0 | 0 | 0 | 10.9261 | 1.2983
2019-03-19 15:55 | 802 | 0.062335 | 0 | 0 | 0 | 10.8772 | 1.2933
2019-03-19 15:56 | 803 | 0.07683 | 0 | 0 | 0 | 10.7997 | 1.2086
2019-03-19 15:56 | 804 | 0.064758 | 0 | 0 | 0 | 10.7298 | 1.3266
2019-03-19 15:56 | 805 | 0.062685 | 0 | 0 | 0 | 10.6687 | 1.2918
2019-03-19 15:56 | 806 | 0.05139 | 0 | 0 | 0 | 10.6064 | 1.275
2019-03-19 15:56 | 807 | 0.042816 | 0 | 0 | 0 | 10.5647 | 1.2812
2019-03-19 15:56 | 808 | 0.060944 | 0 | 0 | 0 | 10.5146 | 1.2421
2019-03-19 15:56 | 809 | 0.056064 | 0 | 0 | 0 | 10.4483 | 1.2873
2019-03-19 15:56 | 810 | 0.034464 | 0 | 0 | 0 | 10.4065 | 1.1918
2019-03-19 15:56 | 811 | 0.043126 | 0 | 0 | 0 | 10.3725 | 1.2612
2019-03-19 15:56 | 812 | 0.043901 | 0 | 0 | 0 | 10.3217 | 1.2162
2019-03-19 15:56 | 813 | 0.030391 | 0 | 0 | 0 | 10.2878 | 1.2068
2019-03-19 15:56 | 814 | 0.040599 | 0 | 0 | 0 | 10.2554 | 1.2359
2019-03-19 15:56 | 815 | 0.049678 | 0 | 0 | 0 | 10.2065 | 1.21
2019-03-19 15:56 | 816 | 0.048701 | 0 | 0 | 0 | 10.1588 | 1.2527
2019-03-19 15:56 | 817 | 0.04066 | 0 | 0 | 0 | 10.1109 | 1.1706
2019-03-19 15:56 | 818 | 0.04101 | 0 | 0 | 0 | 10.0753 | 1.2587
2019-03-19 15:56 | 819 | 0.051161 | 0 | 0 | 0 | 10.0263 | 1.0971
2019-03-19 15:56 | 820 | 0.05615 | 0 | 0 | 0 | 9.9741 | 1.1864
2019-03-19 15:56 | 821 | 0.057529 | 0 | 0 | 0 | 9.9152 | 1.1826
2019-03-19 15:56 | 822 | 0.054633 | 0 | 0 | 0 | 9.8603 | 1.1771
2019-03-19 15:56 | 823 | 0.04955 | 0 | 0 | 0 | 9.8068 | 1.167
2019-03-19 15:56 | 824 | 0.041415 | 0 | 0 | 0 | 9.7619 | 1.139
2019-03-19 15:56 | 825 | 0.03876 | 0 | 0 | 0 | 9.7226 | 1.1506
2019-03-19 15:56 | 826 | 0.035212 | 0 | 0 | 0 | 9.684 | 1.1177
2019-03-19 15:56 | 827 | 0.047958 | 0 | 0 | 0 | 9.6486 | 1.1336
2019-03-19 15:56 | 828 | 0.042182 | 0 | 0 | 0 | 9.5921 | 1.1427
2019-03-19 15:56 | 829 | 0.044339 | 0 | 0 | 0 | 9.563 | 1.1339
2019-03-19 15:56 | 830 | 0.038693 | 0 | 0 | 0 | 9.5057 | 1.1236
2019-03-19 15:56 | 831 | 0.015217 | 0 | 0 | 0 | 9.4892 | 1.1016
2019-03-19 15:56 | 832 | 0.030167 | 0 | 0 | 0 | 9.4662 | 1.0656
2019-03-19 15:56 | 833 | 0.036558 | 0 | 0 | 0 | 9.43 | 1.0973
2019-03-19 15:56 | 834 | 0.040655 | 0 | 0 | 0 | 9.3942 | 1.0869
2019-03-19 15:56 | 835 | 0.03456 | 0 | 0 | 0 | 9.3514 | 1.0881
2019-03-19 15:56 | 836 | 0.030006 | 0 | 0 | 0 | 9.3248 | 1.0574
2019-03-19 15:56 | 837 | 0.030588 | 0 | 0 | 0 | 9.2898 | 1.0484
2019-03-19 15:56 | 838 | 0.038795 | 0 | 0 | 0 | 9.2619 | 1.0217
2019-03-19 15:56 | 839 | 0.034262 | 0 | 0 | 0 | 9.215 | 1.05
2019-03-19 15:56 | 840 | 0.0341 | 0 | 0 | 0 | 9.1919 | 1.0219
2019-03-19 15:56 | 841 | 0.063791 | 0 | 0 | 0 | 9.1408 | 1.048
2019-03-19 15:56 | 842 | 0.03521 | 0 | 0 | 0 | 9.0775 | 0.97697
2019-03-19 15:56 | 843 | 0.028876 | 0 | 0 | 0 | 9.0657 | 0.96371
2019-03-19 15:56 | 844 | 0.048642 | 0 | 0 | 0 | 9.0135 | 0.9444
2019-03-19 15:56 | 845 | 0.038607 | 0 | 0 | 0 | 8.9749 | 0.92957
2019-03-19 15:56 | 846 | 0.038028 | 0 | 0 | 0 | 8.9345 | 0.89844
2019-03-19 15:56 | 847 | 0.043191 | 0 | 0 | 0 | 8.8977 | 0.92467
2019-03-19 15:56 | 848 | 0.029856 | 0 | 0 | 0 | 8.8524 | 0.93111
2019-03-19 15:56 | 849 | 0.023354 | 0 | 0 | 0 | 8.8364 | 0.89638
2019-03-19 15:56 | 850 | 0.030526 | 0 | 0 | 0 | 8.8024 | 0.91543
2019-03-19 15:56 | 851 | 0.024374 | 0 | 0 | 0 | 8.7779 | 0.90595
2019-03-19 15:56 | 852 | 0.033824 | 0 | 0 | 0 | 8.7506 | 0.90368
2019-03-19 15:56 | 853 | 0.025768 | 0 | 0 | 0 | 8.7143 | 0.89498
2019-03-19 15:56 | 854 | 0.016335 | 0 | 0 | 0 | 8.6994 | 0.91862
2019-03-19 15:56 | 855 | 0.024285 | 0 | 0 | 0 | 8.6774 | 0.87575
2019-03-19 15:56 | 856 | 0.027626 | 0 | 0 | 0 | 8.6516 | 0.90128
2019-03-19 15:56 | 857 | 0.024977 | 0 | 0 | 0 | 8.6234 | 0.89998
2019-03-19 15:56 | 858 | 0.033999 | 0 | 0 | 0 | 8.5991 | 0.85262
2019-03-19 15:56 | 859 | 0.035298 | 0 | 0 | 0 | 8.5571 | 0.82139
2019-03-19 15:56 | 860 | 0.035702 | 0 | 0 | 0 | 8.5292 | 0.80234
2019-03-19 15:56 | 861 | 0.026685 | 0 | 0 | 0 | 8.4878 | 0.89487
2019-03-19 15:56 | 862 | 0.029912 | 0 | 0 | 0 | 8.4728 | 0.81641
2019-03-19 15:56 | 863 | 0.042009 | 0 | 0 | 0 | 8.4258 | 0.8276
2019-03-19 15:57 | 864 | 0.038264 | 0 | 0 | 0 | 8.3921 | 0.81975
2019-03-19 15:57 | 865 | 0.049274 | 0 | 0 | 0 | 8.3467 | 0.80405
2019-03-19 15:57 | 866 | 0.037976 | 0 | 0 | 0 | 8.299 | 0.80761
2019-03-19 15:57 | 867 | 0.023277 | 0 | 0 | 0 | 8.2719 | 0.78898
2019-03-19 15:57 | 868 | 0.020032 | 0 | 0 | 0 | 8.2494 | 0.76564
2019-03-19 15:57 | 869 | 0.026664 | 0 | 0 | 0 | 8.2291 | 0.78842
2019-03-19 15:57 | 870 | 0.025965 | 0 | 0 | 0 | 8.1977 | 0.7859
2019-03-19 15:57 | 871 | 0.018645 | 0 | 0 | 0 | 8.1786 | 0.82271
2019-03-19 15:57 | 872 | 0.027745 | 0 | 0 | 0 | 8.1564 | 0.73979
2019-03-19 15:57 | 873 | 0.038826 | 0 | 0 | 0 | 8.1228 | 0.77632
2019-03-19 15:57 | 874 | 0.029554 | 0 | 0 | 0 | 8.084 | 0.82232
2019-03-19 15:57 | 875 | 0.015163 | 0 | 0 | 0 | 8.0647 | 0.74679
2019-03-19 15:57 | 876 | 0.027858 | 0 | 0 | 0 | 8.0469 | 0.72879
2019-03-19 15:57 | 877 | 0.042364 | 0 | 0 | 0 | 8.0085 | 0.70435
2019-03-19 15:57 | 878 | 0.033601 | 0 | 0 | 0 | 7.9681 | 0.68775
2019-03-19 15:57 | 879 | 0.021838 | 0 | 0 | 0 | 7.9421 | 0.721
2019-03-19 15:57 | 880 | 0.030454 | 0 | 0 | 0 | 7.9194 | 0.74859
2019-03-19 15:57 | 881 | 0.037908 | 0 | 0 | 0 | 7.8816 | 0.68688
2019-03-19 15:57 | 882 | 0.019806 | 0 | 0 | 0 | 7.8497 | 0.69644
2019-03-19 15:57 | 883 | 0.015345 | 0 | 0 | 0 | 7.8392 | 0.67502
2019-03-19 15:57 | 884 | 0.017591 | 0 | 0 | 0 | 7.8173 | 0.72266
2019-03-19 15:57 | 885 | 0.0058732 | 0 | 0 | 0 | 7.8066 | 0.72571
2019-03-19 15:57 | 886 | 0.016104 | 0 | 0 | 0 | 7.8006 | 0.7045
2019-03-19 15:57 | 887 | 0.018706 | 0 | 0 | 0 | 7.7758 | 0.76755
2019-03-19 15:57 | 888 | 0.019811 | 0 | 0 | 0 | 7.7637 | 0.76181
2019-03-19 15:57 | 889 | 0.023678 | 0 | 0 | 0 | 7.7359 | 0.74091
2019-03-19 15:57 | 890 | 0.013102 | 0 | 0 | 0 | 7.7195 | 0.69261
2019-03-19 15:57 | 891 | 0.01614 | 0 | 0 | 0 | 7.7066 | 0.59545
2019-03-19 15:57 | 892 | 0.021796 | 0 | 0 | 0 | 7.6865 | 0.61201
2019-03-19 15:57 | 893 | 0.018295 | 0 | 0 | 0 | 7.6651 | 0.63582
2019-03-19 15:57 | 894 | 0.017548 | 0 | 0 | 0 | 7.649 | 0.58316
2019-03-19 15:57 | 895 | 0.034569 | 0 | 0 | 0 | 7.6261 | 0.63012
2019-03-19 15:57 | 896 | 0.03106 | 0 | 0 | 0 | 7.5848 | 0.6012
2019-03-19 15:57 | 897 | 0.016581 | 0 | 0 | 0 | 7.567 | 0.61618
2019-03-19 15:57 | 898 | 0.013347 | 0 | 0 | 0 | 7.5491 | 0.58618
2019-03-19 15:57 | 899 | 0.013768 | 0 | 0 | 0 | 7.5391 | 0.59565
2019-03-19 15:57 | 900 | 0.013179 | 0 | 0 | 0 | 7.5216 | 0.54465
2019-03-19 15:57 | 901 | 0.014265 | 0 | 0 | 0 | 7.5122 | 0.58524
2019-03-19 15:57 | 902 | 0.019533 | 0 | 0 | 0 | 7.4921 | 0.53213
2019-03-19 15:57 | 903 | 0.01816 | 0 | 0 | 0 | 7.4744 | 0.56406
2019-03-19 15:57 | 904 | 0.026808 | 0 | 0 | 0 | 7.4539 | 0.57023
2019-03-19 15:57 | 905 | 0.02109 | 0 | 0 | 0 | 7.4245 | 0.51162
2019-03-19 15:57 | 906 | 0.0026226 | 0 | 0 | 0 | 7.4144 | 0.49822
2019-03-19 15:57 | 907 | 0.017837 | 0 | 0 | 0 | 7.4112 | 0.59913
2019-03-19 15:57 | 908 | 0.030399 | 0 | 0 | 0 | 7.3788 | 0.45423
2019-03-19 15:57 | 909 | 0.027557 | 0 | 0 | 0 | 7.3541 | 0.50706
2019-03-19 15:57 | 910 | 0.033662 | 0 | 0 | 0 | 7.3225 | 0.49103
2019-03-19 15:57 | 911 | 0.019965 | 0 | 0 | 0 | 7.2915 | 0.59174
2019-03-19 15:57 | 912 | 0.0092051 | 0 | 0 | 0 | 7.2821 | 0.3968
2019-03-19 15:57 | 913 | 0.010154 | 0 | 0 | 0 | 7.2704 | 0.3723
2019-03-19 15:57 | 914 | -0.0028996 | 0 | 0 | 0 | 7.2643 | 0.4214
2019-03-19 15:57 | 915 | 0.003987 | 0 | 0 | 0 | 7.2709 | 0.45206
2019-03-19 15:57 | 916 | 0.034771 | 0 | 0 | 0 | 7.2512 | 0.48597
2019-03-19 15:57 | 917 | 0.016014 | 0 | 0 | 0 | 7.2131 | 0.42301
2019-03-19 15:57 | 918 | 0.00039792 | 0 | 0 | 0 | 7.2189 | 0.45637
2019-03-19 15:57 | 919 | 0.01243 | 0 | 0 | 0 | 7.2053 | 0.46699
2019-03-19 15:57 | 920 | 0.023283 | 0 | 0 | 0 | 7.1938 | 0.43885
2019-03-19 15:57 | 921 | 0.025403 | 0 | 0 | 0 | 7.1614 | 0.46857
2019-03-19 15:57 | 922 | 0.011294 | 0 | 0 | 0 | 7.1467 | 0.42958
2019-03-19 15:58 | 923 | 0.020208 | 0 | 0 | 0 | 7.1334 | 0.41665
2019-03-19 15:58 | 924 | 0.032076 | 0 | 0 | 0 | 7.1056 | 0.40998
2019-03-19 15:58 | 925 | 0.021349 | 0 | 0 | 0 | 7.0751 | 0.42912
2019-03-19 15:58 | 926 | -0.0013675 | 0 | 0 | 0 | 7.0654 | 0.40581
2019-03-19 15:58 | 927 | 0.02107 | 0 | 0 | 0 | 7.0673 | 0.41557

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062458 | 0.062687 | 0 | 0.2095 | 2019-03-19 11:25 | Group A | 0.1 | 0.22 | Do Nothing | 0.20947 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025375 | 0.025497 | 0 | 0.0261 | 2019-03-19 11:25 | Group A | -0.01 | 0.1 | Do Nothing | 2.141e-05 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002747 | 0.00276 | 0 | 0.0011 | 2019-03-19 11:25 | Group A | -0.01 | 0.1 | Do Nothing | 9.2638e-07 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.7814 | 185.0045 | 0 | 100 | 2019-03-19 11:25 | None | 0 | 1000 | Do Nothing | 100.6449 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.4711 | 211.1353 | 0 | 100 | 2019-03-19 11:25 | None | 0 | 1000 | Do Nothing | 100.7439 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -17.0283 | 1447 | 0 |  | 2019-03-19 11:14 | None | -10 | 1000 | Do Nothing | 128.2066 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.2612 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017858 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0016674 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -152.2556 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23636 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4028 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0033943 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0047411 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.004523 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 4.7652e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0011949 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 75.0311 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3419 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.548 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 101.21 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.7938 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 17.6124 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 64.798 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 103.5288 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 61.947 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0042781 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0536 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.025737 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 24.4285 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN