## Cone Summary
Date: | 2019-03-21
Time: | 11:36
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | B6S-Test2
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0411 |  | Start Test | 0 | 2019-03-21 11:45
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 7 | 2019-03-21 11:45
Sample Area | 1 | m^2 | Wetting | 221 | 2019-03-21 11:49
Cone Heat Flux | 75 | kW/m^2 | Dripping | 314 | 2019-03-21 11:50
Cone Orientation |  |  | Bottom Ignition | 690 | 2019-03-21 11:57
Sample Grid |  |  | Flame Out | 1011 | 2019-03-21 12:02
Sample Thickness | 106 | mm
Sample Initial Mass | 52.9 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 23.8311 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 4.6589 | kW/m2 | Average Heat Release Rate | 0.85365 | kW/m2
Total Heat Release | 0.98454 | MJ/m2 | Average Mass Loss Rate | 0.04252 | g/m2s
Peak Mass Loss Rate | 0.19849 | g/m2s | Effective Heat of Combustion | 20.1094 | MJ/kg
Total Mass Lost | 45.8993 | g | Average Specific Extinction Area | 57.6153 | m2/kg
Time to Ignition | 7 | s | Peak CO Yield | 10.3219 | kg/kg
Time to Peak Heat Release | 723 | s | Peak CO2 Yield | 154.4699 | kg/kg
Time to Flame Out | 1011 | s | Average CO Yield | 0.06064 | kg/kg
Average Heat Release Rate | 0.97964 | kW/m2 | Average CO2 Yield | 2.3932 | kg/kg
Average Mass Loss Rate | 0.045195 | g/m2s | New Calibration Factor | 24.4106
Effective Heat of Combustion | 21.45 | MJ/kg
Dripping | 314 | s | Time to dripping
HRRD | 0.46772 | kW | HRR at dripping time
THRD | 0.32757 | MJ | Heat released from Ignition to Dripping
AHRRD | 1.067 | kW | AHRR from Ignition to Dripping
Wetting | 221 | s | Time to wetting
HRRW | 0.49595 | kW | HRR at wetting time
THRW | 0.28284 | MJ | Heat released from Ignition to Wetting
AHRRW | 1.2855 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-21 11:36 | 1553182581.3532 | -532 | 0.21043 | 0.000385 | -2e-06 | 100.7519 | 101.1554 | 124.4383 | 20.9346 | 0.00187 | 0.001679 | -151.6876 | 0.23703 | 8.4025 | 0.00342 | 0.005474 | 0.005369 | 6.4e-05 | 0.000973 | 75.1456 | 1.305 | 2.5478 | 99.63 | 21 | 37.152 | 66.4946 | 103.8355 | 61.9093 | 0.009285 | 1.0537 | 0.025047 | 23.7711 | -0.035746 | 0.14101 | 0.0001 | 0.007665 | 0.007665 | 0.75615
2019-03-21 11:36 | 1553182582.3533 | -531 | 0.21043 | 0.000385 | -2e-06 | 100.7519 | 101.1554 | 124.4383 | 20.9346 | 0.00187 | 0.001679 | -151.6876 | 0.23703 | 8.4025 | 0.00342 | 0.005474 | 0.005369 | 6.4e-05 | 0.000973 | 75.1456 | 1.305 | 2.5478 | 99.63 | 21 | 37.153 | 66.4946 | 103.8355 | 61.9093 | 0.009285 | 1.0537 | 0.025047 | 23.7711 | -0.035746 | 0.14101 | 0.0001 | 0.007664 | 0.007664 | 0.75619
2019-03-21 11:36 | 1553182583.3543 | -530 | 0.20937 | 0.000385 | -2e-06 | 100.8369 | 101.0829 | 124.4383 | 20.935 | 0.001844 | 0.001742 | -151.6876 | 0.23729 | 8.4063 | 0.003442 | 0.004906 | 0.004783 | 6.2e-05 | 0.000964 | 75.1541 | 1.297 | 2.5479 | 99.63 | 21 | 37.156 | 65.874 | 104.3746 | 63.4142 | 0.009286 | 1.049 | 0.024991 | 23.8244 | -0.021794 | 0.14116 | 0.0001 | 0.007611 | 0.007611 | 0.76232
2019-03-21 11:36 | 1553182584.3533 | -529 | 0.20938 | 0.000385 | -2e-06 | 100.7001 | 101.0038 | 124.4383 | 20.9355 | 0.001878 | 0.001747 | -151.6876 | 0.23671 | 8.4051 | 0.003471 | 0.00559 | 0.005354 | 6.6e-05 | 0.000968 | 75.1777 | 1.3125 | 2.5479 | 99.63 | 21 | 37.158 | 66.697 | 105.078 | 63.5574 | 0.009286 | 1.0485 | 0.024986 | 23.8295 | -0.026938 | 0.14082 | 0.0001 | 0.00767 | 0.00767 | 0.75457
2019-03-21 11:36 | 1553182585.3545 | -528 | 0.2094 | 0.000387 | -2e-06 | 100.8023 | 101.0376 | 124.4383 | 20.9352 | 0.001884 | 0.001744 | -151.6876 | 0.23753 | 8.4026 | 0.003503 | 0.004951 | 0.004913 | 6.4e-05 | 0.000963 | 75.1916 | 1.3245 | 2.548 | 99.63 | 21 | 37.16 | 66.8525 | 105.8323 | 63.4659 | 0.009287 | 1.0488 | 0.024989 | 23.8263 | -0.020851 | 0.14131 | 9.7e-05 | 0.007382 | 0.007382 | 0.78678
2019-03-21 11:36 | 1553182586.3539 | -527 | 0.20938 | 0.000389 | -2e-06 | 100.8232 | 101.1135 | 124.4383 | 20.9364 | 0.001904 | 0.001765 | -151.6876 | 0.23771 | 8.4028 | 0.003415 | 0.004652 | 0.004426 | 6.6e-05 | 0.000969 | 75.1629 | 1.3665 | 2.5478 | 99.63 | 21 | 37.163 | 67.3351 | 103.7234 | 63.9837 | 0.009287 | 1.0472 | 0.02497 | 23.8446 | -0.025722 | 0.14141 | 9.5e-05 | 0.00723 | 0.00723 | 0.80387
2019-03-21 11:36 | 1553182587.3546 | -526 | 0.20939 | 0.000394 | -2e-06 | 100.8571 | 101.1179 | 124.4383 | 20.9371 | 0.001942 | 0.001751 | -151.6876 | 0.23758 | 8.4001 | 0.003433 | 0.004089 | 0.003913 | 6.5e-05 | 0.000973 | 75.2124 | 1.368 | 2.548 | 99.63 | 21 | 37.164 | 68.2432 | 104.143 | 63.6394 | 0.009288 | 1.0483 | 0.024983 | 23.8324 | -0.0231 | 0.14134 | 8.8e-05 | 0.00676 | 0.00676 | 0.85935
2019-03-21 11:36 | 1553182588.3552 | -525 | 0.20943 | 0.000396 | -2e-06 | 100.8017 | 100.9156 | 124.4383 | 20.9353 | 0.001897 | 0.001738 | -151.6876 | 0.23698 | 8.4033 | 0.003353 | 0.004774 | 0.004588 | 6.5e-05 | 0.000973 | 75.284 | 1.308 | 2.5479 | 99.63 | 21 | 37.167 | 67.1648 | 102.2066 | 63.323 | 0.009288 | 1.0493 | 0.024994 | 23.8212 | -0.010103 | 0.14098 | 8.6e-05 | 0.006598 | 0.006598 | 0.87824
2019-03-21 11:36 | 1553182589.3549 | -524 | 0.20943 | 0.000398 | -2e-06 | 100.8994 | 101.0641 | 124.4383 | 20.937 | 0.001847 | 0.001742 | -151.6876 | 0.23591 | 8.4022 | 0.003344 | 0.003981 | 0.003785 | 6e-05 | 0.000977 | 75.2195 | 1.309 | 2.5479 | 99.63 | 21 | 37.169 | 65.9444 | 102.0041 | 63.4161 | 0.009289 | 1.049 | 0.024991 | 23.8245 | -0.014588 | 0.14034 | 8.3e-05 | 0.006384 | 0.006384 | 0.90348
2019-03-21 11:36 | 1553182590.3539 | -523 | 0.20943 | 0.000408 | -2e-06 | 100.8962 | 101.0798 | 124.4383 | 20.9368 | 0.00195 | 0.001763 | -151.6876 | 0.2365 | 8.3998 | 0.003385 | 0.004906 | 0.004704 | 6.4e-05 | 0.000975 | 75.1645 | 1.2973 | 2.5479 | 99.63 | 21 | 37.17 | 68.4409 | 102.9763 | 63.9225 | 0.009289 | 1.0474 | 0.024972 | 23.8424 | -0.016265 | 0.1407 | 7.1e-05 | 0.00546 | 0.00546 | 1.0592
2019-03-21 11:36 | 1553182591.3537 | -522 | 0.20943 | 0.000416 | -2e-06 | 100.8681 | 100.8349 | 124.4383 | 20.937 | 0.001849 | 0.001717 | -151.6876 | 0.23678 | 8.402 | 0.003331 | 0.004632 | 0.004459 | 6.3e-05 | 0.000978 | 75.1936 | 1.2971 | 2.5479 | 99.63 | 21 | 37.176 | 66.0133 | 101.6694 | 62.8288 | 0.009291 | 1.0508 | 0.025013 | 23.8037 | 0.002948 | 0.14086 | 6.1e-05 | 0.004757 | 0.004757 | 1.2169
2019-03-21 11:37 | 1553182592.3538 | -521 | 0.20943 | 0.000413 | -2e-06 | 100.8729 | 101.0948 | 124.7944 | 20.938 | 0.001923 | 0.001714 | -151.6772 | 0.23626 | 8.4027 | 0.00333 | 0.0047 | 0.004573 | 6.4e-05 | 0.000976 | 75.1765 | 1.28 | 2.5478 | 99.63 | 21 | 37.177 | 67.7853 | 101.6595 | 62.7594 | 0.009291 | 1.051 | 0.025051 | 23.8353 | -0.019651 | 0.14055 | 6.5e-05 | 0.005005 | 0.005005 | 1.1541
2019-03-21 11:37 | 1553182593.3543 | -520 | 0.20943 | 0.000417 | -2e-06 | 100.8571 | 101.0856 | 124.7573 | 20.9371 | 0.001909 | 0.001737 | -151.6851 | 0.23712 | 8.4027 | 0.003378 | 0.004487 | 0.004288 | 6.5e-05 | 0.000975 | 75.1709 | 1.2905 | 2.548 | 99.63 | 21 | 37.183 | 67.4481 | 102.8274 | 63.3005 | 0.009292 | 1.0493 | 0.025027 | 23.8509 | -0.020238 | 0.14106 | 6e-05 | 0.004615 | 0.004615 | 1.2563
2019-03-21 11:37 | 1553182594.354 | -519 | 0.20943 | 0.000416 | -3e-06 | 100.8425 | 100.9194 | 124.5045 | 20.9378 | 0.001848 | 0.001759 | -151.6898 | 0.23666 | 8.4025 | 0.003375 | 0.004377 | 0.004241 | 6.3e-05 | 0.000979 | 75.162 | 1.3122 | 2.548 | 99.63 | 21 | 37.186 | 65.9734 | 102.7448 | 63.8386 | 0.009293 | 1.0476 | 0.024982 | 23.8458 | -0.006814 | 0.14079 | 6.2e-05 | 0.004788 | 0.004788 | 1.2086
2019-03-21 11:37 | 1553182595.355 | -518 | 0.20937 | 0.000423 | -2e-06 | 100.8658 | 100.8659 | 124.5078 | 20.9382 | 0.001898 | 0.001753 | -151.6866 | 0.23686 | 8.402 | 0.003357 | 0.004799 | 0.004553 | 6.5e-05 | 0.000978 | 75.1855 | 1.3107 | 2.548 | 99.63 | 21 | 37.19 | 67.1933 | 102.306 | 63.6984 | 0.009294 | 1.0481 | 0.024988 | 23.8411 | -1.2e-05 | 0.14091 | 5.3e-05 | 0.004105 | 0.004105 | 1.411
2019-03-21 11:37 | 1553182596.3534 | -517 | 0.2094 | 0.000428 | -2e-06 | 100.8866 | 101.0253 | 124.9569 | 20.9391 | 0.001881 | 0.001779 | -151.7081 | 0.23604 | 8.4031 | 0.003415 | 0.005005 | 0.004777 | 6.5e-05 | 0.000982 | 75.1526 | 1.3729 | 2.548 | 99.63 | 21 | 37.197 | 66.7635 | 103.726 | 64.3098 | 0.009296 | 1.0462 | 0.02501 | 23.9057 | -0.012287 | 0.14042 | 4.6e-05 | 0.00358 | 0.00358 | 1.612
2019-03-21 11:37 | 1553182597.3531 | -516 | 0.20936 | 0.000426 | -2e-06 | 100.8699 | 100.9647 | 125.1721 | 20.937 | 0.001852 | 0.001781 | -151.7083 | 0.2367 | 8.4027 | 0.00344 | 0.004416 | 0.00424 | 6.7e-05 | 0.000984 | 75.127 | 1.3293 | 2.5483 | 99.63 | 21 | 37.201 | 66.0678 | 104.3285 | 64.3781 | 0.009297 | 1.046 | 0.025029 | 23.9287 | -0.008397 | 0.14081 | 4.9e-05 | 0.00384 | 0.00384 | 1.5073
2019-03-21 11:37 | 1553182598.3531 | -515 | 0.20942 | 0.000431 | -2e-06 | 100.8521 | 101.0089 | 125.1712 | 20.9379 | 0.001874 | 0.0018 | -151.7115 | 0.23794 | 8.4045 | 0.003497 | 0.005136 | 0.004969 | 6.9e-05 | 0.00098 | 75.0995 | 1.3054 | 2.548 | 99.63 | 21 | 37.204 | 66.6016 | 105.7088 | 64.8207 | 0.009298 | 1.0446 | 0.025012 | 23.9443 | -0.013904 | 0.14155 | 4.1e-05 | 0.003193 | 0.003193 | 1.8222
2019-03-21 11:37 | 1553182599.3536 | -514 | 0.20937 | 0.000432 | -2e-06 | 100.8824 | 100.9851 | 124.8555 | 20.9381 | 0.001934 | 0.001748 | -151.6978 | 0.23721 | 8.4019 | 0.003479 | 0.004802 | 0.004671 | 6.6e-05 | 0.000987 | 75.162 | 1.3274 | 2.5482 | 99.63 | 21 | 37.212 | 68.0374 | 105.2772 | 63.5669 | 0.0093 | 1.0485 | 0.025027 | 23.8697 | -0.009097 | 0.14111 | 4.1e-05 | 0.003196 | 0.003196 | 1.8146
2019-03-21 11:37 | 1553182600.3533 | -513 | 0.20942 | 0.000434 | -2e-06 | 100.8378 | 100.9943 | 125.1263 | 20.9393 | 0.001879 | 0.001809 | -151.7098 | 0.23662 | 8.3978 | 0.003545 | 0.005092 | 0.004926 | 6.3e-05 | 0.00098 | 75.1803 | 1.3027 | 2.5481 | 99.63 | 21 | 37.215 | 66.7358 | 106.8579 | 65.0345 | 0.0093 | 1.0439 | 0.025 | 23.9476 | -0.013875 | 0.14077 | 3.8e-05 | 0.00298 | 0.00298 | 1.9412
2019-03-21 11:37 | 1553182601.3534 | -512 | 0.20946 | 0.000436 | -2e-06 | 100.8932 | 101.0382 | 125.0397 | 20.9384 | 0.001872 | 0.001798 | -151.71 | 0.23694 | 8.4009 | 0.003549 | 0.005048 | 0.00479 | 6.4e-05 | 0.000982 | 75.1192 | 1.3166 | 2.5482 | 99.63 | 21 | 37.221 | 66.5615 | 106.9672 | 64.7822 | 0.009302 | 1.0447 | 0.025001 | 23.9304 | -0.012851 | 0.14095 | 3.5e-05 | 0.002775 | 0.002775 | 2.0874
2019-03-21 11:37 | 1553182602.3538 | -511 | 0.20943 | 0.000439 | -2e-06 | 100.838 | 100.9684 | 124.6928 | 20.9387 | 0.001867 | 0.001737 | -151.701 | 0.2378 | 8.4044 | 0.003508 | 0.004439 | 0.004277 | 6.6e-05 | 0.000984 | 75.0827 | 1.3369 | 2.5481 | 99.63 | 21 | 37.225 | 66.43 | 105.9648 | 63.3044 | 0.009303 | 1.0493 | 0.025021 | 23.8449 | -0.011556 | 0.14147 | 3.2e-05 | 0.00256 | 0.00256 | 2.2712
2019-03-21 11:37 | 1553182603.3534 | -510 | 0.20941 | 0.000438 | -2e-06 | 100.8071 | 101.1608 | 125.0198 | 20.9391 | 0.001901 | 0.001725 | -151.6942 | 0.23735 | 8.4026 | 0.003593 | 0.004876 | 0.004737 | 6.5e-05 | 0.000989 | 75.1776 | 1.3038 | 2.548 | 99.63 | 21 | 37.227 | 67.2442 | 108.036 | 63.0267 | 0.009303 | 1.0502 | 0.025064 | 23.8663 | -0.031336 | 0.1412 | 3.3e-05 | 0.002632 | 0.002632 | 2.2045
2019-03-21 11:37 | 1553182604.3545 | -509 | 0.20943 | 0.000442 | -2e-06 | 100.8905 | 100.9871 | 125.1545 | 20.94 | 0.001909 | 0.001769 | -151.688 | 0.23731 | 8.4028 | 0.00357 | 0.004966 | 0.004661 | 6.5e-05 | 0.000987 | 75.1591 | 1.3333 | 2.548 | 99.63 | 21 | 37.232 | 67.4525 | 107.4767 | 64.0724 | 0.009305 | 1.0469 | 0.025038 | 23.9162 | -0.008565 | 0.14117 | 2.8e-05 | 0.002265 | 0.002265 | 2.5616
2019-03-21 11:37 | 1553182605.3545 | -508 | 0.20945 | 0.000443 | -2e-06 | 100.8532 | 101.0175 | 125.1024 | 20.9405 | 0.001859 | 0.001783 | -151.686 | 0.23694 | 8.4044 | 0.003467 | 0.00524 | 0.005095 | 6.5e-05 | 0.00099 | 75.1747 | 1.2888 | 2.548 | 99.63 | 21 | 37.232 | 66.2405 | 104.9743 | 64.4219 | 0.009305 | 1.0458 | 0.02502 | 23.9236 | -0.014552 | 0.14095 | 2.6e-05 | 0.002106 | 0.002106 | 2.7511
2019-03-21 11:37 | 1553182606.354 | -507 | 0.20942 | 0.000444 | -2e-06 | 100.8285 | 100.9262 | 124.6815 | 20.9408 | 0.001856 | 0.001725 | -151.6802 | 0.2369 | 8.4045 | 0.003422 | 0.004662 | 0.004559 | 6.5e-05 | 0.00099 | 75.1508 | 1.2895 | 2.548 | 99.63 | 21 | 37.234 | 66.1661 | 103.8826 | 63.0292 | 0.009305 | 1.0502 | 0.02503 | 23.834 | -0.00866 | 0.14093 | 2.6e-05 | 0.002101 | 0.002101 | 2.7568
2019-03-21 11:37 | 1553182607.354 | -506 | 0.20948 | 0.000441 | -3e-06 | 100.8431 | 101.1064 | 124.6273 | 20.9399 | 0.001881 | 0.001731 | -151.6723 | 0.23793 | 8.4008 | 0.003531 | 0.004649 | 0.004398 | 6.3e-05 | 0.000989 | 75.1749 | 1.3202 | 2.5476 | 99.63 | 21 | 37.233 | 66.7744 | 106.5225 | 63.1751 | 0.009305 | 1.0497 | 0.025019 | 23.834 | -0.023318 | 0.14155 | 2.9e-05 | 0.002353 | 0.002353 | 2.4721
2019-03-21 11:37 | 1553182608.3547 | -505 | 0.20946 | 0.000452 | -3e-06 | 100.8303 | 101.1428 | 124.4583 | 20.9408 | 0.001875 | 0.00172 | -151.6811 | 0.23671 | 8.405 | 0.003541 | 0.004743 | 0.004585 | 6.7e-05 | 0.000985 | 75.1205 | 1.372 | 2.5479 | 99.63 | 21 | 37.232 | 66.6347 | 106.7744 | 62.8947 | 0.009305 | 1.0506 | 0.025012 | 23.8079 | -0.027678 | 0.14082 | 1.6e-05 | 0.001351 | 0.001351 | 4.2851
2019-03-21 11:37 | 1553182609.3541 | -504 | 0.20945 | 0.000449 | -3e-06 | 100.8592 | 101.1343 | 124.8247 | 20.9407 | 0.001819 | 0.001757 | -151.6867 | 0.23701 | 8.4014 | 0.003546 | 0.004326 | 0.004142 | 6.8e-05 | 0.000989 | 75.205 | 1.3627 | 2.5479 | 99.63 | 21 | 37.228 | 65.2776 | 106.8845 | 63.8045 | 0.009304 | 1.0478 | 0.025015 | 23.8752 | -0.024361 | 0.141 | 2e-05 | 0.001651 | 0.001651 | 3.5093
2019-03-21 11:37 | 1553182610.3547 | -503 | 0.20944 | 0.000445 | -2e-06 | 100.836 | 101.0045 | 124.5689 | 20.9391 | 0.001873 | 0.001736 | -151.6788 | 0.23719 | 8.4042 | 0.003492 | 0.004657 | 0.004476 | 6.7e-05 | 0.000989 | 75.1166 | 1.3459 | 2.548 | 99.63 | 21 | 37.226 | 66.5935 | 105.582 | 63.2916 | 0.009303 | 1.0494 | 0.025009 | 23.8326 | -0.014939 | 0.1411 | 2.4e-05 | 0.001934 | 0.001934 | 2.9986
2019-03-21 11:37 | 1553182611.3548 | -502 | 0.20947 | 0.000446 | -2e-06 | 100.7943 | 101.0185 | 124.6611 | 20.9398 | 0.001906 | 0.00169 | -151.6582 | 0.23684 | 8.4015 | 0.003465 | 0.005209 | 0.005057 | 6.4e-05 | 0.000991 | 75.1227 | 1.3405 | 2.5481 | 99.63 | 21 | 37.224 | 67.3672 | 104.9241 | 62.1884 | 0.009303 | 1.0528 | 0.025059 | 23.8023 | -0.019876 | 0.1409 | 2.3e-05 | 0.001831 | 0.001831 | 3.1625
2019-03-21 11:37 | 1553182612.354 | -501 | 0.20945 | 0.000451 | -2e-06 | 100.8208 | 101.1976 | 124.6679 | 20.9403 | 0.001848 | 0.001751 | -151.6709 | 0.23725 | 8.4029 | 0.003397 | 0.004699 | 0.004506 | 6.7e-05 | 0.000991 | 75.1959 | 1.3741 | 2.5479 | 99.63 | 21 | 37.22 | 65.9765 | 103.2854 | 63.6386 | 0.009302 | 1.0483 | 0.025006 | 23.8543 | -0.033372 | 0.14114 | 1.6e-05 | 0.001362 | 0.001362 | 4.2576
2019-03-21 11:37 | 1553182613.3533 | -500 | 0.20948 | 0.000453 | -2e-06 | 100.796 | 101.0376 | 124.5848 | 20.9405 | 0.001865 | 0.001758 | -151.6938 | 0.23709 | 8.4017 | 0.003404 | 0.005331 | 0.005099 | 6.4e-05 | 0.000989 | 75.1019 | 1.3238 | 2.5479 | 99.63 | 21 | 37.218 | 66.3997 | 103.4466 | 63.8051 | 0.009301 | 1.0478 | 0.024991 | 23.8523 | -0.021406 | 0.14105 | 1.4e-05 | 0.001198 | 0.001198 | 4.8394
2019-03-21 11:37 | 1553182614.354 | -499 | 0.20945 | 0.000453 | -2e-06 | 100.906 | 100.9817 | 124.3942 | 20.9401 | 0.001876 | 0.001731 | -151.6842 | 0.23721 | 8.4038 | 0.003417 | 0.005025 | 0.004946 | 6.3e-05 | 0.000996 | 75.0956 | 1.3571 | 2.548 | 99.63 | 21 | 37.213 | 66.6577 | 103.7718 | 63.1659 | 0.0093 | 1.0497 | 0.024996 | 23.8114 | -0.006709 | 0.14112 | 1.5e-05 | 0.001261 | 0.001261 | 4.5985
2019-03-21 11:37 | 1553182615.3541 | -498 | 0.20946 | 0.000449 | -2e-06 | 100.8166 | 101.123 | 124.7048 | 20.9403 | 0.001915 | 0.00173 | -151.6768 | 0.2379 | 8.4013 | 0.00341 | 0.004766 | 0.004499 | 6.5e-05 | 0.000998 | 75.1067 | 1.3111 | 2.548 | 99.63 | 21 | 37.21 | 67.5915 | 103.5847 | 63.1333 | 0.009299 | 1.0498 | 0.025028 | 23.84 | -0.027147 | 0.14153 | 1.9e-05 | 0.00157 | 0.00157 | 3.7052
2019-03-21 11:37 | 1553182616.3541 | -497 | 0.20946 | 0.000453 | -2e-06 | 100.8711 | 101.0512 | 124.8266 | 20.9412 | 0.001875 | 0.001749 | -151.6728 | 0.23723 | 8.4049 | 0.003487 | 0.004757 | 0.00461 | 6.5e-05 | 0.000997 | 75.1029 | 1.2915 | 2.5478 | 99.63 | 21 | 37.205 | 66.6323 | 105.4589 | 63.5959 | 0.009298 | 1.0484 | 0.025023 | 23.868 | -0.015952 | 0.14113 | 1.3e-05 | 0.001121 | 0.001121 | 5.1756
2019-03-21 11:37 | 1553182617.3542 | -496 | 0.20945 | 0.000455 | -2e-06 | 100.8038 | 100.8996 | 124.5843 | 20.9411 | 0.00187 | 0.001718 | -151.6697 | 0.23731 | 8.4021 | 0.003511 | 0.004646 | 0.004532 | 6.2e-05 | 0.000996 | 75.1028 | 1.3476 | 2.5478 | 99.63 | 21 | 37.202 | 66.5088 | 106.0481 | 62.846 | 0.009297 | 1.0507 | 0.025027 | 23.8183 | -0.008499 | 0.14117 | 1.1e-05 | 0.001021 | 0.001021 | 5.6842
2019-03-21 11:37 | 1553182618.3543 | -495 | 0.20946 | 0.000461 | -2e-06 | 100.8297 | 101.0307 | 124.624 | 20.9428 | 0.001886 | 0.001757 | -151.6472 | 0.23767 | 8.401 | 0.003493 | 0.004516 | 0.004329 | 6.4e-05 | 0.001002 | 75.1569 | 1.3436 | 2.5481 | 99.63 | 21 | 37.199 | 66.9092 | 105.6162 | 63.7934 | 0.009296 | 1.0478 | 0.024996 | 23.8556 | -0.017816 | 0.14139 | 4e-06 | 0.000454 | 0.000454 | 12.8096
2019-03-21 11:37 | 1553182619.3533 | -494 | 0.20945 | 0.000454 | -2e-06 | 100.8678 | 101.045 | 124.5156 | 20.9412 | 0.001921 | 0.00175 | -151.6632 | 0.23733 | 8.4031 | 0.00349 | 0.004381 | 0.004244 | 6e-05 | 0.001001 | 75.1104 | 1.3528 | 2.5479 | 99.63 | 21 | 37.194 | 67.7321 | 105.535 | 63.6137 | 0.009295 | 1.0483 | 0.024991 | 23.8389 | -0.0157 | 0.14119 | 1.3e-05 | 0.001077 | 0.001077 | 5.3855
2019-03-21 11:37 | 1553182620.3548 | -493 | 0.20949 | 0.000458 | -2e-06 | 100.833 | 101.0641 | 124.8522 | 20.9413 | 0.001922 | 0.001719 | -151.6587 | 0.23651 | 8.4012 | 0.003564 | 0.00484 | 0.004626 | 6.3e-05 | 0.001001 | 75.1525 | 1.3325 | 2.548 | 99.63 | 21 | 37.193 | 67.7729 | 107.3293 | 62.8744 | 0.009295 | 1.0507 | 0.025053 | 23.8449 | -0.02048 | 0.1407 | 8e-06 | 0.000769 | 0.000769 | 7.5227
2019-03-21 11:37 | 1553182621.3544 | -492 | 0.20946 | 0.000454 | -3e-06 | 100.8768 | 100.8465 | 124.5642 | 20.9424 | 0.001856 | 0.001745 | -151.6541 | 0.23604 | 8.4048 | 0.003537 | 0.004605 | 0.00442 | 6.1e-05 | 0.001002 | 75.1743 | 1.329 | 2.5479 | 99.63 | 21 | 37.191 | 66.1684 | 106.6814 | 63.5064 | 0.009294 | 1.0487 | 0.025 | 23.8397 | 0.002687 | 0.14042 | 1.3e-05 | 0.001155 | 0.001155 | 4.9961
2019-03-21 11:37 | 1553182622.3544 | -491 | 0.20945 | 0.000458 | -2e-06 | 100.8479 | 101.1285 | 124.6203 | 20.9432 | 0.001866 | 0.001726 | -151.6501 | 0.23647 | 8.4003 | 0.003529 | 0.005291 | 0.005129 | 6.4e-05 | 0.001003 | 75.1248 | 1.3307 | 2.5479 | 99.63 | 21 | 37.188 | 66.418 | 106.4839 | 63.0384 | 0.009294 | 1.0501 | 0.025023 | 23.8285 | -0.024854 | 0.14067 | 8e-06 | 0.000733 | 0.000733 | 7.8853
2019-03-21 11:37 | 1553182623.3539 | -490 | 0.20946 | 0.000457 | -2e-06 | 100.8208 | 101.0863 | 124.7205 | 20.9433 | 0.001903 | 0.001757 | -151.6474 | 0.23667 | 8.402 | 0.003518 | 0.004113 | 0.004019 | 6.2e-05 | 0.001001 | 75.1577 | 1.3407 | 2.5479 | 99.63 | 20.99 | 37.185 | 67.2959 | 106.2249 | 63.7922 | 0.009287 | 1.0478 | 0.025005 | 23.8648 | -0.023522 | 0.1408 | 9e-06 | 0.00084 | 0.00084 | 6.8919
2019-03-21 11:37 | 1553182624.3535 | -489 | 0.20948 | 0.000457 | -3e-06 | 100.8613 | 101.0389 | 125.0235 | 20.9429 | 0.001878 | 0.00171 | -151.6572 | 0.23715 | 8.4018 | 0.003492 | 0.004997 | 0.0048 | 6.4e-05 | 0.001007 | 75.1959 | 1.2997 | 2.5479 | 99.63 | 21 | 37.182 | 66.7071 | 105.5928 | 62.6744 | 0.009292 | 1.0513 | 0.025077 | 23.8541 | -0.015738 | 0.14108 | 1e-05 | 0.000907 | 0.000907 | 6.3946
2019-03-21 11:37 | 1553182625.3536 | -488 | 0.20944 | 0.00046 | -2e-06 | 100.8276 | 100.9521 | 124.7127 | 20.9435 | 0.001876 | 0.001787 | -151.6606 | 0.23587 | 8.4016 | 0.003486 | 0.004652 | 0.004465 | 6.1e-05 | 0.001007 | 75.1772 | 1.3057 | 2.5479 | 99.63 | 20.99 | 37.18 | 66.6509 | 105.4335 | 64.5256 | 0.009286 | 1.0455 | 0.024977 | 23.89 | -0.011033 | 0.14032 | 5e-06 | 0.000534 | 0.000534 | 10.7982
2019-03-21 11:37 | 1553182626.3541 | -487 | 0.20947 | 0.000458 | -3e-06 | 100.8562 | 101.1012 | 124.5175 | 20.9436 | 0.00187 | 0.001728 | -151.6592 | 0.23677 | 8.4023 | 0.003541 | 0.0048 | 0.004658 | 6.1e-05 | 0.001003 | 75.1943 | 1.33 | 2.5479 | 99.63 | 20.99 | 37.18 | 66.5252 | 106.774 | 63.1014 | 0.009286 | 1.0499 | 0.025011 | 23.8209 | -0.021703 | 0.14086 | 8e-06 | 0.000792 | 0.000792 | 7.3121
2019-03-21 11:37 | 1553182627.3537 | -486 | 0.20952 | 0.000464 | -2e-06 | 100.9096 | 101.3601 | 124.7146 | 20.9442 | 0.001937 | 0.001754 | -151.6673 | 0.23649 | 8.4023 | 0.003483 | 0.00508 | 0.004877 | 5.8e-05 | 0.001003 | 75.2362 | 1.3235 | 2.548 | 99.63 | 20.99 | 37.18 | 68.1204 | 105.3704 | 63.7174 | 0.009286 | 1.048 | 0.025008 | 23.8616 | -0.039846 | 0.14069 | -1e-06 | 9.1e-05 | 9.1e-05 | 63.6416
2019-03-21 11:37 | 1553182628.3537 | -485 | 0.20947 | 0.00046 | -3e-06 | 100.8583 | 101.0307 | 124.9888 | 20.9442 | 0.00184 | 0.001826 | -151.688 | 0.23677 | 8.4028 | 0.003433 | 0.004901 | 0.004745 | 6.1e-05 | 0.001003 | 75.2325 | 1.2631 | 2.548 | 99.63 | 20.99 | 37.18 | 65.7932 | 104.1507 | 65.4591 | 0.009286 | 1.0426 | 0.024971 | 23.9495 | -0.015279 | 0.14086 | 6e-06 | 0.000599 | 0.000599 | 9.6676
2019-03-21 11:37 | 1553182629.3532 | -484 | 0.20947 | 0.000463 | -3e-06 | 100.8169 | 101.1176 | 125.0806 | 20.9443 | 0.001903 | 0.00169 | -151.6846 | 0.23649 | 8.4032 | 0.003487 | 0.004927 | 0.004728 | 5.9e-05 | 0.001 | 75.2194 | 1.3085 | 2.548 | 99.63 | 20.99 | 37.181 | 67.3178 | 105.4658 | 62.1858 | 0.009286 | 1.0528 | 0.025101 | 23.8422 | -0.026638 | 0.14069 | 2e-06 | 0.000344 | 0.000344 | 16.8042
2019-03-21 11:37 | 1553182630.3529 | -483 | 0.20947 | 0.000458 | -2e-06 | 100.8747 | 101.1459 | 124.8951 | 20.9451 | 0.001898 | 0.00176 | -151.7021 | 0.23615 | 8.4044 | 0.003508 | 0.00469 | 0.004537 | 5.9e-05 | 0.001003 | 75.1738 | 1.2867 | 2.5478 | 99.63 | 20.99 | 37.183 | 67.1985 | 105.9836 | 63.8651 | 0.009287 | 1.0476 | 0.02502 | 23.8841 | -0.024012 | 0.14048 | 8e-06 | 0.000787 | 0.000787 | 7.336
2019-03-21 11:37 | 1553182631.3529 | -482 | 0.20949 | 0.000458 | -2e-06 | 100.8902 | 101.1091 | 124.8759 | 20.9443 | 0.001906 | 0.001779 | -151.7056 | 0.23682 | 8.4039 | 0.003482 | 0.005178 | 0.005018 | 5.8e-05 | 0.001007 | 75.1397 | 1.3032 | 2.5479 | 99.63 | 20.99 | 37.186 | 67.3865 | 105.3445 | 64.317 | 0.009287 | 1.0462 | 0.025001 | 23.8983 | -0.019385 | 0.14088 | 8e-06 | 0.00072 | 0.00072 | 8.0413
2019-03-21 11:37 | 1553182632.354 | -481 | 0.20946 | 0.00046 | -3e-06 | 100.816 | 101.1023 | 124.6437 | 20.9444 | 0.00186 | 0.001738 | -151.6815 | 0.23758 | 8.4012 | 0.003488 | 0.004688 | 0.004544 | 6e-05 | 0.001007 | 75.172 | 1.2807 | 2.548 | 99.62 | 20.99 | 37.19 | 66.2717 | 105.4816 | 63.327 | 0.009289 | 1.0492 | 0.025015 | 23.841 | -0.025362 | 0.14134 | 6e-06 | 0.000645 | 0.000645 | 9.0007
2019-03-21 11:37 | 1553182633.3535 | -480 | 0.20948 | 0.000465 | -2e-06 | 100.8789 | 101.0634 | 124.4449 | 20.9447 | 0.001883 | 0.001678 | -151.6631 | 0.23717 | 8.4036 | 0.003446 | 0.005584 | 0.005366 | 5.7e-05 | 0.001006 | 75.1534 | 1.3064 | 2.5479 | 99.62 | 20.99 | 37.192 | 66.8165 | 104.4689 | 61.9066 | 0.00929 | 1.0537 | 0.025048 | 23.7716 | -0.01635 | 0.1411 | -1e-06 | 7e-05 | 7e-05 | 83.3052
2019-03-21 11:37 | 1553182634.3535 | -479 | 0.20949 | 0.000458 | -3e-06 | 100.8252 | 101.0219 | 124.2738 | 20.9449 | 0.001813 | 0.001737 | -151.6569 | 0.23727 | 8.404 | 0.003544 | 0.004541 | 0.004357 | 5.6e-05 | 0.001005 | 75.2077 | 1.3089 | 2.5479 | 99.62 | 20.99 | 37.195 | 65.1453 | 106.8437 | 63.309 | 0.009291 | 1.0493 | 0.024978 | 23.8049 | -0.017428 | 0.14115 | 8e-06 | 0.000752 | 0.000752 | 7.7108
2019-03-21 11:37 | 1553182635.3541 | -478 | 0.20946 | 0.000463 | -2e-06 | 100.9116 | 100.9647 | 124.5876 | 20.9448 | 0.001856 | 0.001773 | -151.6529 | 0.23736 | 8.4003 | 0.003441 | 0.004972 | 0.004741 | 6e-05 | 0.001007 | 75.2335 | 1.2753 | 2.5479 | 99.62 | 20.99 | 37.203 | 66.1884 | 104.3565 | 64.1779 | 0.009293 | 1.0466 | 0.024978 | 23.8657 | -0.004699 | 0.14121 | 2e-06 | 0.000266 | 0.000266 | 21.7839
2019-03-21 11:37 | 1553182636.3545 | -477 | 0.20948 | 0.000458 | -3e-06 | 100.8765 | 101.0304 | 125.0113 | 20.9456 | 0.001921 | 0.001794 | -151.6733 | 0.23688 | 8.4017 | 0.003433 | 0.004861 | 0.004682 | 6.2e-05 | 0.001001 | 75.1736 | 1.3218 | 2.5481 | 99.62 | 20.98 | 37.206 | 67.735 | 104.1622 | 64.6852 | 0.009288 | 1.045 | 0.025001 | 23.9243 | -0.013637 | 0.14092 | 8e-06 | 0.000769 | 0.000769 | 7.533
2019-03-21 11:37 | 1553182637.354 | -476 | 0.2095 | 0.000458 | -2e-06 | 100.8074 | 101.0113 | 125.0423 | 20.9454 | 0.001836 | 0.001743 | -151.676 | 0.23699 | 8.4051 | 0.003481 | 0.004583 | 0.004417 | 6.1e-05 | 0.000998 | 75.1623 | 1.3369 | 2.5479 | 99.62 | 20.98 | 37.212 | 65.688 | 105.3291 | 63.449 | 0.009289 | 1.0489 | 0.02505 | 23.8834 | -0.018079 | 0.14099 | 8e-06 | 0.000729 | 0.000729 | 7.944
2019-03-21 11:37 | 1553182638.3533 | -475 | 0.20949 | 0.000463 | -2e-06 | 100.8389 | 101.0498 | 124.8449 | 20.9444 | 0.001847 | 0.001756 | -151.6696 | 0.23687 | 8.4007 | 0.003463 | 0.005289 | 0.005064 | 6e-05 | 0.001001 | 75.1826 | 1.3092 | 2.5481 | 99.62 | 20.98 | 37.214 | 65.9619 | 104.8845 | 63.7824 | 0.00929 | 1.0478 | 0.025018 | 23.8764 | -0.018685 | 0.14091 | 1e-06 | 0.00023 | 0.00023 | 25.1312
2019-03-21 11:37 | 1553182639.3538 | -474 | 0.20949 | 0.00046 | -3e-06 | 100.867 | 101.0508 | 124.5198 | 20.9461 | 0.001849 | 0.001774 | -151.6662 | 0.23738 | 8.4005 | 0.003445 | 0.004921 | 0.004767 | 5.8e-05 | 0.001003 | 75.1821 | 1.3448 | 2.5479 | 99.62 | 20.98 | 37.216 | 66.0166 | 104.4521 | 64.1994 | 0.00929 | 1.0465 | 0.02497 | 23.86 | -0.016291 | 0.14122 | 6e-06 | 0.00059 | 0.00059 | 9.8366
2019-03-21 11:37 | 1553182640.3534 | -473 | 0.20946 | 0.000459 | -3e-06 | 100.9277 | 101.1067 | 124.451 | 20.9462 | 0.001883 | 0.001755 | -151.6521 | 0.23697 | 8.4037 | 0.003499 | 0.004403 | 0.004207 | 5.9e-05 | 0.001003 | 75.1209 | 1.3356 | 2.5479 | 99.62 | 20.98 | 37.22 | 66.8197 | 105.7587 | 63.7529 | 0.009291 | 1.0479 | 0.02498 | 23.8376 | -0.015846 | 0.14097 | 7e-06 | 0.000687 | 0.000687 | 8.4348
2019-03-21 11:37 | 1553182641.3545 | -472 | 0.20948 | 0.000458 | -2e-06 | 100.9528 | 100.9967 | 124.6846 | 20.9461 | 0.001927 | 0.001733 | -151.6773 | 0.23708 | 8.4003 | 0.003439 | 0.005024 | 0.004897 | 5.9e-05 | 0.001005 | 75.1451 | 1.3692 | 2.5478 | 99.62 | 20.98 | 37.222 | 67.8766 | 104.3123 | 63.2249 | 0.009292 | 1.0496 | 0.025023 | 23.8413 | -0.003891 | 0.14104 | 8e-06 | 0.000771 | 0.000771 | 7.5214
2019-03-21 11:37 | 1553182642.3538 | -471 | 0.20947 | 0.000458 | -2e-06 | 100.8979 | 101.1516 | 124.9015 | 20.9468 | 0.001847 | 0.001763 | -151.6793 | 0.23703 | 8.4049 | 0.0034 | 0.004344 | 0.004193 | 5.8e-05 | 0.001007 | 75.1636 | 1.317 | 2.548 | 99.62 | 20.98 | 37.225 | 65.9759 | 103.3572 | 63.9393 | 0.009292 | 1.0473 | 0.025018 | 23.8873 | -0.022463 | 0.14101 | 8e-06 | 0.000742 | 0.000742 | 7.815
2019-03-21 11:37 | 1553182643.3536 | -470 | 0.20949 | 0.000464 | -3e-06 | 100.9429 | 100.9548 | 125.0007 | 20.9464 | 0.001887 | 0.001703 | -151.681 | 0.23686 | 8.402 | 0.003382 | 0.004776 | 0.004532 | 5.7e-05 | 0.001007 | 75.1458 | 1.3097 | 2.548 | 99.62 | 20.98 | 37.226 | 66.9346 | 102.9174 | 62.498 | 0.009293 | 1.0518 | 0.025082 | 23.8457 | -0.001051 | 0.14091 | 0 | 0.000182 | 0.000182 | 31.8852
2019-03-21 11:37 | 1553182644.3526 | -469 | 0.20947 | 0.000465 | -2e-06 | 100.8962 | 100.9044 | 125.1932 | 20.9478 | 0.001836 | 0.001751 | -151.6809 | 0.23679 | 8.4046 | 0.003452 | 0.004773 | 0.004696 | 5.8e-05 | 0.001009 | 75.1722 | 1.3707 | 2.5481 | 99.62 | 20.98 | 37.226 | 65.6997 | 104.6292 | 63.6639 | 0.009293 | 1.0482 | 0.025057 | 23.9054 | -0.00073 | 0.14087 | -1e-06 | 6.9e-05 | 6.9e-05 | 84.2955
2019-03-21 11:37 | 1553182645.3526 | -468 | 0.20949 | 0.000466 | -2e-06 | 100.8988 | 101.2531 | 124.7714 | 20.9475 | 0.001839 | 0.001789 | -151.7064 | 0.23628 | 8.403 | 0.003395 | 0.004349 | 0.004137 | 5.9e-05 | 0.001014 | 75.1451 | 1.3646 | 2.548 | 99.62 | 20.98 | 37.228 | 65.7613 | 103.2507 | 64.5771 | 0.009293 | 1.0454 | 0.024981 | 23.8975 | -0.031354 | 0.14057 | -3e-06 | -8.9e-05 | -8.9e-05 | -64.6974
2019-03-21 11:37 | 1553182646.3541 | -467 | 0.20949 | 0.000462 | -2e-06 | 100.8613 | 100.9664 | 124.4841 | 20.9474 | 0.001863 | 0.001718 | -151.6666 | 0.23658 | 8.4024 | 0.003475 | 0.004546 | 0.004408 | 5.9e-05 | 0.001011 | 75.1582 | 1.3633 | 2.548 | 99.62 | 20.97 | 37.228 | 66.347 | 105.166 | 62.8699 | 0.009287 | 1.0507 | 0.025016 | 23.8095 | -0.009314 | 0.14074 | 3e-06 | 0.000369 | 0.000369 | 15.6715
2019-03-21 11:37 | 1553182647.3537 | -466 | 0.20949 | 0.000468 | -2e-06 | 100.8878 | 100.7927 | 124.4998 | 20.9476 | 0.00184 | 0.001784 | -151.67 | 0.23672 | 8.4033 | 0.003491 | 0.004275 | 0.004033 | 6e-05 | 0.001014 | 75.1684 | 1.3532 | 2.548 | 99.62 | 20.97 | 37.228 | 65.8041 | 105.5706 | 64.4608 | 0.009287 | 1.0457 | 0.024958 | 23.8673 | 0.008439 | 0.14083 | -4e-06 | -0.000184 | -0.000184 | -31.4133
2019-03-21 11:37 | 1553182648.3528 | -465 | 0.2095 | 0.000466 | -3e-06 | 100.9036 | 101.107 | 124.7641 | 20.948 | 0.001827 | 0.001746 | -151.6513 | 0.23633 | 8.4026 | 0.00343 | 0.005078 | 0.004899 | 5.7e-05 | 0.001006 | 75.1589 | 1.3213 | 2.5479 | 99.62 | 20.97 | 37.227 | 65.4745 | 104.0867 | 63.5461 | 0.009287 | 1.0486 | 0.025019 | 23.8603 | -0.018015 | 0.14059 | -1e-06 | 5.2e-05 | 5.2e-05 | 111.1624
2019-03-21 11:37 | 1553182649.3529 | -464 | 0.20948 | 0.000461 | -2e-06 | 100.9289 | 101.0355 | 124.7787 | 20.9483 | 0.001884 | 0.001744 | -151.6373 | 0.23697 | 8.4016 | 0.003363 | 0.004932 | 0.004795 | 5.8e-05 | 0.000998 | 75.1254 | 1.2672 | 2.5479 | 99.62 | 20.97 | 37.229 | 66.8443 | 102.4534 | 63.4901 | 0.009288 | 1.0487 | 0.025022 | 23.8597 | -0.009441 | 0.14098 | 3e-06 | 0.000343 | 0.000343 | 16.8853
2019-03-21 11:37 | 1553182650.353 | -463 | 0.20947 | 0.000465 | -3e-06 | 100.8503 | 101.0437 | 124.7287 | 20.9485 | 0.00189 | 0.001764 | -151.6425 | 0.23758 | 8.3997 | 0.003512 | 0.004388 | 0.004148 | 5.7e-05 | 0.00099 | 75.1826 | 1.3566 | 2.5479 | 99.62 | 20.97 | 37.231 | 66.9895 | 106.0649 | 63.9792 | 0.009288 | 1.0472 | 0.024999 | 23.8722 | -0.017138 | 0.14134 | 0 | 0.000131 | 0.000131 | 44.3785
2019-03-21 11:38 | 1553182651.354 | -462 | 0.20948 | 0.000466 | -3e-06 | 100.8181 | 101.0525 | 124.6184 | 20.9479 | 0.001883 | 0.001735 | -151.6195 | 0.23638 | 8.4031 | 0.00344 | 0.005309 | 0.005227 | 5.9e-05 | 0.000989 | 75.13 | 1.2973 | 2.5479 | 99.62 | 20.97 | 37.231 | 66.8339 | 104.3227 | 63.2618 | 0.009288 | 1.0494 | 0.025015 | 23.8363 | -0.020776 | 0.14062 | -2e-06 | 7e-06 | 7e-06 | 845.3774
2019-03-21 11:38 | 1553182652.3531 | -461 | 0.20951 | 0.000467 | -3e-06 | 100.8425 | 101.0876 | 124.5059 | 20.9485 | 0.001865 | 0.00169 | -151.6436 | 0.2366 | 8.4025 | 0.003467 | 0.004536 | 0.004343 | 6e-05 | 0.000992 | 75.1894 | 1.2903 | 2.548 | 99.62 | 20.97 | 37.232 | 66.3996 | 104.9908 | 62.1978 | 0.009288 | 1.0528 | 0.025043 | 23.7878 | -0.021714 | 0.14075 | -4e-06 | -9.6e-05 | -9.6e-05 | -60.2461
2019-03-21 11:38 | 1553182653.3531 | -460 | 0.20952 | 0.000463 | -3e-06 | 100.8714 | 101.0863 | 124.7186 | 20.9484 | 0.001876 | 0.001788 | -151.6526 | 0.23682 | 8.4042 | 0.003447 | 0.00466 | 0.004487 | 5.8e-05 | 0.000988 | 75.1388 | 1.2803 | 2.5479 | 99.62 | 20.97 | 37.233 | 66.6578 | 104.4875 | 64.5554 | 0.009289 | 1.0454 | 0.024977 | 23.8916 | -0.01903 | 0.14089 | 2e-06 | 0.000335 | 0.000335 | 17.2678
2019-03-21 11:38 | 1553182654.3536 | -459 | 0.20949 | 0.000464 | -3e-06 | 100.8911 | 101.122 | 124.5935 | 20.949 | 0.001822 | 0.00179 | -151.6523 | 0.2377 | 8.4035 | 0.003457 | 0.005045 | 0.004905 | 6.1e-05 | 0.000992 | 75.1512 | 1.3395 | 2.548 | 99.62 | 20.97 | 37.234 | 65.3685 | 104.7506 | 64.5849 | 0.009289 | 1.0453 | 0.024963 | 23.8807 | -0.02045 | 0.14141 | 1e-06 | 0.000251 | 0.000251 | 23.1117
2019-03-21 11:38 | 1553182655.3532 | -458 | 0.20949 | 0.000464 | -3e-06 | 100.9143 | 101.3291 | 124.7247 | 20.9488 | 0.001908 | 0.001717 | -151.6295 | 0.23741 | 8.4006 | 0.003475 | 0.005371 | 0.005197 | 6.1e-05 | 0.00099 | 75.188 | 1.3254 | 2.5479 | 99.62 | 20.97 | 37.235 | 67.433 | 105.1776 | 62.8487 | 0.009289 | 1.0507 | 0.025041 | 23.8318 | -0.036688 | 0.14123 | 1e-06 | 0.000267 | 0.000267 | 21.7757
2019-03-21 11:38 | 1553182656.3537 | -457 | 0.2095 | 0.000466 | -3e-06 | 100.9087 | 101.1554 | 124.7402 | 20.9491 | 0.001862 | 0.001772 | -151.635 | 0.23732 | 8.4019 | 0.003504 | 0.00491 | 0.004757 | 5.9e-05 | 0.000988 | 75.16 | 1.3018 | 2.548 | 99.62 | 20.97 | 37.236 | 66.3299 | 105.8908 | 64.1595 | 0.009289 | 1.0467 | 0.024994 | 23.8797 | -0.021844 | 0.14118 | -1e-06 | 0.000119 | 0.000119 | 48.7639
2019-03-21 11:38 | 1553182657.3537 | -456 | 0.20947 | 0.000463 | -3e-06 | 100.892 | 101.0369 | 124.5224 | 20.9498 | 0.001896 | 0.001753 | -151.624 | 0.2369 | 8.4027 | 0.003522 | 0.004782 | 0.00462 | 6e-05 | 0.000991 | 75.1354 | 1.3156 | 2.5481 | 99.62 | 20.97 | 37.238 | 67.1475 | 106.3203 | 63.7118 | 0.00929 | 1.048 | 0.024988 | 23.843 | -0.012836 | 0.14093 | 2e-06 | 0.000339 | 0.000339 | 17.0985
2019-03-21 11:38 | 1553182658.3538 | -455 | 0.20949 | 0.000465 | -3e-06 | 100.8363 | 101.0989 | 124.1391 | 20.9496 | 0.001841 | 0.001746 | -151.6349 | 0.23673 | 8.4023 | 0.003536 | 0.00445 | 0.004142 | 6.2e-05 | 0.00099 | 75.1907 | 1.3432 | 2.5481 | 99.62 | 20.97 | 37.238 | 65.828 | 106.6688 | 63.5478 | 0.00929 | 1.0486 | 0.024956 | 23.8005 | -0.023263 | 0.14083 | -1e-06 | 0.000119 | 0.000119 | 48.6145
2019-03-21 11:38 | 1553182659.3534 | -454 | 0.20951 | 0.000463 | -2e-06 | 100.8586 | 100.9868 | 124.2264 | 20.9504 | 0.001856 | 0.001747 | -151.6455 | 0.2379 | 8.4008 | 0.003526 | 0.005524 | 0.005417 | 5.8e-05 | 0.000991 | 75.2302 | 1.3666 | 2.5481 | 99.62 | 20.97 | 37.239 | 66.1792 | 106.4097 | 63.5594 | 0.00929 | 1.0485 | 0.024964 | 23.8093 | -0.011362 | 0.14153 | 1e-06 | 0.000236 | 0.000236 | 24.6043
2019-03-21 11:38 | 1553182660.3535 | -453 | 0.2095 | 0.000468 | -2e-06 | 100.9602 | 101.1033 | 124.3984 | 20.9509 | 0.001916 | 0.001769 | -151.6501 | 0.23725 | 8.3986 | 0.003431 | 0.004423 | 0.004294 | 6e-05 | 0.00099 | 75.1438 | 1.3466 | 2.548 | 99.62 | 20.97 | 37.238 | 67.6339 | 104.1093 | 64.0949 | 0.00929 | 1.0469 | 0.024962 | 23.8447 | -0.012667 | 0.14114 | -5e-06 | -0.00024 | -0.00024 | -24.2086
2019-03-21 11:38 | 1553182661.3529 | -452 | 0.2095 | 0.000465 | -3e-06 | 100.9131 | 101.092 | 124.601 | 20.9501 | 0.001875 | 0.001698 | -151.6466 | 0.23757 | 8.4041 | 0.003541 | 0.004594 | 0.004326 | 5.8e-05 | 0.000988 | 75.1535 | 1.2991 | 2.5481 | 99.62 | 20.97 | 37.237 | 66.633 | 106.778 | 62.3804 | 0.00929 | 1.0522 | 0.025046 | 23.8033 | -0.015844 | 0.14133 | -1e-06 | 8.1e-05 | 8.1e-05 | 71.3217
2019-03-21 11:38 | 1553182662.3526 | -451 | 0.20948 | 0.000465 | -3e-06 | 100.878 | 100.9384 | 124.6576 | 20.9508 | 0.001849 | 0.001773 | -151.6413 | 0.23711 | 8.4034 | 0.003525 | 0.004029 | 0.003826 | 6e-05 | 0.00099 | 75.2034 | 1.301 | 2.5479 | 99.62 | 20.97 | 37.236 | 66.0071 | 106.3882 | 64.1939 | 0.009289 | 1.0465 | 0.024984 | 23.873 | -0.00536 | 0.14106 | -1e-06 | 0.000117 | 0.000117 | 49.5991
2019-03-21 11:38 | 1553182663.353 | -450 | 0.2095 | 0.000462 | -2e-06 | 100.9104 | 101.0832 | 124.2966 | 20.9512 | 0.001857 | 0.001718 | -151.631 | 0.23856 | 8.4011 | 0.003496 | 0.004682 | 0.004524 | 5.9e-05 | 0.000989 | 75.1691 | 1.3207 | 2.5481 | 99.62 | 20.97 | 37.236 | 66.2019 | 105.6959 | 62.8541 | 0.009289 | 1.0507 | 0.024998 | 23.791 | -0.015298 | 0.14192 | 3e-06 | 0.000334 | 0.000334 | 17.4616
2019-03-21 11:38 | 1553182664.3531 | -449 | 0.20948 | 0.000462 | -2e-06 | 100.881 | 101.1854 | 124.4644 | 20.9502 | 0.001935 | 0.001772 | -151.6342 | 0.23722 | 8.4022 | 0.003519 | 0.004895 | 0.004671 | 6.1e-05 | 0.000987 | 75.158 | 1.2922 | 2.548 | 99.62 | 20.97 | 37.236 | 68.095 | 106.2412 | 64.1731 | 0.009289 | 1.0466 | 0.024966 | 23.8538 | -0.02695 | 0.14112 | 2e-06 | 0.000258 | 0.000258 | 22.4649
2019-03-21 11:38 | 1553182665.3528 | -448 | 0.20949 | 0.000464 | -3e-06 | 100.8711 | 101.2296 | 124.5792 | 20.9518 | 0.001878 | 0.001757 | -151.6349 | 0.23731 | 8.3993 | 0.00349 | 0.004794 | 0.00465 | 6.4e-05 | 0.000986 | 75.1702 | 1.3341 | 2.5479 | 99.62 | 20.97 | 37.235 | 66.7128 | 105.5495 | 63.7958 | 0.009289 | 1.0478 | 0.024991 | 23.8514 | -0.031735 | 0.14118 | 0 | 0.00017 | 0.00017 | 34.0322
2019-03-21 11:38 | 1553182666.353 | -447 | 0.20947 | 0.000465 | -3e-06 | 100.9435 | 100.9919 | 124.7857 | 20.9516 | 0.001889 | 0.001745 | -151.6412 | 0.23794 | 8.4044 | 0.003451 | 0.005119 | 0.00488 | 5.8e-05 | 0.000988 | 75.1682 | 1.3112 | 2.548 | 99.62 | 20.97 | 37.234 | 66.98 | 104.6108 | 63.5145 | 0.009289 | 1.0487 | 0.025022 | 23.8612 | -0.004287 | 0.14155 | -1e-06 | 0.00011 | 0.00011 | 52.6611
2019-03-21 11:38 | 1553182667.3539 | -446 | 0.20949 | 0.000471 | -2e-06 | 100.8681 | 101.2089 | 124.9604 | 20.9517 | 0.001861 | 0.00173 | -151.6445 | 0.23706 | 8.4019 | 0.003475 | 0.004854 | 0.004655 | 6.2e-05 | 0.000989 | 75.1808 | 1.303 | 2.5479 | 99.62 | 20.97 | 37.231 | 66.3092 | 105.1702 | 63.1514 | 0.009288 | 1.0498 | 0.025053 | 23.865 | -0.030163 | 0.14103 | -9e-06 | -0.000527 | -0.000527 | -10.9983
2019-03-21 11:38 | 1553182668.3538 | -445 | 0.20947 | 0.000461 | -2e-06 | 100.8774 | 101.1377 | 124.7745 | 20.9517 | 0.001863 | 0.001721 | -151.6463 | 0.23702 | 8.4029 | 0.003507 | 0.005297 | 0.005117 | 6.4e-05 | 0.000989 | 75.1395 | 1.3456 | 2.548 | 99.62 | 20.97 | 37.228 | 66.3568 | 105.9537 | 62.9475 | 0.009287 | 1.0504 | 0.025042 | 23.84 | -0.023051 | 0.141 | 4e-06 | 0.000405 | 0.000405 | 14.3111
2019-03-21 11:38 | 1553182669.3533 | -444 | 0.20949 | 0.000469 | -2e-06 | 100.8678 | 101.1006 | 125.0353 | 20.9527 | 0.001892 | 0.001739 | -151.6543 | 0.23772 | 8.4027 | 0.003453 | 0.004487 | 0.004367 | 5.8e-05 | 0.00099 | 75.2151 | 1.3907 | 2.5479 | 99.62 | 20.98 | 37.226 | 67.0419 | 104.6505 | 63.3775 | 0.009293 | 1.0491 | 0.025052 | 23.8802 | -0.020613 | 0.14142 | -6e-06 | -0.000303 | -0.000303 | -19.184
2019-03-21 11:38 | 1553182670.3539 | -443 | 0.20948 | 0.000467 | -2e-06 | 100.9522 | 101.1489 | 124.9284 | 20.9533 | 0.001913 | 0.001736 | -151.6406 | 0.23757 | 8.4022 | 0.00352 | 0.005001 | 0.004796 | 6e-05 | 0.00099 | 75.147 | 1.3447 | 2.5479 | 99.62 | 20.98 | 37.224 | 67.5638 | 106.2757 | 63.3046 | 0.009292 | 1.0493 | 0.025044 | 23.8674 | -0.017417 | 0.14133 | -4e-06 | -0.000132 | -0.000132 | -43.8898
2019-03-21 11:38 | 1553182671.3535 | -442 | 0.20949 | 0.000464 | -3e-06 | 100.872 | 101.354 | 124.9439 | 20.9525 | 0.00191 | 0.001727 | -151.6491 | 0.23822 | 8.4017 | 0.00343 | 0.00429 | 0.004165 | 5.9e-05 | 0.000997 | 75.1135 | 1.356 | 2.5481 | 99.62 | 20.98 | 37.222 | 67.4889 | 104.0943 | 63.071 | 0.009292 | 1.05 | 0.025055 | 23.8606 | -0.042633 | 0.14172 | 1e-06 | 0.000254 | 0.000254 | 22.9412
2019-03-21 11:38 | 1553182672.3541 | -441 | 0.2095 | 0.000467 | -3e-06 | 100.8929 | 101.2187 | 124.7062 | 20.9528 | 0.001823 | 0.001721 | -151.6406 | 0.23701 | 8.4021 | 0.003442 | 0.004604 | 0.004405 | 6.2e-05 | 0.000997 | 75.1733 | 1.3505 | 2.5479 | 99.62 | 20.98 | 37.219 | 65.3863 | 104.3902 | 62.9276 | 0.009291 | 1.0505 | 0.025036 | 23.8328 | -0.028844 | 0.141 | -4e-06 | -8.7e-05 | -8.7e-05 | -66.5518
2019-03-21 11:38 | 1553182673.3545 | -440 | 0.20948 | 0.000462 | -3e-06 | 100.8699 | 101.0335 | 124.5294 | 20.9522 | 0.001861 | 0.00175 | -151.6368 | 0.23701 | 8.4031 | 0.003496 | 0.004619 | 0.004456 | 5.7e-05 | 0.000998 | 75.2228 | 1.3288 | 2.548 | 99.62 | 20.98 | 37.219 | 66.3135 | 105.6859 | 63.6251 | 0.009291 | 1.0483 | 0.024992 | 23.8406 | -0.01449 | 0.141 | 3e-06 | 0.000372 | 0.000372 | 15.5659
2019-03-21 11:38 | 1553182674.3526 | -439 | 0.2095 | 0.000464 | -3e-06 | 100.9182 | 101.0045 | 124.4407 | 20.9533 | 0.001889 | 0.001778 | -151.6509 | 0.23717 | 8.4019 | 0.003441 | 0.00534 | 0.005172 | 6.5e-05 | 0.000999 | 75.1862 | 1.3382 | 2.5481 | 99.62 | 20.98 | 37.219 | 66.9708 | 104.349 | 64.3156 | 0.009291 | 1.0462 | 0.024958 | 23.8565 | -0.007647 | 0.1411 | 0 | 0.000177 | 0.000177 | 32.7055
2019-03-21 11:38 | 1553182675.3527 | -438 | 0.20949 | 0.00047 | -3e-06 | 100.8801 | 100.9878 | 124.5048 | 20.9526 | 0.001885 | 0.001778 | -151.655 | 0.23737 | 8.4033 | 0.003548 | 0.005011 | 0.004788 | 6.2e-05 | 0.001006 | 75.2477 | 1.3344 | 2.548 | 99.62 | 20.98 | 37.217 | 66.8828 | 106.9442 | 64.301 | 0.00929 | 1.0462 | 0.024965 | 23.8622 | -0.00955 | 0.14121 | -6e-06 | -0.000286 | -0.000286 | -20.3196
2019-03-21 11:38 | 1553182676.3532 | -437 | 0.20949 | 0.000462 | -3e-06 | 100.8845 | 101.122 | 124.4491 | 20.9534 | 0.001917 | 0.001713 | -151.656 | 0.23768 | 8.4044 | 0.00351 | 0.005584 | 0.005449 | 6.4e-05 | 0.001 | 75.2014 | 1.3205 | 2.5479 | 99.62 | 20.98 | 37.215 | 67.6519 | 106.0381 | 62.7486 | 0.00929 | 1.051 | 0.025017 | 23.8019 | -0.021031 | 0.14139 | 2e-06 | 0.000334 | 0.000334 | 17.3759
2019-03-21 11:38 | 1553182677.3524 | -436 | 0.20948 | 0.000462 | -3e-06 | 100.9012 | 101.31 | 124.6057 | 20.9533 | 0.00189 | 0.001725 | -151.6495 | 0.23705 | 8.4012 | 0.003504 | 0.004581 | 0.004378 | 6e-05 | 0.001001 | 75.1842 | 1.3072 | 2.548 | 99.62 | 20.98 | 37.214 | 67.0067 | 105.8886 | 63.0437 | 0.00929 | 1.0501 | 0.025022 | 23.8273 | -0.036166 | 0.14102 | 3e-06 | 0.000366 | 0.000366 | 15.851
2019-03-21 11:38 | 1553182678.3538 | -435 | 0.20951 | 0.000466 | -3e-06 | 100.9102 | 101.1527 | 124.9132 | 20.9537 | 0.001855 | 0.001765 | -151.6571 | 0.23822 | 8.4041 | 0.003568 | 0.005175 | 0.005046 | 6.2e-05 | 0.001005 | 75.1826 | 1.288 | 2.5479 | 99.62 | 20.98 | 37.214 | 66.1651 | 107.4516 | 64.0061 | 0.00929 | 1.0471 | 0.025017 | 23.8908 | -0.021471 | 0.14172 | -1e-06 | 9.1e-05 | 9.1e-05 | 64.0766
2019-03-21 11:38 | 1553182679.3525 | -434 | 0.2095 | 0.000463 | -3e-06 | 100.9444 | 101.1425 | 124.7003 | 20.9541 | 0.001827 | 0.00175 | -151.649 | 0.2377 | 8.4027 | 0.003523 | 0.004938 | 0.004831 | 6.2e-05 | 0.001002 | 75.1858 | 1.2875 | 2.5479 | 99.62 | 20.98 | 37.214 | 65.4961 | 106.3575 | 63.6458 | 0.00929 | 1.0482 | 0.025009 | 23.8577 | -0.017531 | 0.14141 | 1e-06 | 0.000266 | 0.000266 | 21.8195
2019-03-21 11:38 | 1553182680.3526 | -433 | 0.2095 | 0.000461 | -3e-06 | 100.8062 | 101.1108 | 124.4698 | 20.9541 | 0.001907 | 0.00174 | -151.6448 | 0.23727 | 8.4025 | 0.003461 | 0.005491 | 0.005237 | 6.4e-05 | 0.000998 | 75.2362 | 1.312 | 2.548 | 99.62 | 20.98 | 37.214 | 67.4203 | 104.8438 | 63.4077 | 0.00929 | 1.049 | 0.024994 | 23.8272 | -0.026987 | 0.14115 | 4e-06 | 0.000473 | 0.000473 | 12.2771
2019-03-21 11:38 | 1553182681.3534 | -432 | 0.20952 | 0.000464 | -2e-06 | 100.9122 | 101.1104 | 123.9896 | 20.9542 | 0.001862 | 0.001753 | -151.6325 | 0.23694 | 8.4027 | 0.003474 | 0.004794 | 0.00468 | 6.2e-05 | 0.000994 | 75.1708 | 1.3082 | 2.5479 | 99.62 | 20.97 | 37.213 | 66.3296 | 105.1583 | 63.7154 | 0.009284 | 1.048 | 0.024935 | 23.7921 | -0.017551 | 0.14096 | 0 | 0.000176 | 0.000176 | 32.933
2019-03-21 11:38 | 1553182682.3541 | -431 | 0.20954 | 0.000468 | -3e-06 | 100.8872 | 101.0362 | 124.4306 | 20.9549 | 0.001827 | 0.001754 | -151.6497 | 0.2374 | 8.4023 | 0.003466 | 0.004477 | 0.004338 | 6.4e-05 | 0.000992 | 75.1589 | 1.3301 | 2.548 | 99.62 | 20.97 | 37.212 | 65.4858 | 104.9665 | 63.7418 | 0.009283 | 1.0479 | 0.024978 | 23.8353 | -0.013199 | 0.14123 | -4e-06 | -0.000116 | -0.000116 | -50.2162
2019-03-21 11:38 | 1553182683.3537 | -430 | 0.20951 | 0.000466 | -2e-06 | 100.9417 | 101.1278 | 124.8463 | 20.9552 | 0.00188 | 0.001756 | -151.6494 | 0.23757 | 8.4017 | 0.003528 | 0.004813 | 0.004648 | 6.4e-05 | 0.000995 | 75.1254 | 1.3276 | 2.5481 | 99.62 | 20.97 | 37.211 | 66.7746 | 106.4656 | 63.7874 | 0.009283 | 1.0478 | 0.025018 | 23.8767 | -0.016473 | 0.14133 | -3e-06 | -6.2e-05 | -6.2e-05 | -93.5708
2019-03-21 11:38 | 1553182684.3538 | -429 | 0.20951 | 0.000464 | -2e-06 | 100.9414 | 101.1029 | 124.9221 | 20.9546 | 0.001873 | 0.001793 | -151.6556 | 0.23702 | 8.4025 | 0.003544 | 0.005106 | 0.004938 | 6.2e-05 | 0.000996 | 75.1574 | 1.3317 | 2.5479 | 99.62 | 20.97 | 37.21 | 66.5866 | 106.8666 | 64.6661 | 0.009283 | 1.0451 | 0.024993 | 23.915 | -0.0143 | 0.141 | 0 | 0.000154 | 0.000154 | 37.7179
2019-03-21 11:38 | 1553182685.3528 | -428 | 0.20956 | 0.000467 | -2e-06 | 100.8431 | 101.0352 | 124.7625 | 20.9552 | 0.001854 | 0.001764 | -151.6593 | 0.23708 | 8.4041 | 0.003507 | 0.004228 | 0.00412 | 6.7e-05 | 0.000999 | 75.1558 | 1.2951 | 2.5477 | 99.62 | 20.97 | 37.205 | 66.1439 | 105.9613 | 63.9731 | 0.009282 | 1.0472 | 0.025003 | 23.8752 | -0.017018 | 0.14104 | -4e-06 | -0.000148 | -0.000148 | -39.147
2019-03-21 11:38 | 1553182686.3538 | -427 | 0.20949 | 0.000463 | -3e-06 | 100.8699 | 101.1496 | 124.7712 | 20.9544 | 0.001898 | 0.001785 | -151.6633 | 0.23798 | 8.4034 | 0.003574 | 0.005159 | 0.005015 | 6.5e-05 | 0.000999 | 75.1589 | 1.3083 | 2.5479 | 99.62 | 20.97 | 37.202 | 67.2027 | 107.5782 | 64.469 | 0.009281 | 1.0457 | 0.024985 | 23.8936 | -0.024766 | 0.14157 | 2e-06 | 0.000348 | 0.000348 | 16.7182
2019-03-21 11:38 | 1553182687.3531 | -426 | 0.2095 | 0.000466 | -2e-06 | 100.9343 | 100.9728 | 124.5346 | 20.956 | 0.001882 | 0.001761 | -151.6586 | 0.23687 | 8.4016 | 0.0035 | 0.004641 | 0.004423 | 6.3e-05 | 0.001002 | 75.2522 | 1.3217 | 2.5478 | 99.62 | 20.97 | 37.202 | 66.8136 | 105.7841 | 63.8991 | 0.009281 | 1.0475 | 0.024983 | 23.8508 | -0.003416 | 0.14091 | -3e-06 | -8.2e-05 | -8.2e-05 | -70.8653
2019-03-21 11:38 | 1553182688.3533 | -425 | 0.20949 | 0.000463 | -2e-06 | 100.8953 | 101.1486 | 124.3423 | 20.9557 | 0.001943 | 0.001804 | -151.6515 | 0.23792 | 8.4024 | 0.003561 | 0.004641 | 0.004543 | 6.5e-05 | 0.001005 | 75.2308 | 1.3081 | 2.5481 | 99.62 | 20.97 | 37.202 | 68.2797 | 107.2809 | 64.9434 | 0.009281 | 1.0442 | 0.024925 | 23.8693 | -0.02243 | 0.14154 | 1e-06 | 0.000241 | 0.000241 | 24.1724
2019-03-21 11:38 | 1553182689.3531 | -424 | 0.20952 | 0.000466 | -3e-06 | 100.8604 | 101.168 | 124.2339 | 20.9557 | 0.001894 | 0.001789 | -151.6464 | 0.23786 | 8.4025 | 0.003477 | 0.004492 | 0.004284 | 6.6e-05 | 0.001002 | 75.2043 | 1.293 | 2.5478 | 99.62 | 20.97 | 37.201 | 67.0949 | 105.2457 | 64.5702 | 0.009281 | 1.0454 | 0.024928 | 23.8457 | -0.027237 | 0.1415 | -2e-06 | 5e-05 | 5e-05 | 116.5602
2019-03-21 11:38 | 1553182690.3535 | -423 | 0.20952 | 0.000463 | -2e-06 | 100.9063 | 101.0668 | 124.4876 | 20.9556 | 0.001857 | 0.001746 | -151.6618 | 0.23705 | 8.4026 | 0.003493 | 0.005278 | 0.005155 | 6.3e-05 | 0.001003 | 75.2086 | 1.2954 | 2.5479 | 99.62 | 20.97 | 37.199 | 66.2162 | 105.6122 | 63.5341 | 0.00928 | 1.0486 | 0.024992 | 23.8334 | -0.014221 | 0.14102 | 1e-06 | 0.000213 | 0.000213 | 27.1849
2019-03-21 11:38 | 1553182691.3536 | -422 | 0.2095 | 0.000463 | -3e-06 | 100.8354 | 101.057 | 124.8803 | 20.9562 | 0.001911 | 0.001738 | -151.6813 | 0.23788 | 8.402 | 0.003583 | 0.00458 | 0.004402 | 6.5e-05 | 0.001007 | 75.1896 | 1.2754 | 2.548 | 99.62 | 20.97 | 37.2 | 67.5048 | 107.799 | 63.3589 | 0.00928 | 1.0491 | 0.025037 | 23.8647 | -0.019635 | 0.14151 | 1e-06 | 0.000268 | 0.000268 | 21.6884
2019-03-21 11:38 | 1553182692.3532 | -421 | 0.20948 | 0.000462 | -3e-06 | 100.7826 | 101.0917 | 125.0756 | 20.9574 | 0.00187 | 0.001732 | -151.6734 | 0.23756 | 8.4021 | 0.003503 | 0.005024 | 0.004829 | 6.5e-05 | 0.001005 | 75.1865 | 1.2823 | 2.548 | 99.62 | 20.97 | 37.199 | 66.5222 | 105.8736 | 63.1952 | 0.00928 | 1.0497 | 0.025063 | 23.8776 | -0.027389 | 0.14132 | 4e-06 | 0.00044 | 0.00044 | 13.2098
2019-03-21 11:38 | 1553182693.3523 | -420 | 0.2095 | 0.000465 | -3e-06 | 100.8309 | 101.2 | 124.9709 | 20.9569 | 0.001837 | 0.001721 | -151.6849 | 0.23759 | 8.4024 | 0.003601 | 0.004863 | 0.00474 | 6.6e-05 | 0.00101 | 75.1754 | 1.2731 | 2.5479 | 99.62 | 20.97 | 37.203 | 65.7415 | 108.2353 | 62.9518 | 0.009281 | 1.0504 | 0.025062 | 23.8589 | -0.032684 | 0.14134 | -1e-06 | 0.000121 | 0.000121 | 47.9215
2019-03-21 11:38 | 1553182694.3523 | -419 | 0.20949 | 0.000458 | -3e-06 | 100.8866 | 100.9354 | 124.9836 | 20.9575 | 0.001863 | 0.001804 | -151.6826 | 0.23616 | 8.4009 | 0.003594 | 0.005101 | 0.00499 | 6.5e-05 | 0.001009 | 75.1336 | 1.2951 | 2.5478 | 99.62 | 20.97 | 37.203 | 66.3648 | 108.0821 | 64.9287 | 0.009281 | 1.0443 | 0.02499 | 23.9302 | -0.004322 | 0.14049 | 8e-06 | 0.000804 | 0.000804 | 7.179
2019-03-21 11:38 | 1553182695.3527 | -418 | 0.2095 | 0.000462 | -3e-06 | 100.8768 | 101.2235 | 124.8996 | 20.9576 | 0.001861 | 0.001732 | -151.6825 | 0.23776 | 8.4005 | 0.00354 | 0.005576 | 0.0054 | 6.1e-05 | 0.001006 | 75.1336 | 1.2724 | 2.5478 | 99.62 | 20.97 | 37.204 | 66.3041 | 106.7616 | 63.2003 | 0.009281 | 1.0496 | 0.025045 | 23.8609 | -0.030691 | 0.14144 | 3e-06 | 0.000391 | 0.000391 | 14.8588
2019-03-21 11:38 | 1553182696.3534 | -417 | 0.20951 | 0.000462 | -3e-06 | 100.8634 | 100.9793 | 124.8996 | 20.9577 | 0.001957 | 0.001741 | -151.698 | 0.23639 | 8.4051 | 0.003523 | 0.005329 | 0.005243 | 6.4e-05 | 0.001013 | 75.1939 | 1.2838 | 2.5479 | 99.62 | 20.97 | 37.207 | 68.6235 | 106.3446 | 63.4259 | 0.009282 | 1.0489 | 0.025037 | 23.8689 | -0.010275 | 0.14063 | 3e-06 | 0.00039 | 0.00039 | 14.8175
2019-03-21 11:38 | 1553182697.3528 | -416 | 0.20952 | 0.000462 | -3e-06 | 100.9003 | 101.166 | 125.032 | 20.9579 | 0.001872 | 0.001772 | -151.7118 | 0.23769 | 8.3987 | 0.003527 | 0.005138 | 0.004927 | 6e-05 | 0.001007 | 75.1636 | 1.3205 | 2.5479 | 99.62 | 20.97 | 37.21 | 66.5703 | 106.4502 | 64.1719 | 0.009283 | 1.0466 | 0.025022 | 23.908 | -0.023517 | 0.1414 | 3e-06 | 0.000412 | 0.000412 | 14.101
2019-03-21 11:38 | 1553182698.3525 | -415 | 0.2095 | 0.000462 | -3e-06 | 100.8964 | 101.0648 | 125.0871 | 20.9587 | 0.001883 | 0.001731 | -151.7053 | 0.23647 | 8.4048 | 0.003526 | 0.004954 | 0.004743 | 6.3e-05 | 0.001012 | 75.2315 | 1.3453 | 2.548 | 99.62 | 20.96 | 37.213 | 66.8364 | 106.4171 | 63.1809 | 0.009278 | 1.0497 | 0.025065 | 23.8782 | -0.014912 | 0.14067 | 2e-06 | 0.000338 | 0.000338 | 17.0959
2019-03-21 11:38 | 1553182699.3535 | -414 | 0.20948 | 0.000463 | -3e-06 | 100.867 | 101.2327 | 125.2078 | 20.9582 | 0.001893 | 0.001765 | -151.7133 | 0.2366 | 8.4038 | 0.003543 | 0.004093 | 0.003965 | 6.5e-05 | 0.001013 | 75.1772 | 1.2995 | 2.548 | 99.62 | 20.96 | 37.214 | 67.0917 | 106.8456 | 64.0081 | 0.009278 | 1.0471 | 0.025046 | 23.919 | -0.032376 | 0.14075 | 1e-06 | 0.000267 | 0.000267 | 21.7043
2019-03-21 11:38 | 1553182700.354 | -413 | 0.20948 | 0.00046 | -3e-06 | 100.8521 | 100.9814 | 125.258 | 20.9593 | 0.001895 | 0.001718 | -151.7148 | 0.23662 | 8.4018 | 0.003544 | 0.004953 | 0.004783 | 6.6e-05 | 0.001016 | 75.1756 | 1.321 | 2.5478 | 99.62 | 20.96 | 37.216 | 67.1328 | 106.861 | 62.8707 | 0.009279 | 1.0507 | 0.025094 | 23.8834 | -0.01146 | 0.14077 | 6e-06 | 0.000614 | 0.000614 | 9.417
2019-03-21 11:38 | 1553182701.3527 | -412 | 0.20949 | 0.000462 | -2e-06 | 100.8738 | 101.2906 | 125.1916 | 20.9592 | 0.001822 | 0.001779 | -151.7162 | 0.2367 | 8.4023 | 0.003515 | 0.00489 | 0.004681 | 6.1e-05 | 0.001016 | 75.1465 | 1.323 | 2.5479 | 99.62 | 20.96 | 37.215 | 65.3752 | 106.1602 | 64.3457 | 0.009278 | 1.0461 | 0.025032 | 23.9295 | -0.036882 | 0.14081 | 2e-06 | 0.000341 | 0.000341 | 16.9953
2019-03-21 11:38 | 1553182702.3537 | -411 | 0.20952 | 0.000463 | -3e-06 | 100.8407 | 101.1394 | 124.6914 | 20.9594 | 0.001888 | 0.001773 | -151.7204 | 0.23602 | 8.4006 | 0.003459 | 0.004544 | 0.004283 | 6.2e-05 | 0.001018 | 75.1079 | 1.3112 | 2.548 | 99.62 | 20.96 | 37.217 | 66.9535 | 104.7994 | 64.1874 | 0.009279 | 1.0466 | 0.024988 | 23.876 | -0.026452 | 0.14041 | 2e-06 | 0.000337 | 0.000337 | 17.1031
2019-03-21 11:38 | 1553182703.3528 | -410 | 0.2095 | 0.000463 | -3e-06 | 100.9295 | 101.1247 | 124.3761 | 20.9599 | 0.001884 | 0.001778 | -151.7063 | 0.23614 | 8.4034 | 0.003468 | 0.004556 | 0.004365 | 6.2e-05 | 0.001019 | 75.1851 | 1.3157 | 2.5479 | 99.62 | 20.96 | 37.216 | 66.857 | 105.0233 | 64.326 | 0.009279 | 1.0461 | 0.024951 | 23.8507 | -0.017284 | 0.14048 | 1e-06 | 0.000234 | 0.000234 | 24.6917
2019-03-21 11:38 | 1553182704.3538 | -409 | 0.20949 | 0.000464 | -3e-06 | 100.9233 | 101.2027 | 123.9619 | 20.9597 | 0.001879 | 0.00179 | -151.6971 | 0.2368 | 8.4036 | 0.00351 | 0.00493 | 0.004721 | 6.1e-05 | 0.001018 | 75.1307 | 1.3311 | 2.5479 | 99.62 | 20.96 | 37.216 | 66.7577 | 106.0404 | 64.5998 | 0.009279 | 1.0453 | 0.024899 | 23.8206 | -0.024735 | 0.14088 | 0 | 0.000152 | 0.000152 | 38.1323
2019-03-21 11:38 | 1553182705.3528 | -408 | 0.20952 | 0.000465 | -2e-06 | 100.9572 | 101.1551 | 124.0619 | 20.9595 | 0.001873 | 0.001756 | -151.6868 | 0.23675 | 8.4028 | 0.003408 | 0.004962 | 0.004832 | 6.2e-05 | 0.001016 | 75.2104 | 1.3049 | 2.5479 | 99.62 | 20.96 | 37.215 | 66.5921 | 103.5682 | 63.7821 | 0.009278 | 1.0478 | 0.02494 | 23.8014 | -0.01751 | 0.14084 | -1e-06 | 5.8e-05 | 5.8e-05 | 99.5208
2019-03-21 11:38 | 1553182706.3519 | -407 | 0.2095 | 0.000461 | -3e-06 | 100.8598 | 101.2109 | 124.1128 | 20.9597 | 0.001898 | 0.001782 | -151.6715 | 0.2379 | 8.4038 | 0.003466 | 0.004794 | 0.004655 | 6.3e-05 | 0.001019 | 75.1876 | 1.2932 | 2.548 | 99.62 | 20.96 | 37.216 | 67.194 | 104.964 | 64.421 | 0.009279 | 1.0458 | 0.024921 | 23.8288 | -0.031084 | 0.14153 | 4e-06 | 0.000509 | 0.000509 | 11.4179
2019-03-21 11:38 | 1553182707.3536 | -406 | 0.20951 | 0.000461 | -3e-06 | 100.8935 | 101.0839 | 124.2553 | 20.9601 | 0.001876 | 0.001739 | -151.6776 | 0.2374 | 8.4037 | 0.00348 | 0.004543 | 0.004322 | 6.3e-05 | 0.001023 | 75.1772 | 1.3023 | 2.5479 | 99.62 | 20.96 | 37.215 | 66.6834 | 105.3208 | 63.3753 | 0.009278 | 1.0491 | 0.024974 | 23.8055 | -0.016864 | 0.14123 | 5e-06 | 0.000581 | 0.000581 | 9.9923
2019-03-21 11:38 | 1553182708.3524 | -405 | 0.20952 | 0.00046 | -3e-06 | 100.9426 | 101.0689 | 124.4015 | 20.9606 | 0.001866 | 0.001761 | -151.6849 | 0.23721 | 8.4038 | 0.00347 | 0.004902 | 0.004719 | 6.3e-05 | 0.001022 | 75.1697 | 1.3083 | 2.548 | 99.62 | 20.96 | 37.213 | 66.4301 | 105.0696 | 63.9175 | 0.009278 | 1.0474 | 0.024969 | 23.8387 | -0.011181 | 0.14112 | 6e-06 | 0.000594 | 0.000594 | 9.7686
2019-03-21 11:38 | 1553182709.3532 | -404 | 0.20946 | 0.000458 | -3e-06 | 100.8541 | 101.1016 | 124.3168 | 20.9605 | 0.001871 | 0.001748 | -151.6822 | 0.23767 | 8.4025 | 0.003475 | 0.004901 | 0.004765 | 6.2e-05 | 0.001019 | 75.1219 | 1.2849 | 2.5479 | 99.62 | 20.96 | 37.212 | 66.544 | 105.2009 | 63.5887 | 0.009278 | 1.0484 | 0.024972 | 23.819 | -0.021919 | 0.14139 | 8e-06 | 0.00075 | 0.00075 | 7.7525
2019-03-21 11:38 | 1553182710.3522 | -403 | 0.2095 | 0.000463 | -3e-06 | 100.8997 | 101.2709 | 124.2219 | 20.9613 | 0.001876 | 0.001769 | -151.6561 | 0.23735 | 8.402 | 0.003459 | 0.00494 | 0.004857 | 6.2e-05 | 0.00102 | 75.1051 | 1.2851 | 2.548 | 99.62 | 20.96 | 37.211 | 66.6766 | 104.8106 | 64.0948 | 0.009277 | 1.0469 | 0.024944 | 23.8278 | -0.03284 | 0.1412 | 2e-06 | 0.000359 | 0.000359 | 16.1637
2019-03-21 11:38 | 1553182711.3536 | -402 | 0.20958 | 0.000464 | -3e-06 | 100.8652 | 101.0784 | 124.1363 | 20.961 | 0.001891 | 0.001769 | -151.6366 | 0.23716 | 8.401 | 0.003467 | 0.004654 | 0.004461 | 6.1e-05 | 0.001021 | 75.1368 | 1.288 | 2.5478 | 99.62 | 20.96 | 37.208 | 67.0385 | 104.9849 | 64.0905 | 0.009277 | 1.0469 | 0.024936 | 23.8194 | -0.018892 | 0.14109 | 0 | 0.000194 | 0.000194 | 29.9559
2019-03-21 11:39 | 1553182712.3522 | -401 | 0.20946 | 0.000459 | -3e-06 | 100.8699 | 101.2211 | 124.2684 | 20.9619 | 0.001821 | 0.001778 | -151.6148 | 0.23789 | 8.4034 | 0.003413 | 0.004857 | 0.004676 | 6.7e-05 | 0.00102 | 75.1772 | 1.262 | 2.5479 | 99.62 | 20.97 | 37.207 | 65.3538 | 103.6948 | 64.3125 | 0.009282 | 1.0462 | 0.024941 | 23.8399 | -0.031088 | 0.14152 | 7e-06 | 0.000728 | 0.000728 | 7.9913
2019-03-21 11:39 | 1553182713.3523 | -400 | 0.20951 | 0.000463 | -3e-06 | 100.9179 | 101.1857 | 124.7491 | 20.9606 | 0.00187 | 0.001773 | -151.6232 | 0.23731 | 8.4041 | 0.003473 | 0.00444 | 0.004307 | 6.4e-05 | 0.001011 | 75.149 | 1.3418 | 2.5479 | 99.62 | 20.97 | 37.205 | 66.5333 | 105.1481 | 64.2003 | 0.009282 | 1.0465 | 0.024993 | 23.882 | -0.023705 | 0.14118 | 1e-06 | 0.000236 | 0.000236 | 24.614
2019-03-21 11:39 | 1553182714.3528 | -399 | 0.20953 | 0.000463 | -2e-06 | 100.9081 | 101.1901 | 125.2019 | 20.9624 | 0.001947 | 0.001753 | -151.6121 | 0.23727 | 8.4028 | 0.003472 | 0.005427 | 0.005306 | 6.1e-05 | 0.001011 | 75.1903 | 1.32 | 2.548 | 99.62 | 20.97 | 37.205 | 68.3814 | 105.1288 | 63.7202 | 0.009282 | 1.048 | 0.025056 | 23.9083 | -0.024968 | 0.14115 | 1e-06 | 0.000253 | 0.000253 | 22.8854
2019-03-21 11:39 | 1553182715.3524 | -398 | 0.20951 | 0.000462 | -3e-06 | 100.8771 | 101.0805 | 125.2045 | 20.9625 | 0.001888 | 0.001798 | -151.624 | 0.23705 | 8.4023 | 0.00353 | 0.004773 | 0.004611 | 6.4e-05 | 0.001011 | 75.1774 | 1.3246 | 2.5478 | 99.62 | 20.96 | 37.204 | 66.9596 | 106.5302 | 64.8025 | 0.009276 | 1.0447 | 0.025016 | 23.9469 | -0.018016 | 0.14102 | 3e-06 | 0.000404 | 0.000404 | 14.3387
2019-03-21 11:39 | 1553182716.3535 | -397 | 0.20951 | 0.00046 | -2e-06 | 100.9262 | 101.1227 | 125.0022 | 20.9624 | 0.001936 | 0.001822 | -151.6059 | 0.2377 | 8.403 | 0.003481 | 0.004603 | 0.004346 | 6.5e-05 | 0.001012 | 75.1557 | 1.2836 | 2.5479 | 99.62 | 20.97 | 37.201 | 68.1249 | 105.3387 | 65.373 | 0.009281 | 1.0429 | 0.024975 | 23.9477 | -0.017394 | 0.14141 | 4e-06 | 0.000481 | 0.000481 | 12.0907
2019-03-21 11:39 | 1553182717.3525 | -396 | 0.20953 | 0.000462 | -3e-06 | 100.8422 | 101.1169 | 124.4567 | 20.9623 | 0.001863 | 0.001737 | -151.6007 | 0.23714 | 8.4018 | 0.003538 | 0.005086 | 0.004986 | 6.3e-05 | 0.001006 | 75.1446 | 1.2849 | 2.5479 | 99.62 | 20.96 | 37.2 | 66.3719 | 106.712 | 63.3259 | 0.009275 | 1.0492 | 0.024996 | 23.823 | -0.024331 | 0.14108 | 3e-06 | 0.000438 | 0.000438 | 13.2339
2019-03-21 11:39 | 1553182718.354 | -395 | 0.20952 | 0.000463 | -3e-06 | 100.8628 | 101.0297 | 124.3949 | 20.9615 | 0.001909 | 0.001775 | -151.5997 | 0.23716 | 8.4005 | 0.003562 | 0.004409 | 0.00422 | 6e-05 | 0.001001 | 75.1279 | 1.2931 | 2.548 | 99.62 | 20.96 | 37.198 | 67.4749 | 107.2978 | 64.2457 | 0.009274 | 1.0464 | 0.024956 | 23.8497 | -0.014792 | 0.14109 | 1e-06 | 0.000274 | 0.000274 | 21.1492
2019-03-21 11:39 | 1553182719.3528 | -394 | 0.20952 | 0.000458 | -2e-06 | 100.8589 | 100.9609 | 124.7247 | 20.9625 | 0.001891 | 0.001799 | -151.601 | 0.23705 | 8.4003 | 0.00358 | 0.005013 | 0.004851 | 6.2e-05 | 0.001003 | 75.1175 | 1.286 | 2.5478 | 99.62 | 20.96 | 37.196 | 67.0274 | 107.7469 | 64.8275 | 0.009274 | 1.0446 | 0.024967 | 23.9018 | -0.009043 | 0.14102 | 8e-06 | 0.00071 | 0.00071 | 8.1636
2019-03-21 11:39 | 1553182720.354 | -393 | 0.20953 | 0.000461 | -3e-06 | 100.892 | 101.0021 | 124.8287 | 20.9635 | 0.001913 | 0.001773 | -151.612 | 0.23593 | 8.4014 | 0.003537 | 0.004593 | 0.004467 | 6.2e-05 | 0.001004 | 75.1352 | 1.3079 | 2.548 | 99.62 | 20.96 | 37.196 | 67.572 | 106.6926 | 64.1937 | 0.009274 | 1.0465 | 0.025001 | 23.8894 | -0.00976 | 0.14035 | 4e-06 | 0.000495 | 0.000495 | 11.6617
2019-03-21 11:39 | 1553182721.3521 | -392 | 0.20951 | 0.000466 | -4e-06 | 100.864 | 101.0617 | 124.9001 | 20.9631 | 0.001877 | 0.001785 | -151.6374 | 0.23576 | 8.4041 | 0.003601 | 0.004782 | 0.004563 | 6e-05 | 0.001005 | 75.0968 | 1.3093 | 2.548 | 99.62 | 20.96 | 37.194 | 66.6978 | 108.2611 | 64.4783 | 0.009273 | 1.0457 | 0.024998 | 23.9063 | -0.01752 | 0.14025 | -1e-06 | 0.000158 | 0.000158 | 36.4473
2019-03-21 11:39 | 1553182722.3517 | -391 | 0.2095 | 0.000462 | -3e-06 | 100.8726 | 101.2211 | 124.7883 | 20.963 | 0.001907 | 0.001746 | -151.6677 | 0.23716 | 8.4032 | 0.003652 | 0.004572 | 0.004404 | 6.1e-05 | 0.001 | 75.0252 | 1.3645 | 2.5479 | 99.62 | 20.96 | 37.194 | 67.4311 | 109.4953 | 63.5576 | 0.009273 | 1.0485 | 0.025021 | 23.863 | -0.03085 | 0.14108 | 3e-06 | 0.000437 | 0.000437 | 13.2574
2019-03-21 11:39 | 1553182723.3527 | -390 | 0.2095 | 0.000458 | -3e-06 | 100.8839 | 101.1588 | 124.3752 | 20.9637 | 0.001875 | 0.001725 | -151.6699 | 0.23719 | 8.4023 | 0.003515 | 0.004947 | 0.004752 | 6.1e-05 | 0.000997 | 75.1036 | 1.3253 | 2.5479 | 99.62 | 20.96 | 37.194 | 66.6431 | 106.1569 | 63.0452 | 0.009273 | 1.0501 | 0.024998 | 23.8053 | -0.024337 | 0.14111 | 7e-06 | 0.000725 | 0.000725 | 8.0026
2019-03-21 11:39 | 1553182724.3529 | -389 | 0.2095 | 0.000457 | -3e-06 | 100.9346 | 101.1782 | 124.1867 | 20.9634 | 0.001821 | 0.001755 | -151.6832 | 0.23712 | 8.4011 | 0.003562 | 0.00571 | 0.00555 | 6.1e-05 | 0.000995 | 75.1196 | 1.3106 | 2.548 | 99.62 | 20.96 | 37.193 | 65.3647 | 107.2927 | 63.7612 | 0.009273 | 1.0479 | 0.024953 | 23.8126 | -0.021564 | 0.14106 | 9e-06 | 0.000853 | 0.000853 | 6.7959
2019-03-21 11:39 | 1553182725.3523 | -388 | 0.20951 | 0.000464 | -3e-06 | 100.8935 | 101.0089 | 124.1867 | 20.9642 | 0.00186 | 0.001761 | -151.6838 | 0.23685 | 8.4027 | 0.003625 | 0.005163 | 0.004982 | 6.2e-05 | 0.000999 | 75.1634 | 1.3053 | 2.5479 | 99.62 | 20.96 | 37.193 | 66.2941 | 108.8397 | 63.9189 | 0.009273 | 1.0474 | 0.024947 | 23.8182 | -0.010232 | 0.1409 | 0 | 0.00021 | 0.00021 | 27.5807
2019-03-21 11:39 | 1553182726.352 | -387 | 0.20952 | 0.000462 | -2e-06 | 100.8684 | 101.1064 | 124.1461 | 20.9642 | 0.001943 | 0.00179 | -151.7181 | 0.2366 | 8.4032 | 0.003605 | 0.005339 | 0.00512 | 6.1e-05 | 0.000997 | 75.1447 | 1.3254 | 2.5479 | 99.62 | 20.96 | 37.193 | 68.2826 | 108.3449 | 64.6152 | 0.009273 | 1.0452 | 0.024917 | 23.8389 | -0.021072 | 0.14075 | 1e-06 | 0.000221 | 0.000221 | 26.1366
2019-03-21 11:39 | 1553182727.3521 | -386 | 0.20952 | 0.000466 | -3e-06 | 100.8869 | 101.0246 | 124.386 | 20.9644 | 0.001838 | 0.001748 | -151.7275 | 0.2357 | 8.4009 | 0.00362 | 0.005315 | 0.005089 | 6.2e-05 | 0.000999 | 75.1044 | 1.3134 | 2.548 | 99.62 | 20.96 | 37.194 | 65.7707 | 108.7211 | 63.6059 | 0.009273 | 1.0484 | 0.024979 | 23.8262 | -0.0122 | 0.14022 | -2e-06 | 7.9e-05 | 7.9e-05 | 72.8329
2019-03-21 11:39 | 1553182728.3522 | -385 | 0.20953 | 0.000459 | -3e-06 | 100.9051 | 101.0519 | 124.382 | 20.9634 | 0.001876 | 0.001743 | -151.7412 | 0.23685 | 8.4024 | 0.003652 | 0.004657 | 0.004571 | 6e-05 | 0.000999 | 75.0988 | 1.2995 | 2.5479 | 99.62 | 20.96 | 37.194 | 66.6724 | 109.4799 | 63.4728 | 0.009273 | 1.0488 | 0.024983 | 23.8211 | -0.013001 | 0.1409 | 7e-06 | 0.000667 | 0.000667 | 8.6779
2019-03-21 11:39 | 1553182729.3526 | -384 | 0.20952 | 0.00046 | -3e-06 | 100.895 | 101.1173 | 124.055 | 20.9648 | 0.001859 | 0.001733 | -151.7418 | 0.23725 | 8.4029 | 0.003582 | 0.004757 | 0.004532 | 6.3e-05 | 0.001001 | 75.1621 | 1.317 | 2.5479 | 99.62 | 20.96 | 37.195 | 66.2766 | 107.7892 | 63.2486 | 0.009273 | 1.0495 | 0.024959 | 23.7818 | -0.019685 | 0.14114 | 5e-06 | 0.000537 | 0.000537 | 10.8099
2019-03-21 11:39 | 1553182730.3522 | -383 | 0.20953 | 0.000462 | -3e-06 | 100.8815 | 100.9987 | 124.171 | 20.9646 | 0.001886 | 0.001741 | -151.7571 | 0.23704 | 8.4034 | 0.003676 | 0.00554 | 0.00538 | 6.3e-05 | 0.001003 | 75.0373 | 1.3401 | 2.548 | 99.62 | 20.96 | 37.195 | 66.9193 | 110.0758 | 63.4201 | 0.009273 | 1.049 | 0.024964 | 23.799 | -0.010383 | 0.14101 | 3e-06 | 0.000436 | 0.000436 | 13.2879
2019-03-21 11:39 | 1553182731.3533 | -382 | 0.2095 | 0.000463 | -2e-06 | 100.8595 | 101.0369 | 124.3705 | 20.9642 | 0.001872 | 0.001762 | -151.7472 | 0.23697 | 8.4037 | 0.003634 | 0.004935 | 0.004743 | 6.3e-05 | 0.000996 | 75.0891 | 1.2911 | 2.5478 | 99.62 | 20.96 | 37.193 | 66.5865 | 109.0486 | 63.9241 | 0.009273 | 1.0474 | 0.024965 | 23.836 | -0.015716 | 0.14097 | 1e-06 | 0.000226 | 0.000226 | 25.5885
2019-03-21 11:39 | 1553182732.3523 | -381 | 0.20954 | 0.000459 | -3e-06 | 100.8226 | 101.0679 | 124.5243 | 20.964 | 0.001858 | 0.001782 | -151.7205 | 0.23652 | 8.4045 | 0.003552 | 0.006246 | 0.006097 | 6.4e-05 | 0.000987 | 75.1693 | 1.3216 | 2.5479 | 99.62 | 20.96 | 37.193 | 66.2564 | 107.0607 | 64.4151 | 0.009273 | 1.0459 | 0.024963 | 23.8681 | -0.021736 | 0.1407 | 6e-06 | 0.000644 | 0.000644 | 8.9845
2019-03-21 11:39 | 1553182733.3534 | -380 | 0.20949 | 0.000465 | -3e-06 | 100.8976 | 100.9694 | 124.4402 | 20.9652 | 0.001912 | 0.001757 | -151.6917 | 0.23723 | 8.402 | 0.003567 | 0.004486 | 0.00443 | 6.2e-05 | 0.000985 | 75.0732 | 1.3258 | 2.5478 | 99.62 | 20.96 | 37.193 | 67.5542 | 107.4194 | 63.8113 | 0.009273 | 1.0477 | 0.024977 | 23.8387 | -0.006362 | 0.14113 | 0 | 0.000167 | 0.000167 | 34.7499
2019-03-21 11:39 | 1553182734.3527 | -379 | 0.20949 | 0.000464 | -3e-06 | 100.8926 | 101.0508 | 124.2891 | 20.9651 | 0.001904 | 0.001764 | -151.6757 | 0.23677 | 8.4019 | 0.003562 | 0.004667 | 0.00441 | 6e-05 | 0.000988 | 75.1139 | 1.2938 | 2.5479 | 99.62 | 20.96 | 37.192 | 67.356 | 107.2943 | 63.9775 | 0.009273 | 1.0472 | 0.024955 | 23.83 | -0.01402 | 0.14086 | 0 | 0.000165 | 0.000165 | 34.9937
2019-03-21 11:39 | 1553182735.3525 | -378 | 0.20949 | 0.000465 | -3e-06 | 100.8139 | 100.9885 | 124.432 | 20.9651 | 0.001914 | 0.001775 | -151.6831 | 0.23704 | 8.4034 | 0.003567 | 0.004893 | 0.004711 | 6.2e-05 | 0.00099 | 75.1754 | 1.271 | 2.5479 | 99.62 | 20.96 | 37.191 | 67.6002 | 107.4224 | 64.2381 | 0.009272 | 1.0464 | 0.02496 | 23.853 | -0.015477 | 0.14101 | -1e-06 | 8.3e-05 | 8.3e-05 | 70.0437
2019-03-21 11:39 | 1553182736.353 | -377 | 0.20947 | 0.000463 | -3e-06 | 100.8869 | 101.1578 | 124.8465 | 20.966 | 0.001897 | 0.001821 | -151.6599 | 0.23732 | 8.4004 | 0.003575 | 0.004771 | 0.004632 | 6e-05 | 0.000988 | 75.1417 | 1.2933 | 2.5478 | 99.62 | 20.96 | 37.19 | 67.1774 | 107.6131 | 65.3686 | 0.009272 | 1.0429 | 0.02496 | 23.9326 | -0.023983 | 0.14118 | 2e-06 | 0.000316 | 0.000316 | 18.3864
2019-03-21 11:39 | 1553182737.3515 | -376 | 0.20946 | 0.000464 | -3e-06 | 100.8944 | 101.1697 | 124.9888 | 20.9665 | 0.001886 | 0.001755 | -151.6337 | 0.23691 | 8.4006 | 0.003632 | 0.005126 | 0.004997 | 6.6e-05 | 0.000989 | 75.1884 | 1.3094 | 2.548 | 99.62 | 20.96 | 37.189 | 66.9141 | 108.9964 | 63.761 | 0.009272 | 1.0479 | 0.025033 | 23.8894 | -0.024376 | 0.14094 | 1e-06 | 0.000287 | 0.000287 | 20.2118
2019-03-21 11:39 | 1553182738.3526 | -375 | 0.20935 | 0.000462 | -3e-06 | 100.8419 | 101.1346 | 124.9824 | 20.9657 | 0.00189 | 0.001767 | -151.6145 | 0.23653 | 8.4027 | 0.003626 | 0.005155 | 0.004946 | 6.2e-05 | 0.000991 | 75.1686 | 1.3032 | 2.5479 | 99.62 | 20.96 | 37.189 | 67.0163 | 108.8601 | 64.0553 | 0.009272 | 1.047 | 0.025022 | 23.8992 | -0.025924 | 0.14071 | 2e-06 | 0.000329 | 0.000329 | 17.5664
2019-03-21 11:39 | 1553182739.3526 | -374 | 0.20947 | 0.000457 | -2e-06 | 100.8872 | 101.0593 | 124.7644 | 20.9669 | 0.001849 | 0.001782 | -151.621 | 0.23736 | 8.4031 | 0.003652 | 0.005472 | 0.005301 | 6.3e-05 | 0.000993 | 75.139 | 1.3102 | 2.5479 | 99.62 | 20.96 | 37.189 | 66.0399 | 109.4984 | 64.4266 | 0.009272 | 1.0458 | 0.024986 | 23.8915 | -0.015249 | 0.1412 | 9e-06 | 0.000839 | 0.000839 | 6.9208
2019-03-21 11:39 | 1553182740.3529 | -373 | 0.20949 | 0.00046 | -3e-06 | 100.8571 | 100.9745 | 124.5578 | 20.9661 | 0.001922 | 0.001787 | -151.6175 | 0.23744 | 8.4013 | 0.003636 | 0.004794 | 0.0046 | 6.5e-05 | 0.000997 | 75.1268 | 1.3188 | 2.548 | 99.62 | 20.96 | 37.188 | 67.7813 | 109.1088 | 64.5295 | 0.009272 | 1.0455 | 0.024962 | 23.8753 | -0.010408 | 0.14125 | 5e-06 | 0.000544 | 0.000544 | 10.6738
2019-03-21 11:39 | 1553182741.3519 | -372 | 0.20955 | 0.000464 | -3e-06 | 100.9256 | 101.2273 | 124.3064 | 20.9665 | 0.001908 | 0.001748 | -151.6041 | 0.23764 | 8.4023 | 0.003661 | 0.005047 | 0.004824 | 6.3e-05 | 0.000996 | 75.1496 | 1.2923 | 2.5479 | 99.62 | 20.96 | 37.188 | 67.4383 | 109.7133 | 63.6007 | 0.009272 | 1.0484 | 0.024971 | 23.8184 | -0.026691 | 0.14137 | 0 | 0.000187 | 0.000187 | 31.1447
2019-03-21 11:39 | 1553182742.3523 | -371 | 0.2094 | 0.000465 | -3e-06 | 100.9104 | 101.1173 | 124.2705 | 20.9667 | 0.0019 | 0.001778 | -151.6339 | 0.23762 | 8.4014 | 0.003663 | 0.004827 | 0.004659 | 6.7e-05 | 0.000994 | 75.1528 | 1.2977 | 2.5479 | 99.62 | 20.96 | 37.189 | 67.2693 | 109.7702 | 64.3117 | 0.009272 | 1.0462 | 0.024941 | 23.8401 | -0.018312 | 0.14136 | -1e-06 | 7e-05 | 7e-05 | 83.3686
2019-03-21 11:39 | 1553182743.353 | -370 | 0.20962 | 0.000459 | -3e-06 | 100.8824 | 101.1091 | 124.5228 | 20.9668 | 0.001937 | 0.001755 | -151.648 | 0.2369 | 8.3992 | 0.003669 | 0.004949 | 0.004781 | 6.4e-05 | 0.000995 | 75.1501 | 1.324 | 2.548 | 99.62 | 20.96 | 37.189 | 68.1448 | 109.9012 | 63.7673 | 0.009272 | 1.0479 | 0.024986 | 23.845 | -0.020071 | 0.14093 | 7e-06 | 0.000663 | 0.000663 | 8.741
2019-03-21 11:39 | 1553182744.3534 | -369 | 0.2096 | 0.000466 | -2e-06 | 100.8667 | 100.919 | 124.5961 | 20.9669 | 0.001911 | 0.001756 | -151.6546 | 0.23705 | 8.4017 | 0.003614 | 0.004559 | 0.004411 | 6.1e-05 | 0.000992 | 75.1219 | 1.2832 | 2.5478 | 99.62 | 20.96 | 37.19 | 67.5239 | 108.573 | 63.7973 | 0.009272 | 1.0478 | 0.024993 | 23.8531 | -0.004643 | 0.14102 | -3e-06 | -0.000111 | -0.000111 | -52.1811
2019-03-21 11:39 | 1553182745.3525 | -368 | 0.20952 | 0.000462 | -3e-06 | 100.9239 | 101.0835 | 124.3257 | 20.967 | 0.001855 | 0.001745 | -151.6341 | 0.23753 | 8.399 | 0.003602 | 0.005147 | 0.004949 | 6e-05 | 0.000996 | 75.1698 | 1.2935 | 2.548 | 99.62 | 20.96 | 37.19 | 66.1773 | 108.2792 | 63.5327 | 0.009272 | 1.0486 | 0.024975 | 23.8178 | -0.01414 | 0.14131 | 4e-06 | 0.000471 | 0.000471 | 12.3407
2019-03-21 11:39 | 1553182746.3522 | -367 | 0.20952 | 0.000463 | -3e-06 | 100.9054 | 101.2143 | 123.8133 | 20.9678 | 0.001922 | 0.001822 | -151.6444 | 0.23697 | 8.4018 | 0.003567 | 0.005314 | 0.005116 | 6.4e-05 | 0.001 | 75.151 | 1.2945 | 2.5479 | 99.62 | 20.96 | 37.189 | 67.7798 | 107.4377 | 65.3797 | 0.009272 | 1.0429 | 0.024856 | 23.8338 | -0.027343 | 0.14098 | 1e-06 | 0.000232 | 0.000232 | 24.9533
2019-03-21 11:39 | 1553182747.3531 | -366 | 0.20954 | 0.00046 | -3e-06 | 100.9 | 101.1656 | 123.718 | 20.9675 | 0.001889 | 0.001734 | -151.6819 | 0.23724 | 8.4022 | 0.003535 | 0.005178 | 0.005103 | 6.4e-05 | 0.00099 | 75.149 | 1.3028 | 2.548 | 99.62 | 20.96 | 37.19 | 66.9996 | 106.6422 | 63.2614 | 0.009272 | 1.0494 | 0.024924 | 23.75 | -0.023513 | 0.14113 | 4e-06 | 0.000487 | 0.000487 | 11.9
2019-03-21 11:39 | 1553182748.3523 | -365 | 0.20955 | 0.000462 | -3e-06 | 100.8839 | 101.1431 | 123.9572 | 20.968 | 0.001923 | 0.001751 | -151.684 | 0.23666 | 8.4029 | 0.003575 | 0.005096 | 0.004924 | 6.2e-05 | 0.000993 | 75.2052 | 1.3109 | 2.5481 | 99.62 | 20.96 | 37.19 | 67.8096 | 107.6326 | 63.6788 | 0.009272 | 1.0481 | 0.024933 | 23.7877 | -0.022952 | 0.14079 | 3e-06 | 0.000362 | 0.000362 | 15.9889
2019-03-21 11:39 | 1553182749.3532 | -364 | 0.20952 | 0.000459 | -3e-06 | 100.8944 | 101.0164 | 124.0917 | 20.9676 | 0.001914 | 0.001694 | -151.6735 | 0.2367 | 8.402 | 0.003547 | 0.00531 | 0.005221 | 6.2e-05 | 0.000994 | 75.1395 | 1.341 | 2.5479 | 99.62 | 20.95 | 37.188 | 67.5854 | 106.9495 | 62.3093 | 0.009266 | 1.0524 | 0.024997 | 23.7521 | -0.010816 | 0.14081 | 7e-06 | 0.000675 | 0.000675 | 8.576
2019-03-21 11:39 | 1553182750.3524 | -363 | 0.20949 | 0.000467 | -3e-06 | 100.8929 | 100.9718 | 123.906 | 20.9684 | 0.001941 | 0.001788 | -151.6725 | 0.23755 | 8.4008 | 0.003521 | 0.004749 | 0.004501 | 6e-05 | 0.000997 | 75.1052 | 1.3149 | 2.548 | 99.62 | 20.95 | 37.188 | 68.2384 | 106.3045 | 64.5772 | 0.009266 | 1.0454 | 0.024895 | 23.8144 | -0.006996 | 0.14132 | -4e-06 | -9.5e-05 | -9.5e-05 | -60.9982
2019-03-21 11:39 | 1553182751.353 | -362 | 0.20952 | 0.000462 | -3e-06 | 100.8973 | 101.1649 | 123.9478 | 20.9683 | 0.001894 | 0.001791 | -151.6801 | 0.237 | 8.4025 | 0.003571 | 0.005937 | 0.005759 | 6.2e-05 | 0.001 | 75.1584 | 1.311 | 2.5479 | 99.62 | 20.95 | 37.188 | 67.1266 | 107.5203 | 64.6269 | 0.009266 | 1.0452 | 0.024897 | 23.8202 | -0.023691 | 0.14099 | -8.3e-05 | -0.005998 | -0.005998 | -0.96618
2019-03-21 11:39 | 1553182752.3529 | -361 | 0.20949 | 0.000458 | -3e-06 | 100.8658 | 101.0399 | 123.7243 | 20.9677 | 0.001839 | 0.001726 | -151.6632 | 0.23736 | 8.4006 | 0.003654 | 0.00494 | 0.004794 | 6.3e-05 | 0.001002 | 75.1103 | 1.2839 | 2.5479 | 99.62 | 20.95 | 37.187 | 65.8011 | 109.5322 | 63.0726 | 0.009266 | 1.05 | 0.024932 | 23.7439 | -0.015432 | 0.14121 | 8.3e-05 | 0.006394 | 0.006394 | 0.90773
2019-03-21 11:39 | 1553182753.352 | -360 | 0.20948 | 0.00046 | -3e-06 | 100.8667 | 101.1258 | 123.7717 | 20.9687 | 0.001909 | 0.001773 | -151.6624 | 0.23694 | 8.4034 | 0.003601 | 0.004513 | 0.004315 | 6.3e-05 | 0.001006 | 75.1806 | 1.3641 | 2.5479 | 99.62 | 20.95 | 37.187 | 67.4819 | 108.2539 | 64.2168 | 0.009266 | 1.0465 | 0.024894 | 23.7888 | -0.022948 | 0.14096 | 0.000138 | 0.010508 | 0.010508 | 0.55134
2019-03-21 11:39 | 1553182754.3526 | -359 | 0.20946 | 0.000464 | -3e-06 | 100.8214 | 101.072 | 123.8961 | 20.9682 | 0.00187 | 0.001811 | -151.6594 | 0.23738 | 8.4028 | 0.003615 | 0.004694 | 0.004499 | 6.3e-05 | 0.001005 | 75.1228 | 1.3462 | 2.548 | 99.62 | 20.95 | 37.186 | 66.5303 | 108.583 | 65.1123 | 0.009266 | 1.0437 | 0.024874 | 23.8324 | -0.022204 | 0.14122 | 0.000259 | 0.019521 | 0.019521 | 0.29732
2019-03-21 11:39 | 1553182755.3521 | -358 | 0.2095 | 0.000462 | -3e-06 | 100.8181 | 101.0689 | 123.971 | 20.9684 | 0.001935 | 0.001745 | -151.6549 | 0.23747 | 8.4032 | 0.003506 | 0.005401 | 0.005254 | 5.9e-05 | 0.001007 | 75.1829 | 1.3098 | 2.5479 | 99.62 | 20.95 | 37.187 | 68.1033 | 105.945 | 63.5323 | 0.009266 | 1.0486 | 0.02494 | 23.7838 | -0.022223 | 0.14127 | 1.1e-05 | 0.000978 | 0.000978 | 5.9339
2019-03-21 11:39 | 1553182756.3517 | -357 | 0.20951 | 0.000457 | -3e-06 | 100.8726 | 100.9977 | 124.1191 | 20.969 | 0.00194 | 0.001783 | -151.6563 | 0.23719 | 8.4029 | 0.003509 | 0.004394 | 0.004104 | 6.2e-05 | 0.001008 | 75.2303 | 1.3022 | 2.548 | 99.62 | 20.95 | 37.188 | 68.2297 | 106.0134 | 64.4489 | 0.009266 | 1.0458 | 0.024921 | 23.8304 | -0.011085 | 0.1411 | 0 | 0.000182 | 0.000182 | 31.9181
2019-03-21 11:39 | 1553182757.3531 | -356 | 0.20947 | 0.000461 | -3e-06 | 100.8419 | 100.9299 | 124.1029 | 20.97 | 0.00193 | 0.001765 | -151.6327 | 0.23676 | 8.4018 | 0.003554 | 0.004665 | 0.004534 | 5.9e-05 | 0.001011 | 75.2102 | 1.3236 | 2.548 | 99.62 | 20.95 | 37.189 | 67.9825 | 107.1114 | 64.0181 | 0.009266 | 1.0471 | 0.024935 | 23.8136 | -0.007802 | 0.14085 | 0.000193 | 0.01465 | 0.01465 | 0.39515
2019-03-21 11:39 | 1553182758.3519 | -355 | 0.2095 | 0.000461 | -3e-06 | 100.8354 | 100.9054 | 123.641 | 20.9704 | 0.001903 | 0.001738 | -151.6316 | 0.23708 | 8.4031 | 0.003539 | 0.00514 | 0.00495 | 6.1e-05 | 0.001007 | 75.1148 | 1.3341 | 2.548 | 99.62 | 20.94 | 37.191 | 67.3392 | 106.7427 | 63.3672 | 0.009261 | 1.0491 | 0.024913 | 23.7463 | -0.00621 | 0.14104 | 5.2e-05 | 0.004098 | 0.004098 | 1.4145
2019-03-21 11:39 | 1553182759.3518 | -354 | 0.20947 | 0.000464 | -4e-06 | 100.7528 | 100.8679 | 123.6114 | 20.9699 | 0.001877 | 0.001776 | -151.6444 | 0.23748 | 8.4041 | 0.003594 | 0.005095 | 0.004961 | 6.3e-05 | 0.001011 | 75.1666 | 1.3294 | 2.548 | 99.62 | 20.94 | 37.194 | 66.7013 | 108.0936 | 64.2748 | 0.009262 | 1.0463 | 0.024876 | 23.7755 | -0.010213 | 0.14128 | 0.000223 | 0.016868 | 0.016868 | 0.34423
2019-03-21 11:39 | 1553182760.3523 | -353 | 0.20952 | 0.000462 | -3e-06 | 100.7728 | 101.0498 | 123.7361 | 20.9701 | 0.001908 | 0.001768 | -151.6464 | 0.23717 | 8.4032 | 0.003653 | 0.004912 | 0.004742 | 6e-05 | 0.001009 | 75.2235 | 1.3016 | 2.548 | 99.62 | 20.94 | 37.195 | 67.4522 | 109.5195 | 64.0767 | 0.009262 | 1.0469 | 0.024896 | 23.7805 | -0.024554 | 0.1411 | -7.3e-05 | -0.005222 | -0.005222 | -1.1104
2019-03-21 11:39 | 1553182761.352 | -352 | 0.20949 | 0.000462 | -3e-06 | 100.7382 | 100.9095 | 123.9156 | 20.9707 | 0.001887 | 0.001794 | -151.6683 | 0.23705 | 8.3989 | 0.003585 | 0.005218 | 0.005029 | 6.1e-05 | 0.001011 | 75.2246 | 1.3313 | 2.5478 | 99.62 | 20.94 | 37.198 | 66.9572 | 107.862 | 64.7019 | 0.009263 | 1.045 | 0.024891 | 23.8198 | -0.015193 | 0.14102 | 7e-05 | 0.005426 | 0.005426 | 1.0683
2019-03-21 11:39 | 1553182762.353 | -351 | 0.2095 | 0.000461 | -3e-06 | 100.7835 | 100.9681 | 124.1928 | 20.9711 | 0.001885 | 0.001752 | -151.6751 | 0.23741 | 8.4021 | 0.003547 | 0.004778 | 0.004544 | 6.5e-05 | 0.001016 | 75.144 | 1.3214 | 2.5479 | 99.62 | 20.94 | 37.201 | 66.9109 | 106.9575 | 63.7085 | 0.009264 | 1.0481 | 0.024956 | 23.8113 | -0.016364 | 0.14123 | 9e-06 | 0.000886 | 0.000886 | 6.5495
2019-03-21 11:39 | 1553182763.3535 | -350 | 0.2095 | 0.000464 | -3e-06 | 100.7651 | 100.9074 | 124.0595 | 20.9718 | 0.001837 | 0.001817 | -151.681 | 0.23757 | 8.404 | 0.003634 | 0.004572 | 0.004485 | 6.2e-05 | 0.001011 | 75.1706 | 1.319 | 2.5478 | 99.62 | 20.94 | 37.203 | 65.7543 | 109.0578 | 65.2744 | 0.009264 | 1.0432 | 0.024884 | 23.8538 | -0.012631 | 0.14133 | 6.3e-05 | 0.004886 | 0.004886 | 1.1888
2019-03-21 11:39 | 1553182764.3521 | -349 | 0.20954 | 0.000462 | -3e-06 | 100.7785 | 101.0273 | 124.0928 | 20.9716 | 0.001892 | 0.001726 | -151.6923 | 0.23705 | 8.4017 | 0.003588 | 0.00484 | 0.004564 | 6.3e-05 | 0.001012 | 75.1691 | 1.3212 | 2.5478 | 99.62 | 20.93 | 37.209 | 67.0682 | 107.9345 | 63.0872 | 0.00926 | 1.05 | 0.024968 | 23.7798 | -0.022061 | 0.14102 | -0.000216 | -0.016033 | -0.016033 | -0.36149
2019-03-21 11:39 | 1553182765.3532 | -348 | 0.20947 | 0.000463 | -3e-06 | 100.7552 | 100.8404 | 123.9762 | 20.9709 | 0.001907 | 0.001734 | -151.6786 | 0.23629 | 8.403 | 0.003671 | 0.005633 | 0.005408 | 6e-05 | 0.001011 | 75.2416 | 1.3329 | 2.548 | 99.62 | 20.93 | 37.212 | 67.4373 | 109.9619 | 63.2748 | 0.009261 | 1.0494 | 0.02495 | 23.7752 | -0.007554 | 0.14057 | 0.000211 | 0.016006 | 0.016006 | 0.36095
2019-03-21 11:39 | 1553182766.3526 | -347 | 0.20949 | 0.000461 | -4e-06 | 100.7761 | 100.9487 | 123.7391 | 20.9717 | 0.001892 | 0.001784 | -151.6657 | 0.23676 | 8.4026 | 0.003609 | 0.004752 | 0.004647 | 6.3e-05 | 0.001004 | 75.2207 | 1.3231 | 2.548 | 99.62 | 20.93 | 37.214 | 67.0753 | 108.4601 | 64.4647 | 0.009261 | 1.0457 | 0.024882 | 23.7944 | -0.015305 | 0.14085 | 0.000137 | 0.010468 | 0.010468 | 0.55298
2019-03-21 11:39 | 1553182767.3517 | -346 | 0.20948 | 0.000461 | -3e-06 | 100.7835 | 100.8945 | 123.694 | 20.9719 | 0.001853 | 0.001726 | -151.6854 | 0.23683 | 8.4008 | 0.00364 | 0.00521 | 0.005056 | 6.1e-05 | 0.001002 | 75.2534 | 1.3129 | 2.548 | 99.62 | 20.93 | 37.219 | 66.1228 | 109.2163 | 63.0687 | 0.009262 | 1.05 | 0.024929 | 23.7409 | -0.009844 | 0.14089 | 0.000149 | 0.011374 | 0.011374 | 0.50909
2019-03-21 11:39 | 1553182768.3524 | -345 | 0.20949 | 0.000457 | -3e-06 | 100.7692 | 101.0198 | 123.9274 | 20.9717 | 0.001875 | 0.001742 | -151.6864 | 0.23716 | 8.4019 | 0.003631 | 0.00519 | 0.004997 | 6.2e-05 | 0.001 | 75.2318 | 1.2902 | 2.5481 | 99.62 | 20.93 | 37.222 | 66.6588 | 108.9836 | 63.4537 | 0.009263 | 1.0488 | 0.024938 | 23.7768 | -0.022218 | 0.14109 | 0.000134 | 0.010198 | 0.010198 | 0.5686
2019-03-21 11:39 | 1553182769.3524 | -344 | 0.20948 | 0.000461 | -3e-06 | 100.6521 | 100.981 | 124.1691 | 20.9715 | 0.00193 | 0.001735 | -151.6713 | 0.23714 | 8.4035 | 0.003537 | 0.004245 | 0.004046 | 6.3e-05 | 0.000996 | 75.1435 | 1.2941 | 2.548 | 99.61 | 20.93 | 37.228 | 67.9894 | 106.7021 | 63.2924 | 0.009265 | 1.0493 | 0.024969 | 23.7943 | -0.029179 | 0.14108 | 0.000155 | 0.011816 | 0.011816 | 0.49072
2019-03-21 11:39 | 1553182770.3519 | -343 | 0.20949 | 0.00046 | -3e-06 | 100.6745 | 100.8513 | 124.201 | 20.9724 | 0.001889 | 0.001749 | -151.6678 | 0.23688 | 8.4021 | 0.003652 | 0.004828 | 0.004682 | 6.1e-05 | 0.001001 | 75.1855 | 1.2818 | 2.5479 | 99.61 | 20.93 | 37.231 | 67.0096 | 109.5014 | 63.6238 | 0.009266 | 1.0483 | 0.024959 | 23.8091 | -0.015692 | 0.14092 | 0.000128 | 0.009784 | 0.009784 | 0.59196
2019-03-21 11:40 | 1553182771.353 | -342 | 0.20949 | 0.000458 | -3e-06 | 100.7269 | 100.9936 | 123.7527 | 20.9717 | 0.001856 | 0.001755 | -151.6538 | 0.23767 | 8.4019 | 0.003681 | 0.00483 | 0.004644 | 6.2e-05 | 0.000998 | 75.2262 | 1.2797 | 2.5479 | 99.61 | 20.93 | 37.233 | 66.2093 | 110.2121 | 63.7644 | 0.009267 | 1.0479 | 0.024909 | 23.7711 | -0.023652 | 0.14139 | 0.000126 | 0.00961 | 0.00961 | 0.60469
2019-03-21 11:40 | 1553182772.3526 | -341 | 0.20948 | 0.000459 | -3e-06 | 100.6447 | 100.6094 | 123.5436 | 20.972 | 0.001885 | 0.001732 | -151.6497 | 0.2372 | 8.4035 | 0.003563 | 0.004808 | 0.004664 | 6.5e-05 | 0.001001 | 75.212 | 1.3167 | 2.5479 | 99.61 | 20.93 | 37.238 | 66.9111 | 107.3246 | 63.2157 | 0.009268 | 1.0496 | 0.024908 | 23.7316 | 0.003133 | 0.14111 | 0.000153 | 0.011656 | 0.011656 | 0.49757
2019-03-21 11:40 | 1553182773.3526 | -340 | 0.20951 | 0.000461 | -3e-06 | 100.5997 | 100.8172 | 123.8236 | 20.9721 | 0.001887 | 0.001842 | -151.6673 | 0.23679 | 8.4017 | 0.00358 | 0.004764 | 0.004555 | 5.8e-05 | 0.001001 | 75.2427 | 1.3098 | 2.548 | 99.61 | 20.92 | 37.24 | 66.9515 | 107.7373 | 65.8755 | 0.009263 | 1.0414 | 0.024839 | 23.8522 | -0.019317 | 0.14087 | -2.9e-05 | -0.00199 | -0.00199 | -2.9093
2019-03-21 11:40 | 1553182774.3516 | -339 | 0.20949 | 0.000463 | -3e-06 | 100.749 | 100.7695 | 124.0856 | 20.9724 | 0.001888 | 0.001768 | -151.655 | 0.23703 | 8.4036 | 0.003619 | 0.004913 | 0.004833 | 6e-05 | 0.001 | 75.1821 | 1.317 | 2.5479 | 99.61 | 20.92 | 37.243 | 66.97 | 108.6887 | 64.0932 | 0.009263 | 1.0469 | 0.02493 | 23.8146 | -0.001824 | 0.14101 | 8.7e-05 | 0.006701 | 0.006701 | 0.86487
2019-03-21 11:40 | 1553182775.3518 | -338 | 0.20952 | 0.00047 | -4e-06 | 100.6909 | 100.8094 | 124.198 | 20.9713 | 0.001821 | 0.00179 | -151.6792 | 0.23736 | 8.4012 | 0.003578 | 0.005085 | 0.00487 | 5.9e-05 | 0.001006 | 75.1988 | 1.3288 | 2.5479 | 99.61 | 20.92 | 37.244 | 65.3643 | 107.7113 | 64.6089 | 0.009264 | 1.0453 | 0.024923 | 23.8436 | -0.01052 | 0.1412 | -0.000116 | -0.008483 | -0.008483 | -0.68412
2019-03-21 11:40 | 1553182776.3518 | -337 | 0.20951 | 0.000459 | -3e-06 | 100.7084 | 101.0287 | 124.3482 | 20.9726 | 0.001938 | 0.001743 | -151.6878 | 0.23677 | 8.4015 | 0.003624 | 0.005197 | 0.004951 | 6.1e-05 | 0.001004 | 75.2188 | 1.3104 | 2.5479 | 99.61 | 20.92 | 37.246 | 68.173 | 108.8188 | 63.488 | 0.009264 | 1.0487 | 0.024979 | 23.8184 | -0.028399 | 0.14085 | -2.9e-05 | -0.002015 | -0.002015 | -2.8735
2019-03-21 11:40 | 1553182777.3528 | -336 | 0.2095 | 0.000458 | -4e-06 | 100.6024 | 100.7273 | 124.3945 | 20.9732 | 0.001893 | 0.001785 | -151.678 | 0.23692 | 8.4013 | 0.003597 | 0.005378 | 0.005227 | 6.1e-05 | 0.001006 | 75.1796 | 1.3221 | 2.548 | 99.61 | 20.92 | 37.249 | 67.0955 | 108.1684 | 64.5029 | 0.009265 | 1.0456 | 0.024946 | 23.8587 | -0.011098 | 0.14094 | 5.9e-05 | 0.004641 | 0.004641 | 1.2481
2019-03-21 11:40 | 1553182778.3523 | -335 | 0.2095 | 0.00046 | -3e-06 | 100.6283 | 100.7031 | 124.1794 | 20.9727 | 0.001953 | 0.001806 | -151.6894 | 0.23678 | 8.4018 | 0.003613 | 0.005127 | 0.004957 | 6.2e-05 | 0.001009 | 75.2025 | 1.3525 | 2.5479 | 99.61 | 20.92 | 37.251 | 68.5398 | 108.5408 | 64.9936 | 0.009265 | 1.0441 | 0.024907 | 23.8554 | -0.006646 | 0.14086 | 6.2e-05 | 0.004809 | 0.004809 | 1.2039
2019-03-21 11:40 | 1553182779.352 | -334 | 0.20951 | 0.000461 | -3e-06 | 100.5496 | 100.8271 | 123.9316 | 20.9727 | 0.001893 | 0.001755 | -151.6816 | 0.23774 | 8.4012 | 0.003561 | 0.005468 | 0.005296 | 6.1e-05 | 0.001011 | 75.2111 | 1.3545 | 2.5478 | 99.61 | 20.92 | 37.256 | 67.0873 | 107.286 | 63.7794 | 0.009267 | 1.0478 | 0.024927 | 23.7888 | -0.024645 | 0.14143 | -7e-06 | -0.000278 | -0.000278 | -20.9434
2019-03-21 11:40 | 1553182780.353 | -333 | 0.20954 | 0.000461 | -3e-06 | 100.5425 | 100.6895 | 124.0297 | 20.974 | 0.001866 | 0.001784 | -151.6789 | 0.23774 | 8.4024 | 0.003631 | 0.005383 | 0.005252 | 6.3e-05 | 0.001017 | 75.139 | 1.3497 | 2.548 | 99.61 | 20.92 | 37.26 | 66.4462 | 109.0007 | 64.4857 | 0.009268 | 1.0456 | 0.02491 | 23.8231 | -0.013066 | 0.14143 | -0.000194 | -0.014331 | -0.014331 | -0.4056
2019-03-21 11:40 | 1553182781.3524 | -332 | 0.20954 | 0.00046 | -3e-06 | 100.5669 | 100.6363 | 124.3308 | 20.9743 | 0.001892 | 0.001782 | -151.6764 | 0.23766 | 8.4026 | 0.003718 | 0.00543 | 0.00521 | 5.8e-05 | 0.001017 | 75.0627 | 1.3659 | 2.5479 | 99.61 | 20.92 | 37.262 | 67.0732 | 111.1068 | 64.4244 | 0.009268 | 1.0458 | 0.024943 | 23.8498 | -0.006172 | 0.14139 | -0.000179 | -0.013184 | -0.013184 | -0.44076
2019-03-21 11:40 | 1553182782.3536 | -331 | 0.2095 | 0.000465 | -3e-06 | 100.5687 | 100.6745 | 124.433 | 20.9745 | 0.001933 | 0.001765 | -151.6844 | 0.23675 | 8.4037 | 0.003597 | 0.004754 | 0.004637 | 5.8e-05 | 0.001023 | 75.1557 | 1.3619 | 2.5479 | 99.61 | 20.91 | 37.269 | 68.0657 | 108.1676 | 64.0162 | 0.009264 | 1.0471 | 0.024968 | 23.8452 | -0.009403 | 0.14084 | 4.3e-05 | 0.003411 | 0.003411 | 1.6971
2019-03-21 11:40 | 1553182783.3521 | -330 | 0.20951 | 0.000461 | -3e-06 | 100.552 | 100.7825 | 124.5393 | 20.9745 | 0.001879 | 0.001786 | -151.664 | 0.23742 | 8.4025 | 0.003571 | 0.005283 | 0.005075 | 5.8e-05 | 0.001022 | 75.2195 | 1.3091 | 2.5479 | 99.61 | 20.91 | 37.273 | 66.762 | 107.5334 | 64.5345 | 0.009265 | 1.0455 | 0.02496 | 23.8737 | -0.020474 | 0.14124 | -3.5e-05 | -0.002409 | -0.002409 | -2.4099
2019-03-21 11:40 | 1553182784.3522 | -329 | 0.20952 | 0.000458 | -3e-06 | 100.5023 | 100.607 | 124.1947 | 20.9751 | 0.001966 | 0.001785 | -151.6635 | 0.23665 | 8.4031 | 0.003591 | 0.004833 | 0.004704 | 6.3e-05 | 0.001021 | 75.1901 | 1.3176 | 2.5479 | 99.61 | 20.91 | 37.276 | 68.8555 | 108.0222 | 64.4954 | 0.009266 | 1.0456 | 0.024927 | 23.8393 | -0.009321 | 0.14078 | -7.4e-05 | -0.005374 | -0.005374 | -1.0766
2019-03-21 11:40 | 1553182785.3513 | -328 | 0.20952 | 0.00046 | -3e-06 | 100.4898 | 100.7297 | 124.1475 | 20.9752 | 0.00186 | 0.001795 | -151.6638 | 0.23672 | 8.4038 | 0.003501 | 0.005019 | 0.004795 | 5.9e-05 | 0.001021 | 75.2276 | 1.3419 | 2.548 | 99.61 | 20.91 | 37.283 | 66.293 | 105.8406 | 64.7327 | 0.009268 | 1.0449 | 0.024913 | 23.8431 | -0.021329 | 0.14083 | -8.1e-05 | -0.005856 | -0.005856 | -0.98844
2019-03-21 11:40 | 1553182786.3523 | -327 | 0.20951 | 0.000461 | -3e-06 | 100.4942 | 100.7419 | 124.2367 | 20.9748 | 0.001935 | 0.001768 | -151.6629 | 0.23744 | 8.4025 | 0.003509 | 0.004814 | 0.004601 | 6.1e-05 | 0.001024 | 75.1828 | 1.3423 | 2.5478 | 99.61 | 20.91 | 37.287 | 68.1165 | 106.0275 | 64.1024 | 0.009269 | 1.0468 | 0.024945 | 23.8294 | -0.02202 | 0.14125 | -3.5e-05 | -0.00237 | -0.00237 | -2.4501
2019-03-21 11:40 | 1553182787.3523 | -326 | 0.20946 | 0.000463 | -4e-06 | 100.4272 | 100.6088 | 124.3773 | 20.9754 | 0.001829 | 0.001768 | -151.6484 | 0.23687 | 8.402 | 0.00358 | 0.004639 | 0.004478 | 5.9e-05 | 0.001027 | 75.1969 | 1.3459 | 2.548 | 99.61 | 20.91 | 37.292 | 65.5527 | 107.7557 | 64.0867 | 0.00927 | 1.0469 | 0.02496 | 23.8424 | -0.016157 | 0.14091 | 0.000266 | 0.020129 | 0.020129 | 0.28772
2019-03-21 11:40 | 1553182788.3514 | -325 | 0.20959 | 0.000459 | -3e-06 | 100.398 | 100.5648 | 124.3165 | 20.9762 | 0.001827 | 0.001764 | -151.6457 | 0.23652 | 8.4041 | 0.003529 | 0.005077 | 0.004933 | 5.5e-05 | 0.001024 | 75.1257 | 1.2992 | 2.548 | 99.61 | 20.91 | 37.295 | 65.5127 | 106.5121 | 64.0039 | 0.009271 | 1.0471 | 0.024957 | 23.8336 | -0.014851 | 0.1407 | -0.000505 | -0.037641 | -0.037641 | -0.15363
2019-03-21 11:40 | 1553182789.3525 | -324 | 0.20951 | 0.000456 | -3e-06 | 100.4272 | 100.575 | 124.6264 | 20.9762 | 0.001881 | 0.001802 | -151.6699 | 0.23712 | 8.4027 | 0.003591 | 0.004849 | 0.004627 | 6.1e-05 | 0.001023 | 75.1378 | 1.2762 | 2.548 | 99.61 | 20.91 | 37.297 | 66.8151 | 108.0179 | 64.9055 | 0.009271 | 1.0443 | 0.024955 | 23.8952 | -0.013158 | 0.14106 | 7e-06 | 0.000732 | 0.000732 | 7.9247
2019-03-21 11:40 | 1553182790.3519 | -323 | 0.20952 | 0.000462 | -3e-06 | 100.4028 | 100.4637 | 124.7444 | 20.9768 | 0.001835 | 0.001756 | -151.683 | 0.23736 | 8.4062 | 0.003566 | 0.005553 | 0.005428 | 5.6e-05 | 0.001021 | 75.1392 | 1.2895 | 2.5478 | 99.61 | 20.91 | 37.301 | 65.6983 | 107.4103 | 63.81 | 0.009272 | 1.0477 | 0.025007 | 23.8677 | -0.005424 | 0.14121 | -7.7e-05 | -0.005557 | -0.005557 | -1.0444
2019-03-21 11:40 | 1553182791.3524 | -322 | 0.20953 | 0.00046 | -4e-06 | 100.4314 | 100.5505 | 124.5379 | 20.9767 | 0.001985 | 0.001807 | -151.6763 | 0.23694 | 8.4011 | 0.003494 | 0.004814 | 0.004659 | 5.8e-05 | 0.001022 | 75.1603 | 1.3181 | 2.5479 | 99.61 | 20.9 | 37.303 | 69.3016 | 105.6753 | 65.0348 | 0.009267 | 1.0439 | 0.024941 | 23.8913 | -0.010606 | 0.14095 | -0.000146 | -0.010696 | -0.010696 | -0.54164
2019-03-21 11:40 | 1553182792.3525 | -321 | 0.20952 | 0.000462 | -3e-06 | 100.4197 | 100.3826 | 124.8536 | 20.978 | 0.001877 | 0.001795 | -151.6511 | 0.23781 | 8.4022 | 0.00354 | 0.005293 | 0.005066 | 6.1e-05 | 0.001027 | 75.1227 | 1.3076 | 2.548 | 99.61 | 20.9 | 37.307 | 66.7084 | 106.7895 | 64.7323 | 0.009268 | 1.0449 | 0.024984 | 23.9108 | 0.003308 | 0.14147 | -4.6e-05 | -0.003281 | -0.003281 | -1.7723
2019-03-21 11:40 | 1553182793.3515 | -320 | 0.20953 | 0.000458 | -3e-06 | 100.3873 | 100.5887 | 124.678 | 20.9779 | 0.001913 | 0.001829 | -151.6496 | 0.2375 | 8.4017 | 0.003532 | 0.005283 | 0.005149 | 5.8e-05 | 0.001031 | 75.1363 | 1.2929 | 2.5478 | 99.61 | 20.9 | 37.308 | 67.5704 | 106.5982 | 65.5659 | 0.009268 | 1.0423 | 0.024936 | 23.9235 | -0.017927 | 0.14129 | -0.000116 | -0.008455 | -0.008455 | -0.68678
2019-03-21 11:40 | 1553182794.3515 | -319 | 0.2095 | 0.000455 | -3e-06 | 100.3599 | 100.4991 | 124.6935 | 20.9782 | 0.001828 | 0.001802 | -151.6843 | 0.23761 | 8.4039 | 0.003589 | 0.004276 | 0.004056 | 6.4e-05 | 0.001033 | 75.1856 | 1.3201 | 2.5479 | 99.61 | 20.9 | 37.309 | 65.539 | 107.9662 | 64.9169 | 0.009268 | 1.0443 | 0.024961 | 23.902 | -0.012401 | 0.14136 | 2.9e-05 | 0.002357 | 0.002357 | 2.4651
2019-03-21 11:40 | 1553182795.3531 | -318 | 0.20954 | 0.000459 | -4e-06 | 100.3542 | 100.3271 | 124.782 | 20.9786 | 0.00186 | 0.001773 | -151.6912 | 0.2367 | 8.4035 | 0.003575 | 0.00499 | 0.004854 | 6e-05 | 0.001037 | 75.1874 | 1.3078 | 2.5479 | 99.61 | 20.9 | 37.31 | 66.3071 | 107.624 | 64.2115 | 0.009269 | 1.0465 | 0.024996 | 23.8855 | 0.002416 | 0.14081 | -0.000205 | -0.015126 | -0.015126 | -0.3826
2019-03-21 11:40 | 1553182796.3528 | -317 | 0.20952 | 0.000462 | -3e-06 | 100.387 | 100.2859 | 124.7749 | 20.9787 | 0.001923 | 0.001841 | -151.7047 | 0.23766 | 8.4015 | 0.003604 | 0.004707 | 0.004508 | 6.1e-05 | 0.001036 | 75.1799 | 1.2991 | 2.5479 | 99.61 | 20.9 | 37.311 | 67.8262 | 108.3261 | 65.8438 | 0.009269 | 1.0415 | 0.024935 | 23.9426 | 0.009011 | 0.14138 | -7.7e-05 | -0.005601 | -0.005601 | -1.0374
2019-03-21 11:40 | 1553182797.3522 | -316 | 0.20954 | 0.000459 | -3e-06 | 100.3253 | 100.5621 | 124.7836 | 20.9786 | 0.001966 | 0.001755 | -151.6984 | 0.23673 | 8.3998 | 0.003524 | 0.004911 | 0.004866 | 6.2e-05 | 0.001032 | 75.1047 | 1.2997 | 2.5479 | 99.61 | 20.9 | 37.313 | 68.8621 | 106.406 | 63.7882 | 0.009269 | 1.0478 | 0.025012 | 23.8707 | -0.021087 | 0.14083 | -0.000183 | -0.013517 | -0.013517 | -0.4282
2019-03-21 11:40 | 1553182798.3512 | -315 | 0.20952 | 0.000463 | -4e-06 | 100.3429 | 100.3751 | 125.0796 | 20.9792 | 0.00197 | 0.001764 | -151.6873 | 0.23677 | 8.4068 | 0.003548 | 0.004613 | 0.004477 | 6.5e-05 | 0.001039 | 75.1417 | 1.2967 | 2.5479 | 99.61 | 20.9 | 37.314 | 68.9467 | 106.9748 | 64.0012 | 0.00927 | 1.0471 | 0.025034 | 23.9066 | -0.002874 | 0.14086 | -8.8e-05 | -0.006358 | -0.006358 | -0.91053
2019-03-21 11:40 | 1553182799.3508 | -314 | 0.20957 | 0.000461 | -3e-06 | 100.3208 | 100.4531 | 124.7968 | 20.979 | 0.001867 | 0.001758 | -151.6676 | 0.23735 | 8.403 | 0.003438 | 0.004765 | 0.004566 | 6.1e-05 | 0.001038 | 75.1498 | 1.3176 | 2.548 | 99.61 | 20.9 | 37.315 | 66.4778 | 104.3105 | 63.8425 | 0.00927 | 1.0476 | 0.025011 | 23.8739 | -0.011787 | 0.1412 | -0.000359 | -0.026805 | -0.026805 | -0.21649
2019-03-21 11:40 | 1553182800.3519 | -313 | 0.20953 | 0.000457 | -4e-06 | 100.3405 | 100.5059 | 124.6083 | 20.9789 | 0.001895 | 0.001798 | -151.6515 | 0.23684 | 8.4063 | 0.003559 | 0.00492 | 0.004778 | 6.5e-05 | 0.001041 | 75.151 | 1.3364 | 2.548 | 99.61 | 20.9 | 37.317 | 67.1612 | 107.2479 | 64.823 | 0.00927 | 1.0446 | 0.024956 | 23.8905 | -0.014733 | 0.1409 | -0.000152 | -0.011185 | -0.011185 | -0.51773
2019-03-21 11:40 | 1553182801.352 | -312 | 0.20951 | 0.000455 | -3e-06 | 100.3441 | 100.4616 | 124.5503 | 20.9783 | 0.001886 | 0.001746 | -151.661 | 0.23789 | 8.3991 | 0.003579 | 0.004467 | 0.004286 | 6e-05 | 0.001039 | 75.1205 | 1.3294 | 2.5479 | 99.61 | 20.9 | 37.318 | 66.935 | 107.734 | 63.5677 | 0.009271 | 1.0485 | 0.024997 | 23.8406 | -0.010473 | 0.14152 | 1.1e-05 | 0.000995 | 0.000995 | 5.8445
2019-03-21 11:40 | 1553182802.3521 | -311 | 0.2095 | 0.000458 | -3e-06 | 100.3223 | 100.2944 | 124.3121 | 20.9809 | 0.001878 | 0.001788 | -151.673 | 0.23649 | 8.4047 | 0.003526 | 0.004767 | 0.00467 | 6.2e-05 | 0.001038 | 75.2136 | 1.3246 | 2.5478 | 99.61 | 20.9 | 37.319 | 66.7502 | 106.4475 | 64.572 | 0.009271 | 1.0454 | 0.024936 | 23.8533 | 0.00249 | 0.14069 | 8.2e-05 | 0.006299 | 0.006299 | 0.91797
2019-03-21 11:40 | 1553182803.3511 | -310 | 0.20948 | 0.000456 | -3e-06 | 100.252 | 100.4259 | 124.4801 | 20.9797 | 0.001845 | 0.001727 | -151.6833 | 0.23824 | 8.4029 | 0.003526 | 0.004898 | 0.004783 | 6.2e-05 | 0.001036 | 75.1454 | 1.3355 | 2.548 | 99.61 | 20.9 | 37.319 | 65.9469 | 106.4401 | 63.103 | 0.009271 | 1.0499 | 0.025007 | 23.8174 | -0.015499 | 0.14173 | 0.000198 | 0.015065 | 0.015065 | 0.38665
2019-03-21 11:40 | 1553182804.3526 | -309 | 0.20947 | 0.000459 | -3e-06 | 100.3342 | 100.1895 | 124.5165 | 20.9802 | 0.001951 | 0.001737 | -151.6374 | 0.23657 | 8.4021 | 0.003558 | 0.005063 | 0.004825 | 6.4e-05 | 0.001039 | 75.1356 | 1.3098 | 2.5479 | 99.61 | 20.9 | 37.319 | 68.4956 | 107.213 | 63.3517 | 0.009271 | 1.0492 | 0.025001 | 23.8297 | 0.012912 | 0.14074 | 0.000214 | 0.016277 | 0.016277 | 0.35535
2019-03-21 11:40 | 1553182805.3512 | -308 | 0.2095 | 0.000462 | -3e-06 | 100.2523 | 100.5209 | 124.2478 | 20.9802 | 0.001858 | 0.00174 | -151.6295 | 0.23807 | 8.4033 | 0.003563 | 0.005024 | 0.004824 | 6.4e-05 | 0.001039 | 75.1238 | 1.2943 | 2.548 | 99.61 | 20.9 | 37.319 | 66.2714 | 107.3555 | 63.4143 | 0.009271 | 1.049 | 0.024972 | 23.8062 | -0.023932 | 0.14163 | 7.2e-05 | 0.005568 | 0.005568 | 1.0454
2019-03-21 11:40 | 1553182806.3523 | -307 | 0.2095 | 0.000461 | -3e-06 | 100.2743 | 100.4303 | 124.2081 | 20.9807 | 0.001895 | 0.001781 | -151.6219 | 0.23714 | 8.4026 | 0.003524 | 0.004924 | 0.004754 | 6.4e-05 | 0.001038 | 75.1326 | 1.3145 | 2.5478 | 99.61 | 20.9 | 37.32 | 67.1465 | 106.408 | 64.4017 | 0.009271 | 1.0459 | 0.024931 | 23.8373 | -0.0139 | 0.14107 | 5.1e-05 | 0.004017 | 0.004017 | 1.4435
2019-03-21 11:40 | 1553182807.3527 | -306 | 0.20952 | 0.000461 | -4e-06 | 100.2541 | 100.3203 | 124.3069 | 20.9804 | 0.001899 | 0.001749 | -151.6326 | 0.23727 | 8.4025 | 0.003555 | 0.004805 | 0.004642 | 6.2e-05 | 0.001036 | 75.1522 | 1.3232 | 2.548 | 99.61 | 20.9 | 37.32 | 67.2444 | 107.1571 | 63.6386 | 0.009271 | 1.0483 | 0.02497 | 23.8198 | -0.005904 | 0.14115 | -8.7e-05 | -0.006278 | -0.006278 | -0.92406
2019-03-21 11:40 | 1553182808.3527 | -305 | 0.20951 | 0.000459 | -4e-06 | 100.2037 | 100.3278 | 124.0675 | 20.9811 | 0.001913 | 0.001792 | -151.6356 | 0.23733 | 8.4031 | 0.00358 | 0.005324 | 0.005118 | 6.2e-05 | 0.001038 | 75.1334 | 1.3297 | 2.548 | 99.61 | 20.9 | 37.321 | 67.5876 | 107.7603 | 64.6769 | 0.009271 | 1.045 | 0.024907 | 23.8335 | -0.011065 | 0.14119 | 2.4e-05 | 0.002031 | 0.002031 | 2.8577
2019-03-21 11:40 | 1553182809.3525 | -304 | 0.20955 | 0.000465 | -3e-06 | 100.2463 | 100.3523 | 124.2064 | 20.9809 | 0.001864 | 0.001765 | -151.6447 | 0.23713 | 8.4016 | 0.003538 | 0.004677 | 0.004521 | 6.5e-05 | 0.001036 | 75.1332 | 1.3207 | 2.5478 | 99.61 | 20.9 | 37.322 | 66.4108 | 106.746 | 64.0217 | 0.009272 | 1.0471 | 0.024945 | 23.8237 | -0.009449 | 0.14107 | -0.00022 | -0.016264 | -0.016264 | -0.35648
2019-03-21 11:40 | 1553182810.3514 | -303 | 0.20952 | 0.000459 | -3e-06 | 100.2013 | 100.2893 | 124.1748 | 20.9817 | 0.001922 | 0.001792 | -151.622 | 0.23744 | 8.4029 | 0.003611 | 0.005182 | 0.005065 | 6.1e-05 | 0.00103 | 75.1894 | 1.339 | 2.5478 | 99.61 | 20.89 | 37.326 | 67.7957 | 108.5158 | 64.6768 | 0.009267 | 1.045 | 0.024918 | 23.8438 | -0.007846 | 0.14125 | -5.3e-05 | -0.003794 | -0.003794 | -1.53
2019-03-21 11:40 | 1553182811.4105 | -302 | 0.2095 | 0.000461 | -3e-06 | 100.2168 | 100.5022 | 124.4902 | 20.9825 | 0.001874 | 0.00178 | -151.6113 | 0.23693 | 8.4036 | 0.003574 | 0.004667 | 0.00443 | 5.7e-05 | 0.001032 | 75.1535 | 1.3166 | 2.548 | 99.62 | 20.89 | 37.327 | 66.6495 | 107.623 | 64.391 | 0.009266 | 1.0459 | 0.02496 | 23.8639 | -0.025429 | 0.14095 | 6.5e-05 | 0.005106 | 0.005106 | 1.1345
2019-03-21 11:40 | 1553182812.3565 | -301 | 0.20948 | 0.000462 | -3e-06 | 100.2246 | 100.215 | 124.5846 | 20.9819 | 0.001895 | 0.001744 | -151.6114 | 0.23822 | 8.4016 | 0.003603 | 0.004413 | 0.004309 | 5.9e-05 | 0.001029 | 75.1451 | 1.3193 | 2.5478 | 99.62 | 20.89 | 37.328 | 67.1491 | 108.318 | 63.5127 | 0.009267 | 1.0487 | 0.025002 | 23.8419 | 0.000852 | 0.14172 | 0.000161 | 0.012302 | 0.012302 | 0.47346
2019-03-21 11:40 | 1553182813.3565 | -300 | 0.20949 | 0.000461 | -3e-06 | 100.0232 | 100.4047 | 124.3569 | 20.9834 | 0.001936 | 0.001744 | -151.6149 | 0.23799 | 8.4033 | 0.003609 | 0.004367 | 0.003846 | 6e-05 | 0.001028 | 75.2041 | 1.306 | 2.548 | 99.62 | 20.89 | 37.333 | 68.1491 | 108.4575 | 63.5267 | 0.009268 | 1.0486 | 0.024979 | 23.8206 | -0.034057 | 0.14158 | 0.000119 | 0.009116 | 0.009116 | 0.63831
2019-03-21 11:40 | 1553182814.3556 | -299 | 0.20949 | 0.000456 | -4e-06 | 99.8218 | 100.4572 | 124.1456 | 20.9827 | 0.001869 | 0.001756 | -151.632 | 0.23754 | 8.4046 | 0.003602 | 0.004841 | 0.005333 | 6.1e-05 | 0.001027 | 75.155 | 1.3211 | 2.548 | 99.62 | 20.89 | 37.333 | 66.5172 | 108.2968 | 63.8016 | 0.009268 | 1.0478 | 0.024947 | 23.8101 | -0.056757 | 0.14131 | 9.9e-05 | 0.007649 | 0.007649 | 0.75929
2019-03-21 11:40 | 1553182815.357 | -298 | 0.20953 | 0.000463 | -3e-06 | 99.9844 | 100.1091 | 124.3543 | 20.9825 | 0.001945 | 0.001792 | -151.6323 | 0.23738 | 8.4026 | 0.003662 | 0.005113 | 0.004889 | 6.1e-05 | 0.001032 | 75.1372 | 1.354 | 2.5479 | 99.61 | 20.89 | 37.335 | 68.3619 | 109.7406 | 64.6859 | 0.009269 | 1.045 | 0.024936 | 23.8613 | -0.011146 | 0.14122 | -0.00011 | -0.008023 | -0.008023 | -0.72348
2019-03-21 11:40 | 1553182816.3567 | -297 | 0.20954 | 0.00046 | -3e-06 | 100.0652 | 100.1456 | 124.1783 | 20.9823 | 0.001884 | 0.001781 | -151.6316 | 0.23669 | 8.4017 | 0.003562 | 0.005552 | 0.005343 | 6.1e-05 | 0.001028 | 75.151 | 1.3009 | 2.548 | 99.61 | 20.88 | 37.336 | 66.8999 | 107.3245 | 64.4047 | 0.009264 | 1.0459 | 0.024928 | 23.8345 | -0.007181 | 0.14081 | -0.000176 | -0.012941 | -0.012941 | -0.44718
2019-03-21 11:40 | 1553182817.3558 | -296 | 0.20952 | 0.000462 | -3e-06 | 100.0196 | 100.29 | 124.0438 | 20.9834 | 0.001928 | 0.001766 | -151.6266 | 0.23733 | 8.4021 | 0.003622 | 0.005313 | 0.005163 | 6.3e-05 | 0.001032 | 75.188 | 1.2778 | 2.5479 | 99.61 | 20.88 | 37.335 | 67.9548 | 108.7689 | 64.0624 | 0.009264 | 1.047 | 0.024927 | 23.8095 | -0.024146 | 0.14119 | -3.6e-05 | -0.002486 | -0.002486 | -2.3341
2019-03-21 11:40 | 1553182818.3572 | -295 | 0.2095 | 0.000463 | -3e-06 | 100.0884 | 100.0883 | 124.3299 | 20.9829 | 0.001904 | 0.001776 | -151.6248 | 0.23732 | 8.4018 | 0.003611 | 0.004978 | 0.004828 | 6.5e-05 | 0.001031 | 75.2133 | 1.2699 | 2.548 | 99.61 | 20.88 | 37.336 | 67.3677 | 108.5002 | 64.2999 | 0.009264 | 1.0462 | 0.024947 | 23.8454 | 8e-06 | 0.14118 | 6.1e-05 | 0.004742 | 0.004742 | 1.2237
2019-03-21 11:40 | 1553182819.3568 | -294 | 0.20949 | 0.000462 | -3e-06 | 100.0279 | 100.1037 | 124.1569 | 20.9838 | 0.001868 | 0.001823 | -151.6214 | 0.237 | 8.4008 | 0.003611 | 0.005155 | 0.005026 | 6.3e-05 | 0.001028 | 75.1984 | 1.2421 | 2.5478 | 99.61 | 20.88 | 37.336 | 66.4971 | 108.506 | 65.421 | 0.009264 | 1.0428 | 0.024889 | 23.8683 | -0.006768 | 0.14099 | 0.000109 | 0.008338 | 0.008338 | 0.69494
2019-03-21 11:40 | 1553182820.357 | -293 | 0.2095 | 0.000462 | -3e-06 | 100.0279 | 100.0696 | 124.1471 | 20.9841 | 0.001903 | 0.001784 | -151.6549 | 0.23682 | 8.4027 | 0.003597 | 0.005633 | 0.005355 | 6.3e-05 | 0.001029 | 75.161 | 1.301 | 2.548 | 99.61 | 20.88 | 37.337 | 67.3511 | 108.1809 | 64.4792 | 0.009264 | 1.0457 | 0.024922 | 23.8341 | -0.003725 | 0.14089 | 3.1e-05 | 0.002472 | 0.002472 | 2.342
2019-03-21 11:40 | 1553182821.3573 | -292 | 0.20952 | 0.000461 | -3e-06 | 99.9395 | 100.1074 | 123.9088 | 20.9842 | 0.001917 | 0.001741 | -151.6639 | 0.23798 | 8.4024 | 0.003603 | 0.005075 | 0.004953 | 6.1e-05 | 0.001028 | 75.1182 | 1.3442 | 2.5479 | 99.61 | 20.88 | 37.336 | 67.689 | 108.3125 | 63.4526 | 0.009264 | 1.0488 | 0.024936 | 23.775 | -0.015019 | 0.14157 | -8.1e-05 | -0.005924 | -0.005924 | -0.98225
2019-03-21 11:40 | 1553182822.357 | -291 | 0.20955 | 0.000458 | -3e-06 | 100.0074 | 100.0757 | 123.9898 | 20.9851 | 0.001899 | 0.00177 | -151.6607 | 0.23744 | 8.4005 | 0.003606 | 0.004863 | 0.004677 | 6.2e-05 | 0.001026 | 75.176 | 1.3411 | 2.548 | 99.61 | 20.87 | 37.337 | 67.2506 | 108.389 | 64.1538 | 0.009258 | 1.0467 | 0.024919 | 23.8076 | -0.006111 | 0.14125 | -0.000232 | -0.017172 | -0.017172 | -0.33809
2019-03-21 11:40 | 1553182823.3561 | -290 | 0.20954 | 0.000461 | -3e-06 | 99.9809 | 100.1118 | 124.2114 | 20.985 | 0.001881 | 0.001742 | -151.6609 | 0.23747 | 8.4006 | 0.003618 | 0.005056 | 0.004853 | 6.2e-05 | 0.001023 | 75.1501 | 1.3134 | 2.5479 | 99.61 | 20.87 | 37.337 | 66.823 | 108.6845 | 63.4743 | 0.009258 | 1.0488 | 0.024966 | 23.8048 | -0.011709 | 0.14127 | -0.000184 | -0.013623 | -0.013623 | -0.42621
2019-03-21 11:40 | 1553182824.3564 | -289 | 0.20954 | 0.000461 | -4e-06 | 99.873 | 99.98 | 124.3036 | 20.9848 | 0.001895 | 0.00178 | -151.6545 | 0.23667 | 8.4015 | 0.003593 | 0.004834 | 0.004713 | 6.5e-05 | 0.001022 | 75.1069 | 1.3361 | 2.5479 | 99.61 | 20.87 | 37.335 | 67.1568 | 108.065 | 64.3853 | 0.009258 | 1.046 | 0.024942 | 23.8458 | -0.009579 | 0.1408 | -0.000193 | -0.014245 | -0.014245 | -0.40624
2019-03-21 11:40 | 1553182825.3571 | -288 | 0.20955 | 0.000461 | -4e-06 | 99.9397 | 99.9283 | 124.3933 | 20.9846 | 0.001837 | 0.001769 | -151.651 | 0.23719 | 8.4024 | 0.003632 | 0.005415 | 0.00521 | 6.3e-05 | 0.001015 | 75.13 | 1.3248 | 2.5479 | 99.61 | 20.87 | 37.333 | 65.7535 | 109.0135 | 64.1189 | 0.009257 | 1.0468 | 0.02496 | 23.845 | 0.001029 | 0.1411 | -0.000252 | -0.018668 | -0.018668 | -0.31065
2019-03-21 11:40 | 1553182826.3576 | -287 | 0.20955 | 0.00046 | -3e-06 | 99.8757 | 99.8881 | 124.5125 | 20.9851 | 0.001862 | 0.001783 | -151.6547 | 0.2373 | 8.403 | 0.003585 | 0.004778 | 0.004635 | 6.3e-05 | 0.001013 | 75.1525 | 1.2925 | 2.5479 | 99.61 | 20.87 | 37.332 | 66.3589 | 107.8934 | 64.4732 | 0.009257 | 1.0457 | 0.024959 | 23.869 | -0.001108 | 0.14117 | -0.000244 | -0.018082 | -0.018082 | -0.32088
2019-03-21 11:40 | 1553182827.3573 | -286 | 0.20955 | 0.00046 | -3e-06 | 99.8825 | 100.0638 | 124.4353 | 20.9854 | 0.001911 | 0.001716 | -151.6689 | 0.23792 | 8.4014 | 0.003564 | 0.005011 | 0.004954 | 6.2e-05 | 0.001003 | 75.1473 | 1.2944 | 2.5479 | 99.61 | 20.87 | 37.333 | 67.5301 | 107.3788 | 62.8563 | 0.009257 | 1.0507 | 0.025012 | 23.8044 | -0.016218 | 0.14154 | -0.00023 | -0.017087 | -0.017087 | -0.34044
2019-03-21 11:40 | 1553182828.3577 | -285 | 0.2096 | 0.000462 | -3e-06 | 99.8441 | 99.9221 | 124.4017 | 20.985 | 0.001952 | 0.001773 | -151.6802 | 0.23638 | 8.4035 | 0.003545 | 0.005255 | 0.004959 | 6e-05 | 0.00102 | 75.137 | 1.3538 | 2.548 | 99.61 | 20.87 | 37.337 | 68.5368 | 106.9216 | 64.2108 | 0.009258 | 1.0465 | 0.024958 | 23.8491 | -0.006986 | 0.14062 | -0.000526 | -0.039213 | -0.039213 | -0.14739
2019-03-21 11:40 | 1553182829.3564 | -284 | 0.20957 | 0.000463 | -3e-06 | 99.8572 | 99.8605 | 124.7461 | 20.9854 | 0.001868 | 0.001721 | -151.6916 | 0.2386 | 8.4023 | 0.003511 | 0.004923 | 0.004746 | 6.1e-05 | 0.001036 | 75.0739 | 1.3581 | 2.548 | 99.61 | 20.87 | 37.339 | 66.4965 | 106.0908 | 62.9757 | 0.009259 | 1.0503 | 0.025038 | 23.8383 | -0.000292 | 0.14194 | -0.000342 | -0.025532 | -0.025532 | -0.22849
2019-03-21 11:40 | 1553182830.3564 | -283 | 0.20957 | 0.000459 | -3e-06 | 99.8128 | 99.6943 | 124.7641 | 20.9862 | 0.001887 | 0.001792 | -151.6916 | 0.2385 | 8.4018 | 0.003567 | 0.004916 | 0.004797 | 6.4e-05 | 0.001045 | 75.0834 | 1.2989 | 2.5479 | 99.61 | 20.87 | 37.341 | 66.9715 | 107.4368 | 64.6811 | 0.009259 | 1.045 | 0.024977 | 23.9004 | 0.01063 | 0.14189 | -0.000348 | -0.025894 | -0.025894 | -0.22521
2019-03-21 11:40 | 1553182831.3559 | -282 | 0.20955 | 0.00046 | -3e-06 | 99.8379 | 99.9075 | 124.7491 | 20.9856 | 0.001842 | 0.0018 | -151.6852 | 0.23781 | 8.4035 | 0.003559 | 0.004476 | 0.004258 | 6.2e-05 | 0.001049 | 75.2913 | 1.3139 | 2.548 | 99.61 | 20.87 | 37.344 | 65.881 | 107.2418 | 64.8733 | 0.00926 | 1.0444 | 0.024968 | 23.9058 | -0.006236 | 0.14147 | -0.000246 | -0.018275 | -0.018275 | -0.31817
2019-03-21 11:41 | 1553182832.3556 | -281 | 0.20955 | 0.000463 | -3e-06 | 99.7369 | 99.8584 | 124.3945 | 20.9866 | 0.001958 | 0.001773 | -151.6792 | 0.23743 | 8.4016 | 0.003552 | 0.004676 | 0.004462 | 6.5e-05 | 0.001058 | 74.6567 | 1.3181 | 2.5478 | 99.61 | 20.87 | 37.347 | 68.6647 | 107.0814 | 64.213 | 0.009261 | 1.0465 | 0.024957 | 23.8485 | -0.010897 | 0.14125 | -0.000223 | -0.016526 | -0.016526 | -0.35128
2019-03-21 11:41 | 1553182833.356 | -280 | 0.20957 | 0.000459 | -3e-06 | 99.7494 | 99.9191 | 123.9961 | 20.9857 | 0.001951 | 0.001758 | -151.6789 | 0.23809 | 8.4013 | 0.003579 | 0.004981 | 0.004872 | 6.3e-05 | 0.001053 | 61.2837 | 1.3418 | 2.5479 | 99.61 | 20.87 | 37.353 | 68.5009 | 107.7367 | 63.8608 | 0.009262 | 1.0476 | 0.02493 | 23.7978 | -0.015204 | 0.14164 | -0.000338 | -0.025094 | -0.025094 | -0.23199
2019-03-21 11:41 | 1553182834.3556 | -279 | 0.20957 | 0.000461 | -3e-06 | 99.6698 | 99.7726 | 123.9515 | 20.9869 | 0.00191 | 0.001773 | -151.6804 | 0.23802 | 8.4038 | 0.00351 | 0.005032 | 0.004748 | 6.4e-05 | 0.001057 | 52.5059 | 1.3235 | 2.5479 | 99.61 | 20.87 | 37.355 | 67.5245 | 106.0701 | 64.2313 | 0.009263 | 1.0464 | 0.024912 | 23.8066 | -0.009221 | 0.14159 | -0.000333 | -0.024705 | -0.024705 | -0.23556
2019-03-21 11:41 | 1553182835.3557 | -278 | 0.20956 | 0.000461 | -3e-06 | 99.7517 | 99.8295 | 123.7271 | 20.9878 | 0.001872 | 0.001743 | -151.699 | 0.23798 | 8.4001 | 0.003562 | 0.004948 | 0.004728 | 6.6e-05 | 0.001065 | 41.0609 | 1.3112 | 2.548 | 99.61 | 20.87 | 37.356 | 66.5955 | 107.3345 | 63.5129 | 0.009263 | 1.0487 | 0.024916 | 23.7597 | -0.006968 | 0.14158 | -0.000299 | -0.022177 | -0.022177 | -0.26238
2019-03-21 11:41 | 1553182836.3561 | -277 | 0.20957 | 0.000462 | -3e-06 | 99.6951 | 99.62 | 123.733 | 20.9874 | 0.001867 | 0.001818 | -151.6984 | 0.23738 | 8.4015 | 0.003615 | 0.0049 | 0.004684 | 6.4e-05 | 0.001064 | 29.9602 | 1.3582 | 2.548 | 99.61 | 20.87 | 37.361 | 66.4725 | 108.6142 | 65.2982 | 0.009264 | 1.0431 | 0.024851 | 23.8232 | 0.006741 | 0.14122 | -0.000365 | -0.02703 | -0.02703 | -0.21473
2019-03-21 11:41 | 1553182837.3568 | -276 | 0.20953 | 0.000458 | -3e-06 | 99.657 | 99.6105 | 123.3976 | 20.9887 | 0.001856 | 0.00177 | -151.6918 | 0.23706 | 8.4027 | 0.00361 | 0.004833 | 0.004711 | 6.3e-05 | 0.001065 | 19.7069 | 1.3391 | 2.5478 | 99.61 | 20.87 | 37.363 | 66.2195 | 108.4918 | 64.1534 | 0.009265 | 1.0467 | 0.024859 | 23.7506 | 0.004175 | 0.14103 | -0.000132 | -0.00962 | -0.00962 | -0.60249
2019-03-21 11:41 | 1553182838.3559 | -275 | 0.20957 | 0.000463 | -3e-06 | 99.604 | 99.6156 | 123.2898 | 20.9882 | 0.001945 | 0.001794 | -151.7027 | 0.23738 | 8.4029 | 0.003581 | 0.0049 | 0.004728 | 6.3e-05 | 0.001066 | 17.9462 | 1.3657 | 2.548 | 99.61 | 20.87 | 37.365 | 68.3651 | 107.7854 | 64.7346 | 0.009265 | 1.0449 | 0.024827 | 23.7607 | -0.001045 | 0.14122 | -0.000341 | -0.025233 | -0.025233 | -0.23002
2019-03-21 11:41 | 1553182839.3573 | -274 | 0.20958 | 0.000463 | -3e-06 | 99.5617 | 99.5744 | 123.3572 | 20.9879 | 0.001949 | 0.001772 | -151.6911 | 0.23711 | 8.4014 | 0.003604 | 0.004595 | 0.004407 | 6.7e-05 | 0.001068 | 8.943 | 1.3214 | 2.548 | 99.61 | 20.87 | 37.366 | 68.4482 | 108.3548 | 64.1886 | 0.009265 | 1.0466 | 0.024854 | 23.748 | -0.001143 | 0.14106 | -0.000413 | -0.030618 | -0.030618 | -0.18935
2019-03-21 11:41 | 1553182840.356 | -273 | 0.20953 | 0.000463 | -3e-06 | 99.6174 | 99.7147 | 123.3922 | 20.9883 | 0.001961 | 0.001768 | -151.6844 | 0.23738 | 8.4027 | 0.00368 | 0.004815 | 0.004618 | 6.3e-05 | 0.001072 | 1.8745 | 1.3557 | 2.5478 | 99.61 | 20.86 | 37.367 | 68.738 | 110.1887 | 64.0998 | 0.00926 | 1.0468 | 0.02486 | 23.7482 | -0.008735 | 0.14122 | -0.000134 | -0.009805 | -0.009805 | -0.59193
2019-03-21 11:41 | 1553182841.3574 | -272 | 0.20956 | 0.000462 | -3e-06 | 99.4994 | 99.4954 | 123.6717 | 20.9887 | 0.00191 | 0.001782 | -151.69 | 0.23792 | 8.4028 | 0.003593 | 0.004185 | 0.004034 | 6.4e-05 | 0.001072 | 0.74285 | 1.3593 | 2.5479 | 99.61 | 20.86 | 37.368 | 67.5311 | 108.0846 | 64.443 | 0.00926 | 1.0458 | 0.024876 | 23.7872 | 0.000361 | 0.14154 | -0.00031 | -0.02299 | -0.02299 | -0.25303
2019-03-21 11:41 | 1553182842.3561 | -271 | 0.20957 | 0.000459 | -3e-06 | 99.4565 | 99.6333 | 123.6802 | 20.9888 | 0.001903 | 0.001746 | -151.7059 | 0.23717 | 8.4023 | 0.00352 | 0.005118 | 0.004909 | 6.4e-05 | 0.001072 | 0.67881 | 1.379 | 2.548 | 99.61 | 20.86 | 37.369 | 67.3619 | 106.3175 | 63.5834 | 0.009261 | 1.0484 | 0.024909 | 23.7577 | -0.015889 | 0.14109 | -0.000319 | -0.023664 | -0.023664 | -0.24505
2019-03-21 11:41 | 1553182843.3561 | -270 | 0.20957 | 0.000459 | -3e-06 | 99.4821 | 99.508 | 123.7445 | 20.989 | 0.001885 | 0.001761 | -151.7235 | 0.23664 | 8.4035 | 0.003608 | 0.00487 | 0.004695 | 6.2e-05 | 0.001067 | 0.57147 | 1.3246 | 2.5479 | 99.61 | 20.86 | 37.369 | 66.9127 | 108.4377 | 63.9374 | 0.009261 | 1.0473 | 0.024902 | 23.7764 | -0.002325 | 0.14078 | -0.000319 | -0.023626 | -0.023626 | -0.2449
2019-03-21 11:41 | 1553182844.3566 | -269 | 0.20955 | 0.000461 | -3e-06 | 99.4389 | 99.2944 | 123.7058 | 20.9894 | 0.001915 | 0.001749 | -151.7302 | 0.23855 | 8.4041 | 0.003642 | 0.005583 | 0.005401 | 6.5e-05 | 0.001071 | 0.56522 | 1.3045 | 2.5479 | 99.61 | 20.86 | 37.37 | 67.6361 | 109.2682 | 63.6459 | 0.009261 | 1.0482 | 0.024909 | 23.7624 | 0.013006 | 0.14191 | -0.00025 | -0.018502 | -0.018502 | -0.31524
2019-03-21 11:41 | 1553182845.3562 | -268 | 0.20956 | 0.000461 | -3e-06 | 99.4443 | 99.4102 | 123.8804 | 20.9905 | 0.001843 | 0.001763 | -151.7157 | 0.23736 | 8.4011 | 0.00361 | 0.00496 | 0.004865 | 6.5e-05 | 0.001066 | 0.3788 | 1.3595 | 2.5479 | 99.61 | 20.86 | 37.369 | 65.9115 | 108.4912 | 63.9952 | 0.009261 | 1.0472 | 0.024913 | 23.7915 | 0.003063 | 0.14121 | -0.00031 | -0.023029 | -0.023029 | -0.25201
2019-03-21 11:41 | 1553182846.3562 | -267 | 0.20958 | 0.000463 | -3e-06 | 99.4356 | 99.527 | 124.0168 | 20.9892 | 0.001912 | 0.001777 | -151.7223 | 0.23717 | 8.4033 | 0.003623 | 0.005681 | 0.005529 | 6.6e-05 | 0.001062 | 0.34164 | 1.3193 | 2.5479 | 99.61 | 20.86 | 37.371 | 67.5761 | 108.8164 | 64.3109 | 0.009261 | 1.0462 | 0.024916 | 23.8157 | -0.008219 | 0.14109 | -0.000401 | -0.029834 | -0.029834 | -0.19437
2019-03-21 11:41 | 1553182847.3567 | -266 | 0.20957 | 0.000458 | -3e-06 | 99.4303 | 99.4528 | 124.3564 | 20.9905 | 0.001855 | 0.001808 | -151.7186 | 0.23784 | 8.4027 | 0.00368 | 0.005083 | 0.004874 | 6.6e-05 | 0.001066 | 0.39119 | 1.3323 | 2.548 | 99.61 | 20.86 | 37.373 | 66.2001 | 110.1887 | 65.0592 | 0.009262 | 1.0439 | 0.024922 | 23.8747 | -0.002027 | 0.14149 | -0.00034 | -0.025298 | -0.025298 | -0.22987
2019-03-21 11:41 | 1553182848.3555 | -265 | 0.20958 | 0.000458 | -3e-06 | 99.3513 | 99.3499 | 124.4815 | 20.9909 | 0.001868 | 0.00174 | -151.7036 | 0.23683 | 8.4038 | 0.003681 | 0.004963 | 0.004732 | 6.7e-05 | 0.001069 | 0.39152 | 1.3572 | 2.548 | 99.61 | 20.86 | 37.377 | 66.5235 | 110.2173 | 63.4394 | 0.009262 | 1.0489 | 0.024994 | 23.8294 | 0.000123 | 0.14089 | -0.000412 | -0.030721 | -0.030721 | -0.18849
2019-03-21 11:41 | 1553182849.3554 | -264 | 0.20953 | 0.000462 | -3e-06 | 99.382 | 99.3019 | 124.3391 | 20.9903 | 0.001926 | 0.001795 | -151.7144 | 0.23805 | 8.4033 | 0.003699 | 0.004862 | 0.004736 | 6.7e-05 | 0.001075 | 0.40678 | 1.328 | 2.548 | 99.61 | 20.86 | 37.378 | 67.8969 | 110.6522 | 64.7568 | 0.009263 | 1.0448 | 0.024931 | 23.8624 | 0.00721 | 0.14161 | -9.4e-05 | -0.00684 | -0.00684 | -0.85098
2019-03-21 11:41 | 1553182850.3564 | -263 | 0.2096 | 0.000457 | -3e-06 | 99.2411 | 99.3125 | 124.2564 | 20.9902 | 0.001934 | 0.001761 | -151.7103 | 0.23783 | 8.4026 | 0.003674 | 0.004857 | 0.00471 | 6.4e-05 | 0.001075 | 0.43205 | 1.3026 | 2.5479 | 99.61 | 20.86 | 37.38 | 68.1067 | 110.044 | 63.933 | 0.009263 | 1.0474 | 0.024954 | 23.8253 | -0.006433 | 0.14149 | -0.000525 | -0.039148 | -0.039148 | -0.14854
2019-03-21 11:41 | 1553182851.3559 | -262 | 0.2096 | 0.000462 | -3e-06 | 99.3504 | 99.3189 | 124.2855 | 20.9905 | 0.001866 | 0.001796 | -151.7089 | 0.23716 | 8.4029 | 0.003612 | 0.005416 | 0.005228 | 6.2e-05 | 0.001068 | 0.37106 | 1.3049 | 2.5479 | 99.61 | 20.86 | 37.383 | 66.471 | 108.5442 | 64.7773 | 0.009264 | 1.0447 | 0.024925 | 23.858 | 0.002834 | 0.14108 | -0.000543 | -0.040423 | -0.040423 | -0.14345
2019-03-21 11:41 | 1553182852.3565 | -261 | 0.20956 | 0.000463 | -3e-06 | 99.312 | 99.4436 | 124.5698 | 20.9905 | 0.001921 | 0.001749 | -151.689 | 0.23716 | 8.4027 | 0.003656 | 0.005339 | 0.005167 | 6.5e-05 | 0.001065 | 0.38455 | 1.3257 | 2.548 | 99.61 | 20.86 | 37.385 | 67.7925 | 109.6153 | 63.6539 | 0.009264 | 1.0482 | 0.024995 | 23.8455 | -0.011846 | 0.14108 | -0.000285 | -0.021217 | -0.021217 | -0.2733
2019-03-21 11:41 | 1553182853.3557 | -260 | 0.20956 | 0.00046 | -3e-06 | 99.2834 | 99.3258 | 124.8106 | 20.9902 | 0.001897 | 0.001755 | -151.6776 | 0.23702 | 8.4005 | 0.003608 | 0.00508 | 0.004922 | 6.3e-05 | 0.001058 | 0.4143 | 1.3226 | 2.548 | 99.61 | 20.86 | 37.389 | 67.201 | 108.4564 | 63.8012 | 0.009265 | 1.0478 | 0.025014 | 23.8737 | -0.003817 | 0.141 | -0.000263 | -0.019577 | -0.019577 | -0.29602
2019-03-21 11:41 | 1553182854.356 | -259 | 0.20957 | 0.000464 | -3e-06 | 99.2408 | 99.4024 | 124.8449 | 20.9913 | 0.001872 | 0.00176 | -151.6813 | 0.23717 | 8.4013 | 0.003611 | 0.00488 | 0.00468 | 6.4e-05 | 0.001059 | 0.3741 | 1.328 | 2.5481 | 99.61 | 20.86 | 37.389 | 66.6035 | 108.5094 | 63.9021 | 0.009265 | 1.0475 | 0.025014 | 23.8806 | -0.014555 | 0.1411 | -0.000349 | -0.025992 | -0.025992 | -0.22311
2019-03-21 11:41 | 1553182855.3568 | -258 | 0.20954 | 0.000463 | -3e-06 | 99.2744 | 99.4637 | 125.0106 | 20.9904 | 0.001851 | 0.001747 | -151.682 | 0.23739 | 8.4013 | 0.003591 | 0.005242 | 0.005131 | 6.8e-05 | 0.001057 | 0.36514 | 1.3198 | 2.548 | 99.61 | 20.86 | 37.389 | 66.0947 | 108.0382 | 63.6045 | 0.009265 | 1.0484 | 0.025041 | 23.8859 | -0.017036 | 0.14122 | -0.000196 | -0.014586 | -0.014586 | -0.39792
2019-03-21 11:41 | 1553182856.3558 | -257 | 0.20957 | 0.000456 | -3e-06 | 99.2816 | 99.1252 | 124.9463 | 20.9911 | 0.001901 | 0.001785 | -151.6712 | 0.23791 | 8.4018 | 0.003664 | 0.004706 | 0.00452 | 6.4e-05 | 0.001056 | 0.35939 | 1.3471 | 2.5479 | 99.61 | 20.86 | 37.389 | 67.3118 | 109.8021 | 64.5242 | 0.009265 | 1.0455 | 0.025001 | 23.9123 | 0.014105 | 0.14153 | -0.000351 | -0.026142 | -0.026142 | -0.22251
2019-03-21 11:41 | 1553182857.3563 | -256 | 0.20955 | 0.000458 | -3e-06 | 99.2807 | 99.2253 | 125.0195 | 20.9912 | 0.001894 | 0.001759 | -151.6614 | 0.23816 | 8.4008 | 0.003603 | 0.004808 | 0.004668 | 6.4e-05 | 0.001057 | 0.37017 | 1.3609 | 2.5479 | 99.61 | 20.86 | 37.389 | 67.1332 | 108.3163 | 63.8957 | 0.009265 | 1.0475 | 0.025031 | 23.8971 | 0.004994 | 0.14168 | -0.00023 | -0.017104 | -0.017104 | -0.34045
2019-03-21 11:41 | 1553182858.3559 | -255 | 0.20955 | 0.000462 | -4e-06 | 99.2771 | 99.3067 | 124.8975 | 20.9918 | 0.001895 | 0.00178 | -151.658 | 0.23774 | 8.4036 | 0.003616 | 0.005114 | 0.00502 | 6.3e-05 | 0.001052 | 0.34219 | 1.3843 | 2.548 | 99.61 | 20.86 | 37.388 | 67.1736 | 108.6367 | 64.3913 | 0.009265 | 1.0459 | 0.025001 | 23.903 | -0.002663 | 0.14143 | -0.000226 | -0.016751 | -0.016751 | -0.347
2019-03-21 11:41 | 1553182859.3565 | -254 | 0.20957 | 0.000461 | -3e-06 | 99.2807 | 99.2311 | 124.9806 | 20.9915 | 0.001915 | 0.001802 | -151.6526 | 0.23721 | 8.4023 | 0.003575 | 0.005491 | 0.005281 | 6.7e-05 | 0.001047 | 0.35939 | 1.3468 | 2.5479 | 99.61 | 20.86 | 37.388 | 67.6398 | 107.6437 | 64.9204 | 0.009265 | 1.0443 | 0.02499 | 23.9296 | 0.004472 | 0.14111 | -0.000364 | -0.027116 | -0.027116 | -0.21389
2019-03-21 11:41 | 1553182860.3561 | -253 | 0.20954 | 0.000459 | -3e-06 | 99.2989 | 99.3166 | 124.7897 | 20.9923 | 0.0019 | 0.001764 | -151.6573 | 0.23803 | 8.4033 | 0.003556 | 0.004738 | 0.004587 | 6.5e-05 | 0.00105 | 0.37985 | 1.3211 | 2.5479 | 99.61 | 20.86 | 37.388 | 67.291 | 107.1807 | 64.0039 | 0.009265 | 1.0471 | 0.025004 | 23.8789 | -0.001593 | 0.1416 | -0.000184 | -0.013631 | -0.013631 | -0.42694
2019-03-21 11:41 | 1553182861.3565 | -252 | 0.20949 | 0.000461 | -4e-06 | 99.3096 | 99.417 | 124.5689 | 20.9922 | 0.001933 | 0.001764 | -151.663 | 0.23774 | 8.4032 | 0.003479 | 0.005125 | 0.00489 | 6.5e-05 | 0.001045 | 0.34562 | 1.3363 | 2.5479 | 99.61 | 20.86 | 37.389 | 68.086 | 105.3263 | 64.006 | 0.009265 | 1.0471 | 0.024982 | 23.8579 | -0.009671 | 0.14143 | 0.000163 | 0.012446 | 0.012446 | 0.46706
2019-03-21 11:41 | 1553182862.3565 | -251 | 0.20954 | 0.000461 | -3e-06 | 99.3558 | 99.3748 | 124.4041 | 20.9927 | 0.001873 | 0.001738 | -151.6772 | 0.23719 | 8.4033 | 0.003562 | 0.00484 | 0.004699 | 6.5e-05 | 0.001054 | 0.35651 | 1.3141 | 2.548 | 99.61 | 20.86 | 37.391 | 66.6431 | 107.3253 | 63.392 | 0.009266 | 1.049 | 0.024988 | 23.8204 | -0.001712 | 0.14111 | -0.000171 | -0.012611 | -0.012611 | -0.45989
2019-03-21 11:41 | 1553182863.3552 | -250 | 0.20953 | 0.000461 | -3e-06 | 99.4091 | 99.3625 | 124.1283 | 20.9922 | 0.001874 | 0.001728 | -151.6631 | 0.23746 | 8.4037 | 0.003526 | 0.004393 | 0.004221 | 6.3e-05 | 0.001054 | 0.34291 | 1.3261 | 2.548 | 99.61 | 20.86 | 37.394 | 66.6582 | 106.4538 | 63.1347 | 0.009267 | 1.0498 | 0.02497 | 23.7848 | 0.004191 | 0.14127 | -8.4e-05 | -0.006106 | -0.006106 | -0.95082
2019-03-21 11:41 | 1553182864.3566 | -249 | 0.20947 | 0.000458 | -3e-06 | 99.4389 | 99.4133 | 124.175 | 20.9928 | 0.001847 | 0.001772 | -151.6447 | 0.23705 | 8.4022 | 0.003529 | 0.004476 | 0.004324 | 6.7e-05 | 0.00105 | 0.33411 | 1.3214 | 2.548 | 99.61 | 20.86 | 37.397 | 66.0059 | 106.5262 | 64.2145 | 0.009267 | 1.0465 | 0.024935 | 23.8275 | 0.002305 | 0.14102 | 0.00023 | 0.017448 | 0.017448 | 0.33218
2019-03-21 11:41 | 1553182865.3563 | -248 | 0.20957 | 0.000458 | -3e-06 | 99.4067 | 99.5526 | 124.083 | 20.9927 | 0.001894 | 0.001746 | -151.6401 | 0.23741 | 8.4029 | 0.003558 | 0.005183 | 0.005038 | 6.6e-05 | 0.001053 | 0.35297 | 1.3721 | 2.5481 | 99.61 | 20.86 | 37.4 | 67.1479 | 107.2362 | 63.5685 | 0.009268 | 1.0485 | 0.02495 | 23.7958 | -0.013115 | 0.14123 | -0.000361 | -0.026877 | -0.026877 | -0.21597
2019-03-21 11:41 | 1553182866.3564 | -247 | 0.2094 | 0.000463 | -3e-06 | 99.5211 | 99.4279 | 124.0313 | 20.9936 | 0.001926 | 0.00169 | -151.6461 | 0.23653 | 8.4029 | 0.003451 | 0.003796 | 0.003601 | 6.6e-05 | 0.001049 | 0.35546 | 1.3851 | 2.5481 | 99.61 | 20.86 | 37.406 | 67.9032 | 104.6495 | 62.2423 | 0.00927 | 1.0526 | 0.024994 | 23.744 | 0.008381 | 0.14071 | 0.000652 | 0.049116 | 0.049116 | 0.11775
2019-03-21 11:41 | 1553182867.3564 | -246 | 0.20952 | 0.000461 | -3e-06 | 99.4967 | 99.6285 | 124.0388 | 20.9938 | 0.001876 | 0.001733 | -151.6654 | 0.23753 | 8.4032 | 0.003574 | 0.00496 | 0.004655 | 6.2e-05 | 0.001046 | 0.3569 | 1.3815 | 2.548 | 99.61 | 20.86 | 37.41 | 66.7074 | 107.6288 | 63.2619 | 0.009271 | 1.0494 | 0.024957 | 23.7808 | -0.011844 | 0.14131 | -6.1e-05 | -0.004378 | -0.004378 | -1.3265
2019-03-21 11:41 | 1553182868.3559 | -245 | 0.20958 | 0.000465 | -3e-06 | 99.5956 | 99.6425 | 124.1837 | 20.993 | 0.001906 | 0.001834 | -151.6639 | 0.23748 | 8.403 | 0.003582 | 0.004921 | 0.004831 | 6.1e-05 | 0.001048 | 0.33483 | 1.3163 | 2.548 | 99.61 | 20.86 | 37.417 | 67.4252 | 107.8161 | 65.7054 | 0.009272 | 1.0419 | 0.024881 | 23.8809 | -0.004209 | 0.14128 | -0.000404 | -0.030023 | -0.030023 | -0.1934
2019-03-21 11:41 | 1553182869.3565 | -244 | 0.20952 | 0.000464 | -3e-06 | 99.6936 | 99.6435 | 123.9609 | 20.9933 | 0.001839 | 0.001778 | -151.6631 | 0.23722 | 8.4023 | 0.00359 | 0.005112 | 0.004996 | 6.4e-05 | 0.001053 | 0.36945 | 1.3207 | 2.5481 | 99.61 | 20.86 | 37.42 | 65.8173 | 108.0151 | 64.3594 | 0.009273 | 1.046 | 0.024908 | 23.812 | 0.004498 | 0.14112 | -7.3e-05 | -0.005224 | -0.005224 | -1.1103
2019-03-21 11:41 | 1553182870.3557 | -243 | 0.2096 | 0.000463 | -3e-06 | 99.7062 | 99.7045 | 124.4038 | 20.9943 | 0.001945 | 0.001812 | -151.67 | 0.23665 | 8.4014 | 0.003591 | 0.004891 | 0.004705 | 6.4e-05 | 0.001049 | 0.37266 | 1.3418 | 2.548 | 99.61 | 20.86 | 37.423 | 68.3741 | 108.041 | 65.1659 | 0.009274 | 1.0435 | 0.024923 | 23.883 | 0.00015 | 0.14078 | -0.000503 | -0.037469 | -0.037469 | -0.15442
2019-03-21 11:41 | 1553182871.3559 | -242 | 0.20955 | 0.000464 | -3e-06 | 99.7643 | 99.761 | 124.3935 | 20.9938 | 0.001927 | 0.001738 | -151.6756 | 0.23757 | 8.4014 | 0.003539 | 0.004793 | 0.004681 | 6.2e-05 | 0.001046 | 0.32244 | 1.315 | 2.5479 | 99.61 | 20.86 | 37.429 | 67.9315 | 106.7853 | 63.3767 | 0.009275 | 1.0491 | 0.024988 | 23.8188 | 0.00029 | 0.14133 | -0.000247 | -0.018334 | -0.018334 | -0.31682
2019-03-21 11:41 | 1553182872.3557 | -241 | 0.20957 | 0.000458 | -3e-06 | 99.8247 | 99.7992 | 124.4088 | 20.9939 | 0.001914 | 0.001765 | -151.6737 | 0.23745 | 8.4013 | 0.003585 | 0.005132 | 0.00484 | 6.1e-05 | 0.001043 | 0.35347 | 1.3126 | 2.548 | 99.61 | 20.86 | 37.431 | 67.6321 | 107.8823 | 64.0428 | 0.009276 | 1.047 | 0.024965 | 23.8438 | 0.002291 | 0.14126 | -0.000313 | -0.023289 | -0.023289 | -0.24929
2019-03-21 11:41 | 1553182873.3567 | -240 | 0.20957 | 0.000458 | -3e-06 | 99.8471 | 99.8237 | 124.3524 | 20.9951 | 0.001884 | 0.001747 | -151.6782 | 0.23714 | 8.4008 | 0.003544 | 0.004215 | 0.004094 | 6.2e-05 | 0.001044 | 0.3533 | 1.2889 | 2.548 | 99.61 | 20.86 | 37.433 | 66.9013 | 106.8953 | 63.5951 | 0.009276 | 1.0484 | 0.024976 | 23.8226 | 0.002096 | 0.14107 | -0.000317 | -0.023596 | -0.023596 | -0.24572
2019-03-21 11:41 | 1553182874.356 | -239 | 0.20956 | 0.000461 | -3e-06 | 99.9257 | 99.8799 | 124.2837 | 20.9943 | 0.001909 | 0.001755 | -151.7028 | 0.23709 | 8.4023 | 0.003503 | 0.00473 | 0.004601 | 6.4e-05 | 0.001042 | 0.32421 | 1.3148 | 2.5478 | 99.61 | 20.86 | 37.436 | 67.5053 | 105.9105 | 63.8083 | 0.009277 | 1.0477 | 0.024961 | 23.8235 | 0.004105 | 0.14105 | -0.000277 | -0.0206 | -0.0206 | -0.2814
2019-03-21 11:41 | 1553182875.3559 | -238 | 0.20958 | 0.00046 | -3e-06 | 99.963 | 100.0696 | 124.2107 | 20.9951 | 0.001919 | 0.001781 | -151.7093 | 0.23792 | 8.3993 | 0.003523 | 0.005106 | 0.004925 | 6.2e-05 | 0.001045 | 0.33771 | 1.3239 | 2.548 | 99.61 | 20.86 | 37.437 | 67.7501 | 106.3938 | 64.4126 | 0.009277 | 1.0459 | 0.024931 | 23.8379 | -0.009535 | 0.14154 | -0.000384 | -0.028536 | -0.028536 | -0.20386
2019-03-21 11:41 | 1553182876.3568 | -237 | 0.20955 | 0.000461 | -3e-06 | 100.0822 | 100.2184 | 124.3757 | 20.9945 | 0.001903 | 0.001769 | -151.7088 | 0.23655 | 8.4025 | 0.003492 | 0.005124 | 0.005002 | 6.2e-05 | 0.001048 | 0.34667 | 1.3237 | 2.548 | 99.61 | 20.86 | 37.441 | 67.3557 | 105.6314 | 64.138 | 0.009278 | 1.0467 | 0.024958 | 23.844 | -0.012171 | 0.14072 | -0.000234 | -0.017357 | -0.017357 | -0.33322
2019-03-21 11:41 | 1553182877.3549 | -236 | 0.20957 | 0.000459 | -4e-06 | 100.1155 | 100.3727 | 124.3358 | 20.9951 | 0.001847 | 0.001757 | -151.7076 | 0.23669 | 8.4009 | 0.003464 | 0.005125 | 0.00496 | 6.3e-05 | 0.001047 | 0.32731 | 1.3266 | 2.5479 | 99.61 | 20.85 | 37.442 | 66.0151 | 104.9555 | 63.8372 | 0.009273 | 1.0477 | 0.024965 | 23.8296 | -0.022948 | 0.14081 | -0.000323 | -0.02401 | -0.02401 | -0.24102
2019-03-21 11:41 | 1553182878.3554 | -235 | 0.20955 | 0.000462 | -4e-06 | 100.1838 | 100.2528 | 124.044 | 20.9945 | 0.001872 | 0.001774 | -151.7051 | 0.23657 | 8.4035 | 0.003467 | 0.00524 | 0.005131 | 6.1e-05 | 0.00105 | 0.32068 | 1.2978 | 2.5481 | 99.61 | 20.85 | 37.443 | 66.6027 | 105.0414 | 64.2659 | 0.009273 | 1.0463 | 0.02492 | 23.8167 | -0.006165 | 0.14074 | -0.000214 | -0.015777 | -0.015777 | -0.36663
2019-03-21 11:41 | 1553182879.356 | -234 | 0.20957 | 0.000464 | -3e-06 | 100.339 | 100.403 | 123.8091 | 20.9954 | 0.001859 | 0.001796 | -151.6949 | 0.23662 | 8.4016 | 0.003486 | 0.004835 | 0.004751 | 6.3e-05 | 0.001051 | 0.3653 | 1.3189 | 2.548 | 99.61 | 20.85 | 37.443 | 66.3102 | 105.4962 | 64.783 | 0.009273 | 1.0447 | 0.024877 | 23.8124 | -0.005707 | 0.14076 | -0.00032 | -0.02372 | -0.02372 | -0.2439
2019-03-21 11:41 | 1553182880.356 | -233 | 0.20957 | 0.000461 | -3e-06 | 100.3217 | 100.4551 | 123.6917 | 20.9958 | 0.001939 | 0.001804 | -151.691 | 0.23793 | 8.4016 | 0.003483 | 0.005024 | 0.004784 | 6.3e-05 | 0.001046 | 0.34291 | 1.3758 | 2.5481 | 99.61 | 20.85 | 37.443 | 68.2189 | 105.4178 | 64.9722 | 0.009273 | 1.0441 | 0.024859 | 23.8078 | -0.011889 | 0.14155 | -0.000332 | -0.024603 | -0.024603 | -0.23646
2019-03-21 11:41 | 1553182881.3554 | -232 | 0.20954 | 0.00046 | -3e-06 | 100.3858 | 100.527 | 123.8234 | 20.9964 | 0.001908 | 0.001793 | -151.6774 | 0.23696 | 8.403 | 0.003507 | 0.004659 | 0.004469 | 6.3e-05 | 0.001044 | 0.32477 | 1.3724 | 2.548 | 99.61 | 20.85 | 37.444 | 67.4715 | 106.0087 | 64.7221 | 0.009273 | 1.0449 | 0.024881 | 23.8116 | -0.012576 | 0.14097 | -0.000167 | -0.012313 | -0.012313 | -0.47052
2019-03-21 11:41 | 1553182882.3562 | -231 | 0.20956 | 0.00046 | -3e-06 | 100.496 | 100.6554 | 124.0388 | 20.9961 | 0.00191 | 0.001813 | -151.6954 | 0.23644 | 8.4035 | 0.003509 | 0.004559 | 0.004329 | 6.5e-05 | 0.001049 | 0.35098 | 1.3058 | 2.548 | 99.61 | 20.85 | 37.444 | 67.5195 | 106.0437 | 65.1933 | 0.009273 | 1.0435 | 0.024885 | 23.8489 | -0.014176 | 0.14066 | -0.000272 | -0.020104 | -0.020104 | -0.28756
2019-03-21 11:41 | 1553182883.3567 | -230 | 0.20956 | 0.000464 | -3e-06 | 100.5443 | 100.8209 | 124.0553 | 20.996 | 0.001897 | 0.001765 | -151.7031 | 0.23691 | 8.4024 | 0.003572 | 0.005027 | 0.004902 | 6.6e-05 | 0.001049 | 0.35563 | 1.2915 | 2.548 | 99.61 | 20.85 | 37.445 | 67.2174 | 107.5671 | 64.0388 | 0.009274 | 1.047 | 0.024929 | 23.8098 | -0.024578 | 0.14094 | -0.000263 | -0.019509 | -0.019509 | -0.29691
2019-03-21 11:41 | 1553182884.3563 | -229 | 0.20955 | 0.000461 | -3e-06 | 100.6363 | 100.7375 | 123.9968 | 20.9958 | 0.001955 | 0.001817 | -151.674 | 0.23749 | 8.4 | 0.003546 | 0.004532 | 0.004414 | 6.6e-05 | 0.001051 | 0.33771 | 1.3236 | 2.5479 | 99.61 | 20.85 | 37.445 | 68.6088 | 106.9575 | 65.2814 | 0.009274 | 1.0432 | 0.024878 | 23.848 | -0.008986 | 0.14128 | -0.000188 | -0.013885 | -0.013885 | -0.4182
2019-03-21 11:41 | 1553182885.3563 | -228 | 0.20954 | 0.000464 | -3e-06 | 100.6414 | 100.7041 | 123.7199 | 20.9965 | 0.001846 | 0.001822 | -151.6784 | 0.23636 | 8.4005 | 0.003486 | 0.004373 | 0.004162 | 6.6e-05 | 0.001056 | 0.3282 | 1.303 | 2.548 | 99.61 | 20.85 | 37.444 | 65.989 | 105.4861 | 65.4107 | 0.009273 | 1.0428 | 0.024845 | 23.8259 | -0.005572 | 0.14061 | -0.000151 | -0.011105 | -0.011105 | -0.52042
2019-03-21 11:41 | 1553182886.3567 | -227 | 0.20952 | 0.000461 | -3e-06 | 100.7716 | 100.9395 | 123.7037 | 20.9969 | 0.001907 | 0.001788 | -151.6649 | 0.23708 | 8.4024 | 0.00349 | 0.005359 | 0.005104 | 6.8e-05 | 0.001055 | 0.33284 | 1.2582 | 2.548 | 99.61 | 20.85 | 37.444 | 67.4501 | 105.5976 | 64.5883 | 0.009273 | 1.0453 | 0.024874 | 23.7954 | -0.014887 | 0.14104 | -3.1e-05 | -0.002118 | -0.002118 | -2.7363
2019-03-21 11:41 | 1553182887.3554 | -226 | 0.20953 | 0.00046 | -3e-06 | 100.7302 | 101.0839 | 123.8914 | 20.9964 | 0.001904 | 0.001781 | -151.6663 | 0.23722 | 8.4024 | 0.003501 | 0.00526 | 0.005214 | 6.3e-05 | 0.001054 | 0.35745 | 1.2694 | 2.5479 | 99.61 | 20.85 | 37.441 | 67.3941 | 105.8544 | 64.424 | 0.009273 | 1.0458 | 0.024899 | 23.8076 | -0.031351 | 0.14112 | -0.000101 | -0.007359 | -0.007359 | -0.78815
2019-03-21 11:41 | 1553182888.3559 | -225 | 0.20959 | 0.000466 | -3e-06 | 100.8458 | 101.154 | 124.3994 | 20.9976 | 0.001863 | 0.001777 | -151.6779 | 0.23755 | 8.4024 | 0.003509 | 0.004447 | 0.004286 | 6.3e-05 | 0.001055 | 0.31813 | 1.2712 | 2.5479 | 99.61 | 20.85 | 37.44 | 66.4045 | 106.0397 | 64.3292 | 0.009272 | 1.0461 | 0.024953 | 23.853 | -0.027298 | 0.14132 | -0.000453 | -0.033788 | -0.033788 | -0.1719
2019-03-21 11:41 | 1553182889.356 | -224 | 0.20955 | 0.000463 | -3e-06 | 100.9706 | 101.0209 | 124.6163 | 20.9978 | 0.00191 | 0.001822 | -151.682 | 0.23591 | 8.4019 | 0.003532 | 0.004626 | 0.004402 | 6.4e-05 | 0.00106 | 0.323 | 1.2985 | 2.5482 | 99.61 | 20.85 | 37.439 | 67.5204 | 106.6192 | 65.4113 | 0.009272 | 1.0428 | 0.024935 | 23.9121 | -0.004448 | 0.14034 | -0.000249 | -0.018461 | -0.018461 | -0.31245
2019-03-21 11:41 | 1553182890.3557 | -223 | 0.20958 | 0.000463 | -3e-06 | 100.9766 | 101.1605 | 124.47 | 20.9979 | 0.001888 | 0.001816 | -151.6835 | 0.23543 | 8.4041 | 0.003432 | 0.004999 | 0.004803 | 6.5e-05 | 0.001058 | 0.34092 | 1.303 | 2.5479 | 99.61 | 20.85 | 37.439 | 67.0033 | 104.1797 | 65.2764 | 0.009272 | 1.0432 | 0.024925 | 23.8933 | -0.016276 | 0.14006 | -0.00038 | -0.028211 | -0.028211 | -0.20405
2019-03-21 11:42 | 1553182891.3549 | -222 | 0.20955 | 0.000461 | -3e-06 | 100.9733 | 101.2378 | 124.1884 | 20.9974 | 0.00188 | 0.001774 | -151.6767 | 0.23699 | 8.4004 | 0.003489 | 0.004663 | 0.00454 | 6.1e-05 | 0.001057 | 0.3366 | 1.3067 | 2.5479 | 99.61 | 20.85 | 37.438 | 66.796 | 105.576 | 64.2531 | 0.009272 | 1.0464 | 0.024935 | 23.8301 | -0.0234 | 0.14098 | -0.000206 | -0.01523 | -0.01523 | -0.38047
2019-03-21 11:42 | 1553182892.3557 | -221 | 0.20956 | 0.000464 | -3e-06 | 101.0651 | 101.186 | 123.9224 | 20.9978 | 0.001883 | 0.001796 | -151.645 | 0.23585 | 8.4047 | 0.00351 | 0.00489 | 0.004744 | 6.5e-05 | 0.001057 | 0.3235 | 1.3136 | 2.5479 | 99.61 | 20.85 | 37.439 | 66.8736 | 106.0644 | 64.7985 | 0.009272 | 1.0447 | 0.024888 | 23.8238 | -0.010699 | 0.14031 | -0.000307 | -0.022782 | -0.022782 | -0.25312
2019-03-21 11:42 | 1553182893.3562 | -220 | 0.20953 | 0.000465 | -3e-06 | 101.0868 | 101.2092 | 124.0935 | 20.9973 | 0.001909 | 0.001779 | -151.6436 | 0.23706 | 8.4031 | 0.003487 | 0.005123 | 0.004937 | 6.2e-05 | 0.001057 | 0.34252 | 1.3619 | 2.548 | 99.61 | 20.85 | 37.439 | 67.5027 | 105.5286 | 64.3686 | 0.009272 | 1.046 | 0.024921 | 23.8251 | -0.010822 | 0.14103 | -0.000122 | -0.00894 | -0.00894 | -0.64835
2019-03-21 11:42 | 1553182894.3558 | -219 | 0.20955 | 0.000462 | -3e-06 | 101.0529 | 101.3557 | 124.3773 | 20.9977 | 0.001867 | 0.001806 | -151.6623 | 0.23792 | 8.4025 | 0.003421 | 0.005048 | 0.004919 | 6.2e-05 | 0.001052 | 0.32388 | 1.3384 | 2.5481 | 99.61 | 20.86 | 37.44 | 66.4842 | 103.9287 | 65.0255 | 0.009278 | 1.044 | 0.024925 | 23.8755 | -0.026761 | 0.14154 | -0.000203 | -0.014993 | -0.014993 | -0.388
2019-03-21 11:42 | 1553182895.3549 | -218 | 0.20951 | 0.000465 | -3e-06 | 101.1113 | 101.3815 | 124.4477 | 20.9976 | 0.001907 | 0.001797 | -151.6556 | 0.23742 | 8.3994 | 0.00337 | 0.005279 | 0.00506 | 6.3e-05 | 0.001051 | 0.33229 | 1.3348 | 2.5481 | 99.61 | 20.85 | 37.442 | 67.4594 | 102.6807 | 64.807 | 0.009273 | 1.0446 | 0.02494 | 23.8746 | -0.023878 | 0.14124 | 1.9e-05 | 0.001646 | 0.001646 | 3.5259
2019-03-21 11:42 | 1553182896.3554 | -217 | 0.20957 | 0.000462 | -3e-06 | 101.2316 | 101.4538 | 124.2151 | 20.9978 | 0.00198 | 0.001795 | -151.6592 | 0.23647 | 8.4032 | 0.003451 | 0.00455 | 0.004366 | 6.4e-05 | 0.001046 | 0.33124 | 1.2898 | 2.548 | 99.61 | 20.86 | 37.443 | 69.2051 | 104.6442 | 64.7658 | 0.009279 | 1.0448 | 0.024919 | 23.8508 | -0.019604 | 0.14068 | -0.000327 | -0.024281 | -0.024281 | -0.23812
2019-03-21 11:42 | 1553182897.356 | -216 | 0.20948 | 0.00046 | -4e-06 | 101.1762 | 101.4769 | 123.9062 | 20.9977 | 0.001907 | 0.001773 | -151.6578 | 0.23729 | 8.4025 | 0.003442 | 0.004368 | 0.004217 | 6.5e-05 | 0.001041 | 0.34722 | 1.3024 | 2.5478 | 99.61 | 20.86 | 37.445 | 67.4679 | 104.4177 | 64.227 | 0.009279 | 1.0464 | 0.024908 | 23.8021 | -0.026544 | 0.14116 | 0.000216 | 0.016414 | 0.016414 | 0.35347
2019-03-21 11:42 | 1553182898.355 | -215 | 0.20952 | 0.000463 | -3e-06 | 101.2111 | 101.37 | 123.5816 | 20.9983 | 0.001962 | 0.001805 | -151.6583 | 0.23733 | 8.4032 | 0.003477 | 0.004545 | 0.004413 | 6.5e-05 | 0.001041 | 0.31924 | 1.2879 | 2.5479 | 99.61 | 20.86 | 37.447 | 68.7672 | 105.2747 | 64.9934 | 0.00928 | 1.0441 | 0.024847 | 23.7979 | -0.014031 | 0.14119 | -3.7e-05 | -0.002612 | -0.002612 | -2.2218
2019-03-21 11:42 | 1553182899.3565 | -214 | 0.20953 | 0.00046 | -3e-06 | 101.2251 | 101.435 | 124.0006 | 20.998 | 0.001953 | 0.001788 | -151.6772 | 0.23728 | 8.4033 | 0.003453 | 0.005203 | 0.005033 | 6.7e-05 | 0.001042 | 0.29463 | 1.3016 | 2.5479 | 99.61 | 20.86 | 37.448 | 68.5626 | 104.7018 | 64.5934 | 0.00928 | 1.0453 | 0.024903 | 23.8241 | -0.018532 | 0.14116 | -0.000121 | -0.008824 | -0.008824 | -0.65746
2019-03-21 11:42 | 1553182900.3562 | -213 | 0.20955 | 0.000466 | -4e-06 | 101.2266 | 101.4592 | 124.4332 | 20.9984 | 0.001959 | 0.001802 | -151.6749 | 0.23772 | 8.4045 | 0.003474 | 0.005452 | 0.005286 | 6.7e-05 | 0.001049 | 0.32493 | 1.2694 | 2.5481 | 99.61 | 20.86 | 37.45 | 68.6981 | 105.1939 | 64.942 | 0.009281 | 1.0442 | 0.024934 | 23.8779 | -0.020532 | 0.14142 | -0.000237 | -0.017543 | -0.017543 | -0.33132
2019-03-21 11:42 | 1553182901.3551 | -212 | 0.20955 | 0.000461 | -3e-06 | 101.2653 | 101.5532 | 124.7482 | 20.9988 | 0.00189 | 0.001716 | -151.6817 | 0.23677 | 8.4017 | 0.003489 | 0.004492 | 0.004449 | 6.7e-05 | 0.001054 | 0.31045 | 1.2718 | 2.548 | 99.61 | 20.85 | 37.451 | 67.0484 | 105.5679 | 62.8645 | 0.009275 | 1.0507 | 0.025043 | 23.8346 | -0.025394 | 0.14085 | -0.000236 | -0.017544 | -0.017544 | -0.32997
2019-03-21 11:42 | 1553182902.3563 | -211 | 0.20953 | 0.00046 | -3e-06 | 101.2188 | 101.5001 | 124.8106 | 20.9993 | 0.001887 | 0.001779 | -151.6906 | 0.23617 | 8.4018 | 0.003426 | 0.004958 | 0.004687 | 7e-05 | 0.001058 | 0.34523 | 1.285 | 2.5478 | 99.61 | 20.86 | 37.453 | 66.9804 | 104.0505 | 64.383 | 0.009281 | 1.046 | 0.024992 | 23.8943 | -0.024819 | 0.1405 | -7.9e-05 | -0.005702 | -0.005702 | -1.0127
2019-03-21 11:42 | 1553182903.3553 | -210 | 0.20955 | 0.000465 | -3e-06 | 101.273 | 101.5392 | 124.8397 | 20.999 | 0.001879 | 0.001792 | -151.6843 | 0.23681 | 8.4031 | 0.003553 | 0.00406 | 0.003951 | 6.6e-05 | 0.001059 | 0.36011 | 1.2921 | 2.5478 | 99.61 | 20.86 | 37.453 | 66.7827 | 107.1163 | 64.6825 | 0.009281 | 1.045 | 0.024984 | 23.9077 | -0.02348 | 0.14088 | -0.000242 | -0.01794 | -0.01794 | -0.32276
2019-03-21 11:42 | 1553182904.3564 | -209 | 0.20957 | 0.000462 | -3e-06 | 101.2623 | 101.5208 | 124.9172 | 21.0002 | 0.001902 | 0.001798 | -151.6797 | 0.2365 | 8.4042 | 0.003416 | 0.004842 | 0.004635 | 6.6e-05 | 0.001059 | 0.3246 | 1.3267 | 2.548 | 99.61 | 20.85 | 37.453 | 67.3504 | 103.7897 | 64.8415 | 0.009276 | 1.0445 | 0.024986 | 23.9208 | -0.022807 | 0.14069 | -0.000311 | -0.023138 | -0.023138 | -0.24991
2019-03-21 11:42 | 1553182905.3554 | -208 | 0.20954 | 0.000465 | -3e-06 | 101.2257 | 101.4411 | 125.2665 | 21.0007 | 0.001917 | 0.001758 | -151.669 | 0.23658 | 8.4003 | 0.003537 | 0.004722 | 0.004568 | 6.4e-05 | 0.001059 | 0.37161 | 1.3049 | 2.548 | 99.61 | 20.85 | 37.453 | 67.7071 | 106.7299 | 63.8823 | 0.009276 | 1.0475 | 0.025057 | 23.9202 | -0.01902 | 0.14074 | -0.000133 | -0.009855 | -0.009855 | -0.58693
2019-03-21 11:42 | 1553182906.3555 | -207 | 0.20953 | 0.00046 | -3e-06 | 101.206 | 101.5008 | 125.4113 | 20.9998 | 0.001894 | 0.001845 | -151.679 | 0.2374 | 8.4019 | 0.003546 | 0.004908 | 0.004755 | 6.6e-05 | 0.001059 | 0.32333 | 1.2933 | 2.5479 | 99.61 | 20.85 | 37.452 | 67.1525 | 106.9424 | 65.9782 | 0.009275 | 1.041 | 0.024994 | 24.0083 | -0.026011 | 0.14123 | -9.7e-05 | -0.007096 | -0.007096 | -0.81803
2019-03-21 11:42 | 1553182907.3557 | -206 | 0.20952 | 0.00046 | -3e-06 | 101.2516 | 101.4415 | 125.1784 | 21.0004 | 0.001936 | 0.001839 | -151.6891 | 0.23692 | 8.4017 | 0.003623 | 0.004678 | 0.004459 | 6.7e-05 | 0.001057 | 0.34307 | 1.3352 | 2.548 | 99.61 | 20.85 | 37.45 | 68.1461 | 108.8281 | 65.8133 | 0.009275 | 1.0415 | 0.024976 | 23.9802 | -0.016759 | 0.14094 | -5e-05 | -0.003519 | -0.003519 | -1.646
2019-03-21 11:42 | 1553182908.3546 | -205 | 0.20952 | 0.000463 | -3e-06 | 101.2424 | 101.4207 | 125.1428 | 21.0007 | 0.001967 | 0.001786 | -151.673 | 0.23686 | 8.3995 | 0.003519 | 0.004898 | 0.004728 | 6.2e-05 | 0.001059 | 0.31454 | 1.3002 | 2.5479 | 99.61 | 20.85 | 37.45 | 68.9047 | 106.3032 | 64.5569 | 0.009275 | 1.0454 | 0.025019 | 23.9323 | -0.015743 | 0.14091 | -4.8e-05 | -0.003426 | -0.003426 | -1.6903
2019-03-21 11:42 | 1553182909.356 | -204 | 0.20952 | 0.000462 | -4e-06 | 101.1771 | 101.4163 | 124.9362 | 20.9997 | 0.001927 | 0.001794 | -151.6648 | 0.23659 | 8.401 | 0.003543 | 0.005088 | 0.005049 | 6.5e-05 | 0.001057 | 0.29623 | 1.3251 | 2.548 | 99.61 | 20.85 | 37.45 | 67.9379 | 106.8777 | 64.7356 | 0.009275 | 1.0449 | 0.024992 | 23.9188 | -0.02112 | 0.14075 | -5e-05 | -0.003539 | -0.003539 | -1.6347
2019-03-21 11:42 | 1553182910.3557 | -203 | 0.20952 | 0.000459 | -4e-06 | 101.2197 | 101.4207 | 124.7477 | 21.0016 | 0.001914 | 0.0018 | -151.6608 | 0.23646 | 8.402 | 0.003536 | 0.004722 | 0.004497 | 6.1e-05 | 0.001057 | 0.31116 | 1.3597 | 2.548 | 99.61 | 20.85 | 37.45 | 67.6357 | 106.7078 | 64.8756 | 0.009275 | 1.0444 | 0.024968 | 23.9058 | -0.017744 | 0.14067 | -9e-06 | -0.00045 | -0.00045 | -12.8569
2019-03-21 11:42 | 1553182911.3546 | -202 | 0.20951 | 0.000463 | -3e-06 | 101.1038 | 101.3632 | 124.7339 | 21.0007 | 0.001957 | 0.001771 | -151.6658 | 0.23659 | 8.4047 | 0.00348 | 0.005255 | 0.005093 | 6.3e-05 | 0.001058 | 0.30973 | 1.3339 | 2.548 | 99.61 | 20.85 | 37.45 | 68.6495 | 105.3566 | 64.1978 | 0.009275 | 1.0465 | 0.024992 | 23.8804 | -0.022914 | 0.14075 | 5.3e-05 | 0.004213 | 0.004213 | 1.3729
2019-03-21 11:42 | 1553182912.3547 | -201 | 0.20952 | 0.00046 | -3e-06 | 101.0678 | 101.3012 | 124.8191 | 21.0009 | 0.001878 | 0.001761 | -151.6613 | 0.23678 | 8.4029 | 0.003575 | 0.004746 | 0.004564 | 6.1e-05 | 0.001059 | 0.3027 | 1.3232 | 2.5479 | 99.61 | 20.85 | 37.449 | 66.7619 | 107.6632 | 63.9567 | 0.009275 | 1.0473 | 0.025009 | 23.8801 | -0.020632 | 0.14086 | -1.6e-05 | -0.001012 | -0.001012 | -5.7235
2019-03-21 11:42 | 1553182913.3562 | -200 | 0.20957 | 0.000463 | -3e-06 | 101.1166 | 101.2279 | 124.5754 | 21.0019 | 0.001872 | 0.001778 | -151.6733 | 0.23611 | 8.4032 | 0.003512 | 0.004981 | 0.004885 | 6.3e-05 | 0.001059 | 0.30574 | 1.3322 | 2.5479 | 99.61 | 20.85 | 37.449 | 66.6155 | 106.1196 | 64.3514 | 0.009275 | 1.0461 | 0.02497 | 23.8707 | -0.009841 | 0.14046 | -0.000344 | -0.025651 | -0.025651 | -0.22506
2019-03-21 11:42 | 1553182914.3551 | -199 | 0.20955 | 0.000463 | -3e-06 | 101.0594 | 101.1288 | 124.336 | 21.0018 | 0.001895 | 0.001734 | -151.6704 | 0.23661 | 8.4029 | 0.003601 | 0.005366 | 0.005139 | 6.3e-05 | 0.001057 | 0.32908 | 1.3156 | 2.548 | 99.61 | 20.85 | 37.448 | 67.1677 | 108.2895 | 63.2947 | 0.009274 | 1.0493 | 0.024985 | 23.8104 | -0.006141 | 0.14076 | -0.000245 | -0.018204 | -0.018204 | -0.3178
2019-03-21 11:42 | 1553182915.355 | -198 | 0.20953 | 0.000468 | -3e-06 | 100.9721 | 101.3509 | 123.8929 | 21.0025 | 0.001882 | 0.001752 | -151.6778 | 0.23777 | 8.4014 | 0.003691 | 0.005268 | 0.00516 | 6.4e-05 | 0.00106 | 0.29336 | 1.3307 | 2.5479 | 99.61 | 20.85 | 37.448 | 66.8601 | 110.475 | 63.7334 | 0.009274 | 1.048 | 0.024924 | 23.7834 | -0.03349 | 0.14145 | -0.000108 | -0.007865 | -0.007865 | -0.73915
2019-03-21 11:42 | 1553182916.3553 | -197 | 0.20954 | 0.000458 | -3e-06 | 100.9164 | 101.1111 | 123.9433 | 21.0019 | 0.001893 | 0.00179 | -151.6829 | 0.23837 | 8.4011 | 0.003678 | 0.005246 | 0.005063 | 6e-05 | 0.001061 | 0.3418 | 1.3381 | 2.548 | 99.61 | 20.85 | 37.447 | 67.1171 | 110.1479 | 64.6464 | 0.009274 | 1.0451 | 0.024896 | 23.8205 | -0.017241 | 0.14181 | -0.000129 | -0.009424 | -0.009424 | -0.61848
2019-03-21 11:42 | 1553182917.3551 | -196 | 0.20954 | 0.00046 | -3e-06 | 100.8964 | 101.1857 | 124.3867 | 21.0015 | 0.001894 | 0.001834 | -151.669 | 0.23682 | 8.3982 | 0.003707 | 0.004952 | 0.004778 | 6.3e-05 | 0.001059 | 0.30038 | 1.339 | 2.548 | 99.61 | 20.85 | 37.449 | 67.1574 | 110.8565 | 65.7028 | 0.009275 | 1.0419 | 0.024901 | 23.9003 | -0.025607 | 0.14088 | -0.000138 | -0.01013 | -0.01013 | -0.57159
2019-03-21 11:42 | 1553182918.3556 | -195 | 0.20954 | 0.000463 | -3e-06 | 100.872 | 101.0774 | 124.5111 | 21.0035 | 0.001947 | 0.001731 | -151.6873 | 0.23579 | 8.4048 | 0.00365 | 0.004881 | 0.004739 | 6.4e-05 | 0.001056 | 0.29479 | 1.298 | 2.5479 | 99.61 | 20.85 | 37.448 | 68.4269 | 109.4908 | 63.2268 | 0.009274 | 1.0496 | 0.025005 | 23.8248 | -0.018194 | 0.14027 | -0.000145 | -0.010677 | -0.010677 | -0.53997
2019-03-21 11:42 | 1553182919.3566 | -194 | 0.20954 | 0.00046 | -4e-06 | 100.8324 | 100.9929 | 124.2954 | 21.0025 | 0.001977 | 0.001822 | -151.669 | 0.23733 | 8.4045 | 0.003642 | 0.005184 | 0.005031 | 6.5e-05 | 0.00106 | 0.30845 | 1.3441 | 2.5478 | 99.61 | 20.85 | 37.448 | 69.1386 | 109.2906 | 65.412 | 0.009274 | 1.0428 | 0.024903 | 23.8813 | -0.01423 | 0.14119 | -0.00017 | -0.01247 | -0.01247 | -0.46533
2019-03-21 11:42 | 1553182920.3552 | -193 | 0.20955 | 0.000463 | -3e-06 | 100.7871 | 100.9912 | 124.3154 | 21.003 | 0.001913 | 0.001822 | -151.6609 | 0.23717 | 8.4021 | 0.003648 | 0.005396 | 0.005215 | 6.3e-05 | 0.00106 | 0.29822 | 1.3562 | 2.548 | 99.61 | 20.86 | 37.448 | 67.6089 | 109.432 | 65.4077 | 0.00928 | 1.0428 | 0.024905 | 23.8831 | -0.018098 | 0.1411 | -0.000225 | -0.016613 | -0.016613 | -0.34907
2019-03-21 11:42 | 1553182921.3563 | -192 | 0.20954 | 0.000463 | -3e-06 | 100.7937 | 100.9558 | 124.343 | 21.002 | 0.00195 | 0.001791 | -151.6666 | 0.23662 | 8.3996 | 0.003611 | 0.00496 | 0.004854 | 6.2e-05 | 0.001065 | 0.31166 | 1.3781 | 2.5479 | 99.61 | 20.86 | 37.447 | 68.4848 | 108.5341 | 64.6722 | 0.00928 | 1.0451 | 0.024935 | 23.8598 | -0.014378 | 0.14076 | -0.000131 | -0.009623 | -0.009623 | -0.60118
2019-03-21 11:42 | 1553182922.3557 | -191 | 0.20955 | 0.000464 | -4e-06 | 100.7659 | 100.8158 | 124.4278 | 21.0036 | 0.001903 | 0.001777 | -151.683 | 0.23697 | 8.4009 | 0.003581 | 0.005378 | 0.005133 | 6.3e-05 | 0.001067 | 0.31509 | 1.3822 | 2.5479 | 99.61 | 20.85 | 37.446 | 67.377 | 107.8084 | 64.3388 | 0.009274 | 1.0461 | 0.024956 | 23.8561 | -0.004427 | 0.14097 | -0.000201 | -0.014876 | -0.014876 | -0.38949
2019-03-21 11:42 | 1553182923.3548 | -190 | 0.20955 | 0.00046 | -3e-06 | 100.6778 | 100.7624 | 124.8001 | 21.0039 | 0.001901 | 0.001756 | -151.6563 | 0.23653 | 8.4024 | 0.003522 | 0.004883 | 0.00474 | 6.8e-05 | 0.001068 | 0.3314 | 1.3384 | 2.5481 | 99.61 | 20.86 | 37.443 | 67.3102 | 106.3779 | 63.8315 | 0.009279 | 1.0477 | 0.025012 | 23.8738 | -0.007513 | 0.14071 | -0.000195 | -0.014448 | -0.014448 | -0.40028
2019-03-21 11:42 | 1553182924.3554 | -189 | 0.20957 | 0.000459 | -3e-06 | 100.6024 | 100.8223 | 125.0061 | 21.0023 | 0.001907 | 0.001804 | -151.6493 | 0.23616 | 8.403 | 0.003573 | 0.004979 | 0.004824 | 6.6e-05 | 0.001068 | 0.27959 | 1.3273 | 2.5479 | 99.61 | 20.86 | 37.443 | 67.4716 | 107.6048 | 64.9886 | 0.009279 | 1.0441 | 0.02499 | 23.9345 | -0.019532 | 0.14049 | -0.000323 | -0.024014 | -0.024014 | -0.24046
2019-03-21 11:42 | 1553182925.3544 | -188 | 0.20958 | 0.000465 | -4e-06 | 100.535 | 100.7504 | 125.3087 | 21.0039 | 0.00192 | 0.001765 | -151.6628 | 0.23663 | 8.4024 | 0.003516 | 0.00499 | 0.004848 | 6.6e-05 | 0.001069 | 0.32029 | 1.3197 | 2.548 | 99.61 | 20.86 | 37.438 | 67.7665 | 106.2348 | 64.0346 | 0.009278 | 1.047 | 0.025055 | 23.9296 | -0.019143 | 0.14077 | -0.000381 | -0.028472 | -0.028472 | -0.20321
2019-03-21 11:42 | 1553182926.3555 | -187 | 0.20955 | 0.000461 | -3e-06 | 100.5032 | 100.7273 | 125.4336 | 21.0034 | 0.001959 | 0.001803 | -151.6527 | 0.23608 | 8.4028 | 0.003624 | 0.004589 | 0.004421 | 6.3e-05 | 0.00107 | 0.30807 | 1.3058 | 2.548 | 99.61 | 20.86 | 37.436 | 68.7193 | 108.8539 | 64.9615 | 0.009277 | 1.0442 | 0.025033 | 23.9744 | -0.019924 | 0.14045 | -0.000233 | -0.017294 | -0.017294 | -0.33377
2019-03-21 11:42 | 1553182927.3556 | -186 | 0.20954 | 0.000454 | -3e-06 | 100.4591 | 100.6047 | 125.3629 | 21.0033 | 0.001882 | 0.001794 | -151.6393 | 0.23604 | 8.4024 | 0.003555 | 0.004403 | 0.00426 | 6.3e-05 | 0.001071 | 0.32532 | 1.3399 | 2.548 | 99.61 | 20.86 | 37.433 | 66.8664 | 107.1844 | 64.735 | 0.009276 | 1.0449 | 0.025035 | 23.9596 | -0.012954 | 0.14042 | -0.000151 | -0.01113 | -0.01113 | -0.51853
2019-03-21 11:42 | 1553182928.3557 | -185 | 0.20958 | 0.000461 | -3e-06 | 100.3813 | 100.6019 | 125.4197 | 21.0027 | 0.001903 | 0.001826 | -151.6355 | 0.23666 | 8.4031 | 0.0037 | 0.004686 | 0.004501 | 6.6e-05 | 0.001072 | 0.29623 | 1.33 | 2.5479 | 99.61 | 20.86 | 37.429 | 67.3761 | 110.6936 | 65.5096 | 0.009275 | 1.0425 | 0.025012 | 23.9925 | -0.019639 | 0.14079 | -0.000417 | -0.031131 | -0.031131 | -0.18587
2019-03-21 11:42 | 1553182929.3552 | -184 | 0.20957 | 0.00046 | -4e-06 | 100.3289 | 100.339 | 125.6913 | 21.0043 | 0.001979 | 0.001796 | -151.6541 | 0.23582 | 8.4016 | 0.003625 | 0.005083 | 0.00495 | 6.8e-05 | 0.00107 | 0.30127 | 1.3105 | 2.5479 | 99.61 | 20.86 | 37.427 | 69.2004 | 108.88 | 64.7961 | 0.009275 | 1.0447 | 0.025065 | 23.9932 | -0.000904 | 0.14029 | -0.000313 | -0.023324 | -0.023324 | -0.24721
2019-03-21 11:42 | 1553182930.3557 | -183 | 0.20958 | 0.000462 | -3e-06 | 100.3125 | 100.3802 | 125.5322 | 21.0039 | 0.001899 | 0.001799 | -151.6718 | 0.23731 | 8.4015 | 0.003618 | 0.00493 | 0.00469 | 6.3e-05 | 0.001069 | 0.32029 | 1.3149 | 2.548 | 99.61 | 20.86 | 37.426 | 67.2609 | 108.7132 | 64.8761 | 0.009275 | 1.0444 | 0.025046 | 23.9808 | -0.006038 | 0.14118 | -0.000377 | -0.028118 | -0.028118 | -0.20636
2019-03-21 11:42 | 1553182931.3561 | -182 | 0.20958 | 0.000462 | -3e-06 | 100.255 | 100.4337 | 125.3951 | 21.0042 | 0.001932 | 0.001763 | -151.6625 | 0.23669 | 8.4036 | 0.003589 | 0.004822 | 0.004658 | 6.5e-05 | 0.001072 | 0.33931 | 1.2996 | 2.548 | 99.61 | 20.86 | 37.424 | 68.057 | 107.9902 | 64.0075 | 0.009274 | 1.0471 | 0.025065 | 23.9369 | -0.015931 | 0.14081 | -0.000394 | -0.029436 | -0.029436 | -0.1966
2019-03-21 11:42 | 1553182932.3548 | -181 | 0.2095 | 0.00046 | -3e-06 | 100.1829 | 100.3543 | 125.0801 | 21.0044 | 0.001921 | 0.001826 | -151.6621 | 0.23595 | 8.401 | 0.003649 | 0.004713 | 0.00455 | 6.7e-05 | 0.001076 | 0.31598 | 1.3729 | 2.548 | 99.61 | 20.86 | 37.423 | 67.8098 | 109.4498 | 65.5237 | 0.009274 | 1.0424 | 0.024977 | 23.9605 | -0.015295 | 0.14036 | 8.2e-05 | 0.006331 | 0.006331 | 0.91118
2019-03-21 11:42 | 1553182933.355 | -180 | 0.20964 | 0.00046 | -4e-06 | 100.1891 | 100.2116 | 124.8282 | 21.0047 | 0.001922 | 0.001755 | -151.6849 | 0.23635 | 8.4012 | 0.003609 | 0.004991 | 0.00488 | 6.7e-05 | 0.001073 | 0.30917 | 1.3751 | 2.548 | 99.61 | 20.86 | 37.424 | 67.8125 | 108.483 | 63.802 | 0.009274 | 1.0478 | 0.025016 | 23.8755 | -0.002008 | 0.1406 | -0.000733 | -0.054886 | -0.054886 | -0.10529
2019-03-21 11:42 | 1553182934.355 | -179 | 0.20957 | 0.000463 | -3e-06 | 100.1471 | 100.3063 | 124.6102 | 21.0042 | 0.001949 | 0.001796 | -151.6804 | 0.23677 | 8.4006 | 0.003565 | 0.005022 | 0.004742 | 6.4e-05 | 0.001075 | 0.31045 | 1.3008 | 2.5478 | 99.61 | 20.86 | 37.425 | 68.4797 | 107.4127 | 64.8016 | 0.009274 | 1.0447 | 0.024957 | 23.89 | -0.014207 | 0.14086 | -0.000352 | -0.026217 | -0.026217 | -0.22082
2019-03-21 11:42 | 1553182935.355 | -178 | 0.20956 | 0.000459 | -3e-06 | 100.1108 | 100.3557 | 124.635 | 21.0035 | 0.001896 | 0.001807 | -151.654 | 0.23639 | 8.4015 | 0.00348 | 0.004488 | 0.004362 | 6.5e-05 | 0.001079 | 0.33395 | 1.3114 | 2.5479 | 99.61 | 20.86 | 37.426 | 67.1965 | 105.3576 | 65.0593 | 0.009275 | 1.0439 | 0.02495 | 23.9014 | -0.021857 | 0.14063 | -0.000253 | -0.018764 | -0.018764 | -0.30803
2019-03-21 11:42 | 1553182936.3561 | -177 | 0.20952 | 0.00046 | -3e-06 | 100.0509 | 100.1469 | 124.4977 | 21.0049 | 0.001911 | 0.001786 | -151.6534 | 0.23664 | 8.4015 | 0.003509 | 0.004838 | 0.004614 | 6.4e-05 | 0.001073 | 0.32693 | 1.3132 | 2.5479 | 99.61 | 20.86 | 37.427 | 67.5576 | 106.0657 | 64.5477 | 0.009275 | 1.0454 | 0.024955 | 23.8702 | -0.008581 | 0.14078 | 2e-06 | 0.000295 | 0.000295 | 19.6171
2019-03-21 11:42 | 1553182937.3561 | -176 | 0.20951 | 0.000461 | -3e-06 | 100.0411 | 99.9896 | 124.3884 | 21.0057 | 0.001905 | 0.001773 | -151.6572 | 0.23692 | 8.3998 | 0.003485 | 0.004967 | 0.004813 | 6.5e-05 | 0.001078 | 0.29645 | 1.3011 | 2.5479 | 99.61 | 20.86 | 37.43 | 67.4187 | 105.4747 | 64.2347 | 0.009276 | 1.0464 | 0.024956 | 23.8487 | 0.004604 | 0.14094 | 2.7e-05 | 0.002226 | 0.002226 | 2.6027
2019-03-21 11:42 | 1553182938.3556 | -175 | 0.20955 | 0.000464 | -3e-06 | 100.0431 | 99.8615 | 124.305 | 21.0056 | 0.00191 | 0.001754 | -151.6441 | 0.23688 | 8.4038 | 0.003457 | 0.004763 | 0.004639 | 6.5e-05 | 0.001078 | 0.31188 | 1.2989 | 2.5479 | 99.61 | 20.86 | 37.433 | 67.5326 | 104.8091 | 63.7949 | 0.009276 | 1.0478 | 0.024964 | 23.8251 | 0.016254 | 0.14092 | -0.000177 | -0.013108 | -0.013108 | -0.44186
2019-03-21 11:42 | 1553182939.3556 | -174 | 0.20952 | 0.000461 | -3e-06 | 99.9263 | 99.9664 | 124.2114 | 21.0058 | 0.001909 | 0.001781 | -151.6336 | 0.23745 | 8.4012 | 0.003584 | 0.004684 | 0.004538 | 6.5e-05 | 0.001074 | 0.30973 | 1.2815 | 2.5479 | 99.61 | 20.86 | 37.436 | 67.5157 | 107.8806 | 64.43 | 0.009277 | 1.0458 | 0.024931 | 23.8386 | -0.003585 | 0.14126 | -4.1e-05 | -0.002864 | -0.002864 | -2.0271
2019-03-21 11:42 | 1553182940.3543 | -173 | 0.20954 | 0.00046 | -3e-06 | 99.8578 | 100.0172 | 124.201 | 21.0057 | 0.001918 | 0.001753 | -151.6302 | 0.23616 | 8.401 | 0.003581 | 0.005068 | 0.004974 | 6.4e-05 | 0.001077 | 0.31957 | 1.2883 | 2.5479 | 99.61 | 20.86 | 37.441 | 67.7183 | 107.8072 | 63.7654 | 0.009278 | 1.0479 | 0.024954 | 23.8141 | -0.014261 | 0.14049 | -0.000119 | -0.008685 | -0.008685 | -0.66484
2019-03-21 11:42 | 1553182941.3547 | -172 | 0.20956 | 0.000463 | -3e-06 | 99.8814 | 100.027 | 124.3186 | 21.0053 | 0.001881 | 0.001791 | -151.6314 | 0.23673 | 8.4019 | 0.003585 | 0.004356 | 0.00415 | 6.5e-05 | 0.001077 | 0.28727 | 1.261 | 2.548 | 99.61 | 20.86 | 37.443 | 66.8318 | 107.8952 | 64.6848 | 0.009279 | 1.045 | 0.024932 | 23.8579 | -0.013036 | 0.14083 | -0.000285 | -0.021094 | -0.021094 | -0.2744
2019-03-21 11:42 | 1553182942.3554 | -171 | 0.20952 | 0.00046 | -4e-06 | 99.803 | 99.9279 | 124.5466 | 21.0069 | 0.001924 | 0.001831 | -151.6236 | 0.23621 | 8.4025 | 0.003547 | 0.004683 | 0.004487 | 6.2e-05 | 0.001076 | 0.29717 | 1.2829 | 2.548 | 99.61 | 20.86 | 37.447 | 67.8748 | 106.9901 | 65.6371 | 0.00928 | 1.0421 | 0.02492 | 23.9134 | -0.011189 | 0.14052 | -4.6e-05 | -0.003244 | -0.003244 | -1.7805
2019-03-21 11:42 | 1553182943.3559 | -170 | 0.20957 | 0.000459 | -3e-06 | 99.8143 | 99.8455 | 124.7799 | 21.006 | 0.001865 | 0.001819 | -151.6098 | 0.23614 | 8.4002 | 0.003499 | 0.005018 | 0.004906 | 6.6e-05 | 0.001078 | 0.31437 | 1.2781 | 2.548 | 99.61 | 20.86 | 37.45 | 66.464 | 105.8094 | 65.3379 | 0.009281 | 1.043 | 0.024954 | 23.9252 | -0.002794 | 0.14048 | -0.000285 | -0.021174 | -0.021174 | -0.27268
2019-03-21 11:42 | 1553182944.3561 | -169 | 0.20952 | 0.000459 | -3e-06 | 99.8098 | 99.9562 | 124.9693 | 21.0071 | 0.001888 | 0.00178 | -151.633 | 0.23599 | 8.4025 | 0.003666 | 0.005582 | 0.005372 | 6.2e-05 | 0.001078 | 0.323 | 1.3068 | 2.548 | 99.61 | 20.86 | 37.454 | 66.9981 | 109.8642 | 64.411 | 0.009282 | 1.0459 | 0.025007 | 23.9105 | -0.013105 | 0.14039 | -4.2e-05 | -0.002975 | -0.002975 | -1.9393
2019-03-21 11:42 | 1553182945.3555 | -168 | 0.20956 | 0.000458 | -4e-06 | 99.8224 | 99.8578 | 124.9308 | 21.0066 | 0.001956 | 0.001764 | -151.6446 | 0.23643 | 8.4025 | 0.003586 | 0.005413 | 0.005334 | 6.4e-05 | 0.001076 | 0.31924 | 1.3176 | 2.5479 | 99.61 | 20.86 | 37.457 | 68.6396 | 107.9396 | 64.024 | 0.009282 | 1.0471 | 0.025018 | 23.8931 | -0.003171 | 0.14066 | -0.000232 | -0.017197 | -0.017197 | -0.33615
2019-03-21 11:42 | 1553182946.3556 | -167 | 0.20955 | 0.000462 | -4e-06 | 99.7747 | 99.9535 | 125.0198 | 21.0064 | 0.001886 | 0.001789 | -151.6325 | 0.23699 | 8.404 | 0.003528 | 0.004769 | 0.004537 | 6.2e-05 | 0.001077 | 0.33284 | 1.3305 | 2.5479 | 99.61 | 20.86 | 37.459 | 66.964 | 106.5291 | 64.6305 | 0.009283 | 1.0452 | 0.025004 | 23.9231 | -0.016012 | 0.14098 | -0.000193 | -0.014238 | -0.014238 | -0.40697
2019-03-21 11:42 | 1553182947.3557 | -166 | 0.20955 | 0.000459 | -4e-06 | 99.7324 | 99.6306 | 125.1982 | 21.0071 | 0.001925 | 0.001823 | -151.6389 | 0.23673 | 8.4036 | 0.003656 | 0.00467 | 0.004567 | 6.4e-05 | 0.001079 | 0.32261 | 1.3355 | 2.5479 | 99.61 | 20.86 | 37.464 | 67.9094 | 109.6348 | 65.4461 | 0.009284 | 1.0427 | 0.024992 | 23.9691 | 0.009134 | 0.14083 | -0.000187 | -0.013767 | -0.013767 | -0.42044
2019-03-21 11:42 | 1553182948.3561 | -165 | 0.20956 | 0.000461 | -3e-06 | 99.6847 | 99.7535 | 125.3507 | 21.0077 | 0.00188 | 0.001824 | -151.6506 | 0.23797 | 8.4007 | 0.003552 | 0.004668 | 0.004419 | 6.5e-05 | 0.001075 | 0.31166 | 1.2989 | 2.5479 | 99.61 | 20.86 | 37.466 | 66.8279 | 107.0933 | 65.47 | 0.009285 | 1.0426 | 0.025006 | 23.9845 | -0.006174 | 0.14157 | -0.000271 | -0.020139 | -0.020139 | -0.28892
2019-03-21 11:42 | 1553182949.3558 | -164 | 0.20954 | 0.00046 | -3e-06 | 99.6761 | 99.7546 | 125.4106 | 21.0086 | 0.001982 | 0.001756 | -151.6569 | 0.23576 | 8.4017 | 0.003507 | 0.004232 | 0.004128 | 6.6e-05 | 0.001079 | 0.3194 | 1.3193 | 2.5479 | 99.61 | 20.86 | 37.467 | 69.2726 | 106.0189 | 63.8291 | 0.009285 | 1.0477 | 0.025073 | 23.932 | -0.007041 | 0.14026 | -0.000113 | -0.008294 | -0.008294 | -0.69504
2019-03-21 11:42 | 1553182950.3548 | -163 | 0.20953 | 0.000461 | -4e-06 | 99.6457 | 99.5768 | 125.6582 | 21.0088 | 0.001956 | 0.001789 | -151.6501 | 0.2367 | 8.4021 | 0.003522 | 0.004389 | 0.004192 | 6.5e-05 | 0.001077 | 0.30646 | 1.359 | 2.5479 | 99.61 | 20.86 | 37.47 | 68.6543 | 106.382 | 64.6328 | 0.009286 | 1.0452 | 0.025068 | 23.9842 | 0.006187 | 0.14081 | -0.000104 | -0.007625 | -0.007625 | -0.75897
2019-03-21 11:42 | 1553182951.3548 | -162 | 0.20955 | 0.00046 | -4e-06 | 99.6209 | 99.6309 | 125.6983 | 21.0085 | 0.001891 | 0.001822 | -151.6489 | 0.23691 | 8.4046 | 0.003445 | 0.00473 | 0.004498 | 6.4e-05 | 0.001083 | 0.33268 | 1.3115 | 2.5479 | 99.61 | 20.86 | 37.471 | 67.0939 | 104.5148 | 65.4325 | 0.009286 | 1.0427 | 0.025042 | 24.0164 | -0.000896 | 0.14094 | -0.000193 | -0.014299 | -0.014299 | -0.4051
2019-03-21 11:43 | 1553182952.356 | -161 | 0.20955 | 0.000453 | -4e-06 | 99.6478 | 99.8094 | 125.4972 | 21.0086 | 0.001905 | 0.001807 | -151.6521 | 0.23675 | 8.4008 | 0.003464 | 0.005307 | 0.005167 | 6.3e-05 | 0.001082 | 0.34987 | 1.3397 | 2.5481 | 99.61 | 20.86 | 37.473 | 67.4146 | 104.9782 | 65.0643 | 0.009286 | 1.0439 | 0.025036 | 23.9841 | -0.014496 | 0.14084 | -0.000189 | -0.014005 | -0.014005 | -0.41333
2019-03-21 11:43 | 1553182953.356 | -160 | 0.20955 | 0.000459 | -4e-06 | 99.5843 | 99.6272 | 125.0904 | 21.0078 | 0.00195 | 0.001847 | -151.6566 | 0.23682 | 8.4025 | 0.003517 | 0.005191 | 0.005028 | 6.5e-05 | 0.001083 | 0.31996 | 1.3162 | 2.548 | 99.61 | 20.86 | 37.472 | 68.4925 | 106.2628 | 66.0343 | 0.009286 | 1.0409 | 0.02496 | 23.9796 | -0.00385 | 0.14088 | -0.00018 | -0.013264 | -0.013264 | -0.43653
2019-03-21 11:43 | 1553182954.3554 | -159 | 0.20957 | 0.000464 | -4e-06 | 99.5205 | 99.445 | 124.8432 | 21.0084 | 0.001955 | 0.001788 | -151.6617 | 0.2365 | 8.4001 | 0.00347 | 0.00463 | 0.004506 | 6.3e-05 | 0.001082 | 0.2891 | 1.3181 | 2.5479 | 99.61 | 20.86 | 37.471 | 68.6234 | 105.1101 | 64.5964 | 0.009286 | 1.0453 | 0.024988 | 23.905 | 0.006795 | 0.14069 | -0.000301 | -0.022329 | -0.022329 | -0.25896
2019-03-21 11:43 | 1553182955.3551 | -158 | 0.20956 | 0.000459 | -4e-06 | 99.5325 | 99.4865 | 124.6062 | 21.0089 | 0.001961 | 0.001805 | -151.6633 | 0.23526 | 8.4053 | 0.003433 | 0.004442 | 0.004134 | 6.5e-05 | 0.001084 | 0.31669 | 1.3238 | 2.5479 | 99.61 | 20.86 | 37.468 | 68.7635 | 104.2125 | 65.0184 | 0.009285 | 1.044 | 0.024949 | 23.8972 | 0.00413 | 0.13995 | -0.000233 | -0.017191 | -0.017191 | -0.33461
2019-03-21 11:43 | 1553182956.3542 | -157 | 0.20957 | 0.000458 | -4e-06 | 99.4472 | 99.6166 | 124.2994 | 21.0094 | 0.001931 | 0.001805 | -151.6718 | 0.2363 | 8.4008 | 0.003459 | 0.004845 | 0.00468 | 6.5e-05 | 0.001084 | 0.29574 | 1.3414 | 2.5479 | 99.61 | 20.86 | 37.469 | 68.0388 | 104.8388 | 65.0025 | 0.009285 | 1.044 | 0.024918 | 23.8672 | -0.015221 | 0.14058 | -0.000328 | -0.024317 | -0.024317 | -0.2376
2019-03-21 11:43 | 1553182957.3555 | -156 | 0.20961 | 0.000463 | -4e-06 | 99.4773 | 99.4753 | 124.1879 | 21.0094 | 0.001925 | 0.001781 | -151.6698 | 0.23627 | 8.4034 | 0.003468 | 0.004851 | 0.004734 | 6.2e-05 | 0.001087 | 0.34899 | 1.3322 | 2.5479 | 99.61 | 20.86 | 37.468 | 67.9061 | 105.0757 | 64.4383 | 0.009285 | 1.0458 | 0.024928 | 23.8366 | 0.000186 | 0.14056 | -0.000532 | -0.039542 | -0.039542 | -0.14609
2019-03-21 11:43 | 1553182958.3543 | -155 | 0.20954 | 0.000462 | -4e-06 | 99.4481 | 99.386 | 123.8858 | 21.0108 | 0.001967 | 0.001763 | -151.6704 | 0.2355 | 8.4041 | 0.003454 | 0.004546 | 0.004433 | 6.4e-05 | 0.001088 | 0.30646 | 1.3228 | 2.5481 | 99.61 | 20.86 | 37.467 | 68.919 | 104.7192 | 64.007 | 0.009285 | 1.0471 | 0.024914 | 23.7924 | 0.005587 | 0.1401 | -0.000124 | -0.009001 | -0.009001 | -0.63968
2019-03-21 11:43 | 1553182959.3543 | -154 | 0.20957 | 0.00046 | -4e-06 | 99.4988 | 99.3837 | 123.6661 | 21.0105 | 0.001941 | 0.001803 | -151.6619 | 0.23588 | 8.4017 | 0.003378 | 0.004985 | 0.004754 | 6.6e-05 | 0.001086 | 0.30055 | 1.3219 | 2.5479 | 99.61 | 20.86 | 37.466 | 68.2862 | 102.8827 | 64.9559 | 0.009285 | 1.0442 | 0.024857 | 23.8047 | 0.010356 | 0.14033 | -0.000286 | -0.021084 | -0.021084 | -0.27354
2019-03-21 11:43 | 1553182960.3544 | -153 | 0.20955 | 0.000455 | -3e-06 | 99.4359 | 99.3057 | 123.3999 | 21.0097 | 0.001915 | 0.00181 | -151.6692 | 0.23675 | 8.4026 | 0.003396 | 0.004281 | 0.004095 | 6.7e-05 | 0.001094 | 0.31421 | 1.309 | 2.5479 | 99.61 | 20.86 | 37.467 | 67.6599 | 103.314 | 65.1457 | 0.009285 | 1.0436 | 0.024823 | 23.7858 | 0.011725 | 0.14084 | -0.000164 | -0.01203 | -0.01203 | -0.4812
2019-03-21 11:43 | 1553182961.3558 | -152 | 0.20956 | 0.000457 | -3e-06 | 99.3862 | 99.3864 | 123.4121 | 21.0105 | 0.001891 | 0.001838 | -151.6679 | 0.23619 | 8.4034 | 0.003437 | 0.004812 | 0.004582 | 6.4e-05 | 0.001095 | 0.30773 | 1.3813 | 2.548 | 99.61 | 20.86 | 37.468 | 67.0889 | 104.3233 | 65.8005 | 0.009285 | 1.0416 | 0.0248 | 23.8099 | -1.9e-05 | 0.14051 | -0.000276 | -0.020344 | -0.020344 | -0.28386
2019-03-21 11:43 | 1553182962.3559 | -151 | 0.20956 | 0.000464 | -4e-06 | 99.3427 | 99.3574 | 123.9093 | 21.0114 | 0.001908 | 0.001808 | -151.6535 | 0.23658 | 8.4048 | 0.003467 | 0.005167 | 0.004977 | 6.4e-05 | 0.001091 | 0.28567 | 1.3431 | 2.548 | 99.61 | 20.86 | 37.469 | 67.4906 | 105.0531 | 65.099 | 0.009285 | 1.0437 | 0.024876 | 23.8331 | -0.001329 | 0.14074 | -0.000285 | -0.021036 | -0.021036 | -0.27497
2019-03-21 11:43 | 1553182963.3545 | -150 | 0.20956 | 0.00046 | -4e-06 | 99.2306 | 99.3247 | 124.2215 | 21.0109 | 0.001953 | 0.001784 | -151.6507 | 0.23744 | 8.4027 | 0.003395 | 0.004747 | 0.00456 | 6.6e-05 | 0.001091 | 0.30596 | 1.3162 | 2.548 | 99.61 | 20.86 | 37.472 | 68.5813 | 103.3018 | 64.502 | 0.009286 | 1.0456 | 0.024929 | 23.8421 | -0.008477 | 0.14125 | -0.000236 | -0.017446 | -0.017446 | -0.33276
2019-03-21 11:43 | 1553182964.356 | -149 | 0.20954 | 0.000461 | -4e-06 | 99.3212 | 99.3496 | 124.5905 | 21.0118 | 0.001962 | 0.001787 | -151.6339 | 0.2369 | 8.4026 | 0.003465 | 0.00448 | 0.00438 | 6.5e-05 | 0.001087 | 0.33339 | 1.2677 | 2.5478 | 99.61 | 20.86 | 37.474 | 68.7945 | 104.9914 | 64.5818 | 0.009287 | 1.0453 | 0.024963 | 23.8803 | -0.002556 | 0.14093 | -0.000147 | -0.010809 | -0.010809 | -0.53585
2019-03-21 11:43 | 1553182965.3546 | -148 | 0.20955 | 0.000463 | -3e-06 | 99.3617 | 99.4208 | 124.8428 | 21.0118 | 0.001929 | 0.001834 | -151.6444 | 0.23614 | 8.4025 | 0.003508 | 0.004684 | 0.004476 | 6.2e-05 | 0.001086 | 0.31492 | 1.2625 | 2.5479 | 99.61 | 20.86 | 37.477 | 67.9872 | 106.0497 | 65.7174 | 0.009287 | 1.0418 | 0.024946 | 23.9446 | -0.005314 | 0.14048 | -0.000196 | -0.014457 | -0.014457 | -0.39937
2019-03-21 11:43 | 1553182966.3547 | -147 | 0.20955 | 0.000463 | -3e-06 | 99.2956 | 99.2658 | 124.9296 | 21.0118 | 0.001914 | 0.001853 | -151.6512 | 0.23641 | 8.4032 | 0.003519 | 0.003986 | 0.003861 | 6.2e-05 | 0.001087 | 0.32892 | 1.2885 | 2.5478 | 99.61 | 20.86 | 37.48 | 67.6268 | 106.3165 | 66.1717 | 0.009288 | 1.0404 | 0.024938 | 23.969 | 0.002683 | 0.14064 | -0.000196 | -0.014446 | -0.014446 | -0.40014
2019-03-21 11:43 | 1553182967.356 | -146 | 0.20957 | 0.000459 | -3e-06 | 99.287 | 99.1766 | 124.8322 | 21.0116 | 0.00193 | 0.001802 | -151.6444 | 0.23641 | 8.4039 | 0.003416 | 0.004707 | 0.004559 | 6.7e-05 | 0.001085 | 0.31813 | 1.3046 | 2.5479 | 99.61 | 20.86 | 37.48 | 68.0166 | 103.8212 | 64.9549 | 0.009288 | 1.0442 | 0.024974 | 23.9166 | 0.009949 | 0.14064 | -0.00029 | -0.021519 | -0.021519 | -0.26862
2019-03-21 11:43 | 1553182968.3548 | -145 | 0.20958 | 0.00046 | -3e-06 | 99.1413 | 99.0475 | 124.8336 | 21.0126 | 0.001951 | 0.001803 | -151.6447 | 0.2372 | 8.4038 | 0.003379 | 0.004674 | 0.004466 | 6.4e-05 | 0.001083 | 0.30663 | 1.308 | 2.5479 | 99.61 | 20.86 | 37.481 | 68.5183 | 102.924 | 64.9589 | 0.009288 | 1.0442 | 0.024973 | 23.9169 | 0.008463 | 0.14111 | -0.00037 | -0.027579 | -0.027579 | -0.21029
2019-03-21 11:43 | 1553182969.3558 | -144 | 0.20955 | 0.000463 | -3e-06 | 99.1711 | 99.2103 | 124.5794 | 21.0127 | 0.001893 | 0.001773 | -151.6587 | 0.23694 | 8.4001 | 0.003457 | 0.004968 | 0.004789 | 6.4e-05 | 0.001081 | 0.32316 | 1.3046 | 2.5479 | 99.61 | 20.86 | 37.482 | 67.1345 | 104.8059 | 64.2502 | 0.009289 | 1.0464 | 0.024974 | 23.8675 | -0.003539 | 0.14096 | -0.000218 | -0.016174 | -0.016174 | -0.35819
2019-03-21 11:43 | 1553182970.3549 | -143 | 0.20957 | 0.000464 | -3e-06 | 99.1073 | 99.1807 | 124.347 | 21.0116 | 0.001903 | 0.001797 | -151.6545 | 0.2364 | 8.4019 | 0.003488 | 0.004582 | 0.004445 | 6.2e-05 | 0.001078 | 0.31581 | 1.2998 | 2.5481 | 99.61 | 20.86 | 37.482 | 67.3753 | 105.5632 | 64.8314 | 0.009289 | 1.0446 | 0.024929 | 23.8658 | -0.00662 | 0.14064 | -0.000323 | -0.023995 | -0.023995 | -0.24089
2019-03-21 11:43 | 1553182971.354 | -142 | 0.20957 | 0.000459 | -3e-06 | 99.0847 | 98.913 | 124.0942 | 21.0125 | 0.001909 | 0.001803 | -151.6547 | 0.23721 | 8.403 | 0.003518 | 0.005255 | 0.005125 | 6e-05 | 0.001077 | 0.32853 | 1.3065 | 2.548 | 99.61 | 20.85 | 37.482 | 67.5104 | 106.2808 | 64.9689 | 0.009283 | 1.0441 | 0.024899 | 23.8463 | 0.015512 | 0.14111 | -0.000286 | -0.021179 | -0.021179 | -0.27385
2019-03-21 11:43 | 1553182972.3554 | -141 | 0.20957 | 0.00046 | -3e-06 | 99.0397 | 99.0611 | 123.5931 | 21.0119 | 0.001964 | 0.001735 | -151.659 | 0.23654 | 8.4035 | 0.003431 | 0.004713 | 0.004416 | 6.1e-05 | 0.001071 | 0.3011 | 1.3181 | 2.548 | 99.61 | 20.85 | 37.482 | 68.8467 | 104.1658 | 63.3233 | 0.009283 | 1.0493 | 0.024909 | 23.7402 | -0.001938 | 0.14072 | -0.000283 | -0.020933 | -0.020933 | -0.27628
2019-03-21 11:43 | 1553182973.354 | -140 | 0.20956 | 0.00046 | -3e-06 | 99.0117 | 98.9981 | 123.4175 | 21.0122 | 0.001952 | 0.001779 | -151.6586 | 0.23688 | 8.4022 | 0.003521 | 0.005171 | 0.00505 | 6e-05 | 0.001068 | 0.32244 | 1.3509 | 2.5479 | 99.61 | 20.85 | 37.48 | 68.5483 | 106.3453 | 64.3996 | 0.009282 | 1.0459 | 0.024852 | 23.7612 | 0.001223 | 0.14092 | -0.000234 | -0.017282 | -0.017282 | -0.33512
2019-03-21 11:43 | 1553182974.3551 | -139 | 0.20956 | 0.000462 | -3e-06 | 98.9038 | 98.9733 | 123.0666 | 21.0124 | 0.001901 | 0.0018 | -151.6597 | 0.23671 | 8.4039 | 0.003519 | 0.004288 | 0.004102 | 6.2e-05 | 0.001067 | 0.31454 | 1.3263 | 2.548 | 99.61 | 20.85 | 37.48 | 67.3209 | 106.3193 | 64.901 | 0.009282 | 1.0444 | 0.024798 | 23.745 | -0.006279 | 0.14082 | -0.000249 | -0.018319 | -0.018319 | -0.31593
2019-03-21 11:43 | 1553182975.3544 | -138 | 0.20956 | 0.000464 | -4e-06 | 98.7742 | 98.71 | 122.9253 | 21.0119 | 0.00199 | 0.001764 | -151.638 | 0.23706 | 8.4007 | 0.003545 | 0.004723 | 0.004521 | 6e-05 | 0.00107 | 0.33804 | 1.3136 | 2.5479 | 99.61 | 20.85 | 37.479 | 69.4629 | 106.9354 | 64.0213 | 0.009282 | 1.0471 | 0.024816 | 23.7005 | 0.005817 | 0.14103 | -0.000262 | -0.019279 | -0.019279 | -0.30064
2019-03-21 11:43 | 1553182976.3552 | -137 | 0.20956 | 0.000462 | -3e-06 | 98.7918 | 98.7042 | 122.9995 | 21.0123 | 0.001913 | 0.001796 | -151.6425 | 0.23653 | 8.4009 | 0.003581 | 0.00453 | 0.004352 | 6.1e-05 | 0.001069 | 0.29927 | 1.2926 | 2.548 | 99.61 | 20.85 | 37.478 | 67.6194 | 107.8013 | 64.7924 | 0.009282 | 1.0447 | 0.024795 | 23.7347 | 0.007933 | 0.14071 | -0.000241 | -0.017761 | -0.017761 | -0.32561
2019-03-21 11:43 | 1553182977.3556 | -136 | 0.20957 | 0.00046 | -3e-06 | 98.7325 | 98.508 | 123.0941 | 21.0129 | 0.001944 | 0.001805 | -151.6502 | 0.23624 | 8.4005 | 0.003554 | 0.00555 | 0.0054 | 5.8e-05 | 0.001067 | 0.31581 | 1.3154 | 2.548 | 99.61 | 20.85 | 37.478 | 68.3602 | 107.1598 | 65.0122 | 0.009282 | 1.044 | 0.024797 | 23.7516 | 0.020359 | 0.14054 | -0.000308 | -0.022738 | -0.022738 | -0.25403
2019-03-21 11:43 | 1553182978.3559 | -135 | 0.20957 | 0.000462 | -3e-06 | 98.6613 | 98.4048 | 123.2129 | 21.0138 | 0.001893 | 0.001793 | -151.6316 | 0.23666 | 8.4019 | 0.003523 | 0.004619 | 0.004493 | 5.8e-05 | 0.001072 | 0.3235 | 1.2644 | 2.548 | 99.61 | 20.85 | 37.477 | 67.1325 | 106.4168 | 64.7337 | 0.009282 | 1.0449 | 0.024819 | 23.7532 | 0.023281 | 0.14079 | -0.000295 | -0.021786 | -0.021786 | -0.2656
2019-03-21 11:43 | 1553182979.3548 | -134 | 0.20958 | 0.00046 | -3e-06 | 98.6154 | 98.4314 | 123.8123 | 21.0116 | 0.001907 | 0.001756 | -151.6237 | 0.23669 | 8.4025 | 0.003497 | 0.005142 | 0.004923 | 6.2e-05 | 0.001069 | 0.32173 | 1.3204 | 2.548 | 99.61 | 20.85 | 37.476 | 67.4814 | 105.7725 | 63.8351 | 0.009281 | 1.0477 | 0.024913 | 23.7793 | 0.016706 | 0.14081 | -0.000348 | -0.02585 | -0.02585 | -0.22387
2019-03-21 11:43 | 1553182980.3548 | -133 | 0.20956 | 0.000462 | -4e-06 | 98.4479 | 98.4351 | 124.2438 | 21.014 | 0.001916 | 0.001859 | -151.6181 | 0.23616 | 8.4025 | 0.003574 | 0.004643 | 0.00454 | 6e-05 | 0.001069 | 0.31758 | 1.3519 | 2.5479 | 99.61 | 20.85 | 37.476 | 67.6797 | 107.6335 | 66.3283 | 0.009281 | 1.04 | 0.024864 | 23.9086 | 0.001165 | 0.14049 | -0.000249 | -0.01841 | -0.01841 | -0.31364
2019-03-21 11:43 | 1553182981.3555 | -132 | 0.20961 | 0.000462 | -3e-06 | 98.3317 | 98.3905 | 124.3254 | 21.013 | 0.001929 | 0.001807 | -151.6285 | 0.23628 | 8.4026 | 0.003478 | 0.004882 | 0.004664 | 6.4e-05 | 0.001069 | 0.32709 | 1.3357 | 2.5479 | 99.61 | 20.85 | 37.477 | 68.004 | 105.3056 | 65.063 | 0.009282 | 1.0439 | 0.024919 | 23.8719 | -0.005343 | 0.14056 | -0.000556 | -0.041411 | -0.041411 | -0.13951
2019-03-21 11:43 | 1553182982.3549 | -131 | 0.20955 | 0.000457 | -3e-06 | 98.2674 | 98.0578 | 124.5228 | 21.0132 | 0.001899 | 0.001732 | -151.6343 | 0.23632 | 8.4018 | 0.003506 | 0.004894 | 0.004734 | 6e-05 | 0.001068 | 0.31421 | 1.3516 | 2.5479 | 99.61 | 20.85 | 37.477 | 67.2897 | 105.9827 | 63.2613 | 0.009282 | 1.0494 | 0.025005 | 23.8271 | 0.019102 | 0.14058 | -0.000193 | -0.014283 | -0.014283 | -0.40454
2019-03-21 11:43 | 1553182983.3546 | -130 | 0.20956 | 0.000458 | -3e-06 | 98.2358 | 98.0683 | 124.6081 | 21.0143 | 0.001909 | 0.001835 | -151.6349 | 0.23587 | 8.4023 | 0.003422 | 0.004976 | 0.004806 | 6.1e-05 | 0.001064 | 0.29159 | 1.3697 | 2.5479 | 99.61 | 20.85 | 37.478 | 67.523 | 103.9512 | 65.7417 | 0.009282 | 1.0418 | 0.024922 | 23.923 | 0.015264 | 0.14032 | -0.000226 | -0.016703 | -0.016703 | -0.34527
2019-03-21 11:43 | 1553182984.3547 | -129 | 0.20958 | 0.000464 | -3e-06 | 98.1604 | 97.9706 | 124.67 | 21.0135 | 0.001907 | 0.001828 | -151.6575 | 0.23647 | 8.403 | 0.003407 | 0.004898 | 0.004677 | 6.2e-05 | 0.001064 | 0.32709 | 1.359 | 2.548 | 99.61 | 20.84 | 37.478 | 67.4645 | 103.5866 | 65.5702 | 0.009276 | 1.0423 | 0.024935 | 23.9228 | 0.017318 | 0.14067 | -0.000342 | -0.025419 | -0.025419 | -0.22745
2019-03-21 11:43 | 1553182985.3547 | -128 | 0.20957 | 0.000461 | -3e-06 | 98.0296 | 97.9913 | 124.8646 | 21.0139 | 0.001918 | 0.001812 | -151.6622 | 0.23643 | 8.4033 | 0.003432 | 0.004495 | 0.004362 | 5.9e-05 | 0.001064 | 0.31437 | 1.2991 | 2.548 | 99.61 | 20.84 | 37.479 | 67.7255 | 104.2 | 65.181 | 0.009276 | 1.0435 | 0.024968 | 23.9278 | 0.003495 | 0.14065 | -0.000319 | -0.02375 | -0.02375 | -0.2434
2019-03-21 11:43 | 1553182986.3551 | -127 | 0.20958 | 0.000462 | -3e-06 | 97.9289 | 97.6923 | 124.7444 | 21.0144 | 0.001953 | 0.00182 | -151.6459 | 0.23712 | 8.3992 | 0.003411 | 0.004905 | 0.004617 | 6.4e-05 | 0.001069 | 0.3079 | 1.3053 | 2.548 | 99.61 | 20.84 | 37.48 | 68.5823 | 103.6872 | 65.3742 | 0.009277 | 1.0429 | 0.024949 | 23.9231 | 0.021639 | 0.14106 | -0.00037 | -0.02757 | -0.02757 | -0.21029
2019-03-21 11:43 | 1553182987.3548 | -126 | 0.20957 | 0.000457 | -3e-06 | 97.8592 | 97.6562 | 124.7541 | 21.0143 | 0.001914 | 0.001797 | -151.6308 | 0.23728 | 8.4005 | 0.003488 | 0.004897 | 0.004777 | 6.2e-05 | 0.001067 | 0.30237 | 1.3488 | 2.5478 | 99.61 | 20.84 | 37.481 | 67.6299 | 105.5547 | 64.82 | 0.009277 | 1.0446 | 0.024971 | 23.9044 | 0.018575 | 0.14116 | -0.000312 | -0.023174 | -0.023174 | -0.25035
2019-03-21 11:43 | 1553182988.3549 | -125 | 0.20954 | 0.000459 | -3e-06 | 97.7192 | 97.3956 | 124.6468 | 21.015 | 0.001936 | 0.00178 | -151.6486 | 0.23697 | 8.4021 | 0.003512 | 0.005402 | 0.005232 | 6.1e-05 | 0.001069 | 0.33179 | 1.3119 | 2.548 | 99.61 | 20.84 | 37.482 | 68.1627 | 106.1511 | 64.4257 | 0.009277 | 1.0458 | 0.024974 | 23.8802 | 0.029663 | 0.14098 | -0.000122 | -0.008956 | -0.008956 | -0.64696
2019-03-21 11:43 | 1553182989.3542 | -124 | 0.20957 | 0.000466 | -3e-06 | 97.5583 | 97.5152 | 124.3632 | 21.0142 | 0.001923 | 0.001837 | -151.6428 | 0.23644 | 8.4038 | 0.003517 | 0.005205 | 0.005061 | 6.2e-05 | 0.001071 | 0.29407 | 1.302 | 2.5478 | 99.61 | 20.84 | 37.484 | 67.8576 | 106.259 | 65.7799 | 0.009278 | 1.0416 | 0.024896 | 23.9008 | 0.003952 | 0.14066 | -0.000298 | -0.022102 | -0.022102 | -0.26156
2019-03-21 11:43 | 1553182990.3539 | -123 | 0.20956 | 0.000458 | -3e-06 | 97.4886 | 97.1487 | 124.2043 | 21.0149 | 0.002028 | 0.001845 | -151.6538 | 0.23694 | 8.4052 | 0.003511 | 0.004851 | 0.004708 | 6.4e-05 | 0.001069 | 0.34363 | 1.3182 | 2.548 | 99.61 | 20.84 | 37.486 | 70.3804 | 106.1211 | 65.9897 | 0.009278 | 1.041 | 0.024873 | 23.8929 | 0.031235 | 0.14095 | -0.000217 | -0.016065 | -0.016065 | -0.3606
2019-03-21 11:43 | 1553182991.3551 | -122 | 0.20959 | 0.000462 | -3e-06 | 97.2797 | 97.1756 | 124.171 | 21.0152 | 0.001901 | 0.001791 | -151.6406 | 0.23624 | 8.4032 | 0.003511 | 0.004716 | 0.004563 | 6.5e-05 | 0.001072 | 0.30237 | 1.3178 | 2.5479 | 99.61 | 20.84 | 37.489 | 67.3375 | 106.1206 | 64.6898 | 0.009279 | 1.045 | 0.024917 | 23.8439 | 0.009575 | 0.14054 | -0.000423 | -0.031521 | -0.031521 | -0.18325
2019-03-21 11:43 | 1553182992.3544 | -121 | 0.20957 | 0.000462 | -3e-06 | 97.1179 | 96.6327 | 124.0635 | 21.0147 | 0.001929 | 0.001779 | -151.639 | 0.23704 | 8.4027 | 0.003503 | 0.004786 | 0.004624 | 6.2e-05 | 0.001074 | 0.32709 | 1.3347 | 2.548 | 99.61 | 20.83 | 37.489 | 68.0127 | 105.9196 | 64.4011 | 0.009273 | 1.0459 | 0.024917 | 23.8234 | 0.044798 | 0.14101 | -0.000295 | -0.021886 | -0.021886 | -0.26482
2019-03-21 11:43 | 1553182993.3551 | -120 | 0.20956 | 0.00046 | -3e-06 | 96.9234 | 96.5854 | 124.156 | 21.0153 | 0.001934 | 0.001807 | -151.6458 | 0.23691 | 8.4034 | 0.003454 | 0.004563 | 0.004439 | 6.3e-05 | 0.001069 | 0.31421 | 1.3343 | 2.548 | 99.61 | 20.83 | 37.491 | 68.1301 | 104.7258 | 65.0645 | 0.009274 | 1.0438 | 0.024902 | 23.8556 | 0.031245 | 0.14094 | -0.000226 | -0.016727 | -0.016727 | -0.34629
2019-03-21 11:43 | 1553182994.3542 | -119 | 0.20956 | 0.000465 | -3e-06 | 96.7664 | 96.3075 | 124.1834 | 21.0149 | 0.001864 | 0.001815 | -151.6194 | 0.23716 | 8.4033 | 0.003516 | 0.003962 | 0.003812 | 6.2e-05 | 0.001072 | 0.33284 | 1.3157 | 2.548 | 99.61 | 20.84 | 37.495 | 66.4383 | 106.2273 | 65.2553 | 0.00928 | 1.0433 | 0.024897 | 23.865 | 0.042517 | 0.14109 | -0.00023 | -0.016962 | -0.016962 | -0.34186
2019-03-21 11:43 | 1553182995.3542 | -118 | 0.20954 | 0.000465 | -3e-06 | 96.6332 | 96.2036 | 124.1501 | 21.0161 | 0.001953 | 0.001864 | -151.6251 | 0.23721 | 8.4035 | 0.003573 | 0.005264 | 0.005084 | 6.6e-05 | 0.001071 | 0.31133 | 1.2953 | 2.5481 | 99.61 | 20.84 | 37.497 | 68.5832 | 107.6166 | 66.4387 | 0.009281 | 1.0396 | 0.024851 | 23.9035 | 0.039851 | 0.14111 | -0.000149 | -0.01088 | -0.01088 | -0.53309
2019-03-21 11:43 | 1553182996.3547 | -117 | 0.20957 | 0.00046 | -3e-06 | 96.4273 | 96.0486 | 124.1433 | 21.0152 | 0.001915 | 0.001751 | -151.6406 | 0.23715 | 8.4012 | 0.003508 | 0.005052 | 0.004914 | 6.4e-05 | 0.001066 | 0.32101 | 1.3234 | 2.548 | 99.61 | 20.84 | 37.499 | 67.6706 | 106.0365 | 63.7099 | 0.009281 | 1.048 | 0.02495 | 23.8066 | 0.035194 | 0.14108 | -0.000329 | -0.024427 | -0.024427 | -0.23738
2019-03-21 11:43 | 1553182997.3556 | -116 | 0.20957 | 0.000463 | -4e-06 | 96.2813 | 95.7911 | 124.1372 | 21.0158 | 0.001947 | 0.001806 | -151.6709 | 0.23649 | 8.4016 | 0.003477 | 0.004254 | 0.004129 | 6.4e-05 | 0.001066 | 0.2896 | 1.3319 | 2.5479 | 99.61 | 20.84 | 37.501 | 68.4349 | 105.2809 | 65.0375 | 0.009282 | 1.0439 | 0.024901 | 23.8529 | 0.045652 | 0.14069 | -0.000309 | -0.022909 | -0.022909 | -0.25239
2019-03-21 11:43 | 1553182998.3554 | -115 | 0.2096 | 0.000462 | -4e-06 | 96.1118 | 95.6801 | 124.1712 | 21.0156 | 0.001871 | 0.001843 | -151.6735 | 0.23748 | 8.4023 | 0.003518 | 0.004636 | 0.004324 | 6.4e-05 | 0.001065 | 0.32333 | 1.3552 | 2.5479 | 99.61 | 20.84 | 37.501 | 66.5975 | 106.2848 | 65.9233 | 0.009282 | 1.0412 | 0.024872 | 23.8874 | 0.040263 | 0.14128 | -0.000455 | -0.033771 | -0.033771 | -0.17194
2019-03-21 11:43 | 1553182999.3559 | -114 | 0.20954 | 0.000461 | -3e-06 | 95.9923 | 95.5824 | 124.2262 | 21.0156 | 0.001953 | 0.001807 | -151.6571 | 0.23667 | 8.4017 | 0.003452 | 0.004377 | 0.004293 | 6.2e-05 | 0.001057 | 0.31924 | 1.3578 | 2.548 | 99.61 | 20.84 | 37.502 | 68.589 | 104.6759 | 65.0734 | 0.009282 | 1.0438 | 0.024908 | 23.8627 | 0.03828 | 0.1408 | -0.000122 | -0.008905 | -0.008905 | -0.64986
2019-03-21 11:43 | 1553183000.3555 | -113 | 0.20953 | 0.000463 | -3e-06 | 95.8174 | 95.4127 | 124.1717 | 21.0164 | 0.001895 | 0.001761 | -151.6476 | 0.2368 | 8.4021 | 0.003437 | 0.004872 | 0.004653 | 6.4e-05 | 0.001058 | 0.27113 | 1.3346 | 2.5481 | 99.61 | 20.84 | 37.503 | 67.1739 | 104.3189 | 63.9553 | 0.009282 | 1.0473 | 0.024944 | 23.818 | 0.037854 | 0.14087 | -6.4e-05 | -0.004566 | -0.004566 | -1.2681
2019-03-21 11:43 | 1553183001.3545 | -112 | 0.2096 | 0.000457 | -3e-06 | 95.6723 | 95.1093 | 124.3069 | 21.0162 | 0.001897 | 0.001801 | -151.6589 | 0.23659 | 8.4031 | 0.003398 | 0.004527 | 0.004346 | 6.3e-05 | 0.001064 | 0.28639 | 1.3666 | 2.5481 | 99.61 | 20.84 | 37.506 | 67.2269 | 103.3778 | 64.9179 | 0.009283 | 1.0443 | 0.024922 | 23.865 | 0.052793 | 0.14075 | -0.000504 | -0.037483 | -0.037483 | -0.15433
2019-03-21 11:43 | 1553183002.3535 | -111 | 0.20957 | 0.000463 | -4e-06 | 95.54 | 95.0841 | 124.1919 | 21.0161 | 0.00189 | 0.001825 | -151.6628 | 0.23651 | 8.4034 | 0.003428 | 0.004215 | 0.004117 | 6e-05 | 0.001071 | 0.29983 | 1.2785 | 2.5479 | 99.61 | 20.84 | 37.507 | 67.0568 | 104.1127 | 65.493 | 0.009283 | 1.0425 | 0.02489 | 23.8742 | 0.042787 | 0.1407 | -0.000285 | -0.021087 | -0.021087 | -0.27423
2019-03-21 11:43 | 1553183003.3536 | -110 | 0.20955 | 0.000463 | -3e-06 | 95.2638 | 94.6454 | 123.891 | 21.0165 | 0.001949 | 0.001792 | -151.6607 | 0.23641 | 8.4022 | 0.003473 | 0.005075 | 0.004936 | 6e-05 | 0.00107 | 0.32084 | 1.3472 | 2.5479 | 99.61 | 20.84 | 37.509 | 68.4784 | 105.1849 | 64.7096 | 0.009284 | 1.0449 | 0.024888 | 23.8177 | 0.058254 | 0.14064 | -0.000184 | -0.013604 | -0.013604 | -0.42489
2019-03-21 11:43 | 1553183004.3541 | -109 | 0.20956 | 0.000463 | -3e-06 | 95.1178 | 94.5991 | 123.7546 | 21.0162 | 0.001935 | 0.001788 | -151.6664 | 0.23717 | 8.402 | 0.003482 | 0.005102 | 0.004947 | 6.3e-05 | 0.001071 | 0.32211 | 1.3552 | 2.548 | 99.61 | 20.84 | 37.512 | 68.1427 | 105.4033 | 64.6025 | 0.009284 | 1.0453 | 0.024878 | 23.8008 | 0.048915 | 0.14109 | -0.000218 | -0.016114 | -0.016114 | -0.35986
2019-03-21 11:43 | 1553183005.3537 | -108 | 0.20957 | 0.00046 | -3e-06 | 94.9757 | 94.5385 | 123.776 | 21.0168 | 0.001905 | 0.001764 | -151.6718 | 0.23642 | 8.4024 | 0.003444 | 0.004249 | 0.004068 | 6.1e-05 | 0.001071 | 0.30414 | 1.3025 | 2.548 | 99.61 | 20.84 | 37.515 | 67.4336 | 104.5055 | 64.0308 | 0.009285 | 1.0471 | 0.024902 | 23.7827 | 0.041274 | 0.14065 | -0.000324 | -0.024005 | -0.024005 | -0.2408
2019-03-21 11:43 | 1553183006.3537 | -107 | 0.20958 | 0.000463 | -3e-06 | 94.8184 | 94.3004 | 123.6849 | 21.017 | 0.00195 | 0.001747 | -151.6556 | 0.23701 | 8.403 | 0.003361 | 0.004625 | 0.004396 | 6.3e-05 | 0.001068 | 0.31708 | 1.3114 | 2.5478 | 99.61 | 20.83 | 37.517 | 68.5178 | 102.482 | 63.6258 | 0.00928 | 1.0483 | 0.024907 | 23.7596 | 0.049 | 0.141 | -0.00033 | -0.02453 | -0.02453 | -0.23624
2019-03-21 11:43 | 1553183007.3543 | -106 | 0.20958 | 0.000465 | -3e-06 | 94.7061 | 93.9203 | 123.1774 | 21.0172 | 0.001933 | 0.001823 | -151.6404 | 0.23713 | 8.4037 | 0.003354 | 0.00506 | 0.004899 | 6.4e-05 | 0.001072 | 0.29103 | 1.3197 | 2.5481 | 99.61 | 20.83 | 37.521 | 68.1031 | 102.3031 | 65.4458 | 0.009281 | 1.0427 | 0.024789 | 23.7748 | 0.074523 | 0.14107 | -0.000376 | -0.027806 | -0.027806 | -0.20852
2019-03-21 11:43 | 1553183008.354 | -105 | 0.20955 | 0.000462 | -3e-06 | 94.502 | 94.0453 | 123.3546 | 21.0177 | 0.001982 | 0.001772 | -151.6294 | 0.23736 | 8.4023 | 0.003439 | 0.004753 | 0.004637 | 6.1e-05 | 0.001073 | 0.32516 | 1.288 | 2.548 | 99.61 | 20.83 | 37.524 | 69.2853 | 104.3673 | 64.2245 | 0.009282 | 1.0465 | 0.024852 | 23.749 | 0.043331 | 0.14121 | -0.000156 | -0.011455 | -0.011455 | -0.50665
2019-03-21 11:43 | 1553183009.3544 | -104 | 0.20957 | 0.00046 | -3e-06 | 94.4218 | 93.6172 | 123.3018 | 21.0184 | 0.00193 | 0.001808 | -151.6463 | 0.23718 | 8.4018 | 0.003421 | 0.004674 | 0.004511 | 6.3e-05 | 0.001073 | 0.29551 | 1.3481 | 2.548 | 99.61 | 20.83 | 37.529 | 68.0319 | 103.9286 | 65.0839 | 0.009283 | 1.0438 | 0.024815 | 23.7741 | 0.076553 | 0.1411 | -0.00028 | -0.020672 | -0.020672 | -0.28053
2019-03-21 11:43 | 1553183010.3543 | -103 | 0.20958 | 0.000464 | -3e-06 | 94.1498 | 93.4632 | 123.1945 | 21.0185 | 0.001909 | 0.001758 | -151.646 | 0.23676 | 8.4032 | 0.003432 | 0.004471 | 0.004314 | 6.2e-05 | 0.001066 | 0.31094 | 1.3385 | 2.548 | 99.61 | 20.83 | 37.531 | 67.5318 | 104.2156 | 63.8878 | 0.009283 | 1.0475 | 0.024848 | 23.7217 | 0.065467 | 0.14085 | -0.000335 | -0.024768 | -0.024768 | -0.23372
2019-03-21 11:44 | 1553183011.3555 | -102 | 0.20956 | 0.000464 | -3e-06 | 94.1012 | 93.3345 | 123.08 | 21.0186 | 0.001915 | 0.001798 | -151.6364 | 0.23693 | 8.4006 | 0.003454 | 0.004608 | 0.004509 | 6.6e-05 | 0.001053 | 0.29302 | 1.3247 | 2.548 | 99.61 | 20.83 | 37.534 | 67.6661 | 104.7447 | 64.8656 | 0.009284 | 1.0445 | 0.024801 | 23.7451 | 0.073182 | 0.14095 | -0.00026 | -0.019232 | -0.019232 | -0.30121
2019-03-21 11:44 | 1553183012.3545 | -101 | 0.20954 | 0.000462 | -3e-06 | 93.9368 | 93.2449 | 123.0457 | 21.0194 | 0.001889 | 0.001798 | -151.6437 | 0.23702 | 8.4028 | 0.003426 | 0.005061 | 0.004895 | 6.3e-05 | 0.001046 | 0.31022 | 1.3439 | 2.5479 | 99.61 | 20.83 | 37.539 | 67.0562 | 104.0718 | 64.8648 | 0.009285 | 1.0445 | 0.024797 | 23.7417 | 0.066123 | 0.141 | -0.000128 | -0.009329 | -0.009329 | -0.6212
2019-03-21 11:44 | 1553183013.3542 | -100 | 0.20954 | 0.000465 | -3e-06 | 93.7419 | 93.0678 | 123.211 | 21.0187 | 0.001903 | 0.00177 | -151.6399 | 0.23682 | 8.4026 | 0.003387 | 0.004693 | 0.004536 | 6.2e-05 | 0.001047 | 0.31094 | 1.3779 | 2.548 | 99.61 | 20.83 | 37.542 | 67.3728 | 103.1106 | 64.1794 | 0.009286 | 1.0466 | 0.024839 | 23.7336 | 0.064555 | 0.14089 | -0.000123 | -0.008954 | -0.008954 | -0.64666
2019-03-21 11:44 | 1553183014.3556 | -99 | 0.20955 | 0.00046 | -3e-06 | 93.6093 | 92.7136 | 123.4828 | 21.0188 | 0.001958 | 0.001797 | -151.648 | 0.23779 | 8.4005 | 0.003397 | 0.004953 | 0.004753 | 6.4e-05 | 0.001051 | 0.31404 | 1.3467 | 2.5479 | 99.61 | 20.83 | 37.545 | 68.7069 | 103.3551 | 64.8322 | 0.009287 | 1.0446 | 0.024843 | 23.7827 | 0.086003 | 0.14146 | -0.000146 | -0.010713 | -0.010713 | -0.54273
2019-03-21 11:44 | 1553183015.3557 | -98 | 0.20956 | 0.000461 | -3e-06 | 93.424 | 92.5518 | 123.3713 | 21.0185 | 0.001942 | 0.001818 | -151.6402 | 0.23652 | 8.4016 | 0.003389 | 0.005006 | 0.004842 | 6.5e-05 | 0.001058 | 0.30093 | 1.3733 | 2.548 | 99.61 | 20.83 | 37.547 | 68.3212 | 103.1748 | 65.3243 | 0.009287 | 1.043 | 0.024813 | 23.7893 | 0.083899 | 0.1407 | -0.000236 | -0.017389 | -0.017389 | -0.33257
2019-03-21 11:44 | 1553183016.3533 | -97 | 0.20954 | 0.000462 | -4e-06 | 93.2655 | 92.4394 | 123.3915 | 21.0187 | 0.0019 | 0.001785 | -151.6495 | 0.23597 | 8.4016 | 0.003405 | 0.004775 | 0.004588 | 6.4e-05 | 0.00106 | 0.31166 | 1.3727 | 2.5482 | 99.61 | 20.83 | 37.547 | 67.3198 | 103.5445 | 64.5385 | 0.009287 | 1.0455 | 0.024844 | 23.7636 | 0.079579 | 0.14038 | -0.000113 | -0.008191 | -0.008191 | -0.70441
2019-03-21 11:44 | 1553183017.3533 | -96 | 0.20955 | 0.000464 | -3e-06 | 92.9881 | 92.2235 | 123.7109 | 21.019 | 0.001961 | 0.001812 | -151.6607 | 0.23733 | 8.4021 | 0.003367 | 0.005249 | 0.004988 | 6e-05 | 0.001063 | 0.30829 | 1.3387 | 2.5478 | 99.61 | 20.83 | 37.548 | 68.7664 | 102.638 | 65.2024 | 0.009288 | 1.0434 | 0.024852 | 23.8177 | 0.073856 | 0.14119 | -0.000159 | -0.011667 | -0.011667 | -0.49738
2019-03-21 11:44 | 1553183018.3535 | -95 | 0.20954 | 0.000462 | -3e-06 | 92.7772 | 91.9238 | 123.303 | 21.0189 | 0.00194 | 0.001829 | -151.6592 | 0.23673 | 8.4018 | 0.003314 | 0.004558 | 0.004393 | 6.4e-05 | 0.001066 | 0.30469 | 1.3559 | 2.548 | 99.61 | 20.83 | 37.548 | 68.2709 | 101.3492 | 65.5941 | 0.009288 | 1.0422 | 0.024797 | 23.7922 | 0.082658 | 0.14083 | -0.000125 | -0.009097 | -0.009097 | -0.6363
2019-03-21 11:44 | 1553183019.3539 | -94 | 0.20956 | 0.000464 | -3e-06 | 92.5305 | 91.5607 | 123.1183 | 21.0191 | 0.001951 | 0.001778 | -151.6536 | 0.23619 | 8.4023 | 0.003303 | 0.004355 | 0.004257 | 6.7e-05 | 0.001069 | 0.32244 | 1.346 | 2.5481 | 99.61 | 20.84 | 37.549 | 68.5302 | 101.0807 | 64.3687 | 0.009294 | 1.046 | 0.024823 | 23.7313 | 0.09424 | 0.14051 | -0.000246 | -0.018127 | -0.018127 | -0.31857
2019-03-21 11:44 | 1553183020.3546 | -93 | 0.20954 | 0.000461 | -4e-06 | 92.2871 | 91.3145 | 123.1845 | 21.0196 | 0.001943 | 0.001761 | -151.6733 | 0.23729 | 8.4028 | 0.003372 | 0.004265 | 0.004078 | 6.5e-05 | 0.00107 | 0.33892 | 1.3282 | 2.548 | 99.61 | 20.84 | 37.549 | 68.3417 | 102.7429 | 63.9584 | 0.009294 | 1.0473 | 0.024845 | 23.7232 | 0.094767 | 0.14116 | -0.000133 | -0.009676 | -0.009676 | -0.59961
2019-03-21 11:44 | 1553183021.3536 | -92 | 0.20954 | 0.000461 | -3e-06 | 92.0228 | 90.9698 | 123.2173 | 21.0202 | 0.001933 | 0.001803 | -151.6708 | 0.23758 | 8.4032 | 0.003326 | 0.005052 | 0.004952 | 6.5e-05 | 0.001071 | 0.29574 | 1.377 | 2.548 | 99.61 | 20.84 | 37.55 | 68.0966 | 101.6394 | 64.9654 | 0.009294 | 1.0442 | 0.024811 | 23.7618 | 0.10294 | 0.14134 | -9.7e-05 | -0.00704 | -0.00704 | -0.82518
2019-03-21 11:44 | 1553183022.354 | -91 | 0.20954 | 0.000462 | -3e-06 | 91.8264 | 90.8427 | 123.3131 | 21.0212 | 0.001911 | 0.001789 | -151.6574 | 0.23656 | 8.4029 | 0.003302 | 0.004555 | 0.004411 | 6.4e-05 | 0.001073 | 0.31332 | 1.3159 | 2.5479 | 99.61 | 20.84 | 37.549 | 67.5789 | 101.0513 | 64.6369 | 0.009294 | 1.0452 | 0.024833 | 23.7595 | 0.096337 | 0.14073 | -0.000116 | -0.008498 | -0.008498 | -0.68059
2019-03-21 11:44 | 1553183023.3551 | -90 | 0.20958 | 0.000465 | -3e-06 | 91.5628 | 90.3547 | 123.1809 | 21.0204 | 0.001918 | 0.001835 | -151.6633 | 0.23658 | 8.4009 | 0.003371 | 0.004878 | 0.00474 | 6.1e-05 | 0.001068 | 0.30685 | 1.3477 | 2.5479 | 99.61 | 20.84 | 37.549 | 67.7537 | 102.7227 | 65.7563 | 0.009294 | 1.0417 | 0.024778 | 23.7861 | 0.1188 | 0.14074 | -0.000357 | -0.026414 | -0.026414 | -0.21899
2019-03-21 11:44 | 1553183024.3548 | -89 | 0.20955 | 0.000467 | -3e-06 | 91.2881 | 90.3254 | 122.9195 | 21.0213 | 0.001903 | 0.001796 | -151.6408 | 0.23673 | 8.4006 | 0.003369 | 0.004436 | 0.004229 | 5.9e-05 | 0.001068 | 0.28943 | 1.3004 | 2.548 | 99.61 | 20.84 | 37.548 | 67.3771 | 102.6748 | 64.798 | 0.009293 | 1.0447 | 0.024787 | 23.7272 | 0.094825 | 0.14083 | -0.000164 | -0.011999 | -0.011999 | -0.48238
2019-03-21 11:44 | 1553183025.3548 | -88 | 0.20955 | 0.000461 | -3e-06 | 90.9836 | 89.8063 | 122.8242 | 21.0218 | 0.001933 | 0.001821 | -151.6355 | 0.23675 | 8.4016 | 0.003277 | 0.005341 | 0.005122 | 6e-05 | 0.001073 | 0.30702 | 1.3152 | 2.548 | 99.61 | 20.84 | 37.547 | 68.1075 | 100.4547 | 65.4127 | 0.009293 | 1.0428 | 0.024755 | 23.7396 | 0.11649 | 0.14084 | -0.000176 | -0.012895 | -0.012895 | -0.44888
2019-03-21 11:44 | 1553183026.3549 | -87 | 0.20957 | 0.000463 | -3e-06 | 90.831 | 89.6803 | 123.1429 | 21.0206 | 0.00193 | 0.001762 | -151.6525 | 0.23664 | 8.3987 | 0.003335 | 0.004724 | 0.004611 | 6.2e-05 | 0.001081 | 0.31797 | 1.3213 | 2.5479 | 99.61 | 20.83 | 37.545 | 68.0362 | 101.8518 | 64.0008 | 0.009287 | 1.0471 | 0.024839 | 23.7207 | 0.11404 | 0.14078 | -0.000279 | -0.02061 | -0.02061 | -0.28074
2019-03-21 11:44 | 1553183027.3548 | -86 | 0.20955 | 0.000462 | -3e-06 | 90.6028 | 89.5322 | 123.3269 | 21.0214 | 0.001887 | 0.001779 | -151.6282 | 0.2366 | 8.4028 | 0.003364 | 0.004509 | 0.004375 | 6.1e-05 | 0.001079 | 0.2975 | 1.3136 | 2.5479 | 99.61 | 20.83 | 37.544 | 67.0057 | 102.5578 | 64.3951 | 0.009287 | 1.0459 | 0.024843 | 23.7523 | 0.10633 | 0.14075 | -0.000173 | -0.012742 | -0.012742 | -0.45399
2019-03-21 11:44 | 1553183028.354 | -85 | 0.20954 | 0.00046 | -3e-06 | 90.4741 | 89.2519 | 123.4746 | 21.0223 | 0.001941 | 0.001827 | -151.6285 | 0.23609 | 8.3997 | 0.003355 | 0.0046 | 0.004463 | 5.8e-05 | 0.001076 | 0.31166 | 1.3593 | 2.5482 | 99.61 | 20.84 | 37.544 | 68.3092 | 102.3339 | 65.5551 | 0.009292 | 1.0423 | 0.024815 | 23.8073 | 0.12166 | 0.14045 | -0.00012 | -0.008746 | -0.008746 | -0.66004
2019-03-21 11:44 | 1553183029.3554 | -84 | 0.20955 | 0.000459 | -3e-06 | 90.2122 | 88.9 | 123.5307 | 21.0214 | 0.00191 | 0.001782 | -151.6343 | 0.23637 | 8.402 | 0.00339 | 0.004584 | 0.004411 | 6.1e-05 | 0.001072 | 0.2672 | 1.3945 | 2.5481 | 99.61 | 20.84 | 37.539 | 67.5534 | 103.2004 | 64.4663 | 0.009291 | 1.0457 | 0.024861 | 23.7744 | 0.13106 | 0.14062 | -0.000192 | -0.014113 | -0.014113 | -0.4095
2019-03-21 11:44 | 1553183030.3545 | -83 | 0.20957 | 0.000459 | -3e-06 | 90.0072 | 88.8135 | 123.5664 | 21.0216 | 0.001942 | 0.001776 | -151.6188 | 0.23773 | 8.4009 | 0.003309 | 0.005331 | 0.0051 | 6.2e-05 | 0.001068 | 0.29927 | 1.403 | 2.548 | 99.61 | 20.84 | 37.537 | 68.3227 | 101.2284 | 64.3385 | 0.009291 | 1.0461 | 0.024869 | 23.7734 | 0.11942 | 0.14142 | -0.000303 | -0.022425 | -0.022425 | -0.25919
2019-03-21 11:44 | 1553183031.3532 | -82 | 0.20949 | 0.000458 | -3e-06 | 89.9363 | 88.665 | 123.475 | 21.0219 | 0.00194 | 0.001822 | -151.6249 | 0.23691 | 8.4037 | 0.003386 | 0.004341 | 0.004285 | 6.1e-05 | 0.001065 | 0.28694 | 1.3469 | 2.5479 | 99.61 | 20.83 | 37.533 | 68.2754 | 103.0902 | 65.433 | 0.009284 | 1.0427 | 0.02482 | 23.8031 | 0.12734 | 0.14094 | 0.000171 | 0.012948 | 0.012948 | 0.44736
2019-03-21 11:44 | 1553183032.3532 | -81 | 0.2095 | 0.000464 | -3e-06 | 89.8746 | 88.5758 | 123.4868 | 21.0217 | 0.001913 | 0.001827 | -151.6278 | 0.23753 | 8.4026 | 0.003327 | 0.004815 | 0.004523 | 5.9e-05 | 0.001067 | 0.30486 | 1.2952 | 2.5479 | 99.61 | 20.83 | 37.531 | 67.6253 | 101.6556 | 65.5492 | 0.009283 | 1.0424 | 0.024817 | 23.8083 | 0.13021 | 0.1413 | 0.000117 | 0.008949 | 0.008949 | 0.64894
2019-03-21 11:44 | 1553183033.3536 | -80 | 0.20956 | 0.000462 | -3e-06 | 89.8469 | 88.5932 | 123.7297 | 21.0221 | 0.001902 | 0.00179 | -151.6202 | 0.2361 | 8.4046 | 0.003357 | 0.004474 | 0.004294 | 6e-05 | 0.001069 | 0.28711 | 1.2988 | 2.548 | 99.61 | 20.83 | 37.529 | 67.3692 | 102.3911 | 64.6597 | 0.009283 | 1.0451 | 0.024874 | 23.8004 | 0.12569 | 0.14046 | -0.000212 | -0.015642 | -0.015642 | -0.36905
2019-03-21 11:44 | 1553183034.3533 | -79 | 0.20955 | 0.000464 | -4e-06 | 89.9345 | 88.5346 | 123.7199 | 21.022 | 0.001952 | 0.001773 | -151.6355 | 0.236 | 8.4035 | 0.003303 | 0.004693 | 0.004626 | 6e-05 | 0.001068 | 0.3011 | 1.2765 | 2.548 | 99.61 | 20.84 | 37.526 | 68.5664 | 101.0916 | 64.2489 | 0.009288 | 1.0464 | 0.024888 | 23.785 | 0.14033 | 0.14039 | -0.000182 | -0.013378 | -0.013378 | -0.43133
2019-03-21 11:44 | 1553183035.3544 | -78 | 0.20957 | 0.00046 | -4e-06 | 89.9843 | 88.6606 | 123.6516 | 21.0226 | 0.001922 | 0.001778 | -151.6376 | 0.23761 | 8.4022 | 0.003328 | 0.00422 | 0.003961 | 6e-05 | 0.001069 | 0.31022 | 1.2848 | 2.5479 | 99.61 | 20.84 | 37.525 | 67.8501 | 101.7005 | 64.369 | 0.009288 | 1.046 | 0.024877 | 23.7826 | 0.13255 | 0.14136 | -0.000258 | -0.019069 | -0.019069 | -0.30467
2019-03-21 11:44 | 1553183036.3544 | -77 | 0.20955 | 0.000461 | -3e-06 | 90.0799 | 88.7801 | 123.3905 | 21.0231 | 0.00199 | 0.001756 | -151.6157 | 0.23658 | 8.4046 | 0.003266 | 0.004412 | 0.004158 | 6.2e-05 | 0.001065 | 0.31885 | 1.3225 | 2.548 | 99.61 | 20.84 | 37.522 | 69.4792 | 100.193 | 63.8386 | 0.009287 | 1.0476 | 0.02487 | 23.7389 | 0.13 | 0.14074 | -0.000137 | -0.010014 | -0.010014 | -0.57764
2019-03-21 11:44 | 1553183037.3535 | -76 | 0.20955 | 0.00046 | -3e-06 | 90.2307 | 88.9232 | 123.3305 | 21.0229 | 0.00197 | 0.001761 | -151.6342 | 0.23737 | 8.4022 | 0.003422 | 0.005061 | 0.004898 | 6.3e-05 | 0.001062 | 0.30381 | 1.2824 | 2.548 | 99.61 | 20.84 | 37.521 | 68.987 | 103.9785 | 63.9758 | 0.009287 | 1.0472 | 0.024859 | 23.7379 | 0.13056 | 0.14121 | -0.000177 | -0.012975 | -0.012975 | -0.4473
2019-03-21 11:44 | 1553183038.3549 | -75 | 0.20954 | 0.000462 | -3e-06 | 90.3737 | 89.1316 | 123.1288 | 21.0229 | 0.001934 | 0.001826 | -151.6311 | 0.23622 | 8.4001 | 0.003376 | 0.005418 | 0.005219 | 6.1e-05 | 0.001063 | 0.32997 | 1.294 | 2.548 | 99.61 | 20.84 | 37.521 | 68.1259 | 102.8437 | 65.5222 | 0.009287 | 1.0424 | 0.024782 | 23.7728 | 0.12378 | 0.14053 | -0.000125 | -0.009123 | -0.009123 | -0.63312
2019-03-21 11:44 | 1553183039.3546 | -74 | 0.20954 | 0.000465 | -3e-06 | 90.5018 | 89.2559 | 123.2185 | 21.0242 | 0.001911 | 0.001788 | -151.6396 | 0.23621 | 8.4021 | 0.003378 | 0.004286 | 0.004164 | 6e-05 | 0.001067 | 0.31924 | 1.3206 | 2.5479 | 99.61 | 20.84 | 37.518 | 67.5771 | 102.8937 | 64.6252 | 0.009286 | 1.0452 | 0.024824 | 23.75 | 0.12399 | 0.14052 | -9.5e-05 | -0.006867 | -0.006867 | -0.84103
2019-03-21 11:44 | 1553183040.3546 | -73 | 0.20954 | 0.000462 | -3e-06 | 90.6603 | 89.4041 | 123.103 | 21.0239 | 0.001963 | 0.001813 | -151.6387 | 0.23683 | 8.4017 | 0.00337 | 0.004833 | 0.004617 | 6.1e-05 | 0.001074 | 0.31725 | 1.2961 | 2.5479 | 99.61 | 20.84 | 37.519 | 68.8226 | 102.7103 | 65.2196 | 0.009286 | 1.0434 | 0.02479 | 23.7597 | 0.1248 | 0.14089 | -9.4e-05 | -0.006771 | -0.006771 | -0.85516
2019-03-21 11:44 | 1553183041.3547 | -72 | 0.20957 | 0.000463 | -3e-06 | 90.8346 | 89.4889 | 123.1145 | 21.0239 | 0.001919 | 0.001824 | -151.6349 | 0.23622 | 8.4022 | 0.003475 | 0.004606 | 0.004418 | 6.2e-05 | 0.001068 | 0.30773 | 1.3092 | 2.5479 | 99.61 | 20.84 | 37.52 | 67.7711 | 105.2409 | 65.4951 | 0.009286 | 1.0425 | 0.024781 | 23.7705 | 0.1335 | 0.14053 | -0.000274 | -0.0202 | -0.0202 | -0.28593
2019-03-21 11:44 | 1553183042.3541 | -71 | 0.20955 | 0.000459 | -4e-06 | 91.0283 | 89.7787 | 123.0347 | 21.0246 | 0.001922 | 0.001806 | -151.6261 | 0.2361 | 8.401 | 0.003411 | 0.004375 | 0.004236 | 6.5e-05 | 0.001066 | 0.32068 | 1.3221 | 2.5478 | 99.6 | 20.84 | 37.52 | 67.8506 | 103.7043 | 65.0625 | 0.009287 | 1.0439 | 0.024789 | 23.7476 | 0.12363 | 0.14046 | -0.000171 | -0.01251 | -0.01251 | -0.46146
2019-03-21 11:44 | 1553183043.3537 | -70 | 0.20955 | 0.000462 | -3e-06 | 91.1558 | 89.9211 | 122.9899 | 21.0239 | 0.001876 | 0.001785 | -151.6237 | 0.2369 | 8.4017 | 0.003457 | 0.00441 | 0.004207 | 6.4e-05 | 0.001066 | 0.3298 | 1.3171 | 2.5478 | 99.6 | 20.84 | 37.521 | 66.7477 | 104.8256 | 64.5436 | 0.009288 | 1.0455 | 0.024804 | 23.7251 | 0.12198 | 0.14093 | -0.000165 | -0.012066 | -0.012066 | -0.48006
2019-03-21 11:44 | 1553183044.3539 | -69 | 0.20959 | 0.000465 | -3e-06 | 91.3894 | 90.34 | 123.0547 | 21.0252 | 0.001983 | 0.001815 | -151.62 | 0.23676 | 8.4029 | 0.003434 | 0.004581 | 0.004452 | 6.5e-05 | 0.001069 | 0.32101 | 1.3775 | 2.548 | 99.6 | 20.84 | 37.523 | 69.3197 | 104.2575 | 65.2755 | 0.009288 | 1.0432 | 0.024783 | 23.757 | 0.10329 | 0.14085 | -0.000393 | -0.02904 | -0.02904 | -0.19934
2019-03-21 11:44 | 1553183045.3553 | -68 | 0.20955 | 0.000459 | -3e-06 | 91.6086 | 90.4391 | 123.2591 | 21.0247 | 0.0019 | 0.00181 | -151.6193 | 0.23622 | 8.4008 | 0.003498 | 0.004544 | 0.004312 | 6.7e-05 | 0.001072 | 0.30021 | 1.3448 | 2.548 | 99.6 | 20.84 | 37.523 | 67.3147 | 105.8064 | 65.1453 | 0.009288 | 1.0436 | 0.024809 | 23.7722 | 0.11492 | 0.14053 | -0.000178 | -0.013081 | -0.013081 | -0.44155
2019-03-21 11:44 | 1553183046.3542 | -67 | 0.20956 | 0.000462 | -3e-06 | 91.7645 | 90.6329 | 123.6131 | 21.0252 | 0.00194 | 0.001775 | -151.645 | 0.23723 | 8.4008 | 0.003504 | 0.005318 | 0.005176 | 6.4e-05 | 0.001071 | 0.33411 | 1.3312 | 2.548 | 99.6 | 20.84 | 37.524 | 68.2669 | 105.9643 | 64.2982 | 0.009288 | 1.0462 | 0.024875 | 23.7765 | 0.11098 | 0.14113 | -0.00021 | -0.015522 | -0.015522 | -0.37369
2019-03-21 11:44 | 1553183047.354 | -66 | 0.20955 | 0.000463 | -4e-06 | 91.9206 | 91.0583 | 123.7093 | 21.0248 | 0.001867 | 0.001841 | -151.6421 | 0.23663 | 8.4012 | 0.003553 | 0.005136 | 0.004993 | 6.6e-05 | 0.00107 | 0.31509 | 1.3437 | 2.548 | 99.6 | 20.84 | 37.525 | 66.5293 | 107.1353 | 65.9018 | 0.009289 | 1.0413 | 0.024826 | 23.8422 | 0.084301 | 0.14077 | -0.000191 | -0.014049 | -0.014049 | -0.41182
2019-03-21 11:44 | 1553183048.3534 | -65 | 0.20957 | 0.00046 | -3e-06 | 92.1306 | 91.1084 | 123.7912 | 21.0251 | 0.001977 | 0.001798 | -151.6512 | 0.23652 | 8.4025 | 0.003465 | 0.004554 | 0.004254 | 6.6e-05 | 0.001074 | 0.32068 | 1.3166 | 2.548 | 99.6 | 20.84 | 37.527 | 69.1553 | 105.003 | 64.8515 | 0.009289 | 1.0445 | 0.024873 | 23.8131 | 0.0998 | 0.1407 | -0.000274 | -0.020268 | -0.020268 | -0.28532
2019-03-21 11:44 | 1553183049.3541 | -64 | 0.20952 | 0.000461 | -4e-06 | 92.3928 | 91.3972 | 124.0187 | 21.0249 | 0.001915 | 0.001817 | -151.6503 | 0.23708 | 8.4002 | 0.00354 | 0.004418 | 0.004258 | 6.2e-05 | 0.001071 | 0.33373 | 1.2863 | 2.548 | 99.6 | 20.84 | 37.527 | 67.6698 | 106.833 | 65.318 | 0.009289 | 1.0431 | 0.024879 | 23.8514 | 0.096909 | 0.14104 | 2.6e-05 | 0.002145 | 0.002145 | 2.7018
2019-03-21 11:44 | 1553183050.3542 | -63 | 0.20954 | 0.00046 | -3e-06 | 92.6127 | 91.7984 | 124.3616 | 21.0253 | 0.001897 | 0.001823 | -151.6484 | 0.23742 | 8.4023 | 0.003552 | 0.004953 | 0.004753 | 6.3e-05 | 0.001069 | 0.30381 | 1.3437 | 2.548 | 99.6 | 20.84 | 37.528 | 67.2334 | 107.1288 | 65.4544 | 0.009289 | 1.0426 | 0.024908 | 23.8892 | 0.078992 | 0.14124 | -0.000126 | -0.009219 | -0.009219 | -0.6297
2019-03-21 11:44 | 1553183051.3533 | -62 | 0.20954 | 0.000459 | -3e-06 | 92.9229 | 91.9984 | 124.4785 | 21.0259 | 0.001849 | 0.001814 | -151.6552 | 0.23574 | 8.4007 | 0.003447 | 0.004925 | 0.004716 | 6.6e-05 | 0.00107 | 0.32621 | 1.3166 | 2.5479 | 99.6 | 20.84 | 37.529 | 66.1003 | 104.5655 | 65.2362 | 0.00929 | 1.0433 | 0.024928 | 23.8927 | 0.089438 | 0.14024 | -0.000111 | -0.008133 | -0.008133 | -0.70869
2019-03-21 11:44 | 1553183052.3543 | -61 | 0.20956 | 0.000458 | -3e-06 | 93.0856 | 92.1918 | 124.6339 | 21.0251 | 0.001999 | 0.001822 | -151.6649 | 0.23666 | 8.4017 | 0.003421 | 0.004718 | 0.00455 | 6.4e-05 | 0.00107 | 0.29734 | 1.3345 | 2.5479 | 99.6 | 20.84 | 37.528 | 69.6963 | 103.9548 | 65.4486 | 0.009289 | 1.0427 | 0.024935 | 23.9151 | 0.086295 | 0.14079 | -0.000201 | -0.014852 | -0.014852 | -0.38959
2019-03-21 11:44 | 1553183053.3542 | -60 | 0.20957 | 0.000464 | -3e-06 | 93.3427 | 92.3158 | 124.7015 | 21.0259 | 0.001942 | 0.001822 | -151.6664 | 0.23544 | 8.4031 | 0.00353 | 0.004851 | 0.004721 | 6.5e-05 | 0.001068 | 0.30469 | 1.3088 | 2.5479 | 99.6 | 20.84 | 37.529 | 68.3169 | 106.5909 | 65.4347 | 0.00929 | 1.0427 | 0.024943 | 23.9211 | 0.098948 | 0.14007 | -0.000278 | -0.020613 | -0.020613 | -0.27927
2019-03-21 11:44 | 1553183054.3534 | -59 | 0.2095 | 0.000462 | -4e-06 | 93.4651 | 92.625 | 124.7491 | 21.0253 | 0.001865 | 0.001799 | -151.674 | 0.23632 | 8.3998 | 0.003564 | 0.004894 | 0.004678 | 6.3e-05 | 0.001069 | 0.28744 | 1.3169 | 2.5479 | 99.6 | 20.84 | 37.531 | 66.472 | 107.4106 | 64.8844 | 0.00929 | 1.0444 | 0.024968 | 23.9062 | 0.08076 | 0.14058 | 0.000151 | 0.011523 | 0.011523 | 0.50145
2019-03-21 11:44 | 1553183055.3538 | -58 | 0.20957 | 0.000461 | -3e-06 | 93.6877 | 92.5869 | 124.6803 | 21.0259 | 0.001946 | 0.001835 | -151.6746 | 0.23695 | 8.4 | 0.003557 | 0.004755 | 0.004614 | 6.4e-05 | 0.001073 | 0.29607 | 1.3471 | 2.548 | 99.6 | 20.84 | 37.534 | 68.411 | 107.2528 | 65.7609 | 0.009291 | 1.0417 | 0.024929 | 23.9306 | 0.10572 | 0.14096 | -0.000286 | -0.021166 | -0.021166 | -0.27372
2019-03-21 11:44 | 1553183056.3535 | -57 | 0.20952 | 0.000462 | -3e-06 | 93.8 | 92.9847 | 124.4844 | 21.0267 | 0.001961 | 0.001765 | -151.6665 | 0.23614 | 8.4023 | 0.003507 | 0.005846 | 0.00569 | 6.2e-05 | 0.00107 | 0.33931 | 1.3452 | 2.5481 | 99.6 | 20.84 | 37.536 | 68.7753 | 106.0313 | 64.0677 | 0.009291 | 1.0469 | 0.024971 | 23.8519 | 0.078088 | 0.14048 | -5e-06 | -0.000181 | -0.000181 | -31.835
2019-03-21 11:44 | 1553183057.3536 | -56 | 0.20955 | 0.000458 | -3e-06 | 93.9016 | 93.2412 | 124.5226 | 21.0265 | 0.001863 | 0.001793 | -151.6712 | 0.23705 | 8.4021 | 0.003526 | 0.005296 | 0.005103 | 6.4e-05 | 0.00107 | 0.3194 | 1.2984 | 2.548 | 99.6 | 20.84 | 37.543 | 66.4371 | 106.481 | 64.7318 | 0.009293 | 1.0449 | 0.024951 | 23.8791 | 0.063135 | 0.14102 | -0.000186 | -0.013718 | -0.013718 | -0.4225
2019-03-21 11:44 | 1553183058.3536 | -55 | 0.20953 | 0.00046 | -4e-06 | 94.0214 | 93.2779 | 124.7618 | 21.0264 | 0.001959 | 0.001808 | -151.677 | 0.23687 | 8.4037 | 0.003599 | 0.004533 | 0.004475 | 6.5e-05 | 0.001075 | 0.30862 | 1.3589 | 2.548 | 99.6 | 20.84 | 37.548 | 68.7438 | 108.2719 | 65.1073 | 0.009294 | 1.0437 | 0.024961 | 23.9153 | 0.071008 | 0.14091 | -1.9e-05 | -0.001177 | -0.001177 | -4.9202
2019-03-21 11:44 | 1553183059.3551 | -54 | 0.20955 | 0.00046 | -3e-06 | 94.1593 | 93.4857 | 124.3351 | 21.0265 | 0.001896 | 0.001843 | -151.6659 | 0.23706 | 8.4031 | 0.003579 | 0.004747 | 0.00455 | 6.3e-05 | 0.001075 | 0.29856 | 1.3245 | 2.5478 | 99.6 | 20.84 | 37.552 | 67.2158 | 107.7813 | 65.9535 | 0.009295 | 1.0411 | 0.024887 | 23.9042 | 0.064222 | 0.14103 | -0.000174 | -0.012791 | -0.012791 | -0.45316
2019-03-21 11:44 | 1553183060.3537 | -53 | 0.20955 | 0.000459 | -3e-06 | 94.2207 | 93.6216 | 123.9945 | 21.0262 | 0.001921 | 0.001754 | -151.6728 | 0.23645 | 8.4029 | 0.003556 | 0.005447 | 0.005274 | 6.6e-05 | 0.001076 | 0.30021 | 1.3454 | 2.5479 | 99.6 | 20.84 | 37.56 | 67.8233 | 107.2131 | 63.794 | 0.009297 | 1.0478 | 0.024932 | 23.7953 | 0.057058 | 0.14067 | -0.000161 | -0.011841 | -0.011841 | -0.48823
2019-03-21 11:44 | 1553183061.3537 | -52 | 0.20955 | 0.000459 | -3e-06 | 94.3891 | 93.7411 | 123.7332 | 21.0282 | 0.001926 | 0.001809 | -151.6553 | 0.23727 | 8.4013 | 0.00358 | 0.004737 | 0.004459 | 6.6e-05 | 0.001078 | 0.34164 | 1.3515 | 2.548 | 99.6 | 20.83 | 37.564 | 67.9337 | 107.805 | 65.1331 | 0.009293 | 1.0436 | 0.024857 | 23.8174 | 0.06161 | 0.14115 | -0.000174 | -0.012787 | -0.012787 | -0.45369
2019-03-21 11:44 | 1553183062.3554 | -51 | 0.20954 | 0.000461 | -3e-06 | 94.4335 | 93.9036 | 123.8783 | 21.0275 | 0.001931 | 0.001725 | -151.657 | 0.23642 | 8.4013 | 0.003616 | 0.005291 | 0.005148 | 6.6e-05 | 0.001072 | 0.32244 | 1.3002 | 2.5479 | 99.6 | 20.84 | 37.572 | 68.057 | 108.6858 | 63.1088 | 0.0093 | 1.0499 | 0.024946 | 23.76 | 0.050328 | 0.14065 | -0.000127 | -0.0093 | -0.0093 | -0.62157
2019-03-21 11:44 | 1553183063.3529 | -50 | 0.20951 | 0.000465 | -3e-06 | 94.58 | 94.0885 | 124.0344 | 21.0272 | 0.001917 | 0.001768 | -151.6657 | 0.23674 | 8.4024 | 0.003603 | 0.005251 | 0.005097 | 6.7e-05 | 0.001068 | 0.30773 | 1.2718 | 2.5479 | 99.6 | 20.84 | 37.576 | 67.7244 | 108.3649 | 64.1321 | 0.009301 | 1.0467 | 0.024924 | 23.8111 | 0.046603 | 0.14084 | 6.2e-05 | 0.00482 | 0.00482 | 1.2008
2019-03-21 11:44 | 1553183064.353 | -49 | 0.20952 | 0.000461 | -3e-06 | 94.6614 | 94.0146 | 124.0799 | 21.0279 | 0.00193 | 0.001829 | -151.6617 | 0.23685 | 8.3994 | 0.003605 | 0.004694 | 0.004581 | 6.7e-05 | 0.001072 | 0.33373 | 1.2799 | 2.5479 | 99.6 | 20.84 | 37.579 | 68.0355 | 108.4081 | 65.6206 | 0.009302 | 1.0421 | 0.024874 | 23.8679 | 0.061321 | 0.1409 | 2.3e-05 | 0.001925 | 0.001925 | 3.0091
2019-03-21 11:44 | 1553183065.3544 | -48 | 0.20957 | 0.000462 | -3e-06 | 94.7978 | 94.2115 | 123.8889 | 21.027 | 0.001886 | 0.001846 | -151.6809 | 0.23605 | 8.404 | 0.00357 | 0.004941 | 0.004787 | 6.8e-05 | 0.001069 | 0.34252 | 1.2584 | 2.5478 | 99.6 | 20.83 | 37.583 | 66.9784 | 107.5651 | 66.0123 | 0.009297 | 1.0409 | 0.02484 | 23.8633 | 0.055496 | 0.14043 | -0.00027 | -0.019959 | -0.019959 | -0.28917
2019-03-21 11:44 | 1553183066.3544 | -47 | 0.20955 | 0.00046 | -3e-06 | 94.9775 | 94.3692 | 123.7464 | 21.0286 | 0.001938 | 0.001812 | -151.6702 | 0.23694 | 8.4008 | 0.003645 | 0.004932 | 0.004714 | 6.7e-05 | 0.001073 | 0.29247 | 1.3082 | 2.5479 | 99.6 | 20.83 | 37.585 | 68.2235 | 109.3839 | 65.196 | 0.009298 | 1.0434 | 0.024856 | 23.8209 | 0.057472 | 0.14095 | -0.000168 | -0.01237 | -0.01237 | -0.46834
2019-03-21 11:44 | 1553183067.3531 | -46 | 0.20956 | 0.000462 | -3e-06 | 94.9674 | 94.281 | 123.5391 | 21.0283 | 0.001898 | 0.001843 | -151.6667 | 0.23752 | 8.4011 | 0.00357 | 0.005056 | 0.004847 | 6.9e-05 | 0.001072 | 0.31813 | 1.302 | 2.548 | 99.6 | 20.84 | 37.585 | 67.2644 | 107.5696 | 65.952 | 0.009303 | 1.0411 | 0.024807 | 23.8275 | 0.064883 | 0.1413 | -0.000236 | -0.017416 | -0.017416 | -0.33345
2019-03-21 11:44 | 1553183068.3532 | -45 | 0.20954 | 0.000459 | -3e-06 | 95.164 | 94.4911 | 123.5189 | 21.0285 | 0.001944 | 0.00185 | -151.6719 | 0.237 | 8.4007 | 0.003594 | 0.00495 | 0.004818 | 7e-05 | 0.001077 | 0.34307 | 1.2856 | 2.5479 | 99.6 | 20.84 | 37.582 | 68.3771 | 108.1472 | 66.1262 | 0.009303 | 1.0406 | 0.024799 | 23.8317 | 0.063471 | 0.14099 | -9.2e-05 | -0.006638 | -0.006638 | -0.87297
2019-03-21 11:44 | 1553183069.3536 | -44 | 0.20954 | 0.000461 | -3e-06 | 95.2179 | 94.66 | 123.8238 | 21.0289 | 0.001976 | 0.001806 | -151.6884 | 0.23675 | 8.4034 | 0.003631 | 0.004908 | 0.004711 | 6.9e-05 | 0.001075 | 0.31454 | 1.3373 | 2.548 | 99.6 | 20.84 | 37.579 | 69.134 | 109.0358 | 65.0653 | 0.009302 | 1.0438 | 0.024868 | 23.8237 | 0.052561 | 0.14084 | -7.6e-05 | -0.005483 | -0.005483 | -1.0557
2019-03-21 11:44 | 1553183070.3537 | -43 | 0.20956 | 0.000462 | -4e-06 | 95.3124 | 94.7224 | 123.9576 | 21.0282 | 0.001958 | 0.001837 | -151.6807 | 0.23637 | 8.4036 | 0.003579 | 0.004839 | 0.004699 | 6.8e-05 | 0.001073 | 0.29662 | 1.3471 | 2.548 | 99.6 | 20.84 | 37.572 | 68.7027 | 107.7742 | 65.7944 | 0.0093 | 1.0416 | 0.024855 | 23.8623 | 0.055542 | 0.14062 | -0.000226 | -0.016655 | -0.016655 | -0.347
2019-03-21 11:44 | 1553183071.3543 | -42 | 0.20954 | 0.000464 | -4e-06 | 95.4387 | 94.7741 | 124.0891 | 21.0284 | 0.00196 | 0.001815 | -151.7071 | 0.23645 | 8.4048 | 0.003681 | 0.004866 | 0.00473 | 6.8e-05 | 0.001075 | 0.31813 | 1.3473 | 2.548 | 99.6 | 20.84 | 37.569 | 68.7606 | 110.2614 | 65.2825 | 0.0093 | 1.0432 | 0.024887 | 23.8569 | 0.062502 | 0.14067 | -8.8e-05 | -0.006356 | -0.006356 | -0.90962
2019-03-21 11:45 | 1553183072.3547 | -41 | 0.20955 | 0.000463 | -4e-06 | 95.4685 | 94.8801 | 123.7555 | 21.0289 | 0.001988 | 0.001779 | -151.7228 | 0.23638 | 8.4026 | 0.003651 | 0.004585 | 0.004297 | 7e-05 | 0.001074 | 0.33754 | 1.3146 | 2.548 | 99.6 | 20.84 | 37.566 | 69.4248 | 109.5262 | 64.4017 | 0.009299 | 1.0459 | 0.024886 | 23.7938 | 0.055303 | 0.14062 | -0.000168 | -0.012287 | -0.012287 | -0.47037
2019-03-21 11:45 | 1553183073.3544 | -40 | 0.20956 | 0.000462 | -3e-06 | 95.4959 | 95.0371 | 123.4333 | 21.0284 | 0.001934 | 0.001817 | -151.7155 | 0.23627 | 8.4042 | 0.003606 | 0.004693 | 0.004526 | 6.7e-05 | 0.001077 | 0.30973 | 1.3267 | 2.5479 | 99.6 | 20.84 | 37.557 | 68.1327 | 108.4237 | 65.3238 | 0.009297 | 1.0431 | 0.02482 | 23.7952 | 0.043081 | 0.14056 | -0.000219 | -0.016132 | -0.016132 | -0.35811
2019-03-21 11:45 | 1553183074.3535 | -39 | 0.20958 | 0.00046 | -3e-06 | 95.5832 | 95.0878 | 123.2859 | 21.0292 | 0.001899 | 0.001802 | -151.7184 | 0.23708 | 8.4017 | 0.003589 | 0.005103 | 0.004985 | 6.8e-05 | 0.001075 | 0.3115 | 1.3169 | 2.5479 | 99.6 | 20.84 | 37.554 | 67.2871 | 108.0233 | 64.9539 | 0.009296 | 1.0442 | 0.024818 | 23.768 | 0.046479 | 0.14104 | -0.000331 | -0.024476 | -0.024476 | -0.23683
2019-03-21 11:45 | 1553183075.3535 | -38 | 0.20955 | 0.000464 | -4e-06 | 95.6035 | 94.9179 | 123.4708 | 21.0303 | 0.001958 | 0.001823 | -151.7085 | 0.23633 | 8.4029 | 0.003579 | 0.005325 | 0.005264 | 6.7e-05 | 0.001078 | 0.32244 | 1.3205 | 2.5479 | 99.6 | 20.84 | 37.547 | 68.7149 | 107.7731 | 65.4717 | 0.009294 | 1.0426 | 0.024818 | 23.8041 | 0.064376 | 0.14059 | -0.000183 | -0.013416 | -0.013416 | -0.43071
2019-03-21 11:45 | 1553183076.3539 | -37 | 0.20954 | 0.000462 | -3e-06 | 95.7227 | 95.1927 | 123.5701 | 21.0303 | 0.001916 | 0.001858 | -151.6838 | 0.23641 | 8.4033 | 0.003554 | 0.004679 | 0.004415 | 7e-05 | 0.001075 | 0.30917 | 1.3268 | 2.548 | 99.6 | 20.84 | 37.543 | 67.6986 | 107.1674 | 66.3184 | 0.009293 | 1.04 | 0.024797 | 23.8434 | 0.049655 | 0.14064 | -8.2e-05 | -0.005871 | -0.005871 | -0.98463
2019-03-21 11:45 | 1553183077.354 | -36 | 0.20953 | 0.000459 | -4e-06 | 95.7933 | 95.3334 | 123.3666 | 21.0288 | 0.001927 | 0.001852 | -151.6954 | 0.23636 | 8.4023 | 0.003568 | 0.005139 | 0.004965 | 6.9e-05 | 0.001073 | 0.33555 | 1.2899 | 2.5479 | 99.6 | 20.84 | 37.54 | 67.9599 | 107.5171 | 66.1757 | 0.009292 | 1.0404 | 0.024782 | 23.8187 | 0.043044 | 0.14061 | -3.4e-05 | -0.002285 | -0.002285 | -2.529
2019-03-21 11:45 | 1553183078.3536 | -35 | 0.20958 | 0.000459 | -4e-06 | 95.8529 | 95.3732 | 123.5016 | 21.0297 | 0.001926 | 0.001787 | -151.7073 | 0.2365 | 8.402 | 0.003476 | 0.004835 | 0.004682 | 6.6e-05 | 0.001072 | 0.32693 | 1.3029 | 2.5479 | 99.6 | 20.84 | 37.533 | 67.9366 | 105.2701 | 64.5976 | 0.009291 | 1.0453 | 0.024853 | 23.7763 | 0.044868 | 0.1407 | -0.000304 | -0.022403 | -0.022403 | -0.25812
2019-03-21 11:45 | 1553183079.3527 | -34 | 0.20956 | 0.000461 | -4e-06 | 95.9893 | 95.5476 | 123.5983 | 21.03 | 0.001947 | 0.001844 | -151.7097 | 0.23666 | 8.4032 | 0.003523 | 0.005053 | 0.004891 | 6.8e-05 | 0.001071 | 0.30182 | 1.2915 | 2.5481 | 99.6 | 20.84 | 37.532 | 68.4508 | 106.4114 | 65.9707 | 0.00929 | 1.0411 | 0.024813 | 23.8339 | 0.041253 | 0.14079 | -0.000214 | -0.015744 | -0.015744 | -0.36753
2019-03-21 11:45 | 1553183080.3541 | -33 | 0.20956 | 0.000461 | -4e-06 | 96.065 | 95.4996 | 123.583 | 21.0301 | 0.001947 | 0.001829 | -151.7167 | 0.23605 | 8.4021 | 0.003494 | 0.004693 | 0.004467 | 6.9e-05 | 0.001076 | 0.30502 | 1.2792 | 2.548 | 99.6 | 20.84 | 37.529 | 68.455 | 105.7236 | 65.6009 | 0.00929 | 1.0422 | 0.024825 | 23.8194 | 0.052799 | 0.14043 | -0.000201 | -0.014762 | -0.014762 | -0.39096
2019-03-21 11:45 | 1553183081.3538 | -32 | 0.20951 | 0.000458 | -4e-06 | 96.1803 | 95.8984 | 123.5018 | 21.0294 | 0.001931 | 0.001848 | -151.6992 | 0.23671 | 8.4004 | 0.003481 | 0.005032 | 0.004813 | 6.8e-05 | 0.001079 | 0.33876 | 1.2853 | 2.5478 | 99.6 | 20.84 | 37.528 | 68.0596 | 105.4012 | 66.0599 | 0.009289 | 1.0408 | 0.0248 | 23.8277 | 0.026252 | 0.14082 | 6.7e-05 | 0.005224 | 0.005224 | 1.108
2019-03-21 11:45 | 1553183082.3539 | -31 | 0.20953 | 0.000458 | -4e-06 | 96.4064 | 95.7441 | 123.452 | 21.0308 | 0.001977 | 0.001845 | -151.6734 | 0.23634 | 8.4012 | 0.003503 | 0.005636 | 0.005483 | 6.8e-05 | 0.001074 | 0.28783 | 1.3027 | 2.5479 | 99.6 | 20.84 | 37.526 | 69.1629 | 105.9414 | 65.9857 | 0.009289 | 1.041 | 0.024797 | 23.8203 | 0.06166 | 0.1406 | -2.1e-05 | -0.001326 | -0.001326 | -4.3582
2019-03-21 11:45 | 1553183083.354 | -30 | 0.20954 | 0.000465 | -4e-06 | 96.4887 | 96.0622 | 123.6933 | 21.0303 | 0.001954 | 0.001832 | -151.6942 | 0.23697 | 8.4037 | 0.003522 | 0.004336 | 0.004183 | 6.5e-05 | 0.001074 | 0.31902 | 1.2947 | 2.548 | 99.6 | 20.85 | 37.522 | 68.6175 | 106.3873 | 65.6793 | 0.009294 | 1.042 | 0.024833 | 23.8328 | 0.039617 | 0.14098 | -9.4e-05 | -0.006758 | -0.006758 | -0.85736
2019-03-21 11:45 | 1553183084.3534 | -29 | 0.20955 | 0.000456 | -4e-06 | 96.581 | 96.1392 | 123.9053 | 21.0296 | 0.001924 | 0.001814 | -151.6879 | 0.2368 | 8.4025 | 0.003497 | 0.004848 | 0.004614 | 6.6e-05 | 0.001077 | 0.35474 | 1.28 | 2.5479 | 99.6 | 20.85 | 37.521 | 67.8982 | 105.7939 | 65.2617 | 0.009293 | 1.0432 | 0.024869 | 23.8385 | 0.041011 | 0.14088 | -0.000115 | -0.008306 | -0.008306 | -0.69706
2019-03-21 11:45 | 1553183085.3541 | -28 | 0.20955 | 0.00046 | -3e-06 | 96.5792 | 96.222 | 124.0459 | 21.0296 | 0.001982 | 0.001833 | -151.6753 | 0.23792 | 8.404 | 0.003511 | 0.004622 | 0.004458 | 6.2e-05 | 0.001079 | 0.29966 | 1.3155 | 2.5481 | 99.6 | 20.85 | 37.519 | 69.3007 | 106.1198 | 65.7176 | 0.009293 | 1.0418 | 0.024867 | 23.8681 | 0.033148 | 0.14154 | -0.000147 | -0.010784 | -0.010784 | -0.53941
2019-03-21 11:45 | 1553183086.3531 | -27 | 0.20953 | 0.00046 | -4e-06 | 96.7166 | 96.365 | 124.4254 | 21.0298 | 0.001971 | 0.001824 | -151.6933 | 0.23757 | 8.399 | 0.00346 | 0.005085 | 0.004821 | 6.2e-05 | 0.00108 | 0.29214 | 1.3815 | 2.548 | 99.6 | 20.85 | 37.518 | 69.017 | 104.8895 | 65.4844 | 0.009293 | 1.0426 | 0.024913 | 23.8963 | 0.032573 | 0.14133 | -5.3e-05 | -0.003745 | -0.003745 | -1.5509
2019-03-21 11:45 | 1553183087.3535 | -26 | 0.20952 | 0.000462 | -3e-06 | 96.8656 | 96.6753 | 124.8505 | 21.0307 | 0.001913 | 0.001806 | -151.6795 | 0.23697 | 8.4036 | 0.003554 | 0.004401 | 0.004316 | 6.6e-05 | 0.00108 | 0.30735 | 1.3129 | 2.548 | 99.6 | 20.85 | 37.516 | 67.6365 | 107.1867 | 65.0506 | 0.009292 | 1.0439 | 0.024972 | 23.9218 | 0.017586 | 0.14097 | 4.7e-05 | 0.003727 | 0.003727 | 1.5546
2019-03-21 11:45 | 1553183088.3532 | -25 | 0.20955 | 0.000456 | -4e-06 | 96.9496 | 96.6641 | 124.6637 | 21.0312 | 0.001958 | 0.001796 | -151.6424 | 0.23684 | 8.4026 | 0.003529 | 0.005215 | 0.005023 | 6.4e-05 | 0.001082 | 0.29479 | 1.2937 | 2.548 | 99.6 | 20.85 | 37.515 | 68.7174 | 106.581 | 64.823 | 0.009292 | 1.0446 | 0.024962 | 23.8958 | 0.026381 | 0.1409 | -0.000118 | -0.008577 | -0.008577 | -0.67518
2019-03-21 11:45 | 1553183089.3542 | -24 | 0.20954 | 0.000459 | -4e-06 | 96.9904 | 96.6872 | 124.647 | 21.0315 | 0.001945 | 0.001782 | -151.6507 | 0.23703 | 8.3998 | 0.003553 | 0.004195 | 0.003947 | 6.4e-05 | 0.001084 | 0.2959 | 1.3042 | 2.5478 | 99.6 | 20.85 | 37.514 | 68.3947 | 107.1592 | 64.4752 | 0.009292 | 1.0457 | 0.024973 | 23.8819 | 0.028003 | 0.14101 | -9.1e-05 | -0.006602 | -0.006602 | -0.87783
2019-03-21 11:45 | 1553183090.3537 | -23 | 0.20955 | 0.000461 | -4e-06 | 96.98 | 96.5994 | 124.8616 | 21.0305 | 0.001952 | 0.001736 | -151.6517 | 0.23703 | 8.403 | 0.00349 | 0.005458 | 0.005299 | 6e-05 | 0.001086 | 0.32621 | 1.2886 | 2.5478 | 99.6 | 20.85 | 37.513 | 68.559 | 105.6329 | 63.3748 | 0.009291 | 1.0491 | 0.025035 | 23.8635 | 0.035174 | 0.14101 | -0.000135 | -0.009878 | -0.009878 | -0.58667
2019-03-21 11:45 | 1553183091.3534 | -22 | 0.20954 | 0.000465 | -4e-06 | 97.0628 | 96.8626 | 125.0411 | 21.0313 | 0.001955 | 0.001823 | -151.6553 | 0.23723 | 8.4007 | 0.003529 | 0.005148 | 0.00503 | 6.4e-05 | 0.001085 | 0.31222 | 1.3036 | 2.548 | 99.6 | 20.85 | 37.509 | 68.6348 | 106.5733 | 65.4758 | 0.00929 | 1.0426 | 0.024975 | 23.9551 | 0.018466 | 0.14113 | -0.000123 | -0.009015 | -0.009015 | -0.64342
2019-03-21 11:45 | 1553183092.3548 | -21 | 0.20954 | 0.000458 | -4e-06 | 97.1373 | 96.7175 | 125.1918 | 21.0304 | 0.001939 | 0.001731 | -151.6628 | 0.23753 | 8.4031 | 0.003491 | 0.004961 | 0.004763 | 6.3e-05 | 0.001086 | 0.33107 | 1.3018 | 2.5478 | 99.6 | 20.85 | 37.507 | 68.25 | 105.6423 | 63.2647 | 0.00929 | 1.0494 | 0.025072 | 23.8911 | 0.038735 | 0.1413 | -7.9e-05 | -0.005704 | -0.005704 | -1.0182
2019-03-21 11:45 | 1553183093.3538 | -20 | 0.20955 | 0.000459 | -4e-06 | 97.1456 | 96.7356 | 125.1864 | 21.0321 | 0.001939 | 0.001799 | -151.6556 | 0.23605 | 8.4008 | 0.00353 | 0.004162 | 0.004002 | 6.2e-05 | 0.001088 | 0.32637 | 1.3344 | 2.548 | 99.6 | 20.85 | 37.505 | 68.2572 | 106.5921 | 64.8903 | 0.009289 | 1.0444 | 0.025011 | 23.9483 | 0.037834 | 0.14043 | -0.000146 | -0.010724 | -0.010724 | -0.53821
2019-03-21 11:45 | 1553183094.3529 | -19 | 0.20955 | 0.000461 | -4e-06 | 97.1424 | 96.7734 | 125.204 | 21.0318 | 0.001916 | 0.001785 | -151.6599 | 0.23563 | 8.4053 | 0.003485 | 0.004956 | 0.004873 | 6.4e-05 | 0.001088 | 0.31653 | 1.2937 | 2.5479 | 99.6 | 20.85 | 37.504 | 67.7165 | 105.4979 | 64.5544 | 0.009289 | 1.0454 | 0.025025 | 23.938 | 0.034038 | 0.14018 | -0.000117 | -0.008565 | -0.008565 | -0.67269
2019-03-21 11:45 | 1553183095.3536 | -18 | 0.20954 | 0.00046 | -4e-06 | 97.1766 | 96.7642 | 125.3183 | 21.032 | 0.001956 | 0.001811 | -151.6679 | 0.23677 | 8.4032 | 0.003428 | 0.005178 | 0.005069 | 6.3e-05 | 0.001084 | 0.31509 | 1.3282 | 2.5481 | 99.6 | 20.85 | 37.502 | 68.6706 | 104.1188 | 65.1775 | 0.009289 | 1.0435 | 0.025014 | 23.9711 | 0.038042 | 0.14085 | -0.000114 | -0.008315 | -0.008315 | -0.69625
2019-03-21 11:45 | 1553183096.3526 | -17 | 0.20954 | 0.000459 | -4e-06 | 97.1397 | 96.9594 | 125.315 | 21.0322 | 0.001896 | 0.001833 | -151.6725 | 0.23708 | 8.4011 | 0.003461 | 0.005039 | 0.004756 | 6e-05 | 0.001084 | 0.31238 | 1.2995 | 2.548 | 99.6 | 20.85 | 37.499 | 67.2307 | 104.9241 | 65.7006 | 0.009288 | 1.0419 | 0.024994 | 23.9893 | 0.016619 | 0.14104 | -7.4e-05 | -0.005291 | -0.005291 | -1.0956
2019-03-21 11:45 | 1553183097.3536 | -16 | 0.20951 | 0.000457 | -4e-06 | 97.2824 | 97.1388 | 125.4392 | 21.0314 | 0.00195 | 0.001754 | -151.6882 | 0.23742 | 8.4028 | 0.003506 | 0.004413 | 0.004256 | 6.5e-05 | 0.001086 | 0.29839 | 1.2863 | 2.5479 | 99.6 | 20.85 | 37.499 | 68.5272 | 106.0169 | 63.799 | 0.009288 | 1.0478 | 0.025077 | 23.9337 | 0.013208 | 0.14124 | 7.3e-05 | 0.00574 | 0.00574 | 1.0114
2019-03-21 11:45 | 1553183098.3526 | -15 | 0.20955 | 0.000459 | -4e-06 | 97.3837 | 96.9594 | 125.6209 | 21.0323 | 0.001913 | 0.001737 | -151.6916 | 0.23703 | 8.4018 | 0.003525 | 0.004775 | 0.004593 | 6.4e-05 | 0.001085 | 0.27472 | 1.3115 | 2.548 | 99.6 | 20.85 | 37.498 | 67.6335 | 106.4695 | 63.4085 | 0.009288 | 1.049 | 0.02511 | 23.9372 | 0.03906 | 0.14101 | -0.000158 | -0.011651 | -0.011651 | -0.49741
2019-03-21 11:45 | 1553183099.3542 | -14 | 0.20955 | 0.00046 | -4e-06 | 97.3992 | 97.1709 | 125.4042 | 21.0332 | 0.001874 | 0.001752 | -151.6656 | 0.2376 | 8.4023 | 0.003572 | 0.00502 | 0.004866 | 6.6e-05 | 0.001085 | 0.30862 | 1.3026 | 2.5479 | 99.6 | 20.85 | 37.497 | 66.6872 | 107.6215 | 63.7688 | 0.009287 | 1.0479 | 0.025075 | 23.9293 | 0.020993 | 0.14135 | -0.000124 | -0.009107 | -0.009107 | -0.63793
2019-03-21 11:45 | 1553183100.3542 | -13 | 0.20955 | 0.000461 | -3e-06 | 97.5002 | 97.1399 | 125.4559 | 21.0322 | 0.001846 | 0.001747 | -151.6722 | 0.23712 | 8.4019 | 0.003595 | 0.004655 | 0.004543 | 6.7e-05 | 0.001084 | 0.30237 | 1.3582 | 2.5478 | 99.6 | 20.85 | 37.497 | 66.0121 | 108.1634 | 63.6482 | 0.009287 | 1.0482 | 0.025084 | 23.9299 | 0.033117 | 0.14106 | -0.000149 | -0.010986 | -0.010986 | -0.52773
2019-03-21 11:45 | 1553183101.3529 | -12 | 0.20957 | 0.000461 | -4e-06 | 97.5005 | 97.1967 | 125.6148 | 21.0334 | 0.001909 | 0.001791 | -151.7174 | 0.23673 | 8.4019 | 0.003653 | 0.005434 | 0.00514 | 6.2e-05 | 0.001074 | 0.31061 | 1.3539 | 2.548 | 99.6 | 20.86 | 37.497 | 67.5394 | 109.5802 | 64.7107 | 0.009293 | 1.0449 | 0.025061 | 23.9828 | 0.027908 | 0.14083 | -0.000247 | -0.01838 | -0.01838 | -0.31491
2019-03-21 11:45 | 1553183102.353 | -11 | 0.20956 | 0.000461 | -4e-06 | 97.488 | 97.2461 | 125.9856 | 21.0331 | 0.001872 | 0.001797 | -124.4204 | 0.23624 | 8.4023 | 0.003617 | 0.004778 | 0.004631 | 6.5e-05 | 0.001073 | 0.31852 | 1.3673 | 2.548 | 99.6 | 20.86 | 37.497 | 66.6583 | 108.7025 | 64.8422 | 0.009293 | 1.0445 | 0.025093 | 24.0229 | 0.022217 | 0.14054 | -0.000199 | -0.014698 | -0.014698 | -0.39299
2019-03-21 11:45 | 1553183103.3529 | -10 | 0.20953 | 0.000462 | -4e-06 | 97.5961 | 97.2955 | 126.0213 | 21.0334 | 0.001892 | 0.001754 | -90.9283 | 0.23706 | 8.403 | 0.00368 | 0.004793 | 0.00459 | 6.5e-05 | 0.001068 | 0.33428 | 1.3203 | 2.548 | 99.6 | 20.86 | 37.497 | 67.1399 | 110.2284 | 63.8173 | 0.009293 | 1.0477 | 0.025134 | 23.9898 | 0.027593 | 0.14103 | -2.7e-05 | -0.001792 | -0.001792 | -3.2352
2019-03-21 11:45 | 1553183104.354 | -9 | 0.20958 | 0.000459 | -4e-06 | 97.6074 | 97.2138 | 126.0356 | 21.0325 | 0.001922 | 0.001717 | -57.4525 | 0.23636 | 8.4014 | 0.003679 | 0.004422 | 0.004232 | 6.2e-05 | 0.001066 | 0.31708 | 1.3266 | 2.5479 | 99.6 | 20.86 | 37.498 | 67.8614 | 110.2226 | 62.9191 | 0.009293 | 1.0505 | 0.025169 | 23.9592 | 0.036148 | 0.14061 | -0.000343 | -0.025721 | -0.025721 | -0.22468
2019-03-21 11:45 | 1553183105.3547 | -8 | 0.20961 | 0.000462 | -3e-06 | 97.6253 | 97.2737 | 126.2408 | 21.0326 | 0.001845 | 0.001756 | -23.9652 | 0.23677 | 8.4026 | 0.003662 | 0.005264 | 0.005041 | 6.4e-05 | 0.001067 | 0.31382 | 1.3059 | 2.548 | 99.6 | 20.86 | 37.498 | 65.9921 | 109.8112 | 63.8462 | 0.009293 | 1.0476 | 0.025155 | 24.0117 | 0.032273 | 0.14086 | -0.000504 | -0.037894 | -0.037894 | -0.15277
2019-03-21 11:45 | 1553183106.3532 | -7 | 0.20952 | 0.000462 | -4e-06 | 97.7153 | 97.4566 | 126.1218 | 21.0325 | 0.001887 | 0.001725 | 9.5288 | 0.23662 | 8.4013 | 0.003616 | 0.004647 | 0.004454 | 6.2e-05 | 0.001065 | 0.30934 | 1.3342 | 2.548 | 99.6 | 20.86 | 37.498 | 67.0141 | 108.6767 | 63.1113 | 0.009293 | 1.0499 | 0.025171 | 23.9742 | 0.023711 | 0.14076 | 4e-06 | 0.000608 | 0.000608 | 9.5154
2019-03-21 11:45 | 1553183107.3542 | -6 | 0.20956 | 0.00046 | -3e-06 | 97.7013 | 97.5091 | 126.3131 | 21.0328 | 0.001924 | 0.001744 | 43.0569 | 0.25796 | 8.4 | 0.003662 | 0.005006 | 0.004963 | 6e-05 | 0.001063 | 0.30469 | 1.3346 | 2.548 | 99.6 | 20.86 | 37.499 | 67.896 | 109.8043 | 63.5628 | 0.009294 | 1.0485 | 0.025173 | 24.0085 | 0.017617 | 0.15346 | -0.000188 | -0.014039 | -0.014039 | -0.44928
2019-03-21 11:45 | 1553183108.3534 | -5 | 0.20954 | 0.000457 | -3e-06 | 97.7713 | 97.5356 | 126.5199 | 21.0338 | 0.001839 | 0.001751 | 183.1563 | 0.26245 | 8.4015 | 0.003682 | 0.005182 | 0.005096 | 6.1e-05 | 0.001057 | 0.29374 | 1.3358 | 2.548 | 99.6 | 20.86 | 37.499 | 65.8527 | 110.2934 | 63.7458 | 0.009294 | 1.0479 | 0.025187 | 24.0347 | 0.021588 | 0.15613 | -9.1e-05 | -0.006701 | -0.006701 | -0.95763
2019-03-21 11:45 | 1553183109.3544 | -4 | 0.20958 | 0.000463 | -3e-06 | 97.8538 | 97.633 | 126.8468 | 21.0331 | 0.001853 | 0.001716 | 183.0809 | 0.262 | 8.3997 | 0.00364 | 0.004912 | 0.004811 | 6.4e-05 | 0.001056 | 0.29789 | 1.3608 | 2.548 | 99.6 | 20.86 | 37.499 | 66.2025 | 109.263 | 62.9002 | 0.009294 | 1.0506 | 0.025251 | 24.0355 | 0.020206 | 0.15586 | -0.000333 | -0.025075 | -0.025075 | -0.25548
2019-03-21 11:45 | 1553183110.3533 | -3 | 0.2095 | 0.000455 | -3e-06 | 97.8538 | 97.5557 | 126.9093 | 21.0333 | 0.001868 | 0.001757 | 183.2444 | 0.26205 | 8.4011 | 0.003584 | 0.005077 | 0.005017 | 6.3e-05 | 0.001049 | 0.33987 | 1.3374 | 2.548 | 99.6 | 20.86 | 37.5 | 66.5633 | 107.9052 | 63.8719 | 0.009294 | 1.0475 | 0.025221 | 24.0762 | 0.027292 | 0.15589 | 0.000162 | 0.01246 | 0.01246 | 0.51422
2019-03-21 11:45 | 1553183111.3528 | -2 | 0.20954 | 0.000458 | -3e-06 | 97.8366 | 97.5111 | 126.8673 | 21.0325 | 0.001875 | 0.00172 | 183.2789 | 0.26222 | 8.4024 | 0.003572 | 0.005491 | 0.005419 | 6.5e-05 | 0.001052 | 0.31741 | 1.3316 | 2.548 | 99.6 | 20.86 | 37.5 | 66.7294 | 107.6099 | 62.9825 | 0.009294 | 1.0503 | 0.02525 | 24.0404 | 0.029804 | 0.15599 | -0.000107 | -0.007915 | -0.007915 | -0.81007
2019-03-21 11:45 | 1553183112.3722 | -1 | 0.20955 | 0.000457 | -4e-06 | 97.8723 | 97.6232 | 126.6882 | 21.0331 | 0.001899 | 0.001701 | 183.467 | 0.26223 | 8.4027 | 0.003635 | 0.005332 | 0.00527 | 6.2e-05 | 0.001047 | 0.29302 | 1.346 | 2.5479 | 99.6 | 20.86 | 37.5 | 67.3066 | 109.1485 | 62.5323 | 0.009294 | 1.0517 | 0.025249 | 24.0073 | 0.022799 | 0.156 | -0.000114 | -0.008408 | -0.008408 | -0.76258
2019-03-21 11:45 | 1553183113.3624 | 0 | 0.20953 | 0.000462 | -4e-06 | 97.8023 | 97.6085 | 126.3535 | 21.031 | 0.001856 | 0.001712 | 183.1417 | 0.26249 | 8.403 | 0.003572 | 0.005066 | 0.005104 | 6.1e-05 | 0.00104 | 0.33716 | 1.2964 | 2.7218 | 99.6 | 20.86 | 37.5 | 66.2707 | 107.6123 | 62.8091 | 0.009294 | 1.0509 | 0.025205 | 23.9855 | 0.01774 | 0.15615 | -1.4e-05 | -0.000845 | -0.000845 | -7.5995
2019-03-21 11:45 | 1553183114.3619 | 1 | 0.20955 | 0.00046 | -3e-06 | 96.3957 | 89.103 | 125.5852 | 21.0234 | 0.001875 | 0.001765 | 183.5429 | 0.26288 | 8.4055 | 0.003642 | 0.005585 | 0.005324 | 6.2e-05 | 0.001024 | 0.34634 | 1.2745 | 2.5479 | 99.6 | 20.86 | 37.5 | 66.7096 | 109.3142 | 64.0762 | 0.009294 | 1.0469 | 0.025081 | 23.9575 | 0.70365 | 0.15639 | -0.000159 | -0.01175 | -0.01175 | -0.54702
2019-03-21 11:45 | 1553183115.3625 | 2 | 0.20952 | 0.00046 | -3e-06 | 95.1932 | 78.5645 | 124.1001 | 21.0192 | 0.001854 | 0.001697 | 183.4754 | 0.26333 | 8.402 | 0.003609 | 0.007461 | 0.00779 | 5.9e-05 | 0.001027 | 0.35242 | 1.2974 | 2.5479 | 99.6 | 20.86 | 37.499 | 66.1929 | 108.506 | 62.4211 | 0.009294 | 1.0521 | 0.024994 | 23.7569 | 1.7172 | 0.15665 | 2.6e-05 | 0.002132 | 0.002132 | 3.0197
2019-03-21 11:45 | 1553183116.3626 | 3 | 0.20952 | 0.000462 | -2e-06 | 95.1837 | 74.7959 | 122.5725 | 21.0212 | 0.00189 | 0.00178 | 183.4706 | 0.26301 | 8.4017 | 0.003588 | 0.006211 | 0.006353 | 6.3e-05 | 0.001032 | 0.3454 | 1.3084 | 2.5479 | 99.6 | 20.86 | 37.498 | 67.0688 | 108.0021 | 64.4161 | 0.009293 | 1.0459 | 0.024766 | 23.6803 | 2.156 | 0.15647 | 6e-06 | 0.000571 | 0.000571 | 11.2664
2019-03-21 11:45 | 1553183117.3617 | 4 | 0.20948 | 0.000463 | 0 | 95.7858 | 75.2976 | 120.4741 | 21.0198 | 0.001861 | 0.001864 | 183.5195 | 0.26149 | 8.4027 | 0.003679 | 0.00771 | 0.008338 | 6.2e-05 | 0.001043 | 0.34612 | 1.2712 | 2.5481 | 99.6 | 20.86 | 37.496 | 66.3628 | 110.2041 | 66.4479 | 0.009293 | 1.0396 | 0.02448 | 23.5473 | 2.1526 | 0.15556 | 0.000251 | 0.018479 | 0.018479 | 0.34598
2019-03-21 11:45 | 1553183118.3617 | 5 | 0.20947 | 0.000464 | 3e-06 | 96.5563 | 81.3145 | 118.1468 | 21.0207 | 0.001905 | 0.001889 | 183.585 | 0.26165 | 8.4025 | 0.003548 | 0.007009 | 0.007534 | 6.3e-05 | 0.001037 | 0.31166 | 1.2908 | 2.5479 | 99.6 | 20.86 | 37.493 | 67.427 | 107.0257 | 67.0418 | 0.009292 | 1.0378 | 0.024221 | 23.3391 | 1.5367 | 0.15566 | 0.000323 | 0.023306 | 0.023306 | 0.2745
2019-03-21 11:45 | 1553183119.3621 | 6 | 0.20905 | 0.000464 | 6e-06 | 97.2651 | 87.9787 | 115.7163 | 21.02 | 0.001983 | 0.002001 | 183.291 | 0.26301 | 8.4019 | 0.003608 | 0.004839 | 0.004941 | 6.6e-05 | 0.001043 | 0.32781 | 1.3212 | 2.5481 | 99.6 | 20.86 | 37.491 | 69.3101 | 108.4643 | 69.743 | 0.009292 | 1.0296 | 0.023876 | 23.1893 | 0.89754 | 0.15647 | 0.002847 | 0.20379 | 0.20379 | 0.031555
2019-03-21 11:45 | 1553183120.3628 | 7 | 0.20813 | 0.000475 | 1.2e-05 | 97.592 | 90.0189 | 113.1934 | 21.0278 | 0.002041 | 0.002029 | 183.2158 | 0.26277 | 8.4027 | 0.003545 | 0.005269 | 0.005463 | 5.9e-05 | 0.001027 | 0.28982 | 1.2899 | 1.8435 | 99.6 | 20.86 | 37.486 | 70.7164 | 106.9628 | 70.4086 | 0.00929 | 1.0276 | 0.023591 | 22.9574 | 0.7225 | 0.15632 | 0.008351 | 0.59054 | 0.59054 | 0.01088
2019-03-21 11:45 | 1553183121.3621 | 8 | 0.20671 | 0.000491 | 1.9e-05 | 97.7111 | 92.9425 | 110.9293 | 21.0345 | 0.002065 | 0.002133 | 183.113 | 0.26313 | 8.4018 | 0.003589 | 0.004981 | 0.004984 | 5.9e-05 | 0.00104 | 0.29103 | 1.2842 | 2.5478 | 99.6 | 20.86 | 37.483 | 71.3009 | 108.0207 | 72.9209 | 0.00929 | 1.0202 | 0.023269 | 22.8096 | 0.44754 | 0.15654 | 0.016869 | 1.1759 | 1.1759 | 0.005471
2019-03-21 11:45 | 1553183122.3628 | 9 | 0.20541 | 0.000554 | 2.9e-05 | 97.855 | 93.4598 | 108.9475 | 21.0348 | 0.002143 | 0.002193 | 183.0033 | 0.2628 | 8.4042 | 0.003672 | 0.005115 | 0.005001 | 6.4e-05 | 0.001039 | 0.23706 | 1.2933 | 2.5479 | 99.6 | 20.86 | 37.48 | 73.1689 | 110.0429 | 74.3663 | 0.009289 | 1.0159 | 0.023012 | 22.6521 | 0.41105 | 0.15634 | 0.024552 | 1.6911 | 1.6911 | 0.0038
2019-03-21 11:45 | 1553183123.3619 | 10 | 0.20403 | 0.000655 | 4e-05 | 97.9039 | 91.1595 | 107.4575 | 21.0352 | 0.002274 | 0.002291 | 182.8962 | 0.26303 | 8.4013 | 0.003628 | 0.005224 | 0.005125 | 6.2e-05 | 0.001044 | 0.29535 | 1.3212 | 2.5479 | 99.6 | 20.86 | 37.475 | 76.3178 | 108.9827 | 76.7168 | 0.009288 | 1.0091 | 0.022778 | 22.5726 | 0.63842 | 0.15648 | 0.032702 | 2.2274 | 2.2274 | 0.002887
2019-03-21 11:45 | 1553183124.362 | 11 | 0.20287 | 0.000977 | 5.2e-05 | 97.928 | 87.1606 | 106.4151 | 21.035 | 0.002346 | 0.00231 | 182.7525 | 0.26249 | 8.402 | 0.003665 | 0.004983 | 0.004839 | 6.1e-05 | 0.001041 | 0.33212 | 1.3464 | 2.548 | 99.6 | 20.86 | 37.473 | 78.0458 | 109.8775 | 77.1751 | 0.009287 | 1.0078 | 0.022652 | 22.4775 | 1.0419 | 0.15615 | 0.03915 | 2.6499 | 2.6499 | 0.002422
2019-03-21 11:45 | 1553183125.362 | 12 | 0.20249 | 0.001453 | 6.3e-05 | 97.9405 | 82.6768 | 105.7154 | 21.0358 | 0.002351 | 0.002365 | 182.6015 | 0.26368 | 8.4012 | 0.003784 | 0.005353 | 0.005228 | 6.2e-05 | 0.001037 | 0.28495 | 1.3218 | 2.5479 | 99.6 | 20.86 | 37.471 | 78.1563 | 112.7826 | 78.4987 | 0.009287 | 1.004 | 0.022535 | 22.4458 | 1.5154 | 0.15687 | 0.040866 | 2.7508 | 2.7508 | 0.002344
2019-03-21 11:45 | 1553183126.3611 | 13 | 0.20185 | 0.001986 | 7.3e-05 | 97.8839 | 77.5578 | 105.4398 | 21.0352 | 0.002439 | 0.002456 | 182.4843 | 0.26347 | 8.4025 | 0.003829 | 0.004888 | 0.00486 | 6.1e-05 | 0.001038 | 0.25625 | 1.3022 | 2.5479 | 99.6 | 20.86 | 37.465 | 80.2843 | 113.8733 | 80.6943 | 0.009285 | 0.99774 | 0.022436 | 22.4864 | 2.0819 | 0.15674 | 0.04401 | 2.9481 | 2.9481 | 0.002185
2019-03-21 11:45 | 1553183127.3622 | 14 | 0.20132 | 0.002507 | 8.2e-05 | 97.9626 | 73.9281 | 105.1211 | 21.036 | 0.002429 | 0.002492 | 182.3638 | 0.2631 | 8.4052 | 0.003825 | 0.004762 | 0.004639 | 6.3e-05 | 0.001034 | 0.32549 | 1.2755 | 2.548 | 99.6 | 20.86 | 37.465 | 80.0253 | 113.7737 | 81.5473 | 0.009285 | 0.99534 | 0.022375 | 22.4794 | 2.5178 | 0.15652 | 0.04648 | 3.1041 | 3.1041 | 0.002072
2019-03-21 11:45 | 1553183128.3613 | 15 | 0.20092 | 0.002975 | 9e-05 | 97.8774 | 72.5208 | 105.2035 | 21.0356 | 0.002534 | 0.002514 | 182.2302 | 0.26307 | 8.4031 | 0.003842 | 0.005426 | 0.005318 | 6.1e-05 | 0.001038 | 0.33323 | 1.3144 | 2.5479 | 99.6 | 20.86 | 37.467 | 82.5702 | 114.178 | 82.0765 | 0.009286 | 0.99386 | 0.022367 | 22.505 | 2.6819 | 0.1565 | 0.048298 | 3.2235 | 3.2235 | 0.001995
2019-03-21 11:45 | 1553183129.3625 | 16 | 0.20087 | 0.003378 | 9.7e-05 | 97.9212 | 72.0757 | 105.1215 | 21.0353 | 0.002567 | 0.002594 | 182.1209 | 0.26376 | 8.4005 | 0.003872 | 0.005279 | 0.00518 | 6.1e-05 | 0.001037 | 0.29856 | 1.2893 | 2.5478 | 99.6 | 20.86 | 37.47 | 83.3461 | 114.9242 | 84.0068 | 0.009286 | 0.98849 | 0.022298 | 22.5573 | 2.741 | 0.15691 | 0.048108 | 3.2005 | 3.2005 | 0.002015
2019-03-21 11:45 | 1553183130.3623 | 17 | 0.20078 | 0.003683 | 0.000102 | 97.8056 | 72.0975 | 104.7537 | 21.0369 | 0.002625 | 0.002694 | 181.9932 | 0.26281 | 8.4026 | 0.003826 | 0.005807 | 0.005748 | 6.2e-05 | 0.001035 | 0.30901 | 1.291 | 2.5479 | 99.6 | 20.86 | 37.477 | 84.7466 | 113.7848 | 86.4161 | 0.009288 | 0.98186 | 0.022184 | 22.5936 | 2.7278 | 0.15635 | 0.048269 | 3.1945 | 3.1945 | 0.002012
2019-03-21 11:46 | 1553183131.3617 | 18 | 0.20044 | 0.003972 | 0.000107 | 97.883 | 72.5869 | 104.3754 | 21.0356 | 0.002664 | 0.002674 | 181.8878 | 0.26255 | 8.4024 | 0.003874 | 0.005847 | 0.005714 | 6e-05 | 0.001032 | 0.32803 | 1.2962 | 2.5481 | 99.6 | 20.86 | 37.482 | 85.695 | 114.9664 | 85.9355 | 0.009289 | 0.98318 | 0.022159 | 22.5377 | 2.6743 | 0.15619 | 0.049919 | 3.2993 | 3.2993 | 0.001946
2019-03-21 11:46 | 1553183132.3616 | 19 | 0.2003 | 0.004168 | 0.000111 | 97.799 | 73.3474 | 103.7029 | 21.0369 | 0.002735 | 0.00271 | 181.7553 | 0.26386 | 8.4027 | 0.003874 | 0.005385 | 0.005362 | 6.2e-05 | 0.001032 | 0.28799 | 1.3244 | 2.548 | 99.6 | 20.87 | 37.486 | 87.4061 | 114.9558 | 86.7901 | 0.009296 | 0.98084 | 0.022061 | 22.4917 | 2.5734 | 0.15697 | 0.05051 | 3.3233 | 3.3233 | 0.001941
2019-03-21 11:46 | 1553183133.3614 | 20 | 0.19998 | 0.004359 | 0.000115 | 97.7472 | 73.6281 | 102.6091 | 21.0367 | 0.002873 | 0.002854 | 181.6096 | 0.26331 | 8.4024 | 0.003902 | 0.004974 | 0.004916 | 6.6e-05 | 0.00103 | 0.30879 | 1.3051 | 2.548 | 99.6 | 20.87 | 37.495 | 90.7251 | 115.6534 | 90.2522 | 0.009298 | 0.9715 | 0.021839 | 22.4801 | 2.5345 | 0.15664 | 0.052171 | 3.3975 | 3.3975 | 0.001895
2019-03-21 11:46 | 1553183134.3615 | 21 | 0.19974 | 0.00456 | 0.000119 | 97.6825 | 72.9575 | 101.5077 | 21.0369 | 0.002948 | 0.002908 | 181.4931 | 0.26294 | 8.4041 | 0.004027 | 0.005126 | 0.00503 | 6.6e-05 | 0.00103 | 0.28838 | 1.3355 | 2.5478 | 99.6 | 20.87 | 37.5 | 92.5237 | 118.6975 | 91.5563 | 0.0093 | 0.96803 | 0.021683 | 22.3992 | 2.6104 | 0.15642 | 0.053383 | 3.451 | 3.451 | 0.001863
2019-03-21 11:46 | 1553183135.3626 | 22 | 0.19918 | 0.004764 | 0.000122 | 97.6408 | 73.374 | 100.413 | 21.0368 | 0.002984 | 0.002987 | 181.3633 | 0.26305 | 8.4035 | 0.003981 | 0.005761 | 0.005707 | 6.6e-05 | 0.001026 | 0.30431 | 1.3116 | 2.5478 | 99.6 | 20.87 | 37.503 | 93.402 | 117.5857 | 93.4763 | 0.0093 | 0.96296 | 0.021509 | 22.3366 | 2.5557 | 0.15649 | 0.056451 | 3.6191 | 3.6191 | 0.001777
2019-03-21 11:46 | 1553183136.3632 | 23 | 0.19828 | 0.004976 | 0.000124 | 97.631 | 72.9595 | 99.4701 | 21.0373 | 0.003025 | 0.002992 | 181.1831 | 0.26348 | 8.4029 | 0.00405 | 0.00542 | 0.005366 | 6.6e-05 | 0.001031 | 0.27936 | 1.3185 | 2.5479 | 99.6 | 20.87 | 37.512 | 94.38 | 119.2679 | 93.5855 | 0.009303 | 0.96267 | 0.021405 | 22.2348 | 2.6055 | 0.15674 | 0.06151 | 3.9227 | 3.9227 | 0.001642
2019-03-21 11:46 | 1553183137.3617 | 24 | 0.19756 | 0.005306 | 0.000125 | 97.572 | 72.2068 | 98.9063 | 21.0372 | 0.003139 | 0.003124 | 181.0317 | 0.2631 | 8.4028 | 0.004083 | 0.005346 | 0.005221 | 6.2e-05 | 0.001033 | 0.30685 | 1.3297 | 2.5479 | 99.6 | 20.87 | 37.516 | 97.1301 | 120.0752 | 96.7838 | 0.009304 | 0.95435 | 0.021252 | 22.2682 | 2.6928 | 0.15652 | 0.065358 | 4.137 | 4.137 | 0.001555
2019-03-21 11:46 | 1553183138.3627 | 25 | 0.19709 | 0.00561 | 0.000126 | 97.4936 | 72.1063 | 98.5061 | 21.0384 | 0.003235 | 0.003166 | 180.8957 | 0.26338 | 8.4033 | 0.004122 | 0.006072 | 0.005999 | 6.4e-05 | 0.001032 | 0.33516 | 1.3755 | 2.5481 | 99.6 | 20.87 | 37.52 | 99.455 | 121.0182 | 97.7793 | 0.009305 | 0.95179 | 0.02118 | 22.253 | 2.6981 | 0.15669 | 0.067804 | 4.2764 | 4.2764 | 0.001506
2019-03-21 11:46 | 1553183139.3618 | 26 | 0.19683 | 0.005955 | 0.000126 | 97.5252 | 73.013 | 98.5695 | 21.0386 | 0.003276 | 0.003242 | 180.7495 | 0.26318 | 8.4005 | 0.004074 | 0.005754 | 0.005627 | 6.6e-05 | 0.001029 | 0.27887 | 1.3098 | 2.548 | 99.6 | 20.87 | 37.521 | 100.4595 | 119.8514 | 99.6254 | 0.009305 | 0.94707 | 0.021134 | 22.3155 | 2.5892 | 0.15657 | 0.068953 | 4.3392 | 4.3392 | 0.001483
2019-03-21 11:46 | 1553183140.3618 | 27 | 0.19663 | 0.00631 | 0.000125 | 97.5997 | 73.1492 | 99.2276 | 21.0385 | 0.003369 | 0.003252 | 180.5965 | 0.26242 | 8.4009 | 0.004125 | 0.005358 | 0.005301 | 6.3e-05 | 0.001025 | 0.30574 | 1.3058 | 2.5479 | 99.6 | 20.87 | 37.523 | 102.6916 | 121.1032 | 99.863 | 0.009305 | 0.94647 | 0.021198 | 22.397 | 2.5794 | 0.15611 | 0.069688 | 4.3984 | 4.3984 | 0.001459
2019-03-21 11:46 | 1553183141.3612 | 28 | 0.19664 | 0.006642 | 0.000124 | 97.6897 | 73.6519 | 99.991 | 21.0384 | 0.003343 | 0.003292 | 180.4423 | 0.26257 | 8.4015 | 0.004175 | 0.005894 | 0.005776 | 6.7e-05 | 0.001021 | 0.29983 | 1.3034 | 2.5478 | 99.6 | 20.87 | 37.523 | 102.0732 | 122.3285 | 100.8241 | 0.009305 | 0.94404 | 0.021252 | 22.5119 | 2.5263 | 0.1562 | 0.06928 | 4.384 | 4.384 | 0.001464
2019-03-21 11:46 | 1553183142.3613 | 29 | 0.1967 | 0.00692 | 0.000123 | 97.6319 | 74.08 | 100.8378 | 21.0381 | 0.003408 | 0.003302 | 180.3003 | 0.26283 | 8.4042 | 0.004205 | 0.005623 | 0.005536 | 6.8e-05 | 0.001022 | 0.27903 | 1.3122 | 2.5479 | 99.6 | 20.87 | 37.524 | 103.638 | 123.0486 | 101.0752 | 0.009306 | 0.9434 | 0.021335 | 22.6146 | 2.4692 | 0.15636 | 0.06859 | 4.3575 | 4.3575 | 0.001475
2019-03-21 11:46 | 1553183143.362 | 30 | 0.19696 | 0.007152 | 0.000122 | 97.7016 | 74.3941 | 102.0769 | 21.0387 | 0.003492 | 0.003297 | 180.1528 | 0.26203 | 8.4049 | 0.004206 | 0.005576 | 0.005508 | 6.9e-05 | 0.001026 | 0.32477 | 1.3204 | 2.5479 | 99.6 | 20.87 | 37.522 | 105.6896 | 123.0771 | 100.9509 | 0.009305 | 0.94372 | 0.021469 | 22.7494 | 2.4378 | 0.15588 | 0.066784 | 4.2702 | 4.2702 | 0.0015
2019-03-21 11:46 | 1553183144.3621 | 31 | 0.19738 | 0.007371 | 0.000121 | 97.723 | 74.4962 | 103.174 | 21.039 | 0.003516 | 0.003277 | 180.0021 | 0.26349 | 8.4008 | 0.004193 | 0.005249 | 0.005129 | 6.7e-05 | 0.001026 | 0.29231 | 1.324 | 2.5481 | 99.6 | 20.87 | 37.522 | 106.2565 | 122.7644 | 100.4637 | 0.009305 | 0.94495 | 0.021598 | 22.8564 | 2.4274 | 0.15675 | 0.064048 | 4.1208 | 4.1208 | 0.001563
2019-03-21 11:46 | 1553183145.3621 | 32 | 0.19781 | 0.007573 | 0.000119 | 97.7963 | 73.9322 | 104.392 | 21.0389 | 0.003462 | 0.003274 | 179.8866 | 0.26249 | 8.4009 | 0.004231 | 0.005919 | 0.005811 | 6.6e-05 | 0.001028 | 0.3131 | 1.3374 | 2.5479 | 99.6 | 20.87 | 37.521 | 104.951 | 123.6899 | 100.4031 | 0.009305 | 0.9451 | 0.021727 | 22.9891 | 2.5021 | 0.15616 | 0.061261 | 3.966 | 3.966 | 0.001618
2019-03-21 11:46 | 1553183146.3621 | 33 | 0.19821 | 0.007747 | 0.000116 | 97.9379 | 73.5988 | 105.4419 | 21.0401 | 0.003534 | 0.00329 | 179.7648 | 0.2625 | 8.4018 | 0.004233 | 0.005671 | 0.005534 | 6.7e-05 | 0.001028 | 0.30127 | 1.3283 | 2.548 | 99.6 | 20.87 | 37.52 | 106.7086 | 123.7492 | 100.7865 | 0.009305 | 0.94413 | 0.021825 | 23.1162 | 2.5555 | 0.15616 | 0.058651 | 3.815 | 3.815 | 0.001682
2019-03-21 11:46 | 1553183147.3626 | 34 | 0.19878 | 0.007875 | 0.000115 | 97.9251 | 72.9554 | 106.3189 | 21.0398 | 0.003504 | 0.00335 | 179.6367 | 0.26336 | 8.4023 | 0.004347 | 0.005696 | 0.005592 | 6.3e-05 | 0.001026 | 0.32156 | 1.2843 | 2.5479 | 99.6 | 20.87 | 37.52 | 105.9614 | 126.5431 | 102.238 | 0.009305 | 0.94048 | 0.021873 | 23.2572 | 2.6329 | 0.15667 | 0.055117 | 3.5941 | 3.5941 | 0.001792
2019-03-21 11:46 | 1553183148.3613 | 35 | 0.19945 | 0.007966 | 0.000113 | 97.9742 | 73.986 | 106.8646 | 21.0397 | 0.003492 | 0.003275 | 179.5121 | 0.26287 | 8.4033 | 0.004328 | 0.005596 | 0.005535 | 6.4e-05 | 0.001022 | 0.30879 | 1.3105 | 2.548 | 99.6 | 20.87 | 37.52 | 105.6756 | 126.078 | 100.4298 | 0.009305 | 0.94503 | 0.021982 | 23.2605 | 2.5119 | 0.15638 | 0.051066 | 3.3476 | 3.3476 | 0.00192
2019-03-21 11:46 | 1553183149.3613 | 36 | 0.19957 | 0.008014 | 0.000111 | 98.1002 | 75.7421 | 107.3319 | 21.0387 | 0.003567 | 0.003276 | 179.4348 | 0.2631 | 8.4032 | 0.004316 | 0.005081 | 0.005023 | 6.4e-05 | 0.001025 | 0.31365 | 1.3234 | 2.548 | 99.6 | 20.87 | 37.521 | 107.4994 | 125.7826 | 100.4445 | 0.009305 | 0.945 | 0.02203 | 23.3118 | 2.3136 | 0.15652 | 0.050258 | 3.302 | 3.302 | 0.001948
2019-03-21 11:46 | 1553183150.3615 | 37 | 0.20031 | 0.008018 | 0.000109 | 98.091 | 75.3092 | 107.7526 | 21.0398 | 0.003535 | 0.003317 | 179.3164 | 0.26332 | 8.4025 | 0.004327 | 0.005973 | 0.005877 | 6.6e-05 | 0.001021 | 0.31404 | 1.3083 | 2.548 | 99.6 | 20.87 | 37.522 | 106.7234 | 126.0448 | 101.4456 | 0.009305 | 0.94247 | 0.022043 | 23.3887 | 2.364 | 0.15665 | 0.045876 | 3.0169 | 3.0169 | 0.002134
2019-03-21 11:46 | 1553183151.3624 | 38 | 0.20105 | 0.007982 | 0.000107 | 98.077 | 75.676 | 108.1868 | 21.039 | 0.003591 | 0.003317 | 179.1976 | 0.26351 | 8.3993 | 0.004268 | 0.006102 | 0.00599 | 6.4e-05 | 0.001023 | 0.33644 | 1.3252 | 2.548 | 99.6 | 20.87 | 37.521 | 108.0936 | 124.6086 | 101.4489 | 0.009305 | 0.94246 | 0.022087 | 23.4359 | 2.3192 | 0.15676 | 0.041519 | 2.7366 | 2.7366 | 0.002354
2019-03-21 11:46 | 1553183152.3643 | 39 | 0.20085 | 0.00791 | 0.000105 | 98.1062 | 76.4754 | 108.5469 | 21.0395 | 0.003572 | 0.003335 | 179.1077 | 0.26283 | 8.4045 | 0.004337 | 0.005534 | 0.005442 | 6.4e-05 | 0.001024 | 0.30381 | 1.315 | 2.5477 | 99.6 | 20.87 | 37.52 | 107.6104 | 126.2856 | 101.8716 | 0.009305 | 0.9414 | 0.022112 | 23.4881 | 2.2279 | 0.15636 | 0.042785 | 2.8231 | 2.8231 | 0.002276
2019-03-21 11:46 | 1553183153.3627 | 40 | 0.20138 | 0.007825 | 0.000101 | 98.0961 | 77.1106 | 108.667 | 21.0399 | 0.003554 | 0.003261 | 179.0126 | 0.26321 | 8.402 | 0.004282 | 0.005091 | 0.00502 | 6.2e-05 | 0.001025 | 0.31022 | 1.3027 | 2.548 | 99.6 | 20.87 | 37.52 | 107.1841 | 124.9356 | 100.0825 | 0.009305 | 0.94591 | 0.022177 | 23.445 | 2.153 | 0.15659 | 0.039722 | 2.6294 | 2.6294 | 0.002448
2019-03-21 11:46 | 1553183154.3631 | 41 | 0.20139 | 0.007714 | 9.8e-05 | 98.0404 | 76.7165 | 108.7654 | 21.0405 | 0.003583 | 0.003312 | 178.9028 | 0.263 | 8.403 | 0.004254 | 0.00625 | 0.006074 | 6.4e-05 | 0.001025 | 0.29574 | 1.3037 | 2.5481 | 99.6 | 20.87 | 37.519 | 107.8918 | 124.2617 | 101.3112 | 0.009304 | 0.94281 | 0.022151 | 23.4942 | 2.1938 | 0.15646 | 0.039783 | 2.6305 | 2.6305 | 0.002445
2019-03-21 11:46 | 1553183155.3626 | 42 | 0.20184 | 0.007594 | 9.4e-05 | 98.2153 | 77.7451 | 108.9717 | 21.0417 | 0.003599 | 0.003319 | 178.8169 | 0.26301 | 8.4037 | 0.004294 | 0.005359 | 0.005362 | 6.1e-05 | 0.001024 | 0.28241 | 1.27 | 2.548 | 99.6 | 20.87 | 37.519 | 108.2798 | 125.2421 | 101.5016 | 0.009304 | 0.94233 | 0.022166 | 23.5224 | 2.0906 | 0.15647 | 0.037274 | 2.4668 | 2.4668 | 0.002607
2019-03-21 11:46 | 1553183156.3616 | 43 | 0.20214 | 0.007471 | 9.1e-05 | 98.1494 | 79.5526 | 109.0168 | 21.0408 | 0.003612 | 0.003368 | 178.7411 | 0.26291 | 8.4034 | 0.004373 | 0.005937 | 0.00579 | 6.3e-05 | 0.001026 | 0.3011 | 1.2761 | 2.5479 | 99.6 | 20.87 | 37.517 | 108.6043 | 127.1636 | 102.668 | 0.009304 | 0.93941 | 0.022136 | 23.5639 | 1.879 | 0.1564 | 0.035636 | 2.3556 | 2.3556 | 0.002729
2019-03-21 11:46 | 1553183157.3621 | 44 | 0.20227 | 0.007341 | 8.6e-05 | 98.2599 | 80.8417 | 109.1478 | 21.0415 | 0.003642 | 0.003323 | 178.6574 | 0.26243 | 8.4025 | 0.004286 | 0.006121 | 0.006041 | 6.1e-05 | 0.001021 | 0.32493 | 1.2947 | 2.5481 | 99.6 | 20.87 | 37.517 | 109.3183 | 125.0477 | 101.5793 | 0.009304 | 0.94213 | 0.022181 | 23.5439 | 1.7453 | 0.15612 | 0.034992 | 2.318 | 2.318 | 0.002768
2019-03-21 11:46 | 1553183158.3619 | 45 | 0.20242 | 0.007211 | 8.3e-05 | 98.2882 | 82.1118 | 109.4177 | 21.041 | 0.003611 | 0.003335 | 178.5666 | 0.26356 | 8.4017 | 0.004311 | 0.005091 | 0.004967 | 6.3e-05 | 0.001025 | 0.34075 | 1.3365 | 2.5478 | 99.6 | 20.87 | 37.516 | 108.5604 | 125.6591 | 101.8662 | 0.009304 | 0.94141 | 0.0222 | 23.582 | 1.6084 | 0.15679 | 0.034281 | 2.2731 | 2.2731 | 0.002835
2019-03-21 11:46 | 1553183159.3623 | 46 | 0.20254 | 0.007081 | 8e-05 | 98.2412 | 82.3028 | 109.7339 | 21.0411 | 0.003644 | 0.003329 | 178.4677 | 0.26318 | 8.4027 | 0.004317 | 0.005552 | 0.005475 | 6.2e-05 | 0.001019 | 0.32012 | 1.324 | 2.5477 | 99.6 | 20.87 | 37.514 | 109.3787 | 125.8096 | 101.7336 | 0.009303 | 0.94175 | 0.022236 | 23.6118 | 1.5834 | 0.15657 | 0.033701 | 2.2385 | 2.2385 | 0.002875
2019-03-21 11:46 | 1553183160.3609 | 47 | 0.20263 | 0.006952 | 7.7e-05 | 98.2087 | 82.5338 | 109.9193 | 21.0415 | 0.003574 | 0.003323 | 178.3892 | 0.26236 | 8.4012 | 0.00437 | 0.005748 | 0.005641 | 6.2e-05 | 0.001021 | 0.30431 | 1.3377 | 2.5479 | 99.6 | 20.87 | 37.513 | 107.6731 | 127.0983 | 101.5762 | 0.009303 | 0.94214 | 0.02226 | 23.6268 | 1.5553 | 0.15608 | 0.033332 | 2.2166 | 2.2166 | 0.002894
2019-03-21 11:46 | 1553183161.361 | 48 | 0.2027 | 0.006827 | 7.4e-05 | 98.2215 | 83.3842 | 109.8346 | 21.0407 | 0.003602 | 0.003305 | 178.316 | 0.26378 | 8.4022 | 0.00437 | 0.005959 | 0.005748 | 6.2e-05 | 0.001023 | 0.31365 | 1.304 | 2.5479 | 99.6 | 20.86 | 37.512 | 108.3498 | 127.1111 | 101.1544 | 0.009297 | 0.9432 | 0.022264 | 23.6044 | 1.4648 | 0.15692 | 0.033063 | 2.1993 | 2.1993 | 0.002933
2019-03-21 11:46 | 1553183162.3614 | 49 | 0.20279 | 0.00671 | 7.2e-05 | 98.2337 | 84.1699 | 110.0005 | 21.0415 | 0.003638 | 0.00331 | 178.2164 | 0.26273 | 8.4027 | 0.004335 | 0.006032 | 0.006 | 6.2e-05 | 0.00102 | 0.29518 | 1.3472 | 2.5479 | 99.6 | 20.86 | 37.511 | 109.2301 | 126.2486 | 101.281 | 0.009297 | 0.94288 | 0.022277 | 23.6262 | 1.382 | 0.1563 | 0.03268 | 2.1752 | 2.1752 | 0.002953
2019-03-21 11:46 | 1553183163.3621 | 50 | 0.20288 | 0.006594 | 7e-05 | 98.2096 | 84.4829 | 110.2775 | 21.0422 | 0.00364 | 0.003296 | 178.1369 | 0.26328 | 8.4043 | 0.00433 | 0.005061 | 0.004971 | 6e-05 | 0.001022 | 0.31404 | 1.3167 | 2.5478 | 99.6 | 20.86 | 37.51 | 109.2782 | 126.1284 | 100.9338 | 0.009296 | 0.94376 | 0.022315 | 23.645 | 1.3466 | 0.15663 | 0.032248 | 2.1503 | 2.1503 | 0.002994
2019-03-21 11:46 | 1553183164.3621 | 51 | 0.20312 | 0.006493 | 6.9e-05 | 98.2948 | 84.2983 | 110.535 | 21.0421 | 0.003588 | 0.003301 | 178.07 | 0.26316 | 8.4023 | 0.004378 | 0.005582 | 0.005531 | 6.2e-05 | 0.001025 | 0.29048 | 1.2895 | 2.5479 | 99.6 | 20.86 | 37.509 | 108.0099 | 127.3044 | 101.0647 | 0.009296 | 0.94343 | 0.022337 | 23.6767 | 1.374 | 0.15655 | 0.030943 | 2.0656 | 2.0656 | 0.003115
2019-03-21 11:46 | 1553183165.3616 | 52 | 0.20316 | 0.006397 | 6.7e-05 | 98.3088 | 83.7582 | 110.5589 | 21.0425 | 0.003625 | 0.003333 | 177.9846 | 0.2627 | 8.4017 | 0.004409 | 0.006171 | 0.006043 | 6e-05 | 0.00102 | 0.3095 | 1.3035 | 2.548 | 99.6 | 20.86 | 37.508 | 108.9031 | 128.0624 | 101.841 | 0.009296 | 0.94148 | 0.022317 | 23.7038 | 1.4327 | 0.15628 | 0.030827 | 2.056 | 2.056 | 0.003124
2019-03-21 11:46 | 1553183166.3626 | 53 | 0.20326 | 0.006302 | 6.6e-05 | 98.3621 | 84.2013 | 110.8286 | 21.0418 | 0.003674 | 0.003291 | 177.9266 | 0.26316 | 8.4009 | 0.004304 | 0.005446 | 0.00535 | 6.2e-05 | 0.001021 | 0.30469 | 1.3045 | 2.5479 | 99.6 | 20.86 | 37.507 | 110.0908 | 125.4887 | 100.8141 | 0.009296 | 0.94406 | 0.022374 | 23.7002 | 1.3904 | 0.15656 | 0.030372 | 2.031 | 2.031 | 0.003168
2019-03-21 11:46 | 1553183167.3612 | 54 | 0.20328 | 0.006208 | 6.5e-05 | 98.3481 | 84.5047 | 111.0514 | 21.0429 | 0.003637 | 0.003312 | 177.846 | 0.26291 | 8.403 | 0.004368 | 0.006277 | 0.006221 | 5.8e-05 | 0.001025 | 0.30917 | 1.3065 | 2.5479 | 99.6 | 20.86 | 37.505 | 109.1953 | 127.0599 | 101.3208 | 0.009295 | 0.94278 | 0.022382 | 23.7401 | 1.3569 | 0.1564 | 0.030337 | 2.0294 | 2.0294 | 0.003168
2019-03-21 11:46 | 1553183168.3626 | 55 | 0.20325 | 0.006116 | 6.4e-05 | 98.4399 | 83.655 | 111.2298 | 21.0424 | 0.003663 | 0.003275 | 177.7634 | 0.26261 | 8.4029 | 0.004439 | 0.006363 | 0.006281 | 6.1e-05 | 0.001021 | 0.3366 | 1.3007 | 2.548 | 99.6 | 20.86 | 37.505 | 109.8296 | 128.802 | 100.4199 | 0.009295 | 0.94506 | 0.022427 | 23.7305 | 1.4557 | 0.15623 | 0.030644 | 2.0541 | 2.0541 | 0.003126
2019-03-21 11:46 | 1553183169.3613 | 56 | 0.20344 | 0.006026 | 6.2e-05 | 98.4718 | 83.623 | 111.4227 | 21.043 | 0.003631 | 0.003219 | 177.6758 | 0.26338 | 8.4013 | 0.004486 | 0.005699 | 0.00564 | 6e-05 | 0.001023 | 0.32892 | 1.2816 | 2.5479 | 99.6 | 20.86 | 37.503 | 109.0613 | 129.9624 | 99.0807 | 0.009295 | 0.94846 | 0.022487 | 23.7085 | 1.462 | 0.15668 | 0.029647 | 1.9927 | 1.9927 | 0.003232
2019-03-21 11:46 | 1553183170.3625 | 57 | 0.20354 | 0.005952 | 6.1e-05 | 98.4095 | 84.1315 | 111.7811 | 21.0435 | 0.00365 | 0.003293 | 177.6003 | 0.26277 | 8.4024 | 0.004465 | 0.005994 | 0.005917 | 5.8e-05 | 0.001026 | 0.32228 | 1.3189 | 2.548 | 99.6 | 20.86 | 37.503 | 109.512 | 129.4372 | 100.8629 | 0.009295 | 0.94394 | 0.022469 | 23.8034 | 1.4021 | 0.15632 | 0.029129 | 1.9565 | 1.9565 | 0.003284
2019-03-21 11:46 | 1553183171.3614 | 58 | 0.20367 | 0.005875 | 5.9e-05 | 98.4271 | 85.2786 | 112.2616 | 21.0424 | 0.003623 | 0.003227 | 177.5195 | 0.2634 | 8.4029 | 0.004547 | 0.005928 | 0.005749 | 5.6e-05 | 0.00102 | 0.30757 | 1.2958 | 2.5479 | 99.6 | 20.86 | 37.503 | 108.8549 | 131.4446 | 99.2674 | 0.009295 | 0.94798 | 0.022565 | 23.8036 | 1.2826 | 0.1567 | 0.028443 | 1.9188 | 1.9188 | 0.003356
2019-03-21 11:46 | 1553183172.3605 | 59 | 0.20396 | 0.005802 | 5.7e-05 | 98.4292 | 86.0292 | 112.6426 | 21.0441 | 0.003616 | 0.003234 | 177.4366 | 0.26277 | 8.4012 | 0.004511 | 0.005466 | 0.005412 | 5.9e-05 | 0.001017 | 0.32421 | 1.2949 | 2.5479 | 99.6 | 20.86 | 37.5 | 108.6928 | 130.5625 | 99.433 | 0.009294 | 0.94756 | 0.022599 | 23.8492 | 1.2044 | 0.15632 | 0.02681 | 1.8115 | 1.8115 | 0.003547
2019-03-21 11:46 | 1553183173.3616 | 60 | 0.20399 | 0.00573 | 5.7e-05 | 98.4897 | 86.8122 | 112.6665 | 21.0427 | 0.003643 | 0.003218 | 177.3725 | 0.26152 | 8.4029 | 0.004448 | 0.005837 | 0.005762 | 6.2e-05 | 0.001013 | 0.27704 | 1.3046 | 2.5479 | 99.6 | 20.86 | 37.498 | 109.3378 | 129.0144 | 99.0545 | 0.009293 | 0.94852 | 0.022612 | 23.8396 | 1.1288 | 0.15558 | 0.026685 | 1.8043 | 1.8043 | 0.003544
2019-03-21 11:46 | 1553183174.362 | 61 | 0.20417 | 0.005646 | 5.6e-05 | 98.5123 | 87.6841 | 112.9698 | 21.043 | 0.003619 | 0.003247 | 177.3136 | 0.2627 | 8.4045 | 0.004369 | 0.005654 | 0.005476 | 6.2e-05 | 0.001009 | 0.30862 | 1.343 | 2.548 | 99.6 | 20.86 | 37.495 | 108.7688 | 127.0694 | 99.7465 | 0.009293 | 0.94676 | 0.022622 | 23.8939 | 1.0415 | 0.15628 | 0.025703 | 1.7386 | 1.7386 | 0.003694
2019-03-21 11:46 | 1553183175.3606 | 62 | 0.20436 | 0.005565 | 5.5e-05 | 98.574 | 88.4624 | 113.3183 | 21.0439 | 0.003572 | 0.0032 | 177.2312 | 0.26288 | 8.4024 | 0.00446 | 0.00563 | 0.005544 | 6.2e-05 | 0.001004 | 0.30342 | 1.3398 | 2.5481 | 99.6 | 20.86 | 37.493 | 107.6266 | 129.3051 | 98.6197 | 0.009292 | 0.94964 | 0.022691 | 23.8945 | 0.96807 | 0.15639 | 0.024647 | 1.6725 | 1.6725 | 0.003843
2019-03-21 11:46 | 1553183176.3607 | 63 | 0.20454 | 0.005475 | 5.4e-05 | 98.6294 | 88.8248 | 113.5908 | 21.0435 | 0.003605 | 0.003214 | 177.1429 | 0.2628 | 8.4035 | 0.004419 | 0.005565 | 0.005519 | 5.9e-05 | 0.00101 | 0.30071 | 1.3068 | 2.5479 | 99.6 | 20.86 | 37.491 | 108.4282 | 128.3186 | 98.9446 | 0.009292 | 0.94881 | 0.022708 | 23.9337 | 0.93653 | 0.15634 | 0.023713 | 1.6104 | 1.6104 | 0.00399
2019-03-21 11:46 | 1553183177.3624 | 64 | 0.20477 | 0.005379 | 5.3e-05 | 98.7221 | 88.9566 | 113.9915 | 21.0435 | 0.003539 | 0.003151 | 177.0803 | 0.26354 | 8.4009 | 0.004519 | 0.005733 | 0.005634 | 5.9e-05 | 0.001005 | 0.30558 | 1.324 | 2.5481 | 99.6 | 20.86 | 37.487 | 106.8127 | 130.7767 | 97.4386 | 0.009291 | 0.95266 | 0.022795 | 23.9273 | 0.93166 | 0.15678 | 0.022443 | 1.53 | 1.53 | 0.004212
2019-03-21 11:46 | 1553183178.3623 | 65 | 0.20494 | 0.005285 | 5.2e-05 | 98.7572 | 89.3687 | 114.5398 | 21.0428 | 0.003559 | 0.003146 | 177.0307 | 0.26133 | 8.4034 | 0.004458 | 0.006403 | 0.006339 | 5.8e-05 | 0.001012 | 0.31852 | 1.272 | 2.5479 | 99.6 | 20.86 | 37.486 | 107.3175 | 129.2546 | 97.3024 | 0.00929 | 0.95301 | 0.022854 | 23.9804 | 0.89351 | 0.15547 | 0.021516 | 1.4707 | 1.4707 | 0.004345
2019-03-21 11:46 | 1553183179.3619 | 66 | 0.20526 | 0.005181 | 5e-05 | 98.7572 | 89.8629 | 114.9837 | 21.0434 | 0.003486 | 0.003128 | 176.9804 | 0.26348 | 8.4034 | 0.004407 | 0.005836 | 0.005746 | 6.3e-05 | 0.001013 | 0.31094 | 1.3097 | 2.548 | 99.6 | 20.86 | 37.481 | 105.5424 | 128.0021 | 96.8774 | 0.009289 | 0.95411 | 0.022911 | 24.013 | 0.84418 | 0.15674 | 0.019736 | 1.3525 | 1.3525 | 0.004763
2019-03-21 11:46 | 1553183180.3613 | 67 | 0.20547 | 0.005077 | 4.7e-05 | 98.8234 | 90.3295 | 115.3945 | 21.0432 | 0.003509 | 0.003079 | 176.9437 | 0.26334 | 8.3997 | 0.004397 | 0.005432 | 0.005316 | 5.9e-05 | 0.001014 | 0.30143 | 1.3184 | 2.548 | 99.6 | 20.86 | 37.48 | 106.0923 | 127.7723 | 95.6924 | 0.009289 | 0.95717 | 0.022989 | 24.0173 | 0.80385 | 0.15666 | 0.018616 | 1.2803 | 1.2803 | 0.005029
2019-03-21 11:46 | 1553183181.3624 | 68 | 0.20568 | 0.004969 | 4.6e-05 | 98.8654 | 90.7811 | 115.6001 | 21.044 | 0.003517 | 0.003137 | 176.9103 | 0.26259 | 8.4009 | 0.004371 | 0.005591 | 0.005565 | 6.3e-05 | 0.001012 | 0.31852 | 1.3201 | 2.548 | 99.6 | 20.86 | 37.48 | 106.292 | 127.1263 | 97.0937 | 0.009289 | 0.95355 | 0.022966 | 24.0843 | 0.76304 | 0.15621 | 0.017494 | 1.2019 | 1.2019 | 0.005342
2019-03-21 11:46 | 1553183182.3611 | 69 | 0.20587 | 0.004858 | 4.5e-05 | 98.8269 | 91.1479 | 115.6907 | 21.0443 | 0.003513 | 0.003062 | 176.8788 | 0.26332 | 8.4045 | 0.004346 | 0.006572 | 0.006404 | 6.5e-05 | 0.001004 | 0.28031 | 1.3209 | 2.548 | 99.6 | 20.86 | 37.481 | 106.1948 | 126.5069 | 95.2846 | 0.009289 | 0.95823 | 0.023031 | 24.0348 | 0.72349 | 0.15665 | 0.016501 | 1.137 | 1.137 | 0.005663
2019-03-21 11:46 | 1553183183.3625 | 70 | 0.20604 | 0.004745 | 4.3e-05 | 98.8821 | 91.8318 | 115.8479 | 21.0442 | 0.003509 | 0.003038 | 176.8103 | 0.26351 | 8.4019 | 0.004384 | 0.005681 | 0.005679 | 6.2e-05 | 0.001005 | 0.31868 | 1.3055 | 2.5479 | 99.6 | 20.86 | 37.483 | 106.1004 | 127.4428 | 94.7107 | 0.00929 | 0.95973 | 0.023065 | 24.0324 | 0.66162 | 0.15676 | 0.015624 | 1.0782 | 1.0782 | 0.005976
2019-03-21 11:46 | 1553183184.3622 | 71 | 0.20612 | 0.004631 | 4.1e-05 | 98.925 | 91.9394 | 115.9022 | 21.0439 | 0.003486 | 0.003051 | 176.7634 | 0.26331 | 8.4018 | 0.004291 | 0.005748 | 0.005633 | 6.4e-05 | 0.001002 | 0.2995 | 1.2778 | 2.548 | 99.6 | 20.86 | 37.485 | 105.5381 | 125.1646 | 95.0257 | 0.00929 | 0.95891 | 0.02306 | 24.0483 | 0.65502 | 0.15665 | 0.015267 | 1.0534 | 1.0534 | 0.006112
2019-03-21 11:46 | 1553183185.3622 | 72 | 0.2062 | 0.004513 | 4e-05 | 98.9407 | 92.5409 | 116.1611 | 21.0447 | 0.003393 | 0.003021 | 176.7097 | 0.26298 | 8.402 | 0.004236 | 0.005208 | 0.005055 | 6.1e-05 | 0.001 | 0.30541 | 1.2823 | 2.5481 | 99.6 | 20.86 | 37.49 | 103.2845 | 123.8078 | 94.2931 | 0.009291 | 0.96082 | 0.023109 | 24.0512 | 0.59812 | 0.15644 | 0.014945 | 1.0335 | 1.0335 | 0.006222
2019-03-21 11:46 | 1553183186.3626 | 73 | 0.20625 | 0.004403 | 3.9e-05 | 98.9568 | 92.5985 | 116.6094 | 21.0451 | 0.003424 | 0.003075 | 176.6683 | 0.26333 | 8.4036 | 0.004266 | 0.005305 | 0.005249 | 6.1e-05 | 0.001001 | 0.31078 | 1.2893 | 2.548 | 99.6 | 20.86 | 37.493 | 104.041 | 124.5636 | 95.5917 | 0.009292 | 0.95743 | 0.023113 | 24.1401 | 0.59402 | 0.15665 | 0.014788 | 1.0228 | 1.0228 | 0.006295
2019-03-21 11:46 | 1553183187.3604 | 74 | 0.20641 | 0.004292 | 3.7e-05 | 99.0772 | 92.6996 | 116.9014 | 21.0452 | 0.00344 | 0.003077 | 176.6433 | 0.2627 | 8.4036 | 0.004224 | 0.005689 | 0.005632 | 6.2e-05 | 0.001006 | 0.33555 | 1.3394 | 2.5478 | 99.59 | 20.85 | 37.5 | 104.4334 | 123.5382 | 95.6543 | 0.009289 | 0.95727 | 0.023139 | 24.1724 | 0.59513 | 0.15628 | 0.013944 | 0.96563 | 0.96563 | 0.006652
2019-03-21 11:46 | 1553183188.3605 | 75 | 0.20642 | 0.004186 | 3.6e-05 | 99.1013 | 92.7691 | 116.9483 | 21.0453 | 0.003383 | 0.002969 | 176.6141 | 0.26271 | 8.4025 | 0.004198 | 0.005891 | 0.005738 | 6.4e-05 | 0.001004 | 0.30021 | 1.3464 | 2.548 | 99.59 | 20.85 | 37.505 | 103.0477 | 122.881 | 93.0362 | 0.00929 | 0.96411 | 0.023227 | 24.0912 | 0.5906 | 0.15629 | 0.014043 | 0.9762 | 0.9762 | 0.00658
2019-03-21 11:46 | 1553183189.3619 | 76 | 0.20644 | 0.004086 | 3.6e-05 | 99.2247 | 93.6035 | 116.8359 | 21.044 | 0.003383 | 0.002994 | 176.5641 | 0.26262 | 8.402 | 0.004145 | 0.0059 | 0.005806 | 6.1e-05 | 0.001005 | 0.33467 | 1.3371 | 2.5479 | 99.59 | 20.85 | 37.509 | 103.0433 | 121.5811 | 93.654 | 0.009291 | 0.96249 | 0.023196 | 24.1 | 0.52163 | 0.15624 | 0.014005 | 0.97232 | 0.97232 | 0.006604
2019-03-21 11:46 | 1553183190.3605 | 77 | 0.20646 | 0.003991 | 3.6e-05 | 99.1395 | 93.8546 | 116.462 | 21.0455 | 0.003354 | 0.002962 | 176.5124 | 0.26262 | 8.4027 | 0.004106 | 0.006273 | 0.006189 | 6.3e-05 | 0.001001 | 0.31686 | 1.3524 | 2.5479 | 99.59 | 20.85 | 37.516 | 102.3395 | 120.6412 | 92.8762 | 0.009293 | 0.96454 | 0.023183 | 24.0358 | 0.48999 | 0.15624 | 0.014039 | 0.9741 | 0.9741 | 0.006592
2019-03-21 11:46 | 1553183191.3605 | 78 | 0.20653 | 0.003895 | 3.6e-05 | 99.2113 | 94.0105 | 116.3772 | 21.0455 | 0.003335 | 0.002955 | 176.449 | 0.26236 | 8.4043 | 0.004028 | 0.006235 | 0.006154 | 6.2e-05 | 0.001005 | 0.30807 | 1.3079 | 2.5479 | 99.59 | 20.86 | 37.521 | 101.8761 | 118.7204 | 92.704 | 0.0093 | 0.96499 | 0.02318 | 24.0214 | 0.48162 | 0.15608 | 0.013755 | 0.95426 | 0.95426 | 0.006722
2019-03-21 11:47 | 1553183192.362 | 79 | 0.20661 | 0.003802 | 3.5e-05 | 99.2515 | 93.9717 | 116.3838 | 21.0454 | 0.00332 | 0.002921 | 176.4025 | 0.26303 | 8.4043 | 0.004115 | 0.005976 | 0.005945 | 6e-05 | 0.001001 | 0.29911 | 1.3134 | 2.5477 | 99.59 | 20.85 | 37.528 | 101.5258 | 120.8498 | 91.8774 | 0.009296 | 0.96718 | 0.023207 | 23.995 | 0.48894 | 0.15648 | 0.013358 | 0.9278 | 0.9278 | 0.006932
2019-03-21 11:47 | 1553183193.3617 | 80 | 0.20663 | 0.003721 | 3.4e-05 | 99.3293 | 93.9905 | 116.5399 | 21.0453 | 0.003349 | 0.002996 | 176.3424 | 0.26316 | 8.4036 | 0.004124 | 0.005707 | 0.00558 | 6.1e-05 | 0.001 | 0.29407 | 1.3292 | 2.548 | 99.59 | 20.85 | 37.531 | 102.2325 | 121.0823 | 93.6958 | 0.009297 | 0.96238 | 0.023165 | 24.0708 | 0.49416 | 0.15655 | 0.013323 | 0.92376 | 0.92376 | 0.006965
2019-03-21 11:47 | 1553183194.3617 | 81 | 0.20668 | 0.003647 | 3.4e-05 | 99.368 | 94.1495 | 116.6969 | 21.0453 | 0.003368 | 0.002934 | 176.2963 | 0.26348 | 8.3998 | 0.004059 | 0.005984 | 0.005932 | 6.3e-05 | 0.000998 | 0.33019 | 1.3655 | 2.5479 | 99.59 | 20.86 | 37.534 | 102.6762 | 119.4906 | 92.2089 | 0.009303 | 0.9663 | 0.023228 | 24.0381 | 0.48252 | 0.15674 | 0.01315 | 0.91421 | 0.91421 | 0.007047
2019-03-21 11:47 | 1553183195.3621 | 82 | 0.20669 | 0.003578 | 3.4e-05 | 99.337 | 94.329 | 116.6415 | 21.0455 | 0.003358 | 0.002945 | 176.2747 | 0.26351 | 8.4017 | 0.004054 | 0.005455 | 0.005369 | 6.5e-05 | 0.000998 | 0.3418 | 1.2957 | 2.548 | 99.59 | 20.85 | 37.539 | 102.4447 | 119.3569 | 92.4576 | 0.009299 | 0.96564 | 0.023215 | 24.0406 | 0.4627 | 0.15676 | 0.013137 | 0.91285 | 0.91285 | 0.007058
2019-03-21 11:47 | 1553183196.3607 | 83 | 0.20672 | 0.003506 | 3.3e-05 | 99.4976 | 94.1716 | 116.2402 | 21.0457 | 0.003312 | 0.002877 | 176.2469 | 0.26282 | 8.4029 | 0.004075 | 0.005289 | 0.005218 | 6.6e-05 | 0.000995 | 0.30773 | 1.3123 | 2.548 | 99.59 | 20.85 | 37.541 | 101.3298 | 119.8758 | 90.8136 | 0.009299 | 0.97 | 0.023227 | 23.9452 | 0.49208 | 0.15635 | 0.013056 | 0.90767 | 0.90767 | 0.00708
2019-03-21 11:47 | 1553183197.3619 | 84 | 0.20674 | 0.00344 | 3.3e-05 | 99.421 | 94.1277 | 116.0918 | 21.0456 | 0.003314 | 0.002887 | 176.1975 | 0.26289 | 8.4028 | 0.004028 | 0.005746 | 0.005586 | 6.2e-05 | 0.000994 | 0.3366 | 1.28 | 2.548 | 99.59 | 20.86 | 37.543 | 101.3627 | 118.7332 | 91.0684 | 0.009306 | 0.96932 | 0.023204 | 23.9383 | 0.48937 | 0.15639 | 0.013059 | 0.90705 | 0.90705 | 0.007086
2019-03-21 11:47 | 1553183198.3619 | 85 | 0.20676 | 0.003379 | 3.3e-05 | 99.5277 | 94.6379 | 116.0221 | 21.0459 | 0.003298 | 0.002908 | 176.1379 | 0.26222 | 8.4036 | 0.004008 | 0.005589 | 0.005404 | 6.6e-05 | 0.000993 | 0.30093 | 1.293 | 2.5479 | 99.59 | 20.85 | 37.55 | 100.9877 | 118.2406 | 91.5713 | 0.009301 | 0.96799 | 0.023181 | 23.9476 | 0.45061 | 0.15599 | 0.013015 | 0.90308 | 0.90308 | 0.007099
2019-03-21 11:47 | 1553183199.362 | 86 | 0.2067 | 0.003325 | 3.4e-05 | 99.4675 | 94.4223 | 116.2481 | 21.0456 | 0.003282 | 0.00284 | 176.0916 | 0.2621 | 8.402 | 0.004036 | 0.005655 | 0.005625 | 6.3e-05 | 0.00099 | 0.32532 | 1.2979 | 2.548 | 99.59 | 20.85 | 37.551 | 100.5957 | 118.9281 | 89.9273 | 0.009302 | 0.97237 | 0.023256 | 23.9168 | 0.4656 | 0.15593 | 0.013403 | 0.93294 | 0.93294 | 0.006869
2019-03-21 11:47 | 1553183200.361 | 87 | 0.20671 | 0.003271 | 3.5e-05 | 99.4994 | 94.7905 | 116.2625 | 21.0463 | 0.003246 | 0.002926 | 176.0616 | 0.26344 | 8.4039 | 0.004028 | 0.005089 | 0.00501 | 6.4e-05 | 0.00099 | 0.32333 | 1.2764 | 2.5479 | 99.59 | 20.85 | 37.554 | 99.7294 | 118.7378 | 91.9946 | 0.009302 | 0.96686 | 0.023192 | 23.9863 | 0.43365 | 0.15672 | 0.013419 | 0.93145 | 0.93145 | 0.006915
2019-03-21 11:47 | 1553183201.362 | 88 | 0.20676 | 0.003221 | 3.6e-05 | 99.5378 | 94.6924 | 116.3291 | 21.0467 | 0.003241 | 0.002879 | 176.0524 | 0.26238 | 8.402 | 0.004032 | 0.004995 | 0.004912 | 6.5e-05 | 0.00099 | 0.29286 | 1.2966 | 2.5479 | 99.59 | 20.85 | 37.556 | 99.5981 | 118.8184 | 90.8724 | 0.009303 | 0.96984 | 0.023234 | 23.9563 | 0.44637 | 0.15609 | 0.013168 | 0.91558 | 0.91558 | 0.007007
2019-03-21 11:47 | 1553183202.3613 | 89 | 0.20675 | 0.003175 | 3.6e-05 | 99.5396 | 94.6665 | 116.0719 | 21.0459 | 0.003282 | 0.002906 | 176.0218 | 0.26237 | 8.4021 | 0.004019 | 0.005063 | 0.004957 | 6.4e-05 | 0.000989 | 0.35314 | 1.2946 | 2.548 | 99.59 | 20.85 | 37.557 | 100.5882 | 118.506 | 91.5166 | 0.009303 | 0.96813 | 0.023188 | 23.9509 | 0.44897 | 0.15608 | 0.013323 | 0.92453 | 0.92453 | 0.006939
2019-03-21 11:47 | 1553183203.3622 | 90 | 0.20673 | 0.003135 | 3.7e-05 | 99.5709 | 94.4768 | 115.9456 | 21.0462 | 0.003202 | 0.00285 | 176.0003 | 0.26227 | 8.4028 | 0.003935 | 0.006029 | 0.005886 | 6.3e-05 | 0.000989 | 0.31404 | 1.3084 | 2.5479 | 99.59 | 20.85 | 37.559 | 98.6649 | 116.4684 | 90.1841 | 0.009304 | 0.97168 | 0.023218 | 23.8941 | 0.46973 | 0.15603 | 0.013488 | 0.93719 | 0.93719 | 0.006842
2019-03-21 11:47 | 1553183204.3613 | 91 | 0.20677 | 0.003099 | 3.8e-05 | 99.6519 | 94.6083 | 115.9578 | 21.0466 | 0.003288 | 0.002872 | 175.9536 | 0.26201 | 8.4027 | 0.004005 | 0.005241 | 0.005169 | 6.6e-05 | 0.000989 | 0.33301 | 1.3006 | 2.548 | 99.59 | 20.85 | 37.56 | 100.7544 | 118.1707 | 90.712 | 0.009304 | 0.97027 | 0.023202 | 23.9128 | 0.46457 | 0.15587 | 0.013268 | 0.9212 | 0.9212 | 0.006954
2019-03-21 11:47 | 1553183205.3612 | 92 | 0.20679 | 0.003066 | 3.9e-05 | 99.7002 | 94.6403 | 116.1848 | 21.0458 | 0.003228 | 0.002877 | 175.9071 | 0.26243 | 8.4022 | 0.004039 | 0.005556 | 0.005522 | 5.9e-05 | 0.000991 | 0.28584 | 1.3039 | 2.5479 | 99.59 | 20.85 | 37.561 | 99.2952 | 118.9997 | 90.8137 | 0.009304 | 0.97 | 0.023221 | 23.9395 | 0.46587 | 0.15612 | 0.013175 | 0.91541 | 0.91541 | 0.007009
2019-03-21 11:47 | 1553183206.3614 | 93 | 0.20674 | 0.003032 | 4.1e-05 | 99.6656 | 94.7207 | 116.3265 | 21.0466 | 0.003222 | 0.002832 | 175.8396 | 0.26237 | 8.4029 | 0.004018 | 0.004796 | 0.004683 | 6.3e-05 | 0.000986 | 0.29856 | 1.3103 | 2.5481 | 99.59 | 20.85 | 37.56 | 99.1453 | 118.4962 | 89.743 | 0.009304 | 0.97286 | 0.02327 | 23.9188 | 0.45518 | 0.15608 | 0.013525 | 0.94165 | 0.94165 | 0.006813
2019-03-21 11:47 | 1553183207.3609 | 94 | 0.20679 | 0.002999 | 4.1e-05 | 99.7071 | 94.6866 | 116.3157 | 21.0463 | 0.003267 | 0.002794 | 175.7697 | 0.26328 | 8.4024 | 0.00395 | 0.005221 | 0.005187 | 6e-05 | 0.000987 | 0.30646 | 1.3086 | 2.548 | 99.59 | 20.85 | 37.559 | 100.2369 | 116.823 | 88.8286 | 0.009304 | 0.97532 | 0.023298 | 23.8876 | 0.46211 | 0.15663 | 0.013293 | 0.92655 | 0.92655 | 0.006948
2019-03-21 11:47 | 1553183208.3615 | 95 | 0.20675 | 0.00297 | 4.1e-05 | 99.7732 | 94.7581 | 116.3585 | 21.0475 | 0.00321 | 0.002852 | 175.7406 | 0.26277 | 8.4009 | 0.003868 | 0.005253 | 0.005134 | 6.3e-05 | 0.000996 | 0.32388 | 1.3321 | 2.548 | 99.59 | 20.85 | 37.557 | 98.8716 | 114.8238 | 90.2152 | 0.009303 | 0.9716 | 0.023258 | 23.9377 | 0.46129 | 0.15632 | 0.01354 | 0.94214 | 0.94214 | 0.006819
2019-03-21 11:47 | 1553183209.3615 | 96 | 0.20676 | 0.002949 | 4.1e-05 | 99.8811 | 94.8399 | 116.1646 | 21.0472 | 0.003185 | 0.002821 | 175.6929 | 0.26249 | 8.4036 | 0.003916 | 0.00518 | 0.005193 | 6.4e-05 | 0.000993 | 0.27151 | 1.3038 | 2.548 | 99.59 | 20.85 | 37.554 | 98.267 | 115.9952 | 89.4679 | 0.009302 | 0.9736 | 0.023262 | 23.8931 | 0.46324 | 0.15616 | 0.013521 | 0.94106 | 0.94106 | 0.00682
2019-03-21 11:47 | 1553183210.3605 | 97 | 0.20678 | 0.002929 | 4.1e-05 | 99.8998 | 94.9458 | 116.4624 | 21.0467 | 0.003219 | 0.002862 | 175.6242 | 0.2628 | 8.4027 | 0.003962 | 0.005431 | 0.005277 | 6.5e-05 | 0.000998 | 0.29789 | 1.2762 | 2.5479 | 99.59 | 20.85 | 37.55 | 99.0739 | 117.1155 | 90.4676 | 0.009301 | 0.97093 | 0.02326 | 23.9567 | 0.45493 | 0.15634 | 0.013404 | 0.93285 | 0.93285 | 0.006888
2019-03-21 11:47 | 1553183211.3605 | 98 | 0.20677 | 0.002906 | 4e-05 | 99.9073 | 94.8167 | 116.4427 | 21.0478 | 0.003147 | 0.002834 | 175.5554 | 0.26276 | 8.4052 | 0.003805 | 0.005657 | 0.0056 | 6.2e-05 | 0.000996 | 0.2995 | 1.2656 | 2.5479 | 99.59 | 20.86 | 37.549 | 97.3386 | 113.2994 | 89.7786 | 0.009307 | 0.97277 | 0.02328 | 23.9319 | 0.46777 | 0.15632 | 0.013532 | 0.94254 | 0.94254 | 0.006816
2019-03-21 11:47 | 1553183212.3607 | 99 | 0.20676 | 0.00288 | 3.9e-05 | 99.9558 | 94.8263 | 116.8385 | 21.0471 | 0.003181 | 0.002749 | 175.5094 | 0.26284 | 8.4021 | 0.003809 | 0.005264 | 0.005181 | 6.1e-05 | 0.000995 | 0.32372 | 1.3062 | 2.5479 | 99.59 | 20.86 | 37.544 | 98.1704 | 113.3998 | 87.736 | 0.009306 | 0.97827 | 0.023386 | 23.905 | 0.47122 | 0.15637 | 0.013599 | 0.95157 | 0.95157 | 0.006754
2019-03-21 11:47 | 1553183213.361 | 100 | 0.20679 | 0.002867 | 4e-05 | 99.9645 | 95.0149 | 117.1333 | 21.0473 | 0.00317 | 0.002782 | 175.4641 | 0.26229 | 8.402 | 0.003842 | 0.005596 | 0.005574 | 6.5e-05 | 0.000996 | 0.3011 | 1.3329 | 2.548 | 99.59 | 20.86 | 37.54 | 97.8937 | 114.1983 | 88.5397 | 0.009305 | 0.9761 | 0.023389 | 23.9618 | 0.45421 | 0.15604 | 0.013454 | 0.94154 | 0.94154 | 0.006811
2019-03-21 11:47 | 1553183214.3617 | 101 | 0.20676 | 0.00285 | 3.9e-05 | 100.0041 | 95.1522 | 117.1281 | 21.0471 | 0.003171 | 0.002802 | 175.4264 | 0.26298 | 8.4018 | 0.00374 | 0.004958 | 0.004872 | 6.6e-05 | 0.001002 | 0.30143 | 1.3295 | 2.5478 | 99.59 | 20.86 | 37.538 | 97.9274 | 111.716 | 89.0069 | 0.009304 | 0.97484 | 0.023374 | 23.9767 | 0.44484 | 0.15645 | 0.013644 | 0.95425 | 0.95425 | 0.006738
2019-03-21 11:47 | 1553183215.3614 | 102 | 0.20679 | 0.002835 | 3.9e-05 | 100.0351 | 95.2363 | 116.8845 | 21.048 | 0.003176 | 0.002789 | 175.3963 | 0.26355 | 8.4017 | 0.003781 | 0.004774 | 0.004701 | 6.6e-05 | 0.001001 | 0.32709 | 1.2906 | 2.5479 | 99.59 | 20.86 | 37.532 | 98.0324 | 112.7014 | 88.7172 | 0.009303 | 0.97562 | 0.023359 | 23.9422 | 0.43971 | 0.15679 | 0.013496 | 0.94334 | 0.94334 | 0.006831
2019-03-21 11:47 | 1553183216.3611 | 103 | 0.20675 | 0.002821 | 3.8e-05 | 100.1415 | 95.442 | 116.876 | 21.047 | 0.003191 | 0.002843 | 175.3631 | 0.26307 | 8.4025 | 0.003737 | 0.005276 | 0.005213 | 6.4e-05 | 0.001002 | 0.33892 | 1.3025 | 2.5479 | 99.59 | 20.86 | 37.53 | 98.3953 | 111.6322 | 90.0084 | 0.009302 | 0.97215 | 0.023316 | 23.984 | 0.42992 | 0.1565 | 0.013716 | 0.95698 | 0.95698 | 0.006721
2019-03-21 11:47 | 1553183217.3619 | 104 | 0.20679 | 0.002809 | 3.9e-05 | 100.1802 | 95.5272 | 116.7594 | 21.0476 | 0.003122 | 0.002828 | 175.3261 | 0.26229 | 8.4025 | 0.003725 | 0.00423 | 0.004174 | 6.5e-05 | 0.001001 | 0.31382 | 1.3267 | 2.5479 | 99.59 | 20.86 | 37.523 | 96.7241 | 111.3518 | 89.6479 | 0.009301 | 0.97312 | 0.023316 | 23.9602 | 0.4254 | 0.15604 | 0.013537 | 0.94442 | 0.94442 | 0.006791
2019-03-21 11:47 | 1553183218.362 | 105 | 0.2068 | 0.00279 | 3.9e-05 | 100.1072 | 95.4441 | 116.6659 | 21.0479 | 0.003171 | 0.002745 | 175.2748 | 0.26286 | 8.4017 | 0.003658 | 0.005131 | 0.005004 | 6.4e-05 | 0.001 | 0.32261 | 1.3156 | 2.548 | 99.59 | 20.86 | 37.519 | 97.918 | 109.7275 | 87.6586 | 0.0093 | 0.97848 | 0.023371 | 23.8848 | 0.42666 | 0.15638 | 0.013517 | 0.94532 | 0.94532 | 0.006799
2019-03-21 11:47 | 1553183219.3599 | 106 | 0.20678 | 0.002789 | 4e-05 | 100.1963 | 95.6522 | 116.8296 | 21.0485 | 0.00316 | 0.002763 | 175.2502 | 0.26294 | 8.4031 | 0.003698 | 0.005082 | 0.005007 | 6.3e-05 | 0.000998 | 0.30685 | 1.3565 | 2.548 | 99.59 | 20.86 | 37.516 | 97.6613 | 110.7016 | 88.0795 | 0.009299 | 0.97734 | 0.023374 | 23.9155 | 0.41514 | 0.15642 | 0.013621 | 0.95262 | 0.95262 | 0.006749
2019-03-21 11:47 | 1553183220.3613 | 107 | 0.20678 | 0.002773 | 3.9e-05 | 100.2454 | 95.5681 | 116.8063 | 21.0486 | 0.003118 | 0.002754 | 175.2267 | 0.26243 | 8.4014 | 0.003638 | 0.005127 | 0.005088 | 6.8e-05 | 0.000999 | 0.31222 | 1.3424 | 2.548 | 99.59 | 20.86 | 37.508 | 96.6326 | 109.2441 | 87.8624 | 0.009297 | 0.97793 | 0.023378 | 23.9059 | 0.4274 | 0.15612 | 0.013639 | 0.95408 | 0.95408 | 0.006725
2019-03-21 11:47 | 1553183221.3611 | 108 | 0.20675 | 0.002766 | 3.9e-05 | 100.2821 | 95.801 | 117.0629 | 21.0487 | 0.003126 | 0.002715 | 175.1895 | 0.26209 | 8.4029 | 0.003578 | 0.005262 | 0.005213 | 6.3e-05 | 0.001 | 0.32964 | 1.3596 | 2.5479 | 99.59 | 20.86 | 37.505 | 96.8239 | 107.784 | 86.9382 | 0.009296 | 0.98044 | 0.023434 | 23.9015 | 0.40889 | 0.15592 | 0.013818 | 0.96893 | 0.96893 | 0.006614
2019-03-21 11:47 | 1553183222.3618 | 109 | 0.20683 | 0.002755 | 3.9e-05 | 100.3518 | 96.0394 | 117.0737 | 21.05 | 0.003085 | 0.002751 | 175.156 | 0.26227 | 8.4033 | 0.003573 | 0.005588 | 0.005456 | 6.3e-05 | 0.000995 | 0.2959 | 1.3306 | 2.5479 | 99.59 | 20.86 | 37.497 | 95.8344 | 107.6642 | 87.7974 | 0.009294 | 0.97811 | 0.023407 | 23.9311 | 0.39287 | 0.15602 | 0.013323 | 0.93313 | 0.93313 | 0.006872
2019-03-21 11:47 | 1553183223.3606 | 110 | 0.20681 | 0.002745 | 3.8e-05 | 100.3262 | 96.3024 | 116.804 | 21.0491 | 0.003088 | 0.002794 | 175.0987 | 0.26261 | 8.4055 | 0.003595 | 0.005714 | 0.005625 | 6.7e-05 | 0.001 | 0.34147 | 1.2989 | 2.5478 | 99.59 | 20.86 | 37.493 | 95.925 | 108.1887 | 88.8165 | 0.009293 | 0.97535 | 0.023347 | 23.9372 | 0.36613 | 0.15623 | 0.013486 | 0.94221 | 0.94221 | 0.006815
2019-03-21 11:47 | 1553183224.3603 | 111 | 0.20684 | 0.002732 | 3.7e-05 | 100.339 | 96.3419 | 116.7246 | 21.0492 | 0.003144 | 0.002724 | 175.0449 | 0.26225 | 8.4033 | 0.003641 | 0.004973 | 0.004964 | 6.5e-05 | 0.000997 | 0.33019 | 1.3256 | 2.5479 | 99.59 | 20.86 | 37.49 | 97.264 | 109.2953 | 87.1369 | 0.009292 | 0.9799 | 0.023394 | 23.8735 | 0.36361 | 0.15601 | 0.013305 | 0.93144 | 0.93144 | 0.006884
2019-03-21 11:47 | 1553183225.3615 | 112 | 0.20681 | 0.00273 | 3.7e-05 | 100.4224 | 96.4832 | 116.8216 | 21.0488 | 0.003069 | 0.002726 | 174.9836 | 0.26212 | 8.4022 | 0.003576 | 0.005616 | 0.005486 | 6.6e-05 | 0.000996 | 0.31116 | 1.3004 | 2.5479 | 99.59 | 20.86 | 37.484 | 95.4481 | 107.7291 | 87.1943 | 0.009291 | 0.97974 | 0.023402 | 23.8854 | 0.35793 | 0.15594 | 0.013501 | 0.94556 | 0.94556 | 0.006778
2019-03-21 11:47 | 1553183226.3618 | 113 | 0.20688 | 0.002718 | 3.7e-05 | 100.4218 | 96.6525 | 116.6598 | 21.0492 | 0.003046 | 0.002715 | 174.9531 | 0.26306 | 8.4024 | 0.003632 | 0.004743 | 0.004628 | 6.8e-05 | 0.000997 | 0.28136 | 1.3046 | 2.5478 | 99.59 | 20.86 | 37.481 | 94.898 | 109.0902 | 86.922 | 0.00929 | 0.98049 | 0.023394 | 23.8598 | 0.3422 | 0.15649 | 0.013076 | 0.91546 | 0.91546 | 0.007026
2019-03-21 11:47 | 1553183227.3615 | 114 | 0.20692 | 0.002705 | 3.7e-05 | 100.5249 | 96.7533 | 116.7422 | 21.0496 | 0.003099 | 0.002736 | 174.9022 | 0.26167 | 8.4012 | 0.003561 | 0.00555 | 0.005428 | 6.5e-05 | 0.000992 | 0.29894 | 1.3119 | 2.5478 | 99.59 | 20.86 | 37.476 | 96.1746 | 107.3739 | 87.4362 | 0.009289 | 0.97909 | 0.023386 | 23.8853 | 0.34205 | 0.15567 | 0.01286 | 0.89999 | 0.89999 | 0.007109
2019-03-21 11:47 | 1553183228.3605 | 115 | 0.20698 | 0.002698 | 3.6e-05 | 100.5106 | 96.804 | 116.8941 | 21.0503 | 0.003071 | 0.002719 | 174.8592 | 0.26225 | 8.4025 | 0.00369 | 0.005225 | 0.005158 | 6.6e-05 | 0.000998 | 0.31974 | 1.284 | 2.5478 | 99.59 | 20.87 | 37.469 | 95.5049 | 110.4917 | 87.0285 | 0.009293 | 0.9802 | 0.023414 | 23.8873 | 0.33609 | 0.15601 | 0.012543 | 0.8789 | 0.8789 | 0.007295
2019-03-21 11:47 | 1553183229.3619 | 116 | 0.20695 | 0.00269 | 3.6e-05 | 100.5058 | 97.085 | 117.0427 | 21.0503 | 0.003049 | 0.002714 | 174.8198 | 0.26181 | 8.4003 | 0.00362 | 0.00508 | 0.005052 | 6.7e-05 | 0.000997 | 0.31653 | 1.2999 | 2.5479 | 99.59 | 20.87 | 37.465 | 94.981 | 108.8036 | 86.9137 | 0.009292 | 0.98051 | 0.023433 | 23.8986 | 0.30974 | 0.15575 | 0.012741 | 0.89349 | 0.89349 | 0.007164
2019-03-21 11:47 | 1553183230.3619 | 117 | 0.207 | 0.002677 | 3.5e-05 | 100.5568 | 97.256 | 116.9971 | 21.0507 | 0.003007 | 0.002729 | 174.7722 | 0.26236 | 8.4041 | 0.003605 | 0.00486 | 0.004865 | 7.1e-05 | 0.001001 | 0.33826 | 1.3152 | 2.548 | 99.59 | 20.87 | 37.461 | 93.9625 | 108.4417 | 87.2602 | 0.009291 | 0.97957 | 0.023417 | 23.9055 | 0.29853 | 0.15608 | 0.012463 | 0.87346 | 0.87346 | 0.007344
2019-03-21 11:47 | 1553183231.3607 | 118 | 0.20697 | 0.002669 | 3.4e-05 | 100.5708 | 97.4583 | 116.8594 | 21.0505 | 0.003007 | 0.002688 | 174.7181 | 0.26226 | 8.4025 | 0.003582 | 0.00484 | 0.004683 | 6.5e-05 | 0.000998 | 0.28296 | 1.2849 | 2.548 | 99.59 | 20.87 | 37.459 | 93.9661 | 107.8812 | 86.284 | 0.00929 | 0.98223 | 0.023435 | 23.859 | 0.28119 | 0.15602 | 0.012629 | 0.8858 | 0.8858 | 0.007239
2019-03-21 11:47 | 1553183232.3621 | 119 | 0.20703 | 0.002653 | 3.4e-05 | 100.6524 | 97.5728 | 116.5753 | 21.0518 | 0.003063 | 0.002749 | 174.6597 | 0.26229 | 8.4019 | 0.00353 | 0.005275 | 0.005215 | 6.6e-05 | 0.000994 | 0.30917 | 1.2923 | 2.548 | 99.59 | 20.86 | 37.457 | 95.3201 | 106.6131 | 87.7358 | 0.009284 | 0.97828 | 0.023359 | 23.8781 | 0.27795 | 0.15604 | 0.012308 | 0.86051 | 0.86051 | 0.007453
2019-03-21 11:47 | 1553183233.3612 | 120 | 0.20703 | 0.002643 | 3.3e-05 | 100.6781 | 97.6991 | 116.7096 | 21.0506 | 0.003053 | 0.002749 | 174.6419 | 0.26253 | 8.402 | 0.003551 | 0.00542 | 0.005305 | 6.8e-05 | 0.000998 | 0.30917 | 1.3106 | 2.5479 | 99.59 | 20.86 | 37.455 | 95.0793 | 107.111 | 87.7409 | 0.009284 | 0.97826 | 0.023373 | 23.892 | 0.26865 | 0.15618 | 0.012319 | 0.86186 | 0.86186 | 0.007448
2019-03-21 11:47 | 1553183234.3618 | 121 | 0.20702 | 0.002625 | 3.2e-05 | 100.7201 | 97.9746 | 116.6232 | 21.0513 | 0.002963 | 0.002666 | 174.5996 | 0.26238 | 8.4034 | 0.003576 | 0.004992 | 0.004997 | 6.4e-05 | 0.001 | 0.31974 | 1.3225 | 2.5479 | 99.59 | 20.86 | 37.453 | 92.8904 | 107.7174 | 85.7492 | 0.009283 | 0.98369 | 0.023429 | 23.8172 | 0.24719 | 0.15609 | 0.012401 | 0.86972 | 0.86972 | 0.007376
2019-03-21 11:47 | 1553183235.3609 | 122 | 0.20704 | 0.002612 | 3.1e-05 | 100.7186 | 98.1082 | 116.5016 | 21.0511 | 0.003059 | 0.002689 | 174.5595 | 0.26236 | 8.4018 | 0.003581 | 0.004765 | 0.004666 | 6.6e-05 | 0.000999 | 0.30381 | 1.3191 | 2.5479 | 99.59 | 20.86 | 37.449 | 95.2263 | 107.8566 | 86.3127 | 0.009282 | 0.98215 | 0.023398 | 23.8234 | 0.23488 | 0.15608 | 0.012284 | 0.86047 | 0.86047 | 0.007455
2019-03-21 11:47 | 1553183236.3603 | 123 | 0.20708 | 0.002596 | 3e-05 | 100.7221 | 98.1116 | 116.9129 | 21.0515 | 0.003028 | 0.002729 | 174.5229 | 0.26114 | 8.4018 | 0.003495 | 0.005353 | 0.00525 | 6.9e-05 | 0.001 | 0.31758 | 1.3025 | 2.5478 | 99.59 | 20.87 | 37.448 | 94.4727 | 105.7754 | 87.2655 | 0.009288 | 0.97955 | 0.023408 | 23.897 | 0.23489 | 0.15535 | 0.012034 | 0.84333 | 0.84333 | 0.007571
2019-03-21 11:47 | 1553183237.361 | 124 | 0.20707 | 0.002585 | 3e-05 | 100.7338 | 98.1708 | 117.126 | 21.0521 | 0.003074 | 0.002673 | 174.4945 | 0.26223 | 8.4018 | 0.003484 | 0.004373 | 0.00428 | 6.6e-05 | 0.001003 | 0.31653 | 1.3493 | 2.5478 | 99.59 | 20.87 | 37.446 | 95.5741 | 105.4949 | 85.9266 | 0.009287 | 0.9832 | 0.023473 | 23.8743 | 0.23052 | 0.156 | 0.012128 | 0.85229 | 0.85229 | 0.007523
2019-03-21 11:47 | 1553183238.3601 | 125 | 0.2071 | 0.002571 | 2.9e-05 | 100.782 | 98.1763 | 117.1075 | 21.0522 | 0.003028 | 0.002687 | 174.4788 | 0.26184 | 8.4036 | 0.00353 | 0.004842 | 0.004833 | 6.5e-05 | 0.001001 | 0.30309 | 1.2986 | 2.5479 | 99.59 | 20.87 | 37.445 | 94.4639 | 106.6062 | 86.2654 | 0.009287 | 0.98228 | 0.02346 | 23.8837 | 0.23431 | 0.15577 | 0.011979 | 0.84141 | 0.84141 | 0.007609
2019-03-21 11:47 | 1553183239.3607 | 126 | 0.2071 | 0.002564 | 2.9e-05 | 100.7481 | 98.4092 | 116.8821 | 21.0517 | 0.002978 | 0.00275 | 174.4397 | 0.26288 | 8.4041 | 0.003551 | 0.005041 | 0.004913 | 6.5e-05 | 0.000996 | 0.31188 | 1.3113 | 2.548 | 99.59 | 20.87 | 37.445 | 93.2547 | 107.1255 | 87.7695 | 0.009287 | 0.97818 | 0.023389 | 23.9106 | 0.21009 | 0.15639 | 0.012008 | 0.84085 | 0.84085 | 0.007644
2019-03-21 11:47 | 1553183240.3601 | 127 | 0.2071 | 0.002547 | 3e-05 | 100.824 | 98.6701 | 116.8063 | 21.0523 | 0.003007 | 0.0027 | 174.3892 | 0.26294 | 8.402 | 0.003445 | 0.004999 | 0.004914 | 6.6e-05 | 0.000997 | 0.32388 | 1.299 | 2.5479 | 99.59 | 20.87 | 37.444 | 93.972 | 104.5632 | 86.5699 | 0.009287 | 0.98145 | 0.02342 | 23.8631 | 0.19315 | 0.15642 | 0.011998 | 0.84125 | 0.84125 | 0.007642
2019-03-21 11:47 | 1553183241.3603 | 128 | 0.20714 | 0.002533 | 3e-05 | 100.8199 | 98.5444 | 116.8725 | 21.0527 | 0.002975 | 0.002696 | 174.3434 | 0.2637 | 8.4034 | 0.003438 | 0.005177 | 0.005129 | 6.7e-05 | 0.000996 | 0.30685 | 1.296 | 2.5479 | 99.59 | 20.87 | 37.442 | 93.1923 | 104.3921 | 86.4837 | 0.009286 | 0.98168 | 0.02343 | 23.867 | 0.20418 | 0.15687 | 0.011811 | 0.82845 | 0.82845 | 0.007783
2019-03-21 11:47 | 1553183242.3603 | 129 | 0.20718 | 0.002526 | 3.1e-05 | 100.7677 | 98.7253 | 117.0176 | 21.0532 | 0.00296 | 0.002752 | 174.3236 | 0.26281 | 8.403 | 0.003446 | 0.005299 | 0.005197 | 6.5e-05 | 0.000997 | 0.30773 | 1.3605 | 2.5479 | 99.59 | 20.87 | 37.442 | 92.8257 | 104.5663 | 87.8115 | 0.009286 | 0.97807 | 0.023401 | 23.9258 | 0.18316 | 0.15635 | 0.011567 | 0.81026 | 0.81026 | 0.007931
2019-03-21 11:47 | 1553183243.3604 | 130 | 0.20715 | 0.00251 | 3.1e-05 | 100.8962 | 98.742 | 117.1448 | 21.052 | 0.002993 | 0.002723 | 174.2902 | 0.26272 | 8.4007 | 0.003412 | 0.004918 | 0.00483 | 6.7e-05 | 0.000999 | 0.30663 | 1.353 | 2.548 | 99.59 | 20.87 | 37.439 | 93.6266 | 103.7514 | 87.1286 | 0.009285 | 0.97992 | 0.023436 | 23.9162 | 0.19304 | 0.15629 | 0.01174 | 0.82361 | 0.82361 | 0.007799
2019-03-21 11:47 | 1553183244.3603 | 131 | 0.20713 | 0.002497 | 3.1e-05 | 100.7686 | 98.8295 | 117.379 | 21.0533 | 0.003002 | 0.002688 | 174.2633 | 0.26292 | 8.4033 | 0.003363 | 0.004843 | 0.004778 | 6.5e-05 | 0.00099 | 0.29695 | 1.2869 | 2.548 | 99.59 | 20.87 | 37.439 | 93.8359 | 102.5654 | 86.2915 | 0.009285 | 0.9822 | 0.023487 | 23.9123 | 0.1738 | 0.15641 | 0.011864 | 0.83416 | 0.83416 | 0.007706
2019-03-21 11:47 | 1553183245.3607 | 132 | 0.20716 | 0.002485 | 3.2e-05 | 100.7844 | 98.7665 | 117.7483 | 21.053 | 0.00297 | 0.002655 | 174.2191 | 0.26295 | 8.4022 | 0.003418 | 0.004926 | 0.004831 | 6.7e-05 | 0.000991 | 0.32189 | 1.3072 | 2.5478 | 99.59 | 20.87 | 37.437 | 93.08 | 103.8949 | 85.4849 | 0.009285 | 0.98442 | 0.02355 | 23.923 | 0.18091 | 0.15643 | 0.0117 | 0.82478 | 0.82478 | 0.007795
2019-03-21 11:47 | 1553183246.3615 | 133 | 0.20721 | 0.002479 | 3.2e-05 | 100.8094 | 98.9256 | 117.4868 | 21.0536 | 0.003032 | 0.002671 | 174.1861 | 0.2626 | 8.4031 | 0.003388 | 0.005108 | 0.005024 | 6.8e-05 | 0.000995 | 0.3178 | 1.2764 | 2.5479 | 99.59 | 20.87 | 37.436 | 94.5633 | 103.1668 | 85.8616 | 0.009285 | 0.98338 | 0.023512 | 23.9089 | 0.16873 | 0.15622 | 0.011419 | 0.80364 | 0.80364 | 0.007989
2019-03-21 11:47 | 1553183247.3606 | 134 | 0.20722 | 0.002466 | 3.2e-05 | 100.788 | 98.9801 | 117.2229 | 21.053 | 0.002975 | 0.002664 | 174.1504 | 0.26238 | 8.4018 | 0.003341 | 0.00437 | 0.004268 | 6.8e-05 | 0.000992 | 0.31653 | 1.2548 | 2.5479 | 99.59 | 20.87 | 37.436 | 93.2034 | 102.0396 | 85.7057 | 0.009285 | 0.98381 | 0.02349 | 23.8769 | 0.1619 | 0.15609 | 0.011387 | 0.80064 | 0.80064 | 0.008013
2019-03-21 11:47 | 1553183248.3606 | 135 | 0.2072 | 0.002457 | 3.1e-05 | 100.8532 | 99.0996 | 117.0164 | 21.0548 | 0.002993 | 0.002677 | 174.0859 | 0.26344 | 8.4022 | 0.003386 | 0.004581 | 0.004535 | 6.6e-05 | 0.000988 | 0.3194 | 1.303 | 2.5479 | 99.59 | 20.88 | 37.435 | 93.6394 | 103.1309 | 86.0259 | 0.00929 | 0.98293 | 0.023459 | 23.8665 | 0.1569 | 0.15672 | 0.011524 | 0.80922 | 0.80922 | 0.00796
2019-03-21 11:47 | 1553183249.3597 | 136 | 0.20721 | 0.002445 | 3.1e-05 | 100.9096 | 99.0972 | 117.1551 | 21.0534 | 0.002924 | 0.002687 | 174.0384 | 0.26218 | 8.4032 | 0.003407 | 0.004609 | 0.004454 | 6.7e-05 | 0.00099 | 0.30502 | 1.3149 | 2.5479 | 99.59 | 20.87 | 37.435 | 91.9531 | 103.6445 | 86.2642 | 0.009284 | 0.98228 | 0.023465 | 23.8885 | 0.1621 | 0.15597 | 0.011451 | 0.80433 | 0.80433 | 0.00797
2019-03-21 11:47 | 1553183250.3596 | 137 | 0.20725 | 0.002434 | 3e-05 | 100.9057 | 99.2866 | 117.4354 | 21.0537 | 0.002967 | 0.002649 | 174.0016 | 0.26289 | 8.4028 | 0.003333 | 0.004512 | 0.004442 | 6.9e-05 | 0.00099 | 0.29607 | 1.3162 | 2.5479 | 99.59 | 20.88 | 37.434 | 93.0061 | 101.8354 | 85.338 | 0.00929 | 0.98482 | 0.023524 | 23.8863 | 0.14468 | 0.15639 | 0.011254 | 0.79251 | 0.79251 | 0.008111
2019-03-21 11:48 | 1553183251.3607 | 138 | 0.20729 | 0.002429 | 3.1e-05 | 100.7838 | 99.3595 | 117.6152 | 21.054 | 0.003015 | 0.002639 | 173.9738 | 0.26292 | 8.4049 | 0.003327 | 0.004597 | 0.004415 | 6.8e-05 | 0.000989 | 0.32333 | 1.317 | 2.548 | 99.59 | 20.88 | 37.432 | 94.1465 | 101.7077 | 85.1103 | 0.009289 | 0.98544 | 0.023549 | 23.8969 | 0.12731 | 0.15641 | 0.010984 | 0.77429 | 0.77429 | 0.008302
2019-03-21 11:48 | 1553183252.3609 | 139 | 0.20731 | 0.002414 | 3.1e-05 | 100.9292 | 99.4521 | 117.7635 | 21.0546 | 0.002945 | 0.002666 | 173.9119 | 0.26245 | 8.4038 | 0.003378 | 0.004758 | 0.0047 | 6.7e-05 | 0.00099 | 0.31078 | 1.3183 | 2.5479 | 99.59 | 20.87 | 37.43 | 92.4769 | 102.9384 | 85.7448 | 0.009283 | 0.9837 | 0.023543 | 23.9332 | 0.13187 | 0.15613 | 0.010916 | 0.76924 | 0.76924 | 0.008342
2019-03-21 11:48 | 1553183253.3608 | 140 | 0.20732 | 0.002405 | 3.2e-05 | 100.8509 | 99.6827 | 117.6483 | 21.0547 | 0.002899 | 0.002687 | 173.8635 | 0.2622 | 8.4053 | 0.003511 | 0.004481 | 0.004402 | 6.6e-05 | 0.000989 | 0.27527 | 1.3603 | 2.5479 | 99.59 | 20.88 | 37.425 | 91.3736 | 106.1664 | 86.2609 | 0.009288 | 0.98229 | 0.023515 | 23.9386 | 0.10421 | 0.15598 | 0.010857 | 0.76413 | 0.76413 | 0.00839
2019-03-21 11:48 | 1553183254.3599 | 141 | 0.20733 | 0.002391 | 3.2e-05 | 100.8145 | 99.589 | 117.7487 | 21.055 | 0.002983 | 0.002627 | 173.8081 | 0.26326 | 8.4033 | 0.003419 | 0.004899 | 0.004813 | 7e-05 | 0.000987 | 0.32278 | 1.3383 | 2.5479 | 99.59 | 20.88 | 37.425 | 93.3907 | 103.9329 | 84.8157 | 0.009288 | 0.98625 | 0.023572 | 23.9007 | 0.10939 | 0.15661 | 0.010832 | 0.76425 | 0.76425 | 0.008422
2019-03-21 11:48 | 1553183255.3599 | 142 | 0.20734 | 0.002386 | 3.1e-05 | 100.8506 | 99.4392 | 117.3809 | 21.0547 | 0.002936 | 0.002671 | 173.7686 | 0.26337 | 8.4025 | 0.003394 | 0.004169 | 0.004034 | 7e-05 | 0.000987 | 0.29623 | 1.2803 | 2.548 | 99.59 | 20.88 | 37.424 | 92.2516 | 103.3149 | 85.8659 | 0.009287 | 0.98337 | 0.023501 | 23.8983 | 0.12606 | 0.15668 | 0.010786 | 0.7587 | 0.7587 | 0.008488
2019-03-21 11:48 | 1553183256.36 | 143 | 0.20739 | 0.002372 | 3.1e-05 | 100.9277 | 99.476 | 117.5692 | 21.0553 | 0.002964 | 0.002666 | 173.7205 | 0.26184 | 8.4028 | 0.003366 | 0.00416 | 0.004111 | 7e-05 | 0.000986 | 0.31492 | 1.2767 | 2.548 | 99.59 | 20.88 | 37.422 | 92.9225 | 102.6373 | 85.7617 | 0.009287 | 0.98365 | 0.023523 | 23.914 | 0.12959 | 0.15577 | 0.010491 | 0.73869 | 0.73869 | 0.008667
2019-03-21 11:48 | 1553183257.361 | 144 | 0.20739 | 0.002359 | 3e-05 | 100.8759 | 99.7307 | 117.8797 | 21.0558 | 0.002929 | 0.002649 | 173.6776 | 0.26342 | 8.4017 | 0.00336 | 0.004871 | 0.004766 | 7e-05 | 0.000987 | 0.32781 | 1.2548 | 2.5479 | 99.59 | 20.88 | 37.42 | 92.0837 | 102.5018 | 85.3565 | 0.009286 | 0.98477 | 0.023567 | 23.932 | 0.10212 | 0.15671 | 0.010467 | 0.73839 | 0.73839 | 0.008723
2019-03-21 11:48 | 1553183258.3601 | 145 | 0.20739 | 0.00235 | 3e-05 | 100.8881 | 99.7011 | 118.1503 | 21.0551 | 0.002907 | 0.002589 | 173.6283 | 0.26219 | 8.4026 | 0.003347 | 0.00472 | 0.004666 | 7.2e-05 | 0.000979 | 0.35834 | 1.3212 | 2.548 | 99.59 | 20.87 | 37.417 | 91.552 | 102.1766 | 83.8911 | 0.00928 | 0.98881 | 0.023643 | 23.9105 | 0.10586 | 0.15598 | 0.010488 | 0.7423 | 0.7423 | 0.008636
2019-03-21 11:48 | 1553183259.3612 | 146 | 0.2074 | 0.002338 | 2.9e-05 | 100.9033 | 99.8387 | 118.1426 | 21.055 | 0.002903 | 0.002621 | 173.6055 | 0.26202 | 8.4012 | 0.003251 | 0.004522 | 0.004407 | 7e-05 | 0.000976 | 0.31133 | 1.303 | 2.5479 | 99.59 | 20.87 | 37.415 | 91.4691 | 99.8469 | 84.669 | 0.009279 | 0.98666 | 0.023616 | 23.9357 | 0.094874 | 0.15588 | 0.010468 | 0.74005 | 0.74005 | 0.008657
2019-03-21 11:48 | 1553183260.3603 | 147 | 0.20742 | 0.002326 | 2.9e-05 | 100.8252 | 99.9715 | 117.9614 | 21.0555 | 0.002917 | 0.002589 | 173.5997 | 0.26327 | 8.4046 | 0.003344 | 0.005113 | 0.004931 | 6.9e-05 | 0.000975 | 0.31116 | 1.3379 | 2.5478 | 99.59 | 20.88 | 37.412 | 91.7934 | 102.1056 | 83.8985 | 0.009284 | 0.98879 | 0.023624 | 23.8916 | 0.076059 | 0.15662 | 0.010342 | 0.73142 | 0.73142 | 0.008801
2019-03-21 11:48 | 1553183261.3615 | 148 | 0.20742 | 0.002316 | 2.9e-05 | 100.8699 | 99.8584 | 117.5194 | 21.0544 | 0.002939 | 0.002666 | 173.5569 | 0.26284 | 8.4004 | 0.003325 | 0.004249 | 0.004265 | 6.8e-05 | 0.000972 | 0.30574 | 1.3368 | 2.5478 | 99.59 | 20.88 | 37.407 | 92.3366 | 101.6476 | 85.7609 | 0.009283 | 0.98366 | 0.023518 | 23.9089 | 0.090146 | 0.15636 | 0.01036 | 0.72939 | 0.72939 | 0.008811
2019-03-21 11:48 | 1553183262.3605 | 149 | 0.20742 | 0.002308 | 2.8e-05 | 100.8568 | 99.9031 | 117.3767 | 21.0557 | 0.002866 | 0.002612 | 173.5171 | 0.26296 | 8.4027 | 0.003327 | 0.004217 | 0.004169 | 6.8e-05 | 0.000971 | 0.31924 | 1.3406 | 2.5478 | 99.59 | 20.88 | 37.404 | 90.5607 | 101.7084 | 84.4652 | 0.009282 | 0.98722 | 0.023546 | 23.8512 | 0.084988 | 0.15644 | 0.01037 | 0.73101 | 0.73101 | 0.008795
2019-03-21 11:48 | 1553183263.3604 | 150 | 0.20748 | 0.002293 | 2.8e-05 | 100.8333 | 99.9545 | 117.6849 | 21.0556 | 0.002928 | 0.002642 | 173.4908 | 0.26301 | 8.4038 | 0.003362 | 0.005027 | 0.004945 | 6.8e-05 | 0.000973 | 0.3178 | 1.2959 | 2.5479 | 99.59 | 20.88 | 37.401 | 92.0625 | 102.5379 | 85.1845 | 0.009282 | 0.98524 | 0.023554 | 23.9065 | 0.078297 | 0.15647 | 0.010059 | 0.70929 | 0.70929 | 0.009067
2019-03-21 11:48 | 1553183264.3604 | 151 | 0.20746 | 0.002281 | 2.8e-05 | 100.8422 | 99.9606 | 118.0058 | 21.0555 | 0.002919 | 0.002653 | 173.4629 | 0.26222 | 8.4051 | 0.003357 | 0.005032 | 0.004916 | 6.8e-05 | 0.000972 | 0.3131 | 1.3106 | 2.5478 | 99.59 | 20.88 | 37.397 | 91.8481 | 102.4206 | 85.4476 | 0.009281 | 0.98452 | 0.023577 | 23.9478 | 0.078541 | 0.15599 | 0.010161 | 0.7172 | 0.7172 | 0.00894
2019-03-21 11:48 | 1553183265.3595 | 152 | 0.20748 | 0.00227 | 3e-05 | 100.9119 | 100.011 | 118.371 | 21.0563 | 0.002903 | 0.00262 | 173.4267 | 0.26279 | 8.4041 | 0.003306 | 0.004938 | 0.004859 | 6.7e-05 | 0.000967 | 0.30973 | 1.3037 | 2.548 | 99.59 | 20.88 | 37.396 | 91.4721 | 101.1917 | 84.639 | 0.009281 | 0.98674 | 0.02364 | 23.9578 | 0.080213 | 0.15634 | 0.010086 | 0.71372 | 0.71372 | 0.009003
2019-03-21 11:48 | 1553183266.3595 | 153 | 0.20749 | 0.002262 | 3e-05 | 100.9078 | 100.1803 | 118.2982 | 21.0562 | 0.002885 | 0.002618 | 173.3767 | 0.26337 | 8.4016 | 0.003333 | 0.004605 | 0.004509 | 7.1e-05 | 0.000971 | 0.29535 | 1.3039 | 2.5479 | 99.59 | 20.88 | 37.391 | 91.0223 | 101.833 | 84.6045 | 0.009279 | 0.98684 | 0.023634 | 23.9493 | 0.064718 | 0.15668 | 0.009999 | 0.70737 | 0.70737 | 0.009103
2019-03-21 11:48 | 1553183267.3607 | 154 | 0.20747 | 0.002252 | 3.1e-05 | 100.8392 | 100.1425 | 118.1684 | 21.0562 | 0.00285 | 0.002558 | 173.3574 | 0.262 | 8.4017 | 0.003247 | 0.004663 | 0.004625 | 6.7e-05 | 0.000967 | 0.30309 | 1.3033 | 2.5479 | 99.59 | 20.88 | 37.387 | 90.1899 | 99.7523 | 83.1566 | 0.009278 | 0.99085 | 0.023669 | 23.8877 | 0.062017 | 0.15587 | 0.010126 | 0.71733 | 0.71733 | 0.00893
2019-03-21 11:48 | 1553183268.3596 | 155 | 0.20751 | 0.002238 | 3.1e-05 | 100.7642 | 100.1272 | 117.9924 | 21.0558 | 0.002882 | 0.002618 | 173.34 | 0.26271 | 8.4014 | 0.003362 | 0.004796 | 0.004757 | 6.8e-05 | 0.000971 | 0.30901 | 1.3274 | 2.5479 | 99.59 | 20.88 | 37.385 | 90.9634 | 102.556 | 84.5971 | 0.009278 | 0.98686 | 0.023604 | 23.9181 | 0.056724 | 0.15629 | 0.009903 | 0.69963 | 0.69963 | 0.009181
2019-03-21 11:48 | 1553183269.3597 | 156 | 0.20751 | 0.002231 | 3.2e-05 | 100.8422 | 100.3373 | 117.9546 | 21.0566 | 0.002897 | 0.002584 | 173.2926 | 0.26276 | 8.4036 | 0.003321 | 0.003999 | 0.00387 | 6.8e-05 | 0.000968 | 0.33035 | 1.307 | 2.5479 | 99.59 | 20.89 | 37.379 | 91.3113 | 101.5609 | 83.7865 | 0.009282 | 0.9891 | 0.023627 | 23.8871 | 0.044898 | 0.15632 | 0.00996 | 0.70427 | 0.70427 | 0.009122
2019-03-21 11:48 | 1553183270.3611 | 157 | 0.20751 | 0.002226 | 3.1e-05 | 100.8512 | 100.2188 | 117.9788 | 21.0556 | 0.002929 | 0.002612 | 173.235 | 0.26212 | 8.4014 | 0.003247 | 0.005099 | 0.005026 | 7e-05 | 0.000965 | 0.30646 | 1.303 | 2.5481 | 99.59 | 20.89 | 37.377 | 92.089 | 99.7651 | 84.4526 | 0.009282 | 0.98726 | 0.023607 | 23.9119 | 0.056263 | 0.15594 | 0.009923 | 0.70108 | 0.70108 | 0.009142
2019-03-21 11:48 | 1553183271.3612 | 158 | 0.20752 | 0.002214 | 3.2e-05 | 100.8348 | 100.2753 | 118.1928 | 21.0565 | 0.002851 | 0.002599 | 173.2034 | 0.26252 | 8.4016 | 0.003265 | 0.004842 | 0.004756 | 6.5e-05 | 0.000965 | 0.31813 | 1.308 | 2.5479 | 99.59 | 20.89 | 37.372 | 90.2137 | 100.1889 | 84.1344 | 0.00928 | 0.98813 | 0.023639 | 23.9229 | 0.049765 | 0.15617 | 0.009904 | 0.70065 | 0.70065 | 0.009161
2019-03-21 11:48 | 1553183272.3609 | 159 | 0.20755 | 0.002201 | 3.1e-05 | 100.7987 | 100.2992 | 118.091 | 21.0573 | 0.0029 | 0.002579 | 173.1943 | 0.26238 | 8.4008 | 0.003229 | 0.005211 | 0.005121 | 6.5e-05 | 0.000968 | 0.34291 | 1.2992 | 2.548 | 99.59 | 20.89 | 37.369 | 91.3775 | 99.3376 | 83.6631 | 0.00928 | 0.98944 | 0.023644 | 23.8968 | 0.04444 | 0.15609 | 0.009735 | 0.68886 | 0.68886 | 0.009313
2019-03-21 11:48 | 1553183273.3599 | 160 | 0.20758 | 0.002195 | 3.2e-05 | 100.925 | 100.2253 | 117.8649 | 21.0563 | 0.002871 | 0.002589 | 173.154 | 0.26277 | 8.4023 | 0.003275 | 0.004533 | 0.00447 | 6.9e-05 | 0.000965 | 0.30735 | 1.2865 | 2.548 | 99.59 | 20.89 | 37.365 | 90.6918 | 100.4484 | 83.911 | 0.009279 | 0.98875 | 0.023614 | 23.8822 | 0.062236 | 0.15632 | 0.009589 | 0.67756 | 0.67756 | 0.009482
2019-03-21 11:48 | 1553183274.3611 | 161 | 0.20756 | 0.002183 | 3.1e-05 | 100.8032 | 100.1738 | 117.4415 | 21.0567 | 0.002846 | 0.0026 | 173.1333 | 0.26219 | 8.401 | 0.003222 | 0.004615 | 0.004506 | 6.9e-05 | 0.000968 | 0.30989 | 1.2651 | 2.5478 | 99.59 | 20.89 | 37.361 | 90.0889 | 99.1504 | 84.158 | 0.009278 | 0.98807 | 0.023563 | 23.8475 | 0.056021 | 0.15598 | 0.009711 | 0.6848 | 0.6848 | 0.009361
2019-03-21 11:48 | 1553183275.3611 | 162 | 0.20758 | 0.002179 | 3.1e-05 | 100.8109 | 100.2157 | 117.2814 | 21.0566 | 0.002872 | 0.002595 | 173.0964 | 0.26329 | 8.4025 | 0.003285 | 0.004613 | 0.004472 | 7e-05 | 0.000967 | 0.2724 | 1.3101 | 2.5478 | 99.59 | 20.89 | 37.359 | 90.7022 | 100.6935 | 84.0517 | 0.009277 | 0.98836 | 0.02355 | 23.8277 | 0.052968 | 0.15663 | 0.009603 | 0.67678 | 0.67678 | 0.009512
2019-03-21 11:48 | 1553183276.361 | 163 | 0.20758 | 0.002169 | 3e-05 | 100.8348 | 100.399 | 117.3171 | 21.0561 | 0.002867 | 0.002574 | 173.093 | 0.26232 | 8.4016 | 0.003216 | 0.004657 | 0.004562 | 6.5e-05 | 0.000967 | 0.32892 | 1.3832 | 2.5479 | 99.59 | 20.89 | 37.355 | 90.5939 | 99.0079 | 83.5309 | 0.009276 | 0.98981 | 0.023571 | 23.814 | 0.038743 | 0.15605 | 0.009597 | 0.67705 | 0.67705 | 0.009473
2019-03-21 11:48 | 1553183277.3602 | 164 | 0.20757 | 0.002159 | 3e-05 | 100.7659 | 100.3945 | 117.2819 | 21.0566 | 0.002864 | 0.002604 | 173.0821 | 0.26249 | 8.4005 | 0.0032 | 0.004991 | 0.004889 | 6.7e-05 | 0.000968 | 0.3079 | 1.327 | 2.5479 | 99.59 | 20.89 | 37.35 | 90.5107 | 98.6328 | 84.2569 | 0.009275 | 0.9878 | 0.023544 | 23.8346 | 0.033031 | 0.15615 | 0.009652 | 0.68015 | 0.68015 | 0.009436
2019-03-21 11:48 | 1553183278.3612 | 165 | 0.20761 | 0.002151 | 3e-05 | 100.7937 | 100.4316 | 117.403 | 21.0581 | 0.002848 | 0.002576 | 173.0508 | 0.26284 | 8.4015 | 0.003195 | 0.004427 | 0.004332 | 6.5e-05 | 0.000967 | 0.32261 | 1.3241 | 2.548 | 99.59 | 20.89 | 37.348 | 90.1309 | 98.5103 | 83.6014 | 0.009274 | 0.98961 | 0.023578 | 23.825 | 0.032183 | 0.15636 | 0.009447 | 0.66666 | 0.66666 | 0.00964
2019-03-21 11:48 | 1553183279.3616 | 166 | 0.2076 | 0.00214 | 2.8e-05 | 100.7752 | 100.5767 | 117.318 | 21.0585 | 0.002863 | 0.002513 | 173.0142 | 0.26293 | 8.4022 | 0.003236 | 0.005089 | 0.004943 | 6.6e-05 | 0.000969 | 0.29983 | 1.2993 | 2.5479 | 99.59 | 20.89 | 37.342 | 90.5011 | 99.4926 | 82.074 | 0.009273 | 0.99387 | 0.02362 | 23.7654 | 0.017631 | 0.15642 | 0.009539 | 0.67442 | 0.67442 | 0.009532
2019-03-21 11:48 | 1553183280.3613 | 167 | 0.20766 | 0.002133 | 2.8e-05 | 100.7484 | 100.5641 | 116.9946 | 21.0573 | 0.00284 | 0.002579 | 172.9887 | 0.26254 | 8.4017 | 0.003276 | 0.004716 | 0.004617 | 6.4e-05 | 0.000969 | 0.3011 | 1.318 | 2.548 | 99.59 | 20.89 | 37.341 | 89.9481 | 100.4582 | 83.6656 | 0.009273 | 0.98943 | 0.023534 | 23.7857 | 0.016371 | 0.15619 | 0.009145 | 0.64422 | 0.64422 | 0.009965
2019-03-21 11:48 | 1553183281.3615 | 168 | 0.20764 | 0.002121 | 2.6e-05 | 100.8148 | 100.5417 | 117.1051 | 21.057 | 0.002777 | 0.002614 | 172.9513 | 0.26331 | 8.4015 | 0.003289 | 0.004733 | 0.004667 | 6.5e-05 | 0.000968 | 0.3194 | 1.2878 | 2.548 | 99.59 | 20.89 | 37.338 | 88.4221 | 100.7849 | 84.4937 | 0.009272 | 0.98714 | 0.023518 | 23.8245 | 0.024268 | 0.15664 | 0.009282 | 0.65358 | 0.65358 | 0.00985
2019-03-21 11:48 | 1553183282.3594 | 169 | 0.20766 | 0.002115 | 2.6e-05 | 100.7183 | 100.2947 | 117.2342 | 21.0565 | 0.002882 | 0.002569 | 172.9073 | 0.26171 | 8.4016 | 0.003277 | 0.004907 | 0.004823 | 6.4e-05 | 0.000965 | 0.29574 | 1.3154 | 2.5479 | 99.59 | 20.89 | 37.336 | 90.9438 | 100.498 | 83.4222 | 0.009271 | 0.99011 | 0.023567 | 23.8019 | 0.037693 | 0.15569 | 0.009195 | 0.64878 | 0.64878 | 0.009863
2019-03-21 11:48 | 1553183283.3595 | 170 | 0.20768 | 0.00211 | 2.6e-05 | 100.7025 | 100.4139 | 117.3192 | 21.0571 | 0.002806 | 0.002602 | 172.8635 | 0.26173 | 8.4021 | 0.003253 | 0.00468 | 0.004571 | 6.6e-05 | 0.000969 | 0.32565 | 1.2975 | 2.5481 | 99.59 | 20.89 | 37.333 | 89.118 | 99.9148 | 84.2271 | 0.009271 | 0.98788 | 0.023548 | 23.8374 | 0.025666 | 0.15571 | 0.009096 | 0.64129 | 0.64129 | 0.009979
2019-03-21 11:48 | 1553183284.3605 | 171 | 0.20766 | 0.002096 | 2.6e-05 | 100.7484 | 100.5502 | 117.4912 | 21.0574 | 0.002876 | 0.002589 | 172.8094 | 0.26225 | 8.4022 | 0.003369 | 0.005294 | 0.005183 | 6.6e-05 | 0.000966 | 0.33356 | 1.348 | 2.548 | 99.59 | 20.89 | 37.33 | 90.8156 | 102.714 | 83.8989 | 0.00927 | 0.98879 | 0.023577 | 23.8439 | 0.017613 | 0.15601 | 0.009198 | 0.64925 | 0.64925 | 0.009876
2019-03-21 11:48 | 1553183285.3595 | 172 | 0.20769 | 0.002085 | 2.7e-05 | 100.6962 | 100.4769 | 117.6905 | 21.0579 | 0.002855 | 0.002587 | 172.7849 | 0.26045 | 8.4035 | 0.003369 | 0.00427 | 0.004239 | 6.4e-05 | 0.000964 | 0.31725 | 1.3738 | 2.5479 | 99.59 | 20.89 | 37.328 | 90.2988 | 102.7169 | 83.8642 | 0.009269 | 0.98888 | 0.023598 | 23.863 | 0.019499 | 0.15494 | 0.009057 | 0.63979 | 0.63979 | 0.009954
2019-03-21 11:48 | 1553183286.36 | 173 | 0.20766 | 0.002076 | 2.7e-05 | 100.7314 | 100.7406 | 117.8562 | 21.0582 | 0.002852 | 0.002558 | 172.7231 | 0.26204 | 8.4028 | 0.003237 | 0.005192 | 0.005056 | 6.1e-05 | 0.000965 | 0.26825 | 1.3 | 2.548 | 99.59 | 20.89 | 37.326 | 90.2279 | 99.5292 | 83.157 | 0.009269 | 0.99085 | 0.023638 | 23.8561 | -0.000815 | 0.15589 | 0.009258 | 0.65508 | 0.65508 | 0.00978
2019-03-21 11:48 | 1553183287.3607 | 174 | 0.20769 | 0.002071 | 2.8e-05 | 100.7195 | 100.5982 | 118.2527 | 21.0584 | 0.002849 | 0.002559 | 172.6688 | 0.26143 | 8.4014 | 0.003282 | 0.005025 | 0.004956 | 6.4e-05 | 0.000967 | 0.30165 | 1.3092 | 2.548 | 99.59 | 20.89 | 37.324 | 90.1695 | 100.6138 | 83.1922 | 0.009268 | 0.99075 | 0.023676 | 23.8974 | 0.010776 | 0.15553 | 0.00908 | 0.6435 | 0.6435 | 0.009933
2019-03-21 11:48 | 1553183288.3597 | 175 | 0.20768 | 0.00206 | 2.9e-05 | 100.723 | 100.4453 | 118.4252 | 21.0584 | 0.002863 | 0.002547 | 172.6171 | 0.26275 | 8.403 | 0.003288 | 0.004428 | 0.004362 | 6.4e-05 | 0.000968 | 0.3235 | 1.3253 | 2.5479 | 99.59 | 20.89 | 37.323 | 90.5094 | 100.7527 | 82.8843 | 0.009268 | 0.9916 | 0.023704 | 23.9045 | 0.024701 | 0.15631 | 0.009123 | 0.64726 | 0.64726 | 0.009925
2019-03-21 11:48 | 1553183289.3612 | 176 | 0.20772 | 0.002054 | 2.8e-05 | 100.7758 | 100.573 | 118.8594 | 21.0584 | 0.002795 | 0.002533 | 172.5713 | 0.26144 | 8.4029 | 0.00323 | 0.004367 | 0.00428 | 6.3e-05 | 0.000971 | 0.28838 | 1.3027 | 2.5479 | 99.59 | 20.89 | 37.319 | 88.8701 | 99.3505 | 82.5472 | 0.009267 | 0.99255 | 0.023758 | 23.9369 | 0.018017 | 0.15553 | 0.008934 | 0.63533 | 0.63533 | 0.010061
2019-03-21 11:48 | 1553183290.3612 | 177 | 0.20769 | 0.002047 | 2.7e-05 | 100.6882 | 100.5801 | 119.1815 | 21.0587 | 0.002812 | 0.002542 | 172.5254 | 0.2622 | 8.4037 | 0.003309 | 0.004634 | 0.004476 | 6.3e-05 | 0.000971 | 0.3413 | 1.346 | 2.5478 | 99.59 | 20.9 | 37.318 | 89.2626 | 101.2562 | 82.7845 | 0.009273 | 0.99188 | 0.023783 | 23.9773 | 0.009603 | 0.15598 | 0.009085 | 0.64678 | 0.64678 | 0.009912
2019-03-21 11:48 | 1553183291.3599 | 178 | 0.20771 | 0.00204 | 2.6e-05 | 100.6655 | 100.7058 | 119.0369 | 21.0581 | 0.002838 | 0.002559 | 172.51 | 0.26198 | 8.4048 | 0.003335 | 0.004911 | 0.004842 | 6.4e-05 | 0.000972 | 0.2943 | 1.3605 | 2.5478 | 99.59 | 20.9 | 37.318 | 89.8941 | 101.9024 | 83.1874 | 0.009273 | 0.99076 | 0.023755 | 23.9763 | -0.003578 | 0.15585 | 0.008959 | 0.63714 | 0.63714 | 0.010054
2019-03-21 11:48 | 1553183292.361 | 179 | 0.20769 | 0.002038 | 2.6e-05 | 100.7034 | 100.6074 | 119.2399 | 21.0587 | 0.002848 | 0.002516 | 172.4692 | 0.26301 | 8.4044 | 0.003333 | 0.005159 | 0.00513 | 6.2e-05 | 0.00097 | 0.31996 | 1.3168 | 2.5479 | 99.59 | 20.89 | 37.315 | 90.1298 | 101.8481 | 82.1395 | 0.009266 | 0.99368 | 0.02381 | 23.9615 | 0.00853 | 0.15647 | 0.009088 | 0.64786 | 0.64786 | 0.009926
2019-03-21 11:48 | 1553183293.3599 | 180 | 0.20773 | 0.002032 | 2.5e-05 | 100.6751 | 100.7944 | 119.3235 | 21.0587 | 0.002792 | 0.002551 | 172.4487 | 0.26205 | 8.4009 | 0.003401 | 0.005265 | 0.005192 | 6.5e-05 | 0.000971 | 0.30469 | 1.329 | 2.5478 | 99.59 | 20.89 | 37.315 | 88.7861 | 103.483 | 82.992 | 0.009266 | 0.9913 | 0.02379 | 23.9986 | -0.010593 | 0.15589 | 0.008874 | 0.63209 | 0.63209 | 0.010137
2019-03-21 11:48 | 1553183294.3611 | 181 | 0.20777 | 0.002022 | 2.6e-05 | 100.6602 | 100.7249 | 119.4519 | 21.0581 | 0.002764 | 0.00253 | 172.4235 | 0.26374 | 8.4034 | 0.0033 | 0.00464 | 0.004568 | 6.6e-05 | 0.000976 | 0.32045 | 1.3139 | 2.5479 | 99.59 | 20.9 | 37.312 | 88.1075 | 101.038 | 82.4805 | 0.009271 | 0.99273 | 0.02382 | 23.9943 | -0.005749 | 0.1569 | 0.008665 | 0.61794 | 0.61794 | 0.010435
2019-03-21 11:48 | 1553183295.3614 | 182 | 0.20776 | 0.002018 | 2.6e-05 | 100.6769 | 100.5032 | 119.5141 | 21.059 | 0.002811 | 0.002558 | 172.3973 | 0.26291 | 8.4014 | 0.003332 | 0.004408 | 0.004304 | 6.6e-05 | 0.000977 | 0.30845 | 1.3124 | 2.548 | 99.59 | 20.9 | 37.311 | 89.2371 | 101.8232 | 83.1695 | 0.009271 | 0.99081 | 0.023803 | 24.0237 | 0.015445 | 0.1564 | 0.008731 | 0.62213 | 0.62213 | 0.010333
2019-03-21 11:48 | 1553183296.3612 | 183 | 0.20778 | 0.002012 | 2.7e-05 | 100.6781 | 100.7344 | 119.3134 | 21.0581 | 0.002788 | 0.002536 | 172.3882 | 0.26228 | 8.4039 | 0.00331 | 0.005485 | 0.005422 | 6.8e-05 | 0.000974 | 0.30718 | 1.2832 | 2.5478 | 99.59 | 20.9 | 37.312 | 88.6887 | 101.3011 | 82.6311 | 0.009271 | 0.99231 | 0.023801 | 23.9854 | -0.005007 | 0.15603 | 0.008596 | 0.61241 | 0.61241 | 0.010472
2019-03-21 11:48 | 1553183297.3592 | 184 | 0.20777 | 0.002 | 2.9e-05 | 100.611 | 100.6758 | 118.89 | 21.0587 | 0.002785 | 0.002543 | 172.3513 | 0.26243 | 8.4025 | 0.003412 | 0.004744 | 0.0046 | 6.4e-05 | 0.000975 | 0.29535 | 1.2985 | 2.548 | 99.59 | 20.9 | 37.313 | 88.6149 | 103.7674 | 82.7925 | 0.009271 | 0.99186 | 0.023753 | 23.9482 | -0.005762 | 0.15612 | 0.008693 | 0.61796 | 0.61796 | 0.010383
2019-03-21 11:48 | 1553183298.3597 | 185 | 0.20776 | 0.001996 | 2.9e-05 | 100.7248 | 100.6265 | 118.8904 | 21.0592 | 0.00279 | 0.002532 | 172.3296 | 0.26158 | 8.4021 | 0.003352 | 0.004746 | 0.004657 | 6.7e-05 | 0.000977 | 0.31094 | 1.2923 | 2.548 | 99.59 | 20.9 | 37.311 | 88.7382 | 102.3056 | 82.54 | 0.009271 | 0.99257 | 0.023762 | 23.9398 | 0.008739 | 0.15561 | 0.008736 | 0.62126 | 0.62126 | 0.010295
2019-03-21 11:48 | 1553183299.3593 | 186 | 0.20778 | 0.00199 | 3e-05 | 100.6652 | 100.6789 | 118.655 | 21.0584 | 0.002796 | 0.002556 | 172.3112 | 0.26286 | 8.4042 | 0.003356 | 0.004943 | 0.004869 | 6.3e-05 | 0.000976 | 0.34451 | 1.2775 | 2.5478 | 99.59 | 20.9 | 37.312 | 88.8905 | 102.4014 | 83.1119 | 0.009271 | 0.99097 | 0.023719 | 23.9353 | -0.001215 | 0.15638 | 0.008644 | 0.61354 | 0.61354 | 0.010475
2019-03-21 11:48 | 1553183300.3594 | 187 | 0.20775 | 0.001984 | 3e-05 | 100.65 | 100.7317 | 118.7888 | 21.0599 | 0.002778 | 0.002523 | 172.2901 | 0.26243 | 8.3996 | 0.003304 | 0.005248 | 0.005048 | 6.3e-05 | 0.000976 | 0.3361 | 1.2746 | 2.5479 | 99.59 | 20.9 | 37.314 | 88.4521 | 101.145 | 82.3193 | 0.009272 | 0.99318 | 0.023759 | 23.9221 | -0.007254 | 0.15612 | 0.008793 | 0.6252 | 0.6252 | 0.010263
2019-03-21 11:48 | 1553183301.3598 | 188 | 0.20777 | 0.001978 | 3.1e-05 | 100.7165 | 100.637 | 118.7024 | 21.059 | 0.002771 | 0.002538 | 172.2457 | 0.26265 | 8.4016 | 0.003327 | 0.004528 | 0.004475 | 6.4e-05 | 0.000979 | 0.30685 | 1.2847 | 2.548 | 99.59 | 20.9 | 37.316 | 88.2942 | 101.7 | 82.6896 | 0.009272 | 0.99215 | 0.023738 | 23.9259 | 0.00706 | 0.15625 | 0.008722 | 0.6195 | 0.6195 | 0.010366
2019-03-21 11:48 | 1553183302.3605 | 189 | 0.20777 | 0.001974 | 3e-05 | 100.6113 | 100.6316 | 118.382 | 21.0599 | 0.002796 | 0.002538 | 172.2004 | 0.26225 | 8.3996 | 0.003353 | 0.00486 | 0.00479 | 6.5e-05 | 0.000978 | 0.34612 | 1.2773 | 2.5479 | 99.59 | 20.9 | 37.316 | 88.8881 | 102.3371 | 82.6835 | 0.009272 | 0.99216 | 0.023706 | 23.8934 | -0.001801 | 0.15601 | 0.008701 | 0.61727 | 0.61727 | 0.010388
2019-03-21 11:48 | 1553183303.3605 | 190 | 0.20779 | 0.001971 | 3e-05 | 100.7186 | 100.6803 | 118.1604 | 21.0599 | 0.002757 | 0.002548 | 172.157 | 0.26245 | 8.4027 | 0.003279 | 0.004755 | 0.004579 | 6.7e-05 | 0.000973 | 0.28241 | 1.3253 | 2.548 | 99.59 | 20.9 | 37.317 | 87.9569 | 100.5359 | 82.9136 | 0.009272 | 0.99152 | 0.023676 | 23.8787 | 0.003402 | 0.15613 | 0.008614 | 0.61029 | 0.61029 | 0.010515
2019-03-21 11:48 | 1553183304.3599 | 191 | 0.20778 | 0.001964 | 3.1e-05 | 100.6432 | 100.6493 | 117.995 | 21.06 | 0.002763 | 0.002577 | 172.1187 | 0.26229 | 8.4026 | 0.003329 | 0.004162 | 0.004064 | 6.5e-05 | 0.000972 | 0.32189 | 1.2947 | 2.548 | 99.59 | 20.9 | 37.318 | 88.0906 | 101.7559 | 83.6275 | 0.009273 | 0.98954 | 0.023636 | 23.8859 | -0.000541 | 0.15604 | 0.008629 | 0.61032 | 0.61032 | 0.010508
2019-03-21 11:48 | 1553183305.3596 | 192 | 0.20778 | 0.001963 | 3.1e-05 | 100.6599 | 100.5784 | 117.7476 | 21.0599 | 0.002734 | 0.002514 | 172.0647 | 0.26292 | 8.4018 | 0.003301 | 0.004255 | 0.004141 | 6.2e-05 | 0.00097 | 0.33229 | 1.2966 | 2.5479 | 99.59 | 20.9 | 37.319 | 87.3923 | 101.0829 | 82.1017 | 0.009273 | 0.99379 | 0.023662 | 23.8098 | 0.00724 | 0.15641 | 0.008673 | 0.61407 | 0.61407 | 0.010469
2019-03-21 11:48 | 1553183306.3606 | 193 | 0.20777 | 0.001961 | 3e-05 | 100.6837 | 100.6387 | 117.7748 | 21.0596 | 0.002757 | 0.002516 | 172.0053 | 0.26317 | 8.4018 | 0.003243 | 0.005063 | 0.004962 | 6.6e-05 | 0.000973 | 0.34971 | 1.3116 | 2.5479 | 99.59 | 20.9 | 37.32 | 87.9584 | 99.6585 | 82.1568 | 0.009273 | 0.99364 | 0.023663 | 23.8144 | 0.003998 | 0.15656 | 0.008702 | 0.61618 | 0.61618 | 0.010443
2019-03-21 11:48 | 1553183307.3602 | 194 | 0.20779 | 0.001956 | 3e-05 | 100.6441 | 100.5417 | 117.7877 | 21.0604 | 0.002814 | 0.002543 | 171.9692 | 0.26284 | 8.4026 | 0.003286 | 0.004527 | 0.004422 | 6.6e-05 | 0.000973 | 0.31868 | 1.3177 | 2.5479 | 99.59 | 20.9 | 37.321 | 89.3189 | 100.699 | 82.8119 | 0.009273 | 0.99181 | 0.023642 | 23.8376 | 0.009108 | 0.15637 | 0.008628 | 0.61042 | 0.61042 | 0.010528
2019-03-21 11:48 | 1553183308.3608 | 195 | 0.20777 | 0.001952 | 3.1e-05 | 100.6063 | 100.681 | 117.5757 | 21.0594 | 0.002767 | 0.00253 | 171.9272 | 0.26286 | 8.4026 | 0.003323 | 0.004368 | 0.004295 | 6e-05 | 0.000968 | 0.33987 | 1.2944 | 2.5479 | 99.59 | 20.9 | 37.321 | 88.1955 | 101.6029 | 82.4866 | 0.009273 | 0.99271 | 0.023632 | 23.8053 | -0.00664 | 0.15637 | 0.008721 | 0.61673 | 0.61673 | 0.010421
2019-03-21 11:48 | 1553183309.3598 | 196 | 0.20777 | 0.001948 | 3e-05 | 100.6176 | 100.6418 | 117.4966 | 21.0599 | 0.00278 | 0.002544 | 171.8837 | 0.26208 | 8.4029 | 0.003201 | 0.005378 | 0.005237 | 6.3e-05 | 0.000967 | 0.30237 | 1.2738 | 2.5479 | 99.59 | 20.9 | 37.323 | 88.5123 | 98.6462 | 82.817 | 0.009274 | 0.99179 | 0.023613 | 23.8083 | -0.002153 | 0.15591 | 0.008719 | 0.61606 | 0.61606 | 0.010401
2019-03-21 11:48 | 1553183310.3599 | 197 | 0.20779 | 0.001942 | 3e-05 | 100.606 | 100.6111 | 117.3518 | 21.0605 | 0.002766 | 0.002557 | 171.8494 | 0.26249 | 8.4001 | 0.003284 | 0.004737 | 0.004704 | 6.1e-05 | 0.000967 | 0.31669 | 1.297 | 2.548 | 99.59 | 20.9 | 37.324 | 88.159 | 100.6707 | 83.1296 | 0.009274 | 0.99092 | 0.023588 | 23.8041 | -0.000461 | 0.15616 | 0.008622 | 0.60855 | 0.60855 | 0.010547
2019-03-21 11:49 | 1553183311.36 | 198 | 0.2078 | 0.001943 | 3.1e-05 | 100.6015 | 100.6391 | 117.3739 | 21.0602 | 0.002789 | 0.002557 | 171.828 | 0.26286 | 8.4016 | 0.003303 | 0.004861 | 0.00476 | 6.3e-05 | 0.000965 | 0.32637 | 1.2742 | 2.5479 | 99.59 | 20.9 | 37.324 | 88.7086 | 101.1279 | 83.1471 | 0.009274 | 0.99087 | 0.02359 | 23.8069 | -0.003341 | 0.15637 | 0.008591 | 0.60638 | 0.60638 | 0.010599
2019-03-21 11:49 | 1553183312.3601 | 199 | 0.20777 | 0.001935 | 3.1e-05 | 100.6185 | 100.6588 | 117.6544 | 21.0606 | 0.00279 | 0.002532 | 171.8102 | 0.26333 | 8.4029 | 0.003265 | 0.004919 | 0.004853 | 6.3e-05 | 0.000963 | 0.32444 | 1.2645 | 2.5479 | 99.59 | 20.9 | 37.323 | 88.7388 | 100.2118 | 82.5337 | 0.009274 | 0.99258 | 0.023638 | 23.8148 | -0.003587 | 0.15666 | 0.008727 | 0.61726 | 0.61726 | 0.010431
2019-03-21 11:49 | 1553183313.3601 | 200 | 0.20779 | 0.00194 | 3.1e-05 | 100.552 | 100.6782 | 117.7912 | 21.0599 | 0.00277 | 0.002511 | 171.7711 | 0.26259 | 8.4048 | 0.003249 | 0.004612 | 0.004442 | 6.4e-05 | 0.000961 | 0.35258 | 1.3109 | 2.5479 | 99.59 | 20.89 | 37.323 | 88.2544 | 99.8171 | 82.0339 | 0.009268 | 0.99398 | 0.023669 | 23.8119 | -0.01122 | 0.15621 | 0.008646 | 0.6123 | 0.6123 | 0.010486
2019-03-21 11:49 | 1553183314.3601 | 201 | 0.2078 | 0.001934 | 3.1e-05 | 100.5961 | 100.5836 | 117.8431 | 21.0608 | 0.002765 | 0.002511 | 171.7389 | 0.26338 | 8.4039 | 0.003266 | 0.005325 | 0.005275 | 6.3e-05 | 0.000956 | 0.3314 | 1.3075 | 2.548 | 99.59 | 20.89 | 37.322 | 88.1516 | 100.2339 | 82.0418 | 0.009268 | 0.99396 | 0.023673 | 23.8174 | 0.001118 | 0.15669 | 0.008595 | 0.60886 | 0.60886 | 0.010577
2019-03-21 11:49 | 1553183315.3595 | 202 | 0.2078 | 0.001928 | 3e-05 | 100.5848 | 100.7038 | 117.9022 | 21.0612 | 0.002826 | 0.002504 | 171.6999 | 0.26115 | 8.4022 | 0.003301 | 0.005066 | 0.005022 | 6.3e-05 | 0.000961 | 0.28816 | 1.293 | 2.548 | 99.59 | 20.89 | 37.323 | 89.6088 | 101.0802 | 81.8737 | 0.009268 | 0.99443 | 0.023685 | 23.8178 | -0.010574 | 0.15536 | 0.008593 | 0.60901 | 0.60901 | 0.010484
2019-03-21 11:49 | 1553183316.3592 | 203 | 0.20782 | 0.001929 | 2.9e-05 | 100.5279 | 100.6636 | 117.8417 | 21.0607 | 0.002802 | 0.002532 | 171.645 | 0.26355 | 8.4023 | 0.003387 | 0.005202 | 0.005088 | 6.3e-05 | 0.000954 | 0.31996 | 1.2942 | 2.548 | 99.59 | 20.89 | 37.323 | 89.034 | 103.1674 | 82.5306 | 0.009268 | 0.99259 | 0.023657 | 23.8337 | -0.012066 | 0.15679 | 0.008463 | 0.59915 | 0.59915 | 0.010755
2019-03-21 11:49 | 1553183317.3592 | 204 | 0.20787 | 0.001923 | 2.9e-05 | 100.5252 | 100.5699 | 118.0163 | 21.0611 | 0.002852 | 0.002577 | 171.6043 | 0.26289 | 8.4032 | 0.003358 | 0.003847 | 0.003821 | 6.2e-05 | 0.000958 | 0.32764 | 1.3474 | 2.5479 | 99.59 | 20.89 | 37.323 | 90.2285 | 102.4639 | 83.6216 | 0.009268 | 0.98956 | 0.023638 | 23.8879 | -0.003978 | 0.15639 | 0.008169 | 0.57788 | 0.57788 | 0.011123
2019-03-21 11:49 | 1553183318.3597 | 205 | 0.20789 | 0.001914 | 2.8e-05 | 100.5505 | 100.465 | 118.2611 | 21.0617 | 0.00274 | 0.002545 | 171.5779 | 0.26363 | 8.4042 | 0.003408 | 0.004865 | 0.004817 | 6.3e-05 | 0.000958 | 0.35081 | 1.3366 | 2.5478 | 99.59 | 20.89 | 37.322 | 87.5362 | 103.6698 | 82.8576 | 0.009268 | 0.99168 | 0.023688 | 23.887 | 0.00761 | 0.15684 | 0.008055 | 0.57102 | 0.57102 | 0.011288
2019-03-21 11:49 | 1553183319.3595 | 206 | 0.20788 | 0.001913 | 2.8e-05 | 100.5687 | 100.3789 | 118.1597 | 21.0608 | 0.0028 | 0.002525 | 171.529 | 0.26355 | 8.4055 | 0.003349 | 0.004914 | 0.004845 | 6.2e-05 | 0.000957 | 0.32493 | 1.3595 | 2.5481 | 99.59 | 20.89 | 37.321 | 88.9802 | 102.2459 | 82.3661 | 0.009268 | 0.99305 | 0.023694 | 23.8603 | 0.016901 | 0.15679 | 0.008117 | 0.57556 | 0.57556 | 0.011196
2019-03-21 11:49 | 1553183320.3604 | 207 | 0.20789 | 0.001904 | 2.8e-05 | 100.5782 | 100.4538 | 118.3452 | 21.0617 | 0.00276 | 0.002495 | 171.4735 | 0.26307 | 8.4037 | 0.003408 | 0.005731 | 0.005616 | 6.5e-05 | 0.000956 | 0.31902 | 1.3395 | 2.548 | 99.59 | 20.89 | 37.321 | 88.0164 | 103.6618 | 81.6424 | 0.009268 | 0.99508 | 0.023737 | 23.8547 | 0.011075 | 0.1565 | 0.008066 | 0.57299 | 0.57299 | 0.011225
2019-03-21 11:49 | 1553183321.3599 | 208 | 0.2079 | 0.001903 | 2.7e-05 | 100.4787 | 100.23 | 118.5062 | 21.0609 | 0.002763 | 0.002508 | 171.4114 | 0.26309 | 8.4018 | 0.003408 | 0.005024 | 0.004959 | 6.2e-05 | 0.000955 | 0.32278 | 1.3314 | 2.5481 | 99.59 | 20.89 | 37.32 | 88.1009 | 103.6596 | 81.9506 | 0.009267 | 0.99421 | 0.023743 | 23.8813 | 0.022167 | 0.15651 | 0.008028 | 0.57045 | 0.57045 | 0.011277
2019-03-21 11:49 | 1553183322.3599 | 209 | 0.2079 | 0.001895 | 2.7e-05 | 100.4978 | 100.4044 | 118.5057 | 21.0611 | 0.002743 | 0.00253 | 171.3637 | 0.26245 | 8.402 | 0.003357 | 0.005012 | 0.004975 | 5.9e-05 | 0.000957 | 0.29806 | 1.2911 | 2.5479 | 99.59 | 20.89 | 37.32 | 87.6021 | 102.4355 | 82.4921 | 0.009267 | 0.9927 | 0.023725 | 23.8994 | 0.008317 | 0.15613 | 0.008015 | 0.56913 | 0.56913 | 0.011275
2019-03-21 11:49 | 1553183323.3596 | 210 | 0.20791 | 0.001886 | 2.8e-05 | 100.4606 | 100.5577 | 118.5855 | 21.0618 | 0.002736 | 0.002537 | 171.3071 | 0.26284 | 8.4033 | 0.003345 | 0.005084 | 0.004956 | 6.2e-05 | 0.000961 | 0.28727 | 1.3256 | 2.5479 | 99.59 | 20.89 | 37.319 | 87.4481 | 102.13 | 82.6527 | 0.009267 | 0.99225 | 0.023728 | 23.9129 | -0.008642 | 0.15636 | 0.007986 | 0.56707 | 0.56707 | 0.011333
2019-03-21 11:49 | 1553183324.361 | 211 | 0.20793 | 0.00188 | 2.8e-05 | 100.5204 | 100.5669 | 118.6156 | 21.0618 | 0.002753 | 0.002499 | 171.2613 | 0.26245 | 8.4042 | 0.003377 | 0.004658 | 0.004591 | 5.8e-05 | 0.000956 | 0.3178 | 1.3244 | 2.548 | 99.59 | 20.89 | 37.32 | 87.8557 | 102.9202 | 81.7331 | 0.009267 | 0.99482 | 0.023761 | 23.885 | -0.004129 | 0.15613 | 0.007869 | 0.55954 | 0.55954 | 0.011469
2019-03-21 11:49 | 1553183325.3604 | 212 | 0.20794 | 0.001871 | 2.8e-05 | 100.4191 | 100.4269 | 118.7468 | 21.0621 | 0.002715 | 0.002563 | 171.2285 | 0.26283 | 8.4011 | 0.003337 | 0.0056 | 0.005576 | 6.1e-05 | 0.000957 | 0.29767 | 1.2785 | 2.548 | 99.59 | 20.89 | 37.322 | 86.9408 | 101.9512 | 83.2841 | 0.009268 | 0.99049 | 0.023723 | 23.9503 | -0.000689 | 0.15636 | 0.007801 | 0.55377 | 0.55377 | 0.011605
2019-03-21 11:49 | 1553183326.3603 | 213 | 0.20795 | 0.001865 | 2.8e-05 | 100.4814 | 100.574 | 118.7719 | 21.0609 | 0.002745 | 0.00252 | 171.2025 | 0.26218 | 8.4017 | 0.003345 | 0.004525 | 0.004429 | 6.3e-05 | 0.000956 | 0.30254 | 1.3172 | 2.5479 | 99.59 | 20.89 | 37.323 | 87.6573 | 102.1274 | 82.2469 | 0.009268 | 0.99338 | 0.02376 | 23.918 | -0.008239 | 0.15597 | 0.007765 | 0.55209 | 0.55209 | 0.011611
2019-03-21 11:49 | 1553183327.3602 | 214 | 0.20798 | 0.001861 | 2.8e-05 | 100.4001 | 100.588 | 118.767 | 21.0622 | 0.002747 | 0.002499 | 171.183 | 0.26314 | 8.4022 | 0.003317 | 0.004918 | 0.004816 | 5.8e-05 | 0.000953 | 0.31238 | 1.2781 | 2.5481 | 99.59 | 20.89 | 37.325 | 87.7187 | 101.466 | 81.742 | 0.009269 | 0.9948 | 0.023776 | 23.9005 | -0.016724 | 0.15654 | 0.007579 | 0.53922 | 0.53922 | 0.011932
2019-03-21 11:49 | 1553183328.3598 | 215 | 0.20796 | 0.001858 | 2.8e-05 | 100.4504 | 100.4851 | 118.7003 | 21.0619 | 0.002743 | 0.002437 | 171.1456 | 0.26269 | 8.4011 | 0.003278 | 0.004714 | 0.00457 | 5.6e-05 | 0.000955 | 0.29789 | 1.296 | 2.548 | 99.59 | 20.89 | 37.326 | 87.6193 | 100.5119 | 80.2539 | 0.009269 | 0.99899 | 0.023819 | 23.8437 | -0.003089 | 0.15627 | 0.007738 | 0.55159 | 0.55159 | 0.011644
2019-03-21 11:49 | 1553183329.3589 | 216 | 0.20793 | 0.001848 | 2.8e-05 | 100.4397 | 100.3983 | 118.817 | 21.0621 | 0.002753 | 0.002509 | 171.1101 | 0.26292 | 8.4031 | 0.003267 | 0.004423 | 0.004558 | 5.8e-05 | 0.000956 | 0.32875 | 1.3393 | 2.5481 | 99.59 | 20.9 | 37.328 | 87.8419 | 100.2454 | 81.9837 | 0.009275 | 0.99412 | 0.023773 | 23.9137 | 0.00369 | 0.15641 | 0.007887 | 0.56114 | 0.56114 | 0.011456
2019-03-21 11:49 | 1553183330.3603 | 217 | 0.20797 | 0.001839 | 2.8e-05 | 100.4516 | 100.51 | 118.6064 | 21.0621 | 0.002715 | 0.002488 | 171.0699 | 0.26234 | 8.4004 | 0.003327 | 0.005108 | 0.004974 | 5.8e-05 | 0.000954 | 0.28799 | 1.3372 | 2.5479 | 99.59 | 20.89 | 37.332 | 86.9463 | 101.71 | 81.4822 | 0.00927 | 0.99552 | 0.023769 | 23.8756 | -0.005196 | 0.15606 | 0.007656 | 0.54454 | 0.54454 | 0.011779
2019-03-21 11:49 | 1553183331.3589 | 218 | 0.208 | 0.001836 | 2.8e-05 | 100.4224 | 100.5355 | 118.7181 | 21.0618 | 0.002732 | 0.002482 | 171.0512 | 0.26325 | 8.4041 | 0.00338 | 0.004854 | 0.004771 | 5.6e-05 | 0.000958 | 0.34202 | 1.3055 | 2.548 | 99.59 | 20.89 | 37.333 | 87.3374 | 102.9869 | 81.3472 | 0.009271 | 0.99591 | 0.023784 | 23.8823 | -0.010069 | 0.15661 | 0.007521 | 0.53524 | 0.53524 | 0.012025
2019-03-21 11:49 | 1553183332.36 | 219 | 0.20801 | 0.001828 | 2.9e-05 | 100.378 | 100.5025 | 118.7935 | 21.0618 | 0.002765 | 0.002491 | 171.0047 | 0.26278 | 8.4033 | 0.003315 | 0.004317 | 0.004229 | 5.9e-05 | 0.000957 | 0.29966 | 1.3079 | 2.548 | 99.59 | 20.89 | 37.337 | 88.1415 | 101.4254 | 81.5522 | 0.009272 | 0.99533 | 0.023785 | 23.8968 | -0.011084 | 0.15633 | 0.007464 | 0.53113 | 0.53113 | 0.012097
2019-03-21 11:49 | 1553183333.3592 | 220 | 0.20804 | 0.001818 | 3e-05 | 100.4528 | 100.4674 | 118.5965 | 21.063 | 0.002683 | 0.0025 | 170.9677 | 0.26265 | 8.4012 | 0.003329 | 0.005019 | 0.004872 | 5.7e-05 | 0.00096 | 0.30414 | 1.3306 | 2.548 | 99.59 | 20.89 | 37.339 | 86.1667 | 101.754 | 81.7601 | 0.009272 | 0.99475 | 0.023758 | 23.884 | -0.0013 | 0.15625 | 0.007275 | 0.517 | 0.517 | 0.012421
2019-03-21 11:49 | 1553183334.3596 | 221 | 0.20809 | 0.001819 | 3e-05 | 100.3977 | 100.481 | 118.6944 | 21.0626 | 0.002681 | 0.002478 | 170.9452 | 0.26252 | 8.4023 | 0.003269 | 0.004557 | 0.004442 | 5.5e-05 | 0.000959 | 0.31708 | 1.3125 | 2.548 | 99.59 | 20.89 | 37.341 | 86.11 | 100.2991 | 81.2458 | 0.009273 | 0.99619 | 0.023786 | 23.8765 | -0.007422 | 0.15617 | 0.006972 | 0.49595 | 0.49595 | 0.012942
2019-03-21 11:49 | 1553183335.3592 | 222 | 0.20802 | 0.001804 | 3.1e-05 | 100.437 | 100.466 | 118.8451 | 21.0627 | 0.002724 | 0.002478 | 170.915 | 0.26242 | 8.4038 | 0.003264 | 0.004899 | 0.004907 | 5.8e-05 | 0.00096 | 0.31509 | 1.2988 | 2.5479 | 99.58 | 20.89 | 37.345 | 87.1616 | 100.1729 | 81.249 | 0.009274 | 0.99618 | 0.023801 | 23.8918 | -0.002585 | 0.15612 | 0.007422 | 0.52835 | 0.52835 | 0.012144
2019-03-21 11:49 | 1553183336.3603 | 223 | 0.20805 | 0.001799 | 3.1e-05 | 100.4275 | 100.3519 | 119.0341 | 21.0632 | 0.002717 | 0.002517 | 170.8765 | 0.2634 | 8.4016 | 0.003282 | 0.00447 | 0.004337 | 5.8e-05 | 0.000953 | 0.31365 | 1.2856 | 2.548 | 99.58 | 20.89 | 37.347 | 86.9864 | 100.6175 | 82.1883 | 0.009275 | 0.99355 | 0.023788 | 23.9424 | 0.00673 | 0.1567 | 0.007229 | 0.51427 | 0.51427 | 0.012523
2019-03-21 11:49 | 1553183337.3598 | 224 | 0.20809 | 0.001796 | 3.1e-05 | 100.4323 | 100.5832 | 119.2035 | 21.0624 | 0.002697 | 0.002464 | 170.8353 | 0.26212 | 8.4014 | 0.003221 | 0.004462 | 0.004375 | 5.8e-05 | 0.000948 | 0.34075 | 1.3104 | 2.5481 | 99.58 | 20.89 | 37.349 | 86.4955 | 99.145 | 80.9023 | 0.009275 | 0.99716 | 0.023848 | 23.9161 | -0.013434 | 0.15593 | 0.007017 | 0.50043 | 0.50043 | 0.012807
2019-03-21 11:49 | 1553183338.3594 | 225 | 0.20807 | 0.001793 | 3.2e-05 | 100.401 | 100.4804 | 119.4617 | 21.063 | 0.002759 | 0.002548 | 170.8019 | 0.26245 | 8.402 | 0.003214 | 0.004335 | 0.004336 | 5.7e-05 | 0.000949 | 0.32211 | 1.2837 | 2.5478 | 99.58 | 20.89 | 37.35 | 87.9983 | 98.977 | 82.9354 | 0.009276 | 0.99146 | 0.023806 | 24.0106 | -0.00707 | 0.15613 | 0.00711 | 0.50614 | 0.50614 | 0.012678
2019-03-21 11:49 | 1553183339.3594 | 226 | 0.20807 | 0.001784 | 3.2e-05 | 100.4349 | 100.621 | 119.6411 | 21.0635 | 0.002649 | 0.002475 | 170.7658 | 0.26342 | 8.4009 | 0.003249 | 0.004843 | 0.004756 | 5.8e-05 | 0.000951 | 0.34667 | 1.2707 | 2.5478 | 99.58 | 20.89 | 37.351 | 85.3624 | 99.8223 | 81.1591 | 0.009276 | 0.99643 | 0.023883 | 23.9686 | -0.016557 | 0.15671 | 0.007145 | 0.51028 | 0.51028 | 0.012622
2019-03-21 11:49 | 1553183340.3598 | 227 | 0.20809 | 0.001774 | 3.2e-05 | 100.4716 | 100.5897 | 119.9955 | 21.0643 | 0.002738 | 0.002513 | 170.7287 | 0.2622 | 8.4024 | 0.003225 | 0.005049 | 0.004899 | 5.9e-05 | 0.000952 | 0.31421 | 1.2634 | 2.5479 | 99.58 | 20.88 | 37.353 | 87.5068 | 99.2473 | 82.0928 | 0.009271 | 0.99381 | 0.023887 | 24.0357 | -0.010508 | 0.15598 | 0.007039 | 0.50277 | 0.50277 | 0.012751
2019-03-21 11:49 | 1553183341.3606 | 228 | 0.20804 | 0.00177 | 3.3e-05 | 100.4588 | 100.4494 | 120.0277 | 21.0636 | 0.002696 | 0.002507 | 170.6752 | 0.26246 | 8.401 | 0.003377 | 0.00411 | 0.004023 | 5.7e-05 | 0.000949 | 0.32565 | 1.3101 | 2.5481 | 99.58 | 20.88 | 37.355 | 86.4918 | 102.9211 | 81.9313 | 0.009271 | 0.99427 | 0.023896 | 24.0334 | 0.000838 | 0.15614 | 0.007356 | 0.52562 | 0.52562 | 0.012209
2019-03-21 11:49 | 1553183342.3596 | 229 | 0.20802 | 0.001764 | 3.2e-05 | 100.493 | 100.5076 | 119.8066 | 21.0633 | 0.002694 | 0.00246 | 170.6497 | 0.263 | 8.3995 | 0.003303 | 0.004206 | 0.0041 | 6.1e-05 | 0.000952 | 0.34938 | 1.2623 | 2.5479 | 99.58 | 20.88 | 37.355 | 86.4355 | 101.1232 | 80.8122 | 0.009271 | 0.99741 | 0.023911 | 23.9734 | -0.001297 | 0.15646 | 0.007475 | 0.53455 | 0.53455 | 0.01203
2019-03-21 11:49 | 1553183343.359 | 230 | 0.20808 | 0.001755 | 3.3e-05 | 100.5023 | 100.5008 | 119.9072 | 21.0633 | 0.002653 | 0.002493 | 170.6235 | 0.2629 | 8.4024 | 0.003222 | 0.005058 | 0.004983 | 5.8e-05 | 0.000949 | 0.31205 | 1.2748 | 2.5479 | 99.58 | 20.88 | 37.359 | 85.4583 | 99.1554 | 81.5919 | 0.009272 | 0.99522 | 0.023895 | 24.0099 | 0.000132 | 0.1564 | 0.007099 | 0.50723 | 0.50723 | 0.012673
2019-03-21 11:49 | 1553183344.3596 | 231 | 0.20807 | 0.001749 | 3.3e-05 | 100.3846 | 100.4994 | 119.7638 | 21.0639 | 0.002637 | 0.002492 | 170.5928 | 0.26223 | 8.4042 | 0.003277 | 0.004261 | 0.004216 | 5.8e-05 | 0.000947 | 0.3079 | 1.3092 | 2.5478 | 99.58 | 20.88 | 37.36 | 85.0747 | 100.4849 | 81.5799 | 0.009273 | 0.99525 | 0.023881 | 23.9951 | -0.010227 | 0.156 | 0.007181 | 0.51278 | 0.51278 | 0.012504
2019-03-21 11:49 | 1553183345.3588 | 232 | 0.20809 | 0.001747 | 3.3e-05 | 100.4174 | 100.418 | 119.6434 | 21.0647 | 0.002661 | 0.002482 | 170.5703 | 0.26264 | 8.4008 | 0.003253 | 0.004693 | 0.004637 | 6e-05 | 0.000946 | 0.32101 | 1.3072 | 2.5481 | 99.58 | 20.88 | 37.362 | 85.6477 | 99.9159 | 81.3343 | 0.009273 | 0.99594 | 0.023877 | 23.9748 | -6e-05 | 0.15624 | 0.007081 | 0.50555 | 0.50555 | 0.012702
2019-03-21 11:49 | 1553183346.3602 | 233 | 0.20808 | 0.00174 | 3.3e-05 | 100.375 | 100.2992 | 119.6847 | 21.0639 | 0.002736 | 0.002497 | 170.5297 | 0.26175 | 8.4055 | 0.003158 | 0.004659 | 0.004565 | 5.9e-05 | 0.000951 | 0.32388 | 1.2803 | 2.5479 | 99.58 | 20.88 | 37.364 | 87.4386 | 97.6121 | 81.7002 | 0.009273 | 0.99491 | 0.023869 | 23.9913 | 0.006765 | 0.15572 | 0.007134 | 0.50914 | 0.50914 | 0.01257
2019-03-21 11:49 | 1553183347.3589 | 234 | 0.20809 | 0.001731 | 3.3e-05 | 100.4519 | 100.5052 | 119.4627 | 21.064 | 0.002717 | 0.00247 | 170.4534 | 0.2626 | 8.4018 | 0.003218 | 0.004958 | 0.004849 | 5.9e-05 | 0.000947 | 0.33157 | 1.2865 | 2.548 | 99.58 | 20.88 | 37.364 | 86.9864 | 99.0806 | 81.0397 | 0.009273 | 0.99677 | 0.023869 | 23.9467 | -0.004745 | 0.15622 | 0.007104 | 0.50702 | 0.50702 | 0.012664
2019-03-21 11:49 | 1553183348.3589 | 235 | 0.20809 | 0.001729 | 3.4e-05 | 100.4194 | 100.5079 | 119.3446 | 21.065 | 0.002735 | 0.002459 | 170.4166 | 0.2623 | 8.4035 | 0.003212 | 0.005172 | 0.005135 | 6.1e-05 | 0.000944 | 0.30989 | 1.3065 | 2.5479 | 99.58 | 20.88 | 37.366 | 87.4192 | 98.9202 | 80.7934 | 0.009274 | 0.99746 | 0.023866 | 23.9265 | -0.007879 | 0.15604 | 0.007086 | 0.50557 | 0.50557 | 0.012685
2019-03-21 11:49 | 1553183349.3604 | 236 | 0.20813 | 0.001722 | 3.5e-05 | 100.4906 | 100.4974 | 119.2125 | 21.0649 | 0.002711 | 0.002457 | 170.3781 | 0.26204 | 8.4022 | 0.003191 | 0.004462 | 0.004384 | 6.3e-05 | 0.000942 | 0.2891 | 1.331 | 2.5479 | 99.58 | 20.88 | 37.367 | 86.8367 | 98.4118 | 80.7404 | 0.009274 | 0.99761 | 0.023854 | 23.9115 | -0.000599 | 0.15589 | 0.006887 | 0.49104 | 0.49104 | 0.013048
2019-03-21 11:49 | 1553183350.3604 | 237 | 0.20812 | 0.001713 | 3.5e-05 | 100.5133 | 100.5754 | 119.2176 | 21.0643 | 0.002664 | 0.002437 | 170.3571 | 0.26175 | 8.4045 | 0.003243 | 0.004032 | 0.004013 | 6.1e-05 | 0.000938 | 0.31758 | 1.3185 | 2.5479 | 99.59 | 20.88 | 37.369 | 85.7253 | 99.6708 | 80.2493 | 0.009274 | 0.999 | 0.023871 | 23.8954 | -0.005523 | 0.15572 | 0.006949 | 0.49584 | 0.49584 | 0.012907
2019-03-21 11:49 | 1553183351.3601 | 238 | 0.20811 | 0.001705 | 3.6e-05 | 100.4451 | 100.4838 | 118.8712 | 21.0656 | 0.002698 | 0.002486 | 170.3129 | 0.26216 | 8.4052 | 0.003144 | 0.004625 | 0.004527 | 6e-05 | 0.000934 | 0.31094 | 1.3422 | 2.548 | 99.59 | 20.88 | 37.37 | 86.54 | 97.2715 | 81.4326 | 0.009274 | 0.99567 | 0.023797 | 23.9006 | -0.003445 | 0.15596 | 0.006998 | 0.49778 | 0.49778 | 0.012877
2019-03-21 11:49 | 1553183352.3595 | 239 | 0.20809 | 0.001704 | 3.5e-05 | 100.4725 | 100.5774 | 118.8376 | 21.0653 | 0.002689 | 0.002465 | 170.2671 | 0.26206 | 8.4018 | 0.003211 | 0.004868 | 0.004768 | 5.9e-05 | 0.000933 | 0.32278 | 1.2935 | 2.548 | 99.59 | 20.88 | 37.371 | 86.3258 | 98.8922 | 80.9317 | 0.009274 | 0.99707 | 0.02381 | 23.8803 | -0.009338 | 0.1559 | 0.007104 | 0.50562 | 0.50562 | 0.012672
2019-03-21 11:49 | 1553183353.3606 | 240 | 0.2081 | 0.0017 | 3.6e-05 | 100.409 | 100.7048 | 118.7393 | 21.0656 | 0.002644 | 0.002482 | 170.2349 | 0.26169 | 8.4025 | 0.003212 | 0.004744 | 0.0046 | 6e-05 | 0.000936 | 0.32372 | 1.2919 | 2.548 | 99.59 | 20.88 | 37.372 | 85.2457 | 98.9214 | 81.3352 | 0.009275 | 0.99594 | 0.023787 | 23.884 | -0.02631 | 0.15568 | 0.007049 | 0.50116 | 0.50116 | 0.012767
2019-03-21 11:49 | 1553183354.3602 | 241 | 0.20814 | 0.00169 | 3.6e-05 | 100.4609 | 100.4316 | 118.8562 | 21.0653 | 0.00263 | 0.002436 | 170.1877 | 0.26372 | 8.4013 | 0.003197 | 0.004488 | 0.004392 | 5.8e-05 | 0.000935 | 0.3095 | 1.271 | 2.5478 | 99.59 | 20.88 | 37.373 | 84.896 | 98.5715 | 80.2385 | 0.009275 | 0.99903 | 0.023836 | 23.8588 | 0.002601 | 0.15688 | 0.006852 | 0.48814 | 0.48814 | 0.013209
2019-03-21 11:49 | 1553183355.3603 | 242 | 0.20809 | 0.001687 | 3.6e-05 | 100.482 | 100.3339 | 118.9085 | 21.0653 | 0.002662 | 0.002409 | 170.1356 | 0.26337 | 8.4007 | 0.003272 | 0.005562 | 0.005569 | 5.9e-05 | 0.000934 | 0.32173 | 1.2794 | 2.548 | 99.59 | 20.88 | 37.375 | 85.6717 | 100.3733 | 79.593 | 0.009275 | 1.0009 | 0.023863 | 23.8422 | 0.013194 | 0.15668 | 0.007148 | 0.50984 | 0.50984 | 0.01263
2019-03-21 11:49 | 1553183356.3597 | 243 | 0.20812 | 0.001686 | 3.6e-05 | 100.4516 | 100.7351 | 118.9465 | 21.0666 | 0.002666 | 0.002443 | 170.119 | 0.26324 | 8.4003 | 0.003178 | 0.004162 | 0.004159 | 5.8e-05 | 0.00093 | 0.31686 | 1.2657 | 2.548 | 99.59 | 20.87 | 37.377 | 85.7571 | 98.1015 | 80.4082 | 0.00927 | 0.99855 | 0.023839 | 23.8736 | -0.025208 | 0.1566 | 0.006973 | 0.49685 | 0.49685 | 0.012954
2019-03-21 11:49 | 1553183357.3604 | 244 | 0.20814 | 0.001682 | 3.5e-05 | 100.4367 | 100.6145 | 119.15 | 21.0654 | 0.002648 | 0.002449 | 170.0931 | 0.26252 | 8.4025 | 0.003179 | 0.004244 | 0.004113 | 5.9e-05 | 0.000934 | 0.33157 | 1.2933 | 2.5479 | 99.59 | 20.87 | 37.377 | 85.322 | 98.1223 | 80.5342 | 0.00927 | 0.99819 | 0.023855 | 23.8983 | -0.015822 | 0.15617 | 0.006832 | 0.48712 | 0.48712 | 0.013177
2019-03-21 11:49 | 1553183358.3604 | 245 | 0.20809 | 0.001673 | 3.5e-05 | 100.4918 | 100.4507 | 118.8184 | 21.0656 | 0.002612 | 0.002483 | 170.0548 | 0.26231 | 8.4031 | 0.003137 | 0.004335 | 0.004194 | 5.7e-05 | 0.000936 | 0.29806 | 1.3383 | 2.5479 | 99.59 | 20.87 | 37.378 | 84.4741 | 97.1029 | 81.371 | 0.00927 | 0.99584 | 0.023794 | 23.8932 | 0.003661 | 0.15605 | 0.00716 | 0.50927 | 0.50927 | 0.012594
2019-03-21 11:49 | 1553183359.3598 | 246 | 0.20813 | 0.001667 | 3.5e-05 | 100.4138 | 100.5056 | 118.7646 | 21.0651 | 0.002637 | 0.002441 | 170.0103 | 0.26279 | 8.4011 | 0.003173 | 0.004291 | 0.004271 | 5.9e-05 | 0.000931 | 0.29446 | 1.2886 | 2.5479 | 99.59 | 20.87 | 37.379 | 85.0744 | 97.9763 | 80.3522 | 0.009271 | 0.99871 | 0.023823 | 23.8534 | -0.008172 | 0.15634 | 0.006956 | 0.49532 | 0.49532 | 0.012972
2019-03-21 11:49 | 1553183360.3596 | 247 | 0.20815 | 0.001662 | 3.4e-05 | 100.4194 | 100.6271 | 118.4846 | 21.0653 | 0.002675 | 0.002469 | 169.9617 | 0.2627 | 8.4008 | 0.003086 | 0.00407 | 0.003979 | 5.6e-05 | 0.000933 | 0.31166 | 1.2964 | 2.548 | 99.59 | 20.87 | 37.379 | 85.9714 | 95.8923 | 81.0277 | 0.009271 | 0.9968 | 0.023772 | 23.848 | -0.018482 | 0.15628 | 0.006842 | 0.48622 | 0.48622 | 0.01321
2019-03-21 11:49 | 1553183361.3597 | 248 | 0.20817 | 0.001657 | 3.4e-05 | 100.3825 | 100.5464 | 118.7611 | 21.0659 | 0.002702 | 0.002454 | 169.9516 | 0.26193 | 8.4031 | 0.003141 | 0.004945 | 0.004764 | 5.5e-05 | 0.000934 | 0.33931 | 1.3172 | 2.548 | 99.59 | 20.87 | 37.381 | 86.6262 | 97.2184 | 80.664 | 0.009271 | 0.99783 | 0.023812 | 23.8636 | -0.014595 | 0.15582 | 0.006727 | 0.47876 | 0.47876 | 0.013377
2019-03-21 11:49 | 1553183362.3597 | 249 | 0.20814 | 0.00165 | 3.4e-05 | 100.4826 | 100.5311 | 118.8064 | 21.0656 | 0.002628 | 0.002461 | 169.9261 | 0.26161 | 8.403 | 0.003228 | 0.004835 | 0.004841 | 5.8e-05 | 0.000935 | 0.32637 | 1.3542 | 2.548 | 99.59 | 20.87 | 37.381 | 84.8469 | 99.309 | 80.8401 | 0.009271 | 0.99733 | 0.02381 | 23.8741 | -0.004316 | 0.15563 | 0.006881 | 0.4898 | 0.4898 | 0.013059
2019-03-21 11:49 | 1553183363.3591 | 250 | 0.20817 | 0.001649 | 3.4e-05 | 100.4454 | 100.4218 | 119.1735 | 21.0652 | 0.002684 | 0.002397 | 169.8848 | 0.26188 | 8.402 | 0.003266 | 0.005748 | 0.005599 | 5.6e-05 | 0.000928 | 0.30829 | 1.3219 | 2.548 | 99.59 | 20.87 | 37.382 | 86.2078 | 100.2436 | 79.3005 | 0.009271 | 1.0017 | 0.023899 | 23.8589 | 0.002101 | 0.15579 | 0.006734 | 0.48109 | 0.48109 | 0.01331
2019-03-21 11:49 | 1553183364.3591 | 251 | 0.20818 | 0.001644 | 3.4e-05 | 100.4108 | 100.4923 | 119.3333 | 21.0657 | 0.002668 | 0.002475 | 169.8518 | 0.26256 | 8.4014 | 0.003319 | 0.004759 | 0.004698 | 5.5e-05 | 0.000925 | 0.29607 | 1.3263 | 2.548 | 99.59 | 20.87 | 37.382 | 85.8014 | 101.5103 | 81.1817 | 0.009271 | 0.99637 | 0.023852 | 23.9385 | -0.007255 | 0.1562 | 0.00665 | 0.47408 | 0.47408 | 0.013542
2019-03-21 11:49 | 1553183365.3589 | 252 | 0.20818 | 0.001639 | 3.4e-05 | 100.4424 | 100.4851 | 119.2824 | 21.066 | 0.002633 | 0.002438 | 169.8299 | 0.26097 | 8.4035 | 0.003262 | 0.004402 | 0.004331 | 5.4e-05 | 0.000929 | 0.3126 | 1.3228 | 2.5479 | 99.59 | 20.87 | 37.382 | 84.9661 | 100.1401 | 80.2742 | 0.009271 | 0.99893 | 0.023877 | 23.9027 | -0.003805 | 0.15525 | 0.006684 | 0.47705 | 0.47705 | 0.013375
2019-03-21 11:49 | 1553183366.3599 | 253 | 0.20817 | 0.001631 | 3.3e-05 | 100.4814 | 100.5774 | 119.0606 | 21.0663 | 0.002603 | 0.002452 | 169.767 | 0.2609 | 8.4001 | 0.00322 | 0.004301 | 0.004247 | 5.7e-05 | 0.000931 | 0.323 | 1.2803 | 2.548 | 99.59 | 20.87 | 37.383 | 84.2548 | 99.1182 | 80.6192 | 0.009272 | 0.99795 | 0.023843 | 23.8922 | -0.008542 | 0.15521 | 0.006753 | 0.4814 | 0.4814 | 0.013251
2019-03-21 11:49 | 1553183367.3603 | 254 | 0.20814 | 0.001625 | 3.2e-05 | 100.4936 | 100.5805 | 118.958 | 21.0661 | 0.002566 | 0.002402 | 169.7361 | 0.26124 | 8.4023 | 0.003304 | 0.004042 | 0.003951 | 5.6e-05 | 0.000932 | 0.30735 | 1.3047 | 2.548 | 99.58 | 20.87 | 37.385 | 83.3697 | 101.1574 | 79.4261 | 0.009273 | 1.0013 | 0.023873 | 23.8416 | -0.007727 | 0.15541 | 0.006919 | 0.49391 | 0.49391 | 0.012933
2019-03-21 11:49 | 1553183368.359 | 255 | 0.20819 | 0.001624 | 3.1e-05 | 100.3938 | 100.2634 | 119.3596 | 21.0658 | 0.002697 | 0.00242 | 169.6902 | 0.2622 | 8.4043 | 0.003222 | 0.004658 | 0.004576 | 5.7e-05 | 0.000935 | 0.29502 | 1.3354 | 2.548 | 99.58 | 20.87 | 37.387 | 86.519 | 99.1578 | 79.8406 | 0.009273 | 1.0002 | 0.0239 | 23.8958 | 0.011627 | 0.15598 | 0.006606 | 0.47206 | 0.47206 | 0.013581
2019-03-21 11:49 | 1553183369.3604 | 256 | 0.20817 | 0.00162 | 3.1e-05 | 100.4486 | 100.4892 | 119.7918 | 21.0664 | 0.002621 | 0.002443 | 169.6785 | 0.26231 | 8.3998 | 0.003257 | 0.005264 | 0.005219 | 5.8e-05 | 0.000933 | 0.30453 | 1.3244 | 2.5479 | 99.58 | 20.86 | 37.389 | 84.6933 | 100.0039 | 80.3941 | 0.009268 | 0.99859 | 0.023924 | 23.9578 | -0.003612 | 0.15604 | 0.006775 | 0.4847 | 0.4847 | 0.013232
2019-03-21 11:49 | 1553183370.3605 | 257 | 0.20821 | 0.001614 | 3.1e-05 | 100.3804 | 100.4334 | 119.9917 | 21.0651 | 0.002593 | 0.002437 | 169.6568 | 0.26128 | 8.4026 | 0.003293 | 0.004919 | 0.004743 | 5.6e-05 | 0.000934 | 0.31797 | 1.3059 | 2.5479 | 99.58 | 20.86 | 37.39 | 83.9968 | 100.8927 | 80.2578 | 0.009269 | 0.99897 | 0.023949 | 23.9731 | -0.004716 | 0.15543 | 0.006517 | 0.46668 | 0.46668 | 0.013689
2019-03-21 11:50 | 1553183371.3591 | 258 | 0.20818 | 0.001609 | 3.1e-05 | 100.4001 | 100.463 | 120.0424 | 21.0658 | 0.002622 | 0.002437 | 169.6153 | 0.26217 | 8.4041 | 0.003297 | 0.004549 | 0.004477 | 5.6e-05 | 0.000933 | 0.32676 | 1.2816 | 2.5479 | 99.58 | 20.86 | 37.393 | 84.7114 | 100.9763 | 80.2639 | 0.009269 | 0.99896 | 0.023953 | 23.9784 | -0.005603 | 0.15597 | 0.006705 | 0.48027 | 0.48027 | 0.013347
2019-03-21 11:50 | 1553183372.3592 | 259 | 0.20823 | 0.001613 | 3.2e-05 | 100.415 | 100.3424 | 120.1105 | 21.0661 | 0.002707 | 0.002501 | 169.5633 | 0.26278 | 8.402 | 0.003363 | 0.005083 | 0.00494 | 5.9e-05 | 0.000936 | 0.3095 | 1.3024 | 2.5479 | 99.58 | 20.87 | 37.394 | 86.7471 | 102.5857 | 81.795 | 0.009275 | 0.99465 | 0.023908 | 24.0371 | 0.006465 | 0.15633 | 0.006403 | 0.45764 | 0.45764 | 0.01404
2019-03-21 11:50 | 1553183373.361 | 260 | 0.20819 | 0.001605 | 3.3e-05 | 100.4141 | 100.464 | 119.9087 | 21.0659 | 0.002579 | 0.002489 | 169.5211 | 0.26221 | 8.4028 | 0.003298 | 0.004359 | 0.004268 | 5.7e-05 | 0.000934 | 0.32372 | 1.3211 | 2.5481 | 99.58 | 20.86 | 37.395 | 83.6597 | 101.0118 | 81.5031 | 0.00927 | 0.99547 | 0.023898 | 24.007 | -0.004446 | 0.15599 | 0.006681 | 0.47734 | 0.47734 | 0.013431
2019-03-21 11:50 | 1553183374.3583 | 261 | 0.20818 | 0.001602 | 3.4e-05 | 100.4272 | 100.51 | 119.7523 | 21.065 | 0.002634 | 0.002434 | 169.4841 | 0.26156 | 8.4045 | 0.003287 | 0.004418 | 0.004263 | 5.7e-05 | 0.000935 | 0.31852 | 1.2915 | 2.548 | 99.58 | 20.86 | 37.4 | 84.9979 | 100.7441 | 80.1907 | 0.009271 | 0.99916 | 0.023927 | 23.9469 | -0.007371 | 0.1556 | 0.006706 | 0.47963 | 0.47963 | 0.013334
2019-03-21 11:50 | 1553183375.3596 | 262 | 0.20814 | 0.001601 | 3.4e-05 | 100.423 | 100.2992 | 119.6221 | 21.0662 | 0.002641 | 0.002482 | 169.429 | 0.26216 | 8.4023 | 0.00317 | 0.00481 | 0.004773 | 5.8e-05 | 0.00093 | 0.31188 | 1.2969 | 2.548 | 99.58 | 20.86 | 37.402 | 85.1676 | 97.9216 | 81.3504 | 0.009271 | 0.99589 | 0.023875 | 23.9732 | 0.011038 | 0.15596 | 0.006987 | 0.4987 | 0.4987 | 0.012853
2019-03-21 11:50 | 1553183376.3598 | 263 | 0.20814 | 0.001589 | 3.4e-05 | 100.4144 | 100.5948 | 119.6376 | 21.0658 | 0.002658 | 0.002444 | 169.3982 | 0.26134 | 8.4 | 0.003299 | 0.004416 | 0.004325 | 5.7e-05 | 0.000933 | 0.3178 | 1.3198 | 2.548 | 99.58 | 20.86 | 37.407 | 85.5675 | 101.0253 | 80.437 | 0.009273 | 0.99847 | 0.023907 | 23.9438 | -0.016056 | 0.15547 | 0.007001 | 0.50031 | 0.50031 | 0.012772
2019-03-21 11:50 | 1553183377.3584 | 264 | 0.20816 | 0.001596 | 3.5e-05 | 100.3974 | 100.4146 | 119.7317 | 21.0661 | 0.002584 | 0.002424 | 169.3674 | 0.26131 | 8.4048 | 0.003327 | 0.005107 | 0.004957 | 6e-05 | 0.000937 | 0.31078 | 1.3442 | 2.548 | 99.58 | 20.86 | 37.409 | 83.8034 | 101.7147 | 79.9547 | 0.009273 | 0.99983 | 0.023933 | 23.9369 | -0.001535 | 0.15545 | 0.00682 | 0.48785 | 0.48785 | 0.013096
2019-03-21 11:50 | 1553183378.3595 | 265 | 0.20816 | 0.001597 | 3.5e-05 | 100.3971 | 100.4221 | 119.8833 | 21.0661 | 0.002654 | 0.002492 | 169.3334 | 0.26203 | 8.4034 | 0.003305 | 0.004828 | 0.004762 | 5.8e-05 | 0.000938 | 0.33716 | 1.3152 | 2.548 | 99.58 | 20.86 | 37.412 | 85.4687 | 101.1808 | 81.5758 | 0.009274 | 0.99526 | 0.023893 | 24.007 | -0.002229 | 0.15588 | 0.006832 | 0.4879 | 0.4879 | 0.013131
2019-03-21 11:50 | 1553183379.3599 | 266 | 0.20818 | 0.001591 | 3.5e-05 | 100.35 | 100.4572 | 119.5955 | 21.0666 | 0.002623 | 0.00241 | 169.3025 | 0.26113 | 8.4035 | 0.003333 | 0.004991 | 0.004993 | 5.8e-05 | 0.000938 | 0.32997 | 1.3269 | 2.5481 | 99.58 | 20.86 | 37.415 | 84.7302 | 101.8526 | 79.6192 | 0.009275 | 1.0008 | 0.023931 | 23.9119 | -0.009548 | 0.15535 | 0.006718 | 0.48052 | 0.48052 | 0.013287
2019-03-21 11:50 | 1553183380.3597 | 267 | 0.20821 | 0.001594 | 3.4e-05 | 100.3837 | 100.4095 | 119.3707 | 21.0657 | 0.002651 | 0.002417 | 169.2547 | 0.26124 | 8.402 | 0.003328 | 0.004178 | 0.004068 | 5.5e-05 | 0.00094 | 0.31006 | 1.3136 | 2.5479 | 99.58 | 20.86 | 37.417 | 85.4012 | 101.7349 | 79.7734 | 0.009275 | 1.0003 | 0.023903 | 23.8946 | -0.002301 | 0.15541 | 0.006575 | 0.46974 | 0.46974 | 0.013598
2019-03-21 11:50 | 1553183381.3597 | 268 | 0.20818 | 0.001595 | 3.4e-05 | 100.4165 | 100.4245 | 119.2397 | 21.0667 | 0.002594 | 0.002374 | 169.2168 | 0.26131 | 8.4013 | 0.00327 | 0.004326 | 0.004233 | 5.8e-05 | 0.00094 | 0.32117 | 1.3373 | 2.548 | 99.58 | 20.86 | 37.423 | 84.0405 | 100.3401 | 78.7392 | 0.009277 | 1.0033 | 0.023925 | 23.8465 | -0.000716 | 0.15545 | 0.006735 | 0.48164 | 0.48164 | 0.013265
2019-03-21 11:50 | 1553183382.3587 | 269 | 0.20813 | 0.001594 | 3.4e-05 | 100.3852 | 100.5648 | 119.2146 | 21.0658 | 0.002638 | 0.002387 | 169.2138 | 0.26148 | 8.4017 | 0.003302 | 0.005293 | 0.005229 | 5.6e-05 | 0.000941 | 0.33069 | 1.2873 | 2.5481 | 99.58 | 20.86 | 37.427 | 85.1023 | 101.1124 | 79.0656 | 0.009278 | 1.0024 | 0.023911 | 23.855 | -0.015992 | 0.15555 | 0.007008 | 0.50096 | 0.50096 | 0.012762
2019-03-21 11:50 | 1553183383.3598 | 270 | 0.20817 | 0.001587 | 3.3e-05 | 100.4138 | 100.341 | 119.1543 | 21.0668 | 0.002629 | 0.002408 | 169.19 | 0.26157 | 8.4022 | 0.003281 | 0.005382 | 0.005266 | 5.6e-05 | 0.000942 | 0.2995 | 1.3093 | 2.5478 | 99.58 | 20.85 | 37.429 | 84.8842 | 100.5976 | 79.5741 | 0.009272 | 1.0009 | 0.023888 | 23.8662 | 0.006481 | 0.15561 | 0.006778 | 0.48407 | 0.48407 | 0.013212
2019-03-21 11:50 | 1553183384.3594 | 271 | 0.20815 | 0.001584 | 3.4e-05 | 100.3756 | 100.4913 | 119.1674 | 21.067 | 0.002633 | 0.002431 | 169.1359 | 0.26058 | 8.4033 | 0.003248 | 0.004637 | 0.004583 | 5.7e-05 | 0.00094 | 0.31332 | 1.3179 | 2.548 | 99.58 | 20.85 | 37.432 | 84.9671 | 99.8006 | 80.1169 | 0.009273 | 0.99937 | 0.023871 | 23.8859 | -0.010296 | 0.15502 | 0.006929 | 0.49447 | 0.49447 | 0.012885
2019-03-21 11:50 | 1553183385.3593 | 272 | 0.20818 | 0.001584 | 3.4e-05 | 100.3616 | 100.3363 | 119.1925 | 21.0671 | 0.0026 | 0.002438 | 169.0966 | 0.26165 | 8.4031 | 0.003334 | 0.004536 | 0.004518 | 6e-05 | 0.000942 | 0.33356 | 1.3101 | 2.5479 | 99.58 | 20.85 | 37.436 | 84.1885 | 101.8893 | 80.2848 | 0.009274 | 0.9989 | 0.023868 | 23.8941 | 0.00226 | 0.15565 | 0.00674 | 0.48088 | 0.48088 | 0.013304
2019-03-21 11:50 | 1553183386.3589 | 273 | 0.20815 | 0.001582 | 3.5e-05 | 100.3298 | 100.3996 | 119.1761 | 21.0668 | 0.002648 | 0.002432 | 169.0717 | 0.26147 | 8.4033 | 0.003323 | 0.004866 | 0.004743 | 6e-05 | 0.000941 | 0.29966 | 1.3322 | 2.5479 | 99.58 | 20.85 | 37.438 | 85.3296 | 101.6016 | 80.1426 | 0.009275 | 0.9993 | 0.023871 | 23.8876 | -0.006227 | 0.15555 | 0.006936 | 0.49492 | 0.49492 | 0.012917
2019-03-21 11:50 | 1553183387.36 | 274 | 0.20819 | 0.001578 | 3.4e-05 | 100.3658 | 100.6145 | 119.3646 | 21.0683 | 0.002619 | 0.002396 | 169.0403 | 0.2618 | 8.4018 | 0.003246 | 0.005213 | 0.005204 | 5.6e-05 | 0.000942 | 0.30917 | 1.3859 | 2.5479 | 99.58 | 20.85 | 37.44 | 84.6412 | 99.7554 | 79.2833 | 0.009275 | 1.0017 | 0.023919 | 23.8774 | -0.02214 | 0.15575 | 0.006712 | 0.47984 | 0.47984 | 0.01334
2019-03-21 11:50 | 1553183388.36 | 275 | 0.20819 | 0.001575 | 3.5e-05 | 100.3727 | 100.3867 | 118.9552 | 21.0672 | 0.002605 | 0.002449 | 169.026 | 0.26185 | 8.4017 | 0.003309 | 0.004793 | 0.004729 | 6e-05 | 0.000944 | 0.3131 | 1.317 | 2.5481 | 99.58 | 20.85 | 37.444 | 84.3016 | 101.2715 | 80.5429 | 0.009276 | 0.99817 | 0.023835 | 23.879 | -0.00125 | 0.15577 | 0.006702 | 0.47749 | 0.47749 | 0.013408
2019-03-21 11:50 | 1553183389.3605 | 276 | 0.20826 | 0.001575 | 3.5e-05 | 100.3968 | 100.3806 | 119.1378 | 21.0671 | 0.002603 | 0.0024 | 168.9918 | 0.26079 | 8.4011 | 0.003223 | 0.005063 | 0.004975 | 5.7e-05 | 0.000938 | 0.30663 | 1.2838 | 2.5479 | 99.58 | 20.86 | 37.447 | 84.2452 | 99.1982 | 79.3638 | 0.009283 | 1.0015 | 0.023893 | 23.8575 | 0.001447 | 0.15515 | 0.006286 | 0.44877 | 0.44877 | 0.014209
2019-03-21 11:50 | 1553183390.3592 | 277 | 0.20825 | 0.001572 | 3.4e-05 | 100.3256 | 100.4763 | 119.4883 | 21.0672 | 0.002634 | 0.002372 | 168.9521 | 0.26106 | 8.4025 | 0.003268 | 0.004872 | 0.004773 | 5.8e-05 | 0.000946 | 0.30934 | 1.3278 | 2.5481 | 99.58 | 20.85 | 37.448 | 85.0024 | 100.2699 | 78.7091 | 0.009277 | 1.0034 | 0.023951 | 23.8703 | -0.013424 | 0.15531 | 0.006327 | 0.45288 | 0.45288 | 0.014095
2019-03-21 11:50 | 1553183391.3582 | 278 | 0.2082 | 0.001575 | 3.4e-05 | 100.2904 | 100.5403 | 119.6554 | 21.0674 | 0.002632 | 0.002374 | 168.9267 | 0.26144 | 8.4021 | 0.003247 | 0.004134 | 0.004061 | 5.9e-05 | 0.000939 | 0.32892 | 1.3053 | 2.5479 | 99.58 | 20.85 | 37.452 | 84.9439 | 99.7696 | 78.7538 | 0.009278 | 1.0032 | 0.023966 | 23.8885 | -0.022257 | 0.15553 | 0.006632 | 0.4751 | 0.4751 | 0.013455
2019-03-21 11:50 | 1553183392.3596 | 279 | 0.20822 | 0.001571 | 3.4e-05 | 100.3259 | 100.3615 | 119.7418 | 21.0679 | 0.002623 | 0.002377 | 168.8914 | 0.26148 | 8.4016 | 0.003263 | 0.005289 | 0.005173 | 5.8e-05 | 0.000939 | 0.31509 | 1.334 | 2.548 | 99.58 | 20.85 | 37.452 | 84.726 | 100.1514 | 78.8221 | 0.009278 | 1.003 | 0.023972 | 23.8995 | -0.003174 | 0.15555 | 0.006546 | 0.46905 | 0.46905 | 0.01363
2019-03-21 11:50 | 1553183393.3594 | 280 | 0.20823 | 0.001564 | 3.3e-05 | 100.3006 | 100.3022 | 119.7028 | 21.0685 | 0.002546 | 0.002432 | 168.8531 | 0.26164 | 8.4007 | 0.003319 | 0.00489 | 0.00486 | 5.6e-05 | 0.000936 | 0.31581 | 1.3058 | 2.5481 | 99.58 | 20.85 | 37.452 | 82.8861 | 101.5225 | 80.152 | 0.009278 | 0.99927 | 0.023923 | 23.9407 | -0.000148 | 0.15565 | 0.006498 | 0.46467 | 0.46467 | 0.013767
2019-03-21 11:50 | 1553183394.3588 | 281 | 0.20822 | 0.001565 | 3.3e-05 | 100.3402 | 100.3342 | 119.7843 | 21.0693 | 0.002543 | 0.002379 | 168.8214 | 0.2618 | 8.4013 | 0.003357 | 0.004269 | 0.004109 | 6e-05 | 0.000939 | 0.29999 | 1.3111 | 2.5477 | 99.58 | 20.85 | 37.452 | 82.8073 | 102.438 | 78.8578 | 0.009278 | 1.0029 | 0.023975 | 23.9049 | 0.00053 | 0.15574 | 0.006503 | 0.46607 | 0.46607 | 0.013734
2019-03-21 11:50 | 1553183395.3587 | 282 | 0.2082 | 0.001565 | 3.3e-05 | 100.2535 | 100.3533 | 119.9051 | 21.0682 | 0.002603 | 0.002489 | 168.7977 | 0.2615 | 8.3992 | 0.003322 | 0.004378 | 0.004271 | 6e-05 | 0.000936 | 0.31996 | 1.2756 | 2.5479 | 99.58 | 20.85 | 37.451 | 84.2612 | 101.5943 | 81.5014 | 0.009278 | 0.99547 | 0.023898 | 24.0066 | -0.008902 | 0.15557 | 0.006626 | 0.47337 | 0.47337 | 0.013507
2019-03-21 11:50 | 1553183396.3595 | 283 | 0.20821 | 0.001564 | 3.3e-05 | 100.2877 | 100.3441 | 120.1502 | 21.068 | 0.002555 | 0.002397 | 168.7791 | 0.26115 | 8.4009 | 0.003321 | 0.004517 | 0.004399 | 5.6e-05 | 0.000934 | 0.35153 | 1.3016 | 2.548 | 99.58 | 20.85 | 37.45 | 83.1056 | 101.5674 | 79.2993 | 0.009278 | 1.0017 | 0.023997 | 23.9564 | -0.005026 | 0.15536 | 0.006597 | 0.47325 | 0.47325 | 0.013492
2019-03-21 11:50 | 1553183397.359 | 284 | 0.20821 | 0.001562 | 3.3e-05 | 100.2475 | 100.3009 | 119.9528 | 21.0681 | 0.002621 | 0.002443 | 168.7408 | 0.26125 | 8.4038 | 0.003348 | 0.005211 | 0.005117 | 5.8e-05 | 0.000934 | 0.30182 | 1.3317 | 2.548 | 99.58 | 20.85 | 37.449 | 84.6723 | 102.2068 | 80.3926 | 0.009277 | 0.99859 | 0.02394 | 23.9738 | -0.004758 | 0.15542 | 0.006572 | 0.47034 | 0.47034 | 0.013581
2019-03-21 11:50 | 1553183398.3589 | 285 | 0.20819 | 0.001559 | 3.3e-05 | 100.1558 | 100.2624 | 120.218 | 21.0688 | 0.002609 | 0.002421 | 168.7047 | 0.26059 | 8.4042 | 0.003274 | 0.005121 | 0.005053 | 6e-05 | 0.000935 | 0.3282 | 1.3466 | 2.548 | 99.58 | 20.85 | 37.447 | 84.4014 | 100.4173 | 79.8684 | 0.009277 | 1.0001 | 0.023984 | 23.9825 | -0.009516 | 0.15503 | 0.00669 | 0.47964 | 0.47964 | 0.013284
2019-03-21 11:50 | 1553183399.3596 | 286 | 0.20822 | 0.001553 | 3.4e-05 | 100.185 | 100.2835 | 120.0047 | 21.0693 | 0.0026 | 0.002471 | 168.6604 | 0.26211 | 8.4017 | 0.003363 | 0.004575 | 0.004407 | 6.2e-05 | 0.000936 | 0.30198 | 1.3352 | 2.5481 | 99.58 | 20.85 | 37.447 | 84.1834 | 102.5772 | 81.0821 | 0.009277 | 0.99665 | 0.023922 | 24.0024 | -0.008793 | 0.15593 | 0.006557 | 0.46884 | 0.46884 | 0.013669
2019-03-21 11:50 | 1553183400.3591 | 287 | 0.20821 | 0.001553 | 3.4e-05 | 100.2034 | 100.1963 | 119.7526 | 21.0693 | 0.002571 | 0.002433 | 168.6462 | 0.26063 | 8.4044 | 0.003223 | 0.004618 | 0.004525 | 5.8e-05 | 0.000943 | 0.30917 | 1.3281 | 2.5478 | 99.58 | 20.85 | 37.446 | 83.4781 | 99.1934 | 80.1614 | 0.009277 | 0.99925 | 0.023928 | 23.946 | 0.000636 | 0.15505 | 0.006604 | 0.4723 | 0.4723 | 0.013492
2019-03-21 11:50 | 1553183401.3597 | 288 | 0.20822 | 0.001548 | 3.5e-05 | 100.2326 | 100.1585 | 119.553 | 21.0701 | 0.002588 | 0.002421 | 168.6186 | 0.26139 | 8.4005 | 0.003324 | 0.004709 | 0.004662 | 5.9e-05 | 0.000947 | 0.31238 | 1.3142 | 2.5479 | 99.58 | 20.86 | 37.442 | 83.9033 | 101.6364 | 79.8683 | 0.009281 | 1.0001 | 0.023918 | 23.9161 | 0.006618 | 0.1555 | 0.006575 | 0.47003 | 0.47003 | 0.013598
2019-03-21 11:50 | 1553183402.3588 | 289 | 0.20824 | 0.001547 | 3.5e-05 | 100.1292 | 100.0301 | 119.5249 | 21.0696 | 0.002594 | 0.002427 | 168.5733 | 0.26184 | 8.4017 | 0.003345 | 0.004083 | 0.00405 | 5.9e-05 | 0.000952 | 0.3298 | 1.2948 | 2.548 | 99.58 | 20.86 | 37.44 | 84.0247 | 102.1523 | 80.0259 | 0.009281 | 0.99963 | 0.02391 | 23.9186 | 0.008861 | 0.15577 | 0.006439 | 0.4601 | 0.4601 | 0.013915
2019-03-21 11:50 | 1553183403.3598 | 290 | 0.20823 | 0.001545 | 3.4e-05 | 100.1447 | 100.3519 | 119.8307 | 21.0683 | 0.002575 | 0.002392 | 168.5574 | 0.2619 | 8.4027 | 0.003332 | 0.00466 | 0.004541 | 5.8e-05 | 0.000948 | 0.33229 | 1.3276 | 2.5481 | 99.58 | 20.86 | 37.438 | 83.5779 | 101.834 | 79.1865 | 0.00928 | 1.002 | 0.023969 | 23.9207 | -0.018489 | 0.1558 | 0.006513 | 0.46661 | 0.46661 | 0.013724
2019-03-21 11:50 | 1553183404.3593 | 291 | 0.20825 | 0.001545 | 3.4e-05 | 100.2112 | 100.1599 | 119.7218 | 21.0703 | 0.002597 | 0.002445 | 168.563 | 0.26212 | 8.3992 | 0.00333 | 0.004917 | 0.004845 | 5.8e-05 | 0.000951 | 0.3194 | 1.317 | 2.548 | 99.58 | 20.86 | 37.436 | 84.1172 | 101.781 | 80.4437 | 0.00928 | 0.99845 | 0.023915 | 23.9525 | 0.004582 | 0.15593 | 0.006408 | 0.45803 | 0.45803 | 0.013992
2019-03-21 11:50 | 1553183405.3582 | 292 | 0.20824 | 0.001546 | 3.4e-05 | 100.1283 | 100.1909 | 119.5223 | 21.0696 | 0.002653 | 0.002393 | 168.5425 | 0.2614 | 8.4022 | 0.003261 | 0.005 | 0.004899 | 5.8e-05 | 0.000952 | 0.33301 | 1.3431 | 2.548 | 99.58 | 20.86 | 37.434 | 85.455 | 100.12 | 79.2082 | 0.009279 | 1.0019 | 0.023937 | 23.8906 | -0.005582 | 0.15551 | 0.006422 | 0.45942 | 0.45942 | 0.013912
2019-03-21 11:50 | 1553183406.358 | 293 | 0.20825 | 0.00154 | 3.4e-05 | 100.1152 | 100.2232 | 119.4136 | 21.0693 | 0.002621 | 0.002421 | 168.5002 | 0.26139 | 8.4035 | 0.003302 | 0.00499 | 0.004917 | 5.9e-05 | 0.000951 | 0.3361 | 1.3336 | 2.5481 | 99.58 | 20.86 | 37.431 | 84.6836 | 101.1007 | 79.8751 | 0.009279 | 1.0001 | 0.023904 | 23.9024 | -0.009641 | 0.1555 | 0.006403 | 0.45746 | 0.45746 | 0.013971
2019-03-21 11:50 | 1553183407.358 | 294 | 0.20821 | 0.00154 | 3.3e-05 | 100.131 | 100.119 | 119.4197 | 21.0703 | 0.002565 | 0.002445 | 168.4752 | 0.26137 | 8.4042 | 0.003298 | 0.004376 | 0.004294 | 5.7e-05 | 0.000952 | 0.32244 | 1.3095 | 2.548 | 99.58 | 20.86 | 37.43 | 83.3473 | 101.0098 | 80.464 | 0.009278 | 0.99839 | 0.023884 | 23.9229 | 0.001076 | 0.15549 | 0.0066 | 0.47118 | 0.47118 | 0.013563
2019-03-21 11:50 | 1553183408.3585 | 295 | 0.2082 | 0.001534 | 3.4e-05 | 100.112 | 100.274 | 119.4516 | 21.0711 | 0.002562 | 0.002437 | 168.4455 | 0.26181 | 8.4032 | 0.003286 | 0.004616 | 0.004517 | 5.6e-05 | 0.000953 | 0.31006 | 1.287 | 2.548 | 99.58 | 20.86 | 37.426 | 83.2709 | 100.7132 | 80.2629 | 0.009277 | 0.99896 | 0.023894 | 23.9193 | -0.014462 | 0.15575 | 0.006679 | 0.47703 | 0.47703 | 0.013419
2019-03-21 11:50 | 1553183409.3592 | 296 | 0.20823 | 0.001535 | 3.4e-05 | 100.0297 | 100.1132 | 119.4831 | 21.0696 | 0.002578 | 0.002387 | 168.4448 | 0.26175 | 8.4037 | 0.003315 | 0.005287 | 0.005172 | 6e-05 | 0.000954 | 0.2912 | 1.2717 | 2.5478 | 99.58 | 20.86 | 37.425 | 83.6431 | 101.4113 | 79.0491 | 0.009277 | 1.0024 | 0.023939 | 23.8813 | -0.00746 | 0.15572 | 0.006505 | 0.46542 | 0.46542 | 0.013751
2019-03-21 11:50 | 1553183410.3593 | 297 | 0.20823 | 0.001531 | 3.4e-05 | 100.1146 | 100.0904 | 119.4925 | 21.0708 | 0.002638 | 0.002394 | 168.4257 | 0.26216 | 8.4032 | 0.003267 | 0.005104 | 0.00507 | 5.5e-05 | 0.000956 | 0.29878 | 1.2624 | 2.5479 | 99.58 | 20.86 | 37.424 | 85.087 | 100.2515 | 79.2343 | 0.009277 | 1.0019 | 0.023933 | 23.8886 | 0.002168 | 0.15596 | 0.006537 | 0.46762 | 0.46762 | 0.013707
2019-03-21 11:50 | 1553183411.3594 | 298 | 0.20824 | 0.001529 | 3.4e-05 | 100.022 | 100.1118 | 119.527 | 21.0703 | 0.002515 | 0.00238 | 168.4039 | 0.26254 | 8.4045 | 0.003351 | 0.004565 | 0.004469 | 5.6e-05 | 0.000958 | 0.30574 | 1.2898 | 2.5479 | 99.58 | 20.86 | 37.422 | 82.1258 | 102.2913 | 78.9032 | 0.009276 | 1.0028 | 0.023948 | 23.8808 | -0.008032 | 0.15619 | 0.006482 | 0.46394 | 0.46394 | 0.013837
2019-03-21 11:50 | 1553183412.3594 | 299 | 0.20829 | 0.001532 | 3.4e-05 | 100.0268 | 100.071 | 119.7317 | 21.0697 | 0.002598 | 0.002434 | 168.3653 | 0.26195 | 8.4021 | 0.003311 | 0.004734 | 0.004648 | 5.5e-05 | 0.000959 | 0.35081 | 1.3362 | 2.5481 | 99.58 | 20.86 | 37.421 | 84.1379 | 101.3202 | 80.1875 | 0.009276 | 0.99917 | 0.023925 | 23.9448 | -0.003953 | 0.15584 | 0.006167 | 0.44089 | 0.44089 | 0.014527
2019-03-21 11:50 | 1553183413.4753 | 300 | 0.2082 | 0.001528 | 3.4e-05 | 99.9576 | 99.9351 | 119.5563 | 21.07 | 0.002581 | 0.002422 | 168.339 | 0.26196 | 8.4025 | 0.003386 | 0.004596 | 0.0046 | 5.9e-05 | 0.00096 | 0.29159 | 1.327 | 2.548 | 99.58 | 20.86 | 37.421 | 83.7224 | 103.1513 | 79.8985 | 0.009276 | 0.99999 | 0.023917 | 23.9174 | 0.002019 | 0.15584 | 0.006679 | 0.4775 | 0.4775 | 0.013414
2019-03-21 11:50 | 1553183414.3867 | 301 | 0.20822 | 0.00153 | 3.3e-05 | 99.9734 | 100.1026 | 119.5422 | 21.0698 | 0.002554 | 0.002391 | 168.3065 | 0.26221 | 8.4015 | 0.003335 | 0.004566 | 0.004467 | 6e-05 | 0.000959 | 0.31188 | 1.3239 | 2.548 | 99.58 | 20.86 | 37.42 | 83.0646 | 101.9006 | 79.1591 | 0.009276 | 1.0021 | 0.023941 | 23.891 | -0.011554 | 0.15599 | 0.006567 | 0.47 | 0.47 | 0.013641
2019-03-21 11:50 | 1553183415.3867 | 302 | 0.20827 | 0.001529 | 3.3e-05 | 99.9886 | 99.9766 | 119.6946 | 21.0703 | 0.002584 | 0.002421 | 168.262 | 0.26249 | 8.4009 | 0.003335 | 0.004602 | 0.00462 | 5.5e-05 | 0.000959 | 0.34346 | 1.3257 | 2.5478 | 99.58 | 20.86 | 37.419 | 83.796 | 101.9035 | 79.8755 | 0.009276 | 1.0001 | 0.023932 | 23.9305 | 0.001073 | 0.15616 | 0.006302 | 0.45082 | 0.45082 | 0.014236
2019-03-21 11:50 | 1553183416.3854 | 303 | 0.20824 | 0.001531 | 3.2e-05 | 99.9639 | 99.9051 | 119.6831 | 21.0708 | 0.002581 | 0.002366 | 168.2276 | 0.26136 | 8.4004 | 0.003327 | 0.004671 | 0.004595 | 5.5e-05 | 0.000958 | 0.31636 | 1.3186 | 2.548 | 99.58 | 20.85 | 37.419 | 83.7246 | 101.7232 | 78.5449 | 0.00927 | 1.0038 | 0.023976 | 23.8842 | 0.005261 | 0.15548 | 0.006447 | 0.46206 | 0.46206 | 0.01383
2019-03-21 11:50 | 1553183417.3855 | 304 | 0.20826 | 0.001525 | 3.2e-05 | 99.9234 | 99.8274 | 119.6007 | 21.0698 | 0.002554 | 0.002416 | 168.1918 | 0.26307 | 8.4014 | 0.003327 | 0.004368 | 0.004276 | 5.7e-05 | 0.000957 | 0.32676 | 1.2863 | 2.5479 | 99.58 | 20.85 | 37.42 | 83.083 | 101.7151 | 79.7579 | 0.00927 | 1.0004 | 0.023926 | 23.9171 | 0.00859 | 0.1565 | 0.006347 | 0.45397 | 0.45397 | 0.014169
2019-03-21 11:50 | 1553183418.3859 | 305 | 0.20823 | 0.00152 | 3.1e-05 | 99.9433 | 99.9405 | 119.4223 | 21.0702 | 0.00255 | 0.002354 | 168.1545 | 0.26034 | 8.4026 | 0.003226 | 0.004645 | 0.004596 | 5.5e-05 | 0.000957 | 0.33179 | 1.2924 | 2.5479 | 99.58 | 20.85 | 37.42 | 82.9844 | 99.2789 | 78.2717 | 0.00927 | 1.0046 | 0.023959 | 23.8489 | 0.000251 | 0.15488 | 0.006514 | 0.46662 | 0.46662 | 0.013642
2019-03-21 11:50 | 1553183419.3866 | 306 | 0.20826 | 0.001523 | 3.1e-05 | 99.9058 | 99.9514 | 119.3092 | 21.0696 | 0.002588 | 0.002443 | 168.095 | 0.26207 | 8.4024 | 0.003226 | 0.0048 | 0.004659 | 5.6e-05 | 0.000956 | 0.31996 | 1.3073 | 2.5479 | 99.58 | 20.85 | 37.422 | 83.8989 | 99.2625 | 80.4032 | 0.009271 | 0.99856 | 0.023875 | 23.9098 | -0.004085 | 0.15591 | 0.006365 | 0.45434 | 0.45434 | 0.014103
2019-03-21 11:50 | 1553183420.3857 | 307 | 0.20825 | 0.001517 | 3.1e-05 | 99.9275 | 99.7907 | 119.3747 | 21.07 | 0.002606 | 0.002359 | 168.0557 | 0.26094 | 8.4024 | 0.003209 | 0.004764 | 0.004649 | 4.9e-05 | 0.000954 | 0.33323 | 1.3201 | 2.5479 | 99.58 | 20.86 | 37.423 | 84.3175 | 98.867 | 78.3993 | 0.009277 | 1.0043 | 0.02395 | 23.8485 | 0.01226 | 0.15524 | 0.006412 | 0.45916 | 0.45916 | 0.013895
2019-03-21 11:50 | 1553183421.385 | 308 | 0.20825 | 0.001517 | 3e-05 | 99.9108 | 99.9467 | 119.5807 | 21.0706 | 0.002638 | 0.002401 | 168.015 | 0.26127 | 8.4015 | 0.003201 | 0.005257 | 0.005096 | 5.6e-05 | 0.000954 | 0.32029 | 1.2995 | 2.5479 | 99.58 | 20.86 | 37.425 | 85.1024 | 98.6723 | 79.3977 | 0.009277 | 1.0014 | 0.023937 | 23.9029 | -0.003205 | 0.15543 | 0.006392 | 0.45752 | 0.45752 | 0.013963
2019-03-21 11:50 | 1553183422.3847 | 309 | 0.20824 | 0.001517 | 3e-05 | 99.8706 | 99.7219 | 119.2413 | 21.0707 | 0.002611 | 0.002424 | 167.9853 | 0.26186 | 8.402 | 0.003239 | 0.004139 | 0.00412 | 5.6e-05 | 0.000951 | 0.30829 | 1.2743 | 2.548 | 99.58 | 20.86 | 37.426 | 84.4392 | 99.5781 | 79.9553 | 0.009277 | 0.99983 | 0.023884 | 23.8878 | 0.013333 | 0.15578 | 0.006494 | 0.46383 | 0.46383 | 0.013804
2019-03-21 11:50 | 1553183423.3847 | 310 | 0.20825 | 0.001512 | 2.9e-05 | 99.8793 | 99.8915 | 119.3829 | 21.0705 | 0.002579 | 0.002385 | 167.9509 | 0.26228 | 8.4032 | 0.003222 | 0.004894 | 0.004827 | 5.5e-05 | 0.000953 | 0.31686 | 1.3025 | 2.5478 | 99.58 | 20.86 | 37.427 | 83.6721 | 99.1663 | 79.0063 | 0.009278 | 1.0025 | 0.02393 | 23.8699 | -0.001093 | 0.15603 | 0.006463 | 0.46252 | 0.46252 | 0.013865
2019-03-21 11:50 | 1553183424.3862 | 311 | 0.20828 | 0.001514 | 2.9e-05 | 99.8852 | 99.7532 | 119.6094 | 21.0709 | 0.002559 | 0.002371 | 167.9101 | 0.26198 | 8.401 | 0.003256 | 0.004886 | 0.004752 | 5.6e-05 | 0.000949 | 0.30143 | 1.2726 | 2.5479 | 99.58 | 20.86 | 37.428 | 83.1912 | 99.9895 | 78.6847 | 0.009278 | 1.0034 | 0.023964 | 23.8816 | 0.011831 | 0.15585 | 0.006267 | 0.44911 | 0.44911 | 0.014263
2019-03-21 11:50 | 1553183425.3849 | 312 | 0.20825 | 0.001511 | 2.9e-05 | 99.8349 | 99.6364 | 119.7547 | 21.0709 | 0.002592 | 0.002376 | 167.9052 | 0.26182 | 8.4007 | 0.003197 | 0.004469 | 0.004445 | 5.6e-05 | 0.000949 | 0.30989 | 1.2765 | 2.5479 | 99.58 | 20.86 | 37.429 | 83.9993 | 98.5754 | 78.7899 | 0.009278 | 1.0031 | 0.023975 | 23.8997 | 0.017802 | 0.15576 | 0.006435 | 0.46139 | 0.46139 | 0.013875
2019-03-21 11:50 | 1553183426.385 | 313 | 0.20827 | 0.001508 | 2.8e-05 | 99.8319 | 100.0407 | 119.5568 | 21.0706 | 0.002527 | 0.002389 | 167.8712 | 0.26235 | 8.3997 | 0.00317 | 0.004383 | 0.004174 | 5.8e-05 | 0.000951 | 0.32316 | 1.3143 | 2.5479 | 99.58 | 20.86 | 37.431 | 82.4201 | 97.9165 | 79.1164 | 0.009279 | 1.0022 | 0.023944 | 23.891 | -0.018684 | 0.15607 | 0.006314 | 0.45215 | 0.45215 | 0.014187
2019-03-21 11:50 | 1553183427.3864 | 314 | 0.20824 | 0.001507 | 2.8e-05 | 99.856 | 99.9276 | 119.4366 | 21.0706 | 0.002535 | 0.002356 | 167.8556 | 0.26137 | 8.4033 | 0.003231 | 0.004724 | 0.004644 | 5.8e-05 | 0.000948 | 0.31725 | 1.3239 | 2.548 | 99.58 | 20.86 | 37.433 | 82.6275 | 99.3788 | 78.3063 | 0.009279 | 1.0045 | 0.023959 | 23.8515 | -0.006407 | 0.15549 | 0.006526 | 0.46772 | 0.46772 | 0.013664
2019-03-21 11:50 | 1553183428.3861 | 315 | 0.20827 | 0.001506 | 2.8e-05 | 99.8882 | 99.9255 | 119.4751 | 21.0704 | 0.002555 | 0.002364 | 167.8329 | 0.26223 | 8.4034 | 0.003201 | 0.004891 | 0.004798 | 5.5e-05 | 0.000952 | 0.33987 | 1.3023 | 2.5479 | 99.58 | 20.85 | 37.434 | 83.1041 | 98.654 | 78.4994 | 0.009274 | 1.004 | 0.023957 | 23.8619 | -0.003342 | 0.156 | 0.006308 | 0.45197 | 0.45197 | 0.014186
2019-03-21 11:50 | 1553183429.3851 | 316 | 0.20825 | 0.001506 | 2.8e-05 | 99.912 | 99.978 | 119.6826 | 21.0709 | 0.002522 | 0.002332 | 167.8262 | 0.26111 | 8.4033 | 0.003208 | 0.005027 | 0.004985 | 5.6e-05 | 0.00095 | 0.30735 | 1.3316 | 2.548 | 99.58 | 20.85 | 37.436 | 82.3161 | 98.8271 | 77.7476 | 0.009274 | 1.0061 | 0.024003 | 23.8571 | -0.005902 | 0.15533 | 0.006411 | 0.4603 | 0.4603 | 0.01387
2019-03-21 11:50 | 1553183430.3865 | 317 | 0.20829 | 0.001508 | 2.8e-05 | 99.901 | 99.7314 | 119.6554 | 21.0713 | 0.002509 | 0.002376 | 167.7921 | 0.26251 | 8.4045 | 0.003212 | 0.004154 | 0.004049 | 5.8e-05 | 0.00095 | 0.33229 | 1.3231 | 2.5481 | 99.58 | 20.85 | 37.437 | 81.9996 | 98.9215 | 78.8083 | 0.009274 | 1.0031 | 0.023964 | 23.8904 | 0.015199 | 0.15616 | 0.006186 | 0.44337 | 0.44337 | 0.014476
2019-03-21 11:50 | 1553183431.3863 | 318 | 0.20828 | 0.001501 | 2.8e-05 | 99.8819 | 99.9511 | 119.4045 | 21.0708 | 0.002555 | 0.002404 | 167.7602 | 0.26203 | 8.402 | 0.003215 | 0.004336 | 0.004163 | 5.8e-05 | 0.000946 | 0.29806 | 1.3028 | 2.5479 | 99.58 | 20.85 | 37.438 | 83.1021 | 99.0058 | 79.4602 | 0.009275 | 1.0012 | 0.023917 | 23.8874 | -0.006189 | 0.15588 | 0.006277 | 0.44903 | 0.44903 | 0.014268
2019-03-21 11:51 | 1553183432.3856 | 319 | 0.20827 | 0.001495 | 2.8e-05 | 99.8572 | 99.698 | 119.4301 | 21.0718 | 0.002565 | 0.002369 | 167.7401 | 0.26223 | 8.4044 | 0.003243 | 0.004688 | 0.004654 | 5.9e-05 | 0.000946 | 0.32565 | 1.2914 | 2.5478 | 99.58 | 20.85 | 37.44 | 83.3473 | 99.6851 | 78.6418 | 0.009275 | 1.0036 | 0.023947 | 23.8622 | 0.014271 | 0.156 | 0.006365 | 0.45591 | 0.45591 | 0.014063
2019-03-21 11:51 | 1553183433.3865 | 320 | 0.20826 | 0.001497 | 2.8e-05 | 99.9168 | 99.7631 | 119.0244 | 21.0708 | 0.002592 | 0.002462 | 167.7016 | 0.26271 | 8.4014 | 0.003181 | 0.004865 | 0.004819 | 6e-05 | 0.000945 | 0.31974 | 1.2977 | 2.5478 | 99.58 | 20.84 | 37.442 | 83.986 | 98.1926 | 80.8605 | 0.00927 | 0.99727 | 0.023832 | 23.8967 | 0.013773 | 0.15629 | 0.006372 | 0.45415 | 0.45415 | 0.014144
2019-03-21 11:51 | 1553183434.3858 | 321 | 0.20828 | 0.001499 | 2.8e-05 | 99.8662 | 99.8394 | 118.8153 | 21.0714 | 0.002612 | 0.002387 | 167.6722 | 0.26225 | 8.4017 | 0.003175 | 0.004517 | 0.004418 | 6.1e-05 | 0.000944 | 0.32421 | 1.2592 | 2.548 | 99.58 | 20.84 | 37.444 | 84.4725 | 98.0286 | 79.0546 | 0.00927 | 1.0024 | 0.023872 | 23.8147 | 0.0024 | 0.15601 | 0.006259 | 0.44685 | 0.44685 | 0.01435
2019-03-21 11:51 | 1553183435.3854 | 322 | 0.20827 | 0.001493 | 2.8e-05 | 99.9147 | 99.776 | 118.7799 | 21.0714 | 0.002616 | 0.002383 | 167.6314 | 0.26167 | 8.3999 | 0.003145 | 0.004609 | 0.004457 | 6e-05 | 0.000943 | 0.30127 | 1.2932 | 2.548 | 99.58 | 20.84 | 37.445 | 84.5772 | 97.3076 | 78.9633 | 0.009271 | 1.0026 | 0.023871 | 23.8081 | 0.012426 | 0.15567 | 0.006366 | 0.45455 | 0.45455 | 0.014075
2019-03-21 11:51 | 1553183436.3865 | 323 | 0.20827 | 0.001495 | 2.8e-05 | 99.8745 | 99.9014 | 119.2028 | 21.0721 | 0.00259 | 0.002434 | 167.6042 | 0.26239 | 8.402 | 0.003205 | 0.004441 | 0.004352 | 5.9e-05 | 0.000945 | 0.3361 | 1.3109 | 2.548 | 99.58 | 20.84 | 37.446 | 83.9334 | 98.7741 | 80.1937 | 0.009271 | 0.99916 | 0.023872 | 23.892 | -0.002404 | 0.1561 | 0.006356 | 0.4538 | 0.4538 | 0.014137
2019-03-21 11:51 | 1553183437.3849 | 324 | 0.20828 | 0.001496 | 2.8e-05 | 99.8557 | 99.9327 | 119.3268 | 21.072 | 0.002549 | 0.002375 | 167.5686 | 0.26216 | 8.4019 | 0.003219 | 0.004295 | 0.004223 | 6.1e-05 | 0.000939 | 0.34324 | 1.3015 | 2.5479 | 99.58 | 20.84 | 37.445 | 82.9636 | 99.1105 | 78.7629 | 0.009271 | 1.0032 | 0.023933 | 23.856 | -0.006891 | 0.15596 | 0.006284 | 0.44978 | 0.44978 | 0.014251
2019-03-21 11:51 | 1553183438.3846 | 325 | 0.20827 | 0.001493 | 2.8e-05 | 99.8632 | 99.9582 | 119.416 | 21.0707 | 0.002544 | 0.002403 | 167.5377 | 0.26224 | 8.3998 | 0.003217 | 0.004558 | 0.00444 | 5.9e-05 | 0.000939 | 0.33019 | 1.3101 | 2.5479 | 99.58 | 20.84 | 37.445 | 82.8399 | 99.0589 | 79.4392 | 0.009271 | 1.0013 | 0.023919 | 23.8878 | -0.00851 | 0.15601 | 0.006366 | 0.45546 | 0.45546 | 0.014078
2019-03-21 11:51 | 1553183439.3846 | 326 | 0.20825 | 0.001496 | 2.8e-05 | 99.8671 | 99.7617 | 119.2327 | 21.0713 | 0.002491 | 0.00239 | 167.4897 | 0.26152 | 8.4038 | 0.003179 | 0.004558 | 0.004389 | 6e-05 | 0.00094 | 0.33107 | 1.295 | 2.5479 | 99.58 | 20.84 | 37.444 | 81.5652 | 98.128 | 79.1365 | 0.00927 | 1.0022 | 0.023911 | 23.8593 | 0.00944 | 0.15558 | 0.006453 | 0.46149 | 0.46149 | 0.013855
2019-03-21 11:51 | 1553183440.3861 | 327 | 0.20828 | 0.001491 | 2.8e-05 | 99.8465 | 99.8421 | 119.277 | 21.071 | 0.002616 | 0.002385 | 167.4579 | 0.26266 | 8.4027 | 0.00319 | 0.005057 | 0.005043 | 6.2e-05 | 0.000942 | 0.2907 | 1.3184 | 2.548 | 99.58 | 20.84 | 37.444 | 84.5628 | 98.3974 | 79.0199 | 0.00927 | 1.0025 | 0.023919 | 23.8597 | 0.000394 | 0.15625 | 0.006275 | 0.44888 | 0.44888 | 0.014307
2019-03-21 11:51 | 1553183441.3848 | 328 | 0.20825 | 0.001492 | 2.8e-05 | 99.8236 | 99.7093 | 119.4033 | 21.0733 | 0.00259 | 0.002352 | 167.4367 | 0.26195 | 8.402 | 0.003282 | 0.004172 | 0.00406 | 6e-05 | 0.000942 | 0.33179 | 1.357 | 2.548 | 99.58 | 20.84 | 37.443 | 83.9518 | 100.615 | 78.2309 | 0.00927 | 1.0047 | 0.023959 | 23.8456 | 0.010247 | 0.15584 | 0.006454 | 0.46247 | 0.46247 | 0.013849
2019-03-21 11:51 | 1553183442.3848 | 329 | 0.20828 | 0.001491 | 2.8e-05 | 99.8456 | 99.8976 | 119.7002 | 21.0721 | 0.002585 | 0.002411 | 167.4263 | 0.26101 | 8.4027 | 0.003191 | 0.004597 | 0.004456 | 5.9e-05 | 0.000941 | 0.32549 | 1.347 | 2.5479 | 99.58 | 20.84 | 37.443 | 83.8155 | 98.4338 | 79.6337 | 0.00927 | 1.0007 | 0.023941 | 23.9229 | -0.004658 | 0.15527 | 0.006279 | 0.44961 | 0.44961 | 0.014194
2019-03-21 11:51 | 1553183443.3859 | 330 | 0.20828 | 0.001488 | 2.8e-05 | 99.8459 | 99.9194 | 119.7533 | 21.0716 | 0.002533 | 0.002411 | 167.4183 | 0.26165 | 8.401 | 0.003156 | 0.004588 | 0.004525 | 6e-05 | 0.000942 | 0.32333 | 1.3321 | 2.5479 | 99.58 | 20.84 | 37.443 | 82.5781 | 97.5774 | 79.6436 | 0.00927 | 1.0007 | 0.023946 | 23.9285 | -0.006583 | 0.15565 | 0.006311 | 0.45199 | 0.45199 | 0.014154
2019-03-21 11:51 | 1553183444.3859 | 331 | 0.20825 | 0.001492 | 2.7e-05 | 99.8244 | 99.8295 | 119.5286 | 21.0721 | 0.002575 | 0.002348 | 167.3871 | 0.26153 | 8.402 | 0.003069 | 0.00501 | 0.004896 | 5.8e-05 | 0.000945 | 0.32421 | 1.3401 | 2.548 | 99.58 | 20.84 | 37.442 | 83.5862 | 95.4872 | 78.1207 | 0.00927 | 1.0051 | 0.023975 | 23.8544 | -0.000452 | 0.15559 | 0.00648 | 0.46476 | 0.46476 | 0.013759
2019-03-21 11:51 | 1553183445.386 | 332 | 0.20829 | 0.001489 | 2.7e-05 | 99.8122 | 99.8445 | 119.3129 | 21.073 | 0.002532 | 0.002406 | 167.3578 | 0.2619 | 8.4015 | 0.003169 | 0.004463 | 0.004385 | 5.8e-05 | 0.000945 | 0.32244 | 1.3139 | 2.5479 | 99.58 | 20.84 | 37.442 | 82.5433 | 97.8898 | 79.5154 | 0.00927 | 1.0011 | 0.023906 | 23.8801 | -0.002889 | 0.1558 | 0.006209 | 0.44402 | 0.44402 | 0.014422
2019-03-21 11:51 | 1553183446.385 | 333 | 0.20828 | 0.001489 | 2.7e-05 | 99.8611 | 99.8414 | 119.07 | 21.0717 | 0.00256 | 0.002416 | 167.3471 | 0.26206 | 8.4002 | 0.003172 | 0.003986 | 0.003943 | 5.9e-05 | 0.000942 | 0.29518 | 1.3211 | 2.5479 | 99.58 | 20.84 | 37.442 | 83.2138 | 97.9661 | 79.7496 | 0.00927 | 1.0004 | 0.023874 | 23.8637 | 0.001763 | 0.1559 | 0.006315 | 0.45092 | 0.45092 | 0.01421
2019-03-21 11:51 | 1553183447.3851 | 334 | 0.20827 | 0.001487 | 2.8e-05 | 99.8295 | 99.6524 | 119.1216 | 21.0723 | 0.002561 | 0.0024 | 167.3467 | 0.26212 | 8.4028 | 0.003262 | 0.004649 | 0.004622 | 5.6e-05 | 0.000936 | 0.28407 | 1.2817 | 2.548 | 99.58 | 20.84 | 37.442 | 83.2565 | 100.1478 | 79.3793 | 0.00927 | 1.0015 | 0.023891 | 23.8564 | 0.015884 | 0.15593 | 0.006354 | 0.45406 | 0.45406 | 0.014115
2019-03-21 11:51 | 1553183448.3867 | 335 | 0.20828 | 0.001486 | 2.8e-05 | 99.7854 | 99.7982 | 119.0913 | 21.0714 | 0.002529 | 0.00237 | 167.3197 | 0.26266 | 8.4012 | 0.003212 | 0.005122 | 0.005046 | 5.5e-05 | 0.000936 | 0.29894 | 1.2797 | 2.5479 | 99.58 | 20.84 | 37.444 | 82.4778 | 98.9256 | 78.6586 | 0.00927 | 1.0035 | 0.023913 | 23.8289 | -0.001142 | 0.15626 | 0.006266 | 0.4481 | 0.4481 | 0.014332
2019-03-21 11:51 | 1553183449.3852 | 336 | 0.20826 | 0.001482 | 2.8e-05 | 99.8045 | 99.7658 | 119.155 | 21.0722 | 0.002563 | 0.002408 | 167.2961 | 0.26195 | 8.402 | 0.003189 | 0.004909 | 0.00472 | 5.8e-05 | 0.000937 | 0.33948 | 1.3207 | 2.5481 | 99.58 | 20.84 | 37.445 | 83.2828 | 98.3861 | 79.5575 | 0.009271 | 1.001 | 0.023889 | 23.8657 | 0.003468 | 0.15584 | 0.006409 | 0.45792 | 0.45792 | 0.013987
2019-03-21 11:51 | 1553183450.3866 | 337 | 0.20827 | 0.001484 | 2.9e-05 | 99.7315 | 99.7403 | 118.9601 | 21.0722 | 0.002574 | 0.002441 | 167.2745 | 0.26239 | 8.4031 | 0.003161 | 0.004747 | 0.004741 | 5.8e-05 | 0.000938 | 0.32189 | 1.3039 | 2.548 | 99.58 | 20.84 | 37.448 | 83.5535 | 97.698 | 80.3622 | 0.009271 | 0.99868 | 0.023842 | 23.8734 | -0.000786 | 0.1561 | 0.006334 | 0.45159 | 0.45159 | 0.014206
2019-03-21 11:51 | 1553183451.3853 | 338 | 0.20824 | 0.001484 | 2.9e-05 | 99.7422 | 99.6514 | 118.9134 | 21.0721 | 0.002606 | 0.002412 | 167.261 | 0.26196 | 8.4043 | 0.003235 | 0.004443 | 0.004334 | 5.4e-05 | 0.000934 | 0.32244 | 1.3781 | 2.5482 | 99.58 | 20.84 | 37.449 | 84.3258 | 99.4949 | 79.6566 | 0.009272 | 1.0007 | 0.023861 | 23.8449 | 0.008151 | 0.15584 | 0.006557 | 0.46793 | 0.46793 | 0.013688
2019-03-21 11:51 | 1553183452.3843 | 339 | 0.20827 | 0.001483 | 2.9e-05 | 99.7321 | 99.5982 | 118.9026 | 21.072 | 0.002652 | 0.002354 | 167.2144 | 0.26156 | 8.4026 | 0.00327 | 0.004955 | 0.004894 | 5.4e-05 | 0.000933 | 0.31437 | 1.2989 | 2.5478 | 99.58 | 20.84 | 37.452 | 85.433 | 100.346 | 78.2616 | 0.009272 | 1.0046 | 0.023907 | 23.7966 | 0.012013 | 0.1556 | 0.006389 | 0.45683 | 0.45683 | 0.013999
2019-03-21 11:51 | 1553183453.3847 | 340 | 0.20829 | 0.001485 | 2.9e-05 | 99.7818 | 99.5911 | 119.1841 | 21.0722 | 0.002539 | 0.00238 | 167.1779 | 0.26254 | 8.4011 | 0.003126 | 0.005352 | 0.005229 | 5.6e-05 | 0.000926 | 0.31669 | 1.3072 | 2.548 | 99.58 | 20.84 | 37.453 | 82.7096 | 96.857 | 78.8981 | 0.009273 | 1.0028 | 0.023914 | 23.8463 | 0.017116 | 0.15618 | 0.006239 | 0.44614 | 0.44614 | 0.014388
2019-03-21 11:51 | 1553183454.3844 | 341 | 0.20828 | 0.001482 | 2.9e-05 | 99.7675 | 99.7835 | 119.3556 | 21.0726 | 0.002514 | 0.002372 | 167.1636 | 0.26182 | 8.4043 | 0.003226 | 0.004329 | 0.00423 | 5.5e-05 | 0.00092 | 0.30287 | 1.3175 | 2.548 | 99.58 | 20.84 | 37.454 | 82.1242 | 99.2696 | 78.6911 | 0.009273 | 1.0034 | 0.023938 | 23.8565 | -0.001432 | 0.15576 | 0.006267 | 0.44864 | 0.44864 | 0.014269
2019-03-21 11:51 | 1553183455.3855 | 342 | 0.20828 | 0.001479 | 2.9e-05 | 99.778 | 99.4903 | 119.4578 | 21.0728 | 0.002543 | 0.002385 | 167.1399 | 0.26149 | 8.4049 | 0.003199 | 0.004755 | 0.004664 | 5.3e-05 | 0.000915 | 0.32997 | 1.2949 | 2.548 | 99.58 | 20.84 | 37.455 | 82.8177 | 98.6196 | 79.0233 | 0.009273 | 1.0025 | 0.023937 | 23.8779 | 0.025828 | 0.15556 | 0.006305 | 0.45133 | 0.45133 | 0.014166
2019-03-21 11:51 | 1553183456.3859 | 343 | 0.20832 | 0.001478 | 3e-05 | 99.7509 | 99.5505 | 119.5768 | 21.0723 | 0.002556 | 0.002369 | 167.1184 | 0.26161 | 8.4019 | 0.003242 | 0.004414 | 0.004343 | 5.2e-05 | 0.000915 | 0.30773 | 1.2937 | 2.5478 | 99.58 | 20.84 | 37.455 | 83.1146 | 99.6614 | 78.6401 | 0.009273 | 1.0036 | 0.023962 | 23.8768 | 0.017979 | 0.15563 | 0.006051 | 0.43346 | 0.43346 | 0.014757
2019-03-21 11:51 | 1553183457.385 | 344 | 0.2083 | 0.001474 | 2.9e-05 | 99.7687 | 99.6933 | 119.5533 | 21.0729 | 0.002602 | 0.002368 | 167.0707 | 0.2623 | 8.4049 | 0.003198 | 0.003991 | 0.00389 | 5.3e-05 | 0.00091 | 0.31598 | 1.3202 | 2.5481 | 99.58 | 20.84 | 37.455 | 84.2251 | 98.5851 | 78.6109 | 0.009273 | 1.0037 | 0.023961 | 23.8735 | 0.006769 | 0.15604 | 0.006211 | 0.44496 | 0.44496 | 0.014413
2019-03-21 11:51 | 1553183458.3852 | 345 | 0.20831 | 0.001475 | 2.9e-05 | 99.7506 | 99.9075 | 119.2106 | 21.0721 | 0.002603 | 0.002429 | 167.0487 | 0.26171 | 8.4031 | 0.003062 | 0.004694 | 0.004595 | 5.5e-05 | 0.00091 | 0.3402 | 1.3059 | 2.5479 | 99.58 | 20.84 | 37.455 | 84.2469 | 95.3054 | 80.0672 | 0.009273 | 0.99951 | 0.023877 | 23.8885 | -0.014061 | 0.15569 | 0.006104 | 0.43576 | 0.43576 | 0.014685
2019-03-21 11:51 | 1553183459.3862 | 346 | 0.20833 | 0.001481 | 2.9e-05 | 99.7675 | 99.634 | 119.3202 | 21.0721 | 0.002549 | 0.002367 | 167.0202 | 0.26097 | 8.4018 | 0.003165 | 0.004336 | 0.004238 | 4.9e-05 | 0.00091 | 0.31974 | 1.3446 | 2.548 | 99.58 | 20.84 | 37.454 | 82.966 | 97.8044 | 78.586 | 0.009273 | 1.0037 | 0.023938 | 23.8494 | 0.011981 | 0.15525 | 0.005992 | 0.4289 | 0.4289 | 0.014877
2019-03-21 11:51 | 1553183460.3848 | 347 | 0.20835 | 0.001473 | 2.9e-05 | 99.7079 | 99.5853 | 119.7441 | 21.073 | 0.002609 | 0.002428 | 167.0083 | 0.26182 | 8.4007 | 0.003118 | 0.004428 | 0.004278 | 5.1e-05 | 0.000912 | 0.3027 | 1.3159 | 2.548 | 99.58 | 20.84 | 37.453 | 84.3976 | 96.6627 | 80.0434 | 0.009273 | 0.99958 | 0.023931 | 23.9411 | 0.011011 | 0.15576 | 0.00587 | 0.41996 | 0.41996 | 0.015243
2019-03-21 11:51 | 1553183461.3857 | 348 | 0.20835 | 0.001468 | 2.9e-05 | 99.736 | 99.6156 | 119.8523 | 21.073 | 0.002568 | 0.0024 | 166.9742 | 0.26126 | 8.4024 | 0.003196 | 0.004918 | 0.004864 | 5.5e-05 | 0.000912 | 0.31061 | 1.3134 | 2.5479 | 99.58 | 20.84 | 37.452 | 83.403 | 98.5412 | 79.3705 | 0.009272 | 1.0015 | 0.023965 | 23.9291 | 0.0108 | 0.15543 | 0.005877 | 0.42107 | 0.42107 | 0.015171
2019-03-21 11:51 | 1553183462.3862 | 349 | 0.2083 | 0.001469 | 2.8e-05 | 99.6695 | 99.8302 | 119.8338 | 21.0723 | 0.002614 | 0.002363 | 166.9401 | 0.26133 | 8.4039 | 0.003061 | 0.00513 | 0.004979 | 5.2e-05 | 0.00091 | 0.31293 | 1.3283 | 2.5479 | 99.58 | 20.84 | 37.451 | 84.5188 | 95.2875 | 78.4928 | 0.009272 | 1.004 | 0.023993 | 23.8975 | -0.014406 | 0.15547 | 0.006202 | 0.44505 | 0.44505 | 0.014357
2019-03-21 11:51 | 1553183463.3853 | 350 | 0.20833 | 0.001467 | 2.8e-05 | 99.691 | 99.621 | 119.5704 | 21.0731 | 0.002576 | 0.002392 | 166.9261 | 0.26116 | 8.4052 | 0.00316 | 0.00487 | 0.004891 | 5.3e-05 | 0.000907 | 0.27793 | 1.362 | 2.5479 | 99.58 | 20.84 | 37.45 | 83.5965 | 97.6667 | 79.1733 | 0.009272 | 1.002 | 0.023943 | 23.8943 | 0.006276 | 0.15536 | 0.006009 | 0.43023 | 0.43023 | 0.014842
2019-03-21 11:51 | 1553183464.386 | 351 | 0.20836 | 0.001463 | 2.9e-05 | 99.6677 | 99.6691 | 119.3134 | 21.0729 | 0.002527 | 0.002382 | 166.9036 | 0.2615 | 8.4025 | 0.003187 | 0.004817 | 0.004706 | 5e-05 | 0.00091 | 0.30093 | 1.3558 | 2.548 | 99.58 | 20.84 | 37.449 | 82.4248 | 98.3273 | 78.9548 | 0.009272 | 1.0027 | 0.023925 | 23.8612 | -0.000121 | 0.15557 | 0.005854 | 0.41873 | 0.41873 | 0.01527
2019-03-21 11:51 | 1553183465.3864 | 352 | 0.20837 | 0.001459 | 2.9e-05 | 99.6883 | 99.6609 | 119.4286 | 21.0729 | 0.0026 | 0.002403 | 166.8717 | 0.2622 | 8.4024 | 0.003194 | 0.004871 | 0.004842 | 5.2e-05 | 0.00091 | 0.30486 | 1.326 | 2.548 | 99.58 | 20.84 | 37.446 | 84.1917 | 98.4933 | 79.4382 | 0.009271 | 1.0013 | 0.02392 | 23.8891 | 0.002458 | 0.15598 | 0.005786 | 0.41377 | 0.41377 | 0.015494
2019-03-21 11:51 | 1553183466.3861 | 353 | 0.20838 | 0.001462 | 2.9e-05 | 99.66 | 99.6718 | 119.1301 | 21.0734 | 0.002614 | 0.00236 | 166.8304 | 0.26162 | 8.4041 | 0.003162 | 0.004606 | 0.004601 | 5.1e-05 | 0.000911 | 0.29623 | 1.35 | 2.548 | 99.58 | 20.84 | 37.446 | 84.53 | 97.7288 | 78.4143 | 0.009271 | 1.0042 | 0.023925 | 23.8245 | -0.00106 | 0.15564 | 0.005737 | 0.41037 | 0.41037 | 0.015588
2019-03-21 11:51 | 1553183467.3862 | 354 | 0.20837 | 0.001458 | 2.9e-05 | 99.6728 | 99.6677 | 119.1402 | 21.0742 | 0.002552 | 0.002384 | 166.7976 | 0.26109 | 8.4016 | 0.003163 | 0.004719 | 0.00455 | 5e-05 | 0.000907 | 0.34235 | 1.3147 | 2.5479 | 99.58 | 20.84 | 37.444 | 83.0258 | 97.7429 | 78.9834 | 0.00927 | 1.0026 | 0.023907 | 23.8448 | 0.000456 | 0.15532 | 0.005807 | 0.41508 | 0.41508 | 0.015379
2019-03-21 11:51 | 1553183468.3852 | 355 | 0.20837 | 0.001454 | 2.8e-05 | 99.6302 | 99.7358 | 119.4803 | 21.072 | 0.002516 | 0.002368 | 166.7579 | 0.26133 | 8.4016 | 0.003111 | 0.004694 | 0.004553 | 5.2e-05 | 0.000909 | 0.33268 | 1.2987 | 2.5479 | 99.58 | 20.84 | 37.443 | 82.1681 | 96.4997 | 78.6086 | 0.00927 | 1.0037 | 0.023953 | 23.8661 | -0.009479 | 0.15547 | 0.005788 | 0.41452 | 0.41452 | 0.015415
2019-03-21 11:51 | 1553183469.3852 | 356 | 0.20839 | 0.001447 | 2.9e-05 | 99.595 | 99.5713 | 119.3512 | 21.0726 | 0.002536 | 0.002345 | 166.7382 | 0.26212 | 8.402 | 0.003149 | 0.004426 | 0.004377 | 5.2e-05 | 0.000907 | 0.34346 | 1.317 | 2.548 | 99.58 | 20.84 | 37.442 | 82.6554 | 97.413 | 78.0658 | 0.00927 | 1.0052 | 0.023959 | 23.8348 | 0.002129 | 0.15594 | 0.005678 | 0.40672 | 0.40672 | 0.015758
2019-03-21 11:51 | 1553183470.3857 | 357 | 0.20839 | 0.001448 | 2.9e-05 | 99.6424 | 99.5693 | 119.3866 | 21.0731 | 0.002551 | 0.00238 | 166.6936 | 0.26186 | 8.4057 | 0.003196 | 0.004448 | 0.004391 | 5.2e-05 | 0.000908 | 0.31293 | 1.3149 | 2.5479 | 99.58 | 20.84 | 37.443 | 83.0045 | 98.5375 | 78.8886 | 0.00927 | 1.0029 | 0.023935 | 23.8663 | 0.006566 | 0.15578 | 0.005709 | 0.40854 | 0.40854 | 0.015672
2019-03-21 11:51 | 1553183471.3844 | 358 | 0.20839 | 0.001447 | 2.9e-05 | 99.6064 | 99.5993 | 119.4261 | 21.0729 | 0.002527 | 0.002363 | 166.655 | 0.26236 | 8.404 | 0.003257 | 0.004867 | 0.004768 | 5.5e-05 | 0.000909 | 0.31332 | 1.2773 | 2.5479 | 99.58 | 20.84 | 37.444 | 82.4365 | 100.0267 | 78.4816 | 0.00927 | 1.004 | 0.023952 | 23.8564 | 0.000638 | 0.15608 | 0.005656 | 0.40501 | 0.40501 | 0.015839
2019-03-21 11:51 | 1553183472.3854 | 359 | 0.20843 | 0.001443 | 2.9e-05 | 99.6055 | 99.681 | 119.1167 | 21.0737 | 0.002553 | 0.00231 | 166.6395 | 0.2618 | 8.4038 | 0.003097 | 0.004771 | 0.004646 | 5e-05 | 0.000908 | 0.32278 | 1.2959 | 2.548 | 99.58 | 20.84 | 37.444 | 83.051 | 96.1481 | 77.2178 | 0.00927 | 1.0076 | 0.023964 | 23.7826 | -0.006781 | 0.15575 | 0.005461 | 0.39116 | 0.39116 | 0.016365
2019-03-21 11:51 | 1553183473.3859 | 360 | 0.20839 | 0.001441 | 2.9e-05 | 99.5664 | 99.574 | 119.2219 | 21.0743 | 0.002561 | 0.002419 | 166.6064 | 0.26138 | 8.4007 | 0.003255 | 0.005049 | 0.00496 | 5.2e-05 | 0.000909 | 0.32068 | 1.2812 | 2.548 | 99.58 | 20.84 | 37.446 | 83.2546 | 99.9725 | 79.8426 | 0.009271 | 1.0002 | 0.023886 | 23.8821 | -0.000685 | 0.1555 | 0.005675 | 0.40523 | 0.40523 | 0.015771
2019-03-21 11:51 | 1553183474.3846 | 361 | 0.2084 | 0.001439 | 3e-05 | 99.6046 | 99.4242 | 119.5045 | 21.0732 | 0.0026 | 0.002401 | 166.5647 | 0.26126 | 8.4006 | 0.003182 | 0.005114 | 0.005037 | 5.2e-05 | 0.000911 | 0.32493 | 1.3135 | 2.5479 | 99.58 | 20.84 | 37.446 | 84.1778 | 98.2208 | 79.4017 | 0.009271 | 1.0014 | 0.023929 | 23.8954 | 0.016212 | 0.15543 | 0.005641 | 0.40349 | 0.40349 | 0.015832
2019-03-21 11:51 | 1553183475.3856 | 362 | 0.20839 | 0.001437 | 3.1e-05 | 99.6829 | 99.4429 | 119.3134 | 21.0739 | 0.002492 | 0.002397 | 166.5469 | 0.26158 | 8.4019 | 0.003183 | 0.004982 | 0.004887 | 5.2e-05 | 0.000911 | 0.32156 | 1.2843 | 2.548 | 99.58 | 20.84 | 37.447 | 81.5951 | 98.2294 | 79.3043 | 0.009271 | 1.0017 | 0.023913 | 23.873 | 0.021561 | 0.15562 | 0.005681 | 0.40605 | 0.40605 | 0.015751
2019-03-21 11:51 | 1553183476.385 | 363 | 0.20841 | 0.001438 | 3.1e-05 | 99.5682 | 99.3697 | 119.5606 | 21.074 | 0.002511 | 0.002378 | 166.5337 | 0.26154 | 8.4019 | 0.003197 | 0.004956 | 0.004888 | 5.3e-05 | 0.000914 | 0.32029 | 1.26 | 2.5477 | 99.58 | 20.84 | 37.449 | 82.046 | 98.5628 | 78.8436 | 0.009272 | 1.003 | 0.023953 | 23.8821 | 0.017851 | 0.15559 | 0.005548 | 0.39716 | 0.39716 | 0.016101
2019-03-21 11:51 | 1553183477.3857 | 364 | 0.2084 | 0.001429 | 3e-05 | 99.5727 | 99.5012 | 119.8263 | 21.0743 | 0.002564 | 0.00239 | 166.5223 | 0.26091 | 8.4023 | 0.003186 | 0.004379 | 0.004224 | 5.2e-05 | 0.000915 | 0.32333 | 1.2991 | 2.5479 | 99.58 | 20.84 | 37.451 | 83.3248 | 98.2949 | 79.1315 | 0.009272 | 1.0022 | 0.02397 | 23.9184 | 0.006427 | 0.15522 | 0.005642 | 0.40421 | 0.40421 | 0.015783
2019-03-21 11:51 | 1553183478.3857 | 365 | 0.20839 | 0.00143 | 3e-05 | 99.4928 | 99.541 | 119.7929 | 21.0751 | 0.002561 | 0.002314 | 166.5059 | 0.26208 | 8.402 | 0.003207 | 0.005018 | 0.00495 | 5.2e-05 | 0.000915 | 0.3418 | 1.3396 | 2.5479 | 99.58 | 20.84 | 37.452 | 83.256 | 98.8077 | 77.305 | 0.009272 | 1.0074 | 0.024029 | 23.853 | -0.00433 | 0.15591 | 0.005703 | 0.40962 | 0.40962 | 0.015643
2019-03-21 11:51 | 1553183479.3851 | 366 | 0.20842 | 0.001423 | 3.1e-05 | 99.5009 | 99.3261 | 120.0152 | 21.0743 | 0.002521 | 0.002417 | 166.4811 | 0.26193 | 8.4022 | 0.003242 | 0.005213 | 0.005095 | 5.2e-05 | 0.000915 | 0.32587 | 1.3566 | 2.548 | 99.58 | 20.84 | 37.452 | 82.2873 | 99.6704 | 79.7748 | 0.009272 | 1.0003 | 0.023967 | 23.9591 | 0.015725 | 0.15582 | 0.00556 | 0.3982 | 0.3982 | 0.016083
2019-03-21 11:51 | 1553183480.3848 | 367 | 0.20838 | 0.00142 | 3e-05 | 99.5813 | 99.4984 | 119.8868 | 21.0748 | 0.002576 | 0.002353 | 166.4707 | 0.26224 | 8.4002 | 0.00322 | 0.005546 | 0.005464 | 5.2e-05 | 0.00092 | 0.31293 | 1.3483 | 2.548 | 99.58 | 20.84 | 37.455 | 83.5974 | 99.1352 | 78.2557 | 0.009273 | 1.0047 | 0.024006 | 23.8947 | 0.007448 | 0.15601 | 0.00575 | 0.41259 | 0.41259 | 0.01554
2019-03-21 11:51 | 1553183481.3849 | 368 | 0.20839 | 0.001418 | 3.1e-05 | 99.4824 | 99.4681 | 119.8899 | 21.0747 | 0.002543 | 0.002436 | 166.4414 | 0.26213 | 8.4044 | 0.003213 | 0.005273 | 0.005198 | 5.4e-05 | 0.000922 | 0.3095 | 1.3613 | 2.548 | 99.58 | 20.84 | 37.455 | 82.8251 | 98.9485 | 80.2344 | 0.009273 | 0.99904 | 0.023939 | 23.9622 | 0.001284 | 0.15594 | 0.005705 | 0.40817 | 0.40817 | 0.015702
2019-03-21 11:51 | 1553183482.385 | 369 | 0.20842 | 0.00141 | 3.2e-05 | 99.4949 | 99.3632 | 120.161 | 21.0753 | 0.00257 | 0.002367 | 166.432 | 0.26297 | 8.4043 | 0.003265 | 0.005426 | 0.005379 | 5.4e-05 | 0.000921 | 0.33196 | 1.3232 | 2.5479 | 99.58 | 20.83 | 37.456 | 83.4755 | 100.2223 | 78.5822 | 0.009268 | 1.0037 | 0.024022 | 23.9331 | 0.011847 | 0.15644 | 0.005566 | 0.39948 | 0.39948 | 0.016095
2019-03-21 11:51 | 1553183483.3859 | 370 | 0.2084 | 0.00141 | 3.2e-05 | 99.5703 | 99.4136 | 120.3504 | 21.075 | 0.002566 | 0.002364 | 166.4213 | 0.26163 | 8.4015 | 0.003222 | 0.004476 | 0.004321 | 5.7e-05 | 0.000924 | 0.31437 | 1.3276 | 2.548 | 99.58 | 20.83 | 37.456 | 83.3737 | 99.1825 | 78.5108 | 0.009268 | 1.0039 | 0.024044 | 23.9495 | 0.014085 | 0.15564 | 0.005653 | 0.40613 | 0.40613 | 0.015751
2019-03-21 11:51 | 1553183484.3851 | 371 | 0.20839 | 0.001408 | 3.3e-05 | 99.5214 | 99.5168 | 120.2725 | 21.0759 | 0.002564 | 0.002385 | 166.3881 | 0.2619 | 8.4045 | 0.003262 | 0.004746 | 0.004685 | 5.8e-05 | 0.000923 | 0.28512 | 1.343 | 2.5479 | 99.58 | 20.83 | 37.457 | 83.3114 | 100.153 | 79.0115 | 0.009268 | 1.0025 | 0.024019 | 23.9588 | 0.000414 | 0.1558 | 0.00575 | 0.41263 | 0.41263 | 0.015519
2019-03-21 11:51 | 1553183485.3841 | 372 | 0.20839 | 0.001406 | 3.3e-05 | 99.4842 | 99.4552 | 120.5659 | 21.076 | 0.002614 | 0.002368 | 166.3652 | 0.26181 | 8.4017 | 0.003326 | 0.004781 | 0.004738 | 5.8e-05 | 0.000926 | 0.32405 | 1.2872 | 2.548 | 99.58 | 20.83 | 37.457 | 84.5215 | 101.6992 | 78.6047 | 0.009268 | 1.0037 | 0.024062 | 23.9742 | 0.002609 | 0.15575 | 0.005762 | 0.41427 | 0.41427 | 0.015452
2019-03-21 11:51 | 1553183486.3852 | 373 | 0.20838 | 0.001403 | 3.4e-05 | 99.4487 | 99.5189 | 120.6121 | 21.0759 | 0.002462 | 0.00234 | 166.3532 | 0.26219 | 8.402 | 0.003243 | 0.004945 | 0.004915 | 5.8e-05 | 0.000927 | 0.31758 | 1.2764 | 2.5479 | 99.58 | 20.83 | 37.457 | 80.8616 | 99.6915 | 77.9295 | 0.009268 | 1.0056 | 0.02409 | 23.9557 | -0.006306 | 0.15598 | 0.005798 | 0.41722 | 0.41722 | 0.015365
2019-03-21 11:51 | 1553183487.3852 | 374 | 0.2084 | 0.001395 | 3.5e-05 | 99.525 | 99.4869 | 120.5948 | 21.0762 | 0.002536 | 0.002342 | 166.3312 | 0.26087 | 8.4019 | 0.00331 | 0.005138 | 0.005065 | 5.4e-05 | 0.000928 | 0.32244 | 1.3079 | 2.548 | 99.58 | 20.83 | 37.456 | 82.6558 | 101.3156 | 77.9774 | 0.009268 | 1.0055 | 0.024086 | 23.9556 | 0.00343 | 0.15519 | 0.005663 | 0.40746 | 0.40746 | 0.015654
2019-03-21 11:51 | 1553183488.3842 | 375 | 0.20842 | 0.001398 | 3.5e-05 | 99.5217 | 99.3847 | 120.8529 | 21.076 | 0.002526 | 0.002373 | 166.2997 | 0.26139 | 8.4004 | 0.003253 | 0.00481 | 0.004721 | 5.7e-05 | 0.000931 | 0.3147 | 1.2691 | 2.5479 | 99.58 | 20.83 | 37.456 | 82.4121 | 99.9287 | 78.7223 | 0.009268 | 1.0033 | 0.024087 | 24.0067 | 0.012326 | 0.1555 | 0.005577 | 0.40118 | 0.40118 | 0.015931
2019-03-21 11:51 | 1553183489.3847 | 376 | 0.20842 | 0.001391 | 3.5e-05 | 99.5256 | 99.4589 | 120.5286 | 21.0758 | 0.002621 | 0.002329 | 166.289 | 0.26142 | 8.4004 | 0.003344 | 0.004747 | 0.004608 | 5.8e-05 | 0.000932 | 0.30326 | 1.287 | 2.5479 | 99.58 | 20.83 | 37.457 | 84.686 | 102.1223 | 77.6643 | 0.009268 | 1.0064 | 0.024091 | 23.9384 | 0.005995 | 0.15552 | 0.005594 | 0.40251 | 0.40251 | 0.01588
2019-03-21 11:51 | 1553183490.3844 | 377 | 0.20839 | 0.001389 | 3.6e-05 | 99.4913 | 99.3023 | 120.2603 | 21.0757 | 0.00251 | 0.002382 | 166.2427 | 0.26228 | 8.4016 | 0.003305 | 0.00596 | 0.005826 | 5.4e-05 | 0.00093 | 0.31868 | 1.3057 | 2.548 | 99.58 | 20.83 | 37.458 | 82.0195 | 101.1973 | 78.9366 | 0.009268 | 1.0027 | 0.02402 | 23.9551 | 0.017016 | 0.15603 | 0.005727 | 0.41087 | 0.41087 | 0.015608
2019-03-21 11:52 | 1553183491.3854 | 378 | 0.20842 | 0.001391 | 3.6e-05 | 99.4267 | 99.3053 | 120.0497 | 21.0757 | 0.002511 | 0.002431 | 166.2171 | 0.26253 | 8.4022 | 0.003433 | 0.005319 | 0.005366 | 5.5e-05 | 0.000931 | 0.32228 | 1.3327 | 2.5479 | 99.58 | 20.83 | 37.458 | 82.0344 | 104.281 | 80.1215 | 0.009268 | 0.99936 | 0.023959 | 23.9743 | 0.010925 | 0.15618 | 0.005571 | 0.39861 | 0.39861 | 0.016103
2019-03-21 11:52 | 1553183492.3855 | 379 | 0.20838 | 0.001382 | 3.6e-05 | 99.4881 | 99.2907 | 120.3394 | 21.0768 | 0.00253 | 0.00235 | 166.1788 | 0.26192 | 8.401 | 0.003335 | 0.004631 | 0.004446 | 5.6e-05 | 0.000929 | 0.30414 | 1.3496 | 2.548 | 99.58 | 20.84 | 37.458 | 82.5088 | 101.9089 | 78.1759 | 0.009274 | 1.0049 | 0.024054 | 23.937 | 0.017764 | 0.15582 | 0.005818 | 0.41796 | 0.41796 | 0.015322
2019-03-21 11:52 | 1553183493.3856 | 380 | 0.20841 | 0.001386 | 3.7e-05 | 99.4666 | 99.2699 | 120.4633 | 21.0761 | 0.002525 | 0.002336 | 166.1523 | 0.26277 | 8.4021 | 0.003378 | 0.005151 | 0.005139 | 5.9e-05 | 0.000934 | 0.33085 | 1.3496 | 2.5478 | 99.58 | 20.84 | 37.458 | 82.3802 | 102.9483 | 77.8453 | 0.009274 | 1.0058 | 0.024078 | 23.9381 | 0.017707 | 0.15632 | 0.005619 | 0.40397 | 0.40397 | 0.015904
2019-03-21 11:52 | 1553183494.3846 | 381 | 0.20841 | 0.001378 | 3.7e-05 | 99.4839 | 99.2757 | 120.5621 | 21.0772 | 0.002569 | 0.002326 | 166.1301 | 0.26247 | 8.4028 | 0.003462 | 0.004779 | 0.004667 | 5.7e-05 | 0.000931 | 0.28296 | 1.3298 | 2.5476 | 99.58 | 20.84 | 37.458 | 83.4437 | 104.9987 | 77.6008 | 0.009274 | 1.0065 | 0.024096 | 23.9396 | 0.018739 | 0.15614 | 0.005675 | 0.40831 | 0.40831 | 0.015717
2019-03-21 11:52 | 1553183495.3857 | 382 | 0.20844 | 0.001379 | 3.8e-05 | 99.4931 | 99.2471 | 120.392 | 21.0767 | 0.002479 | 0.002317 | 166.088 | 0.26273 | 8.4012 | 0.003368 | 0.004832 | 0.004724 | 5.9e-05 | 0.000937 | 0.35977 | 1.317 | 2.5479 | 99.58 | 20.84 | 37.458 | 81.2758 | 102.7053 | 77.3736 | 0.009274 | 1.0072 | 0.024087 | 23.9149 | 0.022147 | 0.1563 | 0.005466 | 0.39296 | 0.39296 | 0.016348
2019-03-21 11:52 | 1553183496.3857 | 383 | 0.20841 | 0.001375 | 3.8e-05 | 99.4481 | 99.54 | 120.4382 | 21.0763 | 0.00257 | 0.002318 | 166.0461 | 0.26171 | 8.4043 | 0.003335 | 0.004731 | 0.0047 | 5.8e-05 | 0.000935 | 0.32925 | 1.3255 | 2.5479 | 99.58 | 20.84 | 37.456 | 83.4772 | 101.9094 | 77.4176 | 0.009273 | 1.0071 | 0.02409 | 23.921 | -0.008257 | 0.15569 | 0.005647 | 0.40609 | 0.40609 | 0.015758
2019-03-21 11:52 | 1553183497.3852 | 384 | 0.20842 | 0.001371 | 3.9e-05 | 99.5483 | 99.4521 | 120.2171 | 21.0776 | 0.002522 | 0.00238 | 166.0102 | 0.26129 | 8.4037 | 0.003379 | 0.005076 | 0.004994 | 5.6e-05 | 0.000936 | 0.31293 | 1.3092 | 2.548 | 99.58 | 20.84 | 37.454 | 82.3021 | 102.9899 | 78.8984 | 0.009273 | 1.0028 | 0.024017 | 23.9495 | 0.008642 | 0.15544 | 0.005581 | 0.40009 | 0.40009 | 0.015968
2019-03-21 11:52 | 1553183498.3838 | 385 | 0.20845 | 0.001364 | 4e-05 | 99.4574 | 99.4129 | 119.7479 | 21.0773 | 0.002484 | 0.002399 | 165.9732 | 0.26219 | 8.4016 | 0.003415 | 0.004634 | 0.00458 | 5.7e-05 | 0.00093 | 0.30525 | 1.2955 | 2.5479 | 99.58 | 20.84 | 37.455 | 81.3879 | 103.8494 | 79.3622 | 0.009273 | 1.0015 | 0.023955 | 23.9184 | 0.003997 | 0.15598 | 0.00543 | 0.38812 | 0.38812 | 0.016517
2019-03-21 11:52 | 1553183499.3848 | 386 | 0.20845 | 0.001362 | 3.9e-05 | 99.4824 | 99.3421 | 119.7265 | 21.0778 | 0.002533 | 0.002314 | 165.9375 | 0.2628 | 8.4035 | 0.003338 | 0.004564 | 0.00445 | 5.7e-05 | 0.00093 | 0.30862 | 1.2898 | 2.5478 | 99.58 | 20.84 | 37.454 | 82.5737 | 101.9791 | 77.3232 | 0.009273 | 1.0073 | 0.024022 | 23.847 | 0.012623 | 0.15634 | 0.005445 | 0.39034 | 0.39034 | 0.016461
2019-03-21 11:52 | 1553183500.3853 | 387 | 0.20841 | 0.001362 | 4e-05 | 99.539 | 99.2873 | 119.8474 | 21.077 | 0.002537 | 0.002372 | 165.9192 | 0.26204 | 8.4039 | 0.003379 | 0.004722 | 0.004694 | 5.7e-05 | 0.000929 | 0.31581 | 1.2967 | 2.5479 | 99.58 | 20.84 | 37.453 | 82.6838 | 102.9853 | 78.7007 | 0.009273 | 1.0034 | 0.023987 | 23.9059 | 0.02265 | 0.15589 | 0.005672 | 0.40609 | 0.40609 | 0.015778
2019-03-21 11:52 | 1553183501.385 | 388 | 0.20844 | 0.001357 | 4.1e-05 | 99.5331 | 99.4177 | 120.1023 | 21.078 | 0.002539 | 0.002382 | 165.9038 | 0.2627 | 8.4024 | 0.003371 | 0.005047 | 0.004914 | 5.6e-05 | 0.00093 | 0.29463 | 1.2869 | 2.5479 | 99.58 | 20.84 | 37.452 | 82.7184 | 102.7814 | 78.9496 | 0.009272 | 1.0027 | 0.024004 | 23.9398 | 0.010371 | 0.15628 | 0.005463 | 0.39126 | 0.39126 | 0.016417
2019-03-21 11:52 | 1553183502.3851 | 389 | 0.20846 | 0.001354 | 4e-05 | 99.5798 | 99.3612 | 119.8507 | 21.0774 | 0.002519 | 0.002337 | 165.9118 | 0.26117 | 8.4042 | 0.00328 | 0.005015 | 0.004962 | 5.8e-05 | 0.000931 | 0.30829 | 1.2926 | 2.5478 | 99.58 | 20.84 | 37.451 | 82.2404 | 100.5859 | 77.8677 | 0.009272 | 1.0058 | 0.024016 | 23.8779 | 0.019662 | 0.15537 | 0.005407 | 0.38742 | 0.38742 | 0.016483
2019-03-21 11:52 | 1553183503.3841 | 390 | 0.20844 | 0.001352 | 4.1e-05 | 99.5748 | 99.3363 | 119.4747 | 21.0782 | 0.002528 | 0.00239 | 165.8755 | 0.26228 | 8.399 | 0.003295 | 0.004662 | 0.004574 | 6.1e-05 | 0.000937 | 0.32117 | 1.3152 | 2.5479 | 99.58 | 20.85 | 37.448 | 82.4542 | 100.9517 | 79.1454 | 0.009277 | 1.0021 | 0.023935 | 23.8838 | 0.021445 | 0.15603 | 0.005491 | 0.39211 | 0.39211 | 0.016354
2019-03-21 11:52 | 1553183504.3841 | 391 | 0.20843 | 0.001352 | 4.1e-05 | 99.5614 | 99.4593 | 119.5376 | 21.0777 | 0.002528 | 0.00237 | 165.8545 | 0.26232 | 8.4028 | 0.003298 | 0.004448 | 0.004352 | 5.9e-05 | 0.00094 | 0.33931 | 1.2988 | 2.5479 | 99.58 | 20.85 | 37.447 | 82.4482 | 101.0282 | 78.6538 | 0.009277 | 1.0035 | 0.023958 | 23.8734 | 0.009177 | 0.15605 | 0.00555 | 0.39676 | 0.39676 | 0.016166
2019-03-21 11:52 | 1553183505.3855 | 392 | 0.20854 | 0.001354 | 4.1e-05 | 99.5328 | 99.399 | 119.5282 | 21.0784 | 0.002556 | 0.002313 | 165.8447 | 0.26298 | 8.4031 | 0.003282 | 0.004109 | 0.004089 | 5.9e-05 | 0.00094 | 0.34274 | 1.2905 | 2.5479 | 99.58 | 20.85 | 37.441 | 83.1219 | 100.6316 | 77.2856 | 0.009275 | 1.0074 | 0.024003 | 23.826 | 0.01203 | 0.15645 | 0.004904 | 0.35099 | 0.35099 | 0.01832
2019-03-21 11:52 | 1553183506.3853 | 393 | 0.20847 | 0.001352 | 4.2e-05 | 99.5176 | 99.3036 | 119.4322 | 21.0778 | 0.00259 | 0.002347 | 165.814 | 0.26286 | 8.4006 | 0.003359 | 0.004414 | 0.004247 | 5.9e-05 | 0.00094 | 0.30718 | 1.3277 | 2.5479 | 99.58 | 20.85 | 37.439 | 83.9365 | 102.4974 | 78.1175 | 0.009275 | 1.0051 | 0.023965 | 23.8446 | 0.01925 | 0.15637 | 0.005311 | 0.37966 | 0.37966 | 0.016928
2019-03-21 11:52 | 1553183507.3853 | 394 | 0.20847 | 0.001347 | 4.2e-05 | 99.4931 | 99.292 | 119.2057 | 21.0777 | 0.002527 | 0.002373 | 165.7862 | 0.26262 | 8.4021 | 0.00331 | 0.004956 | 0.004973 | 6.1e-05 | 0.000944 | 0.32156 | 1.3654 | 2.548 | 99.58 | 20.85 | 37.435 | 82.4395 | 101.3178 | 78.728 | 0.009274 | 1.0033 | 0.023922 | 23.8427 | 0.018097 | 0.15624 | 0.005295 | 0.37776 | 0.37776 | 0.016998
2019-03-21 11:52 | 1553183508.3852 | 395 | 0.2085 | 0.001347 | 4.2e-05 | 99.5527 | 99.5413 | 119.3258 | 21.0781 | 0.002537 | 0.00235 | 165.7931 | 0.26197 | 8.399 | 0.003329 | 0.004734 | 0.004582 | 5.8e-05 | 0.000946 | 0.34379 | 1.3546 | 2.548 | 99.58 | 20.85 | 37.432 | 82.6834 | 101.7564 | 78.1741 | 0.009273 | 1.0049 | 0.023953 | 23.836 | 0.001022 | 0.15584 | 0.005134 | 0.36668 | 0.36668 | 0.017468
2019-03-21 11:52 | 1553183509.3853 | 396 | 0.20844 | 0.001338 | 4.3e-05 | 99.5006 | 99.4593 | 119.566 | 21.078 | 0.002555 | 0.002379 | 165.7823 | 0.26221 | 8.401 | 0.00334 | 0.005244 | 0.00519 | 5.7e-05 | 0.000945 | 0.30055 | 1.3482 | 2.5479 | 99.58 | 20.85 | 37.428 | 83.1076 | 102.0461 | 78.8748 | 0.009272 | 1.0029 | 0.023953 | 23.8837 | 0.003714 | 0.15599 | 0.005547 | 0.39631 | 0.39631 | 0.016177
2019-03-21 11:52 | 1553183510.3854 | 397 | 0.20848 | 0.001339 | 4.3e-05 | 99.6105 | 99.4817 | 119.6596 | 21.0783 | 0.002597 | 0.002391 | 165.7721 | 0.2624 | 8.4021 | 0.003249 | 0.004702 | 0.004624 | 6e-05 | 0.000944 | 0.32173 | 1.3102 | 2.5477 | 99.58 | 20.85 | 37.424 | 84.1065 | 99.8453 | 79.1736 | 0.009271 | 1.002 | 0.023952 | 23.9032 | 0.011571 | 0.1561 | 0.005275 | 0.37671 | 0.37671 | 0.017031
2019-03-21 11:52 | 1553183511.3846 | 398 | 0.20847 | 0.001337 | 4.4e-05 | 99.5626 | 99.4017 | 119.5969 | 21.0785 | 0.002519 | 0.002344 | 165.7388 | 0.26168 | 8.4032 | 0.003303 | 0.00466 | 0.004605 | 5.9e-05 | 0.000944 | 0.32117 | 1.2673 | 2.5478 | 99.58 | 20.85 | 37.418 | 82.2407 | 101.1373 | 78.0372 | 0.00927 | 1.0053 | 0.023985 | 23.8584 | 0.014461 | 0.15567 | 0.005317 | 0.38025 | 0.38025 | 0.016826
2019-03-21 11:52 | 1553183512.3855 | 399 | 0.20844 | 0.001338 | 4.4e-05 | 99.584 | 99.4388 | 119.7892 | 21.0785 | 0.002533 | 0.002366 | 165.6822 | 0.26188 | 8.4021 | 0.003252 | 0.004895 | 0.004844 | 5.9e-05 | 0.000942 | 0.30663 | 1.327 | 2.548 | 99.58 | 20.85 | 37.413 | 82.5761 | 99.8989 | 78.5586 | 0.009268 | 1.0038 | 0.023986 | 23.8953 | 0.013049 | 0.15579 | 0.005525 | 0.39521 | 0.39521 | 0.016202
2019-03-21 11:52 | 1553183513.3856 | 400 | 0.20849 | 0.001329 | 4.4e-05 | 99.5664 | 99.2985 | 119.9047 | 21.0783 | 0.002528 | 0.002398 | 165.6751 | 0.26258 | 8.4011 | 0.003244 | 0.004834 | 0.004749 | 5.8e-05 | 0.000946 | 0.32908 | 1.3072 | 2.5479 | 99.58 | 20.86 | 37.405 | 82.4588 | 99.7159 | 79.3358 | 0.009272 | 1.0016 | 0.023971 | 23.9332 | 0.0241 | 0.15621 | 0.005247 | 0.37495 | 0.37495 | 0.017123
2019-03-21 11:52 | 1553183514.385 | 401 | 0.2085 | 0.001329 | 4.4e-05 | 99.5623 | 99.5294 | 119.6272 | 21.0779 | 0.002556 | 0.002335 | 165.6511 | 0.26295 | 8.4011 | 0.003197 | 0.005062 | 0.004935 | 6e-05 | 0.000944 | 0.34003 | 1.291 | 2.5479 | 99.57 | 20.86 | 37.402 | 83.127 | 98.5712 | 77.8286 | 0.009272 | 1.0059 | 0.023995 | 23.8543 | 0.002949 | 0.15643 | 0.005168 | 0.36962 | 0.36962 | 0.017394
2019-03-21 11:52 | 1553183515.3847 | 402 | 0.2085 | 0.001326 | 4.5e-05 | 99.5882 | 99.4181 | 119.6045 | 21.078 | 0.002517 | 0.002357 | 165.6262 | 0.26216 | 8.4019 | 0.003224 | 0.004697 | 0.004672 | 5.8e-05 | 0.000944 | 0.31741 | 1.293 | 2.5479 | 99.57 | 20.86 | 37.399 | 82.1964 | 99.2166 | 78.3368 | 0.009272 | 1.0044 | 0.023975 | 23.8693 | 0.015292 | 0.15596 | 0.005175 | 0.36976 | 0.36976 | 0.017336
2019-03-21 11:52 | 1553183516.3837 | 403 | 0.20849 | 0.001323 | 4.6e-05 | 99.5381 | 99.3969 | 119.5721 | 21.0792 | 0.002526 | 0.002386 | 165.6132 | 0.26236 | 8.401 | 0.003289 | 0.004347 | 0.004213 | 5.5e-05 | 0.000946 | 0.34938 | 1.3143 | 2.548 | 99.57 | 20.86 | 37.393 | 82.402 | 100.8027 | 79.0522 | 0.00927 | 1.0024 | 0.023948 | 23.8903 | 0.012695 | 0.15608 | 0.00524 | 0.37402 | 0.37402 | 0.017151
2019-03-21 11:52 | 1553183517.384 | 404 | 0.20851 | 0.00132 | 4.5e-05 | 99.5685 | 99.5441 | 119.8941 | 21.0789 | 0.002504 | 0.002341 | 165.6058 | 0.26295 | 8.4014 | 0.003194 | 0.00468 | 0.004633 | 5.8e-05 | 0.000945 | 0.32068 | 1.3154 | 2.5479 | 99.57 | 20.86 | 37.392 | 81.8923 | 98.5118 | 77.9581 | 0.00927 | 1.0055 | 0.024017 | 23.8853 | 0.002195 | 0.15643 | 0.005131 | 0.36729 | 0.36729 | 0.017504
2019-03-21 11:52 | 1553183518.3839 | 405 | 0.20848 | 0.00132 | 4.6e-05 | 99.5691 | 99.5856 | 120.2739 | 21.0796 | 0.002536 | 0.00236 | 165.5672 | 0.26102 | 8.4025 | 0.003248 | 0.00472 | 0.004721 | 5.4e-05 | 0.000945 | 0.32117 | 1.3111 | 2.5479 | 99.58 | 20.86 | 37.389 | 82.6567 | 99.8207 | 78.411 | 0.009268 | 1.0042 | 0.02404 | 23.9385 | -0.001484 | 0.15528 | 0.005321 | 0.38131 | 0.38131 | 0.016738
2019-03-21 11:52 | 1553183519.3849 | 406 | 0.20853 | 0.001313 | 4.6e-05 | 99.5906 | 99.3121 | 120.2323 | 21.0799 | 0.002548 | 0.002308 | 165.5324 | 0.26231 | 8.4035 | 0.003237 | 0.004597 | 0.004595 | 5.7e-05 | 0.000944 | 0.30143 | 1.3343 | 2.548 | 99.58 | 20.86 | 37.387 | 82.9493 | 99.5424 | 77.1693 | 0.009268 | 1.0078 | 0.024078 | 23.8921 | 0.025041 | 0.15604 | 0.004986 | 0.35769 | 0.35769 | 0.01793
2019-03-21 11:52 | 1553183520.3849 | 407 | 0.20851 | 0.001309 | 4.6e-05 | 99.6537 | 99.5516 | 119.8793 | 21.079 | 0.002563 | 0.00233 | 165.5005 | 0.26267 | 8.4025 | 0.003295 | 0.004513 | 0.004404 | 5.7e-05 | 0.000944 | 0.32587 | 1.3515 | 2.5479 | 99.58 | 20.86 | 37.386 | 83.2918 | 100.9406 | 77.6963 | 0.009268 | 1.0063 | 0.024025 | 23.8749 | 0.009174 | 0.15626 | 0.005154 | 0.36901 | 0.36901 | 0.017404
2019-03-21 11:52 | 1553183521.384 | 408 | 0.20847 | 0.001312 | 4.7e-05 | 99.6046 | 99.5649 | 119.903 | 21.0787 | 0.002503 | 0.002327 | 165.4856 | 0.26206 | 8.4017 | 0.003263 | 0.004912 | 0.004776 | 5.8e-05 | 0.000944 | 0.31868 | 1.3558 | 2.5478 | 99.58 | 20.86 | 37.384 | 81.8587 | 100.169 | 77.6211 | 0.009267 | 1.0065 | 0.02403 | 23.8747 | 0.003567 | 0.1559 | 0.005386 | 0.3858 | 0.3858 | 0.016609
2019-03-21 11:52 | 1553183522.3841 | 409 | 0.20851 | 0.001312 | 4.7e-05 | 99.5721 | 99.336 | 119.7983 | 21.0805 | 0.002474 | 0.002342 | 165.464 | 0.26322 | 8.404 | 0.003302 | 0.004612 | 0.004544 | 5.4e-05 | 0.000941 | 0.31868 | 1.3484 | 2.5479 | 99.58 | 20.86 | 37.384 | 81.1704 | 101.1289 | 77.9909 | 0.009267 | 1.0054 | 0.024006 | 23.8769 | 0.021235 | 0.15659 | 0.005137 | 0.36743 | 0.36743 | 0.017516
2019-03-21 11:52 | 1553183523.3847 | 410 | 0.20848 | 0.001301 | 4.7e-05 | 99.5906 | 99.3281 | 120.2903 | 21.0797 | 0.002447 | 0.002328 | 165.4391 | 0.26275 | 8.4031 | 0.003187 | 0.004519 | 0.004467 | 5.9e-05 | 0.000943 | 0.3246 | 1.325 | 2.548 | 99.58 | 20.86 | 37.385 | 80.5113 | 98.3267 | 77.6477 | 0.009267 | 1.0064 | 0.024067 | 23.9142 | 0.0236 | 0.15631 | 0.005306 | 0.3806 | 0.3806 | 0.016879
2019-03-21 11:52 | 1553183524.3842 | 411 | 0.20853 | 0.001307 | 4.7e-05 | 99.6147 | 99.494 | 120.5708 | 21.08 | 0.002508 | 0.002278 | 165.4188 | 0.26228 | 8.4033 | 0.003269 | 0.004853 | 0.004808 | 5.3e-05 | 0.000943 | 0.33356 | 1.3084 | 2.5479 | 99.58 | 20.86 | 37.385 | 81.9667 | 100.3279 | 76.4414 | 0.009267 | 1.0099 | 0.024137 | 23.9008 | 0.010843 | 0.15603 | 0.005019 | 0.36092 | 0.36092 | 0.017768
2019-03-21 11:52 | 1553183525.3853 | 412 | 0.2085 | 0.001301 | 4.7e-05 | 99.6546 | 99.5386 | 120.5365 | 21.0802 | 0.002488 | 0.002336 | 165.4027 | 0.26193 | 8.4024 | 0.003238 | 0.005465 | 0.005301 | 5.3e-05 | 0.000942 | 0.33892 | 1.2944 | 2.5479 | 99.58 | 20.86 | 37.385 | 81.5053 | 99.5811 | 77.8478 | 0.009267 | 1.0058 | 0.024085 | 23.9454 | 0.010417 | 0.15582 | 0.005168 | 0.37091 | 0.37091 | 0.017266
2019-03-21 11:52 | 1553183526.3842 | 413 | 0.20849 | 0.001294 | 4.8e-05 | 99.6567 | 99.6701 | 120.5408 | 21.0804 | 0.002522 | 0.00235 | 165.3822 | 0.26253 | 8.4014 | 0.003238 | 0.00527 | 0.005167 | 5.5e-05 | 0.000939 | 0.34075 | 1.3308 | 2.5479 | 99.58 | 20.86 | 37.384 | 82.319 | 99.5649 | 78.1818 | 0.009267 | 1.0049 | 0.024074 | 23.9572 | -0.001202 | 0.15618 | 0.005254 | 0.37686 | 0.37686 | 0.017033
2019-03-21 11:52 | 1553183527.3844 | 414 | 0.20852 | 0.001298 | 4.7e-05 | 99.6022 | 99.5737 | 120.3835 | 21.0799 | 0.002542 | 0.002329 | 165.3489 | 0.26112 | 8.403 | 0.003259 | 0.00469 | 0.004689 | 5.6e-05 | 0.00094 | 0.30596 | 1.3657 | 2.5479 | 99.58 | 20.87 | 37.384 | 82.7926 | 100.0681 | 77.6815 | 0.009273 | 1.0063 | 0.024076 | 23.9246 | 0.002557 | 0.15534 | 0.005095 | 0.36545 | 0.36545 | 0.01747
2019-03-21 11:52 | 1553183528.3854 | 415 | 0.2085 | 0.001291 | 4.7e-05 | 99.6341 | 99.6595 | 120.4412 | 21.0801 | 0.00251 | 0.002327 | 165.3429 | 0.26234 | 8.4046 | 0.003339 | 0.00476 | 0.004677 | 5.4e-05 | 0.000936 | 0.30254 | 1.3501 | 2.5481 | 99.58 | 20.87 | 37.385 | 82.031 | 102.0119 | 77.6172 | 0.009273 | 1.0065 | 0.024084 | 23.9281 | -0.002287 | 0.15607 | 0.005223 | 0.37481 | 0.37481 | 0.017114
2019-03-21 11:52 | 1553183529.384 | 416 | 0.20853 | 0.001289 | 4.8e-05 | 99.6534 | 99.4194 | 120.5208 | 21.0803 | 0.002447 | 0.002268 | 165.3106 | 0.2618 | 8.4031 | 0.003247 | 0.00497 | 0.004854 | 5.6e-05 | 0.000937 | 0.3115 | 1.302 | 2.548 | 99.58 | 20.87 | 37.386 | 80.501 | 99.7916 | 76.2149 | 0.009273 | 1.0105 | 0.02414 | 23.8881 | 0.021028 | 0.15575 | 0.005059 | 0.36378 | 0.36378 | 0.017596
2019-03-21 11:52 | 1553183530.3836 | 417 | 0.20852 | 0.001286 | 4.8e-05 | 99.6394 | 99.4095 | 120.9421 | 21.0804 | 0.002578 | 0.002332 | 165.3044 | 0.26262 | 8.404 | 0.003211 | 0.005134 | 0.004992 | 5.3e-05 | 0.000934 | 0.29806 | 1.2812 | 2.5479 | 99.58 | 20.87 | 37.387 | 83.6716 | 98.9062 | 77.7397 | 0.009273 | 1.0061 | 0.024129 | 23.982 | 0.02066 | 0.15623 | 0.005106 | 0.367 | 0.367 | 0.017496
2019-03-21 11:52 | 1553183531.3836 | 418 | 0.20854 | 0.001287 | 4.9e-05 | 99.6058 | 99.5781 | 120.8532 | 21.0808 | 0.002514 | 0.00229 | 165.2902 | 0.26126 | 8.4037 | 0.003181 | 0.00456 | 0.004495 | 5.6e-05 | 0.000936 | 0.33467 | 1.3144 | 2.548 | 99.57 | 20.86 | 37.391 | 82.1203 | 98.1939 | 76.736 | 0.00927 | 1.009 | 0.024155 | 23.9389 | 0.002481 | 0.15542 | 0.004986 | 0.35866 | 0.35866 | 0.01781
2019-03-21 11:52 | 1553183532.385 | 419 | 0.20852 | 0.001275 | 5e-05 | 99.6275 | 99.5958 | 120.7384 | 21.0803 | 0.002491 | 0.002359 | 165.2756 | 0.2627 | 8.4015 | 0.00329 | 0.004845 | 0.004698 | 5.6e-05 | 0.000933 | 0.31758 | 1.3081 | 2.548 | 99.57 | 20.86 | 37.392 | 81.5694 | 100.8251 | 78.3936 | 0.00927 | 1.0043 | 0.024087 | 23.9841 | 0.002843 | 0.15628 | 0.005084 | 0.36464 | 0.36464 | 0.017614
2019-03-21 11:52 | 1553183533.3847 | 420 | 0.20851 | 0.001279 | 5e-05 | 99.7005 | 99.5884 | 120.4307 | 21.0812 | 0.002479 | 0.002376 | 165.241 | 0.26143 | 8.4027 | 0.003303 | 0.004257 | 0.004236 | 5.9e-05 | 0.000936 | 0.30502 | 1.2858 | 2.548 | 99.57 | 20.86 | 37.394 | 81.2903 | 101.1453 | 78.8037 | 0.00927 | 1.0031 | 0.024042 | 23.9675 | 0.010067 | 0.15553 | 0.00515 | 0.36873 | 0.36873 | 0.017336
2019-03-21 11:52 | 1553183534.3837 | 421 | 0.20853 | 0.001277 | 5e-05 | 99.6737 | 99.5492 | 120.5337 | 21.0806 | 0.002472 | 0.002315 | 165.2242 | 0.26122 | 8.4017 | 0.003311 | 0.005345 | 0.00523 | 5.5e-05 | 0.000932 | 0.31061 | 1.2759 | 2.5479 | 99.57 | 20.86 | 37.396 | 81.1034 | 101.3395 | 77.3502 | 0.009271 | 1.0073 | 0.024102 | 23.9282 | 0.01118 | 0.1554 | 0.005045 | 0.3621 | 0.3621 | 0.017639
2019-03-21 11:52 | 1553183535.3841 | 422 | 0.20855 | 0.001274 | 5e-05 | 99.6388 | 99.5012 | 120.6006 | 21.0808 | 0.002502 | 0.002314 | 165.1846 | 0.26152 | 8.4014 | 0.003326 | 0.004517 | 0.004471 | 5.3e-05 | 0.000929 | 0.32117 | 1.3055 | 2.548 | 99.57 | 20.86 | 37.398 | 81.8302 | 101.6937 | 77.3129 | 0.009271 | 1.0074 | 0.02411 | 23.9335 | 0.012367 | 0.15558 | 0.004923 | 0.35334 | 0.35334 | 0.018097
2019-03-21 11:52 | 1553183536.3839 | 423 | 0.20852 | 0.001272 | 5.1e-05 | 99.6567 | 99.6401 | 120.5246 | 21.0818 | 0.002478 | 0.002357 | 165.1566 | 0.26218 | 8.4014 | 0.003305 | 0.004349 | 0.004258 | 5.5e-05 | 0.000928 | 0.31078 | 1.298 | 2.5479 | 99.57 | 20.86 | 37.401 | 81.2589 | 101.1987 | 78.3444 | 0.009272 | 1.0044 | 0.024067 | 23.9612 | 0.001488 | 0.15597 | 0.005099 | 0.36542 | 0.36542 | 0.017543
2019-03-21 11:52 | 1553183537.3838 | 424 | 0.20853 | 0.001273 | 5.1e-05 | 99.6129 | 99.4841 | 120.4823 | 21.0814 | 0.002472 | 0.002325 | 165.1383 | 0.26193 | 8.4036 | 0.003248 | 0.00559 | 0.005444 | 5.7e-05 | 0.000926 | 0.3178 | 1.3068 | 2.548 | 99.57 | 20.87 | 37.402 | 81.1111 | 99.8217 | 77.5847 | 0.009278 | 1.0066 | 0.024089 | 23.9311 | 0.011571 | 0.15582 | 0.005048 | 0.36198 | 0.36198 | 0.017692
2019-03-21 11:52 | 1553183538.3853 | 425 | 0.20853 | 0.001265 | 5.1e-05 | 99.7041 | 99.6296 | 120.145 | 21.0813 | 0.002485 | 0.002272 | 165.1065 | 0.26175 | 8.402 | 0.003352 | 0.005017 | 0.00508 | 5.5e-05 | 0.000929 | 0.31957 | 1.3177 | 2.5481 | 99.57 | 20.87 | 37.405 | 81.4269 | 102.325 | 76.3038 | 0.009279 | 1.0103 | 0.024099 | 23.8539 | 0.006687 | 0.15572 | 0.005039 | 0.36153 | 0.36153 | 0.017702
2019-03-21 11:52 | 1553183539.385 | 426 | 0.20852 | 0.001266 | 5.2e-05 | 99.6817 | 99.4354 | 120.1497 | 21.0827 | 0.002552 | 0.002312 | 165.0894 | 0.26157 | 8.4017 | 0.003375 | 0.005303 | 0.005145 | 5.3e-05 | 0.000927 | 0.30973 | 1.3373 | 2.548 | 99.57 | 20.87 | 37.408 | 83.0379 | 102.8745 | 77.2765 | 0.00928 | 1.0075 | 0.024066 | 23.8875 | 0.022128 | 0.15561 | 0.005127 | 0.36738 | 0.36738 | 0.017408
2019-03-21 11:52 | 1553183540.384 | 427 | 0.20855 | 0.001267 | 5.1e-05 | 99.6597 | 99.6037 | 120.2356 | 21.0811 | 0.002549 | 0.002311 | 165.0631 | 0.26231 | 8.404 | 0.003269 | 0.004813 | 0.00472 | 5.5e-05 | 0.000929 | 0.34092 | 1.2785 | 2.5478 | 99.57 | 20.87 | 37.409 | 82.9623 | 100.322 | 77.2462 | 0.00928 | 1.0076 | 0.024076 | 23.895 | 0.005028 | 0.15604 | 0.004965 | 0.35577 | 0.35577 | 0.018027
2019-03-21 11:52 | 1553183541.3855 | 428 | 0.20856 | 0.001261 | 5.2e-05 | 99.7041 | 99.5209 | 120.5178 | 21.0826 | 0.002493 | 0.002284 | 165.0441 | 0.26136 | 8.4015 | 0.00329 | 0.004701 | 0.004589 | 5.4e-05 | 0.000928 | 0.30773 | 1.2816 | 2.5476 | 99.57 | 20.87 | 37.41 | 81.6182 | 100.828 | 76.5866 | 0.00928 | 1.0095 | 0.024127 | 23.9005 | 0.016447 | 0.15548 | 0.004887 | 0.35086 | 0.35086 | 0.018214
2019-03-21 11:52 | 1553183542.3842 | 429 | 0.20858 | 0.001262 | 5.3e-05 | 99.6317 | 99.3445 | 120.4635 | 21.0822 | 0.002463 | 0.002314 | 165.0087 | 0.26218 | 8.4015 | 0.003241 | 0.004839 | 0.004728 | 5.3e-05 | 0.000924 | 0.33052 | 1.2803 | 2.5478 | 99.57 | 20.87 | 37.41 | 80.8944 | 99.6326 | 77.3104 | 0.00928 | 1.0074 | 0.024096 | 23.9199 | 0.025819 | 0.15597 | 0.004762 | 0.34136 | 0.34136 | 0.018779
2019-03-21 11:52 | 1553183543.3851 | 430 | 0.20857 | 0.001258 | 5.3e-05 | 99.6907 | 99.5672 | 120.5067 | 21.0813 | 0.00251 | 0.002346 | 164.9806 | 0.26211 | 8.4001 | 0.003276 | 0.004849 | 0.004815 | 5.4e-05 | 0.000922 | 0.29607 | 1.2699 | 2.548 | 99.57 | 20.87 | 37.409 | 82.0186 | 100.4932 | 78.0757 | 0.00928 | 1.0052 | 0.024074 | 23.9503 | 0.011082 | 0.15593 | 0.004805 | 0.34413 | 0.34413 | 0.018623
2019-03-21 11:52 | 1553183544.3856 | 431 | 0.20858 | 0.001256 | 5.4e-05 | 99.6209 | 99.4429 | 120.5098 | 21.0823 | 0.002513 | 0.002304 | 164.9785 | 0.26235 | 8.3998 | 0.003228 | 0.005145 | 0.005109 | 5.3e-05 | 0.000924 | 0.31238 | 1.289 | 2.548 | 99.57 | 20.87 | 37.409 | 82.093 | 99.3362 | 77.0744 | 0.00928 | 1.0081 | 0.024109 | 23.9164 | 0.015999 | 0.15607 | 0.004756 | 0.34103 | 0.34103 | 0.01881
2019-03-21 11:52 | 1553183545.3852 | 432 | 0.20857 | 0.001253 | 5.4e-05 | 99.6898 | 99.4906 | 120.4626 | 21.0817 | 0.002512 | 0.002321 | 164.9583 | 0.26244 | 8.4045 | 0.003267 | 0.004571 | 0.004422 | 4.9e-05 | 0.000924 | 0.3569 | 1.31 | 2.548 | 99.57 | 20.87 | 37.409 | 82.0666 | 100.2754 | 77.4853 | 0.00928 | 1.0069 | 0.02409 | 23.9257 | 0.017888 | 0.15613 | 0.004851 | 0.34761 | 0.34761 | 0.01846
2019-03-21 11:52 | 1553183546.3834 | 433 | 0.20857 | 0.001252 | 5.5e-05 | 99.6358 | 99.5032 | 120.2175 | 21.0824 | 0.002505 | 0.002349 | 164.9323 | 0.26323 | 8.403 | 0.00323 | 0.004784 | 0.004735 | 5.6e-05 | 0.00092 | 0.32316 | 1.315 | 2.5479 | 99.57 | 20.87 | 37.409 | 81.9051 | 99.3833 | 78.1609 | 0.00928 | 1.0049 | 0.024043 | 23.9244 | 0.011916 | 0.1566 | 0.004833 | 0.34555 | 0.34555 | 0.018626
2019-03-21 11:52 | 1553183547.3834 | 434 | 0.20858 | 0.001253 | 5.4e-05 | 99.6207 | 99.6674 | 120.3997 | 21.0817 | 0.002483 | 0.00237 | 164.9272 | 0.26234 | 8.4015 | 0.003182 | 0.004259 | 0.00413 | 5.3e-05 | 0.00092 | 0.29446 | 1.282 | 2.548 | 99.57 | 20.87 | 37.407 | 81.378 | 98.2175 | 78.6593 | 0.009279 | 1.0035 | 0.024044 | 23.9595 | -0.004193 | 0.15606 | 0.004779 | 0.34174 | 0.34174 | 0.01877
2019-03-21 11:52 | 1553183548.3845 | 435 | 0.20858 | 0.001248 | 5.4e-05 | 99.6433 | 99.6313 | 120.5844 | 21.083 | 0.002512 | 0.00235 | 164.9028 | 0.26228 | 8.4038 | 0.003241 | 0.004614 | 0.004525 | 5e-05 | 0.000916 | 0.32333 | 1.3097 | 2.5479 | 99.57 | 20.87 | 37.407 | 82.0689 | 99.6349 | 78.1896 | 0.009279 | 1.0049 | 0.024078 | 23.9619 | 0.00108 | 0.15603 | 0.004799 | 0.34369 | 0.34369 | 0.018659
2019-03-21 11:52 | 1553183549.3835 | 436 | 0.20858 | 0.00124 | 5.4e-05 | 99.6481 | 99.3792 | 120.6194 | 21.0828 | 0.002501 | 0.002323 | 164.8836 | 0.26114 | 8.4028 | 0.003196 | 0.004525 | 0.004421 | 5.1e-05 | 0.000912 | 0.31166 | 1.3119 | 2.5479 | 99.57 | 20.86 | 37.406 | 81.8158 | 98.5564 | 77.544 | 0.009273 | 1.0067 | 0.024104 | 23.9433 | 0.024163 | 0.15535 | 0.00481 | 0.34483 | 0.34483 | 0.018516
2019-03-21 11:52 | 1553183550.3839 | 437 | 0.2086 | 0.001239 | 5.5e-05 | 99.6311 | 99.6248 | 120.2488 | 21.0832 | 0.002582 | 0.002325 | 164.8616 | 0.26231 | 8.4029 | 0.003164 | 0.004849 | 0.004705 | 5.1e-05 | 0.000916 | 0.30845 | 1.32 | 2.548 | 99.57 | 20.86 | 37.405 | 83.7491 | 97.7901 | 77.5872 | 0.009273 | 1.0066 | 0.024065 | 23.908 | 0.000564 | 0.15605 | 0.004658 | 0.33322 | 0.33322 | 0.019247
2019-03-21 11:52 | 1553183551.3846 | 438 | 0.20859 | 0.001245 | 5.6e-05 | 99.6025 | 99.3554 | 119.8819 | 21.0829 | 0.002511 | 0.002331 | 164.8337 | 0.26214 | 8.4015 | 0.003253 | 0.004781 | 0.004774 | 5e-05 | 0.000913 | 0.33251 | 1.284 | 2.5477 | 99.57 | 20.86 | 37.405 | 82.0452 | 99.9455 | 77.7374 | 0.009273 | 1.0062 | 0.024023 | 23.8766 | 0.022217 | 0.15595 | 0.004707 | 0.33617 | 0.33617 | 0.019066
2019-03-21 11:53 | 1553183552.3837 | 439 | 0.20859 | 0.001238 | 5.7e-05 | 99.5682 | 99.605 | 120.0833 | 21.0828 | 0.002515 | 0.002314 | 164.8176 | 0.26217 | 8.4011 | 0.003134 | 0.004356 | 0.004308 | 5.3e-05 | 0.000912 | 0.31669 | 1.2984 | 2.5479 | 99.57 | 20.86 | 37.405 | 82.1521 | 97.0562 | 77.3211 | 0.009273 | 1.0073 | 0.024058 | 23.8824 | -0.003308 | 0.15597 | 0.004722 | 0.33768 | 0.33768 | 0.018983
2019-03-21 11:53 | 1553183553.3842 | 440 | 0.20859 | 0.001234 | 5.7e-05 | 99.5566 | 99.3636 | 119.9537 | 21.0823 | 0.002471 | 0.002307 | 164.8162 | 0.26216 | 8.4042 | 0.003121 | 0.004115 | 0.004013 | 5.1e-05 | 0.000911 | 0.33732 | 1.298 | 2.5479 | 99.57 | 20.86 | 37.405 | 81.0947 | 96.74 | 77.1413 | 0.009273 | 1.0079 | 0.024051 | 23.8634 | 0.017359 | 0.15596 | 0.00475 | 0.3396 | 0.3396 | 0.018875
2019-03-21 11:53 | 1553183554.3839 | 441 | 0.20862 | 0.001235 | 5.8e-05 | 99.5664 | 99.4211 | 119.8561 | 21.084 | 0.002504 | 0.002345 | 164.7739 | 0.26334 | 8.4041 | 0.003135 | 0.004344 | 0.0042 | 5.3e-05 | 0.000907 | 0.32139 | 1.3178 | 2.5479 | 99.57 | 20.87 | 37.403 | 81.8778 | 97.0786 | 78.0541 | 0.009278 | 1.0052 | 0.02401 | 23.8848 | 0.013063 | 0.15666 | 0.004562 | 0.32534 | 0.32534 | 0.019791
2019-03-21 11:53 | 1553183555.3839 | 442 | 0.20868 | 0.001233 | 5.8e-05 | 99.5611 | 99.462 | 119.8357 | 21.0838 | 0.002408 | 0.002299 | 164.7585 | 0.26267 | 8.4027 | 0.003045 | 0.004669 | 0.004566 | 5.3e-05 | 0.000905 | 0.31166 | 1.3318 | 2.5479 | 99.57 | 20.87 | 37.403 | 79.5843 | 94.9105 | 76.9564 | 0.009278 | 1.0084 | 0.024046 | 23.8454 | 0.008905 | 0.15627 | 0.004217 | 0.30091 | 0.30091 | 0.021343
2019-03-21 11:53 | 1553183556.3849 | 443 | 0.20859 | 0.001231 | 5.9e-05 | 99.5462 | 99.5104 | 120.2305 | 21.0836 | 0.002476 | 0.002245 | 164.7412 | 0.26277 | 8.4029 | 0.003013 | 0.004656 | 0.004571 | 5.3e-05 | 0.000906 | 0.32997 | 1.3182 | 2.5478 | 99.57 | 20.87 | 37.403 | 81.2099 | 94.1462 | 75.66 | 0.009278 | 1.0121 | 0.02413 | 23.8404 | 0.003218 | 0.15632 | 0.00472 | 0.33843 | 0.33843 | 0.018984
2019-03-21 11:53 | 1553183557.3851 | 444 | 0.20866 | 0.001229 | 5.9e-05 | 99.5301 | 99.4266 | 120.1927 | 21.0836 | 0.002469 | 0.00234 | 164.7052 | 0.26272 | 8.403 | 0.003096 | 0.004216 | 0.004176 | 5.3e-05 | 0.000906 | 0.31061 | 1.3464 | 2.548 | 99.57 | 20.87 | 37.403 | 81.0383 | 96.1484 | 77.9495 | 0.009278 | 1.0055 | 0.024047 | 23.9147 | 0.009307 | 0.15629 | 0.004341 | 0.3099 | 0.3099 | 0.020728
2019-03-21 11:53 | 1553183558.384 | 445 | 0.2086 | 0.001222 | 5.9e-05 | 99.525 | 99.2965 | 119.8969 | 21.0847 | 0.002491 | 0.002313 | 164.6808 | 0.26131 | 8.3998 | 0.003075 | 0.004669 | 0.004577 | 5.4e-05 | 0.000905 | 0.31924 | 1.3176 | 2.548 | 99.57 | 20.87 | 37.404 | 81.5713 | 95.638 | 77.3074 | 0.009279 | 1.0074 | 0.02404 | 23.8634 | 0.020563 | 0.15545 | 0.004682 | 0.33437 | 0.33437 | 0.019108
2019-03-21 11:53 | 1553183559.3851 | 446 | 0.20861 | 0.001228 | 6e-05 | 99.4767 | 99.1834 | 119.826 | 21.0831 | 0.002524 | 0.002325 | 164.6418 | 0.26178 | 8.4023 | 0.003072 | 0.004456 | 0.004279 | 5.2e-05 | 0.000903 | 0.31006 | 1.279 | 2.5479 | 99.58 | 20.87 | 37.404 | 82.3762 | 95.5642 | 77.5746 | 0.009278 | 1.0066 | 0.024023 | 23.8655 | 0.026416 | 0.15573 | 0.004642 | 0.33125 | 0.33125 | 0.019323
2019-03-21 11:53 | 1553183560.3844 | 447 | 0.20861 | 0.00122 | 6.1e-05 | 99.4365 | 99.2505 | 120.0366 | 21.0847 | 0.002454 | 0.002316 | 164.6419 | 0.26245 | 8.4048 | 0.003088 | 0.004731 | 0.004752 | 5.3e-05 | 0.000906 | 0.32029 | 1.3446 | 2.548 | 99.58 | 20.87 | 37.404 | 80.6808 | 95.9411 | 77.365 | 0.009278 | 1.0072 | 0.024052 | 23.8793 | 0.01675 | 0.15613 | 0.004601 | 0.3286 | 0.3286 | 0.019528
2019-03-21 11:53 | 1553183561.3842 | 448 | 0.2086 | 0.00122 | 6.1e-05 | 99.4324 | 99.3816 | 120.1612 | 21.0837 | 0.002516 | 0.002379 | 164.6056 | 0.26207 | 8.4028 | 0.003083 | 0.004271 | 0.00414 | 5.1e-05 | 0.000903 | 0.33754 | 1.3642 | 2.548 | 99.58 | 20.87 | 37.404 | 82.1749 | 95.8366 | 78.8929 | 0.009278 | 1.0028 | 0.024012 | 23.9437 | 0.004565 | 0.15591 | 0.004705 | 0.33558 | 0.33558 | 0.019095
2019-03-21 11:53 | 1553183562.3845 | 449 | 0.20863 | 0.001218 | 6.2e-05 | 99.4449 | 99.1466 | 120.1284 | 21.084 | 0.00245 | 0.002315 | 164.5931 | 0.26181 | 8.4022 | 0.0031 | 0.004676 | 0.004563 | 5.1e-05 | 0.000902 | 0.3126 | 1.3197 | 2.548 | 99.58 | 20.87 | 37.405 | 80.5763 | 96.2482 | 77.3372 | 0.009278 | 1.0073 | 0.024062 | 23.8875 | 0.026867 | 0.15575 | 0.004485 | 0.32032 | 0.32032 | 0.019984
2019-03-21 11:53 | 1553183563.3846 | 450 | 0.20861 | 0.00122 | 6.1e-05 | 99.4127 | 99.3115 | 120.3187 | 21.0836 | 0.002527 | 0.002298 | 164.5557 | 0.26302 | 8.4044 | 0.003099 | 0.004353 | 0.00434 | 5.2e-05 | 0.000905 | 0.31332 | 1.3158 | 2.548 | 99.58 | 20.87 | 37.405 | 82.4484 | 96.2109 | 76.9281 | 0.009278 | 1.0085 | 0.024095 | 23.8924 | 0.009113 | 0.15647 | 0.004643 | 0.33223 | 0.33223 | 0.019357
2019-03-21 11:53 | 1553183564.3847 | 451 | 0.20865 | 0.001215 | 6.2e-05 | 99.399 | 99.226 | 120.3509 | 21.0839 | 0.002492 | 0.002322 | 164.534 | 0.26334 | 8.402 | 0.003118 | 0.004477 | 0.00437 | 5.2e-05 | 0.000904 | 0.32709 | 1.3382 | 2.548 | 99.58 | 20.87 | 37.405 | 81.5868 | 96.6631 | 77.5108 | 0.009278 | 1.0068 | 0.024078 | 23.9155 | 0.015583 | 0.15666 | 0.004392 | 0.31381 | 0.31381 | 0.020519
2019-03-21 11:53 | 1553183565.3844 | 452 | 0.20861 | 0.001211 | 6.1e-05 | 99.3737 | 99.3067 | 120.411 | 21.0839 | 0.002494 | 0.002269 | 164.5254 | 0.26312 | 8.4015 | 0.003211 | 0.004885 | 0.004831 | 5.4e-05 | 0.000904 | 0.30326 | 1.3321 | 2.548 | 99.58 | 20.87 | 37.406 | 81.6352 | 98.9205 | 76.241 | 0.009278 | 1.0105 | 0.024128 | 23.8781 | 0.00603 | 0.15653 | 0.004633 | 0.33194 | 0.33194 | 0.019381
2019-03-21 11:53 | 1553183566.3848 | 453 | 0.20862 | 0.001209 | 6.2e-05 | 99.3734 | 99.2008 | 120.2452 | 21.0842 | 0.002471 | 0.002288 | 164.5241 | 0.26338 | 8.4049 | 0.003199 | 0.00454 | 0.004463 | 5.2e-05 | 0.000903 | 0.3418 | 1.2935 | 2.5479 | 99.58 | 20.87 | 37.406 | 81.1007 | 98.63 | 76.7028 | 0.009278 | 1.0091 | 0.024095 | 23.8774 | 0.015549 | 0.15668 | 0.004572 | 0.32709 | 0.32709 | 0.019688
2019-03-21 11:53 | 1553183567.3845 | 454 | 0.20864 | 0.001213 | 6.2e-05 | 99.385 | 99.1578 | 120.2009 | 21.0847 | 0.002482 | 0.002344 | 164.5146 | 0.2633 | 8.4019 | 0.003235 | 0.004558 | 0.004406 | 5.1e-05 | 0.000906 | 0.33157 | 1.3015 | 2.548 | 99.58 | 20.87 | 37.407 | 81.3468 | 99.5009 | 78.0363 | 0.009278 | 1.0053 | 0.024045 | 23.9185 | 0.020465 | 0.15664 | 0.004455 | 0.31793 | 0.31793 | 0.020249
2019-03-21 11:53 | 1553183568.3835 | 455 | 0.20866 | 0.0012 | 6.2e-05 | 99.4106 | 99.1303 | 120.2683 | 21.0843 | 0.002461 | 0.002234 | 164.4947 | 0.26231 | 8.4015 | 0.003151 | 0.004558 | 0.004533 | 5.1e-05 | 0.000902 | 0.30663 | 1.3194 | 2.5479 | 99.58 | 20.87 | 37.407 | 80.8448 | 97.4769 | 75.4104 | 0.009278 | 1.0129 | 0.024142 | 23.8356 | 0.025259 | 0.15605 | 0.004379 | 0.31366 | 0.31366 | 0.020448
2019-03-21 11:53 | 1553183569.3843 | 456 | 0.20864 | 0.001202 | 6.3e-05 | 99.3436 | 99.4708 | 120.4678 | 21.0843 | 0.002526 | 0.002328 | 164.4958 | 0.26259 | 8.4025 | 0.003173 | 0.004579 | 0.004489 | 5.1e-05 | 0.000905 | 0.3454 | 1.276 | 2.548 | 99.58 | 20.87 | 37.407 | 82.4211 | 97.9989 | 77.6678 | 0.009278 | 1.0064 | 0.024084 | 23.9325 | -0.011453 | 0.15621 | 0.004463 | 0.31899 | 0.31899 | 0.020127
2019-03-21 11:53 | 1553183570.3846 | 457 | 0.20862 | 0.001202 | 6.2e-05 | 99.3814 | 99.1391 | 120.6295 | 21.084 | 0.002497 | 0.002325 | 164.4539 | 0.26222 | 8.4043 | 0.003137 | 0.004589 | 0.004545 | 5.3e-05 | 0.000908 | 0.33069 | 1.2861 | 2.548 | 99.58 | 20.87 | 37.407 | 81.7133 | 97.1285 | 77.574 | 0.009278 | 1.0066 | 0.024104 | 23.9453 | 0.021833 | 0.15599 | 0.004578 | 0.32758 | 0.32758 | 0.019572
2019-03-21 11:53 | 1553183571.3837 | 458 | 0.20866 | 0.001196 | 6.2e-05 | 99.3075 | 99.3149 | 120.7677 | 21.0847 | 0.002511 | 0.002302 | 164.4559 | 0.26204 | 8.4013 | 0.00317 | 0.004537 | 0.004427 | 5.8e-05 | 0.000908 | 0.31852 | 1.3244 | 2.5479 | 99.58 | 20.87 | 37.408 | 82.0549 | 97.9312 | 77.0423 | 0.009279 | 1.0081 | 0.024136 | 23.9409 | -0.000662 | 0.15589 | 0.004348 | 0.31134 | 0.31134 | 0.020579
2019-03-21 11:53 | 1553183572.3848 | 459 | 0.20864 | 0.001195 | 6.3e-05 | 99.3772 | 99.1347 | 120.5985 | 21.084 | 0.00247 | 0.002256 | 164.436 | 0.26214 | 8.4023 | 0.003138 | 0.004132 | 0.004056 | 5.6e-05 | 0.000907 | 0.31078 | 1.296 | 2.5479 | 99.58 | 20.87 | 37.409 | 81.0629 | 97.1529 | 75.9347 | 0.009279 | 1.0113 | 0.024157 | 23.8862 | 0.021857 | 0.15595 | 0.004482 | 0.32133 | 0.32133 | 0.019947
2019-03-21 11:53 | 1553183573.3847 | 460 | 0.20862 | 0.001193 | 6.3e-05 | 99.3325 | 99.0734 | 120.6311 | 21.0848 | 0.002435 | 0.002328 | 164.4177 | 0.26249 | 8.4015 | 0.003097 | 0.004837 | 0.004755 | 5.7e-05 | 0.000907 | 0.32764 | 1.3157 | 2.548 | 99.58 | 20.87 | 37.408 | 80.2317 | 96.166 | 77.6513 | 0.009279 | 1.0064 | 0.024101 | 23.9481 | 0.023367 | 0.15616 | 0.004608 | 0.3297 | 0.3297 | 0.019466
2019-03-21 11:53 | 1553183574.3839 | 461 | 0.20863 | 0.001193 | 6.2e-05 | 99.3403 | 99.1398 | 120.602 | 21.085 | 0.002501 | 0.002299 | 164.3908 | 0.26238 | 8.4035 | 0.003063 | 0.004472 | 0.004431 | 5.6e-05 | 0.000906 | 0.32084 | 1.327 | 2.5479 | 99.58 | 20.87 | 37.408 | 81.8117 | 95.3588 | 76.9669 | 0.009279 | 1.0084 | 0.024122 | 23.9219 | 0.01807 | 0.15609 | 0.004575 | 0.32757 | 0.32757 | 0.019584
2019-03-21 11:53 | 1553183575.3849 | 462 | 0.20863 | 0.001187 | 6.3e-05 | 99.3668 | 99.2536 | 120.6293 | 21.0843 | 0.002494 | 0.0023 | 164.3597 | 0.26251 | 8.4008 | 0.003118 | 0.004251 | 0.004119 | 5.2e-05 | 0.000911 | 0.33964 | 1.359 | 2.548 | 99.58 | 20.87 | 37.408 | 81.638 | 96.6855 | 76.9951 | 0.009279 | 1.0083 | 0.024124 | 23.9255 | 0.0102 | 0.15617 | 0.004576 | 0.32764 | 0.32764 | 0.019591
2019-03-21 11:53 | 1553183576.3839 | 463 | 0.20866 | 0.001187 | 6.3e-05 | 99.3287 | 99.102 | 120.576 | 21.0847 | 0.002442 | 0.002333 | 164.3389 | 0.26211 | 8.4022 | 0.003123 | 0.004361 | 0.004375 | 5.2e-05 | 0.000908 | 0.32244 | 1.3373 | 2.548 | 99.58 | 20.87 | 37.406 | 80.3931 | 96.7909 | 77.7741 | 0.009278 | 1.006 | 0.024092 | 23.9468 | 0.020436 | 0.15593 | 0.004399 | 0.31438 | 0.31438 | 0.020385
2019-03-21 11:53 | 1553183577.383 | 464 | 0.20863 | 0.001188 | 6.4e-05 | 99.3418 | 99.0918 | 120.6504 | 21.0849 | 0.002504 | 0.002328 | 164.3086 | 0.2627 | 8.4035 | 0.003079 | 0.004109 | 0.003984 | 5e-05 | 0.000909 | 0.29679 | 1.3264 | 2.5481 | 99.58 | 20.87 | 37.406 | 81.8772 | 95.733 | 77.6528 | 0.009278 | 1.0064 | 0.024103 | 23.9501 | 0.022538 | 0.15628 | 0.004586 | 0.32799 | 0.32799 | 0.019583
2019-03-21 11:53 | 1553183578.3831 | 465 | 0.20862 | 0.001185 | 6.4e-05 | 99.3421 | 99.1364 | 120.9133 | 21.0852 | 0.002523 | 0.002307 | 164.2888 | 0.26136 | 8.4036 | 0.003104 | 0.004493 | 0.004355 | 5.3e-05 | 0.000911 | 0.32659 | 1.2998 | 2.5479 | 99.58 | 20.87 | 37.407 | 82.3337 | 96.3425 | 77.1629 | 0.009278 | 1.0078 | 0.024146 | 23.9594 | 0.018538 | 0.15548 | 0.004594 | 0.32918 | 0.32918 | 0.019412
2019-03-21 11:53 | 1553183579.3841 | 466 | 0.20866 | 0.001185 | 6.5e-05 | 99.2601 | 99.2008 | 121.0257 | 21.0848 | 0.002472 | 0.002304 | 164.2758 | 0.26234 | 8.4047 | 0.003123 | 0.004641 | 0.004528 | 5.1e-05 | 0.000908 | 0.32117 | 1.333 | 2.548 | 99.57 | 20.87 | 37.408 | 81.1058 | 96.8019 | 77.0862 | 0.00928 | 1.008 | 0.02416 | 23.9679 | 0.005352 | 0.15606 | 0.004364 | 0.31263 | 0.31263 | 0.020517
2019-03-21 11:53 | 1553183580.3841 | 467 | 0.20865 | 0.001182 | 6.5e-05 | 99.3668 | 99.1241 | 120.9396 | 21.0856 | 0.002496 | 0.002318 | 164.2543 | 0.26206 | 8.4029 | 0.003087 | 0.004753 | 0.004655 | 5.2e-05 | 0.000908 | 0.30359 | 1.3301 | 2.5479 | 99.57 | 20.87 | 37.409 | 81.6944 | 95.9245 | 77.4071 | 0.00928 | 1.0071 | 0.024141 | 23.9704 | 0.021871 | 0.1559 | 0.004443 | 0.31809 | 0.31809 | 0.020144
2019-03-21 11:53 | 1553183581.3845 | 468 | 0.20866 | 0.001185 | 6.5e-05 | 99.3337 | 99.1071 | 121.0729 | 21.0851 | 0.002467 | 0.002286 | 164.2342 | 0.26225 | 8.4025 | 0.00316 | 0.004534 | 0.004477 | 5.2e-05 | 0.000912 | 0.3314 | 1.3675 | 2.5479 | 99.57 | 20.87 | 37.409 | 81.0039 | 97.6917 | 76.6398 | 0.00928 | 1.0093 | 0.02418 | 23.9573 | 0.02043 | 0.15601 | 0.004349 | 0.31183 | 0.31183 | 0.020563
2019-03-21 11:53 | 1553183582.3847 | 469 | 0.20861 | 0.001177 | 6.5e-05 | 99.3251 | 98.8503 | 120.9443 | 21.0854 | 0.002445 | 0.002299 | 164.2164 | 0.26281 | 8.404 | 0.003128 | 0.005023 | 0.004992 | 5.2e-05 | 0.000911 | 0.28694 | 1.3061 | 2.5479 | 99.57 | 20.87 | 37.408 | 80.4724 | 96.9131 | 76.9534 | 0.00928 | 1.0084 | 0.024157 | 23.9553 | 0.042859 | 0.15634 | 0.004664 | 0.33432 | 0.33432 | 0.01922
2019-03-21 11:53 | 1553183583.3833 | 470 | 0.20864 | 0.001178 | 6.6e-05 | 99.3451 | 99.0598 | 120.5356 | 21.0846 | 0.002461 | 0.00229 | 164.2282 | 0.26323 | 8.4023 | 0.003125 | 0.00463 | 0.004496 | 5.1e-05 | 0.000911 | 0.3126 | 1.3554 | 2.548 | 99.57 | 20.86 | 37.409 | 80.853 | 96.8338 | 76.7529 | 0.009274 | 1.009 | 0.024123 | 23.908 | 0.025723 | 0.1566 | 0.004524 | 0.32375 | 0.32375 | 0.01988
2019-03-21 11:53 | 1553183584.3834 | 471 | 0.20867 | 0.001176 | 6.6e-05 | 99.3319 | 99.1272 | 120.5889 | 21.0854 | 0.002434 | 0.002284 | 164.2132 | 0.26276 | 8.402 | 0.003068 | 0.004776 | 0.004674 | 5.4e-05 | 0.00091 | 0.32189 | 1.3408 | 2.5479 | 99.57 | 20.87 | 37.408 | 80.1965 | 95.4702 | 76.6065 | 0.00928 | 1.0094 | 0.024133 | 23.9082 | 0.018456 | 0.15632 | 0.004338 | 0.31039 | 0.31039 | 0.020699
2019-03-21 11:53 | 1553183585.3834 | 472 | 0.20865 | 0.001179 | 6.6e-05 | 99.3245 | 99.1112 | 120.7736 | 21.0854 | 0.002447 | 0.002278 | 164.199 | 0.26327 | 8.402 | 0.00322 | 0.004778 | 0.004738 | 5.6e-05 | 0.00091 | 0.32549 | 1.3254 | 2.5479 | 99.57 | 20.87 | 37.407 | 80.5085 | 99.1448 | 76.4467 | 0.009279 | 1.0099 | 0.024157 | 23.9211 | 0.01923 | 0.15662 | 0.004433 | 0.31757 | 0.31757 | 0.02027
2019-03-21 11:53 | 1553183586.3835 | 473 | 0.20866 | 0.001176 | 6.6e-05 | 99.3745 | 99.0908 | 121.1604 | 21.0851 | 0.002455 | 0.002312 | 164.1779 | 0.26265 | 8.4046 | 0.003258 | 0.005664 | 0.005604 | 5.3e-05 | 0.000913 | 0.31094 | 1.3159 | 2.548 | 99.57 | 20.86 | 37.407 | 80.7146 | 100.0543 | 77.2649 | 0.009274 | 1.0075 | 0.024167 | 23.9874 | 0.025581 | 0.15625 | 0.004393 | 0.31476 | 0.31476 | 0.020402
2019-03-21 11:53 | 1553183587.3839 | 474 | 0.20869 | 0.001178 | 6.7e-05 | 99.4249 | 99.1997 | 121.1968 | 21.0864 | 0.002424 | 0.002295 | 164.1602 | 0.26228 | 8.4001 | 0.003228 | 0.004879 | 0.004803 | 5.4e-05 | 0.000914 | 0.32637 | 1.2946 | 2.5478 | 99.57 | 20.86 | 37.407 | 79.9752 | 99.3275 | 76.8693 | 0.009274 | 1.0086 | 0.024185 | 23.9774 | 0.020279 | 0.15603 | 0.00421 | 0.30172 | 0.30172 | 0.021255
2019-03-21 11:53 | 1553183588.3836 | 475 | 0.20871 | 0.001176 | 6.7e-05 | 99.3683 | 99.3292 | 121.1886 | 21.0857 | 0.002463 | 0.00227 | 164.1629 | 0.26295 | 8.403 | 0.003266 | 0.004776 | 0.004631 | 5.5e-05 | 0.000914 | 0.32084 | 1.3026 | 2.5478 | 99.57 | 20.86 | 37.407 | 80.9047 | 100.2604 | 76.2716 | 0.009274 | 1.0104 | 0.024205 | 23.9561 | 0.003523 | 0.15643 | 0.004093 | 0.2934 | 0.2934 | 0.021913
2019-03-21 11:53 | 1553183589.3837 | 476 | 0.20866 | 0.001168 | 6.8e-05 | 99.424 | 99.2818 | 120.829 | 21.0857 | 0.002436 | 0.002261 | 164.1476 | 0.26231 | 8.4009 | 0.003193 | 0.004662 | 0.004679 | 5.4e-05 | 0.000917 | 0.335 | 1.2711 | 2.548 | 99.57 | 20.87 | 37.406 | 80.2622 | 98.4792 | 76.0441 | 0.009279 | 1.011 | 0.024177 | 23.9128 | 0.012801 | 0.15604 | 0.004429 | 0.31741 | 0.31741 | 0.020206
2019-03-21 11:53 | 1553183590.3837 | 477 | 0.20871 | 0.001162 | 6.8e-05 | 99.4186 | 99.1572 | 120.3481 | 21.0858 | 0.00249 | 0.00227 | 164.1155 | 0.26308 | 8.4028 | 0.003205 | 0.004905 | 0.004778 | 5.4e-05 | 0.000917 | 0.31957 | 1.3548 | 2.548 | 99.57 | 20.87 | 37.404 | 81.5543 | 98.7672 | 76.2585 | 0.009279 | 1.0104 | 0.024121 | 23.8725 | 0.023556 | 0.15651 | 0.004083 | 0.2916 | 0.2916 | 0.022059
2019-03-21 11:53 | 1553183591.3838 | 478 | 0.20863 | 0.001166 | 7e-05 | 99.4586 | 99.2004 | 120.0288 | 21.0854 | 0.002408 | 0.002326 | 164.0976 | 0.26295 | 8.4021 | 0.003161 | 0.005129 | 0.005034 | 5.4e-05 | 0.000917 | 0.33301 | 1.3217 | 2.548 | 99.57 | 20.87 | 37.403 | 79.5765 | 97.7109 | 77.5988 | 0.009278 | 1.0065 | 0.024043 | 23.8865 | 0.023246 | 0.15643 | 0.004606 | 0.32832 | 0.32832 | 0.019583
2019-03-21 11:53 | 1553183592.3839 | 479 | 0.20863 | 0.001164 | 7e-05 | 99.4884 | 99.4623 | 120.2314 | 21.0873 | 0.002381 | 0.002318 | 164.0775 | 0.2634 | 8.4009 | 0.003268 | 0.004901 | 0.004841 | 5.4e-05 | 0.000917 | 0.3361 | 1.306 | 2.5478 | 99.57 | 20.87 | 37.402 | 78.926 | 100.3046 | 77.4307 | 0.009278 | 1.007 | 0.024069 | 23.9009 | 0.00234 | 0.1567 | 0.004561 | 0.32538 | 0.32538 | 0.019793
2019-03-21 11:53 | 1553183593.3829 | 480 | 0.20868 | 0.001167 | 7e-05 | 99.5328 | 99.2944 | 120.5713 | 21.0852 | 0.00247 | 0.002279 | 164.0415 | 0.2626 | 8.405 | 0.003301 | 0.004988 | 0.004958 | 5.5e-05 | 0.000919 | 0.32764 | 1.3277 | 2.5478 | 99.57 | 20.87 | 37.401 | 81.0742 | 101.095 | 76.491 | 0.009278 | 1.0097 | 0.024135 | 23.9026 | 0.021444 | 0.15622 | 0.004277 | 0.30571 | 0.30571 | 0.021003
2019-03-21 11:53 | 1553183594.3839 | 481 | 0.20859 | 0.001165 | 7.1e-05 | 99.5491 | 99.3751 | 120.7069 | 21.0855 | 0.002428 | 0.002297 | 164.0087 | 0.26225 | 8.403 | 0.003287 | 0.004669 | 0.004606 | 5.9e-05 | 0.000917 | 0.32228 | 1.3344 | 2.548 | 99.57 | 20.87 | 37.401 | 80.0532 | 100.7658 | 76.9098 | 0.009278 | 1.0085 | 0.024134 | 23.9303 | 0.015648 | 0.15601 | 0.004832 | 0.34589 | 0.34589 | 0.018538
2019-03-21 11:53 | 1553183595.383 | 482 | 0.20863 | 0.001167 | 7.1e-05 | 99.556 | 99.2161 | 120.5152 | 21.0862 | 0.002483 | 0.00226 | 163.9993 | 0.26353 | 8.4005 | 0.003264 | 0.004334 | 0.0042 | 5.4e-05 | 0.000917 | 0.30453 | 1.3369 | 2.548 | 99.57 | 20.87 | 37.397 | 81.3824 | 100.2007 | 76.0328 | 0.009277 | 1.0111 | 0.024146 | 23.8813 | 0.030591 | 0.15678 | 0.004597 | 0.32901 | 0.32901 | 0.019585
2019-03-21 11:53 | 1553183596.384 | 483 | 0.20864 | 0.001161 | 7.2e-05 | 99.6272 | 99.5213 | 120.3823 | 21.0852 | 0.002479 | 0.002293 | 163.9704 | 0.26221 | 8.4044 | 0.003197 | 0.00479 | 0.004652 | 5.8e-05 | 0.000917 | 0.30525 | 1.3595 | 2.5479 | 99.57 | 20.87 | 37.396 | 81.2778 | 98.5791 | 76.8206 | 0.009277 | 1.0088 | 0.024105 | 23.8951 | 0.009517 | 0.15599 | 0.004538 | 0.3241 | 0.3241 | 0.019781
2019-03-21 11:53 | 1553183597.3831 | 484 | 0.20868 | 0.001158 | 7.1e-05 | 99.5938 | 99.5175 | 120.1678 | 21.0864 | 0.002462 | 0.002259 | 163.9514 | 0.26253 | 8.4043 | 0.003278 | 0.00473 | 0.004634 | 6e-05 | 0.000921 | 0.31996 | 1.3325 | 2.5479 | 99.57 | 20.87 | 37.395 | 80.8797 | 100.5509 | 75.9924 | 0.009276 | 1.0112 | 0.024112 | 23.8455 | 0.006858 | 0.15618 | 0.004314 | 0.30802 | 0.30802 | 0.020839
2019-03-21 11:53 | 1553183598.384 | 485 | 0.20865 | 0.001163 | 7.1e-05 | 99.6823 | 99.4589 | 119.7828 | 21.0857 | 0.002431 | 0.002265 | 163.9201 | 0.26352 | 8.4025 | 0.003248 | 0.005003 | 0.004926 | 5.8e-05 | 0.000919 | 0.32261 | 1.3323 | 2.548 | 99.57 | 20.87 | 37.393 | 80.1398 | 99.8056 | 76.1532 | 0.009276 | 1.0107 | 0.024068 | 23.8128 | 0.020068 | 0.15677 | 0.004437 | 0.31636 | 0.31636 | 0.020367
2019-03-21 11:53 | 1553183599.3832 | 486 | 0.20866 | 0.00116 | 7.1e-05 | 99.6814 | 99.492 | 120.1772 | 21.0863 | 0.002444 | 0.002327 | 163.8927 | 0.26266 | 8.4056 | 0.003216 | 0.004642 | 0.004598 | 5.9e-05 | 0.000923 | 0.31902 | 1.2886 | 2.5479 | 99.57 | 20.87 | 37.392 | 80.4475 | 99.0318 | 77.6308 | 0.009276 | 1.0065 | 0.024057 | 23.9023 | 0.017017 | 0.15626 | 0.004421 | 0.31507 | 0.31507 | 0.020383
2019-03-21 11:53 | 1553183600.3842 | 487 | 0.20863 | 0.001159 | 7.1e-05 | 99.716 | 99.51 | 120.6194 | 21.0866 | 0.002457 | 0.002293 | 163.8763 | 0.26255 | 8.4032 | 0.003199 | 0.004318 | 0.004175 | 5.6e-05 | 0.000924 | 0.29264 | 1.321 | 2.5479 | 99.57 | 20.87 | 37.389 | 80.7497 | 98.63 | 76.8236 | 0.009275 | 1.0088 | 0.024129 | 23.9187 | 0.018495 | 0.15619 | 0.004562 | 0.32619 | 0.32619 | 0.01968
2019-03-21 11:53 | 1553183601.3837 | 488 | 0.20867 | 0.001153 | 7e-05 | 99.7214 | 99.5478 | 120.5959 | 21.0862 | 0.00244 | 0.002245 | 163.8305 | 0.26322 | 8.4005 | 0.00323 | 0.004974 | 0.004874 | 5.6e-05 | 0.000922 | 0.33251 | 1.3168 | 2.5478 | 99.57 | 20.87 | 37.389 | 80.3492 | 99.3832 | 75.672 | 0.009275 | 1.0121 | 0.024166 | 23.877 | 0.015579 | 0.15659 | 0.004335 | 0.31032 | 0.31032 | 0.020739
2019-03-21 11:53 | 1553183602.3844 | 489 | 0.20869 | 0.001156 | 7e-05 | 99.8259 | 99.6817 | 120.6368 | 21.087 | 0.002468 | 0.002326 | 163.8158 | 0.26312 | 8.4004 | 0.003199 | 0.004485 | 0.004314 | 5.8e-05 | 0.000922 | 0.3361 | 1.3188 | 2.548 | 99.57 | 20.87 | 37.388 | 81.0137 | 98.6366 | 77.6082 | 0.009275 | 1.0065 | 0.024103 | 23.9472 | 0.012935 | 0.15653 | 0.004225 | 0.30153 | 0.30153 | 0.021336
2019-03-21 11:53 | 1553183603.3834 | 490 | 0.20863 | 0.001155 | 7e-05 | 99.8718 | 99.6486 | 120.4755 | 21.0869 | 0.002471 | 0.002299 | 163.8017 | 0.26153 | 8.4033 | 0.003242 | 0.004434 | 0.004425 | 6e-05 | 0.000924 | 0.32372 | 1.312 | 2.548 | 99.57 | 20.87 | 37.387 | 81.0978 | 99.6835 | 76.9573 | 0.009274 | 1.0084 | 0.02411 | 23.909 | 0.020011 | 0.15559 | 0.004611 | 0.3296 | 0.3296 | 0.019401
2019-03-21 11:53 | 1553183604.3848 | 491 | 0.20866 | 0.001153 | 7e-05 | 99.8673 | 99.7328 | 120.0579 | 21.0876 | 0.002422 | 0.002302 | 163.7827 | 0.26359 | 8.401 | 0.003271 | 0.004287 | 0.004167 | 5.7e-05 | 0.00092 | 0.32045 | 1.3141 | 2.5479 | 99.57 | 20.87 | 37.385 | 79.9179 | 100.3661 | 77.024 | 0.009274 | 1.0082 | 0.024066 | 23.8698 | 0.012062 | 0.15681 | 0.004441 | 0.3167 | 0.3167 | 0.020351
2019-03-21 11:53 | 1553183605.3835 | 492 | 0.20868 | 0.001151 | 7e-05 | 99.9377 | 99.6565 | 119.7866 | 21.0877 | 0.002428 | 0.002314 | 163.7543 | 0.26228 | 8.402 | 0.003265 | 0.004338 | 0.004228 | 6.3e-05 | 0.000919 | 0.31797 | 1.3097 | 2.548 | 99.57 | 20.88 | 37.384 | 80.0748 | 100.2414 | 77.3273 | 0.009279 | 1.0073 | 0.024028 | 23.8531 | 0.025203 | 0.15603 | 0.004293 | 0.30547 | 0.30547 | 0.020994
2019-03-21 11:53 | 1553183606.3845 | 493 | 0.20872 | 0.001149 | 7.1e-05 | 99.9222 | 99.7958 | 119.6594 | 21.0874 | 0.002471 | 0.002266 | 163.7261 | 0.26346 | 8.4027 | 0.003227 | 0.004439 | 0.004317 | 6e-05 | 0.00092 | 0.28622 | 1.3105 | 2.5479 | 99.57 | 20.88 | 37.385 | 81.1014 | 99.3019 | 76.182 | 0.00928 | 1.0106 | 0.024054 | 23.8015 | 0.011322 | 0.15673 | 0.004086 | 0.29084 | 0.29084 | 0.022148
2019-03-21 11:53 | 1553183607.384 | 494 | 0.20867 | 0.001149 | 7.1e-05 | 99.9788 | 99.5958 | 119.7962 | 21.088 | 0.002465 | 0.002297 | 163.7097 | 0.26292 | 8.4015 | 0.003201 | 0.004972 | 0.004894 | 6e-05 | 0.000922 | 0.30359 | 1.2859 | 2.5479 | 99.57 | 20.88 | 37.387 | 80.9436 | 98.683 | 76.9184 | 0.00928 | 1.0085 | 0.024043 | 23.8402 | 0.034325 | 0.15641 | 0.004385 | 0.31227 | 0.31227 | 0.020586
2019-03-21 11:53 | 1553183608.3831 | 495 | 0.20869 | 0.001148 | 7.2e-05 | 100.0059 | 99.6061 | 120.1103 | 21.0873 | 0.002469 | 0.002273 | 163.6854 | 0.2623 | 8.4021 | 0.003188 | 0.004742 | 0.004686 | 6e-05 | 0.000919 | 0.31797 | 1.283 | 2.548 | 99.57 | 20.88 | 37.388 | 81.0491 | 98.3616 | 76.3471 | 0.00928 | 1.0102 | 0.024094 | 23.8519 | 0.035833 | 0.15604 | 0.004264 | 0.30412 | 0.30412 | 0.021088
2019-03-21 11:53 | 1553183609.3837 | 496 | 0.20861 | 0.001143 | 7.2e-05 | 99.9576 | 99.5458 | 119.9582 | 21.0879 | 0.002482 | 0.002309 | 163.6576 | 0.26284 | 8.4041 | 0.003143 | 0.004981 | 0.004911 | 6e-05 | 0.00092 | 0.32211 | 1.3231 | 2.5479 | 99.57 | 20.88 | 37.389 | 81.3638 | 97.2806 | 77.1951 | 0.009281 | 1.0077 | 0.02405 | 23.8657 | 0.03693 | 0.15636 | 0.004738 | 0.33775 | 0.33775 | 0.019028
2019-03-21 11:53 | 1553183610.3839 | 497 | 0.20868 | 0.001149 | 7.2e-05 | 100.0357 | 99.6779 | 119.9197 | 21.0882 | 0.002481 | 0.002227 | 163.6458 | 0.26288 | 8.4014 | 0.003204 | 0.004545 | 0.0045 | 5.8e-05 | 0.000924 | 0.31741 | 1.3146 | 2.5479 | 99.57 | 20.87 | 37.39 | 81.3329 | 98.7584 | 75.2412 | 0.009275 | 1.0134 | 0.024113 | 23.7952 | 0.032046 | 0.15639 | 0.004278 | 0.30539 | 0.30539 | 0.021047
2019-03-21 11:54 | 1553183611.3842 | 498 | 0.20865 | 0.00115 | 7.2e-05 | 100.0521 | 99.7106 | 119.873 | 21.089 | 0.002414 | 0.002247 | 163.6338 | 0.26312 | 8.4037 | 0.003216 | 0.004409 | 0.004247 | 5.5e-05 | 0.000924 | 0.33179 | 1.2964 | 2.5479 | 99.57 | 20.87 | 37.391 | 79.7182 | 99.0439 | 75.7088 | 0.009275 | 1.012 | 0.024092 | 23.8066 | 0.030578 | 0.15653 | 0.004474 | 0.31925 | 0.31925 | 0.020151
2019-03-21 11:54 | 1553183612.3829 | 499 | 0.20868 | 0.001143 | 7.3e-05 | 100.0232 | 99.8251 | 119.6446 | 21.0885 | 0.002447 | 0.002304 | 163.6195 | 0.26358 | 8.401 | 0.003302 | 0.004539 | 0.004497 | 6.2e-05 | 0.000919 | 0.32892 | 1.2947 | 2.5479 | 99.57 | 20.87 | 37.391 | 80.5105 | 101.114 | 77.0811 | 0.009275 | 1.008 | 0.024022 | 23.8306 | 0.017734 | 0.1568 | 0.004314 | 0.30679 | 0.30679 | 0.021006
2019-03-21 11:54 | 1553183613.383 | 500 | 0.20864 | 0.001142 | 7.3e-05 | 100.1316 | 99.9071 | 119.4878 | 21.0881 | 0.002426 | 0.00227 | 163.5892 | 0.26229 | 8.4014 | 0.003259 | 0.004741 | 0.004652 | 5.8e-05 | 0.000916 | 0.31133 | 1.2939 | 2.548 | 99.57 | 20.88 | 37.394 | 80.0261 | 100.0925 | 76.2722 | 0.009282 | 1.0104 | 0.024034 | 23.7875 | 0.020075 | 0.15604 | 0.004566 | 0.32513 | 0.32513 | 0.019725
2019-03-21 11:54 | 1553183614.3833 | 501 | 0.20867 | 0.001143 | 7.3e-05 | 100.1236 | 99.9933 | 119.792 | 21.089 | 0.002422 | 0.002274 | 163.5821 | 0.26343 | 8.4036 | 0.003235 | 0.004999 | 0.004949 | 6.1e-05 | 0.000917 | 0.33179 | 1.2608 | 2.5478 | 99.57 | 20.88 | 37.393 | 79.9137 | 99.4981 | 76.3588 | 0.009282 | 1.0101 | 0.024062 | 23.8207 | 0.011645 | 0.15671 | 0.004393 | 0.31299 | 0.31299 | 0.020579
2019-03-21 11:54 | 1553183615.3841 | 502 | 0.20867 | 0.00114 | 7.4e-05 | 100.1298 | 99.9637 | 120.1061 | 21.0894 | 0.00245 | 0.002267 | 163.5723 | 0.26253 | 8.4037 | 0.003245 | 0.004759 | 0.004709 | 6e-05 | 0.000918 | 0.31741 | 1.3236 | 2.548 | 99.57 | 20.87 | 37.397 | 80.5995 | 99.7448 | 76.2074 | 0.009277 | 1.0106 | 0.024098 | 23.8467 | 0.014855 | 0.15618 | 0.004368 | 0.31163 | 0.31163 | 0.020598
2019-03-21 11:54 | 1553183616.3831 | 503 | 0.20862 | 0.001145 | 7.4e-05 | 100.2034 | 100.0049 | 120.4342 | 21.0889 | 0.002492 | 0.002252 | 163.5691 | 0.26321 | 8.402 | 0.003339 | 0.004684 | 0.004624 | 6e-05 | 0.000916 | 0.27599 | 1.3279 | 2.548 | 99.57 | 20.87 | 37.397 | 81.5941 | 102.0323 | 75.8319 | 0.009277 | 1.0116 | 0.024144 | 23.8664 | 0.01774 | 0.15658 | 0.00465 | 0.3326 | 0.3326 | 0.019349
2019-03-21 11:54 | 1553183617.3845 | 504 | 0.20867 | 0.001139 | 7.5e-05 | 100.2145 | 100.1217 | 120.6105 | 21.0886 | 0.002478 | 0.002244 | 163.5373 | 0.26275 | 8.4034 | 0.003323 | 0.004358 | 0.004233 | 6.5e-05 | 0.000923 | 0.32549 | 1.3062 | 2.5479 | 99.57 | 20.87 | 37.398 | 81.2669 | 101.6369 | 75.6413 | 0.009277 | 1.0122 | 0.024169 | 23.8774 | 0.008281 | 0.15631 | 0.00437 | 0.31258 | 0.31258 | 0.020553
2019-03-21 11:54 | 1553183618.3846 | 505 | 0.20863 | 0.00114 | 7.5e-05 | 100.2615 | 100.1656 | 120.4612 | 21.0879 | 0.002487 | 0.00233 | 163.5097 | 0.26368 | 8.4016 | 0.00331 | 0.005466 | 0.00538 | 6.1e-05 | 0.000919 | 0.33179 | 1.3096 | 2.5478 | 99.57 | 20.87 | 37.397 | 81.4924 | 101.3082 | 77.697 | 0.009277 | 1.0063 | 0.024083 | 23.9328 | 0.008558 | 0.15686 | 0.004625 | 0.32995 | 0.32995 | 0.01954
2019-03-21 11:54 | 1553183619.3833 | 506 | 0.20865 | 0.001143 | 7.5e-05 | 100.269 | 100.0413 | 120.2539 | 21.0891 | 0.002459 | 0.002286 | 163.5066 | 0.26282 | 8.4012 | 0.003366 | 0.004371 | 0.004321 | 6.5e-05 | 0.00092 | 0.30862 | 1.2825 | 2.5478 | 99.57 | 20.87 | 37.398 | 80.8136 | 102.6672 | 76.6552 | 0.009277 | 1.0093 | 0.024098 | 23.8767 | 0.02033 | 0.15635 | 0.004474 | 0.31924 | 0.31924 | 0.020129
2019-03-21 11:54 | 1553183620.3843 | 507 | 0.20868 | 0.001142 | 7.4e-05 | 100.2532 | 99.9698 | 120.0194 | 21.0889 | 0.002456 | 0.002262 | 163.4801 | 0.26359 | 8.4002 | 0.003373 | 0.004823 | 0.004703 | 6.2e-05 | 0.000916 | 0.30342 | 1.309 | 2.548 | 99.57 | 20.87 | 37.399 | 80.7387 | 102.8428 | 76.0813 | 0.009277 | 1.0109 | 0.024094 | 23.8338 | 0.025319 | 0.15681 | 0.004292 | 0.30606 | 0.30606 | 0.021058
2019-03-21 11:54 | 1553183621.3847 | 508 | 0.2086 | 0.001145 | 7.5e-05 | 100.2824 | 99.9702 | 120.042 | 21.0886 | 0.002404 | 0.002288 | 163.4629 | 0.26234 | 8.4052 | 0.003296 | 0.004219 | 0.004176 | 6.2e-05 | 0.000918 | 0.32139 | 1.2989 | 2.5479 | 99.57 | 20.87 | 37.398 | 79.4775 | 100.9721 | 76.7103 | 0.009277 | 1.0091 | 0.024075 | 23.8575 | 0.027893 | 0.15607 | 0.004786 | 0.34144 | 0.34144 | 0.018786
2019-03-21 11:54 | 1553183622.3835 | 509 | 0.20869 | 0.001139 | 7.5e-05 | 100.3855 | 100.1227 | 120.4344 | 21.0896 | 0.002407 | 0.002301 | 163.4358 | 0.2633 | 8.4026 | 0.003365 | 0.004307 | 0.004189 | 6.4e-05 | 0.000916 | 0.32228 | 1.3428 | 2.5478 | 99.57 | 20.88 | 37.399 | 79.5486 | 102.6534 | 77.0088 | 0.009283 | 1.0082 | 0.024104 | 23.9067 | 0.023441 | 0.15664 | 0.004262 | 0.30393 | 0.30393 | 0.021182
2019-03-21 11:54 | 1553183623.3824 | 510 | 0.20865 | 0.00114 | 7.5e-05 | 100.3685 | 100.1915 | 120.7748 | 21.0896 | 0.002397 | 0.002287 | 163.4265 | 0.26288 | 8.4001 | 0.003322 | 0.005191 | 0.005035 | 6.5e-05 | 0.000918 | 0.2995 | 1.3513 | 2.5479 | 99.57 | 20.88 | 37.4 | 79.3196 | 101.6128 | 76.6737 | 0.009283 | 1.0092 | 0.024149 | 23.929 | 0.015784 | 0.15639 | 0.00448 | 0.32035 | 0.32035 | 0.020065
2019-03-21 11:54 | 1553183624.3839 | 511 | 0.2087 | 0.001144 | 7.5e-05 | 100.3521 | 100.0621 | 120.6302 | 21.0897 | 0.002456 | 0.002269 | 163.4066 | 0.26414 | 8.4007 | 0.003255 | 0.005414 | 0.005433 | 6.5e-05 | 0.000915 | 0.32621 | 1.2964 | 2.548 | 99.57 | 20.88 | 37.4 | 80.7364 | 100.0001 | 76.2512 | 0.009283 | 1.0104 | 0.024149 | 23.9002 | 0.025885 | 0.15714 | 0.004163 | 0.29728 | 0.29728 | 0.021725
2019-03-21 11:54 | 1553183625.3836 | 512 | 0.20866 | 0.001141 | 7.6e-05 | 100.3384 | 99.993 | 120.8222 | 21.0888 | 0.002452 | 0.002291 | 163.3985 | 0.2621 | 8.4038 | 0.003294 | 0.004854 | 0.004736 | 6.2e-05 | 0.000916 | 0.34036 | 1.2902 | 2.5479 | 99.57 | 20.88 | 37.403 | 80.6386 | 100.9219 | 76.7789 | 0.009284 | 1.0089 | 0.02415 | 23.9373 | 0.030846 | 0.15592 | 0.004459 | 0.31872 | 0.31872 | 0.020107
2019-03-21 11:54 | 1553183626.3836 | 513 | 0.20869 | 0.001138 | 7.7e-05 | 100.4632 | 100.275 | 120.4427 | 21.0896 | 0.002401 | 0.002272 | 163.3815 | 0.26235 | 8.4017 | 0.003314 | 0.005304 | 0.005304 | 6.3e-05 | 0.000917 | 0.30613 | 1.3167 | 2.548 | 99.57 | 20.88 | 37.405 | 79.4075 | 101.4197 | 76.3275 | 0.009285 | 1.0102 | 0.024128 | 23.8842 | 0.016777 | 0.15607 | 0.004247 | 0.30302 | 0.30302 | 0.021169
2019-03-21 11:54 | 1553183627.3826 | 514 | 0.2087 | 0.001134 | 7.8e-05 | 100.4212 | 100.2481 | 120.2805 | 21.0892 | 0.002485 | 0.002272 | 163.3766 | 0.26254 | 8.4024 | 0.003372 | 0.004756 | 0.004609 | 6.1e-05 | 0.000918 | 0.30359 | 1.3228 | 2.5479 | 99.57 | 20.88 | 37.406 | 81.4408 | 102.8158 | 76.3171 | 0.009285 | 1.0102 | 0.024112 | 23.8678 | 0.015436 | 0.15618 | 0.004213 | 0.30035 | 0.30035 | 0.021372
2019-03-21 11:54 | 1553183628.3827 | 515 | 0.2087 | 0.001135 | 7.9e-05 | 100.5163 | 100.1588 | 120.3349 | 21.0895 | 0.002444 | 0.002262 | 163.3594 | 0.26219 | 8.4026 | 0.003286 | 0.004439 | 0.004408 | 6.3e-05 | 0.000914 | 0.30182 | 1.2804 | 2.548 | 99.57 | 20.88 | 37.408 | 80.4423 | 100.7289 | 76.085 | 0.009285 | 1.0109 | 0.024126 | 23.8653 | 0.031863 | 0.15598 | 0.004191 | 0.29887 | 0.29887 | 0.02145
2019-03-21 11:54 | 1553183629.3838 | 516 | 0.20869 | 0.001134 | 8e-05 | 100.4671 | 100.0863 | 120.5084 | 21.09 | 0.002446 | 0.002304 | 163.3136 | 0.26288 | 8.4023 | 0.003316 | 0.004454 | 0.004295 | 5.8e-05 | 0.000918 | 0.2943 | 1.2682 | 2.5479 | 99.57 | 20.88 | 37.409 | 80.5043 | 101.4743 | 77.0794 | 0.009286 | 1.008 | 0.024109 | 23.9164 | 0.033969 | 0.15639 | 0.004275 | 0.30463 | 0.30463 | 0.021099
2019-03-21 11:54 | 1553183630.3843 | 517 | 0.20868 | 0.001131 | 8.1e-05 | 100.4305 | 100.1711 | 120.6102 | 21.0895 | 0.002441 | 0.002298 | 163.3078 | 0.26275 | 8.3989 | 0.003269 | 0.004602 | 0.004469 | 5.6e-05 | 0.000917 | 0.30414 | 1.299 | 2.548 | 99.57 | 20.88 | 37.411 | 80.3821 | 100.3169 | 76.9324 | 0.009286 | 1.0085 | 0.024124 | 23.9215 | 0.023129 | 0.15631 | 0.004333 | 0.30897 | 0.30897 | 0.020793
2019-03-21 11:54 | 1553183631.3829 | 518 | 0.2087 | 0.001127 | 8.2e-05 | 100.4975 | 100.306 | 120.5494 | 21.0902 | 0.002473 | 0.00231 | 163.2693 | 0.26219 | 8.4034 | 0.00323 | 0.004709 | 0.004585 | 6.1e-05 | 0.000917 | 0.30574 | 1.3093 | 2.5479 | 99.57 | 20.87 | 37.411 | 81.1462 | 99.384 | 77.2245 | 0.00928 | 1.0076 | 0.024108 | 23.9255 | 0.017063 | 0.15598 | 0.004187 | 0.29818 | 0.29818 | 0.021499
2019-03-21 11:54 | 1553183632.3839 | 519 | 0.20867 | 0.001124 | 8.2e-05 | 100.5705 | 100.2062 | 120.3774 | 21.0906 | 0.002478 | 0.002302 | 163.2549 | 0.26184 | 8.402 | 0.003192 | 0.005328 | 0.005297 | 6e-05 | 0.000916 | 0.31957 | 1.3318 | 2.5479 | 99.57 | 20.87 | 37.412 | 81.2698 | 98.4589 | 77.0293 | 0.009281 | 1.0082 | 0.024097 | 23.9017 | 0.032461 | 0.15577 | 0.004414 | 0.31438 | 0.31438 | 0.020364
2019-03-21 11:54 | 1553183633.3834 | 520 | 0.20864 | 0.001122 | 8.3e-05 | 100.594 | 100.1963 | 120.4938 | 21.0901 | 0.002463 | 0.002271 | 163.2265 | 0.26288 | 8.4012 | 0.003199 | 0.004168 | 0.004143 | 6.2e-05 | 0.000919 | 0.31902 | 1.3623 | 2.548 | 99.57 | 20.87 | 37.415 | 80.9153 | 98.6372 | 76.2835 | 0.009281 | 1.0103 | 0.024135 | 23.8878 | 0.035435 | 0.15639 | 0.004575 | 0.32657 | 0.32657 | 0.019682
2019-03-21 11:54 | 1553183634.3841 | 521 | 0.2087 | 0.001118 | 8.3e-05 | 100.645 | 100.4562 | 120.7515 | 21.0902 | 0.002414 | 0.002318 | 163.2134 | 0.26276 | 8.4024 | 0.003247 | 0.004552 | 0.004449 | 6.3e-05 | 0.000913 | 0.32156 | 1.3487 | 2.5479 | 99.57 | 20.87 | 37.416 | 79.7215 | 99.802 | 77.4196 | 0.009282 | 1.0071 | 0.024121 | 23.9522 | 0.016796 | 0.15632 | 0.00424 | 0.3021 | 0.3021 | 0.021267
2019-03-21 11:54 | 1553183635.3842 | 522 | 0.20873 | 0.00112 | 8.3e-05 | 100.6548 | 100.4674 | 120.498 | 21.0903 | 0.002457 | 0.002298 | 163.1748 | 0.26261 | 8.4024 | 0.003212 | 0.004563 | 0.004504 | 6e-05 | 0.00091 | 0.33555 | 1.3679 | 2.548 | 99.57 | 20.87 | 37.416 | 80.7564 | 98.9613 | 76.9457 | 0.009282 | 1.0084 | 0.024112 | 23.9108 | 0.016669 | 0.15623 | 0.004026 | 0.28647 | 0.28647 | 0.022414
2019-03-21 11:54 | 1553183636.3845 | 523 | 0.20866 | 0.001118 | 8.3e-05 | 100.5565 | 100.4773 | 120.7628 | 21.0912 | 0.002415 | 0.002262 | 163.1546 | 0.26366 | 8.4007 | 0.003279 | 0.00431 | 0.004128 | 6.2e-05 | 0.000909 | 0.31614 | 1.3927 | 2.548 | 99.57 | 20.87 | 37.418 | 79.7422 | 100.583 | 76.068 | 0.009282 | 1.011 | 0.024169 | 23.9071 | 0.007048 | 0.15685 | 0.004431 | 0.31653 | 0.31653 | 0.020366
2019-03-21 11:54 | 1553183637.3833 | 524 | 0.20871 | 0.001117 | 8.3e-05 | 100.6349 | 100.4575 | 120.952 | 21.0901 | 0.002476 | 0.00229 | 163.1191 | 0.26274 | 8.3994 | 0.003335 | 0.004838 | 0.004811 | 6e-05 | 0.000911 | 0.33107 | 1.3663 | 2.548 | 99.57 | 20.87 | 37.418 | 81.221 | 101.9354 | 76.7537 | 0.009282 | 1.009 | 0.024164 | 23.9493 | 0.015774 | 0.1563 | 0.004159 | 0.29674 | 0.29674 | 0.021649
2019-03-21 11:54 | 1553183638.3833 | 525 | 0.20877 | 0.001113 | 8.3e-05 | 100.6477 | 100.416 | 120.9832 | 21.0903 | 0.002461 | 0.002257 | 163.1162 | 0.26221 | 8.4044 | 0.003196 | 0.004603 | 0.004462 | 6.2e-05 | 0.000909 | 0.32781 | 1.3517 | 2.5478 | 99.57 | 20.87 | 37.417 | 80.8686 | 98.5654 | 75.9493 | 0.009282 | 1.0113 | 0.024195 | 23.9248 | 0.020613 | 0.15599 | 0.003778 | 0.26944 | 0.26944 | 0.023795
2019-03-21 11:54 | 1553183639.3837 | 526 | 0.20866 | 0.001112 | 8.3e-05 | 100.7636 | 100.5185 | 120.7957 | 21.0907 | 0.002411 | 0.002304 | 163.0935 | 0.2623 | 8.4037 | 0.003251 | 0.004826 | 0.004804 | 6.3e-05 | 0.000908 | 0.33212 | 1.3806 | 2.548 | 99.57 | 20.88 | 37.417 | 79.659 | 99.8895 | 77.0918 | 0.009288 | 1.008 | 0.024137 | 23.9453 | 0.02178 | 0.15604 | 0.004464 | 0.31853 | 0.31853 | 0.020134
2019-03-21 11:54 | 1553183640.3824 | 527 | 0.20874 | 0.001112 | 8.2e-05 | 100.7597 | 100.4872 | 121.1417 | 21.0909 | 0.002379 | 0.002305 | 163.0557 | 0.26273 | 8.4037 | 0.003279 | 0.005059 | 0.004937 | 6.2e-05 | 0.000905 | 0.2943 | 1.3178 | 2.5479 | 99.57 | 20.88 | 37.416 | 78.8881 | 100.5802 | 77.1005 | 0.009287 | 1.008 | 0.024171 | 23.9799 | 0.024225 | 0.1563 | 0.003959 | 0.28236 | 0.28236 | 0.022751
2019-03-21 11:54 | 1553183641.3835 | 528 | 0.20865 | 0.00111 | 8.3e-05 | 100.7794 | 100.7579 | 121.3585 | 21.0903 | 0.002448 | 0.002225 | 163.059 | 0.26327 | 8.4065 | 0.003222 | 0.004959 | 0.004898 | 6.4e-05 | 0.000906 | 0.31238 | 1.2911 | 2.5479 | 99.57 | 20.88 | 37.418 | 80.5514 | 99.1887 | 75.1901 | 0.009288 | 1.0135 | 0.024259 | 23.9358 | 0.001902 | 0.15662 | 0.004546 | 0.32615 | 0.32615 | 0.019737
2019-03-21 11:54 | 1553183642.3839 | 529 | 0.20877 | 0.001106 | 8.3e-05 | 100.8696 | 100.6626 | 121.286 | 21.0917 | 0.002479 | 0.002329 | 163.0317 | 0.26329 | 8.4009 | 0.003225 | 0.004736 | 0.004654 | 5.9e-05 | 0.000904 | 0.29048 | 1.2803 | 2.5479 | 99.57 | 20.88 | 37.419 | 81.2865 | 99.2725 | 77.6969 | 0.009288 | 1.0063 | 0.024165 | 24.0146 | 0.01838 | 0.15663 | 0.003817 | 0.27192 | 0.27192 | 0.023674
2019-03-21 11:54 | 1553183643.3831 | 530 | 0.20867 | 0.001105 | 8.4e-05 | 100.8476 | 100.7167 | 120.8635 | 21.0911 | 0.002411 | 0.002254 | 163.0159 | 0.26261 | 8.4012 | 0.003281 | 0.005013 | 0.004882 | 5.6e-05 | 0.000902 | 0.31813 | 1.2784 | 2.548 | 99.57 | 20.88 | 37.419 | 79.6491 | 100.63 | 75.8973 | 0.009288 | 1.0115 | 0.024185 | 23.9112 | 0.011614 | 0.15623 | 0.004399 | 0.31435 | 0.31435 | 0.020426
2019-03-21 11:54 | 1553183644.3837 | 531 | 0.20865 | 0.001105 | 8.4e-05 | 100.8792 | 100.6374 | 120.6182 | 21.0906 | 0.002448 | 0.002266 | 162.9902 | 0.26229 | 8.4041 | 0.003279 | 0.004636 | 0.004564 | 6e-05 | 0.000903 | 0.30055 | 1.2904 | 2.5479 | 99.57 | 20.88 | 37.417 | 80.5431 | 100.5785 | 76.1774 | 0.009288 | 1.0106 | 0.024151 | 23.8965 | 0.021465 | 0.15604 | 0.004524 | 0.32299 | 0.32299 | 0.019856
2019-03-21 11:54 | 1553183645.384 | 532 | 0.20875 | 0.0011 | 8.5e-05 | 100.8684 | 100.7654 | 120.5708 | 21.0916 | 0.002438 | 0.00227 | 162.9777 | 0.26351 | 8.402 | 0.003211 | 0.004897 | 0.004823 | 6.2e-05 | 0.000909 | 0.34346 | 1.3104 | 2.5479 | 99.57 | 20.88 | 37.416 | 80.3054 | 98.935 | 76.2812 | 0.009287 | 1.0103 | 0.024142 | 23.8953 | 0.00914 | 0.15676 | 0.003917 | 0.27883 | 0.27883 | 0.023107
2019-03-21 11:54 | 1553183646.3837 | 533 | 0.2087 | 0.001095 | 8.5e-05 | 100.8815 | 100.8138 | 120.2943 | 21.0906 | 0.002446 | 0.002232 | 162.9415 | 0.26353 | 8.4028 | 0.003276 | 0.00481 | 0.004742 | 6.2e-05 | 0.000907 | 0.30773 | 1.3414 | 2.548 | 99.57 | 20.88 | 37.414 | 80.4961 | 100.5071 | 75.3557 | 0.009287 | 1.013 | 0.024147 | 23.8363 | 0.00601 | 0.15677 | 0.004256 | 0.30344 | 0.30344 | 0.021234
2019-03-21 11:54 | 1553183647.3827 | 534 | 0.20864 | 0.001097 | 8.6e-05 | 100.9102 | 100.8036 | 119.9648 | 21.0909 | 0.002456 | 0.002276 | 162.9144 | 0.26232 | 8.4028 | 0.003307 | 0.004109 | 0.004069 | 6e-05 | 0.000906 | 0.31868 | 1.3435 | 2.548 | 99.57 | 20.88 | 37.413 | 80.7289 | 101.2467 | 76.4239 | 0.009287 | 1.0099 | 0.024077 | 23.8401 | 0.009452 | 0.15606 | 0.004579 | 0.32582 | 0.32582 | 0.019685
2019-03-21 11:54 | 1553183648.3828 | 535 | 0.20871 | 0.001097 | 8.6e-05 | 100.9462 | 100.7341 | 120.1258 | 21.0907 | 0.002475 | 0.00228 | 162.8899 | 0.26321 | 8.4041 | 0.003279 | 0.005015 | 0.004962 | 5.9e-05 | 0.000906 | 0.31293 | 1.3174 | 2.5479 | 99.57 | 20.88 | 37.411 | 81.2005 | 100.566 | 76.5071 | 0.009286 | 1.0097 | 0.02409 | 23.8589 | 0.018815 | 0.15659 | 0.004175 | 0.29682 | 0.29682 | 0.021682
2019-03-21 11:54 | 1553183649.3832 | 536 | 0.20878 | 0.001095 | 8.6e-05 | 100.9784 | 100.856 | 120.1223 | 21.0903 | 0.002412 | 0.00232 | 162.8855 | 0.26235 | 8.4034 | 0.003284 | 0.004932 | 0.004804 | 5.6e-05 | 0.000908 | 0.31924 | 1.3265 | 2.5477 | 99.57 | 20.88 | 37.41 | 79.6874 | 100.6887 | 77.4663 | 0.009286 | 1.0069 | 0.024057 | 23.8913 | 0.010845 | 0.15607 | 0.003779 | 0.26782 | 0.26782 | 0.023951
2019-03-21 11:54 | 1553183650.3829 | 537 | 0.20877 | 0.001091 | 8.6e-05 | 101.0547 | 100.9405 | 119.8918 | 21.0908 | 0.002462 | 0.002319 | 162.8706 | 0.26301 | 8.4012 | 0.003208 | 0.004767 | 0.004726 | 5.8e-05 | 0.000909 | 0.29407 | 1.316 | 2.5479 | 99.57 | 20.88 | 37.409 | 80.8824 | 98.8528 | 77.4529 | 0.009286 | 1.007 | 0.024034 | 23.8679 | 0.010111 | 0.15646 | 0.003846 | 0.27237 | 0.27237 | 0.023609
2019-03-21 11:54 | 1553183651.3829 | 538 | 0.20874 | 0.001089 | 8.7e-05 | 101.0713 | 100.869 | 119.8591 | 21.0915 | 0.00248 | 0.002302 | 162.8502 | 0.26243 | 8.405 | 0.00327 | 0.004397 | 0.004305 | 5.9e-05 | 0.00091 | 0.31188 | 1.3524 | 2.548 | 99.57 | 20.89 | 37.406 | 81.3245 | 100.3548 | 77.0409 | 0.009291 | 1.0082 | 0.024045 | 23.8506 | 0.017928 | 0.15612 | 0.003986 | 0.28254 | 0.28254 | 0.022711
2019-03-21 11:54 | 1553183652.3834 | 539 | 0.20877 | 0.001086 | 8.8e-05 | 101.0648 | 100.8421 | 120.0612 | 21.0914 | 0.002462 | 0.002302 | 162.8119 | 0.26222 | 8.4043 | 0.003198 | 0.003835 | 0.003781 | 5.7e-05 | 0.000909 | 0.2855 | 1.3103 | 2.5479 | 99.57 | 20.88 | 37.404 | 80.8829 | 98.6181 | 77.0324 | 0.009284 | 1.0082 | 0.024066 | 23.8704 | 0.019734 | 0.15599 | 0.003854 | 0.27322 | 0.27322 | 0.023466
2019-03-21 11:54 | 1553183653.383 | 540 | 0.2088 | 0.001085 | 8.8e-05 | 101.1023 | 100.9088 | 120.3136 | 21.0918 | 0.0025 | 0.002275 | 162.7742 | 0.26279 | 8.4046 | 0.003305 | 0.004557 | 0.004434 | 5.8e-05 | 0.000908 | 0.33069 | 1.3465 | 2.548 | 99.57 | 20.88 | 37.404 | 81.8027 | 101.1927 | 76.3852 | 0.009284 | 1.01 | 0.024113 | 23.8734 | 0.017137 | 0.15634 | 0.003631 | 0.25761 | 0.25761 | 0.024942
2019-03-21 11:54 | 1553183654.384 | 541 | 0.20877 | 0.001087 | 8.8e-05 | 101.1291 | 101.025 | 120.3704 | 21.0923 | 0.00238 | 0.00223 | 162.7564 | 0.26377 | 8.405 | 0.003242 | 0.004738 | 0.004692 | 5.8e-05 | 0.000909 | 0.33035 | 1.3219 | 2.5479 | 99.57 | 20.88 | 37.398 | 78.9072 | 99.6824 | 75.3215 | 0.009283 | 1.0131 | 0.024156 | 23.8427 | 0.00922 | 0.15692 | 0.003805 | 0.27064 | 0.27064 | 0.02383
2019-03-21 11:54 | 1553183655.3832 | 542 | 0.2087 | 0.001081 | 8.8e-05 | 101.1017 | 101.135 | 120.4715 | 21.0909 | 0.002436 | 0.00227 | 162.7443 | 0.26273 | 8.4008 | 0.003286 | 0.004994 | 0.004948 | 5.6e-05 | 0.000906 | 0.2912 | 1.3308 | 2.548 | 99.57 | 20.88 | 37.397 | 80.247 | 100.7483 | 76.2665 | 0.009283 | 1.0104 | 0.024133 | 23.885 | -0.00294 | 0.1563 | 0.004236 | 0.30162 | 0.30162 | 0.021298
2019-03-21 11:54 | 1553183656.3823 | 543 | 0.20887 | 0.001081 | 8.9e-05 | 101.1351 | 100.9892 | 120.8574 | 21.0918 | 0.002511 | 0.002229 | 162.7335 | 0.26298 | 8.4012 | 0.003325 | 0.004614 | 0.004523 | 5.9e-05 | 0.000908 | 0.32997 | 1.2801 | 2.5479 | 99.57 | 20.88 | 37.397 | 82.0625 | 101.6929 | 75.2836 | 0.009283 | 1.0132 | 0.024206 | 23.8895 | 0.012914 | 0.15644 | 0.003224 | 0.22897 | 0.22897 | 0.028082
2019-03-21 11:54 | 1553183657.3833 | 544 | 0.20878 | 0.001081 | 8.9e-05 | 101.1217 | 100.9272 | 121.1513 | 21.0925 | 0.002425 | 0.002272 | 162.7297 | 0.26338 | 8.402 | 0.003292 | 0.005127 | 0.005019 | 5.5e-05 | 0.000907 | 0.31725 | 1.2677 | 2.548 | 99.57 | 20.88 | 37.395 | 79.9998 | 100.88 | 76.3084 | 0.009282 | 1.0103 | 0.0242 | 23.9537 | 0.01722 | 0.15668 | 0.003758 | 0.26766 | 0.26766 | 0.024059
2019-03-21 11:54 | 1553183658.3836 | 545 | 0.20867 | 0.001075 | 9e-05 | 101.1697 | 101.0491 | 120.921 | 21.0925 | 0.002414 | 0.002259 | 162.7169 | 0.26293 | 8.4038 | 0.003365 | 0.00487 | 0.004778 | 5.9e-05 | 0.00091 | 0.33179 | 1.2783 | 2.5479 | 99.57 | 20.88 | 37.392 | 79.7355 | 102.6618 | 76.0185 | 0.009281 | 1.0111 | 0.024187 | 23.921 | 0.010662 | 0.15642 | 0.004419 | 0.31549 | 0.31549 | 0.020377
2019-03-21 11:54 | 1553183659.3824 | 546 | 0.20885 | 0.001078 | 9e-05 | 101.164 | 101.0542 | 120.8492 | 21.0929 | 0.002457 | 0.002285 | 162.6994 | 0.26358 | 8.4034 | 0.003234 | 0.005089 | 0.005023 | 5.7e-05 | 0.000909 | 0.33627 | 1.2796 | 2.548 | 99.57 | 20.88 | 37.391 | 80.7683 | 99.476 | 76.6214 | 0.009281 | 1.0094 | 0.024159 | 23.9346 | 0.00971 | 0.1568 | 0.003363 | 0.23856 | 0.23856 | 0.027015
2019-03-21 11:54 | 1553183660.3835 | 547 | 0.20868 | 0.001076 | 9e-05 | 101.2021 | 100.9146 | 120.9213 | 21.0927 | 0.002442 | 0.002317 | 162.6957 | 0.26306 | 8.4013 | 0.003305 | 0.004739 | 0.004653 | 5.6e-05 | 0.000913 | 0.35601 | 1.2929 | 2.5479 | 99.57 | 20.88 | 37.388 | 80.393 | 101.1982 | 77.4126 | 0.00928 | 1.0071 | 0.024139 | 23.9688 | 0.025449 | 0.15649 | 0.004364 | 0.31085 | 0.31085 | 0.020691
2019-03-21 11:54 | 1553183661.3824 | 548 | 0.20875 | 0.001076 | 9e-05 | 101.1783 | 100.9129 | 121.1635 | 21.093 | 0.002416 | 0.00225 | 162.6857 | 0.26327 | 8.403 | 0.003301 | 0.00549 | 0.005341 | 5.6e-05 | 0.000913 | 0.32173 | 1.3013 | 2.5479 | 99.57 | 20.88 | 37.387 | 79.7821 | 101.105 | 75.7946 | 0.00928 | 1.0118 | 0.024219 | 23.9373 | 0.023494 | 0.15662 | 0.003995 | 0.28508 | 0.28508 | 0.022579
2019-03-21 11:54 | 1553183662.3825 | 549 | 0.20876 | 0.001074 | 9e-05 | 101.2209 | 101.0049 | 121.2154 | 21.0928 | 0.002462 | 0.002268 | 162.6832 | 0.26298 | 8.4027 | 0.003279 | 0.005292 | 0.005213 | 5.9e-05 | 0.00091 | 0.32261 | 1.35 | 2.5478 | 99.57 | 20.88 | 37.385 | 80.8961 | 100.5712 | 76.2146 | 0.00928 | 1.0105 | 0.024209 | 23.9568 | 0.019112 | 0.15645 | 0.00393 | 0.28021 | 0.28021 | 0.022947
2019-03-21 11:54 | 1553183663.383 | 550 | 0.20877 | 0.001068 | 9e-05 | 101.2081 | 101.0004 | 121.0729 | 21.0925 | 0.002433 | 0.002284 | 162.6728 | 0.26344 | 8.4052 | 0.003302 | 0.00473 | 0.004611 | 5.5e-05 | 0.000912 | 0.29535 | 1.3651 | 2.548 | 99.57 | 20.88 | 37.383 | 80.186 | 101.122 | 76.6001 | 0.009279 | 1.0094 | 0.024182 | 23.956 | 0.018372 | 0.15672 | 0.003844 | 0.2737 | 0.2737 | 0.023534
2019-03-21 11:54 | 1553183664.383 | 551 | 0.20875 | 0.001069 | 9.1e-05 | 101.259 | 101.1929 | 121.0149 | 21.0927 | 0.002362 | 0.002234 | 162.6505 | 0.26253 | 8.4001 | 0.003331 | 0.005043 | 0.004934 | 6.1e-05 | 0.000911 | 0.30182 | 1.3587 | 2.5479 | 99.57 | 20.88 | 37.382 | 78.4835 | 101.8208 | 75.4022 | 0.009279 | 1.0129 | 0.024217 | 23.9092 | 0.005848 | 0.15618 | 0.004004 | 0.28566 | 0.28566 | 0.02247
2019-03-21 11:54 | 1553183665.3827 | 552 | 0.20892 | 0.001071 | 9e-05 | 101.2534 | 101.1813 | 121.0595 | 21.093 | 0.002449 | 0.0023 | 162.6142 | 0.2627 | 8.4029 | 0.003326 | 0.004765 | 0.004761 | 5.9e-05 | 0.00091 | 0.30879 | 1.3439 | 2.5479 | 99.57 | 20.88 | 37.38 | 80.578 | 101.7105 | 77.0035 | 0.009278 | 1.0083 | 0.024166 | 23.9685 | 0.006372 | 0.15628 | 0.002927 | 0.20696 | 0.20696 | 0.031035
2019-03-21 11:54 | 1553183666.3842 | 553 | 0.20877 | 0.001066 | 9.1e-05 | 101.2647 | 101.2211 | 121.1027 | 21.0928 | 0.002408 | 0.00232 | 162.5943 | 0.2631 | 8.401 | 0.003303 | 0.004893 | 0.00476 | 5.7e-05 | 0.00091 | 0.31061 | 1.3181 | 2.5479 | 99.57 | 20.88 | 37.375 | 79.5846 | 101.1425 | 77.4756 | 0.009277 | 1.0069 | 0.024154 | 23.9889 | 0.00385 | 0.15652 | 0.003856 | 0.27424 | 0.27424 | 0.023457
2019-03-21 11:54 | 1553183667.3838 | 554 | 0.20857 | 0.001061 | 9.2e-05 | 101.2075 | 101.1118 | 121.2529 | 21.0929 | 0.002465 | 0.002295 | 162.565 | 0.26266 | 8.4019 | 0.003314 | 0.0044 | 0.00431 | 5.9e-05 | 0.000908 | 0.35618 | 1.2979 | 2.5479 | 99.57 | 20.88 | 37.374 | 80.9663 | 101.4277 | 76.8794 | 0.009277 | 1.0086 | 0.02419 | 23.9833 | 0.008462 | 0.15625 | 0.005083 | 0.36363 | 0.36363 | 0.017661
2019-03-21 11:54 | 1553183668.3832 | 555 | 0.20898 | 0.001069 | 9.2e-05 | 101.2272 | 101.3124 | 121.2867 | 21.0932 | 0.002465 | 0.002225 | 162.5325 | 0.26282 | 8.4013 | 0.00332 | 0.004843 | 0.004712 | 5.7e-05 | 0.000905 | 0.31758 | 1.3215 | 2.548 | 99.57 | 20.88 | 37.371 | 80.9534 | 101.5592 | 75.1908 | 0.009276 | 1.0135 | 0.024252 | 23.9288 | -0.007529 | 0.15635 | 0.002589 | 0.18296 | 0.18296 | 0.035123
2019-03-21 11:54 | 1553183669.3843 | 556 | 0.20866 | 0.001064 | 9.3e-05 | 101.2671 | 101.1319 | 121.4858 | 21.0939 | 0.002446 | 0.002221 | 162.5195 | 0.26367 | 8.4034 | 0.003434 | 0.005071 | 0.004961 | 5.9e-05 | 0.000904 | 0.28031 | 1.3024 | 2.5479 | 99.57 | 20.88 | 37.371 | 80.5057 | 104.335 | 75.1002 | 0.009276 | 1.0138 | 0.024275 | 23.9453 | 0.011949 | 0.15686 | 0.004546 | 0.32568 | 0.32568 | 0.019795
2019-03-21 11:54 | 1553183670.3823 | 557 | 0.2087 | 0.00107 | 9.4e-05 | 101.2835 | 101.1826 | 121.4344 | 21.0927 | 0.002458 | 0.002311 | 162.4858 | 0.26264 | 8.4019 | 0.003305 | 0.005042 | 0.004922 | 6.2e-05 | 0.000904 | 0.30198 | 1.293 | 2.548 | 99.57 | 20.89 | 37.369 | 80.7782 | 101.2061 | 77.248 | 0.009281 | 1.0076 | 0.024195 | 24.0139 | 0.008909 | 0.15624 | 0.004259 | 0.30375 | 0.30375 | 0.021141
2019-03-21 11:54 | 1553183671.382 | 558 | 0.20905 | 0.001063 | 9.4e-05 | 101.2873 | 101.043 | 121.183 | 21.0936 | 0.002407 | 0.002239 | 162.4748 | 0.26271 | 8.4008 | 0.003341 | 0.00398 | 0.003877 | 6.3e-05 | 0.000902 | 0.31758 | 1.3007 | 2.5479 | 99.57 | 20.88 | 37.369 | 79.5666 | 102.0635 | 75.5395 | 0.009276 | 1.0125 | 0.024229 | 23.9305 | 0.021604 | 0.15629 | 0.002187 | 0.15343 | 0.15343 | 0.041866
2019-03-21 11:55 | 1553183672.3821 | 559 | 0.20849 | 0.001069 | 9.4e-05 | 101.2793 | 101.0869 | 121.152 | 21.0936 | 0.002412 | 0.002317 | 162.4505 | 0.26287 | 8.4038 | 0.003371 | 0.003888 | 0.00374 | 6.3e-05 | 0.000902 | 0.30989 | 1.2754 | 2.5479 | 99.57 | 20.88 | 37.368 | 79.6944 | 102.798 | 77.4103 | 0.009275 | 1.0071 | 0.024162 | 23.9915 | 0.017005 | 0.15638 | 0.00554 | 0.39619 | 0.39619 | 0.016223
2019-03-21 11:55 | 1553183673.3821 | 560 | 0.20879 | 0.001068 | 9.4e-05 | 101.3368 | 101.0733 | 121.2424 | 21.0936 | 0.002429 | 0.00231 | 162.439 | 0.26376 | 8.4013 | 0.003323 | 0.004222 | 0.004133 | 6.2e-05 | 0.000902 | 0.31293 | 1.2738 | 2.5479 | 99.57 | 20.88 | 37.367 | 80.0843 | 101.6286 | 77.2334 | 0.009275 | 1.0076 | 0.024177 | 23.9944 | 0.023288 | 0.15691 | 0.003753 | 0.26676 | 0.26676 | 0.024175
2019-03-21 11:55 | 1553183674.3836 | 561 | 0.20892 | 0.001063 | 9.5e-05 | 101.2912 | 101.23 | 121.5862 | 21.0937 | 0.002377 | 0.002308 | 162.4361 | 0.26253 | 8.4035 | 0.003288 | 0.004661 | 0.004642 | 5.8e-05 | 0.000899 | 0.2907 | 1.2916 | 2.548 | 99.57 | 20.88 | 37.365 | 78.8541 | 100.8022 | 77.1796 | 0.009275 | 1.0078 | 0.024213 | 24.0266 | 0.005409 | 0.15618 | 0.002922 | 0.20671 | 0.20671 | 0.031053
2019-03-21 11:55 | 1553183675.3823 | 562 | 0.20844 | 0.001067 | 9.5e-05 | 101.2873 | 101.0515 | 121.5836 | 21.0943 | 0.002431 | 0.002257 | 162.4233 | 0.26377 | 8.404 | 0.003236 | 0.004278 | 0.004133 | 6e-05 | 0.000902 | 0.33229 | 1.3008 | 2.5479 | 99.57 | 20.88 | 37.365 | 80.1374 | 99.5401 | 75.9556 | 0.009275 | 1.0113 | 0.024255 | 23.9843 | 0.02085 | 0.15692 | 0.005818 | 0.4179 | 0.4179 | 0.015432
2019-03-21 11:55 | 1553183676.3826 | 563 | 0.20917 | 0.001068 | 9.6e-05 | 101.3046 | 101.0716 | 121.3799 | 21.0945 | 0.00245 | 0.002246 | 162.3972 | 0.26315 | 8.4048 | 0.003262 | 0.005029 | 0.004936 | 6.1e-05 | 0.000899 | 0.30165 | 1.3267 | 2.5481 | 99.57 | 20.88 | 37.364 | 80.6021 | 100.153 | 75.6925 | 0.009274 | 1.012 | 0.024244 | 23.9552 | 0.020598 | 0.15655 | 0.001454 | 0.10005 | 0.10005 | 0.064311
2019-03-21 11:55 | 1553183677.3828 | 564 | 0.20869 | 0.001065 | 9.6e-05 | 101.349 | 101.0914 | 121.1738 | 21.0938 | 0.002439 | 0.002293 | 162.3988 | 0.26318 | 8.4022 | 0.003202 | 0.005085 | 0.005068 | 5.9e-05 | 0.000897 | 0.30397 | 1.3293 | 2.5479 | 99.57 | 20.88 | 37.364 | 80.3301 | 98.7146 | 76.8189 | 0.009274 | 1.0088 | 0.024184 | 23.9734 | 0.022768 | 0.15656 | 0.004355 | 0.31039 | 0.31039 | 0.020731
2019-03-21 11:55 | 1553183678.3824 | 565 | 0.20845 | 0.001067 | 9.7e-05 | 101.3457 | 101.1012 | 121.1726 | 21.0936 | 0.002508 | 0.002255 | 162.3902 | 0.26272 | 8.4 | 0.003165 | 0.004526 | 0.004398 | 6e-05 | 0.000901 | 0.32173 | 1.3253 | 2.5478 | 99.57 | 20.88 | 37.363 | 81.9816 | 97.8191 | 75.9129 | 0.009274 | 1.0114 | 0.024215 | 23.9423 | 0.021605 | 0.15629 | 0.005798 | 0.41564 | 0.41564 | 0.015455
2019-03-21 11:55 | 1553183679.3828 | 566 | 0.20896 | 0.001059 | 9.8e-05 | 101.2829 | 101.0117 | 121.3698 | 21.0947 | 0.002448 | 0.002287 | 162.3615 | 0.26269 | 8.4039 | 0.003178 | 0.005038 | 0.004933 | 6e-05 | 0.0009 | 0.31542 | 1.3137 | 2.548 | 99.57 | 20.88 | 37.363 | 80.5588 | 98.1321 | 76.6873 | 0.009274 | 1.0092 | 0.024208 | 23.9883 | 0.023983 | 0.15627 | 0.002746 | 0.1937 | 0.1937 | 0.033159
2019-03-21 11:55 | 1553183680.3829 | 567 | 0.20895 | 0.001057 | 9.9e-05 | 101.2593 | 101.1258 | 121.421 | 21.0945 | 0.0024 | 0.002239 | 162.3407 | 0.26206 | 8.4032 | 0.003116 | 0.004926 | 0.004852 | 6.3e-05 | 0.0009 | 0.2975 | 1.3353 | 2.5481 | 99.57 | 20.88 | 37.362 | 79.3952 | 96.6336 | 75.5202 | 0.009274 | 1.0125 | 0.024254 | 23.9533 | 0.011807 | 0.1559 | 0.00275 | 0.19429 | 0.19429 | 0.032979
2019-03-21 11:55 | 1553183681.3826 | 568 | 0.20879 | 0.001065 | 9.9e-05 | 101.2564 | 101.0307 | 121.643 | 21.0944 | 0.002433 | 0.002243 | 162.3207 | 0.26237 | 8.4025 | 0.003197 | 0.005177 | 0.005081 | 6e-05 | 0.0009 | 0.32964 | 1.3591 | 2.5479 | 99.57 | 20.88 | 37.359 | 80.1833 | 98.5872 | 75.6361 | 0.009273 | 1.0122 | 0.024272 | 23.9792 | 0.019953 | 0.15608 | 0.00374 | 0.26659 | 0.26659 | 0.024064
2019-03-21 11:55 | 1553183682.383 | 569 | 0.20862 | 0.001063 | 9.9e-05 | 101.2573 | 100.8083 | 121.6808 | 21.0952 | 0.002458 | 0.002282 | 162.2932 | 0.2629 | 8.4016 | 0.003175 | 0.003971 | 0.00391 | 6e-05 | 0.000899 | 0.31653 | 1.3061 | 2.548 | 99.57 | 20.88 | 37.358 | 80.7963 | 98.0504 | 76.5535 | 0.009273 | 1.0096 | 0.024244 | 24.0144 | 0.039743 | 0.1564 | 0.004783 | 0.34212 | 0.34212 | 0.018789
2019-03-21 11:55 | 1553183683.3841 | 570 | 0.20823 | 0.001056 | 9.9e-05 | 101.2489 | 100.965 | 121.7557 | 21.0936 | 0.00244 | 0.002326 | 162.2624 | 0.26348 | 8.4028 | 0.00325 | 0.004895 | 0.004815 | 5.7e-05 | 0.000901 | 0.29717 | 1.3398 | 2.5479 | 99.57 | 20.88 | 37.354 | 80.3604 | 99.8798 | 77.6074 | 0.009272 | 1.0065 | 0.024215 | 24.058 | 0.025116 | 0.15675 | 0.007117 | 0.51131 | 0.51131 | 0.0126
2019-03-21 11:55 | 1553183684.3827 | 571 | 0.2092 | 0.001055 | 0.0001 | 101.3133 | 100.9436 | 121.5473 | 21.0947 | 0.002423 | 0.002254 | 162.2397 | 0.26287 | 8.4009 | 0.003241 | 0.004388 | 0.004326 | 5.6e-05 | 0.000902 | 0.32029 | 1.2964 | 2.548 | 99.57 | 20.88 | 37.353 | 79.9544 | 99.6613 | 75.8976 | 0.009272 | 1.0115 | 0.024253 | 23.9787 | 0.032701 | 0.15638 | 0.001273 | 0.086703 | 0.086703 | 0.07413
2019-03-21 11:55 | 1553183685.3821 | 572 | 0.20888 | 0.001053 | 9.9e-05 | 101.2531 | 100.887 | 121.5681 | 21.0956 | 0.002409 | 0.002292 | 162.2259 | 0.26282 | 8.402 | 0.003219 | 0.004818 | 0.004719 | 5.8e-05 | 0.000901 | 0.31598 | 1.2947 | 2.5478 | 99.57 | 20.88 | 37.352 | 79.6239 | 99.1257 | 76.8042 | 0.009271 | 1.0088 | 0.024224 | 24.0119 | 0.032397 | 0.15635 | 0.003185 | 0.22564 | 0.22564 | 0.028479
2019-03-21 11:55 | 1553183686.3832 | 573 | 0.20851 | 0.001052 | 9.9e-05 | 101.2614 | 101.0045 | 121.6198 | 21.0957 | 0.002429 | 0.002245 | 162.2129 | 0.26385 | 8.4049 | 0.003283 | 0.004295 | 0.004238 | 5.8e-05 | 0.000902 | 0.30989 | 1.2982 | 2.548 | 99.57 | 20.88 | 37.352 | 80.0918 | 100.6775 | 75.6739 | 0.009271 | 1.0121 | 0.024268 | 23.9782 | 0.022722 | 0.15696 | 0.005416 | 0.38858 | 0.38858 | 0.016602
2019-03-21 11:55 | 1553183687.3819 | 574 | 0.20901 | 0.001041 | 9.9e-05 | 101.2653 | 100.9619 | 121.3097 | 21.0958 | 0.002439 | 0.002228 | 162.1995 | 0.26351 | 8.4028 | 0.003252 | 0.0045 | 0.004381 | 5.9e-05 | 0.000901 | 0.28584 | 1.2651 | 2.5478 | 99.57 | 20.88 | 37.351 | 80.3351 | 99.9298 | 75.2649 | 0.009271 | 1.0133 | 0.024252 | 23.9336 | 0.026835 | 0.15676 | 0.002432 | 0.17111 | 0.17111 | 0.037653
2019-03-21 11:55 | 1553183688.3819 | 575 | 0.20869 | 0.001048 | 9.9e-05 | 101.2331 | 100.8107 | 121.3182 | 21.095 | 0.002424 | 0.002219 | 162.15 | 0.26307 | 8.4014 | 0.003229 | 0.005051 | 0.005032 | 5.6e-05 | 0.000899 | 0.29717 | 1.3083 | 2.5478 | 99.57 | 20.88 | 37.35 | 79.9758 | 99.3649 | 75.0481 | 0.009271 | 1.0139 | 0.02426 | 23.927 | 0.0374 | 0.1565 | 0.00439 | 0.31376 | 0.31376 | 0.0205
2019-03-21 11:55 | 1553183689.3833 | 576 | 0.20867 | 0.001051 | 9.9e-05 | 101.1923 | 100.886 | 120.8433 | 21.0955 | 0.002504 | 0.002219 | 162.1129 | 0.26335 | 8.4039 | 0.003277 | 0.004322 | 0.004195 | 5.9e-05 | 0.000899 | 0.2855 | 1.3083 | 2.548 | 99.57 | 20.88 | 37.35 | 81.889 | 100.5245 | 75.0573 | 0.009271 | 1.0139 | 0.024212 | 23.8804 | 0.027116 | 0.15667 | 0.004468 | 0.3188 | 0.3188 | 0.020198
2019-03-21 11:55 | 1553183690.3821 | 577 | 0.20878 | 0.001048 | 9.9e-05 | 101.1831 | 100.9541 | 120.8616 | 21.0954 | 0.00244 | 0.002232 | 162.0947 | 0.26282 | 8.402 | 0.003291 | 0.005355 | 0.005294 | 5.6e-05 | 0.000903 | 0.32908 | 1.296 | 2.5478 | 99.57 | 20.88 | 37.35 | 80.3574 | 100.8578 | 75.3566 | 0.009271 | 1.013 | 0.024204 | 23.8925 | 0.020263 | 0.15635 | 0.003789 | 0.26937 | 0.26937 | 0.023856
2019-03-21 11:55 | 1553183691.383 | 578 | 0.20888 | 0.001049 | 9.8e-05 | 101.1899 | 100.9221 | 120.4438 | 21.095 | 0.002434 | 0.002251 | 162.0528 | 0.26329 | 8.4044 | 0.003197 | 0.004314 | 0.004249 | 5.7e-05 | 0.000902 | 0.29911 | 1.3394 | 2.5477 | 99.57 | 20.88 | 37.349 | 80.2253 | 98.5958 | 75.8254 | 0.009271 | 1.0117 | 0.024146 | 23.8672 | 0.023706 | 0.15663 | 0.003207 | 0.22659 | 0.22659 | 0.028411
2019-03-21 11:55 | 1553183692.3825 | 579 | 0.20889 | 0.001048 | 9.9e-05 | 101.2161 | 100.9269 | 120.0032 | 21.0948 | 0.002396 | 0.002227 | 162.0233 | 0.26273 | 8.4025 | 0.003279 | 0.00478 | 0.004744 | 5.6e-05 | 0.000898 | 0.29662 | 1.3095 | 2.5478 | 99.57 | 20.89 | 37.35 | 79.2934 | 100.5801 | 75.2399 | 0.009277 | 1.0134 | 0.024122 | 23.8035 | 0.0256 | 0.1563 | 0.003141 | 0.22155 | 0.22155 | 0.028995
2019-03-21 11:55 | 1553183693.3832 | 580 | 0.20873 | 0.001045 | 9.9e-05 | 101.2036 | 100.9585 | 119.9035 | 21.0951 | 0.002409 | 0.002237 | 162.0011 | 0.26314 | 8.4007 | 0.003219 | 0.004367 | 0.00419 | 5.8e-05 | 0.000898 | 0.3131 | 1.3164 | 2.548 | 99.57 | 20.89 | 37.35 | 79.6032 | 99.1274 | 75.4927 | 0.009277 | 1.0126 | 0.024103 | 23.8022 | 0.021688 | 0.15654 | 0.004139 | 0.29362 | 0.29362 | 0.021913
2019-03-21 11:55 | 1553183694.3832 | 581 | 0.20866 | 0.001044 | 0.0001 | 101.1402 | 100.9701 | 119.7634 | 21.095 | 0.00247 | 0.002238 | 161.9699 | 0.26271 | 8.4035 | 0.003177 | 0.004844 | 0.004703 | 5.3e-05 | 0.000892 | 0.32211 | 1.3408 | 2.5479 | 99.57 | 20.89 | 37.35 | 81.0744 | 98.1089 | 75.4979 | 0.009277 | 1.0126 | 0.024089 | 23.7885 | 0.015051 | 0.15629 | 0.004539 | 0.32227 | 0.32227 | 0.019932
2019-03-21 11:55 | 1553183695.3833 | 582 | 0.20879 | 0.001043 | 0.0001 | 101.1056 | 100.9429 | 119.5937 | 21.095 | 0.002397 | 0.002251 | 161.9298 | 0.2632 | 8.4039 | 0.003218 | 0.004738 | 0.004788 | 5.9e-05 | 0.000894 | 0.29894 | 1.3464 | 2.548 | 99.57 | 20.89 | 37.35 | 79.3248 | 99.0944 | 75.8215 | 0.009277 | 1.0117 | 0.02406 | 23.7827 | 0.014408 | 0.15658 | 0.00377 | 0.26634 | 0.26634 | 0.024163
2019-03-21 11:55 | 1553183696.3833 | 583 | 0.20877 | 0.001042 | 0.000101 | 101.1521 | 100.8983 | 119.3751 | 21.0955 | 0.002375 | 0.002237 | 161.9046 | 0.26375 | 8.4018 | 0.003222 | 0.004377 | 0.004242 | 5.3e-05 | 0.000894 | 0.30541 | 1.2928 | 2.5478 | 99.57 | 20.89 | 37.349 | 78.8051 | 99.2063 | 75.4736 | 0.009276 | 1.0127 | 0.02405 | 23.7491 | 0.022474 | 0.1569 | 0.003858 | 0.27251 | 0.27251 | 0.023664
2019-03-21 11:55 | 1553183697.3838 | 584 | 0.20879 | 0.00104 | 0.000102 | 101.1387 | 100.8914 | 119.358 | 21.0966 | 0.00243 | 0.002204 | 161.865 | 0.26244 | 8.3996 | 0.003222 | 0.004977 | 0.004865 | 5.6e-05 | 0.000894 | 0.30757 | 1.3036 | 2.5478 | 99.57 | 20.89 | 37.349 | 80.1285 | 99.1972 | 74.6873 | 0.009276 | 1.015 | 0.024076 | 23.7206 | 0.021892 | 0.15613 | 0.003791 | 0.26792 | 0.26792 | 0.02395
2019-03-21 11:55 | 1553183698.3835 | 585 | 0.20877 | 0.001038 | 0.000101 | 101.091 | 100.7855 | 119.1561 | 21.0956 | 0.002408 | 0.002324 | 161.8487 | 0.26351 | 8.3992 | 0.003293 | 0.004414 | 0.004283 | 5.5e-05 | 0.000893 | 0.30596 | 1.3787 | 2.5479 | 99.57 | 20.89 | 37.349 | 79.5967 | 100.9026 | 77.5651 | 0.009276 | 1.0066 | 0.023956 | 23.7983 | 0.02707 | 0.15676 | 0.003883 | 0.27326 | 0.27326 | 0.023579
2019-03-21 11:55 | 1553183699.3825 | 586 | 0.20877 | 0.001042 | 0.000101 | 101.0636 | 100.8008 | 119.339 | 21.096 | 0.002419 | 0.002242 | 161.8462 | 0.26363 | 8.402 | 0.003298 | 0.005175 | 0.005154 | 5.6e-05 | 0.000893 | 0.33682 | 1.3373 | 2.5477 | 99.57 | 20.89 | 37.349 | 79.8511 | 101.0453 | 75.6003 | 0.009276 | 1.0123 | 0.024042 | 23.7498 | 0.023285 | 0.15684 | 0.003863 | 0.27284 | 0.27284 | 0.023625
2019-03-21 11:55 | 1553183700.382 | 587 | 0.20891 | 0.001038 | 0.000101 | 101.0317 | 100.5972 | 119.1749 | 21.0952 | 0.002416 | 0.002285 | 161.8251 | 0.2635 | 8.4008 | 0.003282 | 0.004882 | 0.004778 | 5.6e-05 | 0.000891 | 0.31686 | 1.3222 | 2.5481 | 99.57 | 20.89 | 37.349 | 79.7873 | 100.6473 | 76.6339 | 0.009276 | 1.0093 | 0.02399 | 23.7686 | 0.038554 | 0.15676 | 0.003048 | 0.21354 | 0.21354 | 0.030172
2019-03-21 11:55 | 1553183701.3831 | 588 | 0.20879 | 0.001035 | 0.000102 | 101.0493 | 100.7889 | 119.589 | 21.0963 | 0.002409 | 0.002254 | 161.8147 | 0.26249 | 8.4012 | 0.003326 | 0.004701 | 0.004659 | 5.6e-05 | 0.000892 | 0.31741 | 1.3074 | 2.5479 | 99.57 | 20.89 | 37.349 | 79.6138 | 101.7152 | 75.8961 | 0.009276 | 1.0115 | 0.024057 | 23.7847 | 0.023077 | 0.15615 | 0.003779 | 0.26683 | 0.26683 | 0.024052
2019-03-21 11:55 | 1553183702.3827 | 589 | 0.20876 | 0.001036 | 0.000102 | 101.004 | 100.6864 | 119.7577 | 21.0965 | 0.002445 | 0.002245 | 161.7973 | 0.26274 | 8.4017 | 0.003302 | 0.00444 | 0.004259 | 5.6e-05 | 0.000889 | 0.3115 | 1.2958 | 2.5478 | 99.57 | 20.89 | 37.348 | 80.4786 | 101.1329 | 75.6671 | 0.009276 | 1.0121 | 0.024082 | 23.7937 | 0.02817 | 0.1563 | 0.003929 | 0.27798 | 0.27798 | 0.02311
2019-03-21 11:55 | 1553183703.3821 | 590 | 0.20878 | 0.00104 | 0.000102 | 101.0141 | 100.6738 | 119.6638 | 21.0958 | 0.002445 | 0.002259 | 161.7739 | 0.26248 | 8.4016 | 0.003338 | 0.005409 | 0.005302 | 5.5e-05 | 0.000891 | 0.32244 | 1.2791 | 2.5479 | 99.57 | 20.89 | 37.348 | 80.4692 | 101.9949 | 76.0209 | 0.009276 | 1.0111 | 0.02406 | 23.7964 | 0.030186 | 0.15615 | 0.003807 | 0.26891 | 0.26891 | 0.023866
2019-03-21 11:55 | 1553183704.3832 | 591 | 0.20877 | 0.001036 | 0.000102 | 100.9522 | 100.5028 | 119.566 | 21.0965 | 0.002399 | 0.002266 | 161.7556 | 0.26346 | 8.4042 | 0.003261 | 0.004963 | 0.004927 | 5.4e-05 | 0.000888 | 0.31758 | 1.2584 | 2.548 | 99.57 | 20.89 | 37.347 | 79.3714 | 100.1479 | 76.1783 | 0.009276 | 1.0106 | 0.024045 | 23.7921 | 0.039901 | 0.15673 | 0.003866 | 0.27296 | 0.27296 | 0.023599
2019-03-21 11:55 | 1553183705.3818 | 592 | 0.20869 | 0.001035 | 0.000103 | 100.976 | 100.4674 | 119.603 | 21.0963 | 0.002472 | 0.00225 | 161.7447 | 0.26359 | 8.4004 | 0.003275 | 0.004302 | 0.004269 | 5.9e-05 | 0.000884 | 0.32621 | 1.2772 | 2.5478 | 99.57 | 20.89 | 37.347 | 81.12 | 100.4773 | 75.7845 | 0.009276 | 1.0118 | 0.024063 | 23.7823 | 0.045165 | 0.15681 | 0.004403 | 0.31195 | 0.31195 | 0.02066
2019-03-21 11:55 | 1553183706.3828 | 593 | 0.20879 | 0.001034 | 0.000102 | 100.9274 | 100.4173 | 119.4042 | 21.0959 | 0.002435 | 0.0022 | 161.7331 | 0.26253 | 8.4006 | 0.00331 | 0.004772 | 0.004667 | 5.7e-05 | 0.000885 | 0.31686 | 1.3076 | 2.5478 | 99.57 | 20.89 | 37.346 | 80.2307 | 101.3269 | 74.5954 | 0.009276 | 1.0152 | 0.024084 | 23.722 | 0.045321 | 0.15618 | 0.003774 | 0.26678 | 0.26678 | 0.024061
2019-03-21 11:55 | 1553183707.3833 | 594 | 0.20892 | 0.001036 | 0.000102 | 100.925 | 100.6077 | 119.5739 | 21.0957 | 0.002392 | 0.002193 | 161.7073 | 0.26368 | 8.4027 | 0.003296 | 0.004724 | 0.004649 | 5.5e-05 | 0.000882 | 0.33356 | 1.3102 | 2.548 | 99.57 | 20.89 | 37.346 | 79.1974 | 100.9806 | 74.4316 | 0.009276 | 1.0157 | 0.024106 | 23.7333 | 0.028167 | 0.15686 | 0.00301 | 0.21171 | 0.21171 | 0.030452
2019-03-21 11:55 | 1553183708.383 | 595 | 0.20872 | 0.001039 | 0.000103 | 100.9039 | 100.556 | 119.7211 | 21.0971 | 0.002405 | 0.002268 | 161.7001 | 0.26363 | 8.404 | 0.00327 | 0.005205 | 0.005075 | 5.8e-05 | 0.000882 | 0.32477 | 1.2817 | 2.548 | 99.57 | 20.89 | 37.345 | 79.5177 | 100.3618 | 76.2295 | 0.009275 | 1.0105 | 0.024059 | 23.8092 | 0.030896 | 0.15684 | 0.004218 | 0.29851 | 0.29851 | 0.021593
2019-03-21 11:55 | 1553183709.383 | 596 | 0.20891 | 0.001028 | 0.000103 | 100.8351 | 100.6043 | 119.7293 | 21.0968 | 0.00246 | 0.002296 | 161.6736 | 0.2638 | 8.4045 | 0.003332 | 0.005556 | 0.005426 | 5.7e-05 | 0.00088 | 0.3282 | 1.3101 | 2.548 | 99.57 | 20.89 | 37.344 | 80.8516 | 101.8626 | 76.909 | 0.009275 | 1.0085 | 0.024037 | 23.8332 | 0.020492 | 0.15693 | 0.00309 | 0.2169 | 0.2169 | 0.029738
2019-03-21 11:55 | 1553183710.3821 | 597 | 0.20871 | 0.001034 | 0.000103 | 100.8455 | 100.4725 | 119.7704 | 21.0963 | 0.002434 | 0.002213 | 161.6607 | 0.26388 | 8.4036 | 0.00334 | 0.004831 | 0.004785 | 6e-05 | 0.000884 | 0.30574 | 1.3444 | 2.5479 | 99.57 | 20.89 | 37.343 | 80.2126 | 102.056 | 74.9122 | 0.009275 | 1.0143 | 0.024109 | 23.7692 | 0.033143 | 0.15699 | 0.004276 | 0.30333 | 0.30333 | 0.021271
2019-03-21 11:55 | 1553183711.3834 | 598 | 0.20886 | 0.00103 | 0.000103 | 100.8839 | 100.4916 | 119.8246 | 21.0967 | 0.002419 | 0.002205 | 161.6201 | 0.26353 | 8.4011 | 0.003325 | 0.004586 | 0.004532 | 6e-05 | 0.000884 | 0.31238 | 1.3125 | 2.5479 | 99.57 | 20.89 | 37.343 | 79.8559 | 101.6967 | 74.7094 | 0.009275 | 1.0149 | 0.024122 | 23.7676 | 0.034854 | 0.15677 | 0.003336 | 0.23542 | 0.23542 | 0.027369
2019-03-21 11:55 | 1553183712.3822 | 599 | 0.20879 | 0.001027 | 0.000103 | 100.8145 | 100.5059 | 119.703 | 21.0969 | 0.002382 | 0.00222 | 161.6028 | 0.26242 | 8.4027 | 0.003244 | 0.004735 | 0.004645 | 5.8e-05 | 0.000889 | 0.32228 | 1.3146 | 2.5479 | 99.57 | 20.89 | 37.342 | 78.9633 | 99.7379 | 75.0836 | 0.009275 | 1.0138 | 0.024097 | 23.7684 | 0.027423 | 0.15611 | 0.003807 | 0.26921 | 0.26921 | 0.023833
2019-03-21 11:55 | 1553183713.3825 | 600 | 0.20872 | 0.001024 | 0.000104 | 100.8065 | 100.4153 | 119.9178 | 21.0966 | 0.002374 | 0.002246 | 161.5878 | 0.26271 | 8.4004 | 0.003297 | 0.004823 | 0.004897 | 5.6e-05 | 0.000886 | 0.29679 | 1.3074 | 2.5479 | 99.57 | 20.89 | 37.342 | 78.7773 | 101.0096 | 75.6904 | 0.009275 | 1.0121 | 0.024097 | 23.8104 | 0.034776 | 0.15629 | 0.004215 | 0.29868 | 0.29868 | 0.021505
2019-03-21 11:55 | 1553183714.4159 | 601 | 0.20888 | 0.001028 | 0.000105 | 100.7359 | 100.6929 | 120.0507 | 21.0967 | 0.002356 | 0.002247 | 161.572 | 0.26349 | 8.4016 | 0.003265 | 0.004396 | 0.004218 | 5.7e-05 | 0.000885 | 0.29878 | 1.2953 | 2.548 | 99.57 | 20.89 | 37.342 | 78.3387 | 100.2439 | 75.7225 | 0.009275 | 1.012 | 0.02411 | 23.8247 | 0.003817 | 0.15675 | 0.003259 | 0.22957 | 0.22957 | 0.028064
2019-03-21 11:55 | 1553183715.3662 | 602 | 0.20883 | 0.001026 | 0.000106 | 100.7284 | 100.6793 | 120.0014 | 21.0963 | 0.00245 | 0.002254 | 161.5371 | 0.26296 | 8.4033 | 0.003294 | 0.004993 | 0.004869 | 5.5e-05 | 0.000885 | 0.31222 | 1.296 | 2.5479 | 99.57 | 20.89 | 37.342 | 80.6 | 100.9341 | 75.8905 | 0.009275 | 1.0115 | 0.024099 | 23.8255 | 0.004366 | 0.15644 | 0.003527 | 0.24882 | 0.24882 | 0.02584
2019-03-21 11:55 | 1553183716.3656 | 603 | 0.20878 | 0.001025 | 0.000107 | 100.7075 | 100.4626 | 120.2262 | 21.0968 | 0.00241 | 0.00226 | 161.5049 | 0.26353 | 8.4043 | 0.003363 | 0.0053 | 0.00522 | 5.6e-05 | 0.000887 | 0.29302 | 1.2723 | 2.5479 | 99.57 | 20.89 | 37.341 | 79.6431 | 102.6087 | 76.0306 | 0.009274 | 1.0111 | 0.024117 | 23.8526 | 0.021779 | 0.15678 | 0.003826 | 0.27062 | 0.27062 | 0.023811
2019-03-21 11:55 | 1553183717.3667 | 604 | 0.20875 | 0.001026 | 0.000107 | 100.7126 | 100.2331 | 120.4642 | 21.0973 | 0.002388 | 0.00227 | 161.4709 | 0.26374 | 8.4023 | 0.003357 | 0.005353 | 0.005287 | 6.2e-05 | 0.000887 | 0.30934 | 1.3078 | 2.5478 | 99.57 | 20.89 | 37.339 | 79.1041 | 102.4663 | 76.2704 | 0.009274 | 1.0104 | 0.024132 | 23.8844 | 0.04269 | 0.1569 | 0.00405 | 0.28706 | 0.28706 | 0.022464
2019-03-21 11:55 | 1553183718.3663 | 605 | 0.20881 | 0.001022 | 0.000108 | 100.6179 | 100.1149 | 120.5795 | 21.0976 | 0.002444 | 0.002261 | 161.4456 | 0.26381 | 8.4008 | 0.003396 | 0.004938 | 0.00489 | 5.7e-05 | 0.000886 | 0.28351 | 1.2596 | 2.5479 | 99.57 | 20.89 | 37.339 | 80.4654 | 103.4171 | 76.0532 | 0.009274 | 1.011 | 0.024151 | 23.8884 | 0.044824 | 0.15694 | 0.003681 | 0.26044 | 0.26044 | 0.024767
2019-03-21 11:55 | 1553183719.3663 | 606 | 0.20879 | 0.001024 | 0.000108 | 100.6772 | 100.1878 | 120.4272 | 21.0973 | 0.002389 | 0.002285 | 161.4207 | 0.26332 | 8.4038 | 0.003387 | 0.005368 | 0.005305 | 6.2e-05 | 0.000886 | 0.29336 | 1.297 | 2.5478 | 99.57 | 20.89 | 37.338 | 79.1469 | 103.1938 | 76.6392 | 0.009274 | 1.0093 | 0.024116 | 23.8933 | 0.043584 | 0.15665 | 0.003773 | 0.26668 | 0.26668 | 0.024142
2019-03-21 11:55 | 1553183720.3658 | 607 | 0.20878 | 0.001025 | 0.000108 | 100.5982 | 100.2307 | 120.1746 | 21.0978 | 0.002454 | 0.002287 | 161.4064 | 0.26387 | 8.4016 | 0.003351 | 0.004693 | 0.0046 | 6.2e-05 | 0.000886 | 0.33876 | 1.3237 | 2.548 | 99.57 | 20.89 | 37.338 | 80.688 | 102.323 | 76.6863 | 0.009274 | 1.0092 | 0.024089 | 23.8699 | 0.032736 | 0.15697 | 0.003857 | 0.27247 | 0.27247 | 0.023678
2019-03-21 11:55 | 1553183721.3666 | 608 | 0.20882 | 0.001023 | 0.000108 | 100.5213 | 100.1803 | 119.98 | 21.0977 | 0.002385 | 0.002286 | 161.4014 | 0.26327 | 8.4022 | 0.003399 | 0.004537 | 0.004479 | 5.7e-05 | 0.000884 | 0.30596 | 1.3428 | 2.548 | 99.57 | 20.89 | 37.337 | 79.0489 | 103.4848 | 76.658 | 0.009273 | 1.0093 | 0.02407 | 23.8496 | 0.030398 | 0.15662 | 0.003636 | 0.25635 | 0.25635 | 0.02511
2019-03-21 11:55 | 1553183722.3665 | 609 | 0.2088 | 0.001019 | 0.000107 | 100.6057 | 100.2365 | 120.0408 | 21.0976 | 0.00244 | 0.002273 | 161.4029 | 0.26241 | 8.4032 | 0.003333 | 0.004451 | 0.004381 | 5.5e-05 | 0.00088 | 0.29679 | 1.3235 | 2.548 | 99.57 | 20.89 | 37.335 | 80.3706 | 101.876 | 76.3454 | 0.009273 | 1.0102 | 0.024087 | 23.845 | 0.032881 | 0.15611 | 0.003722 | 0.26273 | 0.26273 | 0.024421
2019-03-21 11:55 | 1553183723.3656 | 610 | 0.20886 | 0.001021 | 0.000107 | 100.5833 | 100.2253 | 119.9873 | 21.0975 | 0.002431 | 0.002236 | 161.3971 | 0.26361 | 8.4003 | 0.003392 | 0.00467 | 0.004614 | 5.9e-05 | 0.000881 | 0.30773 | 1.325 | 2.5479 | 99.57 | 20.89 | 37.335 | 80.1488 | 103.322 | 75.4551 | 0.009273 | 1.0127 | 0.024113 | 23.8093 | 0.031898 | 0.15682 | 0.0034 | 0.23969 | 0.23969 | 0.026891
2019-03-21 11:55 | 1553183724.3673 | 611 | 0.20877 | 0.001017 | 0.000107 | 100.5103 | 100.0645 | 120.0298 | 21.0977 | 0.002434 | 0.00227 | 161.3746 | 0.26362 | 8.4023 | 0.003379 | 0.004814 | 0.004745 | 5.6e-05 | 0.00088 | 0.27472 | 1.286 | 2.548 | 99.57 | 20.88 | 37.335 | 80.2084 | 102.9866 | 76.2763 | 0.009267 | 1.0104 | 0.024088 | 23.8415 | 0.039762 | 0.15683 | 0.003927 | 0.27765 | 0.27765 | 0.023215
2019-03-21 11:55 | 1553183725.3657 | 612 | 0.20889 | 0.001023 | 0.000106 | 100.5651 | 100.0407 | 120.5403 | 21.0972 | 0.002367 | 0.002295 | 161.3654 | 0.26316 | 8.4016 | 0.003362 | 0.005267 | 0.005227 | 5.9e-05 | 0.000885 | 0.34274 | 1.3279 | 2.5479 | 99.56 | 20.89 | 37.333 | 78.598 | 102.5912 | 76.8673 | 0.009273 | 1.0087 | 0.024119 | 23.9123 | 0.046771 | 0.15655 | 0.003191 | 0.22468 | 0.22468 | 0.028638
2019-03-21 11:55 | 1553183726.3668 | 613 | 0.20879 | 0.001023 | 0.000106 | 100.4903 | 99.806 | 120.4455 | 21.0976 | 0.002442 | 0.002245 | 161.3608 | 0.26334 | 8.4025 | 0.003388 | 0.00486 | 0.004833 | 5.8e-05 | 0.000886 | 0.29463 | 1.2757 | 2.5479 | 99.56 | 20.89 | 37.333 | 80.4025 | 103.213 | 75.6674 | 0.009273 | 1.0121 | 0.024151 | 23.8619 | 0.061123 | 0.15666 | 0.003804 | 0.26948 | 0.26948 | 0.023893
2019-03-21 11:55 | 1553183727.3668 | 614 | 0.20881 | 0.00102 | 0.000106 | 100.4859 | 99.9967 | 120.2922 | 21.0983 | 0.002466 | 0.002218 | 161.3659 | 0.26264 | 8.4021 | 0.003325 | 0.005602 | 0.005484 | 6e-05 | 0.000885 | 0.30862 | 1.3235 | 2.548 | 99.56 | 20.89 | 37.334 | 80.9841 | 101.6866 | 75.0287 | 0.009274 | 1.014 | 0.024158 | 23.8249 | 0.043648 | 0.15625 | 0.003692 | 0.26139 | 0.26139 | 0.024567
2019-03-21 11:55 | 1553183728.3673 | 615 | 0.20883 | 0.001019 | 0.000105 | 100.4144 | 99.8949 | 120.3239 | 21.0983 | 0.002394 | 0.002292 | 161.3331 | 0.2634 | 8.4002 | 0.003311 | 0.004676 | 0.00452 | 5.7e-05 | 0.000887 | 0.30182 | 1.3234 | 2.548 | 99.56 | 20.89 | 37.335 | 79.2617 | 101.3411 | 76.8076 | 0.009274 | 1.0088 | 0.0241 | 23.8888 | 0.046395 | 0.1567 | 0.003545 | 0.25018 | 0.25018 | 0.025743
2019-03-21 11:55 | 1553183729.367 | 616 | 0.20874 | 0.001022 | 0.000105 | 100.4245 | 99.7968 | 120.4661 | 21.0995 | 0.002373 | 0.002285 | 161.3424 | 0.26327 | 8.4031 | 0.003403 | 0.004624 | 0.00467 | 5.9e-05 | 0.000884 | 0.3147 | 1.2798 | 2.548 | 99.56 | 20.89 | 37.335 | 78.7436 | 103.5779 | 76.6326 | 0.009274 | 1.0093 | 0.02412 | 23.897 | 0.056084 | 0.15662 | 0.004079 | 0.2891 | 0.2891 | 0.022266
2019-03-21 11:55 | 1553183730.3659 | 617 | 0.20886 | 0.001015 | 0.000106 | 100.4692 | 100.1445 | 120.6631 | 21.0991 | 0.002407 | 0.002327 | 161.3111 | 0.26238 | 8.4013 | 0.003272 | 0.004899 | 0.00469 | 5.7e-05 | 0.000884 | 0.32101 | 1.3019 | 2.5479 | 99.56 | 20.89 | 37.336 | 79.579 | 100.4008 | 77.6519 | 0.009274 | 1.0064 | 0.024104 | 23.9513 | 0.028951 | 0.15609 | 0.003373 | 0.23776 | 0.23776 | 0.026982
2019-03-21 11:55 | 1553183731.3661 | 618 | 0.20875 | 0.001018 | 0.000105 | 100.4758 | 100.0039 | 120.79 | 21.0988 | 0.002407 | 0.002285 | 161.289 | 0.2627 | 8.4003 | 0.003252 | 0.004574 | 0.00447 | 5.8e-05 | 0.000883 | 0.32029 | 1.341 | 2.5481 | 99.56 | 20.89 | 37.336 | 79.581 | 99.917 | 76.6281 | 0.009274 | 1.0093 | 0.024152 | 23.9289 | 0.042107 | 0.15628 | 0.004048 | 0.28724 | 0.28724 | 0.022361
2019-03-21 11:56 | 1553183732.3651 | 619 | 0.20882 | 0.001014 | 0.000105 | 100.4471 | 99.8554 | 120.856 | 21.1002 | 0.002492 | 0.002319 | 161.2821 | 0.26276 | 8.405 | 0.003283 | 0.004824 | 0.004773 | 6e-05 | 0.000886 | 0.31581 | 1.2995 | 2.5479 | 99.56 | 20.89 | 37.335 | 81.6137 | 100.6851 | 77.4588 | 0.009274 | 1.0069 | 0.02413 | 23.9639 | 0.052852 | 0.15632 | 0.003623 | 0.25619 | 0.25619 | 0.025078
2019-03-21 11:56 | 1553183733.3652 | 620 | 0.20879 | 0.001021 | 0.000104 | 100.4579 | 99.8332 | 120.6558 | 21.0992 | 0.002445 | 0.002295 | 161.2618 | 0.26246 | 8.4 | 0.003301 | 0.004619 | 0.004534 | 5.8e-05 | 0.000885 | 0.32477 | 1.3133 | 2.5479 | 99.56 | 20.89 | 37.335 | 80.4844 | 101.1034 | 76.8887 | 0.009274 | 1.0086 | 0.02413 | 23.9245 | 0.05579 | 0.15614 | 0.003799 | 0.26894 | 0.26894 | 0.023862
2019-03-21 11:56 | 1553183734.3666 | 621 | 0.20872 | 0.001026 | 0.000104 | 100.4248 | 99.8915 | 120.3692 | 21.0994 | 0.002447 | 0.002258 | 161.2348 | 0.26317 | 8.4028 | 0.0033 | 0.004682 | 0.004569 | 5.9e-05 | 0.000889 | 0.34739 | 1.3471 | 2.548 | 99.56 | 20.89 | 37.335 | 80.543 | 101.0879 | 75.9995 | 0.009274 | 1.0112 | 0.024132 | 23.8657 | 0.047628 | 0.15656 | 0.004231 | 0.30029 | 0.30029 | 0.021428
2019-03-21 11:56 | 1553183735.3658 | 622 | 0.20884 | 0.001026 | 0.000105 | 100.3923 | 99.6108 | 120.1328 | 21.1 | 0.002414 | 0.002284 | 161.2256 | 0.26279 | 8.4012 | 0.003281 | 0.005102 | 0.005002 | 5.9e-05 | 0.000891 | 0.34612 | 1.3283 | 2.5481 | 99.56 | 20.89 | 37.333 | 79.7403 | 100.6285 | 76.6074 | 0.009273 | 1.0094 | 0.024087 | 23.863 | 0.0699 | 0.15633 | 0.003509 | 0.24752 | 0.24752 | 0.025958
2019-03-21 11:56 | 1553183736.3653 | 623 | 0.20873 | 0.001021 | 0.000105 | 100.3867 | 99.914 | 120.4267 | 21.0999 | 0.002469 | 0.002208 | 161.1894 | 0.26334 | 8.4015 | 0.003304 | 0.003856 | 0.003709 | 6.3e-05 | 0.000893 | 0.32029 | 1.3192 | 2.548 | 99.56 | 20.89 | 37.333 | 81.0614 | 101.1912 | 74.7889 | 0.009273 | 1.0147 | 0.02418 | 23.83 | 0.042219 | 0.15666 | 0.004186 | 0.29757 | 0.29757 | 0.021638
2019-03-21 11:56 | 1553183737.3653 | 624 | 0.20875 | 0.001025 | 0.000106 | 100.4293 | 99.5822 | 120.5633 | 21.0999 | 0.002463 | 0.002287 | 161.1542 | 0.26247 | 8.3999 | 0.003276 | 0.004928 | 0.004899 | 5.8e-05 | 0.000893 | 0.2995 | 1.279 | 2.5481 | 99.56 | 20.89 | 37.333 | 80.9239 | 100.4973 | 76.6845 | 0.009273 | 1.0092 | 0.024128 | 23.9084 | 0.075761 | 0.15615 | 0.004051 | 0.28705 | 0.28705 | 0.022357
2019-03-21 11:56 | 1553183738.3658 | 625 | 0.20882 | 0.001027 | 0.000107 | 100.3905 | 99.5662 | 120.5067 | 21.1008 | 0.002431 | 0.002254 | 161.1447 | 0.26274 | 8.4022 | 0.003312 | 0.004554 | 0.004561 | 5.8e-05 | 0.000894 | 0.33539 | 1.295 | 2.548 | 99.56 | 20.89 | 37.334 | 80.1467 | 101.3725 | 75.8934 | 0.009274 | 1.0115 | 0.024149 | 23.8757 | 0.073748 | 0.1563 | 0.003606 | 0.25503 | 0.25503 | 0.025189
2019-03-21 11:56 | 1553183739.3655 | 626 | 0.20867 | 0.001022 | 0.000108 | 100.4349 | 99.5703 | 120.4121 | 21.1003 | 0.002418 | 0.002285 | 161.1133 | 0.26344 | 8.4016 | 0.003411 | 0.004802 | 0.004753 | 5.7e-05 | 0.000897 | 0.32676 | 1.304 | 2.548 | 99.56 | 20.89 | 37.333 | 79.8397 | 103.7839 | 76.649 | 0.009273 | 1.0093 | 0.024114 | 23.8922 | 0.077336 | 0.15672 | 0.004539 | 0.32212 | 0.32212 | 0.019996
2019-03-21 11:56 | 1553183740.3655 | 627 | 0.20879 | 0.001024 | 0.00011 | 100.4388 | 99.6187 | 120.2586 | 21.1011 | 0.002405 | 0.002279 | 161.0889 | 0.26318 | 8.4007 | 0.003346 | 0.004782 | 0.004662 | 5.6e-05 | 0.000895 | 0.33268 | 1.2471 | 2.5479 | 99.56 | 20.89 | 37.334 | 79.5215 | 102.21 | 76.4859 | 0.009274 | 1.0098 | 0.024104 | 23.8714 | 0.073337 | 0.15656 | 0.003776 | 0.2667 | 0.2667 | 0.024128
2019-03-21 11:56 | 1553183741.367 | 628 | 0.20873 | 0.001032 | 0.000111 | 100.37 | 99.746 | 120.0453 | 21.1012 | 0.002427 | 0.002287 | 161.0546 | 0.26292 | 8.4031 | 0.003373 | 0.004524 | 0.004418 | 6.1e-05 | 0.000897 | 0.30143 | 1.2577 | 2.5478 | 99.56 | 20.89 | 37.334 | 80.0496 | 102.8486 | 76.685 | 0.009274 | 1.0092 | 0.024076 | 23.857 | 0.055776 | 0.15641 | 0.004139 | 0.29252 | 0.29252 | 0.021976
2019-03-21 11:56 | 1553183742.3657 | 629 | 0.20877 | 0.001023 | 0.000112 | 100.3292 | 99.745 | 119.9159 | 21.1006 | 0.002444 | 0.002232 | 161.027 | 0.26368 | 8.4032 | 0.003359 | 0.004696 | 0.004664 | 6e-05 | 0.000899 | 0.33019 | 1.2618 | 2.5478 | 99.56 | 20.89 | 37.335 | 80.4498 | 102.5196 | 75.3546 | 0.009274 | 1.013 | 0.024109 | 23.7987 | 0.052229 | 0.15687 | 0.003919 | 0.27698 | 0.27698 | 0.023276
2019-03-21 11:56 | 1553183743.3667 | 630 | 0.20881 | 0.001027 | 0.000113 | 100.345 | 99.8329 | 120.1054 | 21.1002 | 0.002447 | 0.002275 | 161.0211 | 0.2631 | 8.4003 | 0.003325 | 0.00453 | 0.004482 | 6e-05 | 0.000899 | 0.3366 | 1.2878 | 2.5479 | 99.56 | 20.89 | 37.335 | 80.5329 | 101.6875 | 76.4058 | 0.009274 | 1.01 | 0.024092 | 23.8534 | 0.045761 | 0.15652 | 0.003685 | 0.25975 | 0.25975 | 0.024766
2019-03-21 11:56 | 1553183744.3672 | 631 | 0.20877 | 0.001028 | 0.000114 | 100.2991 | 99.9027 | 120.5518 | 21.1004 | 0.002432 | 0.002276 | 161.0196 | 0.26321 | 8.4033 | 0.003253 | 0.005152 | 0.005018 | 5.9e-05 | 0.000899 | 0.3246 | 1.2877 | 2.5479 | 99.56 | 20.9 | 37.338 | 80.1611 | 99.9633 | 76.4255 | 0.00928 | 1.0099 | 0.024136 | 23.8984 | 0.035416 | 0.15659 | 0.003901 | 0.27582 | 0.27582 | 0.023333
2019-03-21 11:56 | 1553183745.3668 | 632 | 0.20872 | 0.001033 | 0.000115 | 100.3447 | 99.808 | 120.2828 | 21.1009 | 0.002413 | 0.002299 | 160.9902 | 0.26236 | 8.4022 | 0.003295 | 0.004626 | 0.004597 | 6e-05 | 0.000899 | 0.29142 | 1.3316 | 2.548 | 99.56 | 20.9 | 37.339 | 79.7155 | 100.9715 | 76.9778 | 0.009281 | 1.0083 | 0.02409 | 23.8906 | 0.047962 | 0.15608 | 0.004205 | 0.29727 | 0.29727 | 0.021579
2019-03-21 11:56 | 1553183746.3658 | 633 | 0.20877 | 0.001036 | 0.000115 | 100.3643 | 99.6844 | 120.1521 | 21.1009 | 0.002459 | 0.002226 | 160.9812 | 0.26317 | 8.4013 | 0.003266 | 0.005244 | 0.005136 | 6e-05 | 0.000899 | 0.31205 | 1.3868 | 2.5481 | 99.56 | 20.9 | 37.341 | 80.8236 | 100.2668 | 75.2254 | 0.009281 | 1.0134 | 0.024137 | 23.8177 | 0.060801 | 0.15656 | 0.00388 | 0.27425 | 0.27425 | 0.023463
2019-03-21 11:56 | 1553183747.3662 | 634 | 0.20878 | 0.001039 | 0.000115 | 100.3244 | 99.8056 | 119.8737 | 21.1007 | 0.002418 | 0.002264 | 160.9577 | 0.2632 | 8.4014 | 0.003233 | 0.005128 | 0.005067 | 5.9e-05 | 0.000899 | 0.33196 | 1.3776 | 2.548 | 99.56 | 20.9 | 37.345 | 79.8418 | 99.4718 | 76.1303 | 0.009282 | 1.0108 | 0.024078 | 23.821 | 0.04637 | 0.15657 | 0.003803 | 0.26801 | 0.26801 | 0.024011
2019-03-21 11:56 | 1553183748.366 | 635 | 0.2087 | 0.001033 | 0.000115 | 100.3023 | 99.6609 | 119.6791 | 21.1013 | 0.002407 | 0.002282 | 160.9317 | 0.26264 | 8.4028 | 0.003264 | 0.004852 | 0.004785 | 6e-05 | 0.000895 | 0.32764 | 1.3129 | 2.5479 | 99.56 | 20.9 | 37.346 | 79.5718 | 100.2232 | 76.5698 | 0.009282 | 1.0095 | 0.024043 | 23.8166 | 0.057386 | 0.15625 | 0.004304 | 0.30377 | 0.30377 | 0.02114
2019-03-21 11:56 | 1553183749.3651 | 636 | 0.20876 | 0.001035 | 0.000115 | 100.3247 | 99.7525 | 119.9544 | 21.1018 | 0.002431 | 0.002263 | 160.8794 | 0.26325 | 8.4029 | 0.003275 | 0.004465 | 0.004332 | 5.5e-05 | 0.000893 | 0.30525 | 1.3124 | 2.548 | 99.56 | 20.9 | 37.348 | 80.1436 | 100.4873 | 76.1024 | 0.009283 | 1.0109 | 0.024087 | 23.8281 | 0.051159 | 0.15661 | 0.003968 | 0.28007 | 0.28007 | 0.022982
2019-03-21 11:56 | 1553183750.3651 | 637 | 0.20874 | 0.00104 | 0.000113 | 100.2836 | 99.6864 | 120.307 | 21.1025 | 0.00247 | 0.002257 | 160.8498 | 0.26408 | 8.4038 | 0.003227 | 0.004654 | 0.004678 | 5.9e-05 | 0.000889 | 0.28279 | 1.2843 | 2.5479 | 99.56 | 20.9 | 37.349 | 81.0771 | 99.3246 | 75.9784 | 0.009283 | 1.0112 | 0.024127 | 23.8588 | 0.053419 | 0.1571 | 0.004045 | 0.28616 | 0.28616 | 0.022564
2019-03-21 11:56 | 1553183751.3656 | 638 | 0.20878 | 0.001043 | 0.000113 | 100.2863 | 99.7576 | 120.6213 | 21.1015 | 0.00239 | 0.002253 | 160.8475 | 0.2632 | 8.405 | 0.003176 | 0.003966 | 0.003774 | 6e-05 | 0.000889 | 0.32781 | 1.264 | 2.5481 | 99.56 | 20.9 | 37.35 | 79.1571 | 98.0831 | 75.883 | 0.009283 | 1.0115 | 0.024161 | 23.8867 | 0.047274 | 0.15658 | 0.003814 | 0.26988 | 0.26988 | 0.023845
2019-03-21 11:56 | 1553183752.3656 | 639 | 0.20874 | 0.001036 | 0.000114 | 100.2809 | 99.586 | 121.0377 | 21.1025 | 0.002385 | 0.002255 | 160.8134 | 0.26266 | 8.4038 | 0.003191 | 0.004531 | 0.004494 | 6e-05 | 0.000892 | 0.33124 | 1.301 | 2.5478 | 99.56 | 20.9 | 37.349 | 79.0397 | 98.4478 | 75.9324 | 0.009283 | 1.0114 | 0.024201 | 23.9296 | 0.0622 | 0.15626 | 0.004105 | 0.29142 | 0.29142 | 0.022038
2019-03-21 11:56 | 1553183753.3666 | 640 | 0.2087 | 0.001032 | 0.000115 | 100.1921 | 99.7927 | 121.1097 | 21.1019 | 0.002398 | 0.002319 | 160.7354 | 0.2632 | 8.4012 | 0.003249 | 0.004058 | 0.004014 | 6.1e-05 | 0.000886 | 0.29374 | 1.2959 | 2.548 | 99.56 | 20.9 | 37.349 | 79.3555 | 99.8651 | 77.4612 | 0.009283 | 1.0069 | 0.024156 | 23.9891 | 0.035728 | 0.15657 | 0.004325 | 0.30673 | 0.30673 | 0.02098
2019-03-21 11:56 | 1553183754.3653 | 641 | 0.20875 | 0.001034 | 0.000115 | 100.1552 | 99.5672 | 120.8504 | 21.1024 | 0.002461 | 0.002252 | 160.6922 | 0.2631 | 8.401 | 0.003254 | 0.004763 | 0.00465 | 6.1e-05 | 0.00089 | 0.3402 | 1.2973 | 2.5479 | 99.56 | 20.9 | 37.349 | 80.8592 | 99.9701 | 75.8544 | 0.009283 | 1.0116 | 0.024185 | 23.9084 | 0.052661 | 0.15652 | 0.004022 | 0.28517 | 0.28517 | 0.022558
2019-03-21 11:56 | 1553183755.3653 | 642 | 0.20875 | 0.001036 | 0.000115 | 100.0926 | 99.3949 | 120.9555 | 21.1025 | 0.002427 | 0.002282 | 160.6728 | 0.26255 | 8.4033 | 0.003228 | 0.004251 | 0.004158 | 6.2e-05 | 0.000888 | 0.31332 | 1.27 | 2.548 | 99.56 | 20.9 | 37.349 | 80.0431 | 99.3536 | 76.5678 | 0.009283 | 1.0095 | 0.024171 | 23.9432 | 0.062566 | 0.15619 | 0.004034 | 0.28582 | 0.28582 | 0.02246
2019-03-21 11:56 | 1553183756.3664 | 643 | 0.20878 | 0.001039 | 0.000115 | 99.9943 | 99.289 | 120.7415 | 21.1022 | 0.002373 | 0.002241 | 160.6343 | 0.26377 | 8.4019 | 0.003196 | 0.004832 | 0.00469 | 6.2e-05 | 0.00089 | 0.28351 | 1.2744 | 2.5479 | 99.56 | 20.9 | 37.348 | 78.7508 | 98.5697 | 75.5736 | 0.009283 | 1.0124 | 0.024184 | 23.888 | 0.063313 | 0.15692 | 0.003812 | 0.26989 | 0.26989 | 0.023896
2019-03-21 11:56 | 1553183757.3664 | 644 | 0.20874 | 0.001034 | 0.000115 | 99.9889 | 99.3714 | 120.7316 | 21.1027 | 0.002378 | 0.002257 | 160.6116 | 0.26341 | 8.4027 | 0.003224 | 0.004343 | 0.004241 | 6.1e-05 | 0.000888 | 0.31924 | 1.3248 | 2.5479 | 99.56 | 20.9 | 37.345 | 78.87 | 99.2511 | 75.977 | 0.009282 | 1.0112 | 0.024169 | 23.9009 | 0.055411 | 0.1567 | 0.004057 | 0.28744 | 0.28744 | 0.022406
2019-03-21 11:56 | 1553183758.3669 | 645 | 0.20873 | 0.001039 | 0.000116 | 99.9293 | 99.1122 | 120.8067 | 21.103 | 0.002465 | 0.002257 | 160.5928 | 0.26306 | 8.4037 | 0.003205 | 0.004354 | 0.004271 | 6e-05 | 0.000888 | 0.26936 | 1.2951 | 2.5479 | 99.56 | 20.9 | 37.344 | 80.9574 | 98.8028 | 75.9727 | 0.009282 | 1.0112 | 0.024177 | 23.9081 | 0.073439 | 0.15649 | 0.004153 | 0.29446 | 0.29446 | 0.021843
2019-03-21 11:56 | 1553183759.3656 | 646 | 0.20872 | 0.00104 | 0.000116 | 99.8909 | 99.0461 | 120.7379 | 21.1022 | 0.002438 | 0.002207 | 160.5589 | 0.26297 | 8.4039 | 0.003112 | 0.004966 | 0.004894 | 6.4e-05 | 0.000889 | 0.30127 | 1.3067 | 2.5479 | 99.56 | 20.9 | 37.343 | 80.3125 | 96.5487 | 74.7769 | 0.009282 | 1.0147 | 0.024211 | 23.8604 | 0.075963 | 0.15644 | 0.004197 | 0.29811 | 0.29811 | 0.021568
2019-03-21 11:56 | 1553183760.367 | 647 | 0.20873 | 0.001036 | 0.000115 | 99.8191 | 98.9521 | 120.4948 | 21.1019 | 0.002406 | 0.002276 | 160.5346 | 0.26262 | 8.4029 | 0.003208 | 0.004749 | 0.004698 | 5.9e-05 | 0.00089 | 0.31454 | 1.3054 | 2.5481 | 99.56 | 20.91 | 37.34 | 79.549 | 98.8621 | 76.419 | 0.009287 | 1.0099 | 0.02413 | 23.8925 | 0.078025 | 0.15623 | 0.004123 | 0.29177 | 0.29177 | 0.022007
2019-03-21 11:56 | 1553183761.3671 | 648 | 0.20875 | 0.001039 | 0.000116 | 99.811 | 99.0948 | 120.4898 | 21.1031 | 0.00244 | 0.002275 | 160.5161 | 0.26286 | 8.4044 | 0.003126 | 0.00545 | 0.005317 | 5.9e-05 | 0.000886 | 0.29789 | 1.2977 | 2.5478 | 99.56 | 20.91 | 37.337 | 80.3586 | 96.884 | 76.4093 | 0.009286 | 1.01 | 0.02413 | 23.8917 | 0.064413 | 0.15637 | 0.003981 | 0.28145 | 0.28145 | 0.022835
2019-03-21 11:56 | 1553183762.3656 | 649 | 0.20873 | 0.001036 | 0.000116 | 99.6972 | 98.8179 | 120.3631 | 21.1025 | 0.00245 | 0.002302 | 160.4779 | 0.26381 | 8.4012 | 0.003145 | 0.00532 | 0.005302 | 6e-05 | 0.000888 | 0.30093 | 1.2957 | 2.5479 | 99.56 | 20.91 | 37.334 | 80.6169 | 97.3324 | 77.0433 | 0.009285 | 1.0081 | 0.024095 | 23.9008 | 0.079236 | 0.15694 | 0.004119 | 0.29107 | 0.29107 | 0.02216
2019-03-21 11:56 | 1553183763.3648 | 650 | 0.20873 | 0.001039 | 0.000115 | 99.6886 | 98.8939 | 120.4351 | 21.1036 | 0.002487 | 0.002281 | 160.4771 | 0.26293 | 8.4035 | 0.00314 | 0.005009 | 0.004892 | 6e-05 | 0.000888 | 0.33859 | 1.3077 | 2.548 | 99.56 | 20.91 | 37.333 | 81.493 | 97.2213 | 76.5345 | 0.009285 | 1.0096 | 0.02412 | 23.8905 | 0.071589 | 0.15642 | 0.004127 | 0.29195 | 0.29195 | 0.02202
2019-03-21 11:56 | 1553183764.3662 | 651 | 0.20873 | 0.001039 | 0.000115 | 99.665 | 99.055 | 120.5987 | 21.1039 | 0.002439 | 0.002243 | 160.4754 | 0.26367 | 8.4032 | 0.003106 | 0.004713 | 0.004638 | 5.9e-05 | 0.000889 | 0.33754 | 1.3309 | 2.548 | 99.56 | 20.91 | 37.332 | 80.3416 | 96.3975 | 75.6447 | 0.009285 | 1.0122 | 0.024167 | 23.8763 | 0.054918 | 0.15686 | 0.00412 | 0.29199 | 0.29199 | 0.022079
2019-03-21 11:56 | 1553183765.3649 | 652 | 0.20879 | 0.001043 | 0.000116 | 99.5301 | 98.8469 | 120.7149 | 21.1033 | 0.002446 | 0.002232 | 160.4431 | 0.26363 | 8.4019 | 0.003114 | 0.005206 | 0.005074 | 5.6e-05 | 0.00089 | 0.3246 | 1.2892 | 2.5481 | 99.56 | 20.91 | 37.328 | 80.5084 | 96.5895 | 75.3759 | 0.009284 | 1.013 | 0.024188 | 23.8786 | 0.061608 | 0.15684 | 0.003787 | 0.26802 | 0.26802 | 0.02405
2019-03-21 11:56 | 1553183766.3649 | 653 | 0.20873 | 0.001038 | 0.000117 | 99.5223 | 98.8459 | 120.9342 | 21.1034 | 0.002415 | 0.002279 | 160.4042 | 0.26297 | 8.4015 | 0.003176 | 0.004768 | 0.004724 | 5.7e-05 | 0.000887 | 0.3126 | 1.3046 | 2.548 | 99.56 | 20.91 | 37.326 | 79.7694 | 98.085 | 76.4985 | 0.009283 | 1.0097 | 0.024171 | 23.9388 | 0.061004 | 0.15644 | 0.004121 | 0.29204 | 0.29204 | 0.022017
2019-03-21 11:56 | 1553183767.3664 | 654 | 0.20873 | 0.001036 | 0.000116 | 99.5176 | 98.7771 | 121.0804 | 21.1036 | 0.002391 | 0.002301 | 160.3819 | 0.26311 | 8.4012 | 0.003133 | 0.004817 | 0.004692 | 5.6e-05 | 0.000888 | 0.33516 | 1.3097 | 2.548 | 99.56 | 20.91 | 37.323 | 79.1988 | 97.046 | 77.02 | 0.009282 | 1.0082 | 0.024168 | 23.9711 | 0.066804 | 0.15652 | 0.004126 | 0.29241 | 0.29241 | 0.022001
2019-03-21 11:56 | 1553183768.366 | 655 | 0.20866 | 0.001039 | 0.000117 | 99.4598 | 98.6024 | 120.9248 | 21.1036 | 0.002433 | 0.002247 | 160.3385 | 0.26249 | 8.4002 | 0.003156 | 0.004648 | 0.004596 | 5.9e-05 | 0.000889 | 0.31669 | 1.3266 | 2.548 | 99.56 | 20.91 | 37.322 | 80.1916 | 97.6117 | 75.7402 | 0.009282 | 1.0119 | 0.024197 | 23.9119 | 0.077442 | 0.15616 | 0.004555 | 0.32389 | 0.32389 | 0.019815
2019-03-21 11:56 | 1553183769.3651 | 656 | 0.20873 | 0.001037 | 0.000116 | 99.3716 | 98.3953 | 120.5736 | 21.104 | 0.00248 | 0.002264 | 160.3489 | 0.26336 | 8.4005 | 0.003091 | 0.004876 | 0.0049 | 5.6e-05 | 0.000886 | 0.34324 | 1.2731 | 2.548 | 99.56 | 20.91 | 37.321 | 81.3249 | 96.0482 | 76.1372 | 0.009282 | 1.0108 | 0.024148 | 23.8907 | 0.088312 | 0.15667 | 0.004148 | 0.29379 | 0.29379 | 0.021917
2019-03-21 11:56 | 1553183770.3655 | 657 | 0.20878 | 0.001043 | 0.000115 | 99.3153 | 98.3033 | 120.5189 | 21.1043 | 0.002466 | 0.002249 | 160.3211 | 0.2633 | 8.4037 | 0.003161 | 0.005262 | 0.005076 | 5.5e-05 | 0.00088 | 0.30558 | 1.2843 | 2.5479 | 99.56 | 20.91 | 37.317 | 80.986 | 97.722 | 75.7767 | 0.009281 | 1.0118 | 0.024155 | 23.8729 | 0.091605 | 0.15664 | 0.003811 | 0.26946 | 0.26946 | 0.023891
2019-03-21 11:56 | 1553183771.3666 | 658 | 0.20871 | 0.001042 | 0.000114 | 99.2283 | 98.1654 | 120.1647 | 21.1038 | 0.002427 | 0.002288 | 160.2952 | 0.26312 | 8.4024 | 0.003134 | 0.004878 | 0.004838 | 5.4e-05 | 0.000884 | 0.28622 | 1.307 | 2.5481 | 99.56 | 20.91 | 37.316 | 80.0483 | 97.0791 | 76.7219 | 0.009281 | 1.0091 | 0.024087 | 23.8701 | 0.096326 | 0.15653 | 0.004233 | 0.29925 | 0.29925 | 0.021499
2019-03-21 11:56 | 1553183772.3652 | 659 | 0.20868 | 0.001045 | 0.000114 | 99.1243 | 98.1463 | 120.1309 | 21.1037 | 0.002455 | 0.002301 | 160.2563 | 0.26234 | 8.4018 | 0.003187 | 0.004525 | 0.004376 | 5.4e-05 | 0.000884 | 0.2896 | 1.3345 | 2.548 | 99.56 | 20.91 | 37.314 | 80.7389 | 98.3542 | 77.0178 | 0.00928 | 1.0082 | 0.024073 | 23.8768 | 0.088686 | 0.15606 | 0.00442 | 0.3126 | 0.3126 | 0.020519
2019-03-21 11:56 | 1553183773.3654 | 660 | 0.20871 | 0.001041 | 0.000113 | 99.099 | 98.3081 | 119.95 | 21.1043 | 0.002466 | 0.002274 | 160.2419 | 0.26296 | 8.4013 | 0.003162 | 0.004795 | 0.00479 | 5.5e-05 | 0.000884 | 0.29894 | 1.3852 | 2.548 | 99.56 | 20.91 | 37.311 | 80.9806 | 97.7448 | 76.3879 | 0.009279 | 1.01 | 0.024077 | 23.8374 | 0.071669 | 0.15644 | 0.004243 | 0.29994 | 0.29994 | 0.021436
2019-03-21 11:56 | 1553183774.3653 | 661 | 0.20876 | 0.001047 | 0.000113 | 99.0578 | 98.1078 | 120.138 | 21.1048 | 0.002446 | 0.00227 | 160.1902 | 0.26307 | 8.4025 | 0.003174 | 0.004621 | 0.004541 | 5.7e-05 | 0.000883 | 0.3282 | 1.3684 | 2.5478 | 99.56 | 20.91 | 37.309 | 80.5138 | 98.0378 | 76.2791 | 0.009279 | 1.0103 | 0.024099 | 23.8523 | 0.086197 | 0.1565 | 0.003936 | 0.278 | 0.278 | 0.023137
2019-03-21 11:56 | 1553183775.3665 | 662 | 0.20873 | 0.001046 | 0.000112 | 98.9342 | 97.7921 | 120.3572 | 21.1055 | 0.002516 | 0.002266 | 160.1433 | 0.26335 | 8.4024 | 0.003172 | 0.004568 | 0.004477 | 5.4e-05 | 0.000884 | 0.32659 | 1.3934 | 2.548 | 99.56 | 20.91 | 37.305 | 82.1849 | 97.9945 | 76.1879 | 0.009278 | 1.0106 | 0.024124 | 23.871 | 0.10386 | 0.15667 | 0.004149 | 0.29374 | 0.29374 | 0.021921
2019-03-21 11:56 | 1553183776.3654 | 663 | 0.20868 | 0.001046 | 0.000112 | 98.8141 | 97.915 | 120.2737 | 21.1051 | 0.002358 | 0.002231 | 160.1207 | 0.26323 | 8.4024 | 0.003133 | 0.004291 | 0.00432 | 5.5e-05 | 0.000881 | 0.32405 | 1.3584 | 2.548 | 99.56 | 20.91 | 37.304 | 78.3874 | 97.0458 | 75.3402 | 0.009278 | 1.0131 | 0.024145 | 23.8337 | 0.081756 | 0.1566 | 0.004426 | 0.31412 | 0.31412 | 0.020489
2019-03-21 11:56 | 1553183777.3665 | 664 | 0.20876 | 0.001045 | 0.000111 | 98.6872 | 97.5179 | 120.4856 | 21.1044 | 0.002458 | 0.002255 | 160.0868 | 0.26238 | 8.4027 | 0.003145 | 0.004971 | 0.004812 | 5.5e-05 | 0.00088 | 0.32117 | 1.2906 | 2.548 | 99.56 | 20.91 | 37.303 | 80.8075 | 97.353 | 75.932 | 0.009277 | 1.0114 | 0.024146 | 23.875 | 0.10661 | 0.15609 | 0.003967 | 0.28091 | 0.28091 | 0.022838
2019-03-21 11:56 | 1553183778.3666 | 665 | 0.20873 | 0.001046 | 0.000111 | 98.6184 | 97.4014 | 120.8065 | 21.1049 | 0.002422 | 0.002266 | 160.0892 | 0.26163 | 8.404 | 0.003127 | 0.004825 | 0.004733 | 5.7e-05 | 0.00088 | 0.33339 | 1.3529 | 2.5481 | 99.56 | 20.91 | 37.301 | 79.9402 | 96.9147 | 76.1819 | 0.009277 | 1.0106 | 0.02417 | 23.9153 | 0.11106 | 0.15564 | 0.004158 | 0.29508 | 0.29508 | 0.021679
2019-03-21 11:56 | 1553183779.3647 | 666 | 0.20866 | 0.001043 | 0.000111 | 98.4366 | 97.1521 | 121.0975 | 21.1047 | 0.002438 | 0.00227 | 160.0531 | 0.26346 | 8.4023 | 0.003094 | 0.004756 | 0.004691 | 5.7e-05 | 0.000877 | 0.29015 | 1.3666 | 2.548 | 99.56 | 20.91 | 37.301 | 80.3236 | 96.1038 | 76.2884 | 0.009277 | 1.0103 | 0.024195 | 23.9477 | 0.11749 | 0.15673 | 0.004558 | 0.32442 | 0.32442 | 0.019856
2019-03-21 11:56 | 1553183780.3657 | 667 | 0.20875 | 0.001051 | 0.000111 | 98.3824 | 97.2611 | 121.3487 | 21.1051 | 0.002493 | 0.002252 | 160.052 | 0.26312 | 8.3999 | 0.003158 | 0.004554 | 0.004448 | 5.6e-05 | 0.000878 | 0.28816 | 1.3352 | 2.5481 | 99.56 | 20.91 | 37.298 | 81.6411 | 97.654 | 75.8608 | 0.009276 | 1.0116 | 0.024235 | 23.9579 | 0.10253 | 0.15653 | 0.004031 | 0.28663 | 0.28663 | 0.022444
2019-03-21 11:56 | 1553183781.3661 | 668 | 0.20864 | 0.00105 | 0.00011 | 98.2564 | 96.8395 | 121.2825 | 21.1046 | 0.002472 | 0.002279 | 160.0085 | 0.26299 | 8.403 | 0.003074 | 0.004548 | 0.004451 | 5.5e-05 | 0.000875 | 0.3079 | 1.3228 | 2.5481 | 99.56 | 20.91 | 37.297 | 81.143 | 95.6409 | 76.4924 | 0.009276 | 1.0097 | 0.024206 | 23.973 | 0.12992 | 0.15645 | 0.00466 | 0.33199 | 0.33199 | 0.019369
2019-03-21 11:56 | 1553183782.3648 | 669 | 0.20869 | 0.001053 | 0.000112 | 98.114 | 96.6876 | 121.0327 | 21.1047 | 0.002449 | 0.002312 | 159.994 | 0.26247 | 8.4031 | 0.003109 | 0.004549 | 0.004443 | 5.5e-05 | 0.000876 | 0.34379 | 1.3501 | 2.548 | 99.56 | 20.91 | 37.297 | 80.5926 | 96.4705 | 77.2862 | 0.009276 | 1.0074 | 0.024154 | 23.9755 | 0.13099 | 0.15615 | 0.004377 | 0.31073 | 0.31073 | 0.020653
2019-03-21 11:56 | 1553183783.3649 | 670 | 0.20872 | 0.001055 | 0.000112 | 98.0502 | 96.5915 | 120.707 | 21.1049 | 0.002435 | 0.00226 | 159.9576 | 0.26249 | 8.4012 | 0.003147 | 0.004077 | 0.003962 | 5.7e-05 | 0.000879 | 0.36497 | 1.3482 | 2.5479 | 99.56 | 20.91 | 37.295 | 80.2513 | 97.3874 | 76.0368 | 0.009275 | 1.0111 | 0.024165 | 23.9005 | 0.13407 | 0.15616 | 0.004207 | 0.29852 | 0.29852 | 0.021499
2019-03-21 11:56 | 1553183784.3663 | 671 | 0.20863 | 0.001061 | 0.000112 | 97.8726 | 96.3524 | 120.3718 | 21.1046 | 0.002436 | 0.002244 | 159.9194 | 0.26353 | 8.4032 | 0.003077 | 0.004745 | 0.004741 | 5.8e-05 | 0.000875 | 0.3115 | 1.3343 | 2.5479 | 99.56 | 20.91 | 37.294 | 80.2603 | 95.695 | 75.6516 | 0.009275 | 1.0122 | 0.024144 | 23.8541 | 0.14002 | 0.15677 | 0.004719 | 0.33536 | 0.33536 | 0.019214
2019-03-21 11:56 | 1553183785.3663 | 672 | 0.20865 | 0.001061 | 0.000112 | 97.8556 | 96.0619 | 120.0363 | 21.1057 | 0.002398 | 0.002251 | 159.919 | 0.2631 | 8.4 | 0.003112 | 0.004379 | 0.004331 | 5.7e-05 | 0.000873 | 0.30237 | 1.338 | 2.5479 | 99.56 | 20.91 | 37.295 | 79.3662 | 96.549 | 75.8381 | 0.009275 | 1.0116 | 0.024104 | 23.8272 | 0.16548 | 0.15652 | 0.004587 | 0.32519 | 0.32519 | 0.019782
2019-03-21 11:56 | 1553183786.366 | 673 | 0.20864 | 0.001064 | 0.000112 | 97.7049 | 96.0221 | 120.2774 | 21.1058 | 0.002453 | 0.002339 | 159.9179 | 0.26264 | 8.4039 | 0.003186 | 0.003386 | 0.003291 | 5.7e-05 | 0.000868 | 0.29087 | 1.3349 | 2.548 | 99.56 | 20.91 | 37.296 | 80.6926 | 98.3335 | 77.9342 | 0.009276 | 1.0056 | 0.024056 | 23.9226 | 0.1554 | 0.15624 | 0.004667 | 0.33037 | 0.33037 | 0.019437
2019-03-21 11:56 | 1553183787.3665 | 674 | 0.20868 | 0.001068 | 0.000112 | 97.6116 | 95.8872 | 120.0831 | 21.1046 | 0.002476 | 0.002258 | 159.9094 | 0.26216 | 8.4042 | 0.003161 | 0.004606 | 0.004505 | 5.7e-05 | 0.000871 | 0.30735 | 1.2689 | 2.5479 | 99.56 | 20.91 | 37.295 | 81.2288 | 97.7254 | 76.003 | 0.009275 | 1.0111 | 0.024103 | 23.8375 | 0.15943 | 0.15596 | 0.004412 | 0.31255 | 0.31255 | 0.020508
2019-03-21 11:56 | 1553183788.3654 | 675 | 0.20859 | 0.001065 | 0.000112 | 97.519 | 95.9444 | 120.0831 | 21.105 | 0.002459 | 0.002274 | 159.8663 | 0.26328 | 8.4021 | 0.003108 | 0.004505 | 0.004425 | 5.7e-05 | 0.000872 | 0.33052 | 1.2981 | 2.548 | 99.56 | 20.91 | 37.296 | 80.8151 | 96.4606 | 76.3737 | 0.009276 | 1.0101 | 0.02409 | 23.8501 | 0.1456 | 0.15663 | 0.004943 | 0.35078 | 0.35078 | 0.018352
2019-03-21 11:56 | 1553183789.3662 | 676 | 0.20863 | 0.001068 | 0.000112 | 97.3417 | 95.7673 | 120.1429 | 21.1059 | 0.002467 | 0.002287 | 159.8243 | 0.26244 | 8.4017 | 0.003152 | 0.004695 | 0.004666 | 5.9e-05 | 0.00087 | 0.27887 | 1.267 | 2.5478 | 99.56 | 20.9 | 37.296 | 81.0196 | 97.5231 | 76.7003 | 0.00927 | 1.0091 | 0.024085 | 23.8672 | 0.14585 | 0.15613 | 0.004686 | 0.3321 | 0.3321 | 0.019322
2019-03-21 11:56 | 1553183790.3662 | 677 | 0.20862 | 0.001072 | 0.000112 | 97.1188 | 95.582 | 120.1828 | 21.1059 | 0.002487 | 0.0023 | 159.7765 | 0.26205 | 8.4016 | 0.003192 | 0.0043 | 0.004228 | 5.7e-05 | 0.000875 | 0.32587 | 1.2891 | 2.5479 | 99.56 | 20.9 | 37.296 | 81.5092 | 98.4731 | 77.0036 | 0.00927 | 1.0083 | 0.024079 | 23.8815 | 0.14267 | 0.15589 | 0.004779 | 0.33879 | 0.33879 | 0.018912
2019-03-21 11:56 | 1553183791.3656 | 678 | 0.20868 | 0.001073 | 0.000112 | 96.9862 | 95.1123 | 120.3074 | 21.1062 | 0.002522 | 0.002325 | 159.739 | 0.26296 | 8.4041 | 0.003143 | 0.004782 | 0.004655 | 5.8e-05 | 0.000874 | 0.30574 | 1.2877 | 2.548 | 99.56 | 20.9 | 37.299 | 82.3308 | 97.3031 | 77.6074 | 0.009271 | 1.0065 | 0.02407 | 23.9145 | 0.17451 | 0.15644 | 0.004407 | 0.31176 | 0.31176 | 0.020623
2019-03-21 11:57 | 1553183792.3663 | 679 | 0.20867 | 0.001077 | 0.000112 | 96.8578 | 95.1066 | 120.2408 | 21.1072 | 0.002425 | 0.002294 | 159.7131 | 0.26312 | 8.4021 | 0.003121 | 0.004259 | 0.004271 | 5.6e-05 | 0.000875 | 0.29927 | 1.3153 | 2.5481 | 99.56 | 20.9 | 37.301 | 80.0117 | 96.7603 | 76.8568 | 0.009271 | 1.0087 | 0.02409 | 23.8823 | 0.1632 | 0.15653 | 0.004446 | 0.31485 | 0.31485 | 0.020433
2019-03-21 11:57 | 1553183793.3653 | 680 | 0.20866 | 0.001073 | 0.000113 | 96.6889 | 94.8893 | 120.4987 | 21.1064 | 0.002417 | 0.002282 | 159.6858 | 0.2629 | 8.4034 | 0.00312 | 0.004413 | 0.004273 | 5.8e-05 | 0.000878 | 0.32803 | 1.3022 | 2.548 | 99.56 | 20.9 | 37.303 | 79.8223 | 96.7377 | 76.58 | 0.009272 | 1.0095 | 0.024125 | 23.8984 | 0.16805 | 0.1564 | 0.00453 | 0.32133 | 0.32133 | 0.020004
2019-03-21 11:57 | 1553183794.3645 | 681 | 0.20862 | 0.001079 | 0.000114 | 96.5468 | 94.6682 | 120.2891 | 21.1063 | 0.002451 | 0.00229 | 159.6642 | 0.26298 | 8.4015 | 0.003137 | 0.00454 | 0.004411 | 5.8e-05 | 0.000878 | 0.31852 | 1.3135 | 2.5479 | 99.56 | 20.9 | 37.305 | 80.6402 | 97.166 | 76.767 | 0.009272 | 1.0089 | 0.024098 | 23.884 | 0.17575 | 0.15645 | 0.004778 | 0.33886 | 0.33886 | 0.018976
2019-03-21 11:57 | 1553183795.3647 | 682 | 0.20864 | 0.00108 | 0.000115 | 96.3129 | 94.3872 | 120.3288 | 21.1059 | 0.002487 | 0.002279 | 159.6272 | 0.26327 | 8.4026 | 0.003211 | 0.004633 | 0.00459 | 5.6e-05 | 0.000874 | 0.31852 | 1.2978 | 2.548 | 99.56 | 20.9 | 37.306 | 81.5092 | 98.9514 | 76.5017 | 0.009272 | 1.0097 | 0.024111 | 23.8789 | 0.18065 | 0.15662 | 0.004619 | 0.32747 | 0.32747 | 0.019657
2019-03-21 11:57 | 1553183796.3655 | 683 | 0.20864 | 0.00108 | 0.000116 | 96.1407 | 93.9956 | 120.2645 | 21.106 | 0.002482 | 0.002358 | 159.5835 | 0.26335 | 8.4017 | 0.003192 | 0.004321 | 0.004075 | 5.8e-05 | 0.000877 | 0.29839 | 1.282 | 2.5479 | 99.56 | 20.9 | 37.31 | 81.3696 | 98.4748 | 78.4022 | 0.009273 | 1.0042 | 0.024039 | 23.9373 | 0.20183 | 0.15667 | 0.004629 | 0.32719 | 0.32719 | 0.01968
2019-03-21 11:57 | 1553183797.365 | 684 | 0.20859 | 0.001089 | 0.000116 | 95.9893 | 93.8076 | 120.0795 | 21.1064 | 0.002469 | 0.002258 | 159.5382 | 0.26275 | 8.4037 | 0.003225 | 0.004764 | 0.004687 | 6.2e-05 | 0.000877 | 0.323 | 1.2944 | 2.5479 | 99.56 | 20.9 | 37.313 | 81.0621 | 99.2687 | 75.9877 | 0.009274 | 1.0112 | 0.024103 | 23.8366 | 0.20565 | 0.15631 | 0.00493 | 0.34977 | 0.34977 | 0.018368
2019-03-21 11:57 | 1553183798.3646 | 685 | 0.20863 | 0.001085 | 0.000117 | 95.804 | 93.5552 | 119.6343 | 21.1055 | 0.002444 | 0.002318 | 159.4801 | 0.26301 | 8.4041 | 0.003223 | 0.004681 | 0.004613 | 5.7e-05 | 0.000875 | 0.31022 | 1.3307 | 2.5479 | 99.56 | 20.9 | 37.316 | 80.4529 | 99.2301 | 77.4359 | 0.009275 | 1.007 | 0.024009 | 23.8416 | 0.21246 | 0.15646 | 0.004699 | 0.33174 | 0.33174 | 0.019385
2019-03-21 11:57 | 1553183799.3657 | 686 | 0.20864 | 0.001088 | 0.000118 | 95.5257 | 93.1805 | 119.2183 | 21.1056 | 0.002454 | 0.002353 | 159.4624 | 0.26385 | 8.4026 | 0.003254 | 0.004617 | 0.004492 | 6.1e-05 | 0.000875 | 0.32659 | 1.3437 | 2.5479 | 99.56 | 20.9 | 37.322 | 80.6956 | 99.9779 | 78.2755 | 0.009276 | 1.0046 | 0.023938 | 23.8286 | 0.22233 | 0.15697 | 0.004597 | 0.3234 | 0.3234 | 0.019948
2019-03-21 11:57 | 1553183800.3658 | 687 | 0.20866 | 0.001091 | 0.000119 | 95.3696 | 92.9758 | 118.7841 | 21.106 | 0.002464 | 0.00231 | 159.4004 | 0.26318 | 8.4061 | 0.003214 | 0.004542 | 0.004515 | 6.1e-05 | 0.000872 | 0.28528 | 1.3204 | 2.5479 | 99.56 | 20.9 | 37.324 | 80.9463 | 99.0167 | 77.2571 | 0.009277 | 1.0075 | 0.02393 | 23.7507 | 0.22737 | 0.15656 | 0.004517 | 0.3174 | 0.3174 | 0.020273
2019-03-21 11:57 | 1553183801.3648 | 688 | 0.20857 | 0.001091 | 0.000122 | 95.1947 | 92.7654 | 118.4806 | 21.107 | 0.002462 | 0.002249 | 159.358 | 0.26309 | 8.403 | 0.00323 | 0.00454 | 0.004435 | 6e-05 | 0.000871 | 0.32997 | 1.3176 | 2.5479 | 99.55 | 20.9 | 37.328 | 80.9011 | 99.4129 | 75.787 | 0.009279 | 1.0118 | 0.023949 | 23.6706 | 0.23123 | 0.15651 | 0.00501 | 0.35296 | 0.35296 | 0.018225
2019-03-21 11:57 | 1553183802.3648 | 689 | 0.2084 | 0.001094 | 0.000126 | 94.9504 | 92.3526 | 118.6165 | 21.1068 | 0.002408 | 0.002315 | 159.312 | 0.26264 | 8.4036 | 0.003179 | 0.004484 | 0.004469 | 5.8e-05 | 0.000875 | 0.31045 | 1.3028 | 2.5481 | 99.55 | 20.9 | 37.329 | 79.6035 | 98.1627 | 77.3607 | 0.009279 | 1.0072 | 0.023909 | 23.7375 | 0.24813 | 0.15624 | 0.006045 | 0.42639 | 0.42639 | 0.01506
2019-03-21 11:57 | 1553183803.3659 | 690 | 0.20792 | 0.001095 | 0.000129 | 94.8106 | 92.061 | 118.5782 | 21.1067 | 0.00244 | 0.002305 | 159.2427 | 0.26313 | 8.4009 | 0.003195 | 0.004887 | 0.004813 | 6.1e-05 | 0.000868 | 0.28385 | 1.2896 | 2.5479 | 99.55 | 20.9 | 37.333 | 80.363 | 98.5608 | 77.1182 | 0.00928 | 1.0079 | 0.023914 | 23.7254 | 0.26324 | 0.15654 | 0.008954 | 0.63491 | 0.63491 | 0.010133
2019-03-21 11:57 | 1553183804.366 | 691 | 0.20757 | 0.001103 | 0.000132 | 94.5902 | 92.4953 | 118.5867 | 21.1066 | 0.002442 | 0.002342 | 159.1962 | 0.26391 | 8.4005 | 0.003177 | 0.005409 | 0.005328 | 6e-05 | 0.000871 | 0.29358 | 1.2855 | 2.5479 | 99.55 | 20.9 | 37.335 | 80.4081 | 98.1178 | 78.013 | 0.009281 | 1.0054 | 0.023884 | 23.7566 | 0.20032 | 0.157 | 0.011007 | 0.78093 | 0.78093 | 0.008263
2019-03-21 11:57 | 1553183805.365 | 692 | 0.20751 | 0.001136 | 0.000135 | 94.3113 | 92.45 | 118.0339 | 21.1064 | 0.002442 | 0.002339 | 159.1387 | 0.26239 | 8.4013 | 0.003147 | 0.004533 | 0.004437 | 5.7e-05 | 0.000871 | 0.28838 | 1.295 | 2.5478 | 99.55 | 20.9 | 37.338 | 80.4226 | 97.3982 | 77.9411 | 0.009281 | 1.0056 | 0.023831 | 23.6987 | 0.17829 | 0.1561 | 0.01136 | 0.80425 | 0.80425 | 0.007977
2019-03-21 11:57 | 1553183806.3661 | 693 | 0.20726 | 0.001173 | 0.000137 | 94.154 | 92.6022 | 117.58 | 21.1056 | 0.002497 | 0.002392 | 159.0561 | 0.26256 | 8.4015 | 0.003227 | 0.004679 | 0.004598 | 6e-05 | 0.00087 | 0.31332 | 1.3284 | 2.5479 | 99.55 | 20.9 | 37.341 | 81.7273 | 99.3346 | 79.2219 | 0.009282 | 1.0019 | 0.023741 | 23.6962 | 0.14864 | 0.1562 | 0.012788 | 0.90272 | 0.90272 | 0.007111
2019-03-21 11:57 | 1553183807.3651 | 694 | 0.20716 | 0.001232 | 0.000138 | 93.8337 | 92.6809 | 117.5438 | 21.1066 | 0.00251 | 0.00238 | 158.9621 | 0.26221 | 8.4011 | 0.003297 | 0.004822 | 0.00471 | 6e-05 | 0.00087 | 0.30414 | 1.3059 | 2.548 | 99.55 | 20.9 | 37.343 | 82.0595 | 101.0179 | 78.9391 | 0.009283 | 1.0027 | 0.023747 | 23.683 | 0.11057 | 0.15599 | 0.013306 | 0.93981 | 0.93981 | 0.006822
2019-03-21 11:57 | 1553183808.3652 | 695 | 0.20705 | 0.001301 | 0.000138 | 93.7086 | 92.5413 | 117.3664 | 21.1073 | 0.002526 | 0.002359 | 158.9183 | 0.26277 | 8.4017 | 0.00338 | 0.004692 | 0.004674 | 6.2e-05 | 0.000874 | 0.3147 | 1.2905 | 2.548 | 99.55 | 20.9 | 37.344 | 82.4379 | 103.0386 | 78.4222 | 0.009283 | 1.0042 | 0.023747 | 23.6478 | 0.11212 | 0.15632 | 0.013912 | 0.98287 | 0.98287 | 0.006537
2019-03-21 11:57 | 1553183809.3647 | 696 | 0.20672 | 0.001376 | 0.000138 | 93.4988 | 92.4524 | 116.9922 | 21.107 | 0.002482 | 0.002387 | 158.824 | 0.26305 | 8.404 | 0.003262 | 0.004148 | 0.003944 | 6e-05 | 0.000863 | 0.28152 | 1.3237 | 2.5479 | 99.55 | 20.9 | 37.345 | 81.3793 | 100.1826 | 79.0909 | 0.009283 | 1.0023 | 0.023686 | 23.6325 | 0.10067 | 0.15649 | 0.015826 | 1.1162 | 1.1162 | 0.005762
2019-03-21 11:57 | 1553183810.3643 | 697 | 0.20646 | 0.001473 | 0.000137 | 93.2786 | 92.1431 | 116.4047 | 21.1073 | 0.002542 | 0.002395 | 158.7442 | 0.26216 | 8.4017 | 0.003357 | 0.004619 | 0.004575 | 5.6e-05 | 0.000858 | 0.33179 | 1.3572 | 2.548 | 99.55 | 20.9 | 37.346 | 82.8146 | 102.4782 | 79.2774 | 0.009283 | 1.0018 | 0.023621 | 23.5793 | 0.10955 | 0.15596 | 0.017253 | 1.214 | 1.214 | 0.00528
2019-03-21 11:57 | 1553183811.3643 | 698 | 0.20616 | 0.001571 | 0.000136 | 93.0367 | 91.9905 | 115.9784 | 21.1073 | 0.002633 | 0.002429 | 158.6424 | 0.26175 | 8.402 | 0.003243 | 0.004624 | 0.0045 | 6e-05 | 0.000854 | 0.30685 | 1.3039 | 2.548 | 99.55 | 20.9 | 37.345 | 85.017 | 99.7159 | 80.1132 | 0.009283 | 0.99938 | 0.023549 | 23.564 | 0.10115 | 0.15572 | 0.018891 | 1.3259 | 1.3259 | 0.004827
2019-03-21 11:57 | 1553183812.3654 | 699 | 0.20593 | 0.001674 | 0.000136 | 92.8135 | 91.652 | 115.3858 | 21.1071 | 0.002561 | 0.002449 | 158.5564 | 0.26204 | 8.4014 | 0.00326 | 0.004747 | 0.004707 | 5.9e-05 | 0.000849 | 0.32693 | 1.2953 | 2.5479 | 99.55 | 20.9 | 37.344 | 83.2871 | 100.1285 | 80.5823 | 0.009283 | 0.99806 | 0.023474 | 23.5193 | 0.11265 | 0.15589 | 0.020172 | 1.4116 | 1.4116 | 0.004539
2019-03-21 11:57 | 1553183813.3654 | 700 | 0.2058 | 0.001783 | 0.000137 | 92.6056 | 91.3999 | 114.8623 | 21.1069 | 0.002589 | 0.002428 | 158.4669 | 0.2621 | 8.4022 | 0.003348 | 0.00449 | 0.004397 | 5.8e-05 | 0.000853 | 0.30414 | 1.3185 | 2.5479 | 99.55 | 20.9 | 37.344 | 83.9601 | 102.2447 | 80.0878 | 0.009283 | 0.99945 | 0.023437 | 23.4495 | 0.11721 | 0.15593 | 0.020819 | 1.4547 | 1.4547 | 0.004406
2019-03-21 11:57 | 1553183814.3659 | 701 | 0.20558 | 0.001902 | 0.000138 | 92.3699 | 91.0468 | 114.5736 | 21.1068 | 0.002618 | 0.00251 | 158.3633 | 0.26166 | 8.4021 | 0.003402 | 0.005065 | 0.005008 | 5.7e-05 | 0.000853 | 0.28512 | 1.28 | 2.548 | 99.55 | 20.9 | 37.344 | 84.6397 | 103.5749 | 82.0434 | 0.009283 | 0.99395 | 0.023343 | 23.4848 | 0.12905 | 0.15566 | 0.021959 | 1.5283 | 1.5283 | 0.004186
2019-03-21 11:57 | 1553183815.366 | 702 | 0.20535 | 0.002033 | 0.000139 | 92.0428 | 90.5304 | 114.2178 | 21.1075 | 0.002626 | 0.002507 | 158.2584 | 0.26368 | 8.4048 | 0.003384 | 0.00526 | 0.005217 | 5.8e-05 | 0.000851 | 0.2724 | 1.311 | 2.5479 | 99.55 | 20.9 | 37.343 | 84.8535 | 103.1299 | 81.9694 | 0.009283 | 0.99416 | 0.023309 | 23.4458 | 0.14819 | 0.15687 | 0.023203 | 1.6128 | 1.6128 | 0.003998
2019-03-21 11:57 | 1553183816.3648 | 703 | 0.2051 | 0.002161 | 0.000142 | 91.7722 | 89.9143 | 113.4221 | 21.1074 | 0.002746 | 0.002568 | 158.1441 | 0.26293 | 8.3992 | 0.003373 | 0.004665 | 0.004544 | 5.5e-05 | 0.000844 | 0.3115 | 1.2803 | 2.5479 | 99.55 | 20.9 | 37.342 | 87.7244 | 102.8634 | 83.4458 | 0.009282 | 0.99004 | 0.023179 | 23.4125 | 0.18294 | 0.15642 | 0.024541 | 1.6964 | 1.6964 | 0.00379
2019-03-21 11:57 | 1553183817.3656 | 704 | 0.20467 | 0.002301 | 0.000144 | 91.4665 | 89.5369 | 112.3909 | 21.1073 | 0.0027 | 0.002612 | 158.0538 | 0.26194 | 8.4033 | 0.003442 | 0.004892 | 0.004833 | 5.7e-05 | 0.00084 | 0.29518 | 1.2758 | 2.5478 | 99.55 | 20.9 | 37.343 | 86.6233 | 104.5249 | 84.5067 | 0.009283 | 0.98711 | 0.02304 | 23.3405 | 0.19072 | 0.15583 | 0.02692 | 1.8498 | 1.8498 | 0.003462
2019-03-21 11:57 | 1553183818.3657 | 705 | 0.20423 | 0.002448 | 0.000147 | 91.14 | 89.306 | 111.5358 | 21.1069 | 0.00272 | 0.002644 | 157.9385 | 0.26245 | 8.4019 | 0.003437 | 0.005558 | 0.005403 | 5.6e-05 | 0.000838 | 0.32012 | 1.3054 | 2.5479 | 99.55 | 20.9 | 37.343 | 87.0991 | 104.4107 | 85.2796 | 0.009283 | 0.98498 | 0.022927 | 23.2766 | 0.18182 | 0.15613 | 0.029397 | 2.0102 | 2.0102 | 0.003192
2019-03-21 11:57 | 1553183819.3647 | 706 | 0.2038 | 0.002603 | 0.00015 | 90.9004 | 88.9583 | 110.5207 | 21.1073 | 0.002731 | 0.00262 | 157.8215 | 0.26294 | 8.4016 | 0.00345 | 0.00428 | 0.004195 | 5.9e-05 | 0.000839 | 0.32421 | 1.3278 | 2.548 | 99.55 | 20.9 | 37.343 | 87.3775 | 104.7365 | 84.7048 | 0.009283 | 0.98656 | 0.022841 | 23.1519 | 0.19318 | 0.15642 | 0.031723 | 2.1611 | 2.1611 | 0.002975
2019-03-21 11:57 | 1553183820.3662 | 707 | 0.20344 | 0.002776 | 0.000153 | 90.6412 | 88.6265 | 109.8303 | 21.1077 | 0.002802 | 0.002672 | 157.7099 | 0.26219 | 8.4003 | 0.003512 | 0.005508 | 0.005485 | 5.9e-05 | 0.000838 | 0.3115 | 1.353 | 2.5479 | 99.55 | 20.9 | 37.344 | 89.0721 | 106.2428 | 85.9525 | 0.009283 | 0.98313 | 0.02273 | 23.1197 | 0.20105 | 0.15598 | 0.033707 | 2.2849 | 2.2849 | 0.002806
2019-03-21 11:57 | 1553183821.3649 | 708 | 0.20301 | 0.002949 | 0.000156 | 90.1866 | 88.3003 | 109.0466 | 21.1079 | 0.002828 | 0.002745 | 157.6058 | 0.26164 | 8.4062 | 0.003487 | 0.004662 | 0.00455 | 5.5e-05 | 0.00084 | 0.32029 | 1.3868 | 2.5478 | 99.55 | 20.9 | 37.343 | 89.7065 | 105.6172 | 87.7037 | 0.009283 | 0.97836 | 0.022593 | 23.0931 | 0.18907 | 0.15565 | 0.036039 | 2.4282 | 2.4282 | 0.002634
2019-03-21 11:57 | 1553183822.3659 | 709 | 0.2026 | 0.003131 | 0.000159 | 90.048 | 88.1409 | 107.9967 | 21.1079 | 0.002848 | 0.002779 | 157.4669 | 0.2616 | 8.4031 | 0.00349 | 0.005486 | 0.005447 | 5.8e-05 | 0.000845 | 0.32388 | 1.3289 | 2.5479 | 99.55 | 20.9 | 37.344 | 90.1912 | 105.6975 | 88.5325 | 0.009283 | 0.97612 | 0.022459 | 23.0081 | 0.19148 | 0.15563 | 0.03827 | 2.5629 | 2.5629 | 0.002496
2019-03-21 11:57 | 1553183823.3649 | 710 | 0.20215 | 0.003319 | 0.000164 | 89.9059 | 87.7771 | 107.0071 | 21.1082 | 0.002925 | 0.002814 | 157.3411 | 0.26261 | 8.4017 | 0.003512 | 0.005022 | 0.00493 | 5.7e-05 | 0.000845 | 0.32012 | 1.3112 | 2.548 | 99.55 | 20.9 | 37.343 | 92.037 | 106.2384 | 89.3622 | 0.009283 | 0.97389 | 0.02233 | 22.9287 | 0.21434 | 0.15623 | 0.040725 | 2.7113 | 2.7113 | 0.002368
2019-03-21 11:57 | 1553183824.366 | 711 | 0.20143 | 0.00361 | 0.000168 | 89.8993 | 87.5258 | 105.7159 | 21.1087 | 0.002942 | 0.002883 | 157.1919 | 0.26272 | 8.4005 | 0.003504 | 0.005335 | 0.005158 | 5.9e-05 | 0.000856 | 0.32925 | 1.3625 | 2.548 | 99.55 | 20.9 | 37.345 | 92.4518 | 106.046 | 91.0336 | 0.009283 | 0.96942 | 0.022144 | 22.8424 | 0.23933 | 0.15629 | 0.044644 | 2.9469 | 2.9469 | 0.00218
2019-03-21 11:57 | 1553183825.366 | 712 | 0.20048 | 0.003877 | 0.000171 | 89.8588 | 87.5615 | 104.8537 | 21.109 | 0.002996 | 0.002913 | 157.0242 | 0.26189 | 8.405 | 0.003587 | 0.004666 | 0.004622 | 5.9e-05 | 0.000856 | 0.30901 | 1.3296 | 2.5479 | 99.55 | 20.9 | 37.346 | 93.7438 | 108.059 | 91.7602 | 0.009283 | 0.96749 | 0.022031 | 22.7717 | 0.23165 | 0.1558 | 0.049929 | 3.2781 | 3.2781 | 0.001953
2019-03-21 11:57 | 1553183826.3651 | 713 | 0.20006 | 0.004186 | 0.000174 | 89.8657 | 87.5472 | 104.0486 | 21.108 | 0.00305 | 0.002971 | 156.8617 | 0.26234 | 8.4009 | 0.003603 | 0.005038 | 0.004998 | 6.1e-05 | 0.000854 | 0.30198 | 1.2934 | 2.5478 | 99.55 | 20.9 | 37.347 | 95.0546 | 108.448 | 93.1575 | 0.009284 | 0.96379 | 0.021905 | 22.7275 | 0.23379 | 0.15606 | 0.052061 | 3.3979 | 3.3979 | 0.001888
2019-03-21 11:57 | 1553183827.3652 | 714 | 0.19954 | 0.004575 | 0.000177 | 89.9774 | 87.5063 | 103.2618 | 21.1088 | 0.00312 | 0.00309 | 156.7145 | 0.26236 | 8.4015 | 0.003582 | 0.005559 | 0.005414 | 6e-05 | 0.00086 | 0.32659 | 1.2843 | 2.5479 | 99.55 | 20.9 | 37.348 | 96.7352 | 107.9296 | 96.0131 | 0.009284 | 0.95634 | 0.021737 | 22.7295 | 0.24908 | 0.15608 | 0.054689 | 3.5415 | 3.5415 | 0.001811
2019-03-21 11:57 | 1553183828.3653 | 715 | 0.19895 | 0.004966 | 0.000179 | 90.146 | 87.8323 | 102.8396 | 21.109 | 0.003159 | 0.003083 | 156.5509 | 0.26188 | 8.4021 | 0.003566 | 0.00486 | 0.004774 | 6e-05 | 0.00086 | 0.32493 | 1.3005 | 2.5479 | 99.55 | 20.9 | 37.349 | 97.6892 | 107.554 | 95.8637 | 0.009284 | 0.95673 | 0.021697 | 22.6784 | 0.23258 | 0.15579 | 0.057737 | 3.7312 | 3.7312 | 0.001716
2019-03-21 11:57 | 1553183829.3643 | 716 | 0.19845 | 0.005383 | 0.000179 | 90.365 | 88.0673 | 102.5877 | 21.1084 | 0.003239 | 0.003196 | 156.377 | 0.26236 | 8.4019 | 0.003661 | 0.005175 | 0.005056 | 6e-05 | 0.000858 | 0.31758 | 1.3259 | 2.548 | 99.55 | 20.9 | 37.35 | 99.6128 | 109.8438 | 98.5758 | 0.009284 | 0.94975 | 0.021591 | 22.7337 | 0.23038 | 0.15608 | 0.060205 | 3.8712 | 3.8712 | 0.001657
2019-03-21 11:57 | 1553183830.3642 | 717 | 0.19778 | 0.005777 | 0.000178 | 90.5241 | 88.24 | 102.0007 | 21.1094 | 0.003218 | 0.00317 | 156.1916 | 0.26208 | 8.4032 | 0.003678 | 0.005409 | 0.005313 | 6.2e-05 | 0.000858 | 0.32676 | 1.3449 | 2.5478 | 99.55 | 20.9 | 37.35 | 99.1102 | 110.2658 | 97.9544 | 0.009284 | 0.95134 | 0.021547 | 22.6496 | 0.22859 | 0.15591 | 0.063678 | 4.0853 | 4.0853 | 0.001569
2019-03-21 11:57 | 1553183831.3656 | 718 | 0.19713 | 0.00616 | 0.000177 | 90.8176 | 88.4699 | 101.4263 | 21.1081 | 0.003395 | 0.003253 | 156.0011 | 0.26146 | 8.401 | 0.00373 | 0.005319 | 0.005219 | 6.1e-05 | 0.000862 | 0.33339 | 1.3857 | 2.548 | 99.55 | 20.9 | 37.35 | 103.3969 | 111.5382 | 99.9624 | 0.009284 | 0.94622 | 0.021429 | 22.6468 | 0.23427 | 0.15554 | 0.0671 | 4.2802 | 4.2802 | 0.001494
2019-03-21 11:57 | 1553183832.3648 | 719 | 0.19659 | 0.006553 | 0.000177 | 91.0199 | 88.6848 | 101.1263 | 21.1078 | 0.003367 | 0.003345 | 155.8073 | 0.26297 | 8.402 | 0.003799 | 0.004857 | 0.004809 | 5.8e-05 | 0.000861 | 0.31725 | 1.3426 | 2.5479 | 99.55 | 20.9 | 37.35 | 102.7059 | 113.198 | 102.1827 | 0.009284 | 0.94062 | 0.021334 | 22.6804 | 0.23247 | 0.15644 | 0.069821 | 4.4331 | 4.4331 | 0.00145
2019-03-21 11:57 | 1553183833.3655 | 720 | 0.19627 | 0.006912 | 0.000175 | 91.303 | 89.0935 | 100.7977 | 21.1086 | 0.003453 | 0.003346 | 155.6078 | 0.26227 | 8.4023 | 0.003809 | 0.005783 | 0.005696 | 6e-05 | 0.000861 | 0.33157 | 1.3178 | 2.5479 | 99.55 | 20.9 | 37.35 | 104.8054 | 113.442 | 102.1977 | 0.009284 | 0.94058 | 0.021299 | 22.644 | 0.21912 | 0.15603 | 0.071274 | 4.5175 | 4.5175 | 0.00142
2019-03-21 11:57 | 1553183834.366 | 721 | 0.19604 | 0.007292 | 0.000175 | 91.5461 | 89.386 | 100.886 | 21.1084 | 0.003475 | 0.003387 | 155.4278 | 0.2622 | 8.4014 | 0.003896 | 0.004968 | 0.004845 | 5.6e-05 | 0.00086 | 0.32261 | 1.306 | 2.548 | 99.55 | 20.9 | 37.351 | 105.3282 | 115.5649 | 103.199 | 0.009285 | 0.93808 | 0.02128 | 22.6841 | 0.21358 | 0.15598 | 0.072194 | 4.5714 | 4.5714 | 0.001402
2019-03-21 11:57 | 1553183835.366 | 722 | 0.19579 | 0.007539 | 0.000175 | 91.7326 | 89.4422 | 100.7876 | 21.108 | 0.003496 | 0.0034 | 155.2426 | 0.26317 | 8.4041 | 0.003855 | 0.005649 | 0.005511 | 5.8e-05 | 0.000864 | 0.33301 | 1.3373 | 2.5479 | 99.55 | 20.9 | 37.352 | 105.8431 | 114.5616 | 103.5123 | 0.009285 | 0.9373 | 0.02126 | 22.6825 | 0.22616 | 0.15656 | 0.0734 | 4.6431 | 4.6431 | 0.001386
2019-03-21 11:57 | 1553183836.3657 | 723 | 0.19567 | 0.007928 | 0.000178 | 91.9948 | 89.7958 | 101.0545 | 21.1095 | 0.003611 | 0.003429 | 155.1049 | 0.26251 | 8.4024 | 0.003849 | 0.004978 | 0.004888 | 5.4e-05 | 0.000862 | 0.31526 | 1.3289 | 2.5478 | 99.55 | 20.9 | 37.351 | 108.6475 | 114.4349 | 104.2133 | 0.009285 | 0.93556 | 0.021269 | 22.7336 | 0.2164 | 0.15616 | 0.073624 | 4.6589 | 4.6589 | 0.001378
2019-03-21 11:57 | 1553183837.3651 | 724 | 0.19568 | 0.008188 | 0.000182 | 92.2814 | 89.9882 | 101.5279 | 21.1082 | 0.003611 | 0.003442 | 154.9119 | 0.26247 | 8.4009 | 0.003876 | 0.005761 | 0.00569 | 5.7e-05 | 0.00086 | 0.30309 | 1.3133 | 2.5479 | 99.55 | 20.9 | 37.349 | 108.6375 | 115.082 | 104.5297 | 0.009284 | 0.93478 | 0.021309 | 22.7964 | 0.22508 | 0.15614 | 0.073261 | 4.6448 | 4.6448 | 0.001382
2019-03-21 11:57 | 1553183838.3657 | 725 | 0.19568 | 0.008424 | 0.000187 | 92.4524 | 90.275 | 101.8504 | 21.1082 | 0.003668 | 0.003424 | 154.7236 | 0.26224 | 8.4027 | 0.003831 | 0.005671 | 0.005559 | 5.6e-05 | 0.00086 | 0.30502 | 1.3075 | 2.5478 | 99.55 | 20.91 | 37.349 | 110.021 | 113.9807 | 104.1083 | 0.00929 | 0.93582 | 0.021355 | 22.8198 | 0.21318 | 0.15601 | 0.072975 | 4.6363 | 4.6363 | 0.001383
2019-03-21 11:57 | 1553183839.3647 | 726 | 0.19577 | 0.008658 | 0.000194 | 92.6723 | 90.3717 | 102.6504 | 21.1088 | 0.003699 | 0.003487 | 154.5469 | 0.26187 | 8.4035 | 0.0039 | 0.004913 | 0.004877 | 5.7e-05 | 0.000862 | 0.31708 | 1.309 | 2.548 | 99.55 | 20.9 | 37.346 | 110.7741 | 115.6648 | 105.6205 | 0.009283 | 0.93208 | 0.021396 | 22.9551 | 0.22485 | 0.15579 | 0.072162 | 4.5934 | 4.5934 | 0.001394
2019-03-21 11:57 | 1553183840.3649 | 727 | 0.19609 | 0.008881 | 0.0002 | 92.8895 | 90.3697 | 103.6308 | 21.1081 | 0.003714 | 0.003434 | 154.4099 | 0.26164 | 8.4022 | 0.003976 | 0.005109 | 0.004977 | 5.7e-05 | 0.000863 | 0.32493 | 1.3054 | 2.5481 | 99.55 | 20.9 | 37.344 | 111.1429 | 117.5342 | 104.3496 | 0.009283 | 0.93522 | 0.021534 | 23.0258 | 0.24599 | 0.15565 | 0.069987 | 4.4841 | 4.4841 | 0.001427
2019-03-21 11:57 | 1553183841.3649 | 728 | 0.19644 | 0.009074 | 0.000207 | 93.1246 | 90.3829 | 104.539 | 21.1095 | 0.003651 | 0.003503 | 154.279 | 0.26199 | 8.4019 | 0.003898 | 0.005563 | 0.005468 | 5.5e-05 | 0.000865 | 0.31526 | 1.2848 | 2.5479 | 99.55 | 20.9 | 37.339 | 109.6091 | 115.6111 | 106.01 | 0.009282 | 0.93112 | 0.021581 | 23.1772 | 0.26729 | 0.15585 | 0.067684 | 4.3463 | 4.3463 | 0.001474
2019-03-21 11:57 | 1553183842.365 | 729 | 0.1969 | 0.009232 | 0.000211 | 93.2673 | 90.6295 | 105.4405 | 21.1086 | 0.003721 | 0.00346 | 154.143 | 0.26228 | 8.4029 | 0.00393 | 0.005874 | 0.005788 | 5.5e-05 | 0.000863 | 0.32836 | 1.3341 | 2.548 | 99.55 | 20.9 | 37.337 | 111.3136 | 116.3905 | 104.9635 | 0.009281 | 0.9337 | 0.021704 | 23.2448 | 0.25661 | 0.15603 | 0.064803 | 4.1855 | 4.1855 | 0.001532
2019-03-21 11:57 | 1553183843.3644 | 730 | 0.19754 | 0.009352 | 0.000213 | 93.5024 | 90.5566 | 106.4301 | 21.1079 | 0.003738 | 0.003458 | 154.0307 | 0.26195 | 8.4027 | 0.003865 | 0.005555 | 0.005465 | 5.5e-05 | 0.00086 | 0.32012 | 1.3365 | 2.548 | 99.55 | 20.9 | 37.335 | 111.7203 | 114.8107 | 104.9134 | 0.009281 | 0.93383 | 0.021807 | 23.3521 | 0.28633 | 0.15583 | 0.060848 | 3.9495 | 3.9495 | 0.001622
2019-03-21 11:57 | 1553183844.3655 | 731 | 0.19834 | 0.009426 | 0.000212 | 93.6195 | 90.6156 | 107.2141 | 21.1085 | 0.003699 | 0.003401 | 153.902 | 0.26274 | 8.4018 | 0.003854 | 0.005365 | 0.005285 | 5.4e-05 | 0.00086 | 0.33445 | 1.3115 | 2.5479 | 99.55 | 20.9 | 37.332 | 110.777 | 114.5382 | 103.53 | 0.00928 | 0.93725 | 0.021927 | 23.395 | 0.2917 | 0.1563 | 0.055984 | 3.6549 | 3.6549 | 0.001758
2019-03-21 11:57 | 1553183845.3652 | 732 | 0.19919 | 0.009461 | 0.000209 | 93.7994 | 90.8141 | 108.2866 | 21.1101 | 0.003682 | 0.003441 | 153.8087 | 0.26156 | 8.402 | 0.00386 | 0.005309 | 0.005075 | 5.1e-05 | 0.000861 | 0.27223 | 1.3269 | 2.5479 | 99.55 | 20.9 | 37.33 | 110.3574 | 114.6877 | 104.501 | 0.009279 | 0.93485 | 0.022008 | 23.542 | 0.2893 | 0.1556 | 0.050894 | 3.3357 | 3.3357 | 0.001917
2019-03-21 11:57 | 1553183846.3652 | 733 | 0.20007 | 0.00943 | 0.000201 | 93.9463 | 90.9286 | 109.1013 | 21.1082 | 0.00367 | 0.003389 | 153.7063 | 0.26213 | 8.4011 | 0.003829 | 0.00486 | 0.004696 | 5.4e-05 | 0.000864 | 0.30038 | 1.3346 | 2.5479 | 99.55 | 20.9 | 37.326 | 110.0604 | 113.9497 | 103.2538 | 0.009278 | 0.93794 | 0.022127 | 23.5914 | 0.29203 | 0.15594 | 0.045713 | 3.0132 | 3.0132 | 0.002127
2019-03-21 11:57 | 1553183847.3646 | 734 | 0.20087 | 0.009356 | 0.000193 | 94.2136 | 91.2739 | 109.5982 | 21.1086 | 0.003703 | 0.003376 | 153.6157 | 0.26029 | 8.4053 | 0.003798 | 0.005305 | 0.005228 | 5.2e-05 | 0.000862 | 0.34346 | 1.3767 | 2.548 | 99.55 | 20.9 | 37.324 | 110.8601 | 113.1805 | 102.9329 | 0.009278 | 0.93874 | 0.022187 | 23.635 | 0.28353 | 0.15485 | 0.041013 | 2.7113 | 2.7113 | 0.002347
2019-03-21 11:57 | 1553183848.3644 | 735 | 0.20141 | 0.009244 | 0.000184 | 94.3268 | 91.6697 | 110.3448 | 21.1086 | 0.0037 | 0.00341 | 153.5406 | 0.26146 | 8.4022 | 0.003779 | 0.005079 | 0.004946 | 5.3e-05 | 0.000862 | 0.33052 | 1.303 | 2.5479 | 99.55 | 20.9 | 37.323 | 110.7954 | 112.7239 | 103.7651 | 0.009278 | 0.93667 | 0.022238 | 23.7416 | 0.25558 | 0.15554 | 0.037928 | 2.5137 | 2.5137 | 0.002543
2019-03-21 11:57 | 1553183849.3654 | 736 | 0.20179 | 0.009087 | 0.000174 | 94.5285 | 92.0283 | 110.8333 | 21.1089 | 0.003722 | 0.00338 | 153.5035 | 0.26156 | 8.402 | 0.003928 | 0.004949 | 0.004884 | 5.6e-05 | 0.00086 | 0.30486 | 1.2685 | 2.5478 | 99.55 | 20.9 | 37.321 | 111.3392 | 116.363 | 103.0377 | 0.009277 | 0.93848 | 0.022309 | 23.7711 | 0.23976 | 0.1556 | 0.035839 | 2.3832 | 2.3832 | 0.002683
2019-03-21 11:57 | 1553183850.3648 | 737 | 0.20216 | 0.008894 | 0.000165 | 94.7016 | 92.3744 | 111.0645 | 21.1091 | 0.003649 | 0.003338 | 153.442 | 0.26171 | 8.4036 | 0.003865 | 0.005627 | 0.005619 | 5.4e-05 | 0.000857 | 0.31653 | 1.2526 | 2.5479 | 99.55 | 20.9 | 37.32 | 109.5599 | 114.8062 | 102.0262 | 0.009277 | 0.94101 | 0.022362 | 23.7638 | 0.22255 | 0.15569 | 0.033877 | 2.2586 | 2.2586 | 0.002833
2019-03-21 11:58 | 1553183851.3654 | 738 | 0.20264 | 0.008671 | 0.000158 | 94.9271 | 92.8651 | 111.4549 | 21.1085 | 0.003677 | 0.003338 | 153.3725 | 0.26193 | 8.4019 | 0.003844 | 0.004831 | 0.004742 | 5.3e-05 | 0.000854 | 0.32565 | 1.2884 | 2.5478 | 99.55 | 20.9 | 37.32 | 110.2501 | 114.2924 | 102.0091 | 0.009277 | 0.94105 | 0.022402 | 23.805 | 0.19643 | 0.15582 | 0.031258 | 2.088 | 2.088 | 0.003067
2019-03-21 11:58 | 1553183852.3645 | 739 | 0.20305 | 0.008435 | 0.00015 | 95.0788 | 93.1788 | 111.8738 | 21.1082 | 0.003702 | 0.003349 | 153.3205 | 0.26171 | 8.4014 | 0.003819 | 0.004918 | 0.004727 | 5.8e-05 | 0.000858 | 0.32587 | 1.2847 | 2.5478 | 99.55 | 20.9 | 37.32 | 110.8445 | 113.6922 | 102.2898 | 0.009277 | 0.94035 | 0.022436 | 23.8586 | 0.18055 | 0.15569 | 0.029083 | 1.9459 | 1.9459 | 0.003288
2019-03-21 11:58 | 1553183853.3659 | 740 | 0.20333 | 0.008184 | 0.000143 | 95.2778 | 93.3174 | 112.0567 | 21.1088 | 0.00361 | 0.003314 | 153.2601 | 0.26088 | 8.4036 | 0.003756 | 0.005554 | 0.005519 | 5.6e-05 | 0.000856 | 0.35386 | 1.3418 | 2.548 | 99.55 | 20.9 | 37.32 | 108.6113 | 112.1625 | 101.4391 | 0.009277 | 0.94249 | 0.022479 | 23.8511 | 0.18596 | 0.1552 | 0.027728 | 1.8591 | 1.8591 | 0.003431
2019-03-21 11:58 | 1553183854.3651 | 741 | 0.20368 | 0.007927 | 0.000137 | 95.5299 | 93.8855 | 112.4006 | 21.1098 | 0.003658 | 0.003291 | 153.1977 | 0.26115 | 8.4026 | 0.003791 | 0.004715 | 0.004582 | 5.8e-05 | 0.00086 | 0.32587 | 1.3428 | 2.5479 | 99.55 | 20.9 | 37.322 | 109.7829 | 113.0087 | 100.879 | 0.009277 | 0.9439 | 0.022531 | 23.8698 | 0.1553 | 0.15536 | 0.025964 | 1.745 | 1.745 | 0.003659
2019-03-21 11:58 | 1553183855.3647 | 742 | 0.20394 | 0.00767 | 0.000131 | 95.6139 | 94.234 | 112.5433 | 21.1087 | 0.003652 | 0.003235 | 153.1301 | 0.26193 | 8.4034 | 0.003802 | 0.005276 | 0.005197 | 6e-05 | 0.00086 | 0.3246 | 1.3094 | 2.548 | 99.55 | 20.9 | 37.322 | 109.6343 | 113.2714 | 99.5258 | 0.009277 | 0.94733 | 0.022586 | 23.8417 | 0.13003 | 0.15582 | 0.0247 | 1.6642 | 1.6642 | 0.003848
2019-03-21 11:58 | 1553183856.3641 | 743 | 0.20404 | 0.007413 | 0.000126 | 95.7504 | 94.2779 | 112.8921 | 21.1091 | 0.003598 | 0.003299 | 153.0493 | 0.26228 | 8.4044 | 0.003854 | 0.005442 | 0.005395 | 5.5e-05 | 0.000856 | 0.32947 | 1.3127 | 2.5479 | 99.55 | 20.9 | 37.323 | 108.3301 | 114.5406 | 101.0683 | 0.009278 | 0.94342 | 0.022574 | 23.9279 | 0.13862 | 0.15603 | 0.024432 | 1.6456 | 1.6456 | 0.003897
2019-03-21 11:58 | 1553183857.3641 | 744 | 0.20419 | 0.007167 | 0.000122 | 95.8129 | 94.4441 | 113.362 | 21.109 | 0.00359 | 0.003259 | 152.981 | 0.26235 | 8.4017 | 0.003744 | 0.005004 | 0.004895 | 5.9e-05 | 0.00086 | 0.32189 | 1.3078 | 2.548 | 99.55 | 20.9 | 37.327 | 108.1158 | 111.8605 | 100.1004 | 0.009279 | 0.94587 | 0.02265 | 23.9467 | 0.12871 | 0.15607 | 0.023821 | 1.61 | 1.61 | 0.003984
2019-03-21 11:58 | 1553183858.3638 | 745 | 0.20437 | 0.006933 | 0.000119 | 95.9086 | 94.6955 | 113.4958 | 21.1092 | 0.003603 | 0.003216 | 152.9186 | 0.26158 | 8.3989 | 0.003759 | 0.00515 | 0.005116 | 5.5e-05 | 0.000859 | 0.32189 | 1.3182 | 2.5481 | 99.55 | 20.9 | 37.329 | 108.4334 | 112.2403 | 99.0674 | 0.009279 | 0.94849 | 0.022695 | 23.9276 | 0.11386 | 0.15562 | 0.023039 | 1.5603 | 1.5603 | 0.004099
2019-03-21 11:58 | 1553183859.3653 | 746 | 0.20456 | 0.006706 | 0.000116 | 96.0841 | 94.9577 | 113.7835 | 21.1094 | 0.003566 | 0.003244 | 152.8404 | 0.26144 | 8.4012 | 0.003665 | 0.005431 | 0.00529 | 5.6e-05 | 0.000861 | 0.33356 | 1.338 | 2.548 | 99.55 | 20.9 | 37.336 | 107.5364 | 109.9536 | 99.7407 | 0.009281 | 0.94678 | 0.022703 | 23.9796 | 0.10547 | 0.15553 | 0.022174 | 1.5024 | 1.5024 | 0.004255
2019-03-21 11:58 | 1553183860.3654 | 747 | 0.20474 | 0.006492 | 0.000113 | 96.2426 | 95.0984 | 114.2415 | 21.1092 | 0.003536 | 0.003237 | 152.7607 | 0.26175 | 8.4025 | 0.003664 | 0.004927 | 0.004793 | 5.8e-05 | 0.00086 | 0.31758 | 1.3231 | 2.548 | 99.55 | 20.9 | 37.341 | 106.8038 | 109.9132 | 99.5648 | 0.009282 | 0.94723 | 0.022754 | 24.0221 | 0.10698 | 0.15572 | 0.021372 | 1.4513 | 1.4513 | 0.00441
2019-03-21 11:58 | 1553183861.364 | 748 | 0.20491 | 0.006292 | 0.00011 | 96.3424 | 95.3232 | 114.6527 | 21.1089 | 0.003572 | 0.003226 | 152.6794 | 0.26155 | 8.4017 | 0.003696 | 0.004969 | 0.004891 | 5.6e-05 | 0.000867 | 0.31509 | 1.3018 | 2.548 | 99.55 | 20.9 | 37.345 | 107.6896 | 110.6961 | 99.3079 | 0.009283 | 0.94788 | 0.022803 | 24.057 | 0.095128 | 0.1556 | 0.02064 | 1.4046 | 1.4046 | 0.004553
2019-03-21 11:58 | 1553183862.3655 | 749 | 0.20484 | 0.006106 | 0.000108 | 96.5435 | 95.4897 | 114.9325 | 21.1094 | 0.003544 | 0.00319 | 152.6246 | 0.26205 | 8.403 | 0.003734 | 0.004698 | 0.004635 | 6e-05 | 0.000867 | 0.34523 | 1.343 | 2.5481 | 99.55 | 20.9 | 37.352 | 107.001 | 111.6295 | 98.4272 | 0.009285 | 0.95013 | 0.022858 | 24.0579 | 0.098166 | 0.15589 | 0.021244 | 1.4494 | 1.4494 | 0.004421
2019-03-21 11:58 | 1553183863.3651 | 750 | 0.20497 | 0.005925 | 0.000106 | 96.6719 | 95.7411 | 115.1794 | 21.1096 | 0.003543 | 0.003193 | 152.5441 | 0.26164 | 8.4027 | 0.003642 | 0.004984 | 0.004897 | 5.7e-05 | 0.000865 | 0.31349 | 1.3311 | 2.548 | 99.55 | 20.89 | 37.357 | 106.9917 | 109.3956 | 98.5214 | 0.00928 | 0.94989 | 0.02288 | 24.0868 | 0.086542 | 0.15565 | 0.020728 | 1.4156 | 1.4156 | 0.004519
2019-03-21 11:58 | 1553183864.3642 | 751 | 0.20499 | 0.005754 | 0.000105 | 96.8194 | 95.7475 | 115.2435 | 21.1096 | 0.003596 | 0.003174 | 152.4833 | 0.26186 | 8.4017 | 0.003591 | 0.005335 | 0.005226 | 5.5e-05 | 0.000873 | 0.32084 | 1.2999 | 2.5481 | 99.55 | 20.9 | 37.361 | 108.2842 | 108.1526 | 98.0456 | 0.009287 | 0.9511 | 0.022901 | 24.078 | 0.099573 | 0.15578 | 0.020819 | 1.4232 | 1.4232 | 0.004499
2019-03-21 11:58 | 1553183865.3656 | 752 | 0.20509 | 0.005601 | 0.000104 | 97.0128 | 95.9826 | 115.2182 | 21.1087 | 0.00351 | 0.003189 | 152.4295 | 0.26138 | 8.3985 | 0.0036 | 0.005242 | 0.005115 | 5.6e-05 | 0.000868 | 0.29806 | 1.3236 | 2.548 | 99.54 | 20.89 | 37.369 | 106.176 | 108.3676 | 98.4146 | 0.009284 | 0.95016 | 0.022887 | 24.0873 | 0.095492 | 0.1555 | 0.020399 | 1.3936 | 1.3936 | 0.004586
2019-03-21 11:58 | 1553183866.3652 | 753 | 0.2051 | 0.005456 | 0.000103 | 97.1397 | 96.1433 | 114.9471 | 21.1089 | 0.003512 | 0.003186 | 152.3731 | 0.26122 | 8.4037 | 0.003588 | 0.004994 | 0.004913 | 5.9e-05 | 0.00087 | 0.27328 | 1.2958 | 2.5479 | 99.54 | 20.9 | 37.374 | 106.2408 | 108.0796 | 98.3443 | 0.009291 | 0.95034 | 0.022862 | 24.0567 | 0.092218 | 0.1554 | 0.020511 | 1.3999 | 1.3999 | 0.004562
2019-03-21 11:58 | 1553183867.3644 | 754 | 0.20511 | 0.005324 | 0.000101 | 97.2633 | 96.2383 | 114.7842 | 21.1085 | 0.003554 | 0.003186 | 152.2938 | 0.26177 | 8.4014 | 0.003622 | 0.004705 | 0.004701 | 5.5e-05 | 0.000869 | 0.30917 | 1.2778 | 2.5479 | 99.54 | 20.89 | 37.383 | 107.2401 | 108.8969 | 98.3352 | 0.009288 | 0.95036 | 0.022846 | 24.0394 | 0.094761 | 0.15573 | 0.0206 | 1.405 | 1.405 | 0.004555
2019-03-21 11:58 | 1553183868.3644 | 755 | 0.2052 | 0.005204 | 0.000101 | 97.404 | 96.349 | 114.653 | 21.1092 | 0.003524 | 0.003137 | 152.2236 | 0.26146 | 8.403 | 0.003601 | 0.005052 | 0.004937 | 5.4e-05 | 0.000871 | 0.30359 | 1.308 | 2.5479 | 99.54 | 20.89 | 37.385 | 106.5354 | 108.3909 | 97.1469 | 0.009288 | 0.95341 | 0.02287 | 23.9872 | 0.097402 | 0.15554 | 0.020259 | 1.3832 | 1.3832 | 0.004622
2019-03-21 11:58 | 1553183869.3645 | 756 | 0.20527 | 0.005089 | 0.0001 | 97.5282 | 96.7891 | 114.3948 | 21.1093 | 0.003512 | 0.003154 | 152.1604 | 0.26201 | 8.4051 | 0.003554 | 0.004742 | 0.004691 | 5.6e-05 | 0.000868 | 0.33915 | 1.2802 | 2.5479 | 99.54 | 20.89 | 37.388 | 106.2363 | 107.2511 | 97.5616 | 0.009289 | 0.95234 | 0.022831 | 23.9736 | 0.068046 | 0.15587 | 0.019936 | 1.3589 | 1.3589 | 0.004714
2019-03-21 11:58 | 1553183870.3645 | 757 | 0.20528 | 0.00499 | 9.9e-05 | 97.6304 | 96.8425 | 114.3077 | 21.1096 | 0.003512 | 0.003199 | 152.0853 | 0.26188 | 8.4021 | 0.003578 | 0.005272 | 0.00514 | 5.5e-05 | 0.00087 | 0.34418 | 1.2986 | 2.5478 | 99.54 | 20.89 | 37.39 | 106.2304 | 107.8321 | 98.6618 | 0.009289 | 0.94953 | 0.022789 | 24 | 0.072473 | 0.15579 | 0.020044 | 1.3637 | 1.3637 | 0.004695
2019-03-21 11:58 | 1553183871.3649 | 758 | 0.20536 | 0.004899 | 9.8e-05 | 97.7984 | 97.1215 | 114.1603 | 21.1093 | 0.003499 | 0.003082 | 152.0356 | 0.26149 | 8.4021 | 0.003634 | 0.005253 | 0.005233 | 5.7e-05 | 0.000872 | 0.35059 | 1.3017 | 2.5479 | 99.54 | 20.89 | 37.391 | 105.9114 | 109.199 | 95.835 | 0.00929 | 0.9568 | 0.022861 | 23.8931 | 0.062129 | 0.15556 | 0.019662 | 1.342 | 1.342 | 0.004764
2019-03-21 11:58 | 1553183872.3646 | 759 | 0.20541 | 0.004809 | 9.7e-05 | 97.8276 | 96.835 | 113.7648 | 21.1091 | 0.003475 | 0.003136 | 151.968 | 0.26225 | 8.403 | 0.003591 | 0.005536 | 0.005462 | 5.4e-05 | 0.000875 | 0.32925 | 1.3145 | 2.5479 | 99.54 | 20.89 | 37.391 | 105.3313 | 108.1412 | 97.1415 | 0.00929 | 0.95343 | 0.022781 | 23.8939 | 0.091217 | 0.15601 | 0.019478 | 1.3248 | 1.3248 | 0.00484
2019-03-21 11:58 | 1553183873.3636 | 760 | 0.20536 | 0.004725 | 9.6e-05 | 97.9829 | 97.1109 | 113.749 | 21.1103 | 0.003464 | 0.003143 | 151.8946 | 0.26224 | 8.4025 | 0.003554 | 0.004926 | 0.004842 | 5.6e-05 | 0.00087 | 0.35635 | 1.3278 | 2.5479 | 99.54 | 20.89 | 37.392 | 105.0777 | 107.2563 | 97.3095 | 0.00929 | 0.95299 | 0.022774 | 23.8977 | 0.079953 | 0.15601 | 0.019892 | 1.3527 | 1.3527 | 0.00474
2019-03-21 11:58 | 1553183874.3651 | 761 | 0.20547 | 0.004652 | 9.5e-05 | 98.0657 | 97.288 | 113.8169 | 21.1108 | 0.003471 | 0.003141 | 151.836 | 0.26179 | 8.4033 | 0.003527 | 0.005597 | 0.005534 | 5.6e-05 | 0.000871 | 0.29983 | 1.3345 | 2.5479 | 99.54 | 20.89 | 37.392 | 105.244 | 106.5938 | 97.2482 | 0.00929 | 0.95315 | 0.022783 | 23.9028 | 0.071213 | 0.15574 | 0.019315 | 1.3139 | 1.3139 | 0.004871
2019-03-21 11:58 | 1553183875.3638 | 762 | 0.20539 | 0.00458 | 9.5e-05 | 98.3139 | 97.5469 | 113.9441 | 21.1106 | 0.003465 | 0.003104 | 151.7503 | 0.26152 | 8.4006 | 0.003497 | 0.005142 | 0.005095 | 5.5e-05 | 0.000872 | 0.30646 | 1.3207 | 2.548 | 99.54 | 20.89 | 37.394 | 105.1007 | 105.8757 | 96.3694 | 0.00929 | 0.95542 | 0.022823 | 23.8878 | 0.070055 | 0.15558 | 0.019851 | 1.3529 | 1.3529 | 0.004727
2019-03-21 11:58 | 1553183876.3637 | 763 | 0.20547 | 0.004512 | 9.5e-05 | 98.2882 | 97.7393 | 114.2668 | 21.1102 | 0.00341 | 0.003094 | 151.6594 | 0.26197 | 8.4031 | 0.003558 | 0.004944 | 0.004849 | 5.6e-05 | 0.000871 | 0.31365 | 1.3059 | 2.548 | 99.54 | 20.89 | 37.395 | 103.7502 | 107.353 | 96.1241 | 0.009291 | 0.95605 | 0.022863 | 23.9136 | 0.050096 | 0.15585 | 0.019493 | 1.3307 | 1.3307 | 0.004813
2019-03-21 11:58 | 1553183877.3648 | 764 | 0.20556 | 0.004454 | 9.5e-05 | 98.4789 | 97.8558 | 114.3821 | 21.11 | 0.003415 | 0.003093 | 151.6142 | 0.26163 | 8.4004 | 0.003545 | 0.004629 | 0.004539 | 5.8e-05 | 0.000868 | 0.30901 | 1.3005 | 2.5479 | 99.54 | 20.89 | 37.395 | 103.8718 | 107.029 | 96.0862 | 0.009291 | 0.95615 | 0.022875 | 23.9245 | 0.056779 | 0.15564 | 0.018988 | 1.2969 | 1.2969 | 0.004933
2019-03-21 11:58 | 1553183878.364 | 765 | 0.20552 | 0.004391 | 9.5e-05 | 98.5704 | 97.9174 | 114.445 | 21.1108 | 0.003402 | 0.003077 | 151.5521 | 0.26175 | 8.404 | 0.003523 | 0.005378 | 0.005334 | 5.6e-05 | 0.000867 | 0.32333 | 1.2838 | 2.5479 | 99.54 | 20.89 | 37.395 | 103.5715 | 106.4983 | 95.7073 | 0.009291 | 0.95713 | 0.022893 | 23.9188 | 0.059449 | 0.15571 | 0.019345 | 1.3224 | 1.3224 | 0.00484
2019-03-21 11:58 | 1553183879.3649 | 766 | 0.20552 | 0.004337 | 9.5e-05 | 98.6789 | 97.9753 | 114.3537 | 21.1114 | 0.003419 | 0.003053 | 151.4846 | 0.26111 | 8.4023 | 0.003505 | 0.005228 | 0.005152 | 5.9e-05 | 0.000866 | 0.3126 | 1.2982 | 2.548 | 99.54 | 20.89 | 37.395 | 103.9905 | 106.0645 | 95.1316 | 0.009291 | 0.95863 | 0.022902 | 23.8906 | 0.063998 | 0.15534 | 0.019397 | 1.3265 | 1.3265 | 0.004813
2019-03-21 11:58 | 1553183880.3639 | 767 | 0.20553 | 0.004292 | 9.4e-05 | 98.7596 | 97.9985 | 114.2924 | 21.1113 | 0.003402 | 0.003097 | 151.4263 | 0.26262 | 8.4034 | 0.003488 | 0.005 | 0.0049 | 6e-05 | 0.000864 | 0.32421 | 1.264 | 2.5478 | 99.54 | 20.89 | 37.395 | 103.558 | 105.656 | 96.2003 | 0.009291 | 0.95586 | 0.022863 | 23.9188 | 0.069199 | 0.15623 | 0.019413 | 1.3253 | 1.3253 | 0.004845
2019-03-21 11:58 | 1553183881.3645 | 768 | 0.20545 | 0.004242 | 9.4e-05 | 98.8424 | 98.2655 | 114.2288 | 21.1114 | 0.003407 | 0.003087 | 151.3522 | 0.26165 | 8.4015 | 0.003488 | 0.004839 | 0.004748 | 5.8e-05 | 0.000862 | 0.30326 | 1.2925 | 2.548 | 99.54 | 20.89 | 37.395 | 103.6838 | 105.6539 | 95.9421 | 0.009291 | 0.95652 | 0.022865 | 23.9038 | 0.05236 | 0.15566 | 0.019963 | 1.363 | 1.363 | 0.004694
2019-03-21 11:58 | 1553183882.3651 | 769 | 0.20557 | 0.004206 | 9.3e-05 | 98.9059 | 98.3292 | 114.5945 | 21.1111 | 0.00339 | 0.003106 | 151.2746 | 0.2629 | 8.4005 | 0.003543 | 0.005561 | 0.005466 | 6.3e-05 | 0.000864 | 0.30558 | 1.293 | 2.548 | 99.54 | 20.89 | 37.394 | 103.2889 | 106.9978 | 96.4031 | 0.00929 | 0.95533 | 0.022887 | 23.957 | 0.052305 | 0.1564 | 0.019288 | 1.3182 | 1.3182 | 0.004876
2019-03-21 11:58 | 1553183883.3642 | 770 | 0.20564 | 0.00417 | 9.2e-05 | 98.9604 | 98.4004 | 114.6065 | 21.1105 | 0.003405 | 0.003045 | 151.2283 | 0.26199 | 8.4028 | 0.003501 | 0.005059 | 0.004926 | 5.9e-05 | 0.000868 | 0.30254 | 1.3042 | 2.548 | 99.54 | 20.89 | 37.394 | 103.6475 | 105.9756 | 94.9319 | 0.00929 | 0.95915 | 0.022934 | 23.9105 | 0.050762 | 0.15586 | 0.018907 | 1.2948 | 1.2948 | 0.004947
2019-03-21 11:58 | 1553183884.3653 | 771 | 0.20555 | 0.00413 | 9.2e-05 | 99.0706 | 98.5874 | 114.3713 | 21.1126 | 0.003415 | 0.003063 | 151.172 | 0.26191 | 8.4023 | 0.00354 | 0.005123 | 0.005094 | 6e-05 | 0.000873 | 0.33323 | 1.3005 | 2.5481 | 99.54 | 20.89 | 37.394 | 103.8909 | 106.9207 | 95.3652 | 0.00929 | 0.95802 | 0.022897 | 23.9 | 0.04374 | 0.15581 | 0.019498 | 1.3333 | 1.3333 | 0.004803
2019-03-21 11:58 | 1553183885.3643 | 772 | 0.20554 | 0.004096 | 9.1e-05 | 99.1234 | 98.7212 | 114.4441 | 21.1113 | 0.003415 | 0.003005 | 151.1248 | 0.26158 | 8.4019 | 0.003514 | 0.005446 | 0.005405 | 6.3e-05 | 0.00087 | 0.31349 | 1.3156 | 2.548 | 99.54 | 20.89 | 37.392 | 103.8913 | 106.288 | 93.9785 | 0.00929 | 0.96164 | 0.022947 | 23.8626 | 0.036364 | 0.15562 | 0.019592 | 1.3427 | 1.3427 | 0.004764
2019-03-21 11:58 | 1553183886.3653 | 773 | 0.20569 | 0.004065 | 9e-05 | 99.2074 | 98.8438 | 114.311 | 21.1105 | 0.003451 | 0.003017 | 151.056 | 0.26227 | 8.406 | 0.003566 | 0.005229 | 0.0051 | 6.3e-05 | 0.000874 | 0.31349 | 1.3239 | 2.548 | 99.54 | 20.89 | 37.391 | 104.7467 | 107.5438 | 94.2616 | 0.00929 | 0.9609 | 0.022925 | 23.8579 | 0.03284 | 0.15603 | 0.018705 | 1.2806 | 1.2806 | 0.005008
2019-03-21 11:58 | 1553183887.3654 | 774 | 0.20554 | 0.004025 | 9e-05 | 99.2309 | 98.7641 | 114.3504 | 21.1114 | 0.003387 | 0.003075 | 151.0004 | 0.26139 | 8.4012 | 0.003549 | 0.005599 | 0.005462 | 6.2e-05 | 0.000877 | 0.34844 | 1.3241 | 2.5479 | 99.54 | 20.89 | 37.391 | 103.1998 | 107.1367 | 95.6671 | 0.00929 | 0.95724 | 0.022885 | 23.9076 | 0.042177 | 0.1555 | 0.019665 | 1.3442 | 1.3442 | 0.004755
2019-03-21 11:58 | 1553183888.3644 | 775 | 0.20564 | 0.003995 | 8.9e-05 | 99.3692 | 98.9712 | 114.6762 | 21.1113 | 0.003375 | 0.003071 | 150.9399 | 0.26148 | 8.4022 | 0.00354 | 0.004682 | 0.00464 | 6.2e-05 | 0.000871 | 0.27759 | 1.3459 | 2.548 | 99.54 | 20.89 | 37.389 | 102.9169 | 106.9194 | 95.5575 | 0.009289 | 0.95752 | 0.022921 | 23.9381 | 0.035896 | 0.15555 | 0.01914 | 1.3103 | 1.3103 | 0.004879
2019-03-21 11:58 | 1553183889.3639 | 776 | 0.20563 | 0.00396 | 8.8e-05 | 99.4404 | 99.0747 | 114.7316 | 21.1105 | 0.003362 | 0.003065 | 150.8581 | 0.2618 | 8.4024 | 0.003543 | 0.005161 | 0.005122 | 6.1e-05 | 0.000869 | 0.31868 | 1.3388 | 2.5479 | 99.54 | 20.89 | 37.388 | 102.601 | 106.9783 | 95.4188 | 0.009289 | 0.95788 | 0.022931 | 23.9393 | 0.032951 | 0.15575 | 0.019202 | 1.3151 | 1.3151 | 0.004867
2019-03-21 11:58 | 1553183890.3636 | 777 | 0.20574 | 0.003934 | 8.7e-05 | 99.4371 | 98.9334 | 114.8126 | 21.1108 | 0.003296 | 0.003028 | 150.8094 | 0.26164 | 8.4051 | 0.003609 | 0.00519 | 0.005096 | 6.1e-05 | 0.00087 | 0.31454 | 1.3288 | 2.5479 | 99.54 | 20.89 | 37.387 | 101.0071 | 108.5919 | 94.5322 | 0.009289 | 0.96019 | 0.022967 | 23.919 | 0.045425 | 0.15565 | 0.018622 | 1.2774 | 1.2774 | 0.005008
2019-03-21 11:58 | 1553183891.3646 | 778 | 0.20565 | 0.003908 | 8.7e-05 | 99.4309 | 98.9504 | 114.8795 | 21.1113 | 0.003316 | 0.003023 | 150.7615 | 0.26273 | 8.4009 | 0.00355 | 0.004675 | 0.004632 | 6.1e-05 | 0.000869 | 0.33019 | 1.3026 | 2.548 | 99.54 | 20.89 | 37.387 | 101.4832 | 107.1632 | 94.412 | 0.009289 | 0.96051 | 0.022977 | 23.922 | 0.043323 | 0.1563 | 0.019141 | 1.3137 | 1.3137 | 0.00489
2019-03-21 11:58 | 1553183892.3646 | 779 | 0.20564 | 0.003887 | 8.6e-05 | 99.5727 | 99.133 | 114.8351 | 21.1114 | 0.003393 | 0.003057 | 150.6839 | 0.26046 | 8.4013 | 0.003519 | 0.005231 | 0.005102 | 6e-05 | 0.000873 | 0.31653 | 1.3135 | 2.548 | 99.54 | 20.89 | 37.387 | 103.342 | 106.3983 | 95.2216 | 0.009289 | 0.95839 | 0.022948 | 23.9437 | 0.039585 | 0.15495 | 0.019264 | 1.3205 | 1.3205 | 0.004823
2019-03-21 11:58 | 1553183893.3647 | 780 | 0.2056 | 0.003867 | 8.6e-05 | 99.6186 | 99.2508 | 114.8851 | 21.1117 | 0.003376 | 0.00302 | 150.6078 | 0.2623 | 8.4031 | 0.003572 | 0.005441 | 0.005343 | 5.9e-05 | 0.00087 | 0.3115 | 1.3119 | 2.5481 | 99.54 | 20.89 | 37.383 | 102.9494 | 107.6945 | 94.3395 | 0.009288 | 0.96069 | 0.02298 | 23.9202 | 0.033079 | 0.15604 | 0.019517 | 1.3397 | 1.3397 | 0.004787
2019-03-21 11:58 | 1553183894.3637 | 781 | 0.20554 | 0.003844 | 8.6e-05 | 99.6933 | 99.1984 | 115.1344 | 21.1114 | 0.00335 | 0.003033 | 150.5464 | 0.26193 | 8.4022 | 0.003558 | 0.004612 | 0.004494 | 6.1e-05 | 0.000873 | 0.35419 | 1.2829 | 2.5479 | 99.54 | 20.89 | 37.382 | 102.3167 | 107.3469 | 94.6461 | 0.009287 | 0.95989 | 0.022995 | 23.9562 | 0.04452 | 0.15582 | 0.019891 | 1.3664 | 1.3664 | 0.004687
2019-03-21 11:58 | 1553183895.3644 | 782 | 0.20562 | 0.003829 | 8.5e-05 | 99.6692 | 99.3775 | 115.0886 | 21.1111 | 0.003353 | 0.003031 | 150.464 | 0.26175 | 8.4055 | 0.003563 | 0.00524 | 0.005159 | 5.7e-05 | 0.000875 | 0.30397 | 1.3213 | 2.5479 | 99.54 | 20.89 | 37.38 | 102.3706 | 107.4648 | 94.6036 | 0.009287 | 0.96 | 0.022992 | 23.95 | 0.026215 | 0.15572 | 0.01945 | 1.3359 | 1.3359 | 0.004791
2019-03-21 11:58 | 1553183896.3649 | 783 | 0.20555 | 0.003809 | 8.5e-05 | 99.7425 | 99.304 | 115.041 | 21.1113 | 0.003357 | 0.002963 | 150.4285 | 0.26167 | 8.3993 | 0.003534 | 0.00526 | 0.005213 | 5.9e-05 | 0.000873 | 0.34346 | 1.2744 | 2.548 | 99.54 | 20.89 | 37.378 | 102.4783 | 106.7759 | 92.97 | 0.009286 | 0.96429 | 0.023039 | 23.8918 | 0.039414 | 0.15567 | 0.019904 | 1.3699 | 1.3699 | 0.00467
2019-03-21 11:58 | 1553183897.3649 | 784 | 0.20558 | 0.003798 | 8.4e-05 | 99.7398 | 99.4385 | 115.2611 | 21.1114 | 0.003324 | 0.002992 | 150.3592 | 0.26175 | 8.4011 | 0.003537 | 0.004933 | 0.004821 | 6.1e-05 | 0.000872 | 0.30414 | 1.3087 | 2.5479 | 99.54 | 20.89 | 37.376 | 101.6805 | 106.8297 | 93.6639 | 0.009286 | 0.96247 | 0.023039 | 23.9373 | 0.027064 | 0.15571 | 0.019705 | 1.3562 | 1.3562 | 0.004719
2019-03-21 11:58 | 1553183898.3649 | 785 | 0.20567 | 0.003785 | 8.5e-05 | 99.738 | 99.2825 | 115.1196 | 21.1117 | 0.003348 | 0.003014 | 150.2796 | 0.262 | 8.4031 | 0.003592 | 0.0052 | 0.005134 | 6.5e-05 | 0.000876 | 0.32068 | 1.3606 | 2.5479 | 99.54 | 20.89 | 37.37 | 102.2652 | 108.1768 | 94.1792 | 0.009284 | 0.96111 | 0.023009 | 23.9394 | 0.040946 | 0.15586 | 0.019192 | 1.3191 | 1.3191 | 0.004856
2019-03-21 11:58 | 1553183899.3641 | 786 | 0.20577 | 0.003774 | 8.4e-05 | 99.7812 | 99.4882 | 115.3308 | 21.1117 | 0.003337 | 0.003005 | 150.2428 | 0.26162 | 8.4004 | 0.003567 | 0.004866 | 0.004859 | 5.9e-05 | 0.000873 | 0.35098 | 1.346 | 2.548 | 99.54 | 20.89 | 37.366 | 102.0013 | 107.5765 | 93.9615 | 0.009283 | 0.96168 | 0.023036 | 23.9543 | 0.026306 | 0.15564 | 0.018607 | 1.2804 | 1.2804 | 0.004996
2019-03-21 11:58 | 1553183900.3651 | 787 | 0.20554 | 0.003757 | 8.5e-05 | 99.7902 | 99.5907 | 115.2848 | 21.1116 | 0.003318 | 0.002996 | 150.1751 | 0.26256 | 8.4059 | 0.003573 | 0.004839 | 0.004742 | 6.1e-05 | 0.000873 | 0.3079 | 1.3493 | 2.548 | 99.54 | 20.89 | 37.363 | 101.523 | 107.7267 | 93.7582 | 0.009283 | 0.96222 | 0.023038 | 23.9429 | 0.017895 | 0.1562 | 0.020041 | 1.3793 | 1.3793 | 0.004654
2019-03-21 11:58 | 1553183901.3641 | 788 | 0.20562 | 0.003752 | 8.5e-05 | 99.9189 | 99.5775 | 115.2135 | 21.1114 | 0.003366 | 0.003044 | 150.0889 | 0.26178 | 8.4036 | 0.003596 | 0.005154 | 0.005121 | 6.4e-05 | 0.000874 | 0.31421 | 1.3598 | 2.5479 | 99.54 | 20.89 | 37.358 | 102.6993 | 108.2828 | 94.9147 | 0.009281 | 0.95919 | 0.022995 | 23.9731 | 0.030617 | 0.15574 | 0.019526 | 1.3413 | 1.3413 | 0.004772
2019-03-21 11:58 | 1553183902.3642 | 789 | 0.2058 | 0.003735 | 8.5e-05 | 99.9698 | 99.5226 | 114.9255 | 21.112 | 0.003313 | 0.002993 | 150.0321 | 0.26205 | 8.4028 | 0.003621 | 0.005016 | 0.004989 | 5.8e-05 | 0.000875 | 0.31238 | 1.3345 | 2.5479 | 99.54 | 20.89 | 37.354 | 101.4181 | 108.8705 | 93.6952 | 0.00928 | 0.96238 | 0.023004 | 23.9035 | 0.040104 | 0.15589 | 0.018517 | 1.2724 | 1.2724 | 0.005036
2019-03-21 11:58 | 1553183903.3642 | 790 | 0.20539 | 0.003726 | 8.5e-05 | 100.0053 | 99.5059 | 114.617 | 21.1108 | 0.003301 | 0.003021 | 149.9812 | 0.26204 | 8.4015 | 0.003572 | 0.005243 | 0.005029 | 5.8e-05 | 0.000879 | 0.34202 | 1.3618 | 2.548 | 99.54 | 20.89 | 37.348 | 101.1249 | 107.6937 | 94.3529 | 0.009279 | 0.96066 | 0.022953 | 23.8928 | 0.044776 | 0.15589 | 0.020934 | 1.4356 | 1.4356 | 0.004463
2019-03-21 11:58 | 1553183904.3632 | 791 | 0.20587 | 0.003719 | 8.5e-05 | 99.9961 | 99.6602 | 114.4814 | 21.1122 | 0.00333 | 0.002982 | 149.9182 | 0.2613 | 8.4039 | 0.003583 | 0.005144 | 0.005103 | 6.3e-05 | 0.00088 | 0.3314 | 1.3623 | 2.5481 | 99.54 | 20.89 | 37.345 | 101.8179 | 107.9495 | 93.4081 | 0.009278 | 0.96314 | 0.022969 | 23.8479 | 0.030092 | 0.15545 | 0.018081 | 1.2404 | 1.2404 | 0.005151
2019-03-21 11:58 | 1553183905.3637 | 792 | 0.20572 | 0.003709 | 8.5e-05 | 99.9579 | 99.7913 | 114.3666 | 21.1127 | 0.00332 | 0.002985 | 149.8683 | 0.26094 | 8.4033 | 0.003572 | 0.004448 | 0.004367 | 5.8e-05 | 0.000881 | 0.3079 | 1.3703 | 2.548 | 99.54 | 20.89 | 37.338 | 101.5741 | 107.6994 | 93.4901 | 0.009276 | 0.96292 | 0.022955 | 23.8386 | 0.014919 | 0.15523 | 0.01901 | 1.3035 | 1.3035 | 0.004895
2019-03-21 11:58 | 1553183906.3633 | 793 | 0.20544 | 0.003703 | 8.6e-05 | 100.0935 | 99.7801 | 114.2443 | 21.1123 | 0.003324 | 0.003068 | 149.8131 | 0.26156 | 8.4033 | 0.003649 | 0.00461 | 0.004614 | 5.9e-05 | 0.000879 | 0.3011 | 1.3075 | 2.548 | 99.54 | 20.89 | 37.336 | 101.6878 | 109.5599 | 95.4949 | 0.009276 | 0.95768 | 0.02288 | 23.8909 | 0.028049 | 0.1556 | 0.020719 | 1.4162 | 1.4162 | 0.004516
2019-03-21 11:58 | 1553183907.3634 | 794 | 0.20605 | 0.003692 | 8.7e-05 | 100.1206 | 99.7205 | 113.8274 | 21.1125 | 0.003315 | 0.003034 | 149.7555 | 0.26158 | 8.4031 | 0.003615 | 0.004734 | 0.004624 | 6.2e-05 | 0.000879 | 0.28423 | 1.3015 | 2.548 | 99.54 | 20.89 | 37.333 | 101.4697 | 108.7291 | 94.6706 | 0.009275 | 0.95983 | 0.022864 | 23.8206 | 0.035816 | 0.15562 | 0.017048 | 1.164 | 1.164 | 0.005495
2019-03-21 11:58 | 1553183908.3635 | 795 | 0.20558 | 0.003682 | 8.8e-05 | 100.1012 | 99.6166 | 113.6444 | 21.1113 | 0.003337 | 0.002991 | 149.7111 | 0.26199 | 8.402 | 0.003498 | 0.005074 | 0.00496 | 6.1e-05 | 0.000879 | 0.32156 | 1.2853 | 2.5478 | 99.54 | 20.89 | 37.328 | 101.9883 | 105.9048 | 93.6341 | 0.009274 | 0.96254 | 0.022878 | 23.7679 | 0.043408 | 0.15586 | 0.019876 | 1.3582 | 1.3582 | 0.004716
2019-03-21 11:58 | 1553183909.3636 | 796 | 0.20471 | 0.003677 | 8.7e-05 | 100.1873 | 99.6473 | 113.369 | 21.1123 | 0.003277 | 0.003002 | 149.6577 | 0.26208 | 8.4037 | 0.003552 | 0.005007 | 0.004889 | 6.2e-05 | 0.000872 | 0.29302 | 1.2752 | 2.5479 | 99.54 | 20.89 | 37.327 | 100.5321 | 107.2052 | 93.9049 | 0.009274 | 0.96183 | 0.022841 | 23.7478 | 0.048347 | 0.15591 | 0.02511 | 1.7137 | 1.7137 | 0.003739
2019-03-21 11:58 | 1553183910.365 | 797 | 0.20546 | 0.003665 | 8.8e-05 | 100.1608 | 99.8077 | 113.1 | 21.1137 | 0.003352 | 0.002941 | 149.5989 | 0.26199 | 8.4025 | 0.003517 | 0.004861 | 0.00479 | 6.1e-05 | 0.000875 | 0.31421 | 1.3058 | 2.5479 | 99.54 | 20.89 | 37.324 | 102.3715 | 106.3627 | 92.4404 | 0.009273 | 0.96569 | 0.02286 | 23.6723 | 0.031591 | 0.15585 | 0.020608 | 1.4073 | 1.4073 | 0.004552
2019-03-21 11:58 | 1553183911.3646 | 798 | 0.20493 | 0.003659 | 8.8e-05 | 100.2195 | 99.8315 | 113.1925 | 21.1126 | 0.003302 | 0.002989 | 149.5429 | 0.26198 | 8.4013 | 0.003514 | 0.004915 | 0.004819 | 5.9e-05 | 0.000874 | 0.35474 | 1.3055 | 2.5481 | 99.54 | 20.89 | 37.32 | 101.1392 | 106.2758 | 93.5978 | 0.009272 | 0.96264 | 0.022833 | 23.7194 | 0.034695 | 0.15585 | 0.023802 | 1.6237 | 1.6237 | 0.003945
2019-03-21 11:59 | 1553183912.3637 | 799 | 0.20555 | 0.003654 | 8.8e-05 | 100.2305 | 99.8234 | 113.3655 | 21.1118 | 0.003309 | 0.002973 | 149.5059 | 0.26061 | 8.4044 | 0.003505 | 0.005789 | 0.005769 | 6.1e-05 | 0.000876 | 0.31669 | 1.2843 | 2.5479 | 99.54 | 20.89 | 37.318 | 101.3064 | 106.0719 | 93.2098 | 0.009272 | 0.96366 | 0.022863 | 23.725 | 0.036411 | 0.15504 | 0.020087 | 1.3718 | 1.3718 | 0.004645
2019-03-21 11:59 | 1553183913.3642 | 800 | 0.20575 | 0.003646 | 8.8e-05 | 100.2156 | 100.0318 | 113.5901 | 21.1132 | 0.003284 | 0.002996 | 149.4475 | 0.26165 | 8.4031 | 0.003587 | 0.005037 | 0.004962 | 6.2e-05 | 0.000879 | 0.31349 | 1.3155 | 2.5477 | 99.54 | 20.88 | 37.312 | 100.7229 | 108.0459 | 93.7681 | 0.009264 | 0.96219 | 0.022868 | 23.7666 | 0.016424 | 0.15565 | 0.018872 | 1.289 | 1.289 | 0.004963
2019-03-21 11:59 | 1553183914.3643 | 801 | 0.20608 | 0.003644 | 8.9e-05 | 100.2246 | 100.1023 | 113.7174 | 21.1126 | 0.003287 | 0.003017 | 149.3868 | 0.26173 | 8.4031 | 0.003526 | 0.005106 | 0.005007 | 6.2e-05 | 0.000881 | 0.3095 | 1.3298 | 2.5481 | 99.54 | 20.88 | 37.31 | 100.7811 | 106.5647 | 94.2741 | 0.009264 | 0.96087 | 0.022865 | 23.7962 | 0.01092 | 0.1557 | 0.016898 | 1.1537 | 1.1537 | 0.005547
2019-03-21 11:59 | 1553183915.3649 | 802 | 0.2061 | 0.003635 | 8.9e-05 | 100.2425 | 100.0621 | 113.9087 | 21.1122 | 0.003307 | 0.002917 | 149.3218 | 0.26193 | 8.4023 | 0.003566 | 0.005022 | 0.004938 | 6.3e-05 | 0.000879 | 0.29607 | 1.3124 | 2.5479 | 99.54 | 20.88 | 37.308 | 101.2794 | 107.5399 | 91.8445 | 0.009263 | 0.96726 | 0.02296 | 23.7374 | 0.016107 | 0.15582 | 0.016815 | 1.1528 | 1.1528 | 0.005556
2019-03-21 11:59 | 1553183916.364 | 803 | 0.20631 | 0.003635 | 8.9e-05 | 100.2755 | 100.0584 | 114.2694 | 21.1134 | 0.003305 | 0.002989 | 149.2369 | 0.26274 | 8.4016 | 0.003433 | 0.005343 | 0.00528 | 6e-05 | 0.000875 | 0.30359 | 1.322 | 2.548 | 99.54 | 20.88 | 37.307 | 101.2232 | 104.3305 | 93.5783 | 0.009263 | 0.96269 | 0.022942 | 23.8313 | 0.019392 | 0.1563 | 0.015538 | 1.0641 | 1.0641 | 0.006037
2019-03-21 11:59 | 1553183917.3644 | 804 | 0.20552 | 0.003637 | 9e-05 | 100.3214 | 100.0941 | 114.5201 | 21.1131 | 0.003308 | 0.002981 | 149.1639 | 0.26106 | 8.4018 | 0.003372 | 0.005501 | 0.005412 | 6e-05 | 0.000879 | 0.31166 | 1.3556 | 2.5481 | 99.54 | 20.88 | 37.306 | 101.3051 | 102.8486 | 93.3918 | 0.009263 | 0.96318 | 0.022973 | 23.8514 | 0.020288 | 0.1553 | 0.020296 | 1.3927 | 1.3927 | 0.004583
2019-03-21 11:59 | 1553183918.364 | 805 | 0.20552 | 0.003632 | 8.9e-05 | 100.3181 | 99.978 | 114.6659 | 21.1127 | 0.003315 | 0.002959 | 149.1072 | 0.26265 | 8.4033 | 0.00349 | 0.004808 | 0.004669 | 6.3e-05 | 0.000879 | 0.32173 | 1.309 | 2.548 | 99.54 | 20.88 | 37.303 | 101.4541 | 105.7102 | 92.8586 | 0.009262 | 0.96458 | 0.023005 | 23.8492 | 0.03038 | 0.15625 | 0.020325 | 1.3967 | 1.3967 | 0.004598
2019-03-21 11:59 | 1553183919.3631 | 806 | 0.20524 | 0.003625 | 8.9e-05 | 100.3977 | 99.8578 | 114.595 | 21.1128 | 0.003298 | 0.002973 | 149.0199 | 0.2618 | 8.4006 | 0.003502 | 0.005461 | 0.005321 | 6e-05 | 0.000874 | 0.32731 | 1.2886 | 2.5478 | 99.54 | 20.88 | 37.302 | 101.0618 | 105.988 | 93.2045 | 0.009262 | 0.96367 | 0.022987 | 23.8531 | 0.048233 | 0.15575 | 0.022012 | 1.5115 | 1.5115 | 0.004235
2019-03-21 11:59 | 1553183920.3642 | 807 | 0.20501 | 0.003626 | 8.8e-05 | 100.3762 | 99.9371 | 114.695 | 21.1132 | 0.003311 | 0.002998 | 148.9279 | 0.26136 | 8.4025 | 0.003475 | 0.005182 | 0.005076 | 5.9e-05 | 0.000876 | 0.32228 | 1.285 | 2.5478 | 99.54 | 20.88 | 37.301 | 101.3769 | 105.3472 | 93.8121 | 0.009262 | 0.96208 | 0.022978 | 23.8833 | 0.039216 | 0.15548 | 0.023382 | 1.6051 | 1.6051 | 0.003981
2019-03-21 11:59 | 1553183921.3646 | 808 | 0.20599 | 0.003627 | 8.9e-05 | 100.3905 | 100.0853 | 114.6466 | 21.1139 | 0.003327 | 0.003004 | 148.8712 | 0.26217 | 8.4016 | 0.003516 | 0.004771 | 0.004646 | 6.2e-05 | 0.000876 | 0.33411 | 1.3087 | 2.5479 | 99.54 | 20.88 | 37.3 | 101.7481 | 106.3229 | 93.9396 | 0.009261 | 0.96174 | 0.022969 | 23.8824 | 0.02724 | 0.15597 | 0.017474 | 1.1985 | 1.1985 | 0.005349
2019-03-21 11:59 | 1553183922.3633 | 809 | 0.20543 | 0.003625 | 8.8e-05 | 100.4245 | 99.9 | 114.5656 | 21.1135 | 0.003337 | 0.003046 | 148.798 | 0.26181 | 8.4022 | 0.003605 | 0.005551 | 0.005461 | 6e-05 | 0.000875 | 0.3027 | 1.266 | 2.5479 | 99.54 | 20.88 | 37.299 | 101.9999 | 108.5054 | 94.955 | 0.009261 | 0.95909 | 0.022929 | 23.907 | 0.04684 | 0.15575 | 0.020861 | 1.4289 | 1.4289 | 0.00448
2019-03-21 11:59 | 1553183923.3633 | 810 | 0.2051 | 0.003629 | 8.8e-05 | 100.3759 | 100.043 | 114.306 | 21.1137 | 0.00331 | 0.002992 | 148.7251 | 0.26082 | 8.4008 | 0.003476 | 0.004714 | 0.004587 | 5.9e-05 | 0.000877 | 0.30702 | 1.2667 | 2.5479 | 99.54 | 20.88 | 37.299 | 101.3334 | 105.3593 | 93.6671 | 0.009261 | 0.96246 | 0.022943 | 23.838 | 0.029714 | 0.15516 | 0.022802 | 1.563 | 1.563 | 0.00408
2019-03-21 11:59 | 1553183924.3637 | 811 | 0.20544 | 0.003632 | 8.7e-05 | 100.4325 | 99.9065 | 114.1436 | 21.1132 | 0.003301 | 0.002978 | 148.6568 | 0.26192 | 8.4002 | 0.003472 | 0.004793 | 0.004698 | 5.6e-05 | 0.000877 | 0.33356 | 1.2907 | 2.5478 | 99.54 | 20.88 | 37.299 | 101.1248 | 105.274 | 93.3209 | 0.009261 | 0.96337 | 0.022938 | 23.8099 | 0.046977 | 0.15582 | 0.02079 | 1.4246 | 1.4246 | 0.004495
2019-03-21 11:59 | 1553183925.3635 | 812 | 0.20569 | 0.003629 | 8.6e-05 | 100.4895 | 100.105 | 114.1039 | 21.1137 | 0.003238 | 0.003 | 148.6024 | 0.26253 | 8.4031 | 0.003562 | 0.005694 | 0.005554 | 5.6e-05 | 0.000881 | 0.31885 | 1.3617 | 2.5477 | 99.54 | 20.88 | 37.299 | 99.5982 | 107.4555 | 93.8635 | 0.009261 | 0.96194 | 0.022917 | 23.8233 | 0.034284 | 0.15618 | 0.019282 | 1.3199 | 1.3199 | 0.004863
2019-03-21 11:59 | 1553183926.3645 | 813 | 0.2051 | 0.003638 | 8.7e-05 | 100.4281 | 100.0941 | 113.9028 | 21.113 | 0.003251 | 0.002979 | 148.5619 | 0.26216 | 8.3997 | 0.00356 | 0.004423 | 0.004296 | 5.7e-05 | 0.00088 | 0.3131 | 1.3095 | 2.5479 | 99.54 | 20.88 | 37.299 | 99.906 | 107.4123 | 93.3495 | 0.009261 | 0.96329 | 0.022912 | 23.7857 | 0.029793 | 0.15596 | 0.02279 | 1.5601 | 1.5601 | 0.004109
2019-03-21 11:59 | 1553183927.3639 | 814 | 0.20566 | 0.003641 | 8.7e-05 | 100.4683 | 100.1258 | 113.703 | 21.1134 | 0.003344 | 0.003013 | 148.5173 | 0.26201 | 8.4012 | 0.003522 | 0.00486 | 0.004818 | 5.8e-05 | 0.000883 | 0.30541 | 1.2847 | 2.5479 | 99.54 | 20.88 | 37.299 | 102.1654 | 106.4767 | 94.1738 | 0.009261 | 0.96113 | 0.022867 | 23.7915 | 0.030545 | 0.15587 | 0.019418 | 1.3264 | 1.3264 | 0.00483
2019-03-21 11:59 | 1553183928.3636 | 815 | 0.20509 | 0.003646 | 8.7e-05 | 100.4415 | 100.1204 | 113.3409 | 21.1129 | 0.003242 | 0.002975 | 148.4419 | 0.26128 | 8.4027 | 0.003492 | 0.00506 | 0.004946 | 5.7e-05 | 0.000882 | 0.31332 | 1.308 | 2.548 | 99.54 | 20.88 | 37.299 | 99.7095 | 105.7623 | 93.2624 | 0.009261 | 0.96352 | 0.022859 | 23.7241 | 0.028644 | 0.15543 | 0.022878 | 1.5624 | 1.5624 | 0.004089
2019-03-21 11:59 | 1553183929.3636 | 816 | 0.20527 | 0.003648 | 8.7e-05 | 100.4555 | 100.0781 | 113.2801 | 21.114 | 0.003295 | 0.00299 | 148.3946 | 0.26128 | 8.4028 | 0.003574 | 0.005304 | 0.005339 | 6e-05 | 0.000879 | 0.33212 | 1.3241 | 2.548 | 99.54 | 20.88 | 37.3 | 100.9758 | 107.7528 | 93.6133 | 0.009261 | 0.9626 | 0.022842 | 23.7291 | 0.033665 | 0.15543 | 0.021806 | 1.4881 | 1.4881 | 0.004293
2019-03-21 11:59 | 1553183930.3651 | 817 | 0.20538 | 0.003658 | 8.7e-05 | 100.5473 | 100.057 | 113.3824 | 21.1141 | 0.003223 | 0.003022 | 148.3682 | 0.26254 | 8.4027 | 0.003487 | 0.005092 | 0.004989 | 5.4e-05 | 0.000887 | 0.27505 | 1.3539 | 2.5481 | 99.54 | 20.88 | 37.3 | 99.2354 | 105.6422 | 94.3959 | 0.009261 | 0.96055 | 0.022828 | 23.7651 | 0.043719 | 0.15619 | 0.021141 | 1.4417 | 1.4417 | 0.004453
2019-03-21 11:59 | 1553183931.3642 | 818 | 0.20594 | 0.003666 | 8.8e-05 | 100.4752 | 100.1319 | 113.3758 | 21.1149 | 0.003321 | 0.003006 | 148.3253 | 0.2621 | 8.4034 | 0.003502 | 0.00497 | 0.004916 | 6e-05 | 0.00088 | 0.31564 | 1.3298 | 2.5481 | 99.54 | 20.88 | 37.3 | 101.6014 | 106.0034 | 94.0086 | 0.009261 | 0.96156 | 0.022839 | 23.7519 | 0.030607 | 0.15593 | 0.017744 | 1.2103 | 1.2103 | 0.005295
2019-03-21 11:59 | 1553183932.3638 | 819 | 0.20539 | 0.003676 | 8.9e-05 | 100.5607 | 100.151 | 113.4134 | 21.1153 | 0.003323 | 0.00305 | 148.2508 | 0.26134 | 8.4017 | 0.003417 | 0.005086 | 0.005033 | 5.5e-05 | 0.000879 | 0.29391 | 1.3152 | 2.5479 | 99.54 | 20.88 | 37.301 | 101.6522 | 103.9432 | 95.061 | 0.009262 | 0.95881 | 0.02281 | 23.7899 | 0.036512 | 0.15547 | 0.021006 | 1.4314 | 1.4314 | 0.004464
2019-03-21 11:59 | 1553183933.3648 | 820 | 0.20518 | 0.003679 | 8.9e-05 | 100.5833 | 100.2801 | 113.6997 | 21.1145 | 0.003357 | 0.003002 | 148.1854 | 0.26154 | 8.401 | 0.003384 | 0.005317 | 0.005127 | 5.7e-05 | 0.000881 | 0.31404 | 1.3143 | 2.5479 | 99.54 | 20.87 | 37.301 | 102.4744 | 103.1297 | 93.897 | 0.009256 | 0.96185 | 0.022875 | 23.7822 | 0.027005 | 0.15559 | 0.022308 | 1.5245 | 1.5245 | 0.004195
2019-03-21 11:59 | 1553183934.3652 | 821 | 0.20507 | 0.003694 | 9e-05 | 100.5169 | 100.2082 | 113.7413 | 21.1146 | 0.003332 | 0.003006 | 148.115 | 0.26182 | 8.4045 | 0.003458 | 0.004738 | 0.004603 | 6e-05 | 0.000881 | 0.30973 | 1.3135 | 2.548 | 99.54 | 20.87 | 37.303 | 101.8863 | 104.9323 | 94.009 | 0.009256 | 0.96156 | 0.022876 | 23.7902 | 0.027507 | 0.15576 | 0.02295 | 1.5684 | 1.5684 | 0.004082
2019-03-21 11:59 | 1553183935.363 | 822 | 0.20466 | 0.003708 | 9e-05 | 100.6197 | 100.1776 | 113.6425 | 21.1152 | 0.003336 | 0.003039 | 148.0347 | 0.26124 | 8.4019 | 0.003369 | 0.005154 | 0.005044 | 5.6e-05 | 0.000883 | 0.31006 | 1.2946 | 2.5479 | 99.54 | 20.88 | 37.307 | 101.981 | 102.7814 | 94.7914 | 0.009263 | 0.95952 | 0.022841 | 23.8052 | 0.039386 | 0.15541 | 0.025353 | 1.7301 | 1.7301 | 0.003692
2019-03-21 11:59 | 1553183936.363 | 823 | 0.2051 | 0.003718 | 9.1e-05 | 100.6068 | 100.137 | 113.673 | 21.1143 | 0.003358 | 0.00299 | 147.9493 | 0.26223 | 8.403 | 0.003489 | 0.004895 | 0.004834 | 5.6e-05 | 0.000878 | 0.33445 | 1.3276 | 2.5479 | 99.54 | 20.88 | 37.31 | 102.5088 | 105.6753 | 93.6262 | 0.009264 | 0.96256 | 0.022881 | 23.7706 | 0.041866 | 0.156 | 0.022703 | 1.5517 | 1.5517 | 0.004132
2019-03-21 11:59 | 1553183937.3645 | 824 | 0.20514 | 0.003732 | 9.1e-05 | 100.5911 | 100.18 | 113.7305 | 21.1142 | 0.003314 | 0.003049 | 147.8743 | 0.26119 | 8.4037 | 0.003518 | 0.004963 | 0.004862 | 5.8e-05 | 0.000879 | 0.30198 | 1.3054 | 2.548 | 99.54 | 20.88 | 37.313 | 101.4375 | 106.3916 | 95.0395 | 0.009265 | 0.95887 | 0.022843 | 23.8224 | 0.03663 | 0.15538 | 0.022464 | 1.5329 | 1.5329 | 0.004166
2019-03-21 11:59 | 1553183938.3644 | 825 | 0.20528 | 0.003748 | 9e-05 | 100.6176 | 100.2961 | 113.8934 | 21.1147 | 0.003322 | 0.003067 | 147.8065 | 0.26201 | 8.4019 | 0.003488 | 0.005135 | 0.004981 | 5.8e-05 | 0.000882 | 0.28871 | 1.3039 | 2.548 | 99.54 | 20.88 | 37.32 | 101.6335 | 105.6453 | 95.4613 | 0.009266 | 0.95777 | 0.022846 | 23.8531 | 0.028624 | 0.15587 | 0.021624 | 1.4757 | 1.4757 | 0.004341
2019-03-21 11:59 | 1553183939.3643 | 826 | 0.20541 | 0.003759 | 9e-05 | 100.6229 | 100.2205 | 113.9615 | 21.1151 | 0.003341 | 0.003068 | 147.7082 | 0.26113 | 8.401 | 0.003481 | 0.00495 | 0.004799 | 5.9e-05 | 0.00088 | 0.29999 | 1.3143 | 2.548 | 99.54 | 20.88 | 37.324 | 102.0963 | 105.4984 | 95.4938 | 0.009267 | 0.95769 | 0.022852 | 23.8613 | 0.035846 | 0.15535 | 0.020787 | 1.4189 | 1.4189 | 0.0045
2019-03-21 11:59 | 1553183940.3642 | 827 | 0.20558 | 0.003776 | 8.9e-05 | 100.6334 | 100.0645 | 113.8572 | 21.1154 | 0.003362 | 0.003025 | 147.639 | 0.26194 | 8.4027 | 0.003524 | 0.004942 | 0.004916 | 5.5e-05 | 0.000875 | 0.32173 | 1.3104 | 2.5481 | 99.54 | 20.88 | 37.334 | 102.6041 | 106.5391 | 94.4568 | 0.00927 | 0.96039 | 0.022873 | 23.8168 | 0.050706 | 0.15583 | 0.019737 | 1.3484 | 1.3484 | 0.00475
2019-03-21 11:59 | 1553183941.3636 | 828 | 0.20498 | 0.003793 | 8.9e-05 | 100.6208 | 100.2825 | 113.5876 | 21.1147 | 0.003421 | 0.002976 | 147.5616 | 0.26166 | 8.403 | 0.003477 | 0.005089 | 0.004929 | 5.7e-05 | 0.000879 | 0.34468 | 1.3333 | 2.5479 | 99.54 | 20.88 | 37.34 | 104.0299 | 105.3957 | 93.276 | 0.009271 | 0.96348 | 0.022883 | 23.7503 | 0.03013 | 0.15566 | 0.023326 | 1.5947 | 1.5947 | 0.004012
2019-03-21 11:59 | 1553183942.3643 | 829 | 0.20537 | 0.003809 | 8.9e-05 | 100.6804 | 100.134 | 113.2524 | 21.1149 | 0.003351 | 0.003051 | 147.4976 | 0.26171 | 8.4033 | 0.003468 | 0.005651 | 0.005585 | 5.7e-05 | 0.000875 | 0.32781 | 1.3139 | 2.5478 | 99.54 | 20.88 | 37.345 | 102.3366 | 105.1639 | 95.0855 | 0.009273 | 0.95875 | 0.022793 | 23.7738 | 0.048681 | 0.15569 | 0.020983 | 1.4287 | 1.4287 | 0.004479
2019-03-21 11:59 | 1553183943.3634 | 830 | 0.20525 | 0.003819 | 8.9e-05 | 100.7043 | 100.2348 | 112.9461 | 21.1147 | 0.003324 | 0.003052 | 147.426 | 0.26173 | 8.4023 | 0.003525 | 0.005309 | 0.005256 | 5.8e-05 | 0.000874 | 0.30574 | 1.3067 | 2.5479 | 99.54 | 20.89 | 37.355 | 101.6845 | 106.5488 | 95.1018 | 0.009281 | 0.95871 | 0.022762 | 23.7421 | 0.041797 | 0.1557 | 0.02171 | 1.4762 | 1.4762 | 0.004335
2019-03-21 11:59 | 1553183944.3644 | 831 | 0.20512 | 0.003835 | 8.9e-05 | 100.6733 | 100.3707 | 112.7069 | 21.1155 | 0.003287 | 0.002982 | 147.3676 | 0.26205 | 8.4052 | 0.003602 | 0.004945 | 0.004812 | 5.9e-05 | 0.000872 | 0.31653 | 1.317 | 2.548 | 99.54 | 20.89 | 37.36 | 100.7864 | 108.4318 | 93.4145 | 0.009282 | 0.96312 | 0.02279 | 23.6626 | 0.026926 | 0.15589 | 0.02245 | 1.5285 | 1.5285 | 0.004192
2019-03-21 11:59 | 1553183945.3637 | 832 | 0.20506 | 0.003855 | 8.9e-05 | 100.6441 | 100.23 | 112.4381 | 21.1159 | 0.003314 | 0.003027 | 147.3195 | 0.26195 | 8.4019 | 0.003575 | 0.004928 | 0.00477 | 6e-05 | 0.000877 | 0.2975 | 1.3376 | 2.548 | 99.54 | 20.89 | 37.367 | 101.4423 | 107.7601 | 94.5183 | 0.009284 | 0.96023 | 0.022728 | 23.6699 | 0.036875 | 0.15583 | 0.022811 | 1.5488 | 1.5488 | 0.004135
2019-03-21 11:59 | 1553183946.3635 | 833 | 0.20505 | 0.003873 | 8.9e-05 | 100.6664 | 100.3642 | 112.3057 | 21.1148 | 0.00333 | 0.003048 | 147.2646 | 0.26039 | 8.4034 | 0.003526 | 0.005244 | 0.00517 | 5.9e-05 | 0.000873 | 0.29159 | 1.3172 | 2.548 | 99.54 | 20.89 | 37.376 | 101.8196 | 106.578 | 95.0096 | 0.009286 | 0.95895 | 0.0227 | 23.6717 | 0.026894 | 0.1549 | 0.022848 | 1.5494 | 1.5494 | 0.004109
2019-03-21 11:59 | 1553183947.3636 | 834 | 0.20507 | 0.003896 | 8.9e-05 | 100.6602 | 100.2477 | 112.3515 | 21.1155 | 0.003338 | 0.003065 | 147.2192 | 0.26107 | 8.4026 | 0.003551 | 0.004582 | 0.004582 | 6.2e-05 | 0.000875 | 0.32333 | 1.3005 | 2.548 | 99.54 | 20.89 | 37.379 | 102.0261 | 107.1975 | 95.4314 | 0.009287 | 0.95785 | 0.022692 | 23.6901 | 0.036725 | 0.15531 | 0.022662 | 1.5362 | 1.5362 | 0.004155
2019-03-21 11:59 | 1553183948.3636 | 835 | 0.20505 | 0.003905 | 8.9e-05 | 100.7326 | 100.1857 | 112.3456 | 21.1148 | 0.003369 | 0.003027 | 147.1488 | 0.26055 | 8.4029 | 0.00351 | 0.004931 | 0.004763 | 6e-05 | 0.000875 | 0.32101 | 1.3 | 2.5479 | 99.54 | 20.89 | 37.384 | 102.7621 | 106.1852 | 94.4977 | 0.009288 | 0.96028 | 0.02272 | 23.6595 | 0.048688 | 0.155 | 0.022801 | 1.5475 | 1.5475 | 0.004117
2019-03-21 11:59 | 1553183949.3647 | 836 | 0.20559 | 0.003923 | 9e-05 | 100.6792 | 100.1748 | 112.3008 | 21.1148 | 0.003383 | 0.003054 | 147.091 | 0.26157 | 8.4036 | 0.003503 | 0.004658 | 0.004602 | 5.7e-05 | 0.000872 | 0.31205 | 1.3327 | 2.5478 | 99.54 | 20.89 | 37.385 | 103.1198 | 106.0222 | 95.167 | 0.009288 | 0.95854 | 0.022695 | 23.6763 | 0.044925 | 0.15561 | 0.019495 | 1.3214 | 1.3214 | 0.00484
2019-03-21 11:59 | 1553183950.3628 | 837 | 0.20626 | 0.003945 | 9.1e-05 | 100.7227 | 100.2433 | 112.2046 | 21.1155 | 0.003377 | 0.003083 | 147.008 | 0.26141 | 8.4028 | 0.003528 | 0.004427 | 0.004319 | 5.8e-05 | 0.000874 | 0.31437 | 1.31 | 2.5477 | 99.54 | 20.89 | 37.386 | 102.964 | 106.6267 | 95.863 | 0.009288 | 0.95673 | 0.022663 | 23.6885 | 0.042678 | 0.15551 | 0.015464 | 1.0461 | 1.0461 | 0.00611
2019-03-21 11:59 | 1553183951.3628 | 838 | 0.20591 | 0.003966 | 9.1e-05 | 100.7427 | 100.309 | 112.1454 | 21.1162 | 0.003365 | 0.003012 | 146.9186 | 0.26145 | 8.4012 | 0.003508 | 0.00468 | 0.0046 | 5.9e-05 | 0.000871 | 0.27776 | 1.2803 | 2.548 | 99.54 | 20.89 | 37.385 | 102.6794 | 106.1368 | 94.1459 | 0.009288 | 0.9612 | 0.02271 | 23.6271 | 0.038587 | 0.15554 | 0.017569 | 1.1913 | 1.1913 | 0.005366
2019-03-21 11:59 | 1553183952.3633 | 839 | 0.20481 | 0.003986 | 9.1e-05 | 100.8047 | 100.5383 | 112.3806 | 21.1155 | 0.003386 | 0.003058 | 146.8557 | 0.26173 | 8.4025 | 0.003543 | 0.005075 | 0.004988 | 5.5e-05 | 0.000873 | 0.30989 | 1.2951 | 2.548 | 99.54 | 20.89 | 37.384 | 103.1731 | 106.9888 | 95.2616 | 0.009288 | 0.95829 | 0.0227 | 23.6877 | 0.023672 | 0.1557 | 0.024128 | 1.6362 | 1.6362 | 0.003911
2019-03-21 11:59 | 1553183953.3629 | 840 | 0.20458 | 0.004008 | 9.1e-05 | 100.7147 | 100.3547 | 112.6116 | 21.1155 | 0.003333 | 0.003065 | 146.7871 | 0.26098 | 8.4054 | 0.003499 | 0.004132 | 0.004039 | 6.1e-05 | 0.000871 | 0.3366 | 1.2774 | 2.548 | 99.54 | 20.89 | 37.379 | 101.898 | 105.9348 | 95.4189 | 0.009287 | 0.95788 | 0.022718 | 23.7171 | 0.032031 | 0.15526 | 0.02545 | 1.7273 | 1.7273 | 0.003694
2019-03-21 11:59 | 1553183954.363 | 841 | 0.20454 | 0.00403 | 9.2e-05 | 100.7969 | 100.4163 | 112.639 | 21.1159 | 0.003351 | 0.003069 | 146.7038 | 0.26141 | 8.4028 | 0.003475 | 0.005223 | 0.005122 | 5.8e-05 | 0.000878 | 0.3366 | 1.3046 | 2.548 | 99.54 | 20.9 | 37.376 | 102.3274 | 105.3467 | 95.5236 | 0.009292 | 0.95761 | 0.022718 | 23.7234 | 0.033839 | 0.15551 | 0.025672 | 1.7423 | 1.7423 | 0.003668
2019-03-21 11:59 | 1553183955.364 | 842 | 0.20371 | 0.004052 | 9.2e-05 | 100.712 | 100.4923 | 112.6557 | 21.1157 | 0.003364 | 0.003063 | 146.6267 | 0.26165 | 8.4028 | 0.003564 | 0.004974 | 0.004882 | 5.8e-05 | 0.000873 | 0.30989 | 1.3151 | 2.5479 | 99.54 | 20.9 | 37.373 | 102.6601 | 107.4907 | 95.3761 | 0.009291 | 0.95799 | 0.022724 | 23.7204 | 0.019538 | 0.15566 | 0.03063 | 2.0792 | 2.0792 | 0.003077
2019-03-21 11:59 | 1553183956.364 | 843 | 0.20411 | 0.004076 | 9.3e-05 | 100.7323 | 100.4361 | 112.3656 | 21.1161 | 0.003395 | 0.00313 | 146.5587 | 0.26079 | 8.4044 | 0.003476 | 0.004822 | 0.004698 | 5.9e-05 | 0.000874 | 0.31564 | 1.2923 | 2.5478 | 99.54 | 20.9 | 37.367 | 103.4016 | 105.3609 | 96.9947 | 0.009289 | 0.9538 | 0.022645 | 23.7418 | 0.02634 | 0.15515 | 0.028181 | 1.9064 | 1.9064 | 0.003345
2019-03-21 11:59 | 1553183957.3645 | 844 | 0.20447 | 0.004093 | 9.4e-05 | 100.7639 | 100.2804 | 112.1163 | 21.116 | 0.003421 | 0.003108 | 146.457 | 0.26236 | 8.4016 | 0.003515 | 0.005041 | 0.004962 | 5.9e-05 | 0.00087 | 0.30685 | 1.282 | 2.5478 | 99.54 | 20.9 | 37.364 | 104.0451 | 106.3176 | 96.4626 | 0.009289 | 0.95518 | 0.022636 | 23.6984 | 0.043017 | 0.15608 | 0.026039 | 1.7607 | 1.7607 | 0.003643
2019-03-21 11:59 | 1553183958.3631 | 845 | 0.20452 | 0.004117 | 9.4e-05 | 100.7862 | 100.2627 | 112.0311 | 21.1169 | 0.003382 | 0.003136 | 146.3941 | 0.26201 | 8.4029 | 0.003599 | 0.005058 | 0.004993 | 5.9e-05 | 0.000873 | 0.33229 | 1.2856 | 2.5479 | 99.54 | 20.9 | 37.356 | 103.0872 | 108.352 | 97.1326 | 0.009287 | 0.95345 | 0.022607 | 23.7109 | 0.04658 | 0.15587 | 0.025691 | 1.7349 | 1.7349 | 0.003693
2019-03-21 11:59 | 1553183959.3631 | 846 | 0.20425 | 0.004134 | 9.5e-05 | 100.8017 | 100.3547 | 112.1914 | 21.1163 | 0.003359 | 0.003142 | 146.3504 | 0.2624 | 8.4028 | 0.003586 | 0.004815 | 0.004713 | 6.2e-05 | 0.000871 | 0.3115 | 1.3057 | 2.5479 | 99.54 | 20.9 | 37.353 | 102.5311 | 108.0458 | 97.2871 | 0.009286 | 0.95305 | 0.022619 | 23.7328 | 0.039755 | 0.1561 | 0.027268 | 1.8423 | 1.8423 | 0.003482
2019-03-21 11:59 | 1553183960.3636 | 847 | 0.20413 | 0.004156 | 9.5e-05 | 100.7963 | 100.1977 | 112.3031 | 21.1162 | 0.003467 | 0.003083 | 146.2956 | 0.26223 | 8.401 | 0.003453 | 0.005704 | 0.005654 | 6.2e-05 | 0.000873 | 0.33035 | 1.3032 | 2.5478 | 99.54 | 20.9 | 37.35 | 105.1567 | 104.8097 | 95.8661 | 0.009285 | 0.95672 | 0.022673 | 23.699 | 0.053284 | 0.156 | 0.027994 | 1.8959 | 1.8959 | 0.003382
2019-03-21 11:59 | 1553183961.3633 | 848 | 0.2044 | 0.004169 | 9.6e-05 | 100.7913 | 100.3833 | 112.497 | 21.1152 | 0.003388 | 0.003062 | 146.2232 | 0.26178 | 8.4007 | 0.003575 | 0.004597 | 0.004483 | 5.9e-05 | 0.000872 | 0.31581 | 1.2576 | 2.5479 | 99.54 | 20.9 | 37.343 | 103.2348 | 107.7657 | 95.3575 | 0.009283 | 0.95804 | 0.022709 | 23.7031 | 0.03628 | 0.15573 | 0.026357 | 1.7878 | 1.7878 | 0.00358
2019-03-21 11:59 | 1553183962.3634 | 849 | 0.20432 | 0.004185 | 9.6e-05 | 100.7701 | 100.4276 | 112.3503 | 21.1162 | 0.003403 | 0.00307 | 146.1475 | 0.26197 | 8.4013 | 0.003601 | 0.0051 | 0.005025 | 5.8e-05 | 0.000872 | 0.30558 | 1.2768 | 2.5479 | 99.54 | 20.9 | 37.34 | 103.5942 | 108.4027 | 95.5521 | 0.009283 | 0.95754 | 0.022688 | 23.6939 | 0.030458 | 0.15584 | 0.026823 | 1.8178 | 1.8178 | 0.003524
2019-03-21 11:59 | 1553183963.3634 | 850 | 0.20478 | 0.004206 | 9.6e-05 | 100.7925 | 100.4323 | 112.1327 | 21.1155 | 0.003436 | 0.003141 | 146.0933 | 0.26136 | 8.4032 | 0.003625 | 0.005989 | 0.005804 | 6.1e-05 | 0.000872 | 0.31454 | 1.3133 | 2.5479 | 99.54 | 20.9 | 37.338 | 104.4033 | 108.9709 | 97.2581 | 0.009282 | 0.95312 | 0.022614 | 23.7256 | 0.032016 | 0.15548 | 0.024009 | 1.6216 | 1.6216 | 0.003941
2019-03-21 11:59 | 1553183964.3636 | 851 | 0.20462 | 0.004225 | 9.6e-05 | 100.7978 | 100.2889 | 112.0177 | 21.1173 | 0.003397 | 0.003092 | 146.0122 | 0.26184 | 8.4034 | 0.003611 | 0.005418 | 0.005373 | 5.8e-05 | 0.000875 | 0.32565 | 1.3145 | 2.5478 | 99.54 | 20.9 | 37.333 | 103.4577 | 108.651 | 96.0755 | 0.009281 | 0.95618 | 0.022638 | 23.6756 | 0.045272 | 0.15577 | 0.024975 | 1.6888 | 1.6888 | 0.003791
2019-03-21 11:59 | 1553183965.3649 | 852 | 0.20428 | 0.004244 | 9.6e-05 | 100.8643 | 100.3032 | 112.1884 | 21.1162 | 0.0034 | 0.003102 | 145.9287 | 0.26235 | 8.4019 | 0.003589 | 0.005654 | 0.005556 | 5.9e-05 | 0.000876 | 0.32947 | 1.3353 | 2.5478 | 99.54 | 20.9 | 37.33 | 103.5306 | 108.104 | 96.3076 | 0.00928 | 0.95558 | 0.022648 | 23.701 | 0.04989 | 0.15607 | 0.026959 | 1.8238 | 1.8238 | 0.003517
2019-03-21 11:59 | 1553183966.3637 | 853 | 0.20387 | 0.004265 | 9.6e-05 | 100.8205 | 100.2212 | 112.2853 | 21.1171 | 0.00343 | 0.003192 | 145.843 | 0.26113 | 8.4005 | 0.003598 | 0.005235 | 0.00514 | 5.8e-05 | 0.000874 | 0.3533 | 1.3075 | 2.5479 | 99.54 | 20.9 | 37.325 | 104.2459 | 108.3246 | 98.5021 | 0.009279 | 0.94993 | 0.022591 | 23.7816 | 0.053328 | 0.15535 | 0.029436 | 1.9863 | 1.9863 | 0.003214
2019-03-21 11:59 | 1553183967.3626 | 854 | 0.20352 | 0.004286 | 9.7e-05 | 100.8127 | 100.4221 | 112.1855 | 21.117 | 0.003381 | 0.003106 | 145.7601 | 0.261 | 8.4024 | 0.003652 | 0.005029 | 0.004948 | 6e-05 | 0.000875 | 0.3147 | 1.3071 | 2.5479 | 99.54 | 20.9 | 37.322 | 103.0652 | 109.6361 | 96.422 | 0.009278 | 0.95528 | 0.022644 | 23.7044 | 0.034724 | 0.15527 | 0.031484 | 2.1294 | 2.1294 | 0.002997
2019-03-21 11:59 | 1553183968.3641 | 855 | 0.20428 | 0.004305 | 9.7e-05 | 100.8112 | 100.2712 | 111.9105 | 21.1171 | 0.003457 | 0.003128 | 145.6979 | 0.26202 | 8.4028 | 0.003603 | 0.00511 | 0.004998 | 5.7e-05 | 0.000872 | 0.30718 | 1.3036 | 2.548 | 99.54 | 20.9 | 37.321 | 104.9081 | 108.4538 | 96.951 | 0.009278 | 0.95392 | 0.0226 | 23.6923 | 0.048041 | 0.15588 | 0.026894 | 1.8155 | 1.8155 | 0.003529
2019-03-21 11:59 | 1553183969.363 | 856 | 0.20428 | 0.004328 | 9.9e-05 | 100.875 | 100.165 | 111.6516 | 21.1174 | 0.003407 | 0.003083 | 145.6153 | 0.26214 | 8.4031 | 0.003685 | 0.004356 | 0.00433 | 5.7e-05 | 0.000878 | 0.31813 | 1.2769 | 2.5481 | 99.54 | 20.91 | 37.316 | 103.6882 | 110.4307 | 95.8641 | 0.009282 | 0.95673 | 0.022608 | 23.6301 | 0.06318 | 0.15594 | 0.026854 | 1.8133 | 1.8133 | 0.003535
2019-03-21 11:59 | 1553183970.3638 | 857 | 0.20471 | 0.004348 | 9.9e-05 | 100.8392 | 100.1571 | 111.4898 | 21.117 | 0.003417 | 0.003163 | 145.5067 | 0.26104 | 8.4018 | 0.003635 | 0.00418 | 0.0041 | 5.7e-05 | 0.000869 | 0.30182 | 1.3012 | 2.5483 | 99.54 | 20.91 | 37.313 | 103.94 | 109.216 | 97.8027 | 0.009282 | 0.95173 | 0.022532 | 23.6749 | 0.060709 | 0.15529 | 0.024249 | 1.6318 | 1.6318 | 0.003911
2019-03-21 12:00 | 1553183971.3639 | 858 | 0.20523 | 0.004361 | 0.0001 | 100.7823 | 100.1265 | 111.1746 | 21.1165 | 0.003467 | 0.003138 | 145.4287 | 0.26186 | 8.4022 | 0.003586 | 0.005212 | 0.005096 | 5.6e-05 | 0.000873 | 0.2808 | 1.3173 | 2.548 | 99.54 | 20.91 | 37.308 | 105.1584 | 108.0333 | 97.1996 | 0.00928 | 0.95328 | 0.022518 | 23.6222 | 0.058398 | 0.15578 | 0.02112 | 1.4201 | 1.4201 | 0.004509
2019-03-21 12:00 | 1553183972.3633 | 859 | 0.20472 | 0.004375 | 0.0001 | 100.7996 | 100.3012 | 111.145 | 21.1175 | 0.003384 | 0.003171 | 145.3619 | 0.26197 | 8.4024 | 0.003604 | 0.004815 | 0.004772 | 5.7e-05 | 0.00087 | 0.3178 | 1.3176 | 2.548 | 99.54 | 20.91 | 37.306 | 103.1427 | 108.4707 | 97.9873 | 0.00928 | 0.95125 | 0.022492 | 23.6441 | 0.044337 | 0.15585 | 0.024193 | 1.625 | 1.625 | 0.003942
2019-03-21 12:00 | 1553183973.3644 | 860 | 0.20459 | 0.00439 | 0.0001 | 100.8598 | 100.1466 | 110.807 | 21.1173 | 0.003446 | 0.00315 | 145.2748 | 0.26143 | 8.4045 | 0.003624 | 0.004766 | 0.004658 | 5.2e-05 | 0.000872 | 0.30469 | 1.3136 | 2.5482 | 99.54 | 20.91 | 37.304 | 104.6465 | 108.9639 | 97.4847 | 0.009279 | 0.95254 | 0.022473 | 23.5922 | 0.063476 | 0.15553 | 0.024969 | 1.6758 | 1.6758 | 0.003814
2019-03-21 12:00 | 1553183974.363 | 861 | 0.20448 | 0.004401 | 0.0001 | 100.8255 | 100.2099 | 111.0631 | 21.1169 | 0.00351 | 0.003135 | 145.2174 | 0.26138 | 8.4015 | 0.003627 | 0.004873 | 0.00478 | 5.7e-05 | 0.00087 | 0.29927 | 1.2843 | 2.5482 | 99.54 | 20.91 | 37.299 | 106.2038 | 109.0361 | 97.1131 | 0.009278 | 0.9535 | 0.02251 | 23.6076 | 0.05478 | 0.1555 | 0.025593 | 1.7205 | 1.7205 | 0.003715
2019-03-21 12:00 | 1553183975.3645 | 862 | 0.20434 | 0.004412 | 0.0001 | 100.8369 | 100.1691 | 111.0777 | 21.1178 | 0.003379 | 0.003174 | 145.1633 | 0.26231 | 8.4029 | 0.003578 | 0.00504 | 0.005011 | 5.3e-05 | 0.000872 | 0.31382 | 1.2871 | 2.5479 | 99.54 | 20.91 | 37.296 | 103.0244 | 107.8486 | 98.0683 | 0.009277 | 0.95105 | 0.022482 | 23.6396 | 0.059433 | 0.15604 | 0.026411 | 1.7734 | 1.7734 | 0.003617
2019-03-21 12:00 | 1553183976.3631 | 863 | 0.20507 | 0.004429 | 0.0001 | 100.8535 | 100.3652 | 110.977 | 21.1175 | 0.003487 | 0.003193 | 145.1067 | 0.26165 | 8.4029 | 0.003575 | 0.00435 | 0.004202 | 5.5e-05 | 0.000872 | 0.32676 | 1.2829 | 2.5477 | 99.54 | 20.91 | 37.294 | 105.6339 | 107.7602 | 98.5078 | 0.009277 | 0.94992 | 0.022459 | 23.6428 | 0.043413 | 0.15566 | 0.022001 | 1.4755 | 1.4755 | 0.004336
2019-03-21 12:00 | 1553183977.3641 | 864 | 0.2045 | 0.004445 | 0.0001 | 100.8303 | 100.2052 | 110.9577 | 21.117 | 0.003514 | 0.003186 | 145.0425 | 0.26173 | 8.4035 | 0.003551 | 0.005094 | 0.004938 | 5.5e-05 | 0.00087 | 0.30071 | 1.3003 | 2.5478 | 99.54 | 20.91 | 37.29 | 106.2983 | 107.1927 | 98.3435 | 0.009276 | 0.95034 | 0.022462 | 23.6355 | 0.055629 | 0.15571 | 0.025404 | 1.7042 | 1.7042 | 0.003755
2019-03-21 12:00 | 1553183978.3633 | 865 | 0.20416 | 0.004462 | 0.0001 | 100.8124 | 100.3833 | 111.2328 | 21.1182 | 0.00347 | 0.003193 | 144.9662 | 0.26177 | 8.4023 | 0.003583 | 0.005087 | 0.004942 | 5.7e-05 | 0.000869 | 0.32333 | 1.3102 | 2.5481 | 99.54 | 20.91 | 37.287 | 105.2245 | 107.9546 | 98.5157 | 0.009275 | 0.9499 | 0.022484 | 23.6703 | 0.038157 | 0.15573 | 0.027432 | 1.8421 | 1.8421 | 0.003474
2019-03-21 12:00 | 1553183979.3632 | 866 | 0.2035 | 0.004474 | 9.9e-05 | 100.8288 | 100.3039 | 111.1467 | 21.1182 | 0.003489 | 0.003165 | 144.8808 | 0.26152 | 8.4023 | 0.003675 | 0.005826 | 0.005766 | 5.5e-05 | 0.000871 | 0.31653 | 1.3653 | 2.5479 | 99.54 | 20.91 | 37.282 | 105.6778 | 110.2003 | 97.8487 | 0.009274 | 0.95161 | 0.022496 | 23.6399 | 0.046684 | 0.15558 | 0.031395 | 2.1093 | 2.1093 | 0.003031
2019-03-21 12:00 | 1553183980.3637 | 867 | 0.20347 | 0.004489 | 9.9e-05 | 100.7966 | 100.073 | 111.0699 | 21.1185 | 0.003521 | 0.003127 | 144.8181 | 0.26187 | 8.4035 | 0.003674 | 0.005369 | 0.005305 | 5.6e-05 | 0.000873 | 0.29894 | 1.4015 | 2.5482 | 99.54 | 20.91 | 37.28 | 106.454 | 110.1794 | 96.9289 | 0.009274 | 0.95397 | 0.022516 | 23.6024 | 0.064445 | 0.15579 | 0.031531 | 2.1204 | 2.1204 | 0.00302
2019-03-21 12:00 | 1553183981.3634 | 868 | 0.20415 | 0.0045 | 0.0001 | 100.7919 | 100.3901 | 111.4811 | 21.1172 | 0.003421 | 0.003235 | 144.7569 | 0.26131 | 8.4038 | 0.00366 | 0.004831 | 0.004636 | 5.3e-05 | 0.000872 | 0.31188 | 1.3882 | 2.5482 | 99.54 | 20.91 | 37.279 | 104.0259 | 109.8429 | 99.5365 | 0.009273 | 0.9473 | 0.022479 | 23.7292 | 0.035726 | 0.15545 | 0.027427 | 1.8414 | 1.8414 | 0.00347
2019-03-21 12:00 | 1553183982.3635 | 869 | 0.20461 | 0.004516 | 9.9e-05 | 100.8115 | 100.2627 | 111.6722 | 21.1171 | 0.003486 | 0.00316 | 144.6867 | 0.2606 | 8.4044 | 0.003726 | 0.005783 | 0.005688 | 5.3e-05 | 0.000872 | 0.30917 | 1.429 | 2.548 | 99.54 | 20.91 | 37.276 | 105.6066 | 111.4295 | 97.7111 | 0.009273 | 0.95196 | 0.022553 | 23.6913 | 0.048827 | 0.15503 | 0.024642 | 1.6598 | 1.6598 | 0.003839
2019-03-21 12:00 | 1553183983.3635 | 870 | 0.20412 | 0.004526 | 9.9e-05 | 100.8068 | 100.307 | 111.7947 | 21.1176 | 0.003466 | 0.003177 | 144.6119 | 0.26246 | 8.4025 | 0.003701 | 0.005343 | 0.00526 | 5.4e-05 | 0.000874 | 0.32333 | 1.3747 | 2.548 | 99.54 | 20.91 | 37.275 | 105.1163 | 110.8231 | 98.1317 | 0.009272 | 0.95088 | 0.022553 | 23.7178 | 0.044455 | 0.15614 | 0.027577 | 1.8576 | 1.8576 | 0.003455
2019-03-21 12:00 | 1553183984.364 | 871 | 0.2035 | 0.00453 | 9.8e-05 | 100.8139 | 100.3431 | 111.6478 | 21.1172 | 0.003466 | 0.003187 | 144.5454 | 0.26165 | 8.4014 | 0.003635 | 0.004608 | 0.004532 | 5.5e-05 | 0.000871 | 0.30702 | 1.3685 | 2.5481 | 99.54 | 20.91 | 37.272 | 105.1324 | 109.2279 | 98.3611 | 0.009272 | 0.9503 | 0.022531 | 23.7095 | 0.041871 | 0.15566 | 0.031327 | 2.1082 | 2.1082 | 0.003035
2019-03-21 12:00 | 1553183985.3626 | 872 | 0.20409 | 0.004542 | 9.8e-05 | 100.8345 | 100.3193 | 111.4978 | 21.1179 | 0.003504 | 0.003166 | 144.4755 | 0.26251 | 8.4004 | 0.003606 | 0.005251 | 0.005018 | 5.2e-05 | 0.00087 | 0.31365 | 1.3265 | 2.5479 | 99.54 | 20.91 | 37.27 | 106.0544 | 108.5294 | 97.8749 | 0.009271 | 0.95154 | 0.022531 | 23.6781 | 0.04582 | 0.15616 | 0.027765 | 1.8684 | 1.8684 | 0.003435
2019-03-21 12:00 | 1553183986.3627 | 873 | 0.20453 | 0.004552 | 9.8e-05 | 100.8446 | 100.3829 | 111.5131 | 21.118 | 0.003428 | 0.003144 | 144.3933 | 0.26177 | 8.402 | 0.003567 | 0.005256 | 0.005238 | 5.5e-05 | 0.000866 | 0.28335 | 1.3485 | 2.5478 | 99.54 | 20.91 | 37.268 | 104.2083 | 107.574 | 97.3303 | 0.009271 | 0.95294 | 0.022549 | 23.6623 | 0.041042 | 0.15572 | 0.025116 | 1.6915 | 1.6915 | 0.003784
2019-03-21 12:00 | 1553183987.3631 | 874 | 0.20369 | 0.004562 | 9.7e-05 | 100.7946 | 100.434 | 111.5666 | 21.1187 | 0.003547 | 0.003187 | 144.3216 | 0.26156 | 8.4029 | 0.003567 | 0.005428 | 0.00528 | 5.3e-05 | 0.000867 | 0.2959 | 1.3219 | 2.5479 | 99.54 | 20.91 | 37.265 | 107.0894 | 107.5754 | 98.3657 | 0.00927 | 0.95028 | 0.022523 | 23.701 | 0.03205 | 0.1556 | 0.030134 | 2.0272 | 2.0272 | 0.003155
2019-03-21 12:00 | 1553183988.3628 | 875 | 0.20435 | 0.004574 | 9.7e-05 | 100.7543 | 100.278 | 111.5119 | 21.1181 | 0.00351 | 0.003205 | 144.2425 | 0.26216 | 8.4056 | 0.003548 | 0.00522 | 0.005168 | 5.1e-05 | 0.000869 | 0.31045 | 1.2984 | 2.5479 | 99.54 | 20.91 | 37.263 | 106.1906 | 107.1158 | 98.8064 | 0.009269 | 0.94916 | 0.022504 | 23.7093 | 0.042383 | 0.15596 | 0.026159 | 1.7584 | 1.7584 | 0.003645
2019-03-21 12:00 | 1553183989.3628 | 876 | 0.20388 | 0.004586 | 9.6e-05 | 100.83 | 100.4003 | 111.652 | 21.1185 | 0.003508 | 0.003189 | 144.1634 | 0.26205 | 8.4036 | 0.003591 | 0.005164 | 0.00501 | 5.3e-05 | 0.000867 | 0.32372 | 1.2857 | 2.548 | 99.54 | 20.91 | 37.262 | 106.1454 | 108.1494 | 98.4204 | 0.009269 | 0.95014 | 0.02253 | 23.7118 | 0.038199 | 0.15589 | 0.028956 | 1.9486 | 1.9486 | 0.003288
2019-03-21 12:00 | 1553183990.3638 | 877 | 0.20435 | 0.004592 | 9.6e-05 | 100.7639 | 100.1633 | 111.7886 | 21.1184 | 0.003493 | 0.003219 | 144.0905 | 0.26145 | 8.3991 | 0.003572 | 0.00504 | 0.004989 | 5.2e-05 | 0.000864 | 0.26233 | 1.2902 | 2.5481 | 99.54 | 20.91 | 37.26 | 105.7771 | 107.7074 | 99.1543 | 0.009269 | 0.94827 | 0.022521 | 23.7498 | 0.053473 | 0.15553 | 0.026125 | 1.7575 | 1.7575 | 0.003637
2019-03-21 12:00 | 1553183991.3633 | 878 | 0.20382 | 0.004601 | 9.6e-05 | 100.836 | 100.3819 | 111.4959 | 21.1183 | 0.003487 | 0.003274 | 144.0349 | 0.26024 | 8.4036 | 0.003676 | 0.004877 | 0.004819 | 5.3e-05 | 0.000863 | 0.31382 | 1.3081 | 2.5479 | 99.54 | 20.91 | 37.26 | 105.6408 | 110.2328 | 100.4769 | 0.009269 | 0.94491 | 0.022452 | 23.7607 | 0.040366 | 0.15482 | 0.029335 | 1.9674 | 1.9674 | 0.003234
2019-03-21 12:00 | 1553183992.3639 | 879 | 0.20425 | 0.004604 | 9.6e-05 | 100.7478 | 100.3039 | 111.4483 | 21.1193 | 0.003588 | 0.003218 | 143.9504 | 0.26143 | 8.4037 | 0.003627 | 0.004584 | 0.004457 | 5.7e-05 | 0.000862 | 0.33339 | 1.2918 | 2.548 | 99.54 | 20.91 | 37.255 | 108.0779 | 109.0409 | 99.1346 | 0.009267 | 0.94832 | 0.022487 | 23.7129 | 0.039492 | 0.15553 | 0.026733 | 1.7957 | 1.7957 | 0.00356
2019-03-21 12:00 | 1553183993.364 | 880 | 0.2038 | 0.004611 | 9.7e-05 | 100.8372 | 100.1173 | 111.5356 | 21.1189 | 0.003445 | 0.003252 | 143.8945 | 0.26119 | 8.4033 | 0.003709 | 0.005361 | 0.005266 | 5.6e-05 | 0.00086 | 0.33052 | 1.3003 | 2.5478 | 99.54 | 20.91 | 37.253 | 104.6298 | 111.0357 | 99.939 | 0.009267 | 0.94628 | 0.022472 | 23.7478 | 0.064084 | 0.15538 | 0.029401 | 1.9735 | 1.9735 | 0.003236
2019-03-21 12:00 | 1553183994.3631 | 881 | 0.20429 | 0.004619 | 9.7e-05 | 100.8443 | 100.466 | 111.4837 | 21.1194 | 0.003521 | 0.003246 | 143.8106 | 0.26122 | 8.3981 | 0.003666 | 0.004814 | 0.004735 | 5.5e-05 | 0.000863 | 0.32516 | 1.3223 | 2.5479 | 99.54 | 20.92 | 37.251 | 106.4518 | 109.98 | 99.7991 | 0.009272 | 0.94663 | 0.022471 | 23.7379 | 0.033613 | 0.1554 | 0.026481 | 1.7774 | 1.7774 | 0.003593
2019-03-21 12:00 | 1553183995.3631 | 882 | 0.20369 | 0.004626 | 9.8e-05 | 100.8372 | 100.2515 | 111.4635 | 21.118 | 0.003599 | 0.003173 | 143.7239 | 0.26144 | 8.4034 | 0.003754 | 0.005627 | 0.005515 | 5e-05 | 0.00086 | 0.29502 | 1.3004 | 2.5481 | 99.54 | 20.92 | 37.246 | 108.3412 | 112.1251 | 98.0381 | 0.009271 | 0.95112 | 0.022522 | 23.6796 | 0.052103 | 0.15553 | 0.030035 | 2.0204 | 2.0204 | 0.003164
2019-03-21 12:00 | 1553183996.3635 | 883 | 0.20423 | 0.004633 | 9.9e-05 | 100.8196 | 100.2045 | 111.8126 | 21.1203 | 0.003507 | 0.003151 | 143.6544 | 0.26193 | 8.4027 | 0.00371 | 0.005261 | 0.005236 | 5.4e-05 | 0.000861 | 0.33467 | 1.3091 | 2.548 | 99.54 | 20.92 | 37.245 | 106.1125 | 111.0506 | 97.4971 | 0.009271 | 0.95251 | 0.022574 | 23.6994 | 0.054738 | 0.15582 | 0.026839 | 1.8096 | 1.8096 | 0.003539
2019-03-21 12:00 | 1553183997.3623 | 884 | 0.20382 | 0.004645 | 9.9e-05 | 100.8962 | 100.3005 | 111.7703 | 21.1189 | 0.003527 | 0.003216 | 143.6002 | 0.26087 | 8.4022 | 0.003768 | 0.00511 | 0.005056 | 5.5e-05 | 0.000863 | 0.33179 | 1.2945 | 2.5478 | 99.53 | 20.92 | 37.241 | 106.6111 | 112.4733 | 99.0769 | 0.00927 | 0.94847 | 0.022522 | 23.7453 | 0.05296 | 0.15519 | 0.029237 | 1.9667 | 1.9667 | 0.003243
2019-03-21 12:00 | 1553183998.3626 | 885 | 0.20402 | 0.00465 | 9.9e-05 | 100.7424 | 100.2849 | 111.7304 | 21.1191 | 0.003493 | 0.003236 | 143.5248 | 0.261 | 8.4 | 0.003724 | 0.005438 | 0.005373 | 5.6e-05 | 0.000863 | 0.32421 | 1.3335 | 2.5481 | 99.53 | 20.92 | 37.239 | 105.7911 | 111.384 | 99.5579 | 0.00927 | 0.94724 | 0.022503 | 23.7564 | 0.040717 | 0.15527 | 0.028081 | 1.8874 | 1.8874 | 0.003381
2019-03-21 12:00 | 1553183999.3623 | 886 | 0.20402 | 0.004658 | 9.9e-05 | 100.8336 | 100.2358 | 111.6628 | 21.12 | 0.003549 | 0.003195 | 143.4645 | 0.26183 | 8.4021 | 0.003785 | 0.005565 | 0.005427 | 5.3e-05 | 0.000862 | 0.32731 | 1.3045 | 2.548 | 99.53 | 20.92 | 37.237 | 107.1531 | 112.8851 | 98.5654 | 0.009269 | 0.94977 | 0.022526 | 23.7176 | 0.053184 | 0.15576 | 0.02804 | 1.8866 | 1.8866 | 0.003393
2019-03-21 12:00 | 1553184000.3624 | 887 | 0.204 | 0.004665 | 9.9e-05 | 100.8345 | 100.2753 | 111.9975 | 21.1192 | 0.003528 | 0.003205 | 143.4059 | 0.26139 | 8.4019 | 0.00375 | 0.005646 | 0.005576 | 5.6e-05 | 0.000866 | 0.29717 | 1.2629 | 2.5478 | 99.53 | 20.92 | 37.235 | 106.628 | 112.0197 | 98.8092 | 0.009269 | 0.94915 | 0.022553 | 23.7609 | 0.049738 | 0.1555 | 0.028157 | 1.8966 | 1.8966 | 0.00337
2019-03-21 12:00 | 1553184001.3634 | 888 | 0.20403 | 0.004668 | 9.8e-05 | 100.8181 | 100.2014 | 111.978 | 21.1196 | 0.003482 | 0.003252 | 143.3073 | 0.26111 | 8.4037 | 0.003819 | 0.005626 | 0.005569 | 5.4e-05 | 0.000867 | 0.26034 | 1.2711 | 2.5479 | 99.53 | 20.92 | 37.233 | 105.5045 | 113.7161 | 99.9516 | 0.009268 | 0.94625 | 0.022516 | 23.7953 | 0.054879 | 0.15534 | 0.027953 | 1.8799 | 1.8799 | 0.003396
2019-03-21 12:00 | 1553184002.3625 | 889 | 0.20402 | 0.004674 | 9.8e-05 | 100.7999 | 100.2436 | 111.6431 | 21.1195 | 0.003535 | 0.00324 | 143.227 | 0.26104 | 8.4017 | 0.003884 | 0.005162 | 0.005047 | 5.7e-05 | 0.000865 | 0.33771 | 1.2639 | 2.5478 | 99.53 | 20.92 | 37.232 | 106.8078 | 115.287 | 99.6565 | 0.009268 | 0.94699 | 0.022491 | 23.7503 | 0.049498 | 0.15529 | 0.028005 | 1.8814 | 1.8814 | 0.003393
2019-03-21 12:00 | 1553184003.3636 | 890 | 0.20416 | 0.004677 | 9.7e-05 | 100.8148 | 100.3264 | 111.6271 | 21.1194 | 0.003581 | 0.003241 | 143.1364 | 0.26108 | 8.4021 | 0.003794 | 0.005581 | 0.005533 | 5.6e-05 | 0.000866 | 0.33699 | 1.2681 | 2.548 | 99.53 | 20.92 | 37.231 | 107.9109 | 113.1034 | 99.6743 | 0.009268 | 0.94695 | 0.022489 | 23.7492 | 0.043438 | 0.15532 | 0.027181 | 1.8259 | 1.8259 | 0.003496
2019-03-21 12:00 | 1553184004.3635 | 891 | 0.20407 | 0.004681 | 9.6e-05 | 100.805 | 100.4051 | 111.8396 | 21.1205 | 0.003499 | 0.003259 | 143.0873 | 0.26175 | 8.4015 | 0.003921 | 0.004531 | 0.004463 | 5.4e-05 | 0.000867 | 0.31996 | 1.3274 | 2.5479 | 99.53 | 20.92 | 37.231 | 105.9349 | 116.2003 | 100.1116 | 0.009268 | 0.94584 | 0.022497 | 23.7857 | 0.035554 | 0.15572 | 0.027736 | 1.8639 | 1.8639 | 0.003434
2019-03-21 12:00 | 1553184005.3637 | 892 | 0.20403 | 0.004681 | 9.6e-05 | 100.7636 | 100.431 | 111.5635 | 21.1195 | 0.003557 | 0.003207 | 143.0107 | 0.26083 | 8.4003 | 0.003871 | 0.005481 | 0.005227 | 5.6e-05 | 0.000865 | 0.33987 | 1.3215 | 2.548 | 99.53 | 20.92 | 37.231 | 107.3249 | 114.9665 | 98.8689 | 0.009268 | 0.949 | 0.022507 | 23.7167 | 0.029573 | 0.15517 | 0.027947 | 1.8789 | 1.8789 | 0.003394
2019-03-21 12:00 | 1553184006.3628 | 893 | 0.2042 | 0.004686 | 9.6e-05 | 100.7823 | 100.3567 | 111.2004 | 21.1207 | 0.003536 | 0.003238 | 142.9399 | 0.2612 | 8.4026 | 0.003805 | 0.00518 | 0.005069 | 6e-05 | 0.000863 | 0.32836 | 1.3323 | 2.548 | 99.53 | 20.92 | 37.232 | 106.838 | 113.3764 | 99.6199 | 0.009268 | 0.94709 | 0.022448 | 23.702 | 0.037854 | 0.15539 | 0.026959 | 1.8077 | 1.8077 | 0.003533
2019-03-21 12:00 | 1553184007.3628 | 894 | 0.20407 | 0.004679 | 9.5e-05 | 100.7928 | 100.3959 | 111.226 | 21.1206 | 0.003532 | 0.003264 | 142.8697 | 0.26064 | 8.4018 | 0.003821 | 0.005342 | 0.005249 | 5.6e-05 | 0.000864 | 0.29806 | 1.3661 | 2.548 | 99.53 | 20.92 | 37.233 | 106.73 | 113.7565 | 100.235 | 0.009268 | 0.94553 | 0.022432 | 23.7243 | 0.035289 | 0.15505 | 0.027704 | 1.8564 | 1.8564 | 0.003433
2019-03-21 12:00 | 1553184008.3638 | 895 | 0.20408 | 0.004674 | 9.5e-05 | 100.8872 | 100.3809 | 111.3218 | 21.1205 | 0.003561 | 0.003287 | 142.789 | 0.26142 | 8.4013 | 0.00376 | 0.004582 | 0.004477 | 5.9e-05 | 0.000866 | 0.33644 | 1.3568 | 2.5479 | 99.53 | 20.92 | 37.233 | 107.4213 | 112.2806 | 100.7874 | 0.009268 | 0.94413 | 0.022425 | 23.752 | 0.045002 | 0.15552 | 0.027695 | 1.8552 | 1.8552 | 0.003445
2019-03-21 12:00 | 1553184009.3629 | 896 | 0.20414 | 0.004671 | 9.5e-05 | 100.7701 | 100.3768 | 111.1992 | 21.1213 | 0.003582 | 0.00326 | 142.7356 | 0.26102 | 8.3998 | 0.003795 | 0.004911 | 0.004798 | 5.6e-05 | 0.000859 | 0.32493 | 1.3307 | 2.548 | 99.53 | 20.92 | 37.233 | 107.9387 | 113.1313 | 100.1532 | 0.009268 | 0.94573 | 0.022432 | 23.7188 | 0.034979 | 0.15528 | 0.027304 | 1.8296 | 1.8296 | 0.003488
2019-03-21 12:00 | 1553184010.3629 | 897 | 0.20419 | 0.00467 | 9.4e-05 | 100.7737 | 100.3428 | 110.9371 | 21.1206 | 0.00355 | 0.003247 | 142.6713 | 0.26172 | 8.4021 | 0.003807 | 0.004932 | 0.004867 | 5.5e-05 | 0.000856 | 0.30182 | 1.329 | 2.548 | 99.53 | 20.92 | 37.234 | 107.1734 | 113.4263 | 99.8239 | 0.009269 | 0.94657 | 0.022415 | 23.6804 | 0.038332 | 0.1557 | 0.027007 | 1.8084 | 1.8084 | 0.003539
2019-03-21 12:00 | 1553184011.3629 | 898 | 0.20413 | 0.004668 | 9.4e-05 | 100.8282 | 100.2021 | 111.3448 | 21.121 | 0.003527 | 0.003237 | 142.6103 | 0.2613 | 8.4021 | 0.003692 | 0.005036 | 0.004884 | 5.5e-05 | 0.000849 | 0.28279 | 1.316 | 2.5481 | 99.53 | 20.92 | 37.234 | 106.6146 | 110.6212 | 99.5961 | 0.009269 | 0.94715 | 0.022463 | 23.7166 | 0.055717 | 0.15544 | 0.027365 | 1.8363 | 1.8363 | 0.003479
2019-03-21 12:00 | 1553184012.364 | 899 | 0.20407 | 0.004665 | 9.3e-05 | 100.8416 | 100.2232 | 111.5091 | 21.121 | 0.003549 | 0.003241 | 142.5351 | 0.26055 | 8.401 | 0.003704 | 0.005597 | 0.005439 | 5.4e-05 | 0.000849 | 0.31222 | 1.3349 | 2.5479 | 99.53 | 20.92 | 37.234 | 107.1436 | 110.906 | 99.672 | 0.009269 | 0.94695 | 0.022477 | 23.7365 | 0.055022 | 0.155 | 0.027771 | 1.8648 | 1.8648 | 0.003416
2019-03-21 12:00 | 1553184013.3625 | 900 | 0.20415 | 0.004669 | 9.3e-05 | 100.777 | 100.4439 | 111.4523 | 21.1208 | 0.003547 | 0.003213 | 142.4935 | 0.26078 | 8.4015 | 0.003686 | 0.005375 | 0.005296 | 5.7e-05 | 0.000848 | 0.31526 | 1.2982 | 2.5479 | 99.53 | 20.92 | 37.234 | 107.0939 | 110.4671 | 99.0091 | 0.009269 | 0.94864 | 0.022492 | 23.7094 | 0.02961 | 0.15514 | 0.027279 | 1.8329 | 1.8329 | 0.003479
2019-03-21 12:00 | 1553184014.3626 | 901 | 0.20419 | 0.004665 | 9.2e-05 | 100.7329 | 100.4882 | 111.1284 | 21.1205 | 0.003543 | 0.003284 | 142.3921 | 0.26156 | 8.402 | 0.003702 | 0.004865 | 0.004778 | 5.2e-05 | 0.000847 | 0.31725 | 1.308 | 2.5479 | 99.53 | 20.91 | 37.232 | 107.0046 | 110.8552 | 100.7192 | 0.009262 | 0.9443 | 0.022408 | 23.7292 | 0.021754 | 0.1556 | 0.026998 | 1.8073 | 1.8073 | 0.003539
2019-03-21 12:00 | 1553184015.5006 | 902 | 0.20419 | 0.004668 | 9.2e-05 | 100.8246 | 100.4245 | 111.4044 | 21.1205 | 0.003562 | 0.003239 | 142.3323 | 0.26169 | 8.4012 | 0.003646 | 0.00512 | 0.004999 | 5.7e-05 | 0.000849 | 0.30397 | 1.2963 | 2.548 | 99.53 | 20.92 | 37.233 | 107.4536 | 109.4959 | 99.6386 | 0.009268 | 0.94704 | 0.022468 | 23.7243 | 0.035569 | 0.15568 | 0.027001 | 1.8124 | 1.8124 | 0.00353
2019-03-21 12:00 | 1553184016.3995 | 903 | 0.20424 | 0.004666 | 9.2e-05 | 100.8488 | 100.3741 | 111.4649 | 21.1211 | 0.003617 | 0.003248 | 142.2742 | 0.26145 | 8.4029 | 0.003739 | 0.004921 | 0.004892 | 5.6e-05 | 0.000852 | 0.31061 | 1.3213 | 2.548 | 99.53 | 20.92 | 37.231 | 108.7849 | 111.7665 | 99.8604 | 0.009268 | 0.94648 | 0.022467 | 23.7378 | 0.042201 | 0.15553 | 0.026709 | 1.7927 | 1.7927 | 0.003566
2019-03-21 12:00 | 1553184017.4 | 904 | 0.20425 | 0.004665 | 9.2e-05 | 100.7892 | 100.3983 | 111.4358 | 21.1209 | 0.003576 | 0.003297 | 142.1967 | 0.26161 | 8.403 | 0.003688 | 0.00499 | 0.004934 | 5.9e-05 | 0.000851 | 0.29087 | 1.2945 | 2.5479 | 99.53 | 20.92 | 37.229 | 107.8077 | 110.5083 | 101.0413 | 0.009267 | 0.94349 | 0.022429 | 23.7723 | 0.03476 | 0.15563 | 0.026666 | 1.7867 | 1.7867 | 0.00358
2019-03-21 12:00 | 1553184018.3986 | 905 | 0.20421 | 0.004658 | 9.2e-05 | 100.7791 | 100.4691 | 111.5201 | 21.1208 | 0.003569 | 0.003179 | 142.1472 | 0.26134 | 8.4025 | 0.003716 | 0.004969 | 0.004818 | 5.5e-05 | 0.000855 | 0.27815 | 1.3069 | 2.5479 | 99.53 | 20.92 | 37.228 | 107.6224 | 111.1884 | 98.1803 | 0.009267 | 0.95076 | 0.022524 | 23.6902 | 0.027551 | 0.15547 | 0.026913 | 1.8109 | 1.8109 | 0.003529
2019-03-21 12:00 | 1553184019.3987 | 906 | 0.20425 | 0.004646 | 9.2e-05 | 100.8127 | 100.3748 | 111.8626 | 21.1209 | 0.003526 | 0.003293 | 142.0553 | 0.26077 | 8.4032 | 0.003634 | 0.005397 | 0.00524 | 5.3e-05 | 0.00086 | 0.33467 | 1.3195 | 2.548 | 99.53 | 20.92 | 37.227 | 106.5924 | 109.211 | 100.9385 | 0.009267 | 0.94375 | 0.022475 | 23.8145 | 0.038942 | 0.15513 | 0.026676 | 1.7911 | 1.7911 | 0.00356
2019-03-21 12:00 | 1553184020.3987 | 907 | 0.20428 | 0.004642 | 9.2e-05 | 100.7478 | 100.3319 | 111.629 | 21.1213 | 0.00355 | 0.003231 | 141.9749 | 0.26158 | 8.4028 | 0.003703 | 0.005146 | 0.005088 | 5.5e-05 | 0.000859 | 0.31509 | 1.2725 | 2.5479 | 99.53 | 20.92 | 37.227 | 107.1629 | 110.8779 | 99.4354 | 0.009267 | 0.94756 | 0.022497 | 23.7418 | 0.037002 | 0.15562 | 0.026505 | 1.7813 | 1.7813 | 0.00359
2019-03-21 12:00 | 1553184021.3987 | 908 | 0.20432 | 0.004635 | 9.2e-05 | 100.701 | 100.4592 | 111.2734 | 21.1204 | 0.003618 | 0.003276 | 141.9038 | 0.26096 | 8.4008 | 0.003664 | 0.005119 | 0.005123 | 5.4e-05 | 0.00086 | 0.31526 | 1.2827 | 2.5479 | 99.53 | 20.91 | 37.228 | 108.807 | 109.9416 | 100.5337 | 0.009261 | 0.94477 | 0.022428 | 23.7388 | 0.0215 | 0.15524 | 0.026305 | 1.7625 | 1.7625 | 0.00362
2019-03-21 12:00 | 1553184022.3991 | 909 | 0.20423 | 0.004627 | 9.1e-05 | 100.7427 | 100.3877 | 111.3516 | 21.1218 | 0.003618 | 0.003273 | 141.8717 | 0.26104 | 8.4018 | 0.003671 | 0.004623 | 0.004541 | 5.4e-05 | 0.000861 | 0.28998 | 1.2749 | 2.5479 | 99.53 | 20.91 | 37.228 | 108.8257 | 110.1085 | 100.4544 | 0.009261 | 0.94497 | 0.022438 | 23.7446 | 0.031574 | 0.15529 | 0.026837 | 1.799 | 1.799 | 0.003548
2019-03-21 12:00 | 1553184023.3989 | 910 | 0.20433 | 0.004617 | 9.1e-05 | 100.7296 | 100.3179 | 111.6922 | 21.1205 | 0.003555 | 0.003243 | 141.7892 | 0.26082 | 8.4031 | 0.003622 | 0.004813 | 0.004769 | 5.3e-05 | 0.000858 | 0.29479 | 1.2983 | 2.548 | 99.53 | 20.91 | 37.229 | 107.2968 | 108.9072 | 99.7201 | 0.009262 | 0.94683 | 0.022494 | 23.7575 | 0.036633 | 0.15516 | 0.026239 | 1.7633 | 1.7633 | 0.003616
2019-03-21 12:00 | 1553184024.3999 | 911 | 0.20438 | 0.004613 | 9.1e-05 | 100.7293 | 100.3424 | 111.8414 | 21.1216 | 0.003581 | 0.003279 | 141.7229 | 0.26046 | 8.4014 | 0.003579 | 0.004905 | 0.004724 | 5.4e-05 | 0.000857 | 0.29878 | 1.3145 | 2.548 | 99.53 | 20.91 | 37.229 | 107.9138 | 107.8603 | 100.5923 | 0.009262 | 0.94462 | 0.022483 | 23.8012 | 0.034421 | 0.15495 | 0.025958 | 1.7436 | 1.7436 | 0.003652
2019-03-21 12:00 | 1553184025.4 | 912 | 0.20431 | 0.004603 | 9e-05 | 100.76 | 100.3584 | 111.9863 | 21.1209 | 0.003556 | 0.003268 | 141.6566 | 0.26122 | 8.4039 | 0.003661 | 0.005563 | 0.005432 | 5.5e-05 | 0.000862 | 0.30807 | 1.3128 | 2.5479 | 99.53 | 20.91 | 37.231 | 107.3193 | 109.8505 | 100.3388 | 0.009262 | 0.94526 | 0.022505 | 23.8085 | 0.035719 | 0.1554 | 0.026363 | 1.7726 | 1.7726 | 0.003603
2019-03-21 12:00 | 1553184026.3999 | 913 | 0.20445 | 0.00459 | 9.1e-05 | 100.7075 | 100.3468 | 112.3062 | 21.1205 | 0.003602 | 0.003235 | 141.5884 | 0.26097 | 8.4009 | 0.003683 | 0.005149 | 0.004991 | 5e-05 | 0.00086 | 0.32659 | 1.344 | 2.548 | 99.53 | 20.92 | 37.233 | 108.4382 | 110.3982 | 99.539 | 0.009268 | 0.94729 | 0.022562 | 23.817 | 0.032095 | 0.15525 | 0.025592 | 1.725 | 1.725 | 0.003699
2019-03-21 12:00 | 1553184027.4 | 914 | 0.20442 | 0.004583 | 9.1e-05 | 100.8029 | 100.4041 | 112.5552 | 21.1216 | 0.003553 | 0.003274 | 141.5235 | 0.26082 | 8.4044 | 0.003671 | 0.004958 | 0.004882 | 5.2e-05 | 0.000862 | 0.28584 | 1.3071 | 2.548 | 99.53 | 20.92 | 37.234 | 107.2372 | 110.1178 | 100.4878 | 0.009269 | 0.94489 | 0.022558 | 23.8737 | 0.03546 | 0.15516 | 0.025733 | 1.7342 | 1.7342 | 0.003677
2019-03-21 12:00 | 1553184028.3998 | 915 | 0.20441 | 0.004573 | 9e-05 | 100.6876 | 100.3124 | 112.3926 | 21.1219 | 0.003531 | 0.003189 | 141.4614 | 0.26097 | 8.4011 | 0.003705 | 0.004833 | 0.004778 | 4.8e-05 | 0.000857 | 0.30845 | 1.3136 | 2.5478 | 99.53 | 20.91 | 37.237 | 106.7077 | 110.9283 | 98.4299 | 0.009264 | 0.95012 | 0.022604 | 23.7906 | 0.033388 | 0.15525 | 0.025861 | 1.7465 | 1.7465 | 0.003654
2019-03-21 12:00 | 1553184029.3982 | 916 | 0.20449 | 0.004559 | 9e-05 | 100.6935 | 100.2699 | 112.3517 | 21.1214 | 0.003564 | 0.003247 | 141.4196 | 0.25988 | 8.4026 | 0.003631 | 0.00465 | 0.004584 | 5.1e-05 | 0.000858 | 0.34916 | 1.3181 | 2.5481 | 99.53 | 20.91 | 37.237 | 107.5103 | 109.1233 | 99.8278 | 0.009264 | 0.94656 | 0.022557 | 23.831 | 0.037714 | 0.1546 | 0.025362 | 1.7092 | 1.7092 | 0.003718
2019-03-21 12:00 | 1553184030.3996 | 917 | 0.20451 | 0.004552 | 9e-05 | 100.7335 | 100.4173 | 112.4897 | 21.1217 | 0.003547 | 0.003232 | 141.3458 | 0.2611 | 8.4039 | 0.003664 | 0.00468 | 0.0046 | 5.4e-05 | 0.000858 | 0.32748 | 1.3101 | 2.548 | 99.53 | 20.91 | 37.237 | 107.1043 | 109.9373 | 99.4594 | 0.009264 | 0.94749 | 0.022582 | 23.8339 | 0.028115 | 0.15533 | 0.025248 | 1.7034 | 1.7034 | 0.003748
2019-03-21 12:00 | 1553184031.3983 | 918 | 0.20453 | 0.004538 | 9e-05 | 100.6641 | 100.3826 | 112.4149 | 21.1213 | 0.003575 | 0.003258 | 141.3012 | 0.26188 | 8.4028 | 0.003659 | 0.005328 | 0.005253 | 5e-05 | 0.000862 | 0.33411 | 1.3641 | 2.5479 | 99.53 | 20.91 | 37.237 | 107.778 | 109.8115 | 100.0882 | 0.009264 | 0.9459 | 0.022556 | 23.846 | 0.025043 | 0.15579 | 0.025177 | 1.6966 | 1.6966 | 0.003774
2019-03-21 12:01 | 1553184032.3993 | 919 | 0.2046 | 0.004521 | 9e-05 | 100.7251 | 100.2235 | 112.2247 | 21.1215 | 0.003559 | 0.003242 | 141.2118 | 0.2607 | 8.4021 | 0.003593 | 0.005789 | 0.005662 | 4.9e-05 | 0.000862 | 0.33085 | 1.3296 | 2.5479 | 99.53 | 20.91 | 37.237 | 107.3837 | 108.2043 | 99.7014 | 0.009264 | 0.94688 | 0.022549 | 23.8135 | 0.044652 | 0.15509 | 0.024738 | 1.6665 | 1.6665 | 0.003825
2019-03-21 12:01 | 1553184033.3993 | 920 | 0.2046 | 0.004503 | 9e-05 | 100.6813 | 100.2416 | 112.1806 | 21.121 | 0.003586 | 0.003214 | 141.1363 | 0.26156 | 8.402 | 0.003695 | 0.00497 | 0.004931 | 5.1e-05 | 0.000865 | 0.33052 | 1.317 | 2.5478 | 99.53 | 20.91 | 37.238 | 108.0424 | 110.693 | 99.0314 | 0.009264 | 0.94858 | 0.022564 | 23.7874 | 0.039151 | 0.1556 | 0.024773 | 1.6701 | 1.6701 | 0.003829
2019-03-21 12:01 | 1553184034.3986 | 921 | 0.20467 | 0.004483 | 9e-05 | 100.6307 | 100.2975 | 112.0398 | 21.1209 | 0.003541 | 0.003257 | 141.0722 | 0.26079 | 8.4031 | 0.003655 | 0.00506 | 0.00484 | 5.1e-05 | 0.000867 | 0.31868 | 1.3269 | 2.5479 | 99.53 | 20.91 | 37.239 | 106.9516 | 109.7093 | 100.0745 | 0.009264 | 0.94593 | 0.022519 | 23.8058 | 0.029668 | 0.15514 | 0.024357 | 1.6386 | 1.6386 | 0.003891
2019-03-21 12:01 | 1553184035.3989 | 922 | 0.2047 | 0.004468 | 9e-05 | 100.6286 | 100.3172 | 111.8433 | 21.1224 | 0.003566 | 0.003266 | 141.0273 | 0.26102 | 8.403 | 0.003573 | 0.004945 | 0.00489 | 4.9e-05 | 0.000865 | 0.32676 | 1.3151 | 2.5479 | 99.53 | 20.91 | 37.239 | 107.5671 | 107.7239 | 100.2806 | 0.009264 | 0.94541 | 0.022493 | 23.7915 | 0.027721 | 0.15528 | 0.024215 | 1.6272 | 1.6272 | 0.003922
2019-03-21 12:01 | 1553184036.3986 | 923 | 0.20479 | 0.004449 | 8.9e-05 | 100.6331 | 100.2031 | 111.5955 | 21.1221 | 0.003578 | 0.003215 | 140.9549 | 0.26091 | 8.3992 | 0.003636 | 0.005143 | 0.005071 | 5.4e-05 | 0.00087 | 0.33323 | 1.2992 | 2.548 | 99.53 | 20.91 | 37.239 | 107.8535 | 109.2516 | 99.0621 | 0.009264 | 0.94851 | 0.022505 | 23.7263 | 0.038297 | 0.15522 | 0.023678 | 1.592 | 1.592 | 0.004007
2019-03-21 12:01 | 1553184037.3995 | 924 | 0.20474 | 0.004431 | 8.9e-05 | 100.6116 | 100.2477 | 111.6654 | 21.1216 | 0.003554 | 0.003266 | 140.8936 | 0.26106 | 8.4044 | 0.003546 | 0.00516 | 0.005119 | 5.1e-05 | 0.00087 | 0.33555 | 1.2741 | 2.548 | 99.53 | 20.91 | 37.24 | 107.2632 | 107.0639 | 100.2776 | 0.009264 | 0.94542 | 0.022475 | 23.7724 | 0.032408 | 0.15531 | 0.024018 | 1.6127 | 1.6127 | 0.003958
2019-03-21 12:01 | 1553184038.3991 | 925 | 0.2047 | 0.004402 | 8.9e-05 | 100.6635 | 100.2467 | 111.6701 | 21.1218 | 0.003554 | 0.003215 | 140.8295 | 0.25976 | 8.4016 | 0.0036 | 0.004644 | 0.004459 | 5.2e-05 | 0.00087 | 0.30829 | 1.3604 | 2.5478 | 99.53 | 20.91 | 37.24 | 107.2612 | 108.3819 | 99.0486 | 0.009264 | 0.94854 | 0.022512 | 23.7338 | 0.037107 | 0.15453 | 0.024305 | 1.6347 | 1.6347 | 0.003885
2019-03-21 12:01 | 1553184039.4001 | 926 | 0.20474 | 0.004387 | 8.8e-05 | 100.6128 | 100.1326 | 111.3535 | 21.1213 | 0.003513 | 0.003278 | 140.7306 | 0.26086 | 8.4042 | 0.003679 | 0.005291 | 0.005235 | 5.3e-05 | 0.000872 | 0.32764 | 1.3421 | 2.5479 | 99.53 | 20.9 | 37.241 | 106.2804 | 110.2918 | 100.584 | 0.009259 | 0.94464 | 0.022434 | 23.749 | 0.042792 | 0.15519 | 0.024108 | 1.6159 | 1.6159 | 0.003947
2019-03-21 12:01 | 1553184040.3988 | 927 | 0.20478 | 0.004365 | 8.8e-05 | 100.5565 | 100.4296 | 111.534 | 21.1221 | 0.003477 | 0.00327 | 140.6626 | 0.26071 | 8.4016 | 0.003651 | 0.005263 | 0.005092 | 5.3e-05 | 0.000872 | 0.33699 | 1.3539 | 2.5479 | 99.53 | 20.9 | 37.242 | 105.4085 | 109.6088 | 100.3886 | 0.009259 | 0.94514 | 0.022458 | 23.762 | 0.011294 | 0.1551 | 0.023883 | 1.6026 | 1.6026 | 0.003978
2019-03-21 12:01 | 1553184041.3999 | 928 | 0.20474 | 0.004352 | 8.8e-05 | 100.5505 | 100.2583 | 112.1172 | 21.1222 | 0.003605 | 0.003266 | 140.5963 | 0.26104 | 8.4021 | 0.003709 | 0.005164 | 0.00506 | 5.4e-05 | 0.000871 | 0.29623 | 1.3793 | 2.548 | 99.53 | 20.9 | 37.243 | 108.4943 | 111.0327 | 100.2782 | 0.00926 | 0.94542 | 0.02252 | 23.8205 | 0.026035 | 0.15529 | 0.024147 | 1.6248 | 1.6248 | 0.003928
2019-03-21 12:01 | 1553184042.3993 | 929 | 0.20472 | 0.004331 | 8.7e-05 | 100.5639 | 100.2913 | 112.3625 | 21.1216 | 0.003567 | 0.003191 | 140.5519 | 0.2603 | 8.4021 | 0.003719 | 0.005251 | 0.005225 | 6e-05 | 0.000875 | 0.347 | 1.387 | 2.548 | 99.53 | 20.9 | 37.244 | 107.5884 | 111.2665 | 98.4664 | 0.00926 | 0.95003 | 0.0226 | 23.7886 | 0.02428 | 0.15485 | 0.024252 | 1.6376 | 1.6376 | 0.003886
2019-03-21 12:01 | 1553184043.398 | 930 | 0.20477 | 0.004319 | 8.6e-05 | 100.6006 | 100.2276 | 112.305 | 21.1229 | 0.003545 | 0.003228 | 140.4865 | 0.2611 | 8.4013 | 0.00362 | 0.005346 | 0.005148 | 5.5e-05 | 0.000875 | 0.3115 | 1.4036 | 2.548 | 99.53 | 20.9 | 37.244 | 107.0434 | 108.8749 | 99.3579 | 0.00926 | 0.94775 | 0.022567 | 23.8111 | 0.033221 | 0.15532 | 0.023999 | 1.6183 | 1.6183 | 0.003945
2019-03-21 12:01 | 1553184044.3981 | 931 | 0.20486 | 0.004307 | 8.5e-05 | 100.62 | 100.3724 | 112.5282 | 21.1224 | 0.003534 | 0.003197 | 140.4587 | 0.26049 | 8.4009 | 0.003653 | 0.005364 | 0.005253 | 5.6e-05 | 0.000874 | 0.33948 | 1.3912 | 2.548 | 99.53 | 20.9 | 37.243 | 106.7688 | 109.6604 | 98.6295 | 0.00926 | 0.94961 | 0.022612 | 23.8114 | 0.022035 | 0.15496 | 0.023464 | 1.5853 | 1.5853 | 0.004018
2019-03-21 12:01 | 1553184045.3984 | 932 | 0.20483 | 0.004292 | 8.4e-05 | 100.474 | 100.3884 | 112.7256 | 21.1222 | 0.003586 | 0.003218 | 140.3872 | 0.26103 | 8.4029 | 0.003592 | 0.004951 | 0.004881 | 5.7e-05 | 0.000875 | 0.32781 | 1.362 | 2.548 | 99.53 | 20.9 | 37.243 | 108.0365 | 108.1902 | 99.1375 | 0.00926 | 0.94831 | 0.022616 | 23.8485 | 0.007621 | 0.15529 | 0.023662 | 1.5991 | 1.5991 | 0.003991
2019-03-21 12:01 | 1553184046.3985 | 933 | 0.20499 | 0.00428 | 8.3e-05 | 100.4799 | 100.3877 | 112.8407 | 21.1234 | 0.003564 | 0.003226 | 140.3325 | 0.25954 | 8.4015 | 0.003622 | 0.005487 | 0.005362 | 5.5e-05 | 0.000872 | 0.32836 | 1.3073 | 2.5479 | 99.53 | 20.91 | 37.244 | 107.5013 | 108.9069 | 99.3128 | 0.009265 | 0.94787 | 0.022622 | 23.8663 | 0.008212 | 0.1544 | 0.022714 | 1.5354 | 1.5354 | 0.004133
2019-03-21 12:01 | 1553184047.3992 | 934 | 0.205 | 0.004261 | 8.3e-05 | 100.5636 | 100.3285 | 112.578 | 21.1234 | 0.003532 | 0.003205 | 140.2701 | 0.26071 | 8.4009 | 0.003584 | 0.004975 | 0.004898 | 5.8e-05 | 0.000872 | 0.33268 | 1.2756 | 2.5479 | 99.53 | 20.91 | 37.244 | 106.7281 | 108.0046 | 98.8062 | 0.009265 | 0.94916 | 0.022611 | 23.8223 | 0.020944 | 0.1551 | 0.022658 | 1.5309 | 1.5309 | 0.004164
2019-03-21 12:01 | 1553184048.3988 | 935 | 0.20504 | 0.004248 | 8.3e-05 | 100.5082 | 100.4361 | 112.7801 | 21.1231 | 0.00356 | 0.003179 | 140.1951 | 0.26043 | 8.4035 | 0.003611 | 0.004848 | 0.00478 | 5.8e-05 | 0.000872 | 0.33876 | 1.2926 | 2.548 | 99.53 | 20.91 | 37.244 | 107.4151 | 108.6458 | 98.1731 | 0.009265 | 0.95078 | 0.022651 | 23.8234 | 0.006423 | 0.15493 | 0.022446 | 1.5192 | 1.5192 | 0.004191
2019-03-21 12:01 | 1553184049.3993 | 936 | 0.2051 | 0.004225 | 8.3e-05 | 100.5047 | 100.4027 | 113.3411 | 21.1233 | 0.003552 | 0.003214 | 140.1681 | 0.26044 | 8.4033 | 0.003559 | 0.005092 | 0.004983 | 5.8e-05 | 0.000873 | 0.30021 | 1.2814 | 2.5479 | 99.53 | 20.91 | 37.243 | 107.209 | 107.3792 | 99.0201 | 0.009265 | 0.94861 | 0.022681 | 23.9098 | 0.009078 | 0.15494 | 0.022141 | 1.5006 | 1.5006 | 0.004244
2019-03-21 12:01 | 1553184050.3993 | 937 | 0.20514 | 0.004198 | 8.2e-05 | 100.389 | 100.1527 | 113.5638 | 21.1229 | 0.003565 | 0.003249 | 140.125 | 0.2611 | 8.4017 | 0.003468 | 0.00529 | 0.005156 | 5.5e-05 | 0.000875 | 0.2959 | 1.2827 | 2.5479 | 99.53 | 20.91 | 37.244 | 107.5434 | 105.1727 | 99.8825 | 0.009265 | 0.94642 | 0.022677 | 23.961 | 0.021083 | 0.15533 | 0.021912 | 1.4849 | 1.4849 | 0.004299
2019-03-21 12:01 | 1553184051.3984 | 938 | 0.20516 | 0.004178 | 8.1e-05 | 100.4617 | 100.2409 | 113.2263 | 21.1231 | 0.003498 | 0.003188 | 140.0684 | 0.26203 | 8.4032 | 0.003495 | 0.004971 | 0.005034 | 5.4e-05 | 0.000874 | 0.30862 | 1.3012 | 2.5479 | 99.53 | 20.91 | 37.243 | 105.8996 | 105.8367 | 98.4004 | 0.009265 | 0.9502 | 0.022689 | 23.8778 | 0.019683 | 0.15588 | 0.021839 | 1.4807 | 1.4807 | 0.004327
2019-03-21 12:01 | 1553184052.3988 | 939 | 0.20521 | 0.004152 | 8.1e-05 | 100.4302 | 100.3312 | 113.2559 | 21.1232 | 0.00359 | 0.003228 | 139.9981 | 0.261 | 8.4026 | 0.003613 | 0.004994 | 0.004858 | 5.4e-05 | 0.000871 | 0.30342 | 1.3063 | 2.548 | 99.53 | 20.91 | 37.244 | 108.1344 | 108.6978 | 99.3613 | 0.009265 | 0.94774 | 0.022662 | 23.9118 | 0.008821 | 0.15527 | 0.021538 | 1.4585 | 1.4585 | 0.004375
2019-03-21 12:01 | 1553184053.3997 | 940 | 0.20525 | 0.004127 | 8e-05 | 100.4078 | 100.2399 | 113.2423 | 21.1229 | 0.003539 | 0.003223 | 139.9479 | 0.26202 | 8.4024 | 0.003596 | 0.005235 | 0.00504 | 5.9e-05 | 0.000875 | 0.31365 | 1.3081 | 2.5479 | 99.53 | 20.91 | 37.244 | 106.9087 | 108.2956 | 99.2503 | 0.009265 | 0.94803 | 0.022664 | 23.9068 | 0.014971 | 0.15588 | 0.021349 | 1.4459 | 1.4459 | 0.004431
2019-03-21 12:01 | 1553184054.3995 | 941 | 0.20525 | 0.004105 | 8.1e-05 | 100.3986 | 100.2416 | 113.0878 | 21.1225 | 0.003538 | 0.003179 | 139.8897 | 0.26034 | 8.4033 | 0.003484 | 0.005231 | 0.005136 | 5.7e-05 | 0.000876 | 0.31277 | 1.3062 | 2.5479 | 99.53 | 20.91 | 37.244 | 106.8847 | 105.5789 | 98.1867 | 0.009265 | 0.95074 | 0.022681 | 23.8563 | 0.013997 | 0.15488 | 0.02139 | 1.4498 | 1.4498 | 0.004391
2019-03-21 12:01 | 1553184055.3998 | 942 | 0.20535 | 0.004076 | 8e-05 | 100.3563 | 100.1895 | 112.8536 | 21.1225 | 0.003524 | 0.003212 | 139.8135 | 0.26133 | 8.4006 | 0.003522 | 0.005475 | 0.00546 | 5.6e-05 | 0.000877 | 0.32333 | 1.3213 | 2.5478 | 99.53 | 20.91 | 37.244 | 106.5256 | 106.498 | 98.9794 | 0.009265 | 0.94872 | 0.022634 | 23.857 | 0.014878 | 0.15547 | 0.020844 | 1.4098 | 1.4098 | 0.004532
2019-03-21 12:01 | 1553184056.4 | 943 | 0.20536 | 0.004049 | 7.9e-05 | 100.3697 | 100.2549 | 112.8946 | 21.1235 | 0.003462 | 0.00332 | 139.7665 | 0.26159 | 8.4019 | 0.003598 | 0.004687 | 0.004582 | 5.3e-05 | 0.00087 | 0.31133 | 1.3406 | 2.548 | 99.53 | 20.9 | 37.244 | 105.0394 | 108.3396 | 101.6052 | 0.00926 | 0.94207 | 0.022558 | 23.9454 | 0.010236 | 0.15562 | 0.020821 | 1.4036 | 1.4036 | 0.004557
2019-03-21 12:01 | 1553184057.3997 | 944 | 0.20533 | 0.004021 | 7.9e-05 | 100.2782 | 100.0522 | 113.1819 | 21.1229 | 0.00352 | 0.003221 | 139.7216 | 0.26124 | 8.4032 | 0.00357 | 0.004518 | 0.004422 | 5.7e-05 | 0.000877 | 0.33987 | 1.3103 | 2.5479 | 99.53 | 20.9 | 37.243 | 106.4489 | 107.6465 | 99.2048 | 0.00926 | 0.94814 | 0.02266 | 23.8989 | 0.02018 | 0.15541 | 0.021018 | 1.4233 | 1.4233 | 0.004488
2019-03-21 12:01 | 1553184058.3987 | 945 | 0.20541 | 0.003998 | 7.8e-05 | 100.3029 | 100.2447 | 113.2737 | 21.1224 | 0.003506 | 0.003196 | 139.6935 | 0.26172 | 8.4029 | 0.0036 | 0.005645 | 0.005565 | 5.2e-05 | 0.000876 | 0.33373 | 1.2786 | 2.548 | 99.53 | 20.9 | 37.242 | 106.0927 | 108.392 | 98.588 | 0.009259 | 0.94972 | 0.022688 | 23.8888 | 0.005198 | 0.1557 | 0.020569 | 1.3947 | 1.3947 | 0.004588
2019-03-21 12:01 | 1553184059.3978 | 946 | 0.20544 | 0.00397 | 7.7e-05 | 100.3098 | 99.9722 | 113.254 | 21.1224 | 0.003544 | 0.003212 | 139.6482 | 0.26158 | 8.4041 | 0.003637 | 0.005184 | 0.005089 | 5.4e-05 | 0.00088 | 0.31454 | 1.2811 | 2.5479 | 99.53 | 20.9 | 37.243 | 107.0254 | 109.2842 | 98.9762 | 0.00926 | 0.94873 | 0.022674 | 23.8992 | 0.030154 | 0.15561 | 0.020391 | 1.3818 | 1.3818 | 0.004629
2019-03-21 12:01 | 1553184060.3989 | 947 | 0.20547 | 0.003944 | 7.7e-05 | 100.2699 | 100.0969 | 113.1632 | 21.1233 | 0.003489 | 0.003116 | 139.616 | 0.2613 | 8.3993 | 0.00358 | 0.005147 | 0.004964 | 5.5e-05 | 0.000875 | 0.30574 | 1.299 | 2.5481 | 99.53 | 20.9 | 37.243 | 105.6992 | 107.9013 | 96.6639 | 0.00926 | 0.95466 | 0.022735 | 23.8153 | 0.015448 | 0.15544 | 0.020282 | 1.3782 | 1.3782 | 0.004636
2019-03-21 12:01 | 1553184061.3989 | 948 | 0.20547 | 0.00392 | 7.6e-05 | 100.2085 | 100.2167 | 113.3254 | 21.1241 | 0.003481 | 0.003229 | 139.5549 | 0.26074 | 8.4017 | 0.003521 | 0.004874 | 0.004796 | 5.2e-05 | 0.000866 | 0.30342 | 1.2996 | 2.5479 | 99.53 | 20.9 | 37.243 | 105.4932 | 106.4729 | 99.3832 | 0.00926 | 0.94769 | 0.022669 | 23.9198 | -0.000736 | 0.15512 | 0.020285 | 1.3743 | 1.3743 | 0.004639
2019-03-21 12:01 | 1553184062.3979 | 949 | 0.20557 | 0.003892 | 7.5e-05 | 100.1683 | 100.0733 | 113.6164 | 21.1234 | 0.003444 | 0.00316 | 139.5087 | 0.26086 | 8.4006 | 0.003613 | 0.004746 | 0.004612 | 4.8e-05 | 0.000861 | 0.31349 | 1.28 | 2.548 | 99.53 | 20.9 | 37.244 | 104.5996 | 108.6934 | 97.7344 | 0.00926 | 0.9519 | 0.022748 | 23.8974 | 0.00848 | 0.15519 | 0.019724 | 1.341 | 1.341 | 0.004756
2019-03-21 12:01 | 1553184063.398 | 950 | 0.20557 | 0.003867 | 7.5e-05 | 100.2153 | 100.0008 | 113.896 | 21.1233 | 0.003499 | 0.003157 | 139.4613 | 0.26097 | 8.4015 | 0.003565 | 0.005178 | 0.005083 | 5.2e-05 | 0.000861 | 0.34578 | 1.2757 | 2.5478 | 99.53 | 20.9 | 37.244 | 105.9384 | 107.5234 | 97.6521 | 0.00926 | 0.95211 | 0.022778 | 23.9242 | 0.019169 | 0.15525 | 0.019776 | 1.3463 | 1.3463 | 0.004739
2019-03-21 12:01 | 1553184064.399 | 951 | 0.2056 | 0.003838 | 7.5e-05 | 100.1903 | 100.2253 | 114.4154 | 21.1232 | 0.003464 | 0.003106 | 139.3906 | 0.26043 | 8.4053 | 0.003497 | 0.005016 | 0.004913 | 5.2e-05 | 0.000862 | 0.31332 | 1.3312 | 2.548 | 99.53 | 20.91 | 37.245 | 105.0721 | 105.8902 | 96.4312 | 0.009266 | 0.95526 | 0.022868 | 23.9391 | -0.003118 | 0.15493 | 0.019608 | 1.3401 | 1.3401 | 0.004752
2019-03-21 12:01 | 1553184065.3995 | 952 | 0.20563 | 0.003816 | 7.5e-05 | 100.1558 | 100.1421 | 114.7034 | 21.1227 | 0.003495 | 0.003128 | 139.3397 | 0.26055 | 8.4025 | 0.003537 | 0.005301 | 0.005208 | 5e-05 | 0.00086 | 0.33212 | 1.2987 | 2.5479 | 99.53 | 20.91 | 37.245 | 105.8245 | 106.8439 | 96.962 | 0.009266 | 0.95389 | 0.02288 | 23.9865 | 0.001216 | 0.155 | 0.01946 | 1.3307 | 1.3307 | 0.004787
2019-03-21 12:01 | 1553184066.3981 | 953 | 0.20565 | 0.003789 | 7.5e-05 | 100.1745 | 99.965 | 114.6572 | 21.1222 | 0.003478 | 0.003183 | 139.2906 | 0.26145 | 8.3993 | 0.003539 | 0.004574 | 0.004527 | 4.9e-05 | 0.000863 | 0.3235 | 1.3356 | 2.548 | 99.53 | 20.91 | 37.246 | 105.4244 | 106.8936 | 98.2734 | 0.009266 | 0.95052 | 0.022835 | 24.0241 | 0.018725 | 0.15554 | 0.019366 | 1.3217 | 1.3217 | 0.004837
2019-03-21 12:01 | 1553184067.3982 | 954 | 0.20569 | 0.003761 | 7.5e-05 | 100.0875 | 100.2399 | 114.6041 | 21.1225 | 0.003493 | 0.003187 | 139.2303 | 0.26098 | 8.4011 | 0.003497 | 0.004611 | 0.004526 | 5e-05 | 0.00086 | 0.33085 | 1.3486 | 2.5481 | 99.53 | 20.91 | 37.246 | 105.7913 | 105.8881 | 98.3865 | 0.009266 | 0.95023 | 0.022827 | 24.0222 | -0.013606 | 0.15526 | 0.019205 | 1.3102 | 1.3102 | 0.00487
2019-03-21 12:01 | 1553184068.3997 | 955 | 0.20571 | 0.00374 | 7.5e-05 | 100.0384 | 99.8758 | 114.3431 | 21.1234 | 0.003471 | 0.003136 | 139.1628 | 0.26141 | 8.403 | 0.003538 | 0.004729 | 0.004677 | 5.2e-05 | 0.00086 | 0.32892 | 1.3534 | 2.5481 | 99.53 | 20.91 | 37.247 | 105.2541 | 106.8846 | 97.1506 | 0.009266 | 0.9534 | 0.022839 | 23.9549 | 0.014547 | 0.15551 | 0.019078 | 1.3022 | 1.3022 | 0.004908
2019-03-21 12:01 | 1553184069.3993 | 956 | 0.20575 | 0.003715 | 7.5e-05 | 100.084 | 99.8949 | 114.1234 | 21.1225 | 0.00352 | 0.003167 | 139.1051 | 0.26126 | 8.4046 | 0.003632 | 0.004949 | 0.004859 | 5.3e-05 | 0.000859 | 0.30215 | 1.3209 | 2.5481 | 99.53 | 20.91 | 37.249 | 106.4453 | 109.1621 | 97.8865 | 0.009267 | 0.95151 | 0.022794 | 23.9556 | 0.016914 | 0.15542 | 0.018895 | 1.2872 | 1.2872 | 0.004963
2019-03-21 12:01 | 1553184070.3998 | 957 | 0.20583 | 0.003685 | 7.5e-05 | 99.9678 | 99.9037 | 114.042 | 21.1222 | 0.003448 | 0.003181 | 139.051 | 0.26018 | 8.4022 | 0.003578 | 0.004929 | 0.004847 | 5.2e-05 | 0.000861 | 0.2808 | 1.2987 | 2.5479 | 99.53 | 20.91 | 37.249 | 104.692 | 107.8473 | 98.2294 | 0.009267 | 0.95063 | 0.022775 | 23.9581 | 0.00573 | 0.15478 | 0.018444 | 1.2554 | 1.2554 | 0.005067
2019-03-21 12:01 | 1553184071.3988 | 958 | 0.20586 | 0.003664 | 7.4e-05 | 100.0226 | 99.7443 | 114.065 | 21.1232 | 0.003495 | 0.003132 | 139.0085 | 0.26082 | 8.402 | 0.003576 | 0.004915 | 0.004887 | 4.6e-05 | 0.000861 | 0.30845 | 1.3031 | 2.5478 | 99.53 | 20.91 | 37.248 | 105.832 | 107.7893 | 97.0383 | 0.009266 | 0.95369 | 0.022814 | 23.9221 | 0.024916 | 0.15516 | 0.018285 | 1.2467 | 1.2467 | 0.005115
2019-03-21 12:01 | 1553184072.3986 | 959 | 0.20593 | 0.003639 | 7.4e-05 | 99.9949 | 99.7341 | 113.7488 | 21.1228 | 0.003467 | 0.003088 | 138.9723 | 0.26098 | 8.4018 | 0.003526 | 0.004552 | 0.00441 | 5e-05 | 0.000862 | 0.28174 | 1.3016 | 2.5479 | 99.53 | 20.91 | 37.244 | 105.1621 | 106.5778 | 95.9977 | 0.009265 | 0.95638 | 0.022815 | 23.8553 | 0.023354 | 0.15526 | 0.01789 | 1.2198 | 1.2198 | 0.005231
2019-03-21 12:01 | 1553184073.3997 | 960 | 0.20594 | 0.003608 | 7.3e-05 | 99.9123 | 99.7712 | 113.6892 | 21.1227 | 0.003497 | 0.003136 | 138.9427 | 0.26078 | 8.3998 | 0.003601 | 0.00482 | 0.004642 | 5.1e-05 | 0.000861 | 0.27201 | 1.3232 | 2.548 | 99.53 | 20.91 | 37.243 | 105.8711 | 108.4034 | 97.1477 | 0.009265 | 0.95341 | 0.022773 | 23.8862 | 0.01264 | 0.15514 | 0.017855 | 1.2153 | 1.2153 | 0.005247
2019-03-21 12:01 | 1553184074.3979 | 961 | 0.20594 | 0.003584 | 7.3e-05 | 99.9451 | 99.7876 | 113.6148 | 21.1224 | 0.003375 | 0.003066 | 138.8904 | 0.26046 | 8.4015 | 0.003575 | 0.005092 | 0.005024 | 5.1e-05 | 0.000865 | 0.33445 | 1.268 | 2.5479 | 99.53 | 20.91 | 37.242 | 102.9354 | 107.7822 | 95.4483 | 0.009265 | 0.95781 | 0.022818 | 23.8235 | 0.014108 | 0.15495 | 0.017901 | 1.2208 | 1.2208 | 0.005217
2019-03-21 12:01 | 1553184075.3983 | 962 | 0.20604 | 0.003553 | 7.3e-05 | 99.9064 | 99.6636 | 113.4336 | 21.123 | 0.003491 | 0.003159 | 138.8663 | 0.26086 | 8.4024 | 0.003527 | 0.00452 | 0.00444 | 5.1e-05 | 0.000865 | 0.30431 | 1.2754 | 2.5479 | 99.53 | 20.91 | 37.242 | 105.7484 | 106.6183 | 97.6948 | 0.009265 | 0.952 | 0.022731 | 23.8769 | 0.02176 | 0.15519 | 0.017348 | 1.1785 | 1.1785 | 0.005412
2019-03-21 12:01 | 1553184076.3977 | 963 | 0.20607 | 0.003528 | 7.3e-05 | 99.9192 | 99.837 | 113.4254 | 21.1222 | 0.003397 | 0.003144 | 138.8361 | 0.25995 | 8.4046 | 0.003602 | 0.005096 | 0.005083 | 5.3e-05 | 0.000867 | 0.30807 | 1.2864 | 2.5478 | 99.53 | 20.91 | 37.241 | 103.4548 | 108.4353 | 97.3422 | 0.009265 | 0.95291 | 0.022741 | 23.8647 | 0.007362 | 0.15464 | 0.017195 | 1.1686 | 1.1686 | 0.005439
2019-03-21 12:01 | 1553184077.3982 | 964 | 0.20602 | 0.003499 | 7.3e-05 | 99.8107 | 99.7324 | 113.3683 | 21.1233 | 0.003415 | 0.003136 | 138.7981 | 0.26015 | 8.4024 | 0.003595 | 0.004985 | 0.004853 | 5.2e-05 | 0.000863 | 0.3178 | 1.3026 | 2.548 | 99.53 | 20.91 | 37.24 | 103.896 | 108.259 | 97.1466 | 0.009264 | 0.95341 | 0.022741 | 23.8524 | 0.007021 | 0.15476 | 0.017513 | 1.1902 | 1.1902 | 0.005344
2019-03-21 12:01 | 1553184078.3978 | 965 | 0.20605 | 0.003477 | 7.3e-05 | 99.8706 | 99.6033 | 113.4925 | 21.1232 | 0.003444 | 0.003101 | 138.7567 | 0.2604 | 8.4002 | 0.003623 | 0.00545 | 0.005344 | 5.2e-05 | 0.000861 | 0.3115 | 1.3353 | 2.5478 | 99.53 | 20.91 | 37.239 | 104.5878 | 108.9485 | 96.3106 | 0.009264 | 0.95557 | 0.022779 | 23.8385 | 0.023971 | 0.15491 | 0.017403 | 1.1847 | 1.1847 | 0.005374
2019-03-21 12:01 | 1553184079.3989 | 966 | 0.20603 | 0.003451 | 7.3e-05 | 99.7836 | 99.8632 | 113.7098 | 21.1244 | 0.003454 | 0.003176 | 138.7244 | 0.26025 | 8.4049 | 0.003671 | 0.005177 | 0.005172 | 5.3e-05 | 0.000855 | 0.2896 | 1.3337 | 2.5478 | 99.53 | 20.91 | 37.237 | 104.8311 | 110.1181 | 98.1251 | 0.009264 | 0.9509 | 0.022745 | 23.9198 | -0.007131 | 0.15482 | 0.017535 | 1.192 | 1.192 | 0.005338
2019-03-21 12:01 | 1553184080.398 | 967 | 0.20613 | 0.003423 | 7.2e-05 | 99.8569 | 99.6807 | 113.942 | 21.1238 | 0.003439 | 0.003134 | 138.674 | 0.26073 | 8.4031 | 0.003596 | 0.005192 | 0.005071 | 5.3e-05 | 0.00086 | 0.31957 | 1.3089 | 2.5479 | 99.53 | 20.91 | 37.236 | 104.4749 | 108.2855 | 97.1056 | 0.009263 | 0.95352 | 0.0228 | 23.9113 | 0.015803 | 0.15511 | 0.016993 | 1.1579 | 1.1579 | 0.005506
2019-03-21 12:01 | 1553184081.399 | 968 | 0.20612 | 0.003406 | 7.2e-05 | 99.8462 | 99.7474 | 114.1319 | 21.1231 | 0.003384 | 0.003103 | 138.6308 | 0.26133 | 8.4021 | 0.003659 | 0.004667 | 0.004534 | 5e-05 | 0.000856 | 0.32405 | 1.2665 | 2.5479 | 99.53 | 20.91 | 37.235 | 103.156 | 109.8147 | 96.3433 | 0.009263 | 0.95549 | 0.022842 | 23.9066 | 0.008854 | 0.15546 | 0.017016 | 1.1616 | 1.1616 | 0.005501
2019-03-21 12:01 | 1553184082.3991 | 969 | 0.20617 | 0.003382 | 7.1e-05 | 99.7649 | 99.6285 | 114.4354 | 21.1239 | 0.003428 | 0.003118 | 138.6046 | 0.26153 | 8.4022 | 0.003609 | 0.004616 | 0.004528 | 5.7e-05 | 0.000861 | 0.29806 | 1.2933 | 2.5479 | 99.53 | 20.91 | 37.234 | 104.1999 | 108.6125 | 96.7038 | 0.009263 | 0.95455 | 0.022862 | 23.9501 | 0.01223 | 0.15559 | 0.016772 | 1.1459 | 1.1459 | 0.00558
2019-03-21 12:01 | 1553184083.3993 | 970 | 0.20618 | 0.00336 | 7e-05 | 99.7604 | 99.6554 | 114.6999 | 21.123 | 0.003388 | 0.003099 | 138.5369 | 0.26149 | 8.4022 | 0.003564 | 0.004657 | 0.004612 | 5.4e-05 | 0.000864 | 0.30414 | 1.3124 | 2.548 | 99.53 | 20.91 | 37.234 | 103.2448 | 107.4993 | 96.2468 | 0.009263 | 0.95574 | 0.022902 | 23.9629 | 0.009414 | 0.15556 | 0.016764 | 1.1475 | 1.1475 | 0.005572
2019-03-21 12:01 | 1553184084.3981 | 971 | 0.2063 | 0.003337 | 7e-05 | 99.6948 | 99.5669 | 114.7184 | 21.1241 | 0.003381 | 0.003081 | 138.502 | 0.26066 | 8.4016 | 0.003586 | 0.005365 | 0.005239 | 5.7e-05 | 0.000866 | 0.29214 | 1.3197 | 2.5479 | 99.53 | 20.91 | 37.235 | 103.0786 | 108.0423 | 95.8236 | 0.009263 | 0.95683 | 0.022917 | 23.9511 | 0.011486 | 0.15507 | 0.016063 | 1.1002 | 1.1002 | 0.005793
2019-03-21 12:01 | 1553184085.3983 | 972 | 0.20634 | 0.003312 | 6.9e-05 | 99.7845 | 99.7627 | 114.8612 | 21.1238 | 0.003354 | 0.003094 | 138.465 | 0.26149 | 8.3984 | 0.003667 | 0.005383 | 0.005216 | 5.9e-05 | 0.000868 | 0.30381 | 1.3091 | 2.5479 | 99.53 | 20.91 | 37.234 | 102.4261 | 110.0033 | 96.1349 | 0.009263 | 0.95602 | 0.022922 | 23.9761 | 0.001953 | 0.15556 | 0.015835 | 1.0848 | 1.0848 | 0.005894
2019-03-21 12:01 | 1553184086.3986 | 973 | 0.20641 | 0.003286 | 6.9e-05 | 99.7002 | 99.7116 | 114.9433 | 21.1236 | 0.003379 | 0.003035 | 138.4113 | 0.26206 | 8.4049 | 0.003642 | 0.005444 | 0.005371 | 5.5e-05 | 0.000867 | 0.30453 | 1.3251 | 2.548 | 99.53 | 20.91 | 37.233 | 103.0199 | 109.395 | 94.7113 | 0.009263 | 0.95972 | 0.022974 | 23.9384 | -0.001026 | 0.1559 | 0.01543 | 1.0594 | 1.0594 | 0.006048
2019-03-21 12:01 | 1553184087.3983 | 974 | 0.20639 | 0.003262 | 6.8e-05 | 99.7383 | 99.6343 | 115.1517 | 21.124 | 0.003317 | 0.003077 | 138.3709 | 0.26049 | 8.4027 | 0.003572 | 0.005109 | 0.005048 | 5.8e-05 | 0.000868 | 0.33179 | 1.3199 | 2.548 | 99.53 | 20.91 | 37.231 | 101.5114 | 107.7114 | 95.7342 | 0.009262 | 0.95706 | 0.022963 | 23.9934 | 0.009332 | 0.15496 | 0.015638 | 1.0732 | 1.0732 | 0.005935
2019-03-21 12:01 | 1553184088.3983 | 975 | 0.20644 | 0.003238 | 6.8e-05 | 99.7312 | 99.6486 | 115.3609 | 21.1249 | 0.003349 | 0.003075 | 138.3462 | 0.26098 | 8.4025 | 0.003603 | 0.004625 | 0.004553 | 5.9e-05 | 0.00087 | 0.32211 | 1.3254 | 2.5479 | 99.53 | 20.91 | 37.23 | 102.2928 | 108.4677 | 95.6826 | 0.009262 | 0.9572 | 0.022986 | 24.0135 | 0.007407 | 0.15526 | 0.015303 | 1.0512 | 1.0512 | 0.00607
2019-03-21 12:01 | 1553184089.3977 | 976 | 0.20648 | 0.003214 | 6.9e-05 | 99.7121 | 99.695 | 115.1832 | 21.1244 | 0.003365 | 0.003003 | 138.3004 | 0.26068 | 8.4038 | 0.003596 | 0.004941 | 0.004805 | 5.9e-05 | 0.00087 | 0.29999 | 1.2893 | 2.5479 | 99.53 | 20.91 | 37.23 | 102.6734 | 108.2987 | 93.9349 | 0.009262 | 0.96175 | 0.023023 | 23.9381 | 0.00154 | 0.15508 | 0.015145 | 1.042 | 1.042 | 0.006117
2019-03-21 12:01 | 1553184090.3975 | 977 | 0.20649 | 0.003179 | 6.8e-05 | 99.7836 | 99.3983 | 114.776 | 21.1243 | 0.003367 | 0.003027 | 138.2755 | 0.26102 | 8.4015 | 0.003581 | 0.005501 | 0.005395 | 6.1e-05 | 0.000869 | 0.32156 | 1.3007 | 2.548 | 99.53 | 20.91 | 37.23 | 102.7392 | 107.9259 | 94.507 | 0.009262 | 0.96026 | 0.022964 | 23.9143 | 0.034607 | 0.15528 | 0.015106 | 1.0366 | 1.0366 | 0.006156
2019-03-21 12:02 | 1553184091.3974 | 978 | 0.20652 | 0.003156 | 6.8e-05 | 99.733 | 99.7471 | 114.5462 | 21.1239 | 0.003354 | 0.003041 | 138.2604 | 0.26042 | 8.4016 | 0.003616 | 0.005222 | 0.005111 | 5.9e-05 | 0.000868 | 0.32068 | 1.2826 | 2.5479 | 99.53 | 20.91 | 37.229 | 102.4269 | 108.7772 | 94.865 | 0.009262 | 0.95932 | 0.02293 | 23.902 | -0.001263 | 0.15492 | 0.014987 | 1.027 | 1.027 | 0.0062
2019-03-21 12:02 | 1553184092.3976 | 979 | 0.20656 | 0.003128 | 6.8e-05 | 99.7148 | 99.6081 | 114.6757 | 21.1247 | 0.003363 | 0.003056 | 138.1943 | 0.26139 | 8.4005 | 0.003551 | 0.0053 | 0.005189 | 6.1e-05 | 0.000867 | 0.3131 | 1.299 | 2.5479 | 99.53 | 20.91 | 37.229 | 102.6377 | 107.1963 | 95.2259 | 0.009262 | 0.95838 | 0.022931 | 23.9272 | 0.009576 | 0.1555 | 0.014749 | 1.0107 | 1.0107 | 0.006324
2019-03-21 12:02 | 1553184093.3986 | 980 | 0.2066 | 0.003107 | 6.8e-05 | 99.6719 | 99.6319 | 114.4776 | 21.1249 | 0.003352 | 0.003043 | 138.1815 | 0.26003 | 8.4048 | 0.003593 | 0.005078 | 0.004981 | 6.1e-05 | 0.000871 | 0.31868 | 1.306 | 2.5479 | 99.53 | 20.9 | 37.229 | 102.3778 | 108.2196 | 94.9145 | 0.009256 | 0.95919 | 0.022921 | 23.8965 | 0.003586 | 0.15469 | 0.014541 | 0.99597 | 0.99597 | 0.006383
2019-03-21 12:02 | 1553184094.3977 | 981 | 0.20665 | 0.003078 | 6.8e-05 | 99.6731 | 99.7621 | 114.2173 | 21.1248 | 0.003362 | 0.003029 | 138.1619 | 0.26078 | 8.4028 | 0.003611 | 0.005497 | 0.005367 | 5.9e-05 | 0.000868 | 0.29878 | 1.3257 | 2.548 | 99.53 | 20.9 | 37.23 | 102.6077 | 108.6632 | 94.5583 | 0.009256 | 0.96012 | 0.022906 | 23.8577 | -0.00798 | 0.15514 | 0.014287 | 0.9779 | 0.9779 | 0.00652
2019-03-21 12:02 | 1553184095.3987 | 982 | 0.20668 | 0.003053 | 6.8e-05 | 99.688 | 99.7311 | 114.4682 | 21.1242 | 0.003305 | 0.003038 | 138.1334 | 0.26009 | 8.4009 | 0.003592 | 0.005142 | 0.005043 | 6.1e-05 | 0.000866 | 0.29645 | 1.3315 | 2.548 | 99.53 | 20.91 | 37.232 | 101.2299 | 108.196 | 94.7736 | 0.009262 | 0.95956 | 0.022925 | 23.8909 | -0.003864 | 0.15473 | 0.01412 | 0.96724 | 0.96724 | 0.006575
2019-03-21 12:02 | 1553184096.3981 | 983 | 0.20672 | 0.003033 | 6.7e-05 | 99.6677 | 99.6796 | 114.7424 | 21.124 | 0.003295 | 0.002953 | 138.0881 | 0.26175 | 8.4034 | 0.003699 | 0.005038 | 0.0049 | 5.9e-05 | 0.000867 | 0.3079 | 1.3592 | 2.5481 | 99.53 | 20.91 | 37.233 | 100.9996 | 110.7929 | 92.7256 | 0.009263 | 0.96493 | 0.023016 | 23.8528 | -0.001068 | 0.15572 | 0.013904 | 0.95623 | 0.95623 | 0.006693
2019-03-21 12:02 | 1553184097.398 | 984 | 0.20675 | 0.003001 | 6.8e-05 | 99.6999 | 99.6486 | 114.7567 | 21.1255 | 0.003349 | 0.003039 | 138.0571 | 0.26098 | 8.4018 | 0.003666 | 0.005709 | 0.005581 | 5.9e-05 | 0.000867 | 0.32516 | 1.3595 | 2.5479 | 99.53 | 20.91 | 37.233 | 102.3036 | 109.9797 | 94.7952 | 0.009263 | 0.95951 | 0.022953 | 23.9217 | 0.004601 | 0.15526 | 0.013791 | 0.94581 | 0.94581 | 0.006747
2019-03-21 12:02 | 1553184098.3989 | 985 | 0.20677 | 0.002974 | 6.7e-05 | 99.7121 | 99.5448 | 115.0083 | 21.125 | 0.003328 | 0.002995 | 138.039 | 0.26132 | 8.4006 | 0.003733 | 0.004935 | 0.004756 | 6e-05 | 0.000864 | 0.29983 | 1.291 | 2.5479 | 99.53 | 20.91 | 37.235 | 101.7855 | 111.6258 | 93.7538 | 0.009263 | 0.96223 | 0.023011 | 23.914 | 0.015026 | 0.15546 | 0.013693 | 0.94148 | 0.94148 | 0.006786
2019-03-21 12:02 | 1553184099.3984 | 986 | 0.20686 | 0.002947 | 6.8e-05 | 99.6805 | 99.6101 | 115.7113 | 21.1238 | 0.003285 | 0.003012 | 138.0021 | 0.26209 | 8.3991 | 0.003672 | 0.004829 | 0.004734 | 5.9e-05 | 0.000865 | 0.33124 | 1.3122 | 2.5478 | 99.53 | 20.9 | 37.236 | 100.7451 | 110.1238 | 94.1518 | 0.009258 | 0.96119 | 0.023068 | 24 | 0.006318 | 0.15592 | 0.013153 | 0.90645 | 0.90645 | 0.00707
2019-03-21 12:02 | 1553184100.398 | 987 | 0.20683 | 0.002919 | 6.8e-05 | 99.6436 | 99.4974 | 116.1996 | 21.1246 | 0.003253 | 0.002978 | 137.9661 | 0.26048 | 8.403 | 0.003646 | 0.005312 | 0.005201 | 6.3e-05 | 0.000861 | 0.29894 | 1.3209 | 2.548 | 99.53 | 20.91 | 37.236 | 99.9713 | 109.4906 | 93.3334 | 0.009263 | 0.96333 | 0.023143 | 24.0237 | 0.013131 | 0.15496 | 0.01341 | 0.92718 | 0.92718 | 0.006869
2019-03-21 12:02 | 1553184101.398 | 988 | 0.20689 | 0.002896 | 6.9e-05 | 99.6284 | 99.5046 | 116.3279 | 21.1248 | 0.003297 | 0.003018 | 137.9151 | 0.26073 | 8.402 | 0.003647 | 0.005701 | 0.005612 | 6.1e-05 | 0.000862 | 0.31758 | 1.3233 | 2.548 | 99.53 | 20.91 | 37.238 | 101.0465 | 109.5201 | 94.2999 | 0.009264 | 0.9608 | 0.023125 | 24.0687 | 0.011124 | 0.15511 | 0.013048 | 0.90134 | 0.90134 | 0.007073
2019-03-21 12:02 | 1553184102.3995 | 989 | 0.20686 | 0.002871 | 6.9e-05 | 99.6078 | 99.5683 | 116.1777 | 21.1252 | 0.003291 | 0.002979 | 137.8946 | 0.26084 | 8.4029 | 0.003601 | 0.005257 | 0.005172 | 6.1e-05 | 0.000862 | 0.32731 | 1.3271 | 2.5481 | 99.53 | 20.91 | 37.238 | 100.8931 | 108.402 | 93.366 | 0.009264 | 0.96325 | 0.02314 | 24.0226 | 0.003555 | 0.15517 | 0.013268 | 0.91718 | 0.91718 | 0.006954
2019-03-21 12:02 | 1553184103.3982 | 990 | 0.20695 | 0.002844 | 7e-05 | 99.6755 | 99.6125 | 116.2291 | 21.1251 | 0.003289 | 0.003011 | 137.8527 | 0.26137 | 8.4025 | 0.003627 | 0.0047 | 0.004581 | 6e-05 | 0.000863 | 0.31404 | 1.3722 | 2.5481 | 99.53 | 20.91 | 37.238 | 100.8577 | 109.0441 | 94.1391 | 0.009264 | 0.96122 | 0.02312 | 24.0532 | 0.00565 | 0.15549 | 0.012787 | 0.88302 | 0.88302 | 0.007237
2019-03-21 12:02 | 1553184104.3992 | 991 | 0.20695 | 0.00281 | 7e-05 | 99.6087 | 99.57 | 116.3679 | 21.1246 | 0.003298 | 0.002996 | 137.8334 | 0.26136 | 8.4019 | 0.00362 | 0.005166 | 0.005061 | 5.9e-05 | 0.000864 | 0.34274 | 1.3624 | 2.5479 | 99.53 | 20.91 | 37.242 | 101.0582 | 108.867 | 93.7628 | 0.009265 | 0.9622 | 0.023146 | 24.0552 | 0.003483 | 0.15548 | 0.01282 | 0.88631 | 0.88631 | 0.00721
2019-03-21 12:02 | 1553184105.3973 | 992 | 0.20697 | 0.002783 | 7.1e-05 | 99.5393 | 99.5199 | 116.5411 | 21.1241 | 0.003318 | 0.00299 | 137.7907 | 0.26097 | 8.4007 | 0.003616 | 0.004894 | 0.004821 | 5.9e-05 | 0.00086 | 0.30879 | 1.3267 | 2.548 | 99.53 | 20.91 | 37.243 | 101.5352 | 108.7672 | 93.6246 | 0.009265 | 0.96257 | 0.023168 | 24.0686 | 0.001745 | 0.15525 | 0.012724 | 0.8804 | 0.8804 | 0.007247
2019-03-21 12:02 | 1553184106.3972 | 993 | 0.20706 | 0.002763 | 7.1e-05 | 99.5974 | 99.5247 | 116.3129 | 21.125 | 0.003316 | 0.002962 | 137.7746 | 0.26038 | 8.4032 | 0.003653 | 0.004769 | 0.004715 | 5.8e-05 | 0.000861 | 0.31116 | 1.3055 | 2.548 | 99.53 | 20.91 | 37.243 | 101.5109 | 109.6684 | 92.941 | 0.009265 | 0.96437 | 0.023167 | 24.0226 | 0.006536 | 0.1549 | 0.012214 | 0.84493 | 0.84493 | 0.007535
2019-03-21 12:02 | 1553184107.3974 | 994 | 0.207 | 0.00273 | 7.2e-05 | 99.5977 | 99.4225 | 116.3704 | 21.1241 | 0.003237 | 0.002945 | 137.7629 | 0.26064 | 8.4034 | 0.003587 | 0.005273 | 0.005206 | 6.1e-05 | 0.000856 | 0.32084 | 1.2928 | 2.548 | 99.53 | 20.91 | 37.245 | 99.5866 | 108.0644 | 92.5452 | 0.009266 | 0.96541 | 0.023185 | 24.0155 | 0.01575 | 0.15505 | 0.012574 | 0.87054 | 0.87054 | 0.00732
2019-03-21 12:02 | 1553184108.3984 | 995 | 0.20711 | 0.002707 | 7.4e-05 | 99.6361 | 99.4974 | 116.2826 | 21.1255 | 0.003254 | 0.002917 | 137.6929 | 0.26159 | 8.4013 | 0.003641 | 0.005539 | 0.005388 | 5.9e-05 | 0.000858 | 0.33699 | 1.2927 | 2.5479 | 99.53 | 20.91 | 37.246 | 99.9926 | 109.3708 | 91.863 | 0.009266 | 0.96721 | 0.023198 | 23.9841 | 0.012463 | 0.15562 | 0.011972 | 0.82913 | 0.82913 | 0.007714
2019-03-21 12:02 | 1553184109.3984 | 996 | 0.20707 | 0.002677 | 7.5e-05 | 99.525 | 99.5502 | 115.9108 | 21.1255 | 0.003202 | 0.002937 | 137.6572 | 0.26079 | 8.4022 | 0.003594 | 0.004892 | 0.00483 | 6.1e-05 | 0.000855 | 0.33339 | 1.3083 | 2.5479 | 99.53 | 20.91 | 37.249 | 98.7424 | 108.2336 | 92.3362 | 0.009267 | 0.96596 | 0.023146 | 23.9612 | -0.002264 | 0.15514 | 0.012229 | 0.84503 | 0.84503 | 0.007546
2019-03-21 12:02 | 1553184110.3977 | 997 | 0.20715 | 0.002657 | 7.6e-05 | 99.5512 | 99.5396 | 115.9714 | 21.1245 | 0.003248 | 0.002941 | 137.6249 | 0.26027 | 8.4027 | 0.003601 | 0.005457 | 0.005383 | 5.9e-05 | 0.000857 | 0.33069 | 1.3236 | 2.548 | 99.53 | 20.91 | 37.25 | 99.8504 | 108.4083 | 92.4501 | 0.009267 | 0.96566 | 0.023148 | 23.9712 | 0.001041 | 0.15483 | 0.011777 | 0.8137 | 0.8137 | 0.007821
2019-03-21 12:02 | 1553184111.3986 | 998 | 0.20714 | 0.002628 | 7.6e-05 | 99.5658 | 99.3568 | 115.8458 | 21.1258 | 0.003229 | 0.002966 | 137.6013 | 0.26036 | 8.4006 | 0.003568 | 0.005328 | 0.005304 | 6.1e-05 | 0.000853 | 0.31924 | 1.3078 | 2.548 | 99.53 | 20.91 | 37.25 | 99.4076 | 107.6115 | 93.0396 | 0.009267 | 0.9641 | 0.023117 | 23.9775 | 0.018802 | 0.15489 | 0.011891 | 0.82047 | 0.82047 | 0.007759
2019-03-21 12:02 | 1553184112.3986 | 999 | 0.20719 | 0.002602 | 7.7e-05 | 99.5533 | 99.3772 | 115.8946 | 21.1252 | 0.003194 | 0.002962 | 137.5693 | 0.26074 | 8.4029 | 0.003613 | 0.004882 | 0.004818 | 5.9e-05 | 0.000862 | 0.30735 | 1.3247 | 2.5479 | 99.53 | 20.91 | 37.253 | 98.5383 | 108.6903 | 92.9545 | 0.009268 | 0.96433 | 0.023124 | 23.9798 | 0.015839 | 0.15511 | 0.011595 | 0.80016 | 0.80016 | 0.007967
2019-03-21 12:02 | 1553184113.3987 | 1000 | 0.20722 | 0.002577 | 7.9e-05 | 99.531 | 99.4456 | 115.8554 | 21.1255 | 0.003227 | 0.002986 | 137.5558 | 0.26065 | 8.3998 | 0.003559 | 0.005273 | 0.005189 | 6e-05 | 0.000857 | 0.31188 | 1.2992 | 2.548 | 99.53 | 20.91 | 37.253 | 99.3429 | 107.3859 | 93.5252 | 0.009268 | 0.96283 | 0.023103 | 23.9944 | 0.007671 | 0.15506 | 0.011439 | 0.78853 | 0.78853 | 0.008082
2019-03-21 12:02 | 1553184114.3981 | 1001 | 0.20723 | 0.002552 | 8e-05 | 99.5989 | 99.556 | 116.0744 | 21.125 | 0.003169 | 0.002943 | 137.506 | 0.26059 | 8.4036 | 0.003569 | 0.00493 | 0.00485 | 5.9e-05 | 0.000852 | 0.29574 | 1.3046 | 2.548 | 99.53 | 20.91 | 37.254 | 97.957 | 107.6289 | 92.4858 | 0.009268 | 0.96557 | 0.023157 | 23.983 | 0.003854 | 0.15502 | 0.01143 | 0.78965 | 0.78965 | 0.008069
2019-03-21 12:02 | 1553184115.3988 | 1002 | 0.20726 | 0.002526 | 8.1e-05 | 99.5774 | 99.5093 | 116.2298 | 21.1247 | 0.003189 | 0.002904 | 137.4796 | 0.25999 | 8.4015 | 0.003541 | 0.005312 | 0.005233 | 5.9e-05 | 0.000859 | 0.31332 | 1.2865 | 2.5479 | 99.53 | 20.91 | 37.255 | 98.4233 | 106.9451 | 91.5424 | 0.009268 | 0.96806 | 0.023203 | 23.9681 | 0.00612 | 0.15467 | 0.011286 | 0.78116 | 0.78116 | 0.008138
2019-03-21 12:02 | 1553184116.3979 | 1003 | 0.20728 | 0.002504 | 8.3e-05 | 99.5048 | 99.4691 | 116.5178 | 21.1245 | 0.003203 | 0.002898 | 137.4572 | 0.26039 | 8.4003 | 0.003601 | 0.004823 | 0.004686 | 6.2e-05 | 0.000862 | 0.33373 | 1.2895 | 2.5479 | 99.53 | 20.91 | 37.256 | 98.7735 | 108.413 | 91.4109 | 0.009268 | 0.96841 | 0.023236 | 23.9934 | 0.003201 | 0.15491 | 0.011172 | 0.77426 | 0.77426 | 0.008223
2019-03-21 12:02 | 1553184117.3988 | 1004 | 0.2073 | 0.002478 | 8.4e-05 | 99.5548 | 99.4508 | 116.4289 | 21.1251 | 0.003199 | 0.002937 | 137.4149 | 0.26053 | 8.4006 | 0.003638 | 0.004914 | 0.004806 | 5.9e-05 | 0.000859 | 0.32532 | 1.2836 | 2.5477 | 99.53 | 20.91 | 37.256 | 98.6801 | 109.3083 | 92.3358 | 0.009268 | 0.96596 | 0.023197 | 24.0147 | 0.009354 | 0.15499 | 0.011111 | 0.76864 | 0.76864 | 0.008287
2019-03-21 12:02 | 1553184118.3979 | 1005 | 0.20734 | 0.002453 | 8.5e-05 | 99.4902 | 99.398 | 116.1648 | 21.1242 | 0.003171 | 0.002893 | 137.4037 | 0.25987 | 8.4028 | 0.003585 | 0.004472 | 0.004355 | 5.9e-05 | 0.000862 | 0.28208 | 1.2689 | 2.5479 | 99.53 | 20.91 | 37.256 | 97.9908 | 108.0149 | 91.28 | 0.009268 | 0.96876 | 0.023204 | 23.9528 | 0.008292 | 0.1546 | 0.010883 | 0.75293 | 0.75293 | 0.008439
2019-03-21 12:02 | 1553184119.399 | 1006 | 0.20742 | 0.002427 | 8.7e-05 | 99.5611 | 99.3663 | 115.9538 | 21.1257 | 0.003163 | 0.002937 | 137.3821 | 0.25986 | 8.4029 | 0.003597 | 0.005034 | 0.00501 | 5.9e-05 | 0.000863 | 0.33301 | 1.3221 | 2.5479 | 99.53 | 20.91 | 37.254 | 97.8019 | 108.311 | 92.3419 | 0.009268 | 0.96595 | 0.02315 | 23.9658 | 0.017516 | 0.15459 | 0.010449 | 0.72091 | 0.72091 | 0.008813
2019-03-21 12:02 | 1553184120.398 | 1007 | 0.20744 | 0.002399 | 8.9e-05 | 99.4907 | 99.3918 | 116.0198 | 21.1258 | 0.003188 | 0.002884 | 137.3548 | 0.26054 | 8.4028 | 0.00362 | 0.00454 | 0.004493 | 5.9e-05 | 0.000857 | 0.30934 | 1.3377 | 2.548 | 99.53 | 20.92 | 37.252 | 98.3968 | 108.8604 | 91.076 | 0.009273 | 0.9693 | 0.023196 | 23.9311 | 0.008897 | 0.155 | 0.010322 | 0.71346 | 0.71346 | 0.008929
2019-03-21 12:02 | 1553184121.3971 | 1008 | 0.20745 | 0.002376 | 9.1e-05 | 99.5241 | 99.5764 | 116.3322 | 21.1257 | 0.003133 | 0.002868 | 137.3261 | 0.2596 | 8.401 | 0.003583 | 0.005169 | 0.005085 | 5.7e-05 | 0.000857 | 0.33516 | 1.3051 | 2.5481 | 99.53 | 20.92 | 37.252 | 97.0666 | 107.969 | 90.6859 | 0.009273 | 0.97034 | 0.02324 | 23.9505 | -0.0047 | 0.15443 | 0.010289 | 0.71236 | 0.71236 | 0.00891
2019-03-21 12:02 | 1553184122.3971 | 1009 | 0.20742 | 0.002356 | 9.3e-05 | 99.5521 | 99.351 | 116.6178 | 21.1252 | 0.003139 | 0.002902 | 137.3057 | 0.26097 | 8.4008 | 0.003601 | 0.005256 | 0.005165 | 6.4e-05 | 0.000859 | 0.3178 | 1.3344 | 2.5479 | 99.53 | 20.92 | 37.248 | 97.2103 | 108.4098 | 91.4962 | 0.009272 | 0.96819 | 0.023243 | 24.0065 | 0.018092 | 0.15525 | 0.01052 | 0.72844 | 0.72844 | 0.008759
2019-03-21 12:02 | 1553184123.3983 | 1010 | 0.2075 | 0.002333 | 9.5e-05 | 99.4764 | 99.4906 | 116.6732 | 21.1247 | 0.003125 | 0.002869 | 137.2633 | 0.2607 | 8.4042 | 0.00358 | 0.004966 | 0.004816 | 6.2e-05 | 0.000859 | 0.31094 | 1.3606 | 2.548 | 99.53 | 20.92 | 37.247 | 96.8823 | 107.901 | 90.7006 | 0.009272 | 0.9703 | 0.023274 | 23.986 | -0.001273 | 0.15509 | 0.010047 | 0.69627 | 0.69627 | 0.009155
2019-03-21 12:02 | 1553184124.3986 | 1011 | 0.2075 | 0.002305 | 9.7e-05 | 99.4496 | 99.3993 | 116.47 | 21.1255 | 0.003198 | 0.002867 | 137.2424 | 0.26177 | 8.4023 | 0.003563 | 0.004773 | 0.004704 | 6.1e-05 | 0.000858 | 0.3449 | 1.341 | 2.5479 | 99.53 | 20.92 | 37.244 | 98.6403 | 107.489 | 90.656 | 0.009271 | 0.97042 | 0.023255 | 23.9637 | 0.004526 | 0.15572 | 0.010117 | 0.70043 | 0.70043 | 0.009138
2019-03-21 12:02 | 1553184125.3986 | 1012 | 0.20749 | 0.002285 | 0.0001 | 99.4312 | 99.3612 | 115.9012 | 21.1258 | 0.00309 | 0.002868 | 137.2284 | 0.26127 | 8.4012 | 0.003544 | 0.005564 | 0.005506 | 5.9e-05 | 0.000858 | 0.3465 | 1.3054 | 2.548 | 99.53 | 20.92 | 37.241 | 96.0444 | 107.0178 | 90.6946 | 0.00927 | 0.97032 | 0.023197 | 23.9063 | 0.006297 | 0.15543 | 0.010183 | 0.70312 | 0.70312 | 0.009085
2019-03-21 12:02 | 1553184126.3984 | 1013 | 0.20753 | 0.002264 | 0.000103 | 99.4815 | 99.3264 | 115.8688 | 21.1259 | 0.003124 | 0.002815 | 137.19 | 0.26072 | 8.4029 | 0.003596 | 0.004817 | 0.004696 | 5.8e-05 | 0.000858 | 0.29214 | 1.3094 | 2.5478 | 99.53 | 20.92 | 37.239 | 96.8521 | 108.2978 | 89.417 | 0.00927 | 0.97374 | 0.023234 | 23.861 | 0.013954 | 0.1551 | 0.009977 | 0.68979 | 0.68979 | 0.009242
2019-03-21 12:02 | 1553184127.3988 | 1014 | 0.20759 | 0.002242 | 0.000106 | 99.5313 | 99.4112 | 115.9245 | 21.1262 | 0.003135 | 0.002891 | 137.1698 | 0.26043 | 8.4034 | 0.003587 | 0.004849 | 0.0048 | 6.4e-05 | 0.000856 | 0.30757 | 1.3141 | 2.548 | 99.53 | 20.92 | 37.235 | 97.1229 | 108.0633 | 91.2335 | 0.009269 | 0.96888 | 0.023182 | 23.9264 | 0.010792 | 0.15493 | 0.009648 | 0.66512 | 0.66512 | 0.009573
2019-03-21 12:02 | 1553184128.3984 | 1015 | 0.20764 | 0.002216 | 0.00011 | 99.4833 | 99.3241 | 116.4035 | 21.1264 | 0.003131 | 0.00287 | 137.1695 | 0.26159 | 8.4002 | 0.003559 | 0.004712 | 0.004622 | 6.2e-05 | 0.000858 | 0.33516 | 1.2817 | 2.548 | 99.53 | 20.92 | 37.234 | 97.0173 | 107.3876 | 90.7374 | 0.009269 | 0.9702 | 0.023246 | 23.9595 | 0.014329 | 0.15562 | 0.009369 | 0.64727 | 0.64727 | 0.009881
2019-03-21 12:02 | 1553184129.3989 | 1016 | 0.20767 | 0.002194 | 0.000114 | 99.4869 | 99.5039 | 116.8361 | 21.1249 | 0.00314 | 0.002861 | 137.1737 | 0.26097 | 8.4046 | 0.003542 | 0.004399 | 0.004278 | 6e-05 | 0.000855 | 0.32189 | 1.3369 | 2.5479 | 99.53 | 20.92 | 37.232 | 97.2494 | 106.9776 | 90.5194 | 0.009268 | 0.97079 | 0.023296 | 23.9968 | -0.001529 | 0.15525 | 0.009177 | 0.63506 | 0.63506 | 0.010048
2019-03-21 12:02 | 1553184130.399 | 1017 | 0.20764 | 0.002172 | 0.000117 | 99.5182 | 99.4743 | 116.842 | 21.1257 | 0.003091 | 0.002879 | 137.1517 | 0.26063 | 8.3998 | 0.003484 | 0.005489 | 0.005409 | 5.7e-05 | 0.00086 | 0.31758 | 1.285 | 2.5479 | 99.53 | 20.92 | 37.229 | 96.0516 | 105.5782 | 90.9555 | 0.009267 | 0.96962 | 0.023282 | 24.0118 | 0.003947 | 0.15505 | 0.009405 | 0.65044 | 0.65044 | 0.009797
2019-03-21 12:02 | 1553184131.3977 | 1018 | 0.20766 | 0.002146 | 0.000121 | 99.4598 | 99.2249 | 116.7406 | 21.126 | 0.003099 | 0.00283 | 137.1235 | 0.26079 | 8.4009 | 0.003552 | 0.004349 | 0.004289 | 6.2e-05 | 0.000858 | 0.28694 | 1.2836 | 2.548 | 99.53 | 20.92 | 37.229 | 96.2648 | 107.2257 | 89.7593 | 0.009267 | 0.97282 | 0.023311 | 23.9619 | 0.021142 | 0.15515 | 0.009271 | 0.64163 | 0.64163 | 0.009938
2019-03-21 12:02 | 1553184132.3987 | 1019 | 0.20772 | 0.002126 | 0.000124 | 99.4604 | 99.2614 | 116.4974 | 21.126 | 0.003118 | 0.002784 | 137.0933 | 0.26031 | 8.4018 | 0.003479 | 0.004953 | 0.004801 | 5.9e-05 | 0.000859 | 0.30359 | 1.2794 | 2.548 | 99.53 | 20.92 | 37.227 | 96.723 | 105.4516 | 88.6562 | 0.009267 | 0.97579 | 0.023322 | 23.9005 | 0.017911 | 0.15486 | 0.008986 | 0.62178 | 0.62178 | 0.010236
2019-03-21 12:02 | 1553184133.3991 | 1020 | 0.20778 | 0.0021 | 0.000127 | 99.4318 | 99.5826 | 116.5293 | 21.1252 | 0.003097 | 0.002808 | 137.0659 | 0.26088 | 8.4032 | 0.003527 | 0.004696 | 0.004692 | 6e-05 | 0.000857 | 0.30989 | 1.2908 | 2.5479 | 99.53 | 20.93 | 37.225 | 96.1991 | 106.603 | 89.2367 | 0.009272 | 0.97422 | 0.023306 | 23.923 | -0.013556 | 0.1552 | 0.008624 | 0.59592 | 0.59592 | 0.010704
2019-03-21 12:02 | 1553184134.3978 | 1021 | 0.20782 | 0.002078 | 0.00013 | 99.4824 | 99.4777 | 116.8594 | 21.1256 | 0.003096 | 0.002838 | 137.0603 | 0.26066 | 8.4051 | 0.003539 | 0.004935 | 0.004851 | 5.9e-05 | 0.000858 | 0.27312 | 1.2933 | 2.548 | 99.53 | 20.93 | 37.224 | 96.1838 | 106.8931 | 89.9723 | 0.009272 | 0.97225 | 0.023316 | 23.9811 | 0.000427 | 0.15507 | 0.0084 | 0.58031 | 0.58031 | 0.010982
2019-03-21 12:02 | 1553184135.3979 | 1022 | 0.20789 | 0.002054 | 0.000132 | 99.4276 | 99.4794 | 117.2457 | 21.1269 | 0.003011 | 0.002851 | 137.0515 | 0.26089 | 8.4011 | 0.003563 | 0.004785 | 0.004628 | 6.1e-05 | 0.00086 | 0.29966 | 1.323 | 2.548 | 99.53 | 20.93 | 37.224 | 94.1315 | 107.4841 | 90.2782 | 0.009272 | 0.97143 | 0.023344 | 24.0308 | -0.004657 | 0.1552 | 0.008036 | 0.55542 | 0.55542 | 0.011485
2019-03-21 12:02 | 1553184136.3973 | 1023 | 0.20788 | 0.002031 | 0.000134 | 99.5077 | 99.2658 | 117.4783 | 21.1259 | 0.003037 | 0.002795 | 137.054 | 0.26122 | 8.4018 | 0.003628 | 0.004922 | 0.004818 | 5.9e-05 | 0.000861 | 0.31133 | 1.3071 | 2.548 | 99.53 | 20.93 | 37.224 | 94.7534 | 109.0576 | 88.921 | 0.009272 | 0.97507 | 0.023411 | 24.0097 | 0.021772 | 0.1554 | 0.008128 | 0.56336 | 0.56336 | 0.011337
2019-03-21 12:02 | 1553184137.3973 | 1024 | 0.20789 | 0.002009 | 0.000137 | 99.337 | 99.367 | 117.8795 | 21.1253 | 0.003079 | 0.002779 | 137.0438 | 0.26168 | 8.4004 | 0.003595 | 0.004919 | 0.004848 | 5.7e-05 | 0.00086 | 0.31349 | 1.316 | 2.548 | 99.53 | 20.92 | 37.224 | 95.7839 | 108.2673 | 88.5431 | 0.009266 | 0.97609 | 0.023463 | 24.0381 | -0.002697 | 0.15567 | 0.008063 | 0.55985 | 0.55985 | 0.011428
2019-03-21 12:02 | 1553184138.397 | 1025 | 0.20791 | 0.001981 | 0.00014 | 99.458 | 99.4232 | 117.5783 | 21.1261 | 0.00302 | 0.002811 | 136.9956 | 0.2605 | 8.4031 | 0.003635 | 0.005101 | 0.005031 | 5.5e-05 | 0.000864 | 0.31708 | 1.3467 | 2.5479 | 99.53 | 20.92 | 37.222 | 94.3458 | 109.2351 | 89.3093 | 0.009266 | 0.97403 | 0.023409 | 24.0328 | 0.003131 | 0.15497 | 0.007978 | 0.55238 | 0.55238 | 0.011531
2019-03-21 12:02 | 1553184139.3984 | 1026 | 0.20796 | 0.001952 | 0.000144 | 99.4374 | 99.2273 | 117.2457 | 21.1258 | 0.003074 | 0.002778 | 136.9499 | 0.26078 | 8.4013 | 0.003613 | 0.005443 | 0.005342 | 5.8e-05 | 0.000863 | 0.31454 | 1.3122 | 2.548 | 99.53 | 20.92 | 37.222 | 95.6591 | 108.6901 | 88.5099 | 0.009266 | 0.97618 | 0.023401 | 23.9723 | 0.018918 | 0.15514 | 0.007698 | 0.53236 | 0.53236 | 0.011977
2019-03-21 12:02 | 1553184140.3976 | 1027 | 0.20797 | 0.001933 | 0.000147 | 99.4181 | 99.3721 | 117.218 | 21.1251 | 0.003056 | 0.002806 | 136.9285 | 0.261 | 8.4035 | 0.003543 | 0.005221 | 0.005135 | 6.2e-05 | 0.000866 | 0.29192 | 1.3021 | 2.548 | 99.53 | 20.92 | 37.221 | 95.2248 | 106.9974 | 89.184 | 0.009265 | 0.97436 | 0.023377 | 23.9918 | 0.004137 | 0.15527 | 0.007693 | 0.53124 | 0.53124 | 0.012012
2019-03-21 12:02 | 1553184141.3982 | 1028 | 0.20796 | 0.00191 | 0.00015 | 99.4234 | 99.2467 | 117.2941 | 21.1264 | 0.003016 | 0.002823 | 136.9383 | 0.26117 | 8.4037 | 0.003596 | 0.004185 | 0.004157 | 6.2e-05 | 0.000861 | 0.29894 | 1.3225 | 2.548 | 99.53 | 20.92 | 37.221 | 94.2638 | 108.2943 | 89.5971 | 0.009265 | 0.97325 | 0.023371 | 24.0133 | 0.015908 | 0.15537 | 0.007743 | 0.53445 | 0.53445 | 0.011948
2019-03-21 12:02 | 1553184142.3971 | 1029 | 0.20804 | 0.001888 | 0.000153 | 99.4121 | 99.1963 | 117.6426 | 21.1257 | 0.003047 | 0.002756 | 136.9428 | 0.26087 | 8.4037 | 0.003648 | 0.004862 | 0.004779 | 6e-05 | 0.000862 | 0.33516 | 1.2912 | 2.5479 | 99.53 | 20.92 | 37.221 | 94.9889 | 109.5579 | 87.9802 | 0.009265 | 0.97761 | 0.023458 | 23.9953 | 0.019433 | 0.15519 | 0.007277 | 0.50351 | 0.50351 | 0.012668
2019-03-21 12:02 | 1553184143.3973 | 1030 | 0.20806 | 0.001859 | 0.000155 | 99.4094 | 99.3006 | 117.6001 | 21.1266 | 0.002997 | 0.00277 | 136.9562 | 0.26117 | 8.4029 | 0.003623 | 0.005077 | 0.005009 | 5.7e-05 | 0.000862 | 0.32211 | 1.2943 | 2.5479 | 99.53 | 20.92 | 37.223 | 93.7892 | 108.947 | 88.328 | 0.009266 | 0.97667 | 0.023443 | 24.0025 | 0.0098 | 0.15537 | 0.007237 | 0.50025 | 0.50025 | 0.012765
2019-03-21 12:02 | 1553184144.3983 | 1031 | 0.20807 | 0.001842 | 0.000158 | 99.3343 | 99.3404 | 117.5241 | 21.1259 | 0.003003 | 0.002749 | 136.936 | 0.26167 | 8.4001 | 0.003612 | 0.005092 | 0.004914 | 6e-05 | 0.000862 | 0.30431 | 1.2931 | 2.548 | 99.53 | 20.92 | 37.224 | 93.9402 | 108.6864 | 87.8227 | 0.009266 | 0.97804 | 0.023451 | 23.9779 | -0.000547 | 0.15567 | 0.00719 | 0.49695 | 0.49695 | 0.012875
2019-03-21 12:02 | 1553184145.3988 | 1032 | 0.20811 | 0.001811 | 0.000161 | 99.3477 | 99.3864 | 117.3272 | 21.1255 | 0.003003 | 0.002794 | 136.9085 | 0.26143 | 8.402 | 0.003574 | 0.004607 | 0.004591 | 5.8e-05 | 0.000858 | 0.32117 | 1.3237 | 2.5479 | 99.53 | 20.92 | 37.224 | 93.9398 | 107.761 | 88.9042 | 0.009266 | 0.97512 | 0.023397 | 23.9937 | -0.003479 | 0.15552 | 0.006952 | 0.4789 | 0.4789 | 0.013347
2019-03-21 12:02 | 1553184146.3978 | 1033 | 0.20809 | 0.001787 | 0.000164 | 99.407 | 99.384 | 117.1852 | 21.1255 | 0.002997 | 0.002785 | 136.8814 | 0.26044 | 8.4022 | 0.003565 | 0.004827 | 0.004705 | 5.7e-05 | 0.000858 | 0.31885 | 1.3151 | 2.5479 | 99.53 | 20.92 | 37.226 | 93.7825 | 107.5301 | 88.6954 | 0.009267 | 0.97568 | 0.023389 | 23.9723 | 0.002072 | 0.15494 | 0.007087 | 0.48804 | 0.48804 | 0.013048
2019-03-21 12:02 | 1553184147.3975 | 1034 | 0.20815 | 0.00177 | 0.000168 | 99.376 | 99.4191 | 117.0488 | 21.1249 | 0.00301 | 0.002732 | 136.8525 | 0.26079 | 8.4019 | 0.003524 | 0.004645 | 0.004496 | 5.8e-05 | 0.000855 | 0.30829 | 1.2937 | 2.5479 | 99.53 | 20.92 | 37.227 | 94.1131 | 106.5335 | 87.4078 | 0.009267 | 0.97916 | 0.023417 | 23.9156 | -0.003873 | 0.15514 | 0.006763 | 0.46561 | 0.46561 | 0.013695
2019-03-21 12:02 | 1553184148.3979 | 1035 | 0.20817 | 0.001738 | 0.000172 | 99.3597 | 99.1217 | 117.3309 | 21.1255 | 0.003019 | 0.00281 | 136.8319 | 0.26124 | 8.4029 | 0.003492 | 0.005744 | 0.005724 | 6e-05 | 0.000855 | 0.32405 | 1.2902 | 2.5478 | 99.53 | 20.92 | 37.229 | 94.3156 | 105.7588 | 89.2993 | 0.009267 | 0.97406 | 0.023384 | 24.0072 | 0.021442 | 0.15541 | 0.006676 | 0.45861 | 0.45861 | 0.013928
2019-03-21 12:02 | 1553184149.398 | 1036 | 0.20816 | 0.001717 | 0.000175 | 99.3748 | 99.2328 | 117.6785 | 21.1257 | 0.003023 | 0.00275 | 136.8329 | 0.25999 | 8.4035 | 0.003471 | 0.005577 | 0.005479 | 5.8e-05 | 0.000855 | 0.30198 | 1.2843 | 2.5479 | 99.53 | 20.92 | 37.23 | 94.4277 | 105.243 | 87.8545 | 0.009268 | 0.97795 | 0.023466 | 23.9948 | 0.012797 | 0.15467 | 0.00675 | 0.46526 | 0.46526 | 0.013663
2019-03-21 12:02 | 1553184150.3987 | 1037 | 0.20817 | 0.001698 | 0.000178 | 99.3811 | 99.2328 | 117.9163 | 21.1245 | 0.003018 | 0.002756 | 136.8201 | 0.26098 | 8.4004 | 0.003611 | 0.004918 | 0.004756 | 5.9e-05 | 0.000858 | 0.29717 | 1.2997 | 2.548 | 99.53 | 20.92 | 37.231 | 94.2909 | 108.6432 | 87.9894 | 0.009268 | 0.97759 | 0.023485 | 24.0235 | 0.01336 | 0.15526 | 0.006722 | 0.46352 | 0.46352 | 0.013766
2019-03-21 12:02 | 1553184151.3978 | 1038 | 0.20824 | 0.001678 | 0.00018 | 99.3218 | 99.2144 | 118.0088 | 21.1256 | 0.003005 | 0.002781 | 136.7792 | 0.26059 | 8.4039 | 0.00364 | 0.004479 | 0.004426 | 5.6e-05 | 0.000853 | 0.31492 | 1.351 | 2.548 | 99.53 | 20.92 | 37.233 | 93.9956 | 109.3607 | 88.5863 | 0.009268 | 0.97598 | 0.023475 | 24.0527 | 0.00968 | 0.15502 | 0.006347 | 0.43676 | 0.43676 | 0.014588
2019-03-21 12:03 | 1553184152.3978 | 1039 | 0.20821 | 0.001657 | 0.000182 | 99.3748 | 99.1711 | 117.8297 | 21.1252 | 0.002943 | 0.002695 | 136.7706 | 0.26064 | 8.4016 | 0.003581 | 0.004802 | 0.00477 | 5.8e-05 | 0.000855 | 0.30381 | 1.3252 | 2.548 | 99.53 | 20.92 | 37.234 | 92.4814 | 107.9275 | 86.5197 | 0.009269 | 0.98158 | 0.023524 | 23.9657 | 0.018355 | 0.15505 | 0.00651 | 0.44909 | 0.44909 | 0.01419
2019-03-21 12:03 | 1553184153.3982 | 1040 | 0.20827 | 0.001638 | 0.000184 | 99.4344 | 99.2893 | 117.9318 | 21.1259 | 0.002966 | 0.002772 | 136.7466 | 0.26093 | 8.4018 | 0.003643 | 0.005263 | 0.005089 | 5.9e-05 | 0.000856 | 0.29407 | 1.3193 | 2.5479 | 99.53 | 20.92 | 37.234 | 93.0429 | 109.4321 | 88.3618 | 0.009269 | 0.97658 | 0.023474 | 24.0374 | 0.013064 | 0.15523 | 0.006212 | 0.427 | 0.427 | 0.014941
2019-03-21 12:03 | 1553184154.3969 | 1041 | 0.2083 | 0.001617 | 0.000185 | 99.3162 | 99.2028 | 117.7579 | 21.1261 | 0.002943 | 0.002712 | 136.7317 | 0.26083 | 8.4038 | 0.003637 | 0.004504 | 0.004503 | 6e-05 | 0.000859 | 0.31725 | 1.3017 | 2.5481 | 99.53 | 20.91 | 37.234 | 92.4915 | 109.2801 | 86.94 | 0.009263 | 0.98044 | 0.023503 | 23.9724 | 0.010214 | 0.15517 | 0.006028 | 0.41451 | 0.41451 | 0.015386
2019-03-21 12:03 | 1553184155.398 | 1042 | 0.20834 | 0.001598 | 0.000188 | 99.4088 | 99.2416 | 117.7804 | 21.1247 | 0.002949 | 0.002662 | 136.7232 | 0.26124 | 8.4011 | 0.00363 | 0.00497 | 0.004855 | 5.7e-05 | 0.000856 | 0.29302 | 1.2927 | 2.5479 | 99.53 | 20.91 | 37.234 | 92.6288 | 109.1147 | 85.7277 | 0.009263 | 0.98375 | 0.023545 | 23.9343 | 0.015055 | 0.15541 | 0.005842 | 0.40194 | 0.40194 | 0.015891
2019-03-21 12:03 | 1553184156.3973 | 1043 | 0.20828 | 0.001574 | 0.00019 | 99.36 | 99.2937 | 117.6966 | 21.1265 | 0.002964 | 0.002736 | 136.7193 | 0.26136 | 8.4041 | 0.003559 | 0.004822 | 0.004668 | 5.8e-05 | 0.000861 | 0.33682 | 1.3011 | 2.5479 | 99.53 | 20.91 | 37.235 | 92.9964 | 107.3981 | 87.5025 | 0.009263 | 0.97891 | 0.023479 | 23.9849 | 0.005962 | 0.15548 | 0.006189 | 0.42511 | 0.42511 | 0.015032
2019-03-21 12:03 | 1553184157.398 | 1044 | 0.20833 | 0.001558 | 0.000192 | 99.2881 | 99.182 | 117.8684 | 21.1262 | 0.00294 | 0.002721 | 136.7064 | 0.26118 | 8.4046 | 0.00356 | 0.005219 | 0.005 | 5.7e-05 | 0.000858 | 0.30342 | 1.2668 | 2.548 | 99.53 | 20.91 | 37.236 | 92.4173 | 107.4198 | 87.146 | 0.009263 | 0.97988 | 0.023508 | 23.9905 | 0.009565 | 0.15538 | 0.005938 | 0.40787 | 0.40787 | 0.015657
2019-03-21 12:03 | 1553184158.3971 | 1045 | 0.20838 | 0.00154 | 0.000194 | 99.2712 | 98.9981 | 117.8208 | 21.1267 | 0.002971 | 0.002674 | 136.73 | 0.26092 | 8.4027 | 0.003614 | 0.004743 | 0.004804 | 6e-05 | 0.000858 | 0.28799 | 1.2819 | 2.5479 | 99.53 | 20.91 | 37.235 | 93.1719 | 108.7185 | 86.0074 | 0.009263 | 0.98298 | 0.02354 | 23.9477 | 0.024636 | 0.15522 | 0.00561 | 0.38514 | 0.38514 | 0.016564
2019-03-21 12:03 | 1553184159.3978 | 1046 | 0.20838 | 0.001521 | 0.000198 | 99.3534 | 99.1711 | 117.8647 | 21.1256 | 0.002981 | 0.002751 | 136.7124 | 0.26154 | 8.402 | 0.003608 | 0.004828 | 0.004791 | 5.6e-05 | 0.000859 | 0.28313 | 1.2945 | 2.5479 | 99.53 | 20.91 | 37.235 | 93.411 | 108.5765 | 87.8607 | 0.009263 | 0.97794 | 0.023484 | 24.0139 | 0.016424 | 0.15559 | 0.005631 | 0.38541 | 0.38541 | 0.016592
2019-03-21 12:03 | 1553184160.3982 | 1047 | 0.20836 | 0.001502 | 0.000201 | 99.3036 | 99.0049 | 117.8022 | 21.1263 | 0.002939 | 0.0027 | 136.6975 | 0.26111 | 8.4016 | 0.003673 | 0.005317 | 0.005209 | 5.9e-05 | 0.000858 | 0.29839 | 1.264 | 2.5479 | 99.53 | 20.91 | 37.235 | 92.3892 | 110.1635 | 86.6434 | 0.009263 | 0.98125 | 0.023518 | 23.9671 | 0.026947 | 0.15534 | 0.00579 | 0.39697 | 0.39697 | 0.016083
2019-03-21 12:03 | 1553184161.3982 | 1048 | 0.20841 | 0.001489 | 0.000206 | 99.3641 | 99.2045 | 117.8715 | 21.1261 | 0.002937 | 0.002685 | 136.7068 | 0.26089 | 8.3998 | 0.003662 | 0.005235 | 0.005131 | 5.8e-05 | 0.000858 | 0.29607 | 1.3023 | 2.5481 | 99.53 | 20.91 | 37.234 | 92.3454 | 109.9037 | 86.2802 | 0.009263 | 0.98224 | 0.023536 | 23.962 | 0.01438 | 0.15521 | 0.005493 | 0.37606 | 0.37606 | 0.016963
2019-03-21 12:03 | 1553184162.3973 | 1049 | 0.20848 | 0.001471 | 0.000211 | 99.2831 | 99.2001 | 118.0358 | 21.1265 | 0.002946 | 0.002754 | 136.6957 | 0.26052 | 8.4031 | 0.003662 | 0.005337 | 0.005282 | 6e-05 | 0.000861 | 0.30829 | 1.287 | 2.5479 | 99.53 | 20.91 | 37.234 | 92.5608 | 109.8837 | 87.9427 | 0.009263 | 0.97771 | 0.023498 | 24.0341 | 0.007481 | 0.15499 | 0.005107 | 0.34793 | 0.34793 | 0.018308
2019-03-21 12:03 | 1553184163.3974 | 1050 | 0.20849 | 0.001452 | 0.000215 | 99.2995 | 99.1882 | 118.2804 | 21.127 | 0.002976 | 0.002714 | 136.6774 | 0.26132 | 8.4033 | 0.003685 | 0.005613 | 0.005497 | 6.1e-05 | 0.000859 | 0.30254 | 1.2889 | 2.5478 | 99.53 | 20.91 | 37.234 | 93.2959 | 110.4606 | 86.9712 | 0.009263 | 0.98035 | 0.023554 | 24.0266 | 0.010032 | 0.15546 | 0.005083 | 0.3468 | 0.3468 | 0.018424
2019-03-21 12:03 | 1553184164.3974 | 1051 | 0.20857 | 0.001429 | 0.000219 | 99.329 | 99.3125 | 118.7282 | 21.1267 | 0.002916 | 0.002679 | 136.6662 | 0.26103 | 8.4013 | 0.003781 | 0.005298 | 0.005211 | 5.8e-05 | 0.000858 | 0.31653 | 1.2972 | 2.5479 | 99.53 | 20.91 | 37.236 | 91.8352 | 112.7755 | 86.1403 | 0.009263 | 0.98262 | 0.023626 | 24.0442 | 0.001485 | 0.15529 | 0.004616 | 0.31449 | 0.31449 | 0.020294
2019-03-21 12:03 | 1553184165.3988 | 1052 | 0.20861 | 0.00141 | 0.000223 | 99.2676 | 98.9831 | 118.9632 | 21.1272 | 0.002927 | 0.002681 | 136.6537 | 0.2618 | 8.4039 | 0.003737 | 0.004853 | 0.004666 | 5.9e-05 | 0.000858 | 0.30038 | 1.3237 | 2.548 | 99.53 | 20.91 | 37.237 | 92.1047 | 111.7236 | 86.1929 | 0.009264 | 0.98247 | 0.023648 | 24.0698 | 0.025668 | 0.15574 | 0.004369 | 0.29705 | 0.29705 | 0.021549
2019-03-21 12:03 | 1553184166.3974 | 1053 | 0.20863 | 0.001389 | 0.000226 | 99.3051 | 99.1756 | 118.986 | 21.1268 | 0.002934 | 0.00272 | 136.6301 | 0.26204 | 8.4002 | 0.003724 | 0.004958 | 0.004883 | 6.3e-05 | 0.000853 | 0.29479 | 1.3075 | 2.5479 | 99.53 | 20.91 | 37.237 | 92.2841 | 111.4011 | 87.1303 | 0.009264 | 0.97992 | 0.023619 | 24.1035 | 0.011678 | 0.15589 | 0.004255 | 0.28841 | 0.28841 | 0.022215
2019-03-21 12:03 | 1553184167.3977 | 1054 | 0.20866 | 0.001367 | 0.000229 | 99.2688 | 99.1006 | 118.6566 | 21.1266 | 0.002903 | 0.002721 | 136.627 | 0.26131 | 8.4015 | 0.003697 | 0.005867 | 0.005772 | 5.7e-05 | 0.000855 | 0.29302 | 1.2897 | 2.5478 | 99.53 | 20.91 | 37.238 | 91.5177 | 110.7349 | 87.1402 | 0.009264 | 0.97989 | 0.023586 | 24.0704 | 0.015164 | 0.15545 | 0.004112 | 0.27771 | 0.27771 | 0.023006
2019-03-21 12:03 | 1553184168.3977 | 1055 | 0.20859 | 0.00134 | 0.000231 | 99.2759 | 99.1228 | 118.3776 | 21.1268 | 0.002892 | 0.002685 | 136.5949 | 0.26098 | 8.4037 | 0.003698 | 0.004869 | 0.004795 | 6.1e-05 | 0.000857 | 0.31349 | 1.316 | 2.5478 | 99.53 | 20.9 | 37.238 | 91.263 | 110.7587 | 86.2854 | 0.009258 | 0.98222 | 0.023587 | 24.0136 | 0.01381 | 0.15526 | 0.004575 | 0.31038 | 0.31038 | 0.020559
2019-03-21 12:03 | 1553184169.3977 | 1056 | 0.20864 | 0.001322 | 0.000233 | 99.292 | 99.0764 | 118.1602 | 21.1269 | 0.00301 | 0.002648 | 136.589 | 0.26147 | 8.4013 | 0.003807 | 0.005122 | 0.004955 | 6.1e-05 | 0.000856 | 0.27831 | 1.3755 | 2.548 | 99.53 | 20.9 | 37.239 | 94.114 | 113.4212 | 85.398 | 0.009259 | 0.98465 | 0.023594 | 23.9619 | 0.019441 | 0.15555 | 0.004316 | 0.29202 | 0.29202 | 0.021892
2019-03-21 12:03 | 1553184170.3977 | 1057 | 0.20865 | 0.001298 | 0.000235 | 99.2789 | 99.1912 | 117.9325 | 21.1268 | 0.002857 | 0.002621 | 136.5956 | 0.26085 | 8.4037 | 0.003797 | 0.005364 | 0.005232 | 6.1e-05 | 0.000853 | 0.29031 | 1.307 | 2.548 | 99.53 | 20.9 | 37.241 | 90.4157 | 113.1857 | 84.7317 | 0.009259 | 0.98649 | 0.023593 | 23.9165 | 0.007903 | 0.15518 | 0.004277 | 0.28916 | 0.28916 | 0.022057
2019-03-21 12:03 | 1553184171.3982 | 1058 | 0.20865 | 0.001279 | 0.000235 | 99.2652 | 99.2205 | 117.7321 | 21.1269 | 0.002871 | 0.002696 | 136.5652 | 0.26103 | 8.402 | 0.003765 | 0.005355 | 0.005336 | 5.7e-05 | 0.000862 | 0.32444 | 1.2905 | 2.5479 | 99.53 | 20.9 | 37.243 | 90.7682 | 112.3952 | 86.5512 | 0.00926 | 0.9815 | 0.023514 | 23.9568 | 0.004028 | 0.15529 | 0.00427 | 0.28766 | 0.28766 | 0.022187
2019-03-21 12:03 | 1553184172.3978 | 1059 | 0.20869 | 0.001259 | 0.000236 | 99.3191 | 99.1224 | 117.4774 | 21.1268 | 0.002893 | 0.002692 | 136.5401 | 0.26035 | 8.4018 | 0.003718 | 0.005485 | 0.005368 | 5.6e-05 | 0.000858 | 0.27489 | 1.3017 | 2.5479 | 99.53 | 20.9 | 37.244 | 91.2818 | 111.2636 | 86.4541 | 0.00926 | 0.98176 | 0.023491 | 23.9277 | 0.017733 | 0.15488 | 0.004102 | 0.27551 | 0.27551 | 0.023105
2019-03-21 12:03 | 1553184173.3969 | 1060 | 0.20867 | 0.001237 | 0.000236 | 99.2211 | 99.072 | 117.4246 | 21.1264 | 0.002875 | 0.002685 | 136.5441 | 0.26118 | 8.4048 | 0.003656 | 0.004505 | 0.004442 | 5.8e-05 | 0.000856 | 0.28407 | 1.3014 | 2.5481 | 99.53 | 20.9 | 37.245 | 90.843 | 109.748 | 86.2861 | 0.00926 | 0.98222 | 0.023491 | 23.9167 | 0.01345 | 0.15538 | 0.004196 | 0.28208 | 0.28208 | 0.022638
2019-03-21 12:03 | 1553184174.3973 | 1061 | 0.20873 | 0.001221 | 0.000238 | 99.3159 | 99.3459 | 117.2624 | 21.1264 | 0.002867 | 0.002683 | 136.5151 | 0.26166 | 8.4028 | 0.003761 | 0.005775 | 0.005674 | 5.5e-05 | 0.000862 | 0.32421 | 1.2961 | 2.5479 | 99.53 | 20.9 | 37.245 | 90.6725 | 112.3015 | 86.2383 | 0.00926 | 0.98235 | 0.023477 | 23.8986 | -0.002701 | 0.15566 | 0.003888 | 0.26007 | 0.26007 | 0.0246
2019-03-21 12:03 | 1553184175.3984 | 1062 | 0.20877 | 0.001201 | 0.000239 | 99.3653 | 99.2403 | 117.2037 | 21.1273 | 0.002878 | 0.002586 | 136.5059 | 0.26075 | 8.4007 | 0.003636 | 0.005301 | 0.005243 | 5.7e-05 | 0.000859 | 0.32444 | 1.2928 | 2.548 | 99.53 | 20.9 | 37.245 | 90.9267 | 109.2528 | 83.9035 | 0.00926 | 0.98877 | 0.023548 | 23.8149 | 0.011263 | 0.15512 | 0.00367 | 0.24541 | 0.24541 | 0.025978
2019-03-21 12:03 | 1553184176.397 | 1063 | 0.20876 | 0.001184 | 0.000241 | 99.2503 | 99.1252 | 117.3232 | 21.126 | 0.002878 | 0.00269 | 136.5043 | 0.26161 | 8.4017 | 0.003656 | 0.005363 | 0.005222 | 5.7e-05 | 0.000861 | 0.33301 | 1.2587 | 2.5479 | 99.53 | 20.9 | 37.246 | 90.9285 | 109.7356 | 86.4098 | 0.00926 | 0.98188 | 0.023477 | 23.9105 | 0.011286 | 0.15563 | 0.00377 | 0.25161 | 0.25161 | 0.025422
2019-03-21 12:03 | 1553184177.3975 | 1064 | 0.20873 | 0.00117 | 0.000241 | 99.3763 | 99.01 | 117.395 | 21.126 | 0.002845 | 0.002628 | 136.4898 | 0.26063 | 8.4024 | 0.003606 | 0.005184 | 0.005056 | 5.4e-05 | 0.000856 | 0.31526 | 1.3001 | 2.5479 | 99.53 | 20.9 | 37.248 | 90.1312 | 108.5335 | 84.9118 | 0.009261 | 0.98599 | 0.023534 | 23.8679 | 0.033031 | 0.15505 | 0.003944 | 0.26448 | 0.26448 | 0.024094
2019-03-21 12:03 | 1553184178.3975 | 1065 | 0.20873 | 0.001153 | 0.000241 | 99.3707 | 99.179 | 117.7459 | 21.1259 | 0.002916 | 0.002626 | 136.4863 | 0.26052 | 8.4011 | 0.003629 | 0.005694 | 0.005675 | 5.6e-05 | 0.00086 | 0.34164 | 1.3379 | 2.5479 | 99.53 | 20.9 | 37.249 | 91.8306 | 109.0843 | 84.8682 | 0.009261 | 0.98611 | 0.02357 | 23.9021 | 0.017273 | 0.15499 | 0.003948 | 0.2652 | 0.2652 | 0.024019
2019-03-21 12:03 | 1553184179.3982 | 1066 | 0.2088 | 0.001136 | 0.000241 | 99.2905 | 99.2127 | 118.1147 | 21.1267 | 0.00286 | 0.002674 | 136.499 | 0.2618 | 8.4012 | 0.003668 | 0.00525 | 0.005126 | 5.5e-05 | 0.000856 | 0.32676 | 1.3501 | 2.548 | 99.53 | 20.89 | 37.25 | 90.4891 | 110.029 | 86.0116 | 0.009256 | 0.98297 | 0.023569 | 23.9777 | 0.007015 | 0.15574 | 0.00357 | 0.23841 | 0.23841 | 0.026849
2019-03-21 12:03 | 1553184180.3976 | 1067 | 0.20883 | 0.001122 | 0.000242 | 99.3531 | 98.9515 | 118.2264 | 21.1268 | 0.002859 | 0.002649 | 136.4689 | 0.26231 | 8.402 | 0.003601 | 0.005225 | 0.005191 | 5.8e-05 | 0.00086 | 0.30757 | 1.314 | 2.5478 | 99.53 | 20.89 | 37.251 | 90.4784 | 108.4216 | 85.4262 | 0.009256 | 0.98458 | 0.0236 | 23.9695 | 0.036233 | 0.15605 | 0.003423 | 0.22828 | 0.22828 | 0.028095
2019-03-21 12:03 | 1553184181.3983 | 1068 | 0.20878 | 0.001108 | 0.000243 | 99.3561 | 99.2031 | 118.6172 | 21.1262 | 0.002873 | 0.002616 | 136.476 | 0.26217 | 8.4025 | 0.003673 | 0.005641 | 0.005537 | 5.6e-05 | 0.000857 | 0.32045 | 1.2793 | 2.5479 | 99.53 | 20.89 | 37.253 | 90.8068 | 110.1571 | 84.6102 | 0.009256 | 0.98682 | 0.023666 | 23.9818 | 0.013778 | 0.15597 | 0.003698 | 0.24845 | 0.24845 | 0.025801
2019-03-21 12:03 | 1553184182.3984 | 1069 | 0.20881 | 0.001102 | 0.000245 | 99.2995 | 99.0775 | 118.8651 | 21.1264 | 0.002869 | 0.002618 | 136.4642 | 0.26077 | 8.403 | 0.003558 | 0.004582 | 0.004512 | 5.4e-05 | 0.000859 | 0.29214 | 1.309 | 2.548 | 99.53 | 20.89 | 37.253 | 90.721 | 107.3616 | 84.6736 | 0.009256 | 0.98665 | 0.023688 | 24.0089 | 0.020019 | 0.15513 | 0.003521 | 0.23597 | 0.23597 | 0.02702
2019-03-21 12:03 | 1553184183.3968 | 1070 | 0.2088 | 0.001092 | 0.000245 | 99.3787 | 99.3217 | 118.9867 | 21.1266 | 0.002862 | 0.002632 | 136.4354 | 0.26152 | 8.4042 | 0.003629 | 0.00494 | 0.004861 | 5.7e-05 | 0.000854 | 0.32405 | 1.325 | 2.5479 | 99.53 | 20.89 | 37.254 | 90.5359 | 109.1015 | 84.9968 | 0.009257 | 0.98576 | 0.02369 | 24.0321 | 0.005136 | 0.15558 | 0.003618 | 0.24286 | 0.24286 | 0.026329
2019-03-21 12:03 | 1553184184.3965 | 1071 | 0.20879 | 0.001078 | 0.000245 | 99.4139 | 99.2144 | 119.1151 | 21.1268 | 0.002805 | 0.002629 | 136.4455 | 0.26069 | 8.4009 | 0.003554 | 0.004506 | 0.00442 | 5.5e-05 | 0.000859 | 0.31332 | 1.3187 | 2.5479 | 99.52 | 20.89 | 37.256 | 89.1694 | 107.2715 | 84.9425 | 0.009258 | 0.98591 | 0.023704 | 24.0432 | 0.017967 | 0.15509 | 0.003678 | 0.24727 | 0.24727 | 0.025777
2019-03-21 12:03 | 1553184185.3978 | 1072 | 0.20879 | 0.001067 | 0.000243 | 99.3808 | 99.211 | 119.1662 | 21.1272 | 0.002845 | 0.002626 | 136.4397 | 0.26152 | 8.4031 | 0.003538 | 0.005353 | 0.005195 | 5.5e-05 | 0.00086 | 0.32101 | 1.308 | 2.548 | 99.52 | 20.89 | 37.257 | 90.1254 | 106.8757 | 84.8733 | 0.009258 | 0.9861 | 0.023712 | 24.046 | 0.015298 | 0.15558 | 0.003708 | 0.24962 | 0.24962 | 0.025616
2019-03-21 12:03 | 1553184186.3976 | 1073 | 0.2088 | 0.001056 | 0.000241 | 99.4026 | 99.3922 | 119.1484 | 21.127 | 0.002859 | 0.002613 | 136.4216 | 0.2611 | 8.4007 | 0.003578 | 0.004853 | 0.004797 | 5.4e-05 | 0.000859 | 0.31974 | 1.2877 | 2.5479 | 99.52 | 20.89 | 37.259 | 90.4667 | 107.8492 | 84.546 | 0.009259 | 0.987 | 0.023721 | 24.0332 | 0.000934 | 0.15532 | 0.003683 | 0.24806 | 0.24806 | 0.025735
2019-03-21 12:03 | 1553184187.3966 | 1074 | 0.20881 | 0.001048 | 0.000237 | 99.4371 | 99.1878 | 119.2467 | 21.1263 | 0.002839 | 0.002645 | 136.4183 | 0.26086 | 8.402 | 0.003616 | 0.005486 | 0.005332 | 5.1e-05 | 0.000859 | 0.31741 | 1.3004 | 2.5481 | 99.52 | 20.89 | 37.259 | 89.9831 | 108.782 | 85.3141 | 0.009259 | 0.98488 | 0.023705 | 24.069 | 0.022453 | 0.15519 | 0.003603 | 0.24243 | 0.24243 | 0.02631
2019-03-21 12:03 | 1553184188.3977 | 1075 | 0.2088 | 0.001041 | 0.000233 | 99.4192 | 99.3104 | 119.2836 | 21.1257 | 0.002756 | 0.002613 | 136.4103 | 0.26139 | 8.401 | 0.003614 | 0.005008 | 0.004982 | 5.5e-05 | 0.000856 | 0.33964 | 1.3336 | 2.548 | 99.52 | 20.89 | 37.26 | 87.9763 | 108.7294 | 84.5401 | 0.009259 | 0.98701 | 0.023734 | 24.0467 | 0.009795 | 0.1555 | 0.003711 | 0.25066 | 0.25066 | 0.025497
2019-03-21 12:03 | 1553184189.3967 | 1076 | 0.20882 | 0.001035 | 0.000229 | 99.4148 | 99.3424 | 119.1003 | 21.1268 | 0.002836 | 0.002628 | 136.3784 | 0.26122 | 8.4023 | 0.003679 | 0.004916 | 0.004863 | 5.3e-05 | 0.000862 | 0.29679 | 1.3215 | 2.5479 | 99.53 | 20.89 | 37.261 | 89.925 | 110.3003 | 84.9142 | 0.009258 | 0.98598 | 0.023704 | 24.0408 | 0.00651 | 0.1554 | 0.003609 | 0.24329 | 0.24329 | 0.026252
2019-03-21 12:03 | 1553184190.3978 | 1077 | 0.20887 | 0.001024 | 0.000226 | 99.4055 | 99.229 | 119.2127 | 21.1261 | 0.002829 | 0.002575 | 136.394 | 0.26185 | 8.4054 | 0.003603 | 0.004897 | 0.00475 | 5.4e-05 | 0.000854 | 0.31078 | 1.3645 | 2.5481 | 99.53 | 20.89 | 37.262 | 89.7562 | 108.4563 | 83.6347 | 0.009259 | 0.98952 | 0.023757 | 24.0091 | 0.015896 | 0.15577 | 0.003281 | 0.22067 | 0.22067 | 0.029014
2019-03-21 12:03 | 1553184191.3978 | 1078 | 0.20884 | 0.001017 | 0.000221 | 99.4249 | 99.3496 | 118.9275 | 21.1269 | 0.002798 | 0.002633 | 136.3661 | 0.26005 | 8.4028 | 0.003563 | 0.004986 | 0.004883 | 5.2e-05 | 0.000856 | 0.32587 | 1.2989 | 2.548 | 99.53 | 20.89 | 37.263 | 89.0007 | 107.4929 | 85.0228 | 0.009259 | 0.98568 | 0.023683 | 24.027 | 0.006777 | 0.15471 | 0.003524 | 0.23756 | 0.23756 | 0.026766
2019-03-21 12:03 | 1553184192.3969 | 1079 | 0.20884 | 0.00101 | 0.000217 | 99.4472 | 99.3224 | 118.8217 | 21.1263 | 0.002804 | 0.002599 | 136.3479 | 0.26135 | 8.4009 | 0.003569 | 0.005401 | 0.005272 | 5.5e-05 | 0.000861 | 0.32012 | 1.3513 | 2.548 | 99.53 | 20.89 | 37.265 | 89.1493 | 107.6372 | 84.2161 | 0.009259 | 0.98791 | 0.023699 | 23.9892 | 0.011241 | 0.15548 | 0.003503 | 0.2364 | 0.2364 | 0.027031
2019-03-21 12:03 | 1553184193.3983 | 1080 | 0.20884 | 0.001004 | 0.000213 | 99.3892 | 99.3472 | 118.7503 | 21.1271 | 0.002822 | 0.002635 | 136.3571 | 0.26122 | 8.4045 | 0.003527 | 0.004916 | 0.004865 | 5.6e-05 | 0.000859 | 0.30342 | 1.3162 | 2.5478 | 99.53 | 20.89 | 37.267 | 89.5706 | 106.615 | 85.0752 | 0.00926 | 0.98554 | 0.023664 | 24.0108 | 0.003775 | 0.1554 | 0.003535 | 0.2386 | 0.2386 | 0.026768
2019-03-21 12:03 | 1553184194.3971 | 1081 | 0.20887 | 0.000999 | 0.000211 | 99.4252 | 99.57 | 118.6869 | 21.1263 | 0.002795 | 0.00258 | 136.3682 | 0.26137 | 8.405 | 0.003502 | 0.005185 | 0.005106 | 5.7e-05 | 0.000861 | 0.30663 | 1.2879 | 2.5479 | 99.53 | 20.89 | 37.269 | 88.9165 | 106.0153 | 83.759 | 0.00926 | 0.98918 | 0.023701 | 23.9603 | -0.013013 | 0.15549 | 0.003361 | 0.2267 | 0.2267 | 0.02819
2019-03-21 12:03 | 1553184195.3982 | 1082 | 0.20885 | 0.000992 | 0.000208 | 99.4815 | 99.3915 | 118.7648 | 21.1269 | 0.002789 | 0.002621 | 136.3412 | 0.2611 | 8.4034 | 0.003451 | 0.005613 | 0.005531 | 5.6e-05 | 0.000865 | 0.30596 | 1.2884 | 2.5479 | 99.53 | 20.89 | 37.27 | 88.7743 | 104.7627 | 84.7332 | 0.009261 | 0.98648 | 0.023676 | 24.0008 | 0.008097 | 0.15532 | 0.003493 | 0.23606 | 0.23606 | 0.027043
2019-03-21 12:03 | 1553184196.3981 | 1083 | 0.20888 | 0.000987 | 0.000204 | 99.5393 | 99.2941 | 118.9813 | 21.1272 | 0.002757 | 0.002554 | 136.3361 | 0.26023 | 8.4024 | 0.003422 | 0.005229 | 0.005137 | 5.4e-05 | 0.000864 | 0.30989 | 1.3005 | 2.5479 | 99.53 | 20.89 | 37.276 | 88.0247 | 104.0631 | 83.1203 | 0.009262 | 0.99095 | 0.023752 | 23.9685 | 0.022063 | 0.15481 | 0.003348 | 0.22668 | 0.22668 | 0.028069
2019-03-21 12:03 | 1553184197.3984 | 1084 | 0.20879 | 0.000982 | 0.000202 | 99.4458 | 99.4726 | 119.1733 | 21.1266 | 0.002759 | 0.002573 | 136.3434 | 0.26199 | 8.4047 | 0.003535 | 0.004871 | 0.004861 | 5.4e-05 | 0.000866 | 0.31133 | 1.2742 | 2.548 | 99.53 | 20.89 | 37.278 | 88.0607 | 106.8084 | 83.5868 | 0.009263 | 0.98965 | 0.023755 | 24.0035 | -0.002409 | 0.15586 | 0.003856 | 0.26311 | 0.26311 | 0.024346
2019-03-21 12:03 | 1553184198.3986 | 1085 | 0.20886 | 0.000973 | 0.000199 | 99.5 | 99.4395 | 119.1437 | 21.128 | 0.002811 | 0.002643 | 136.3489 | 0.26082 | 8.4033 | 0.003505 | 0.005079 | 0.004978 | 5.7e-05 | 0.000868 | 0.30093 | 1.2662 | 2.5481 | 99.53 | 20.89 | 37.283 | 89.3059 | 106.0902 | 85.2767 | 0.009264 | 0.98499 | 0.023696 | 24.0573 | 0.005438 | 0.15516 | 0.003448 | 0.23357 | 0.23357 | 0.027303
2019-03-21 12:03 | 1553184199.3973 | 1086 | 0.20885 | 0.000969 | 0.000196 | 99.5337 | 99.308 | 119.1242 | 21.1275 | 0.002714 | 0.002554 | 136.3542 | 0.2604 | 8.4037 | 0.003504 | 0.004702 | 0.004652 | 5.7e-05 | 0.000866 | 0.29878 | 1.2679 | 2.548 | 99.53 | 20.88 | 37.286 | 86.9733 | 106.0645 | 83.1391 | 0.009259 | 0.9909 | 0.023765 | 23.9835 | 0.020297 | 0.15491 | 0.003525 | 0.23996 | 0.23996 | 0.026533
2019-03-21 12:03 | 1553184200.3963 | 1087 | 0.20882 | 0.000965 | 0.000192 | 99.5536 | 99.3401 | 118.9686 | 21.1273 | 0.002778 | 0.002604 | 136.3308 | 0.26141 | 8.402 | 0.003496 | 0.004142 | 0.004037 | 5.8e-05 | 0.000867 | 0.34219 | 1.2764 | 2.5479 | 99.53 | 20.88 | 37.289 | 88.5278 | 105.8669 | 84.3311 | 0.00926 | 0.98759 | 0.02371 | 24.0079 | 0.019207 | 0.15551 | 0.003737 | 0.25467 | 0.25467 | 0.025097
2019-03-21 12:03 | 1553184201.3974 | 1088 | 0.20886 | 0.000967 | 0.000189 | 99.5319 | 99.4378 | 119.1043 | 21.1274 | 0.002783 | 0.002571 | 136.3333 | 0.26188 | 8.4016 | 0.003464 | 0.00503 | 0.004913 | 6e-05 | 0.000869 | 0.31205 | 1.2905 | 2.5479 | 99.53 | 20.88 | 37.293 | 88.6427 | 105.0779 | 83.5299 | 0.009261 | 0.98981 | 0.02375 | 23.9946 | 0.008456 | 0.15579 | 0.003469 | 0.23617 | 0.23617 | 0.027111
2019-03-21 12:03 | 1553184202.3965 | 1089 | 0.20886 | 0.000963 | 0.000185 | 99.4726 | 99.6803 | 119.4258 | 21.1269 | 0.002783 | 0.002616 | 136.3413 | 0.26141 | 8.4028 | 0.00347 | 0.00472 | 0.004632 | 5.7e-05 | 0.000863 | 0.29999 | 1.2854 | 2.5479 | 99.53 | 20.88 | 37.297 | 88.6408 | 105.2418 | 84.6296 | 0.009262 | 0.98677 | 0.023746 | 24.064 | -0.01866 | 0.15551 | 0.003475 | 0.23677 | 0.23677 | 0.026995
2019-03-21 12:03 | 1553184203.3964 | 1090 | 0.20884 | 0.000962 | 0.000183 | 99.6159 | 99.508 | 119.803 | 21.1274 | 0.002752 | 0.00257 | 136.3397 | 0.26051 | 8.4015 | 0.003539 | 0.004712 | 0.004623 | 6e-05 | 0.00086 | 0.32045 | 1.3023 | 2.5479 | 99.53 | 20.89 | 37.3 | 87.8812 | 106.9113 | 83.5049 | 0.009268 | 0.98988 | 0.023821 | 24.0641 | 0.009695 | 0.15498 | 0.003614 | 0.24763 | 0.24763 | 0.025722
2019-03-21 12:03 | 1553184204.3975 | 1091 | 0.20885 | 0.000953 | 0.00018 | 99.4925 | 99.5369 | 119.9016 | 21.127 | 0.002809 | 0.002543 | 136.2956 | 0.26211 | 8.4044 | 0.003498 | 0.004645 | 0.004559 | 6.2e-05 | 0.000862 | 0.31205 | 1.3277 | 2.548 | 99.52 | 20.89 | 37.306 | 89.2697 | 105.9139 | 82.864 | 0.00927 | 0.99166 | 0.023852 | 24.0523 | -0.00399 | 0.15593 | 0.003562 | 0.24436 | 0.24436 | 0.026226
2019-03-21 12:03 | 1553184205.3976 | 1092 | 0.20888 | 0.000947 | 0.000178 | 99.6201 | 99.5352 | 120.1718 | 21.1275 | 0.002719 | 0.00255 | 136.3011 | 0.26093 | 8.4015 | 0.003551 | 0.004986 | 0.004875 | 6e-05 | 0.00086 | 0.30127 | 1.3419 | 2.5479 | 99.52 | 20.89 | 37.308 | 87.1084 | 107.1903 | 83.0401 | 0.009271 | 0.99117 | 0.023873 | 24.0854 | 0.00762 | 0.15523 | 0.003437 | 0.23578 | 0.23578 | 0.027059
2019-03-21 12:03 | 1553184206.397 | 1093 | 0.20889 | 0.000941 | 0.000174 | 99.5536 | 99.5393 | 120.1699 | 21.1283 | 0.002711 | 0.002558 | 136.2839 | 0.26152 | 8.3994 | 0.003434 | 0.004819 | 0.004684 | 6e-05 | 0.000864 | 0.31564 | 1.3222 | 2.5479 | 99.52 | 20.89 | 37.313 | 86.9139 | 104.363 | 83.2194 | 0.009272 | 0.99067 | 0.023867 | 24.0912 | 0.001286 | 0.15558 | 0.003332 | 0.22845 | 0.22845 | 0.027989
2019-03-21 12:03 | 1553184207.3977 | 1094 | 0.20889 | 0.000945 | 0.00017 | 99.6028 | 99.4627 | 120.0335 | 21.1279 | 0.002786 | 0.002611 | 136.2585 | 0.26089 | 8.402 | 0.003532 | 0.004496 | 0.004434 | 5.9e-05 | 0.000863 | 0.31293 | 1.337 | 2.5481 | 99.52 | 20.89 | 37.315 | 88.701 | 106.7284 | 84.5118 | 0.009273 | 0.98709 | 0.02381 | 24.1212 | 0.01259 | 0.15521 | 0.003347 | 0.22918 | 0.22918 | 0.027834
2019-03-21 12:03 | 1553184208.3977 | 1095 | 0.20887 | 0.000939 | 0.000167 | 99.6528 | 99.4014 | 120.1453 | 21.1278 | 0.002775 | 0.002509 | 136.2304 | 0.26063 | 8.4043 | 0.003584 | 0.005361 | 0.005272 | 6e-05 | 0.000867 | 0.31133 | 1.292 | 2.5479 | 99.52 | 20.89 | 37.316 | 88.444 | 107.9971 | 82.0506 | 0.009273 | 0.99393 | 0.023903 | 24.0492 | 0.022599 | 0.15505 | 0.003498 | 0.24108 | 0.24108 | 0.026433
2019-03-21 12:03 | 1553184209.3981 | 1096 | 0.20892 | 0.000942 | 0.000165 | 99.6329 | 99.4824 | 120.1213 | 21.1282 | 0.002757 | 0.002576 | 136.197 | 0.26176 | 8.4026 | 0.003562 | 0.004911 | 0.00482 | 5.9e-05 | 0.000867 | 0.3131 | 1.3325 | 2.5479 | 99.52 | 20.89 | 37.317 | 88.0188 | 107.4728 | 83.6688 | 0.009273 | 0.98942 | 0.023847 | 24.1016 | 0.013516 | 0.15572 | 0.003217 | 0.22056 | 0.22056 | 0.029017
2019-03-21 12:03 | 1553184210.3969 | 1097 | 0.20895 | 0.000933 | 0.000162 | 99.6558 | 99.4981 | 120.3607 | 21.1273 | 0.002795 | 0.002559 | 136.1964 | 0.26083 | 8.4037 | 0.003591 | 0.004581 | 0.004463 | 6e-05 | 0.000867 | 0.30359 | 1.3223 | 2.5477 | 99.52 | 20.89 | 37.319 | 88.9276 | 108.1777 | 83.2514 | 0.009274 | 0.99058 | 0.023884 | 24.1115 | 0.014167 | 0.15517 | 0.003018 | 0.20681 | 0.20681 | 0.030836
2019-03-21 12:04 | 1553184211.3979 | 1098 | 0.20889 | 0.000927 | 0.00016 | 99.5995 | 99.6527 | 120.3882 | 21.1269 | 0.002728 | 0.002567 | 136.1987 | 0.26209 | 8.4011 | 0.003617 | 0.005083 | 0.004904 | 5.9e-05 | 0.000869 | 0.31365 | 1.3428 | 2.5479 | 99.52 | 20.89 | 37.32 | 87.3178 | 108.7896 | 83.4484 | 0.009274 | 0.99004 | 0.023881 | 24.1209 | -0.004779 | 0.15592 | 0.003387 | 0.23337 | 0.23337 | 0.027459
2019-03-21 12:04 | 1553184212.3974 | 1099 | 0.20896 | 0.000924 | 0.000158 | 99.6603 | 99.5049 | 120.403 | 21.1276 | 0.00273 | 0.002534 | 136.2023 | 0.26133 | 8.4001 | 0.003597 | 0.00527 | 0.00529 | 5.9e-05 | 0.00087 | 0.33555 | 1.3459 | 2.5479 | 99.52 | 20.89 | 37.321 | 87.3652 | 108.3152 | 82.6386 | 0.009274 | 0.99229 | 0.023909 | 24.0949 | 0.013956 | 0.15547 | 0.002976 | 0.20425 | 0.20425 | 0.031285
2019-03-21 12:04 | 1553184213.397 | 1100 | 0.20894 | 0.000918 | 0.000156 | 99.6764 | 99.5131 | 119.9572 | 21.1276 | 0.002756 | 0.002551 | 136.1829 | 0.26213 | 8.4032 | 0.00356 | 0.005178 | 0.004987 | 5.8e-05 | 0.000871 | 0.31996 | 1.2898 | 2.5478 | 99.52 | 20.89 | 37.322 | 87.9838 | 107.4045 | 83.0644 | 0.009274 | 0.9911 | 0.023851 | 24.0647 | 0.014664 | 0.15594 | 0.003095 | 0.21243 | 0.21243 | 0.030171
2019-03-21 12:04 | 1553184214.3974 | 1101 | 0.20895 | 0.000918 | 0.000155 | 99.6448 | 99.3574 | 120.0399 | 21.1273 | 0.002709 | 0.002616 | 136.1917 | 0.26126 | 8.4028 | 0.003617 | 0.005451 | 0.005392 | 6e-05 | 0.000871 | 0.32139 | 1.3369 | 2.5479 | 99.52 | 20.89 | 37.321 | 86.8653 | 108.8114 | 84.6259 | 0.009274 | 0.98678 | 0.023807 | 24.1257 | 0.025831 | 0.15542 | 0.003042 | 0.20829 | 0.20829 | 0.030668
2019-03-21 12:04 | 1553184215.3965 | 1102 | 0.20894 | 0.00091 | 0.000155 | 99.6433 | 99.557 | 119.5735 | 21.1275 | 0.002693 | 0.002523 | 136.17 | 0.26079 | 8.4021 | 0.003477 | 0.004775 | 0.004677 | 6e-05 | 0.000872 | 0.28926 | 1.3012 | 2.548 | 99.52 | 20.89 | 37.321 | 86.4721 | 105.4138 | 82.3935 | 0.009274 | 0.99297 | 0.023835 | 24.0035 | 0.007748 | 0.15514 | 0.003097 | 0.21243 | 0.21243 | 0.030016
2019-03-21 12:04 | 1553184216.3972 | 1103 | 0.20903 | 0.000905 | 0.000154 | 99.7023 | 99.7648 | 119.4054 | 21.127 | 0.002727 | 0.002488 | 136.1596 | 0.26128 | 8.402 | 0.003548 | 0.00547 | 0.005379 | 5.9e-05 | 0.000874 | 0.30397 | 1.2998 | 2.548 | 99.52 | 20.89 | 37.321 | 87.2992 | 107.1141 | 81.5359 | 0.009274 | 0.99538 | 0.023847 | 23.9577 | -0.005604 | 0.15544 | 0.002608 | 0.17761 | 0.17761 | 0.035969
2019-03-21 12:04 | 1553184217.3962 | 1104 | 0.20897 | 0.000902 | 0.000153 | 99.663 | 99.5863 | 119.5735 | 21.1278 | 0.002738 | 0.002539 | 136.1547 | 0.2611 | 8.402 | 0.003439 | 0.005267 | 0.005161 | 5.9e-05 | 0.000875 | 0.31708 | 1.2892 | 2.5478 | 99.52 | 20.89 | 37.321 | 87.5591 | 104.4883 | 82.7807 | 0.009274 | 0.99189 | 0.023822 | 24.0166 | 0.006882 | 0.15533 | 0.002949 | 0.20186 | 0.20186 | 0.031626
2019-03-21 12:04 | 1553184218.3972 | 1105 | 0.20892 | 0.000897 | 0.000152 | 99.6165 | 99.5928 | 119.3639 | 21.1266 | 0.002691 | 0.002568 | 136.1393 | 0.26115 | 8.4018 | 0.003477 | 0.004884 | 0.004793 | 6.2e-05 | 0.000878 | 0.31045 | 1.3177 | 2.5479 | 99.52 | 20.89 | 37.32 | 86.4313 | 105.3894 | 83.4596 | 0.009274 | 0.99001 | 0.023778 | 24.0184 | 0.002128 | 0.15536 | 0.003259 | 0.22367 | 0.22367 | 0.028548
2019-03-21 12:04 | 1553184219.3964 | 1106 | 0.20896 | 0.00089 | 0.000151 | 99.7246 | 99.4807 | 119.362 | 21.1281 | 0.002748 | 0.002529 | 136.1298 | 0.26039 | 8.4037 | 0.003435 | 0.004782 | 0.004627 | 6e-05 | 0.000877 | 0.26488 | 1.2951 | 2.5479 | 99.52 | 20.9 | 37.321 | 87.7998 | 104.3754 | 82.5384 | 0.00928 | 0.99257 | 0.023809 | 23.9872 | 0.021904 | 0.1549 | 0.003041 | 0.20845 | 0.20845 | 0.030543
2019-03-21 12:04 | 1553184220.3964 | 1107 | 0.20897 | 0.000886 | 0.00015 | 99.6936 | 99.7699 | 119.0061 | 21.1272 | 0.002762 | 0.002543 | 136.1429 | 0.26147 | 8.4026 | 0.003459 | 0.005118 | 0.004986 | 5.9e-05 | 0.000877 | 0.29159 | 1.3157 | 2.5478 | 99.52 | 20.9 | 37.32 | 88.1403 | 104.9691 | 82.8589 | 0.00928 | 0.99168 | 0.023763 | 23.9622 | -0.006837 | 0.15555 | 0.002977 | 0.20356 | 0.20356 | 0.031406
2019-03-21 12:04 | 1553184221.3967 | 1108 | 0.20901 | 0.000881 | 0.000148 | 99.6793 | 99.6435 | 118.882 | 21.1278 | 0.00271 | 0.002549 | 136.126 | 0.26112 | 8.4019 | 0.003432 | 0.005165 | 0.005221 | 6.3e-05 | 0.00088 | 0.34418 | 1.341 | 2.548 | 99.52 | 20.89 | 37.321 | 86.8791 | 104.3177 | 83.0108 | 0.009274 | 0.99125 | 0.023745 | 23.9548 | 0.003215 | 0.15534 | 0.002772 | 0.18889 | 0.18889 | 0.033799
2019-03-21 12:04 | 1553184222.3975 | 1109 | 0.20896 | 0.000881 | 0.000147 | 99.7196 | 99.8605 | 118.3888 | 21.1273 | 0.002725 | 0.002561 | 136.1046 | 0.26044 | 8.4012 | 0.003381 | 0.004823 | 0.004649 | 5.8e-05 | 0.000879 | 0.2995 | 1.3178 | 2.548 | 99.52 | 20.9 | 37.321 | 87.2549 | 103.0659 | 83.2928 | 0.00928 | 0.99047 | 0.023687 | 23.9145 | -0.012631 | 0.15494 | 0.003068 | 0.20957 | 0.20957 | 0.030386
2019-03-21 12:04 | 1553184223.3964 | 1110 | 0.20899 | 0.00088 | 0.000145 | 99.6701 | 99.6582 | 118.2412 | 21.1286 | 0.002724 | 0.002523 | 136.0912 | 0.26149 | 8.4034 | 0.003393 | 0.004792 | 0.004675 | 6.1e-05 | 0.000879 | 0.29302 | 1.3238 | 2.548 | 99.52 | 20.9 | 37.321 | 87.2117 | 103.3581 | 82.3877 | 0.00928 | 0.99299 | 0.023702 | 23.8692 | 0.001071 | 0.15556 | 0.002873 | 0.1959 | 0.1959 | 0.032637
2019-03-21 12:04 | 1553184224.3966 | 1111 | 0.20898 | 0.000873 | 0.000145 | 99.6647 | 99.6834 | 117.9978 | 21.1278 | 0.002729 | 0.002497 | 136.0861 | 0.26179 | 8.4028 | 0.003462 | 0.004951 | 0.004871 | 6.3e-05 | 0.000878 | 0.2975 | 1.3207 | 2.5479 | 99.52 | 20.9 | 37.322 | 87.3389 | 105.046 | 81.7698 | 0.00928 | 0.99472 | 0.023698 | 23.8239 | -0.001672 | 0.15574 | 0.00293 | 0.19995 | 0.19995 | 0.032013
2019-03-21 12:04 | 1553184225.3976 | 1112 | 0.209 | 0.000866 | 0.000144 | 99.7246 | 99.5669 | 118.0459 | 21.1275 | 0.002723 | 0.002528 | 136.0819 | 0.26197 | 8.4024 | 0.003486 | 0.004746 | 0.004638 | 6e-05 | 0.000883 | 0.32587 | 1.3079 | 2.5479 | 99.52 | 20.9 | 37.323 | 87.2058 | 105.6224 | 82.4957 | 0.00928 | 0.99269 | 0.023679 | 23.8531 | 0.014159 | 0.15585 | 0.002822 | 0.1922 | 0.1922 | 0.033326
2019-03-21 12:04 | 1553184226.3967 | 1113 | 0.20904 | 0.000864 | 0.000143 | 99.7842 | 99.8905 | 118.0128 | 21.1283 | 0.002704 | 0.002512 | 136.0985 | 0.26121 | 8.401 | 0.00335 | 0.005346 | 0.005192 | 6e-05 | 0.00088 | 0.33339 | 1.319 | 2.5479 | 99.52 | 20.9 | 37.323 | 86.7508 | 102.3363 | 82.1119 | 0.00928 | 0.99376 | 0.023688 | 23.8369 | -0.009517 | 0.15539 | 0.002625 | 0.17831 | 0.17831 | 0.035818
2019-03-21 12:04 | 1553184227.3977 | 1114 | 0.20901 | 0.00086 | 0.000141 | 99.6996 | 99.7052 | 117.6715 | 21.1281 | 0.002716 | 0.00254 | 136.0755 | 0.26118 | 8.4021 | 0.003446 | 0.004984 | 0.004964 | 6.1e-05 | 0.000879 | 0.29192 | 1.3244 | 2.548 | 99.52 | 20.89 | 37.324 | 87.0379 | 104.657 | 82.7848 | 0.009275 | 0.99188 | 0.023632 | 23.825 | -0.000499 | 0.15538 | 0.002773 | 0.18845 | 0.18845 | 0.033886
2019-03-21 12:04 | 1553184228.3968 | 1115 | 0.20901 | 0.000856 | 0.00014 | 99.7377 | 99.6061 | 117.8297 | 21.1278 | 0.002677 | 0.002597 | 136.0975 | 0.2616 | 8.4049 | 0.003309 | 0.005059 | 0.004939 | 6e-05 | 0.000879 | 0.29286 | 1.3121 | 2.5482 | 99.52 | 20.89 | 37.325 | 86.086 | 101.3353 | 84.1685 | 0.009275 | 0.98804 | 0.023602 | 23.8873 | 0.011817 | 0.15562 | 0.002774 | 0.18835 | 0.18835 | 0.033959
2019-03-21 12:04 | 1553184229.3974 | 1116 | 0.20906 | 0.00085 | 0.000139 | 99.7377 | 99.6803 | 117.7349 | 21.1286 | 0.002678 | 0.002457 | 136.1309 | 0.26149 | 8.4026 | 0.003433 | 0.005247 | 0.005091 | 5.9e-05 | 0.000876 | 0.34202 | 1.3334 | 2.5479 | 99.52 | 20.89 | 37.324 | 86.1062 | 104.3436 | 80.8022 | 0.009275 | 0.99744 | 0.023704 | 23.7649 | 0.005152 | 0.15556 | 0.002487 | 0.16882 | 0.16882 | 0.037872
2019-03-21 12:04 | 1553184230.3968 | 1117 | 0.209 | 0.000847 | 0.000138 | 99.7789 | 99.7886 | 117.9274 | 21.128 | 0.002722 | 0.002471 | 136.1024 | 0.26147 | 8.4031 | 0.003378 | 0.004886 | 0.004809 | 6.2e-05 | 0.000879 | 0.31404 | 1.3328 | 2.5479 | 99.52 | 20.89 | 37.324 | 87.1743 | 102.9935 | 81.1283 | 0.009275 | 0.99652 | 0.023712 | 23.7953 | -0.000875 | 0.15555 | 0.002857 | 0.19525 | 0.19525 | 0.032743
2019-03-21 12:04 | 1553184231.396 | 1118 | 0.20902 | 0.000843 | 0.000138 | 99.7434 | 99.6589 | 118.0607 | 21.1277 | 0.002682 | 0.002559 | 136.0897 | 0.26188 | 8.4039 | 0.003496 | 0.004228 | 0.004081 | 5.4e-05 | 0.000876 | 0.28407 | 1.3289 | 2.548 | 99.52 | 20.89 | 37.324 | 86.2076 | 105.8672 | 83.2556 | 0.009275 | 0.99057 | 0.023655 | 23.8801 | 0.007585 | 0.15579 | 0.002731 | 0.18586 | 0.18586 | 0.03445
2019-03-21 12:04 | 1553184232.397 | 1119 | 0.20897 | 0.000846 | 0.000138 | 99.747 | 99.6943 | 118.1299 | 21.1281 | 0.002715 | 0.002506 | 136.0937 | 0.26203 | 8.4002 | 0.003435 | 0.005408 | 0.005269 | 6e-05 | 0.000874 | 0.31022 | 1.3519 | 2.548 | 99.52 | 20.9 | 37.325 | 87.0049 | 104.3873 | 81.9706 | 0.009281 | 0.99416 | 0.023705 | 23.844 | 0.004728 | 0.15588 | 0.003027 | 0.20734 | 0.20734 | 0.0309
2019-03-21 12:04 | 1553184233.397 | 1120 | 0.20904 | 0.000836 | 0.000137 | 99.7753 | 99.9289 | 118.2642 | 21.1275 | 0.002755 | 0.002534 | 136.0917 | 0.26193 | 8.4016 | 0.003469 | 0.005094 | 0.005057 | 5.6e-05 | 0.000877 | 0.31061 | 1.3041 | 2.5479 | 99.52 | 20.89 | 37.325 | 87.9695 | 105.2077 | 82.6563 | 0.009275 | 0.99224 | 0.023695 | 23.8806 | -0.013764 | 0.15582 | 0.002642 | 0.17987 | 0.17987 | 0.035606
2019-03-21 12:04 | 1553184234.3971 | 1121 | 0.20902 | 0.000833 | 0.000137 | 99.7976 | 99.7304 | 118.2963 | 21.128 | 0.002681 | 0.002538 | 136.0943 | 0.26141 | 8.4022 | 0.003427 | 0.005184 | 0.005125 | 5.7e-05 | 0.000875 | 0.31421 | 1.3026 | 2.548 | 99.52 | 20.9 | 37.324 | 86.1743 | 104.1897 | 82.7566 | 0.009281 | 0.99196 | 0.023695 | 23.8872 | 0.00603 | 0.15551 | 0.002794 | 0.19068 | 0.19068 | 0.03352
2019-03-21 12:04 | 1553184235.3976 | 1122 | 0.20905 | 0.000832 | 0.000137 | 99.809 | 99.649 | 118.4731 | 21.1273 | 0.002737 | 0.002506 | 136.1115 | 0.26191 | 8.4001 | 0.003469 | 0.004991 | 0.004866 | 5.6e-05 | 0.000873 | 0.31094 | 1.3115 | 2.5479 | 99.52 | 20.9 | 37.325 | 87.5438 | 105.2076 | 81.9855 | 0.009281 | 0.99411 | 0.023739 | 23.8791 | 0.014348 | 0.15581 | 0.002569 | 0.17503 | 0.17503 | 0.036587
2019-03-21 12:04 | 1553184236.3972 | 1123 | 0.20902 | 0.000829 | 0.000137 | 99.8289 | 99.6057 | 118.2961 | 21.1277 | 0.002713 | 0.002513 | 136.0906 | 0.26139 | 8.4039 | 0.003421 | 0.00466 | 0.00462 | 5.4e-05 | 0.000873 | 0.31957 | 1.3114 | 2.5479 | 99.52 | 20.9 | 37.325 | 86.9656 | 104.0465 | 82.1356 | 0.009281 | 0.99369 | 0.023716 | 23.8663 | 0.02002 | 0.1555 | 0.002748 | 0.18762 | 0.18762 | 0.034063
2019-03-21 12:04 | 1553184237.3966 | 1124 | 0.20903 | 0.000827 | 0.000136 | 99.8018 | 99.6807 | 118.3841 | 21.1276 | 0.002714 | 0.002508 | 136.1107 | 0.2617 | 8.4041 | 0.003428 | 0.005466 | 0.005396 | 5.6e-05 | 0.000873 | 0.32565 | 1.3039 | 2.5479 | 99.52 | 20.9 | 37.324 | 86.9835 | 104.2145 | 82.0357 | 0.009281 | 0.99397 | 0.023728 | 23.8718 | 0.010864 | 0.15568 | 0.002724 | 0.18602 | 0.18602 | 0.034398
2019-03-21 12:04 | 1553184238.3976 | 1125 | 0.209 | 0.000822 | 0.000137 | 99.811 | 99.7375 | 118.2762 | 21.1276 | 0.002738 | 0.002478 | 136.0906 | 0.26168 | 8.4047 | 0.003459 | 0.005057 | 0.004937 | 5.8e-05 | 0.000875 | 0.31813 | 1.3024 | 2.5479 | 99.52 | 20.9 | 37.323 | 87.5527 | 104.9545 | 81.294 | 0.00928 | 0.99605 | 0.023742 | 23.836 | 0.00659 | 0.15567 | 0.00291 | 0.19934 | 0.19934 | 0.032096
2019-03-21 12:04 | 1553184239.3974 | 1126 | 0.20896 | 0.00082 | 0.000137 | 99.7929 | 99.7508 | 118.4649 | 21.1275 | 0.002651 | 0.002492 | 136.0714 | 0.26172 | 8.4033 | 0.00348 | 0.00518 | 0.005056 | 5.7e-05 | 0.000874 | 0.30021 | 1.2811 | 2.5479 | 99.52 | 20.9 | 37.322 | 85.4619 | 105.4696 | 81.6448 | 0.00928 | 0.99507 | 0.023749 | 23.8668 | 0.00377 | 0.1557 | 0.00315 | 0.21652 | 0.21652 | 0.029555
2019-03-21 12:04 | 1553184240.3968 | 1127 | 0.20902 | 0.000819 | 0.000138 | 99.8426 | 99.7781 | 118.6468 | 21.1284 | 0.002691 | 0.002477 | 136.0833 | 0.26098 | 8.4018 | 0.003414 | 0.004356 | 0.004282 | 5.9e-05 | 0.000878 | 0.30735 | 1.2838 | 2.5479 | 99.52 | 20.9 | 37.322 | 86.4354 | 103.8639 | 81.2901 | 0.00928 | 0.99606 | 0.023779 | 23.8732 | 0.005786 | 0.15526 | 0.002786 | 0.19073 | 0.19073 | 0.033456
2019-03-21 12:04 | 1553184241.3969 | 1128 | 0.20904 | 0.000823 | 0.000139 | 99.8319 | 99.7883 | 118.6489 | 21.1285 | 0.002673 | 0.002496 | 136.1014 | 0.26155 | 8.4041 | 0.003342 | 0.004642 | 0.004435 | 5.9e-05 | 0.000877 | 0.27992 | 1.267 | 2.5479 | 99.52 | 20.9 | 37.322 | 85.9851 | 102.1259 | 81.7371 | 0.00928 | 0.99481 | 0.023764 | 23.8885 | 0.003909 | 0.1556 | 0.002685 | 0.18334 | 0.18334 | 0.034881
2019-03-21 12:04 | 1553184242.3966 | 1129 | 0.209 | 0.000816 | 0.000139 | 99.865 | 99.8625 | 119.0446 | 21.1276 | 0.002674 | 0.002514 | 136.0845 | 0.26169 | 8.4027 | 0.003381 | 0.00447 | 0.004464 | 6.2e-05 | 0.000879 | 0.29767 | 1.2761 | 2.5479 | 99.52 | 20.9 | 37.32 | 86.0264 | 103.0686 | 82.1681 | 0.00928 | 0.9936 | 0.02379 | 23.9428 | 0.000219 | 0.15568 | 0.002914 | 0.19994 | 0.19994 | 0.032002
2019-03-21 12:04 | 1553184243.3965 | 1130 | 0.20896 | 0.000814 | 0.000138 | 99.8864 | 99.7018 | 119.235 | 21.1281 | 0.002643 | 0.002494 | 136.0967 | 0.26172 | 8.4038 | 0.003337 | 0.005264 | 0.00516 | 5.8e-05 | 0.000873 | 0.30038 | 1.3433 | 2.5479 | 99.52 | 20.9 | 37.319 | 85.2722 | 102.0132 | 81.6898 | 0.009279 | 0.99494 | 0.023825 | 23.9458 | 0.01655 | 0.1557 | 0.003152 | 0.21728 | 0.21728 | 0.029452
2019-03-21 12:04 | 1553184244.397 | 1131 | 0.20908 | 0.000813 | 0.000138 | 99.8906 | 99.7222 | 119.3974 | 21.1279 | 0.002691 | 0.0025 | 136.0922 | 0.26183 | 8.4023 | 0.003351 | 0.005031 | 0.005057 | 5.8e-05 | 0.000869 | 0.28711 | 1.2843 | 2.5478 | 99.52 | 20.89 | 37.318 | 86.4161 | 102.3516 | 81.8268 | 0.009273 | 0.99456 | 0.023836 | 23.9667 | 0.015091 | 0.15576 | 0.002434 | 0.16604 | 0.16604 | 0.038556
2019-03-21 12:04 | 1553184245.3968 | 1132 | 0.20903 | 0.000813 | 0.000137 | 99.9064 | 99.8223 | 119.5772 | 21.1285 | 0.002658 | 0.00254 | 136.0699 | 0.26184 | 8.4013 | 0.003385 | 0.004199 | 0.004065 | 5.7e-05 | 0.000872 | 0.30431 | 1.308 | 2.5479 | 99.52 | 20.89 | 37.316 | 85.6239 | 103.1638 | 82.7939 | 0.009273 | 0.99186 | 0.023822 | 24.0174 | 0.007527 | 0.15577 | 0.002746 | 0.18833 | 0.18833 | 0.033994
2019-03-21 12:04 | 1553184246.3971 | 1133 | 0.20901 | 0.000814 | 0.000137 | 99.8644 | 99.7699 | 119.8413 | 21.1283 | 0.002681 | 0.00254 | 136.0719 | 0.26236 | 8.4034 | 0.003444 | 0.004826 | 0.004737 | 5.6e-05 | 0.000872 | 0.27865 | 1.3004 | 2.548 | 99.52 | 20.89 | 37.315 | 86.1801 | 104.6076 | 82.793 | 0.009273 | 0.99186 | 0.023848 | 24.0439 | 0.008467 | 0.15608 | 0.002844 | 0.19555 | 0.19555 | 0.032803
2019-03-21 12:04 | 1553184247.3958 | 1134 | 0.20904 | 0.000806 | 0.000136 | 99.8915 | 99.9593 | 119.9969 | 21.1284 | 0.002692 | 0.002544 | 136.0689 | 0.26104 | 8.4034 | 0.003393 | 0.004847 | 0.00485 | 5.6e-05 | 0.000875 | 0.31852 | 1.3011 | 2.548 | 99.52 | 20.89 | 37.314 | 86.4486 | 103.3743 | 82.8942 | 0.009272 | 0.99158 | 0.02386 | 24.0629 | -0.006066 | 0.15529 | 0.002685 | 0.18432 | 0.18432 | 0.034627
2019-03-21 12:04 | 1553184248.3959 | 1135 | 0.20901 | 0.000809 | 0.000135 | 99.9409 | 100.0025 | 119.7307 | 21.1275 | 0.002671 | 0.002512 | 136.066 | 0.26209 | 8.4026 | 0.003413 | 0.005097 | 0.004922 | 5.8e-05 | 0.000872 | 0.30165 | 1.3075 | 2.548 | 99.52 | 20.89 | 37.314 | 85.9467 | 103.8616 | 82.1229 | 0.009272 | 0.99373 | 0.02386 | 24.0102 | -0.005508 | 0.15592 | 0.002858 | 0.19674 | 0.19674 | 0.032572
2019-03-21 12:04 | 1553184249.3963 | 1136 | 0.20905 | 0.000807 | 0.000134 | 100.0184 | 99.9569 | 119.6552 | 21.1267 | 0.002622 | 0.002482 | 136.0531 | 0.2616 | 8.4002 | 0.003412 | 0.004864 | 0.004776 | 5.8e-05 | 0.000873 | 0.32316 | 1.3341 | 2.5479 | 99.52 | 20.89 | 37.314 | 84.7738 | 103.8245 | 81.4063 | 0.009272 | 0.99574 | 0.023876 | 23.9784 | 0.005505 | 0.15562 | 0.00266 | 0.18279 | 0.18279 | 0.034991
2019-03-21 12:04 | 1553184250.396 | 1137 | 0.20904 | 0.000802 | 0.000132 | 100.0321 | 99.9419 | 119.407 | 21.1281 | 0.002704 | 0.002489 | 136.0352 | 0.2616 | 8.4007 | 0.003396 | 0.005009 | 0.004977 | 5.7e-05 | 0.000874 | 0.33339 | 1.3302 | 2.5479 | 99.52 | 20.89 | 37.314 | 86.7326 | 103.4511 | 81.5745 | 0.009272 | 0.99527 | 0.023846 | 23.9592 | 0.008072 | 0.15563 | 0.002701 | 0.18554 | 0.18554 | 0.034473
2019-03-21 12:04 | 1553184251.397 | 1138 | 0.20907 | 0.000796 | 0.000131 | 99.9925 | 99.9698 | 119.4286 | 21.1275 | 0.002669 | 0.002532 | 136.0194 | 0.2626 | 8.4012 | 0.003322 | 0.005746 | 0.005626 | 5.7e-05 | 0.000874 | 0.29087 | 1.3061 | 2.5478 | 99.52 | 20.89 | 37.315 | 85.9007 | 101.6593 | 82.5987 | 0.009273 | 0.9924 | 0.023814 | 23.9959 | 0.002029 | 0.15622 | 0.002529 | 0.17307 | 0.17307 | 0.037099
2019-03-21 12:04 | 1553184252.3975 | 1139 | 0.20908 | 0.000803 | 0.000131 | 99.9967 | 99.8653 | 119.0193 | 21.1281 | 0.002685 | 0.002476 | 136.0439 | 0.26076 | 8.4028 | 0.003315 | 0.004927 | 0.004841 | 5.8e-05 | 0.000872 | 0.29966 | 1.3208 | 2.5479 | 99.52 | 20.89 | 37.314 | 86.2859 | 101.4674 | 81.2586 | 0.009272 | 0.99615 | 0.023818 | 23.9096 | 0.011762 | 0.15512 | 0.002428 | 0.16588 | 0.16588 | 0.038435
2019-03-21 12:04 | 1553184253.3966 | 1140 | 0.20907 | 0.000794 | 0.00013 | 99.9448 | 100.0812 | 118.973 | 21.1279 | 0.002676 | 0.002531 | 136.0234 | 0.25983 | 8.4025 | 0.003358 | 0.005495 | 0.005399 | 5.6e-05 | 0.000874 | 0.32388 | 1.3014 | 2.5479 | 99.52 | 20.89 | 37.315 | 86.0664 | 102.5173 | 82.567 | 0.009273 | 0.99249 | 0.023769 | 23.949 | -0.012196 | 0.15457 | 0.002521 | 0.17223 | 0.17223 | 0.036888
2019-03-21 12:04 | 1553184254.3972 | 1141 | 0.20904 | 0.000792 | 0.00013 | 100.0229 | 99.9657 | 118.9179 | 21.1278 | 0.002639 | 0.002468 | 136.0121 | 0.26 | 8.402 | 0.003417 | 0.004707 | 0.004684 | 6e-05 | 0.000874 | 0.29087 | 1.2844 | 2.5479 | 99.52 | 20.89 | 37.316 | 85.1688 | 103.9411 | 81.0617 | 0.009273 | 0.99671 | 0.023814 | 23.8928 | 0.005112 | 0.15468 | 0.002721 | 0.1869 | 0.1869 | 0.034015
2019-03-21 12:04 | 1553184255.3976 | 1142 | 0.20911 | 0.000788 | 0.000129 | 100.0607 | 99.9415 | 119.0216 | 21.128 | 0.002734 | 0.002481 | 136.0121 | 0.26074 | 8.4029 | 0.003363 | 0.004684 | 0.00452 | 5.7e-05 | 0.000876 | 0.32803 | 1.2696 | 2.548 | 99.52 | 20.9 | 37.317 | 87.4532 | 102.6479 | 81.3785 | 0.009279 | 0.99582 | 0.023814 | 23.9139 | 0.010659 | 0.15512 | 0.002292 | 0.15622 | 0.15622 | 0.04081
2019-03-21 12:04 | 1553184256.3977 | 1143 | 0.20906 | 0.000788 | 0.000128 | 100.0452 | 100.1633 | 119.1944 | 21.1293 | 0.002635 | 0.002479 | 135.9963 | 0.26091 | 8.401 | 0.003409 | 0.004875 | 0.004823 | 5.9e-05 | 0.000872 | 0.30718 | 1.2735 | 2.5478 | 99.52 | 20.9 | 37.316 | 85.0916 | 103.7657 | 81.3386 | 0.009279 | 0.99593 | 0.023832 | 23.9299 | -0.010547 | 0.15522 | 0.002592 | 0.17786 | 0.17786 | 0.035867
2019-03-21 12:04 | 1553184257.3964 | 1144 | 0.209 | 0.000786 | 0.000129 | 100.0634 | 100.0325 | 119.2381 | 21.1288 | 0.002656 | 0.002452 | 135.9773 | 0.26156 | 8.4042 | 0.003361 | 0.004779 | 0.004743 | 5.9e-05 | 0.000872 | 0.30685 | 1.3147 | 2.5479 | 99.52 | 20.9 | 37.316 | 85.5904 | 102.6022 | 80.6868 | 0.009279 | 0.99776 | 0.023859 | 23.9122 | 0.002764 | 0.1556 | 0.002956 | 0.20413 | 0.20413 | 0.031329
2019-03-21 12:04 | 1553184258.3968 | 1145 | 0.20908 | 0.000783 | 0.000128 | 100.1394 | 100.1456 | 119.7526 | 21.1295 | 0.002665 | 0.002508 | 135.9601 | 0.2615 | 8.4023 | 0.003452 | 0.004403 | 0.004348 | 5.7e-05 | 0.00087 | 0.32189 | 1.3528 | 2.5481 | 99.52 | 20.9 | 37.313 | 85.8018 | 104.7897 | 82.0209 | 0.009278 | 0.99401 | 0.023865 | 24.0089 | -0.000552 | 0.15557 | 0.002481 | 0.17022 | 0.17022 | 0.037562
2019-03-21 12:04 | 1553184259.3975 | 1146 | 0.20914 | 0.000776 | 0.000127 | 100.109 | 100.0424 | 119.7181 | 21.1289 | 0.002622 | 0.00251 | 135.9358 | 0.26035 | 8.404 | 0.003394 | 0.005561 | 0.005469 | 5.6e-05 | 0.000869 | 0.27743 | 1.2768 | 2.5479 | 99.52 | 20.9 | 37.312 | 84.7703 | 103.3952 | 82.079 | 0.009278 | 0.99385 | 0.02386 | 24.0074 | 0.005955 | 0.15488 | 0.002134 | 0.14531 | 0.14531 | 0.043806
2019-03-21 12:04 | 1553184260.3964 | 1147 | 0.20906 | 0.000778 | 0.000126 | 100.1221 | 100.1043 | 119.6176 | 21.1279 | 0.00262 | 0.002488 | 135.9467 | 0.26141 | 8.4004 | 0.003395 | 0.005296 | 0.005184 | 5.8e-05 | 0.000872 | 0.33395 | 1.3273 | 2.5479 | 99.52 | 20.9 | 37.311 | 84.7211 | 103.4076 | 81.5399 | 0.009277 | 0.99536 | 0.023868 | 23.9791 | 0.001586 | 0.15551 | 0.002602 | 0.179 | 0.179 | 0.035707
2019-03-21 12:04 | 1553184261.397 | 1148 | 0.20908 | 0.000777 | 0.000126 | 100.1581 | 99.9916 | 119.6251 | 21.1287 | 0.002662 | 0.002508 | 135.9374 | 0.26065 | 8.4025 | 0.003477 | 0.004411 | 0.004286 | 6.1e-05 | 0.000874 | 0.31885 | 1.3136 | 2.5479 | 99.52 | 20.89 | 37.311 | 85.7294 | 105.3992 | 82.0344 | 0.009272 | 0.99398 | 0.023852 | 23.9966 | 0.014885 | 0.15506 | 0.002514 | 0.1726 | 0.1726 | 0.036923
2019-03-21 12:04 | 1553184262.3966 | 1149 | 0.20901 | 0.000774 | 0.000126 | 100.1531 | 100.0338 | 119.7549 | 21.1286 | 0.002638 | 0.002493 | 135.9187 | 0.26116 | 8.4032 | 0.003529 | 0.004237 | 0.004126 | 5.8e-05 | 0.000873 | 0.31797 | 1.3268 | 2.548 | 99.52 | 20.89 | 37.312 | 85.1602 | 106.6645 | 81.6716 | 0.009272 | 0.99499 | 0.023877 | 23.9973 | 0.010655 | 0.15536 | 0.002917 | 0.20164 | 0.20164 | 0.031667
2019-03-21 12:04 | 1553184263.3966 | 1150 | 0.2091 | 0.000772 | 0.000126 | 100.1421 | 100.0941 | 119.6901 | 21.1282 | 0.002671 | 0.002527 | 135.9251 | 0.26057 | 8.4029 | 0.003441 | 0.005389 | 0.005311 | 5.6e-05 | 0.000872 | 0.29878 | 1.2915 | 2.5479 | 99.52 | 20.89 | 37.313 | 85.9411 | 104.5384 | 82.4721 | 0.009272 | 0.99275 | 0.023844 | 24.0179 | 0.004282 | 0.15502 | 0.002369 | 0.16216 | 0.16216 | 0.03929
2019-03-21 12:04 | 1553184264.3972 | 1151 | 0.20911 | 0.000772 | 0.000127 | 100.092 | 100.2222 | 119.7101 | 21.1292 | 0.002648 | 0.002484 | 135.9087 | 0.26095 | 8.4023 | 0.003488 | 0.004688 | 0.004563 | 6e-05 | 0.000874 | 0.32333 | 1.3171 | 2.5479 | 99.52 | 20.89 | 37.314 | 85.3957 | 105.6689 | 81.4484 | 0.009272 | 0.99562 | 0.02388 | 23.9853 | -0.011626 | 0.15524 | 0.002327 | 0.15932 | 0.15932 | 0.040048
2019-03-21 12:04 | 1553184265.3962 | 1152 | 0.20899 | 0.00077 | 0.000127 | 100.1135 | 100.1295 | 119.4068 | 21.1293 | 0.00267 | 0.002459 | 135.9094 | 0.26204 | 8.4008 | 0.003445 | 0.004762 | 0.004631 | 5.5e-05 | 0.000874 | 0.29407 | 1.283 | 2.5478 | 99.52 | 20.89 | 37.317 | 85.9328 | 104.6291 | 80.8401 | 0.009273 | 0.99733 | 0.02387 | 23.9343 | -0.001438 | 0.15589 | 0.003018 | 0.20874 | 0.20874 | 0.030695
2019-03-21 12:04 | 1553184266.3968 | 1153 | 0.20918 | 0.000771 | 0.000127 | 100.0881 | 100.0934 | 119.4052 | 21.1285 | 0.002648 | 0.002479 | 135.9206 | 0.26191 | 8.4025 | 0.003485 | 0.005134 | 0.005009 | 5.7e-05 | 0.000875 | 0.31741 | 1.3077 | 2.5479 | 99.52 | 20.89 | 37.32 | 85.3842 | 105.5935 | 81.3392 | 0.009274 | 0.99593 | 0.023853 | 23.951 | -0.000475 | 0.15581 | 0.001903 | 0.12881 | 0.12881 | 0.049717
2019-03-21 12:04 | 1553184267.3973 | 1154 | 0.20907 | 0.000766 | 0.000126 | 100.0899 | 100.1033 | 119.5331 | 21.1293 | 0.002629 | 0.002463 | 135.9167 | 0.26119 | 8.4016 | 0.003537 | 0.005042 | 0.004943 | 5.6e-05 | 0.000878 | 0.32156 | 1.2953 | 2.548 | 99.52 | 20.89 | 37.322 | 84.9463 | 106.87 | 80.937 | 0.009274 | 0.99706 | 0.02388 | 23.9503 | -0.001198 | 0.15538 | 0.002584 | 0.17782 | 0.17782 | 0.035915
2019-03-21 12:04 | 1553184268.3973 | 1155 | 0.20896 | 0.000767 | 0.000125 | 100.0729 | 100.1926 | 119.477 | 21.1288 | 0.002711 | 0.002427 | 135.9421 | 0.26156 | 8.4052 | 0.003474 | 0.004673 | 0.004592 | 5.5e-05 | 0.000879 | 0.28351 | 1.2805 | 2.5479 | 99.52 | 20.89 | 37.328 | 86.8981 | 105.3239 | 80.0677 | 0.009276 | 0.99951 | 0.023904 | 23.9152 | -0.010686 | 0.1556 | 0.003263 | 0.22676 | 0.22676 | 0.028203
2019-03-21 12:04 | 1553184269.3959 | 1156 | 0.20906 | 0.000765 | 0.000126 | 100.1352 | 100.262 | 119.5732 | 21.1297 | 0.00264 | 0.002459 | 135.9551 | 0.26106 | 8.4035 | 0.003441 | 0.00431 | 0.004266 | 5.9e-05 | 0.000881 | 0.31238 | 1.2808 | 2.5479 | 99.52 | 20.89 | 37.333 | 85.1894 | 104.5297 | 80.8415 | 0.009277 | 0.99733 | 0.023887 | 23.9511 | -0.011322 | 0.15531 | 0.002613 | 0.17999 | 0.17999 | 0.035463
2019-03-21 12:04 | 1553184270.3965 | 1157 | 0.20904 | 0.000759 | 0.000126 | 100.1522 | 100.2014 | 120.0528 | 21.129 | 0.002668 | 0.002452 | 135.9386 | 0.2615 | 8.4032 | 0.00333 | 0.005096 | 0.004986 | 5.6e-05 | 0.000878 | 0.33843 | 1.2766 | 2.5479 | 99.52 | 20.89 | 37.336 | 85.8834 | 101.8398 | 80.6863 | 0.009278 | 0.99777 | 0.02394 | 23.9938 | -0.004395 | 0.15557 | 0.002753 | 0.19041 | 0.19041 | 0.033579
2019-03-21 12:05 | 1553184271.3961 | 1158 | 0.20907 | 0.000759 | 0.000126 | 100.1811 | 100.1544 | 119.9842 | 21.129 | 0.002672 | 0.002436 | 135.9589 | 0.26201 | 8.4007 | 0.003429 | 0.00486 | 0.004742 | 5.8e-05 | 0.000872 | 0.32211 | 1.2955 | 2.548 | 99.52 | 20.89 | 37.341 | 85.97 | 104.2396 | 80.2933 | 0.009279 | 0.99887 | 0.023947 | 23.9736 | 0.002382 | 0.15587 | 0.002589 | 0.17865 | 0.17865 | 0.035859
2019-03-21 12:05 | 1553184272.3965 | 1159 | 0.20906 | 0.000761 | 0.000127 | 100.1283 | 100.1667 | 119.8983 | 21.1291 | 0.002612 | 0.002506 | 135.9522 | 0.26199 | 8.4012 | 0.003413 | 0.004986 | 0.004919 | 5.7e-05 | 0.000868 | 0.28174 | 1.3057 | 2.548 | 99.52 | 20.89 | 37.344 | 84.5279 | 103.8481 | 81.9701 | 0.00928 | 0.99416 | 0.023881 | 24.0218 | -0.003423 | 0.15585 | 0.00265 | 0.18249 | 0.18249 | 0.035102
2019-03-21 12:05 | 1553184273.3972 | 1160 | 0.20913 | 0.000765 | 0.000127 | 100.1981 | 100.2607 | 119.6899 | 21.1289 | 0.00261 | 0.002455 | 135.9388 | 0.26206 | 8.4009 | 0.003448 | 0.004769 | 0.004742 | 5.4e-05 | 0.000866 | 0.32228 | 1.3053 | 2.5479 | 99.52 | 20.89 | 37.347 | 84.4902 | 104.7066 | 80.7612 | 0.009281 | 0.99755 | 0.023901 | 23.96 | -0.005587 | 0.1559 | 0.00219 | 0.14966 | 0.14966 | 0.042813
2019-03-21 12:05 | 1553184274.3963 | 1161 | 0.20923 | 0.000756 | 0.000126 | 100.1668 | 100.0907 | 119.2751 | 21.1295 | 0.002645 | 0.002468 | 135.9469 | 0.26284 | 8.4043 | 0.003459 | 0.005069 | 0.004983 | 5.6e-05 | 0.000866 | 0.31421 | 1.3144 | 2.5479 | 99.52 | 20.89 | 37.348 | 85.3282 | 104.977 | 81.0634 | 0.009281 | 0.9967 | 0.02385 | 23.9287 | 0.006795 | 0.15636 | 0.001603 | 0.10733 | 0.10733 | 0.059875
2019-03-21 12:05 | 1553184275.3963 | 1162 | 0.20897 | 0.000755 | 0.000125 | 100.19 | 100.1891 | 119.016 | 21.1289 | 0.002641 | 0.002431 | 135.9318 | 0.26229 | 8.402 | 0.003375 | 0.004707 | 0.004645 | 5.5e-05 | 0.000868 | 0.32012 | 1.3283 | 2.548 | 99.52 | 20.89 | 37.351 | 85.2237 | 102.9362 | 80.1724 | 0.009282 | 0.99922 | 0.023854 | 23.8726 | 7.8e-05 | 0.15604 | 0.003195 | 0.2214 | 0.2214 | 0.028967
2019-03-21 12:05 | 1553184276.3964 | 1163 | 0.20903 | 0.000751 | 0.000123 | 100.1993 | 100.2518 | 118.7158 | 21.1292 | 0.002635 | 0.002474 | 135.9086 | 0.26137 | 8.4024 | 0.003443 | 0.004614 | 0.004492 | 5.3e-05 | 0.000869 | 0.31332 | 1.2882 | 2.5478 | 99.52 | 20.89 | 37.352 | 85.0905 | 104.5708 | 81.2121 | 0.009282 | 0.99628 | 0.023789 | 23.8775 | -0.004691 | 0.15549 | 0.002822 | 0.19428 | 0.19428 | 0.032895
2019-03-21 12:05 | 1553184277.3964 | 1164 | 0.2093 | 0.000757 | 0.000123 | 100.1948 | 100.3295 | 118.3442 | 21.1287 | 0.002585 | 0.002462 | 135.9194 | 0.26217 | 8.4017 | 0.003404 | 0.005441 | 0.005347 | 5.6e-05 | 0.000871 | 0.31188 | 1.2631 | 2.5479 | 99.52 | 20.89 | 37.353 | 83.8885 | 103.6425 | 80.9285 | 0.009282 | 0.99708 | 0.023761 | 23.8306 | -0.012015 | 0.15597 | 0.001185 | 0.077304 | 0.077304 | 0.082922
2019-03-21 12:05 | 1553184278.3969 | 1165 | 0.2089 | 0.000759 | 0.000122 | 100.2585 | 100.4044 | 118.1588 | 21.1283 | 0.002563 | 0.002436 | 135.8912 | 0.26225 | 8.4021 | 0.003272 | 0.004985 | 0.004884 | 5.5e-05 | 0.000869 | 0.31277 | 1.3156 | 2.548 | 99.52 | 20.89 | 37.355 | 83.3522 | 100.4457 | 80.3059 | 0.009283 | 0.99884 | 0.023763 | 23.791 | -0.013003 | 0.15601 | 0.003611 | 0.25042 | 0.25042 | 0.025606
2019-03-21 12:05 | 1553184279.397 | 1166 | 0.20932 | 0.000753 | 0.000122 | 100.2454 | 100.1408 | 118.0609 | 21.1285 | 0.002605 | 0.002481 | 135.8759 | 0.26191 | 8.4016 | 0.003421 | 0.004658 | 0.004596 | 5.6e-05 | 0.000871 | 0.2907 | 1.3232 | 2.548 | 99.52 | 20.89 | 37.357 | 84.3587 | 104.0518 | 81.3744 | 0.009283 | 0.99583 | 0.023718 | 23.817 | 0.009343 | 0.15581 | 0.001104 | 0.071411 | 0.071411 | 0.089677
2019-03-21 12:05 | 1553184280.3956 | 1167 | 0.20867 | 0.000749 | 0.000122 | 100.2973 | 100.3748 | 117.7013 | 21.1284 | 0.002632 | 0.002476 | 135.8717 | 0.26262 | 8.4014 | 0.003466 | 0.004685 | 0.004672 | 5.3e-05 | 0.000873 | 0.34827 | 1.2836 | 2.548 | 99.52 | 20.89 | 37.358 | 85.0024 | 105.1371 | 81.2479 | 0.009283 | 0.99618 | 0.023686 | 23.7765 | -0.006908 | 0.15623 | 0.005017 | 0.34956 | 0.34956 | 0.018369
2019-03-21 12:05 | 1553184281.3966 | 1168 | 0.20914 | 0.000745 | 0.000122 | 100.2869 | 100.3577 | 117.6842 | 21.1286 | 0.00263 | 0.002458 | 135.8443 | 0.26177 | 8.4034 | 0.003326 | 0.004944 | 0.00484 | 5.7e-05 | 0.000874 | 0.33268 | 1.3035 | 2.5479 | 99.52 | 20.89 | 37.359 | 84.9636 | 101.7491 | 80.8325 | 0.009284 | 0.99735 | 0.023698 | 23.7608 | -0.00632 | 0.15573 | 0.002215 | 0.15043 | 0.15043 | 0.042548
2019-03-21 12:05 | 1553184282.3957 | 1169 | 0.20986 | 0.000746 | 0.000124 | 100.3545 | 100.4671 | 117.6713 | 21.1291 | 0.002637 | 0.002385 | 135.8255 | 0.26173 | 8.4021 | 0.003393 | 0.004576 | 0.004457 | 5.6e-05 | 0.000873 | 0.32156 | 1.3034 | 2.548 | 99.52 | 20.89 | 37.359 | 85.1286 | 103.3656 | 79.0606 | 0.009284 | 1.0024 | 0.023756 | 23.7 | -0.010029 | 0.1557 | -0.002164 | -0.16187 | -0.16187 | -0.039533
2019-03-21 12:05 | 1553184283.3961 | 1170 | 0.20906 | 0.000745 | 0.000123 | 100.2609 | 100.4173 | 117.9473 | 21.1291 | 0.002649 | 0.002479 | 135.8309 | 0.26169 | 8.4031 | 0.003418 | 0.004835 | 0.004753 | 5.6e-05 | 0.000872 | 0.32875 | 1.3087 | 2.5478 | 99.52 | 20.89 | 37.359 | 85.4231 | 103.9644 | 81.3173 | 0.009284 | 0.99599 | 0.023708 | 23.8036 | -0.013943 | 0.15568 | 0.002665 | 0.18248 | 0.18248 | 0.035064
2019-03-21 12:05 | 1553184284.3958 | 1171 | 0.20909 | 0.000747 | 0.000124 | 100.2577 | 100.4276 | 118.3621 | 21.1293 | 0.002582 | 0.002454 | 135.8394 | 0.26231 | 8.4034 | 0.003439 | 0.004351 | 0.00422 | 5.7e-05 | 0.000872 | 0.33445 | 1.2859 | 2.5479 | 99.52 | 20.89 | 37.362 | 83.8088 | 104.4733 | 80.7397 | 0.009284 | 0.99761 | 0.023769 | 23.826 | -0.015145 | 0.15605 | 0.002491 | 0.1705 | 0.1705 | 0.037617
2019-03-21 12:05 | 1553184285.3969 | 1172 | 0.20908 | 0.000745 | 0.000123 | 100.336 | 100.434 | 118.5719 | 21.1293 | 0.002637 | 0.002491 | 135.8584 | 0.26212 | 8.4011 | 0.003476 | 0.005036 | 0.004916 | 5.2e-05 | 0.000862 | 0.32101 | 1.2529 | 2.548 | 99.52 | 20.89 | 37.363 | 85.1391 | 105.3731 | 81.6099 | 0.009285 | 0.99517 | 0.023761 | 23.8764 | -0.008733 | 0.15594 | 0.002539 | 0.17385 | 0.17385 | 0.036865
2019-03-21 12:05 | 1553184286.3973 | 1173 | 0.20909 | 0.000742 | 0.000123 | 100.3694 | 100.4238 | 118.9355 | 21.1284 | 0.002571 | 0.002502 | 135.8779 | 0.26234 | 8.4019 | 0.003395 | 0.004649 | 0.004577 | 5.6e-05 | 0.000869 | 0.31614 | 1.2872 | 2.5479 | 99.52 | 20.89 | 37.364 | 83.5448 | 103.4175 | 81.8726 | 0.009285 | 0.99443 | 0.023789 | 23.9219 | -0.004849 | 0.15607 | 0.002511 | 0.17208 | 0.17208 | 0.037277
2019-03-21 12:05 | 1553184287.397 | 1174 | 0.20908 | 0.000745 | 0.000124 | 100.3238 | 100.4565 | 119.3296 | 21.1293 | 0.002664 | 0.002449 | 135.8637 | 0.26201 | 8.4016 | 0.003454 | 0.004611 | 0.004453 | 5.6e-05 | 0.000864 | 0.31094 | 1.3027 | 2.5479 | 99.52 | 20.89 | 37.366 | 85.7821 | 104.8422 | 80.6007 | 0.009285 | 0.99801 | 0.023871 | 23.9185 | -0.011824 | 0.15587 | 0.002542 | 0.17485 | 0.17485 | 0.036639
2019-03-21 12:05 | 1553184288.397 | 1175 | 0.20943 | 0.000741 | 0.000123 | 100.4084 | 100.4991 | 119.5714 | 21.1301 | 0.002645 | 0.00243 | 135.8713 | 0.26198 | 8.4006 | 0.003422 | 0.004484 | 0.004503 | 5.8e-05 | 0.000867 | 0.31116 | 1.3137 | 2.548 | 99.52 | 20.89 | 37.366 | 85.3139 | 104.0732 | 80.1502 | 0.009285 | 0.99928 | 0.02391 | 23.9275 | -0.008073 | 0.15585 | 0.000457 | 0.025404 | 0.025404 | 0.25215
2019-03-21 12:05 | 1553184289.3971 | 1176 | 0.20913 | 0.000747 | 0.000123 | 100.4081 | 100.4759 | 119.5995 | 21.1292 | 0.00259 | 0.002471 | 135.89 | 0.26169 | 8.403 | 0.003521 | 0.005185 | 0.005051 | 5.5e-05 | 0.000864 | 0.32781 | 1.2738 | 2.5479 | 99.52 | 20.89 | 37.367 | 83.9967 | 106.4669 | 81.1318 | 0.009286 | 0.99651 | 0.02388 | 23.9635 | -0.006038 | 0.15568 | 0.002245 | 0.15361 | 0.15361 | 0.041653
2019-03-21 12:05 | 1553184290.3961 | 1177 | 0.20878 | 0.00074 | 0.000122 | 100.3867 | 100.384 | 119.6171 | 21.129 | 0.002649 | 0.0024 | 135.8691 | 0.26212 | 8.4012 | 0.003493 | 0.005067 | 0.004912 | 5.4e-05 | 0.000863 | 0.3314 | 1.3371 | 2.5478 | 99.52 | 20.89 | 37.367 | 85.4275 | 105.7851 | 79.4232 | 0.009286 | 1.0013 | 0.023939 | 23.9074 | 0.000241 | 0.15594 | 0.004356 | 0.30584 | 0.30584 | 0.020955
2019-03-21 12:05 | 1553184291.3965 | 1178 | 0.20913 | 0.000745 | 0.000121 | 100.4817 | 100.619 | 119.4716 | 21.1302 | 0.002543 | 0.002441 | 135.8613 | 0.26158 | 8.4022 | 0.003496 | 0.004676 | 0.004708 | 5.5e-05 | 0.000867 | 0.31598 | 1.3406 | 2.548 | 99.52 | 20.89 | 37.366 | 82.8578 | 105.8563 | 80.4074 | 0.009285 | 0.99855 | 0.023892 | 23.9262 | -0.01221 | 0.15561 | 0.002274 | 0.15595 | 0.15595 | 0.04101
2019-03-21 12:05 | 1553184292.3963 | 1179 | 0.2093 | 0.000742 | 0.00012 | 100.5136 | 100.6786 | 119.7465 | 21.1292 | 0.002608 | 0.002443 | 135.8858 | 0.26248 | 8.4027 | 0.003536 | 0.004867 | 0.004781 | 5.6e-05 | 0.000867 | 0.32731 | 1.3426 | 2.548 | 99.52 | 20.89 | 37.366 | 84.4219 | 106.844 | 80.4526 | 0.009285 | 0.99842 | 0.023917 | 23.9552 | -0.01467 | 0.15615 | 0.001261 | 0.083379 | 0.083379 | 0.076971
2019-03-21 12:05 | 1553184293.3963 | 1180 | 0.20913 | 0.00074 | 0.000119 | 100.4758 | 100.6292 | 119.5753 | 21.1302 | 0.002593 | 0.00242 | 135.8493 | 0.26196 | 8.4018 | 0.00349 | 0.004719 | 0.004595 | 5.4e-05 | 0.000865 | 0.31708 | 1.3383 | 2.5476 | 99.52 | 20.88 | 37.367 | 84.0679 | 105.728 | 79.9155 | 0.00928 | 0.99994 | 0.023919 | 23.9199 | -0.013649 | 0.15584 | 0.002289 | 0.1573 | 0.1573 | 0.040718
2019-03-21 12:05 | 1553184294.3963 | 1181 | 0.20897 | 0.000743 | 0.000118 | 100.5663 | 100.5413 | 119.219 | 21.1285 | 0.002582 | 0.002453 | 135.8392 | 0.26167 | 8.4029 | 0.0035 | 0.005377 | 0.005341 | 5.6e-05 | 0.000868 | 0.30879 | 1.3021 | 2.548 | 99.52 | 20.89 | 37.368 | 83.8143 | 105.9561 | 80.6976 | 0.009286 | 0.99773 | 0.023856 | 23.9107 | 0.002225 | 0.15567 | 0.003203 | 0.22243 | 0.22243 | 0.028764
2019-03-21 12:05 | 1553184295.3964 | 1182 | 0.20909 | 0.000741 | 0.000118 | 100.5151 | 100.6629 | 119.0237 | 21.1296 | 0.00256 | 0.002458 | 135.8238 | 0.261 | 8.4059 | 0.003435 | 0.004994 | 0.004911 | 5.5e-05 | 0.000868 | 0.32228 | 1.3282 | 2.5479 | 99.52 | 20.88 | 37.37 | 83.2878 | 104.3863 | 80.821 | 0.009281 | 0.99738 | 0.023833 | 23.8953 | -0.013145 | 0.15527 | 0.002495 | 0.17153 | 0.17153 | 0.037204
2019-03-21 12:05 | 1553184296.3958 | 1183 | 0.20915 | 0.000736 | 0.000118 | 100.5356 | 100.5934 | 118.7341 | 21.1296 | 0.002574 | 0.002449 | 135.8209 | 0.26253 | 8.4004 | 0.003459 | 0.004877 | 0.00479 | 5.2e-05 | 0.000868 | 0.31813 | 1.3271 | 2.548 | 99.52 | 20.88 | 37.37 | 83.6023 | 104.9715 | 80.6112 | 0.009281 | 0.99798 | 0.023811 | 23.8591 | -0.00514 | 0.15618 | 0.002169 | 0.14809 | 0.14809 | 0.043347
2019-03-21 12:05 | 1553184297.3964 | 1184 | 0.20912 | 0.000733 | 0.000117 | 100.5699 | 100.6915 | 118.3435 | 21.1295 | 0.0026 | 0.002413 | 135.8244 | 0.26173 | 8.4031 | 0.003468 | 0.004857 | 0.00476 | 5.4e-05 | 0.00087 | 0.32421 | 1.2961 | 2.548 | 99.52 | 20.88 | 37.371 | 84.2348 | 105.1844 | 79.7339 | 0.009281 | 1.0005 | 0.023801 | 23.7903 | -0.01081 | 0.15571 | 0.002327 | 0.15939 | 0.15939 | 0.040149
2019-03-21 12:05 | 1553184298.3956 | 1185 | 0.20915 | 0.000736 | 0.000116 | 100.65 | 100.6844 | 117.9942 | 21.1286 | 0.002604 | 0.002422 | 135.7928 | 0.26113 | 8.4031 | 0.003462 | 0.004875 | 0.004931 | 5.2e-05 | 0.00087 | 0.27975 | 1.2776 | 2.5481 | 99.52 | 20.89 | 37.372 | 84.3441 | 105.0367 | 79.9654 | 0.009287 | 0.9998 | 0.023758 | 23.7629 | -0.003049 | 0.15535 | 0.002147 | 0.14626 | 0.14626 | 0.043654
2019-03-21 12:05 | 1553184299.3956 | 1186 | 0.20925 | 0.000737 | 0.000117 | 100.6292 | 100.8461 | 118.2222 | 21.1288 | 0.00259 | 0.002406 | 135.7688 | 0.26186 | 8.402 | 0.003413 | 0.004934 | 0.004817 | 5.2e-05 | 0.000868 | 0.32516 | 1.2751 | 2.5481 | 99.52 | 20.88 | 37.373 | 84.0057 | 103.8461 | 79.5702 | 0.009281 | 1.0009 | 0.023794 | 23.7726 | -0.019263 | 0.15578 | 0.001538 | 0.10293 | 0.10293 | 0.062202
2019-03-21 12:05 | 1553184300.3959 | 1187 | 0.20904 | 0.000734 | 0.000117 | 100.5937 | 100.7314 | 118.083 | 21.1298 | 0.002591 | 0.002443 | 135.7751 | 0.26191 | 8.4026 | 0.003412 | 0.005364 | 0.005176 | 5.2e-05 | 0.000869 | 0.32908 | 1.3135 | 2.548 | 99.52 | 20.88 | 37.373 | 84.0182 | 103.8236 | 80.4647 | 0.009281 | 0.99839 | 0.02375 | 23.7887 | -0.012229 | 0.15581 | 0.002814 | 0.19376 | 0.19376 | 0.03305
2019-03-21 12:05 | 1553184301.3971 | 1188 | 0.20923 | 0.000731 | 0.000116 | 100.7183 | 100.8877 | 118.2377 | 21.1291 | 0.002607 | 0.002426 | 135.7724 | 0.26137 | 8.4021 | 0.003443 | 0.005157 | 0.005089 | 5.1e-05 | 0.000869 | 0.32532 | 1.2824 | 2.5479 | 99.52 | 20.88 | 37.374 | 84.4124 | 104.5771 | 80.0603 | 0.009282 | 0.99953 | 0.02378 | 23.7906 | -0.015033 | 0.15549 | 0.001689 | 0.1137 | 0.1137 | 0.056208
2019-03-21 12:05 | 1553184302.3958 | 1189 | 0.20908 | 0.000733 | 0.000117 | 100.639 | 100.731 | 118.1567 | 21.1294 | 0.002589 | 0.002439 | 135.737 | 0.26233 | 8.4024 | 0.003431 | 0.0051 | 0.004971 | 5.4e-05 | 0.000871 | 0.30917 | 1.2815 | 2.5479 | 99.52 | 20.89 | 37.376 | 83.9851 | 104.2933 | 80.3701 | 0.009288 | 0.99866 | 0.023761 | 23.7929 | -0.008173 | 0.15606 | 0.002603 | 0.17878 | 0.17878 | 0.035878
2019-03-21 12:05 | 1553184303.3969 | 1190 | 0.20921 | 0.00073 | 0.000116 | 100.6682 | 100.7263 | 118.4484 | 21.1296 | 0.002556 | 0.002448 | 135.7672 | 0.26237 | 8.4033 | 0.00343 | 0.005565 | 0.005493 | 5.7e-05 | 0.000871 | 0.34058 | 1.3019 | 2.5478 | 99.52 | 20.89 | 37.378 | 83.1841 | 104.262 | 80.5746 | 0.009288 | 0.99808 | 0.023783 | 23.8292 | -0.005155 | 0.15608 | 0.001801 | 0.12175 | 0.12175 | 0.052692
2019-03-21 12:05 | 1553184304.3962 | 1191 | 0.20904 | 0.000737 | 0.000116 | 100.6822 | 100.6891 | 118.9017 | 21.1301 | 0.00261 | 0.002448 | 135.7543 | 0.26232 | 8.4012 | 0.003404 | 0.004589 | 0.004517 | 5.7e-05 | 0.000871 | 0.33339 | 1.3131 | 2.5479 | 99.52 | 20.89 | 37.377 | 84.4882 | 103.6408 | 80.59 | 0.009288 | 0.99804 | 0.023828 | 23.8752 | -0.000614 | 0.15605 | 0.002791 | 0.19281 | 0.19281 | 0.033264
2019-03-21 12:05 | 1553184305.396 | 1192 | 0.20917 | 0.000733 | 0.000116 | 100.6709 | 100.9453 | 119.0744 | 21.1288 | 0.002566 | 0.002394 | 135.7443 | 0.26106 | 8.3995 | 0.003353 | 0.005087 | 0.004987 | 5.3e-05 | 0.000871 | 0.30143 | 1.3309 | 2.5479 | 99.52 | 20.89 | 37.377 | 83.4236 | 102.387 | 79.2803 | 0.009288 | 1.0017 | 0.02389 | 23.8483 | -0.024343 | 0.15531 | 0.002014 | 0.13759 | 0.13759 | 0.046394
2019-03-21 12:05 | 1553184306.3969 | 1193 | 0.20916 | 0.000735 | 0.000116 | 100.7213 | 100.8993 | 119.4197 | 21.13 | 0.002599 | 0.002445 | 135.7414 | 0.26238 | 8.4042 | 0.00347 | 0.004712 | 0.004625 | 5.7e-05 | 0.000871 | 0.3147 | 1.3123 | 2.5479 | 99.52 | 20.89 | 37.376 | 84.2252 | 105.2431 | 80.5024 | 0.009288 | 0.99828 | 0.023883 | 23.9242 | -0.015795 | 0.15609 | 0.00212 | 0.14514 | 0.14514 | 0.044202
2019-03-21 12:05 | 1553184307.396 | 1194 | 0.20907 | 0.000731 | 0.000115 | 100.723 | 100.8758 | 119.7671 | 21.1291 | 0.002652 | 0.002422 | 135.733 | 0.26237 | 8.4014 | 0.003456 | 0.005217 | 0.005122 | 5.4e-05 | 0.00087 | 0.33268 | 1.3209 | 2.5478 | 99.52 | 20.89 | 37.377 | 85.4892 | 104.9001 | 79.9565 | 0.009288 | 0.99983 | 0.023936 | 23.9405 | -0.013553 | 0.15608 | 0.002646 | 0.18335 | 0.18335 | 0.034987
2019-03-21 12:05 | 1553184308.3954 | 1195 | 0.20911 | 0.000739 | 0.000115 | 100.8303 | 100.8853 | 119.9779 | 21.1308 | 0.002617 | 0.00245 | 135.7606 | 0.26195 | 8.403 | 0.003407 | 0.00527 | 0.005234 | 6e-05 | 0.000874 | 0.30663 | 1.347 | 2.5479 | 99.52 | 20.89 | 37.378 | 84.638 | 103.6981 | 80.6421 | 0.009288 | 0.99789 | 0.023934 | 23.9848 | -0.004878 | 0.15584 | 0.002372 | 0.16363 | 0.16363 | 0.039142
2019-03-21 12:05 | 1553184309.3962 | 1196 | 0.2092 | 0.000735 | 0.000114 | 100.8762 | 100.8203 | 119.7666 | 21.1291 | 0.002595 | 0.002459 | 135.7751 | 0.26197 | 8.402 | 0.003452 | 0.004865 | 0.004774 | 5.3e-05 | 0.000874 | 0.31205 | 1.3388 | 2.5479 | 99.52 | 20.89 | 37.378 | 84.1145 | 104.7894 | 80.8417 | 0.009288 | 0.99733 | 0.023906 | 23.9704 | 0.00496 | 0.15585 | 0.001878 | 0.128 | 0.128 | 0.050042
2019-03-21 12:05 | 1553184310.3953 | 1197 | 0.20912 | 0.000731 | 0.000114 | 100.8282 | 101.0389 | 120.3182 | 21.1295 | 0.002562 | 0.00241 | 135.7716 | 0.26184 | 8.4044 | 0.003434 | 0.005436 | 0.005293 | 5.6e-05 | 0.000876 | 0.31797 | 1.3189 | 2.548 | 99.52 | 20.89 | 37.378 | 83.3214 | 104.3564 | 79.6654 | 0.009288 | 1.0007 | 0.024001 | 23.9856 | -0.018672 | 0.15577 | 0.002329 | 0.16104 | 0.16104 | 0.039753
2019-03-21 12:05 | 1553184311.3956 | 1198 | 0.20913 | 0.000728 | 0.000115 | 100.7829 | 100.947 | 120.4901 | 21.13 | 0.00259 | 0.002414 | 135.7735 | 0.26128 | 8.4042 | 0.003475 | 0.004481 | 0.00443 | 5.2e-05 | 0.000876 | 0.31061 | 1.3134 | 2.548 | 99.52 | 20.89 | 37.382 | 83.9934 | 105.3471 | 79.7594 | 0.009289 | 1.0004 | 0.024015 | 24.0059 | -0.014546 | 0.15544 | 0.002261 | 0.15615 | 0.15615 | 0.040912
2019-03-21 12:05 | 1553184312.3963 | 1199 | 0.20912 | 0.000728 | 0.000115 | 100.7943 | 100.9204 | 120.6558 | 21.129 | 0.002557 | 0.002438 | 120.7461 | 0.26247 | 8.4014 | 0.003486 | 0.004889 | 0.004781 | 5.4e-05 | 0.000873 | 0.32781 | 1.3136 | 2.5479 | 99.52 | 20.89 | 37.382 | 83.1953 | 105.6158 | 80.333 | 0.009289 | 0.99876 | 0.024012 | 24.0419 | -0.011187 | 0.15615 | 0.002373 | 0.16428 | 0.16428 | 0.039065
2019-03-21 12:05 | 1553184313.3953 | 1200 | 0.20911 | 0.000735 | 0.000115 | 100.816 | 101.045 | 120.8025 | 21.1293 | 0.002597 | 0.002469 | 91.9625 | 0.26224 | 8.4011 | 0.003443 | 0.004352 | 0.004262 | 5.6e-05 | 0.000872 | 0.28479 | 1.2902 | 2.548 | 99.52 | 20.89 | 37.383 | 84.163 | 104.57 | 81.1007 | 0.00929 | 0.9966 | 0.024001 | 24.0827 | -0.020298 | 0.15601 | 0.002403 | 0.16637 | 0.16637 | 0.03854
2019-03-21 12:05 | 1553184314.3954 | 1201 | 0.20916 | 0.000728 | 0.000114 | 100.8538 | 101.0341 | 121.0273 | 21.1296 | 0.002612 | 0.002439 | 63.196 | 0.26242 | 8.4007 | 0.003365 | 0.004724 | 0.004578 | 5.1e-05 | 0.000877 | 0.33699 | 1.2901 | 2.5479 | 99.52 | 20.89 | 37.384 | 84.5339 | 102.6918 | 80.3749 | 0.00929 | 0.99864 | 0.024048 | 24.0804 | -0.015977 | 0.15612 | 0.002086 | 0.14382 | 0.14382 | 0.044614
2019-03-21 12:05 | 1553184315.397 | 1202 | 0.20917 | 0.00073 | 0.000113 | 100.8896 | 101.0386 | 121.416 | 21.1296 | 0.002575 | 0.002423 | 34.4124 | 0.26118 | 8.4012 | 0.003422 | 0.005211 | 0.00514 | 5.3e-05 | 0.000875 | 0.2808 | 1.2906 | 2.548 | 99.52 | 20.88 | 37.385 | 83.6388 | 104.0601 | 79.9726 | 0.009284 | 0.99978 | 0.0241 | 24.1053 | -0.013198 | 0.15538 | 0.002062 | 0.14249 | 0.14249 | 0.044816
2019-03-21 12:05 | 1553184316.4316 | 1203 | 0.20915 | 0.000731 | 0.000109 | 100.9936 | 101.0369 | 121.5773 | 21.1284 | 0.002577 | 0.002392 | 5.6075 | 0.26228 | 8.4043 | 0.003418 | 0.005268 | 0.005179 | 5e-05 | 0.000874 | 0.34844 | 1.3333 | 2.5479 | 99.52 | 20.88 | 37.39 | 83.6814 | 103.9788 | 79.2308 | 0.009286 | 1.0019 | 0.024141 | 24.0959 | -0.003834 | 0.15603 | 0.002173 | 0.15097 | 0.15097 | 0.042479
2019-03-21 12:05 | 1553184317.3763 | 1204 | 0.20921 | 0.000725 | 0.000103 | 100.9259 | 101.1585 | 121.7306 | 21.1285 | 0.002568 | 0.002397 | -23.1938 | 0.24987 | 8.4012 | 0.003417 | 0.005022 | 0.004911 | 5.3e-05 | 0.000869 | 0.32045 | 1.3266 | 2.5479 | 99.52 | 20.88 | 37.392 | 83.4677 | 103.9372 | 79.347 | 0.009286 | 1.0016 | 0.024153 | 24.1151 | -0.020583 | 0.14865 | 0.001848 | 0.12785 | 0.12785 | 0.047785
2019-03-21 12:05 | 1553184318.3778 | 1205 | 0.20921 | 0.00073 | 9.7e-05 | 101.001 | 100.994 | 122.1402 | 21.1285 | 0.002593 | 0.002448 | -51.9944 | 0.23912 | 8.3999 | 0.003388 | 0.004676 | 0.00453 | 5.4e-05 | 0.000867 | 0.3063 | 1.3541 | 2.5479 | 99.52 | 20.88 | 37.396 | 84.0614 | 103.2526 | 80.5812 | 0.009287 | 0.99806 | 0.024151 | 24.1979 | 0.000626 | 0.14225 | 0.00183 | 0.12686 | 0.12686 | 0.046089
2019-03-21 12:05 | 1553184319.3765 | 1206 | 0.20925 | 0.00073 | 9.2e-05 | 101.0058 | 101.1534 | 122.6704 | 21.1285 | 0.002544 | 0.00237 | -80.7969 | 0.23983 | 8.4015 | 0.003391 | 0.003845 | 0.003656 | 5.3e-05 | 0.00087 | 0.32587 | 1.3395 | 2.5479 | 99.52 | 20.88 | 37.397 | 82.8953 | 103.3154 | 78.7079 | 0.009287 | 1.0034 | 0.024268 | 24.186 | -0.013058 | 0.14267 | 0.001608 | 0.11163 | 0.11163 | 0.052531
2019-03-21 12:05 | 1553184320.3766 | 1207 | 0.20927 | 0.000725 | 8.6e-05 | 101.0013 | 101.01 | 122.5561 | 21.1283 | 0.002569 | 0.002404 | -109.5797 | 0.23915 | 8.4016 | 0.003461 | 0.004563 | 0.004384 | 5.2e-05 | 0.000867 | 0.31758 | 1.3247 | 2.5478 | 99.52 | 20.88 | 37.398 | 83.4933 | 105.0133 | 79.5347 | 0.009288 | 1.001 | 0.024228 | 24.2032 | -0.000765 | 0.14227 | 0.001464 | 0.10131 | 0.10131 | 0.057719
2019-03-21 12:05 | 1553184321.3767 | 1208 | 0.20935 | 0.000721 | 8.1e-05 | 101.0597 | 101.0018 | 122.846 | 21.1292 | 0.002491 | 0.002341 | -138.3653 | 0.23856 | 8.4026 | 0.003462 | 0.004589 | 0.004409 | 5.6e-05 | 0.000869 | 0.3501 | 1.3379 | 2.5479 | 99.52 | 20.89 | 37.4 | 81.6161 | 105.0316 | 78.0131 | 0.009294 | 1.0054 | 0.024309 | 24.1794 | 0.005129 | 0.14192 | 0.001045 | 0.071414 | 0.071414 | 0.081678
2019-03-21 12:05 | 1553184322.3766 | 1209 | 0.20929 | 0.000718 | 7.5e-05 | 101.0261 | 101.2123 | 123.1654 | 21.1289 | 0.002559 | 0.002375 | -152.1187 | 0.23932 | 8.4026 | 0.003389 | 0.00524 | 0.005094 | 5.3e-05 | 0.000866 | 0.30574 | 1.3084 | 2.5481 | 99.52 | 20.89 | 37.4 | 83.2552 | 103.271 | 78.8315 | 0.009294 | 1.003 | 0.024312 | 24.2391 | -0.016472 | 0.14237 | 0.001367 | 0.095283 | 0.095283 | 0.06141
2019-03-21 12:05 | 1553184323.3777 | 1210 | 0.20939 | 0.000711 | 6.8e-05 | 101.102 | 101.3073 | 123.3185 | 21.1287 | 0.002494 | 0.00234 | -152.1404 | 0.23924 | 8.4029 | 0.003362 | 0.005073 | 0.004938 | 5.3e-05 | 0.000864 | 0.34794 | 1.2961 | 2.5479 | 99.52 | 20.89 | 37.401 | 81.6858 | 102.6084 | 77.9768 | 0.009294 | 1.0055 | 0.024357 | 24.2246 | -0.018142 | 0.14232 | 0.000798 | 0.054206 | 0.054206 | 0.10791
2019-03-21 12:05 | 1553184324.3768 | 1211 | 0.20936 | 0.000704 | 6.1e-05 | 101.0952 | 101.0849 | 123.741 | 21.1287 | 0.002562 | 0.002315 | -152.1635 | 0.23941 | 8.4049 | 0.003502 | 0.004481 | 0.004312 | 5.4e-05 | 0.000864 | 0.34363 | 1.3158 | 2.5479 | 99.52 | 20.89 | 37.401 | 83.3167 | 106.0028 | 77.3774 | 0.009294 | 1.0072 | 0.02442 | 24.2454 | 0.00091 | 0.14243 | 0.001 | 0.069598 | 0.069598 | 0.084108
2019-03-21 12:05 | 1553184325.3772 | 1212 | 0.20943 | 0.000692 | 5.4e-05 | 101.1583 | 101.3645 | 124.1564 | 21.1282 | 0.002516 | 0.00227 | -152.1612 | 0.23959 | 8.4011 | 0.003493 | 0.004623 | 0.004413 | 5.3e-05 | 0.000859 | 0.34396 | 1.2844 | 2.5479 | 99.52 | 20.89 | 37.403 | 82.2118 | 105.7811 | 76.3135 | 0.009295 | 1.0103 | 0.024498 | 24.2492 | -0.018212 | 0.14253 | 0.000588 | 0.040006 | 0.040006 | 0.14643
2019-03-21 12:05 | 1553184326.3758 | 1213 | 0.20941 | 0.000681 | 4.7e-05 | 101.0734 | 101.2269 | 124.3761 | 21.129 | 0.002493 | 0.002345 | -152.1585 | 0.23949 | 8.4032 | 0.003507 | 0.004996 | 0.004867 | 5.6e-05 | 0.000856 | 0.32659 | 1.2816 | 2.5478 | 99.52 | 20.89 | 37.403 | 81.6664 | 106.1196 | 78.104 | 0.009295 | 1.0051 | 0.024457 | 24.3327 | -0.013573 | 0.14247 | 0.000724 | 0.050335 | 0.050335 | 0.11633
2019-03-21 12:05 | 1553184327.376 | 1214 | 0.20946 | 0.000672 | 4.1e-05 | 101.1181 | 101.432 | 124.3808 | 21.1289 | 0.002458 | 0.002315 | -152.1279 | 0.2403 | 8.4036 | 0.003454 | 0.004717 | 0.00463 | 5.5e-05 | 0.000856 | 0.33212 | 1.2879 | 2.5478 | 99.52 | 20.89 | 37.403 | 80.8196 | 104.8552 | 77.387 | 0.009295 | 1.0072 | 0.024482 | 24.3083 | -0.027718 | 0.14295 | 0.000476 | 0.032491 | 0.032491 | 0.18083
2019-03-21 12:05 | 1553184328.3765 | 1215 | 0.20948 | 0.000664 | 3.4e-05 | 101.1589 | 101.2218 | 124.6029 | 21.1289 | 0.0025 | 0.00229 | -152.1239 | 0.23971 | 8.4025 | 0.003493 | 0.004205 | 0.004116 | 5.3e-05 | 0.000858 | 0.31957 | 1.28 | 2.548 | 99.52 | 20.89 | 37.403 | 81.8216 | 105.799 | 76.7796 | 0.009295 | 1.0089 | 0.024525 | 24.3089 | -0.005558 | 0.1426 | 0.000349 | 0.023622 | 0.023622 | 0.24812
2019-03-21 12:05 | 1553184329.3761 | 1216 | 0.20949 | 0.000656 | 2.9e-05 | 101.178 | 101.3962 | 124.6705 | 21.1299 | 0.002442 | 0.002253 | -152.1155 | 0.23905 | 8.4035 | 0.003452 | 0.004913 | 0.004643 | 5.1e-05 | 0.000858 | 0.32604 | 1.2989 | 2.548 | 99.51 | 20.89 | 37.401 | 80.4384 | 104.7986 | 75.8873 | 0.009295 | 1.0115 | 0.024563 | 24.2845 | -0.019268 | 0.14221 | 0.000314 | 0.021404 | 0.021404 | 0.27309
2019-03-21 12:05 | 1553184330.3774 | 1217 | 0.20949 | 0.000643 | 2.4e-05 | 101.2358 | 101.2995 | 124.9752 | 21.1287 | 0.002421 | 0.002266 | -152.1239 | 0.23909 | 8.3996 | 0.00339 | 0.004773 | 0.004657 | 5.5e-05 | 0.000863 | 0.31332 | 1.3014 | 2.5479 | 99.51 | 20.89 | 37.4 | 79.9241 | 103.2952 | 76.2049 | 0.009295 | 1.0106 | 0.024582 | 24.3252 | -0.005623 | 0.14223 | 0.000369 | 0.025762 | 0.025762 | 0.22691
2019-03-21 12:06 | 1553184331.377 | 1218 | 0.20949 | 0.000634 | 2e-05 | 101.2322 | 101.4074 | 125.0012 | 21.1289 | 0.002442 | 0.002256 | -152.1077 | 0.23984 | 8.4007 | 0.003398 | 0.004509 | 0.004277 | 5.5e-05 | 0.000859 | 0.31957 | 1.3027 | 2.548 | 99.51 | 20.89 | 37.398 | 80.4499 | 103.4957 | 75.9627 | 0.009294 | 1.0113 | 0.024593 | 24.3193 | -0.015467 | 0.14268 | 0.000387 | 0.027333 | 0.027333 | 0.21454
2019-03-21 12:06 | 1553184332.3762 | 1219 | 0.20956 | 0.000618 | 1.6e-05 | 101.2006 | 101.388 | 124.9843 | 21.1291 | 0.002415 | 0.002257 | -152.1141 | 0.23901 | 8.4023 | 0.003449 | 0.00478 | 0.004645 | 5.4e-05 | 0.000859 | 0.32333 | 1.3215 | 2.5478 | 99.51 | 20.89 | 37.398 | 79.7967 | 104.7193 | 76.0026 | 0.009294 | 1.0111 | 0.02459 | 24.3191 | -0.016545 | 0.14219 | -2.9e-05 | -0.00316 | -0.00316 | -1.8493
2019-03-21 12:06 | 1553184333.3776 | 1220 | 0.20956 | 0.000609 | 1.3e-05 | 101.29 | 101.3383 | 124.9773 | 21.1284 | 0.002471 | 0.002256 | -152.1018 | 0.23911 | 8.4008 | 0.003386 | 0.004993 | 0.004681 | 5.1e-05 | 0.000863 | 0.31365 | 1.3339 | 2.5479 | 99.52 | 20.89 | 37.396 | 81.1395 | 103.1997 | 75.9803 | 0.009293 | 1.0112 | 0.02459 | 24.3176 | -0.004261 | 0.14225 | 1e-06 | -0.000751 | -0.000751 | -7.7857
2019-03-21 12:06 | 1553184334.3766 | 1221 | 0.20955 | 0.000605 | 1.1e-05 | 101.3031 | 101.3461 | 124.844 | 21.1276 | 0.002415 | 0.002204 | -152.088 | 0.24019 | 8.4024 | 0.003409 | 0.005092 | 0.00501 | 5.1e-05 | 0.00086 | 0.31061 | 1.334 | 2.5478 | 99.52 | 20.88 | 37.395 | 79.7976 | 103.7493 | 74.7275 | 0.009287 | 1.0149 | 0.024621 | 24.261 | -0.003795 | 0.14289 | 3e-05 | 0.001545 | 0.001545 | 3.8003
2019-03-21 12:06 | 1553184335.3763 | 1222 | 0.20955 | 0.000593 | 8e-06 | 101.2582 | 101.5815 | 124.4757 | 21.1272 | 0.002419 | 0.002252 | -152.0875 | 0.23923 | 8.4011 | 0.003424 | 0.004485 | 0.004263 | 5.1e-05 | 0.000858 | 0.32101 | 1.3375 | 2.5479 | 99.52 | 20.88 | 37.393 | 79.8886 | 104.1269 | 75.8614 | 0.009286 | 1.0116 | 0.024545 | 24.2646 | -0.028515 | 0.14232 | 4.1e-05 | 0.002531 | 0.002531 | 2.3114
2019-03-21 12:06 | 1553184336.3763 | 1223 | 0.20954 | 0.000582 | 6e-06 | 101.2343 | 101.3213 | 124.3818 | 21.1269 | 0.002389 | 0.002205 | -152.0825 | 0.23939 | 8.402 | 0.003449 | 0.004757 | 0.004622 | 5.1e-05 | 0.000857 | 0.33251 | 1.371 | 2.548 | 99.52 | 20.88 | 37.392 | 79.1631 | 104.7141 | 74.7519 | 0.009286 | 1.0148 | 0.024575 | 24.2169 | -0.007679 | 0.14241 | 0.00013 | 0.009228 | 0.009228 | 0.63427
2019-03-21 12:06 | 1553184337.3777 | 1224 | 0.20957 | 0.000577 | 4e-06 | 101.2951 | 101.4953 | 124.5229 | 21.1272 | 0.002392 | 0.002254 | -152.0903 | 0.24022 | 8.4019 | 0.003383 | 0.004675 | 0.004489 | 5.3e-05 | 0.000865 | 0.29264 | 1.351 | 2.548 | 99.51 | 20.88 | 37.388 | 79.2249 | 103.1351 | 75.9128 | 0.009286 | 1.0114 | 0.024548 | 24.271 | -0.017661 | 0.14291 | -4.7e-05 | -0.003712 | -0.003712 | -1.5825
2019-03-21 12:06 | 1553184338.3775 | 1225 | 0.20959 | 0.000567 | 3e-06 | 101.2567 | 101.3921 | 124.2764 | 21.1281 | 0.002373 | 0.002218 | -152.0974 | 0.24023 | 8.4039 | 0.003503 | 0.004537 | 0.004404 | 5.1e-05 | 0.000866 | 0.32477 | 1.3178 | 2.5478 | 99.51 | 20.88 | 37.387 | 78.776 | 106.0258 | 75.0563 | 0.009286 | 1.0139 | 0.024554 | 24.2172 | -0.011957 | 0.14291 | -0.000148 | -0.011081 | -0.011081 | -0.53007
2019-03-21 12:06 | 1553184339.3765 | 1226 | 0.20957 | 0.000556 | 2e-06 | 101.2564 | 101.4718 | 124.2672 | 21.1268 | 0.002408 | 0.002164 | -152.0996 | 0.23958 | 8.3986 | 0.003414 | 0.005091 | 0.004871 | 5.2e-05 | 0.000865 | 0.33323 | 1.3165 | 2.548 | 99.51 | 20.88 | 37.383 | 79.6096 | 103.8795 | 73.7539 | 0.009285 | 1.0177 | 0.024599 | 24.171 | -0.019011 | 0.14253 | -1.1e-05 | -0.000963 | -0.000963 | -6.0856
2019-03-21 12:06 | 1553184340.377 | 1227 | 0.20954 | 0.000557 | 2e-06 | 101.3219 | 101.5106 | 124.0468 | 21.1279 | 0.002399 | 0.002213 | -152.1049 | 0.24018 | 8.4016 | 0.003452 | 0.005404 | 0.00525 | 5.2e-05 | 0.000871 | 0.33467 | 1.3217 | 2.548 | 99.51 | 20.88 | 37.382 | 79.4113 | 104.7984 | 74.9438 | 0.009284 | 1.0142 | 0.024535 | 24.1909 | -0.016644 | 0.14288 | 0.000171 | 0.01253 | 0.01253 | 0.46867
2019-03-21 12:06 | 1553184341.3757 | 1228 | 0.20956 | 0.000551 | 0 | 101.2611 | 101.4136 | 123.9029 | 21.1276 | 0.00234 | 0.002182 | -152.1212 | 0.23983 | 8.4015 | 0.003392 | 0.004504 | 0.004321 | 5.1e-05 | 0.000872 | 0.34235 | 1.3191 | 2.548 | 99.51 | 20.88 | 37.38 | 77.9928 | 103.353 | 74.1795 | 0.009284 | 1.0165 | 0.024548 | 24.1503 | -0.013454 | 0.14267 | 6.1e-05 | 0.004516 | 0.004516 | 1.2986
2019-03-21 12:06 | 1553184342.3767 | 1229 | 0.20956 | 0.000543 | 0 | 101.3502 | 101.357 | 123.7781 | 21.1274 | 0.002368 | 0.002181 | -152.0999 | 0.23995 | 8.4004 | 0.00339 | 0.004192 | 0.004034 | 5.4e-05 | 0.000875 | 0.30486 | 1.3148 | 2.548 | 99.51 | 20.88 | 37.379 | 78.6578 | 103.2915 | 74.1592 | 0.009284 | 1.0165 | 0.024536 | 24.1374 | -0.000601 | 0.14274 | 8.2e-05 | 0.006028 | 0.006028 | 0.97332
2019-03-21 12:06 | 1553184343.3761 | 1230 | 0.20954 | 0.000531 | -1e-06 | 101.2641 | 101.3798 | 123.5774 | 21.1279 | 0.002381 | 0.002148 | -152.1164 | 0.23872 | 8.4049 | 0.003382 | 0.005368 | 0.005255 | 5.5e-05 | 0.00088 | 0.31996 | 1.2645 | 2.548 | 99.51 | 20.88 | 37.378 | 78.9792 | 103.1081 | 73.3754 | 0.009283 | 1.0188 | 0.024544 | 24.0906 | -0.010217 | 0.14201 | 0.000227 | 0.016805 | 0.016805 | 0.34733
2019-03-21 12:06 | 1553184344.3758 | 1231 | 0.20958 | 0.00053 | -1e-06 | 101.2885 | 101.5369 | 123.2683 | 21.1279 | 0.002325 | 0.002176 | -152.1223 | 0.23963 | 8.4021 | 0.003393 | 0.004614 | 0.004376 | 5.2e-05 | 0.000886 | 0.31006 | 1.3291 | 2.548 | 99.51 | 20.88 | 37.377 | 77.619 | 103.3564 | 74.045 | 0.009283 | 1.0169 | 0.02449 | 24.0837 | -0.021901 | 0.14255 | -4e-05 | -0.002829 | -0.002829 | -2.0708
2019-03-21 12:06 | 1553184345.3758 | 1232 | 0.20959 | 0.000528 | -1e-06 | 101.2817 | 101.5147 | 123.2213 | 21.1265 | 0.002287 | 0.00215 | -152.1253 | 0.23975 | 8.4028 | 0.0034 | 0.004878 | 0.004746 | 5.8e-05 | 0.00089 | 0.3194 | 1.3128 | 2.5481 | 99.51 | 20.88 | 37.377 | 76.7025 | 103.5265 | 73.4161 | 0.009283 | 1.0187 | 0.024507 | 24.0573 | -0.020556 | 0.14262 | -5.2e-05 | -0.003753 | -0.003753 | -1.5618
2019-03-21 12:06 | 1553184346.3773 | 1233 | 0.20957 | 0.000521 | -2e-06 | 101.279 | 101.4156 | 123.0321 | 21.1279 | 0.002286 | 0.00219 | -152.1272 | 0.23943 | 8.4016 | 0.003422 | 0.004293 | 0.004076 | 5.6e-05 | 0.000887 | 0.35187 | 1.3394 | 2.548 | 99.51 | 20.89 | 37.377 | 76.6813 | 104.075 | 74.3896 | 0.009289 | 1.0158 | 0.024454 | 24.0726 | -0.012056 | 0.14244 | 6.3e-05 | 0.004726 | 0.004726 | 1.2386
2019-03-21 12:06 | 1553184347.377 | 1234 | 0.2096 | 0.000514 | -2e-06 | 101.3088 | 101.3717 | 122.7739 | 21.1265 | 0.002287 | 0.00214 | -152.1244 | 0.23897 | 8.403 | 0.003385 | 0.004801 | 0.004612 | 5.2e-05 | 0.000882 | 0.33085 | 1.3582 | 2.548 | 99.51 | 20.89 | 37.379 | 76.722 | 103.1775 | 73.1738 | 0.009289 | 1.0194 | 0.024471 | 24.0052 | -0.005549 | 0.14216 | -0.000126 | -0.009131 | -0.009131 | -0.63987
2019-03-21 12:06 | 1553184348.376 | 1235 | 0.20955 | 0.000513 | -2e-06 | 101.2284 | 101.2831 | 122.9077 | 21.1281 | 0.00237 | 0.002121 | -152.1192 | 0.23914 | 8.4023 | 0.003374 | 0.004725 | 0.004448 | 5.6e-05 | 0.000881 | 0.31614 | 1.3341 | 2.548 | 99.51 | 20.89 | 37.38 | 78.712 | 102.9042 | 72.7223 | 0.00929 | 1.0207 | 0.0245 | 24.0026 | -0.004837 | 0.14226 | 0.000165 | 0.012267 | 0.012267 | 0.47665
2019-03-21 12:06 | 1553184349.3775 | 1236 | 0.20956 | 0.000513 | -3e-06 | 101.3431 | 101.4125 | 122.8507 | 21.1275 | 0.002283 | 0.002137 | -152.1318 | 0.23911 | 8.3993 | 0.003397 | 0.004576 | 0.004405 | 5.4e-05 | 0.000885 | 0.34092 | 1.3295 | 2.5479 | 99.51 | 20.89 | 37.381 | 76.6184 | 103.4621 | 73.117 | 0.00929 | 1.0196 | 0.024481 | 24.0107 | -0.00613 | 0.14225 | 0.000101 | 0.007593 | 0.007593 | 0.76999
2019-03-21 12:06 | 1553184350.3771 | 1237 | 0.20957 | 0.000507 | -3e-06 | 101.29 | 101.5277 | 122.7436 | 21.1287 | 0.002334 | 0.002101 | -152.1068 | 0.23861 | 8.4022 | 0.00342 | 0.004613 | 0.00445 | 5.6e-05 | 0.000882 | 0.3115 | 1.3672 | 2.548 | 99.51 | 20.89 | 37.382 | 77.8449 | 104.0166 | 72.2474 | 0.00929 | 1.0221 | 0.024501 | 23.9701 | -0.02096 | 0.14195 | 7e-05 | 0.005337 | 0.005337 | 1.0931
2019-03-21 12:06 | 1553184351.3762 | 1238 | 0.2096 | 0.000508 | -3e-06 | 101.2292 | 101.5416 | 122.7664 | 21.1279 | 0.002298 | 0.002112 | -152.1037 | 0.24038 | 8.4018 | 0.003427 | 0.00474 | 0.004538 | 5.6e-05 | 0.000888 | 0.33732 | 1.352 | 2.5479 | 99.51 | 20.89 | 37.383 | 76.9766 | 104.1811 | 72.511 | 0.00929 | 1.0214 | 0.024494 | 23.9815 | -0.027558 | 0.143 | -8.8e-05 | -0.0063 | -0.0063 | -0.93293
2019-03-21 12:06 | 1553184352.3766 | 1239 | 0.20958 | 0.000503 | -2e-06 | 101.2871 | 101.4888 | 122.8457 | 21.1277 | 0.00229 | 0.002123 | -152.1186 | 0.23945 | 8.4027 | 0.003391 | 0.004304 | 0.004143 | 5.4e-05 | 0.00089 | 0.32565 | 1.3563 | 2.5479 | 99.51 | 20.89 | 37.383 | 76.7878 | 103.3124 | 72.7773 | 0.00929 | 1.0206 | 0.024492 | 23.9985 | -0.017801 | 0.14245 | -1e-06 | 7.3e-05 | 7.3e-05 | 80.4554
2019-03-21 12:06 | 1553184353.3764 | 1240 | 0.20956 | 0.000497 | -3e-06 | 101.3076 | 101.5031 | 123.099 | 21.1287 | 0.002293 | 0.002116 | -152.0992 | 0.23992 | 8.4019 | 0.003366 | 0.004881 | 0.004744 | 5.4e-05 | 0.000892 | 0.35297 | 1.3464 | 2.5481 | 99.51 | 20.89 | 37.384 | 76.8575 | 102.7043 | 72.6098 | 0.009291 | 1.0211 | 0.024523 | 24.0174 | -0.017247 | 0.14273 | 0.000175 | 0.013079 | 0.013079 | 0.44853
2019-03-21 12:06 | 1553184354.3763 | 1241 | 0.20959 | 0.000502 | -3e-06 | 101.2543 | 101.4292 | 123.337 | 21.1279 | 0.002232 | 0.002129 | -152.1111 | 0.23947 | 8.402 | 0.003365 | 0.004626 | 0.004402 | 5.5e-05 | 0.000894 | 0.34827 | 1.3447 | 2.548 | 99.51 | 20.89 | 37.385 | 75.3974 | 102.6895 | 72.9194 | 0.009291 | 1.0202 | 0.024536 | 24.0514 | -0.015441 | 0.14246 | -4.2e-05 | -0.002932 | -0.002932 | -1.9971
2019-03-21 12:06 | 1553184355.3763 | 1242 | 0.20956 | 0.000493 | -3e-06 | 101.2829 | 101.5099 | 123.6093 | 21.1288 | 0.002251 | 0.002107 | -152.1184 | 0.23873 | 8.403 | 0.003442 | 0.005304 | 0.005166 | 5.3e-05 | 0.000892 | 0.33644 | 1.2912 | 2.548 | 99.51 | 20.89 | 37.385 | 75.8464 | 104.5618 | 72.3973 | 0.009291 | 1.0217 | 0.024582 | 24.0597 | -0.020031 | 0.14202 | 0.000141 | 0.010593 | 0.010593 | 0.551
2019-03-21 12:06 | 1553184356.3764 | 1243 | 0.20957 | 0.000496 | -3e-06 | 101.3148 | 101.5495 | 123.6795 | 21.1289 | 0.00231 | 0.002108 | -152.1199 | 0.23981 | 8.4037 | 0.003408 | 0.004512 | 0.004382 | 5.5e-05 | 0.000898 | 0.32892 | 1.3033 | 2.5479 | 99.51 | 20.89 | 37.384 | 77.2582 | 103.7411 | 72.4161 | 0.009291 | 1.0216 | 0.024588 | 24.0672 | -0.020696 | 0.14266 | 0.000106 | 0.008058 | 0.008058 | 0.72762
2019-03-21 12:06 | 1553184357.3759 | 1244 | 0.2096 | 0.000494 | -3e-06 | 101.2385 | 101.3931 | 124.0403 | 21.1278 | 0.002279 | 0.002058 | -152.1329 | 0.23881 | 8.4019 | 0.003396 | 0.00479 | 0.004589 | 5.7e-05 | 0.000895 | 0.30901 | 1.3065 | 2.548 | 99.51 | 20.89 | 37.382 | 76.5337 | 103.4484 | 71.223 | 0.00929 | 1.0252 | 0.024667 | 24.0606 | -0.013652 | 0.14207 | -7.8e-05 | -0.00559 | -0.00559 | -1.0446
2019-03-21 12:06 | 1553184358.3765 | 1245 | 0.2096 | 0.000492 | -3e-06 | 101.3168 | 101.5525 | 123.9196 | 21.129 | 0.002264 | 0.002076 | -152.1369 | 0.23938 | 8.4023 | 0.00335 | 0.00445 | 0.004303 | 6e-05 | 0.000897 | 0.34562 | 1.3182 | 2.548 | 99.51 | 20.89 | 37.382 | 76.1663 | 102.3346 | 71.6466 | 0.00929 | 1.0239 | 0.024639 | 24.0637 | -0.020782 | 0.14241 | -5.9e-05 | -0.004186 | -0.004186 | -1.3981
2019-03-21 12:06 | 1553184359.3755 | 1246 | 0.2096 | 0.000488 | -3e-06 | 101.2438 | 101.3972 | 124.0891 | 21.1278 | 0.002207 | 0.002114 | -152.1335 | 0.23934 | 8.4042 | 0.00341 | 0.005528 | 0.005484 | 5.8e-05 | 0.000899 | 0.34219 | 1.3338 | 2.548 | 99.51 | 20.89 | 37.381 | 74.7994 | 103.7902 | 72.5491 | 0.00929 | 1.0213 | 0.024624 | 24.1117 | -0.013539 | 0.14238 | -4.8e-05 | -0.003331 | -0.003331 | -1.7567
2019-03-21 12:06 | 1553184360.3771 | 1247 | 0.20959 | 0.000491 | -3e-06 | 101.2975 | 101.404 | 124.1921 | 21.1277 | 0.002269 | 0.00212 | -152.1494 | 0.23992 | 8.402 | 0.003424 | 0.003754 | 0.003457 | 5.8e-05 | 0.000901 | 0.32388 | 1.3133 | 2.5481 | 99.51 | 20.89 | 37.377 | 76.2859 | 104.1094 | 72.7033 | 0.009289 | 1.0208 | 0.024629 | 24.1271 | -0.009403 | 0.14273 | -4.7e-05 | -0.003257 | -0.003257 | -1.8012
2019-03-21 12:06 | 1553184361.3767 | 1248 | 0.20964 | 0.000488 | -3e-06 | 101.2736 | 101.3219 | 124.0593 | 21.1281 | 0.002212 | 0.002053 | -152.1431 | 0.23894 | 8.4033 | 0.003445 | 0.004652 | 0.004504 | 6e-05 | 0.000902 | 0.33892 | 1.3264 | 2.548 | 99.51 | 20.89 | 37.376 | 74.9229 | 104.6193 | 71.1024 | 0.009289 | 1.0255 | 0.024673 | 24.0583 | -0.004265 | 0.14214 | -0.000317 | -0.023301 | -0.023301 | -0.25072
2019-03-21 12:06 | 1553184362.3757 | 1249 | 0.20959 | 0.000482 | -3e-06 | 101.2972 | 101.4439 | 123.986 | 21.1279 | 0.002199 | 0.002115 | -152.1296 | 0.24057 | 8.4024 | 0.003449 | 0.004573 | 0.004348 | 5.7e-05 | 0.000903 | 0.32621 | 1.3334 | 2.548 | 99.51 | 20.89 | 37.374 | 74.6025 | 104.7284 | 72.5725 | 0.009288 | 1.0212 | 0.024613 | 24.1025 | -0.012944 | 0.14311 | -1.9e-05 | -0.001184 | -0.001184 | -4.9671
2019-03-21 12:06 | 1553184363.3772 | 1250 | 0.20958 | 0.000489 | -3e-06 | 101.2573 | 101.4003 | 123.8501 | 21.1276 | 0.0023 | 0.002031 | -152.127 | 0.23995 | 8.4021 | 0.00344 | 0.004652 | 0.00446 | 5.7e-05 | 0.000907 | 0.3314 | 1.3395 | 2.548 | 99.51 | 20.89 | 37.374 | 77.027 | 104.5063 | 70.5638 | 0.009288 | 1.0271 | 0.024671 | 24.0192 | -0.012625 | 0.14274 | 4.6e-05 | 0.003611 | 0.003611 | 1.6246
2019-03-21 12:06 | 1553184364.3768 | 1251 | 0.20957 | 0.000484 | -4e-06 | 101.2668 | 101.3305 | 123.6361 | 21.1277 | 0.002262 | 0.002085 | -152.1265 | 0.23941 | 8.4027 | 0.003405 | 0.004163 | 0.004042 | 5.9e-05 | 0.000908 | 0.29574 | 1.3185 | 2.548 | 99.51 | 20.89 | 37.373 | 76.1043 | 103.6629 | 71.8544 | 0.009288 | 1.0233 | 0.024604 | 24.0434 | -0.005622 | 0.14243 | 9.6e-05 | 0.007302 | 0.007302 | 0.8017
2019-03-21 12:06 | 1553184365.3769 | 1252 | 0.20958 | 0.000482 | -3e-06 | 101.2859 | 101.2954 | 123.5473 | 21.1275 | 0.002182 | 0.002088 | -152.1221 | 0.23912 | 8.4022 | 0.003379 | 0.004584 | 0.004401 | 5.8e-05 | 0.000908 | 0.3178 | 1.3002 | 2.5478 | 99.51 | 20.89 | 37.37 | 74.1888 | 103.0324 | 71.9226 | 0.009287 | 1.0231 | 0.024593 | 24.0372 | -0.000841 | 0.14225 | 7.1e-05 | 0.005459 | 0.005459 | 1.071
2019-03-21 12:06 | 1553184366.376 | 1253 | 0.20958 | 0.000474 | -3e-06 | 101.2927 | 101.4422 | 123.5455 | 21.1275 | 0.002208 | 0.002032 | -152.1164 | 0.23839 | 8.4019 | 0.003361 | 0.004068 | 0.00387 | 5.6e-05 | 0.00091 | 0.31509 | 1.3179 | 2.5478 | 99.51 | 20.89 | 37.369 | 74.8209 | 102.5963 | 70.5918 | 0.009287 | 1.0271 | 0.02464 | 23.9906 | -0.013188 | 0.14182 | 5.2e-05 | 0.004061 | 0.004061 | 1.4354
2019-03-21 12:06 | 1553184367.377 | 1254 | 0.20957 | 0.000478 | -3e-06 | 101.2829 | 101.3941 | 123.3549 | 21.1278 | 0.002243 | 0.002057 | -152.1231 | 0.23935 | 8.4037 | 0.00335 | 0.004157 | 0.004079 | 5.6e-05 | 0.00091 | 0.31758 | 1.3328 | 2.548 | 99.51 | 20.89 | 37.367 | 75.6529 | 102.3373 | 71.1963 | 0.009286 | 1.0253 | 0.024599 | 23.9931 | -0.009821 | 0.14239 | 0.000103 | 0.007846 | 0.007846 | 0.74591
2019-03-21 12:06 | 1553184368.3773 | 1255 | 0.20959 | 0.000477 | -4e-06 | 101.2582 | 101.4599 | 123.5126 | 21.1269 | 0.002202 | 0.002019 | -152.1043 | 0.24017 | 8.4013 | 0.003433 | 0.004581 | 0.004355 | 5.6e-05 | 0.000912 | 0.34346 | 1.3209 | 2.5479 | 99.51 | 20.89 | 37.366 | 74.6591 | 104.3342 | 70.2817 | 0.009286 | 1.028 | 0.024648 | 23.9766 | -0.017802 | 0.14288 | 1.5e-05 | 0.001344 | 0.001344 | 4.369
2019-03-21 12:06 | 1553184369.3761 | 1256 | 0.20958 | 0.000479 | -3e-06 | 101.2522 | 101.5045 | 123.2952 | 21.1274 | 0.002206 | 0.002056 | -152.1058 | 0.23921 | 8.4023 | 0.003349 | 0.004828 | 0.004634 | 5.5e-05 | 0.000908 | 0.32947 | 1.314 | 2.5479 | 99.51 | 20.89 | 37.365 | 74.7632 | 102.304 | 71.1539 | 0.009286 | 1.0254 | 0.024595 | 23.9859 | -0.022261 | 0.14231 | 6.6e-05 | 0.005065 | 0.005065 | 1.1548
2019-03-21 12:06 | 1553184370.3772 | 1257 | 0.20958 | 0.000475 | -3e-06 | 101.3201 | 101.4272 | 123.6548 | 21.1275 | 0.002158 | 0.002108 | -152.0978 | 0.23863 | 8.4035 | 0.003431 | 0.004428 | 0.004263 | 5.6e-05 | 0.000913 | 0.31686 | 1.3492 | 2.548 | 99.51 | 20.89 | 37.364 | 73.6258 | 104.2851 | 72.413 | 0.009286 | 1.0217 | 0.024586 | 24.0647 | -0.009446 | 0.14196 | 9.2e-05 | 0.006945 | 0.006945 | 0.84008
2019-03-21 12:06 | 1553184371.3757 | 1258 | 0.20957 | 0.000475 | -3e-06 | 101.2173 | 101.4882 | 124.0025 | 21.1269 | 0.002151 | 0.002044 | -152.1077 | 0.24017 | 8.4003 | 0.003414 | 0.004273 | 0.004159 | 5.6e-05 | 0.00091 | 0.35187 | 1.3508 | 2.5479 | 99.51 | 20.89 | 37.363 | 73.4392 | 103.879 | 70.8674 | 0.009285 | 1.0262 | 0.024676 | 24.0446 | -0.0239 | 0.14288 | 0.000146 | 0.011003 | 0.011003 | 0.5337
2019-03-21 12:06 | 1553184372.3753 | 1259 | 0.2096 | 0.000476 | -4e-06 | 101.318 | 101.5334 | 124.1623 | 21.1282 | 0.002185 | 0.002005 | -152.1262 | 0.23933 | 8.4007 | 0.003358 | 0.004741 | 0.004502 | 5.6e-05 | 0.000911 | 0.29142 | 1.2928 | 2.5478 | 99.51 | 20.89 | 37.364 | 74.2621 | 102.5309 | 69.9376 | 0.009286 | 1.029 | 0.024725 | 24.0275 | -0.018997 | 0.14238 | -3.7e-05 | -0.00251 | -0.00251 | -2.331
2019-03-21 12:06 | 1553184373.3757 | 1260 | 0.20957 | 0.000474 | -4e-06 | 101.1822 | 101.6305 | 124.2855 | 21.1276 | 0.002149 | 0.002034 | -152.1337 | 0.23974 | 8.4019 | 0.003356 | 0.004975 | 0.004721 | 5.6e-05 | 0.000911 | 0.31614 | 1.3231 | 2.548 | 99.51 | 20.89 | 37.364 | 73.3884 | 102.4754 | 70.6239 | 0.009286 | 1.027 | 0.024712 | 24.0635 | -0.039546 | 0.14262 | 0.000155 | 0.011756 | 0.011756 | 0.49861
2019-03-21 12:06 | 1553184374.3767 | 1261 | 0.20963 | 0.000474 | -3e-06 | 101.3022 | 101.4813 | 124.4231 | 21.1272 | 0.002217 | 0.001998 | -152.1321 | 0.23869 | 8.4008 | 0.003272 | 0.004965 | 0.004797 | 5.7e-05 | 0.000909 | 0.3115 | 1.3707 | 2.5481 | 99.51 | 20.89 | 37.363 | 75.0296 | 100.4447 | 69.7752 | 0.009285 | 1.0295 | 0.024757 | 24.047 | -0.015799 | 0.14199 | -0.000257 | -0.01889 | -0.01889 | -0.30895
2019-03-21 12:06 | 1553184375.377 | 1262 | 0.20961 | 0.000472 | -4e-06 | 101.2927 | 101.42 | 124.3327 | 21.1271 | 0.002164 | 0.002031 | -152.1318 | 0.23915 | 8.4028 | 0.003347 | 0.005593 | 0.005361 | 5.7e-05 | 0.00091 | 0.29839 | 1.3289 | 2.548 | 99.51 | 20.89 | 37.362 | 73.7464 | 102.2456 | 70.5518 | 0.009285 | 1.0272 | 0.02472 | 24.0655 | -0.011236 | 0.14227 | -0.000138 | -0.010021 | -0.010021 | -0.58352
2019-03-21 12:06 | 1553184376.3759 | 1263 | 0.20963 | 0.000471 | -3e-06 | 101.2462 | 101.6401 | 124.3445 | 21.1266 | 0.002186 | 0.002032 | -152.1357 | 0.23911 | 8.4035 | 0.003382 | 0.00417 | 0.003972 | 5.5e-05 | 0.000905 | 0.3147 | 1.3091 | 2.548 | 99.51 | 20.89 | 37.361 | 74.2939 | 103.0951 | 70.591 | 0.009285 | 1.0271 | 0.024719 | 24.068 | -0.034724 | 0.14225 | -0.000243 | -0.017839 | -0.017839 | -0.32773
2019-03-21 12:06 | 1553184377.3759 | 1264 | 0.20962 | 0.000466 | -3e-06 | 101.2716 | 101.6871 | 124.1107 | 21.1265 | 0.002167 | 0.002052 | -152.1309 | 0.23891 | 8.4042 | 0.003294 | 0.00444 | 0.004312 | 5.7e-05 | 0.000905 | 0.32477 | 1.339 | 2.548 | 99.51 | 20.89 | 37.36 | 73.8277 | 100.9674 | 71.0758 | 0.009285 | 1.0256 | 0.024679 | 24.0623 | -0.036623 | 0.14213 | -0.000151 | -0.010977 | -0.010977 | -0.53217
2019-03-21 12:06 | 1553184378.3765 | 1265 | 0.20961 | 0.000469 | -3e-06 | 101.2835 | 101.2589 | 123.7924 | 21.1276 | 0.002228 | 0.002026 | -152.158 | 0.23879 | 8.4031 | 0.003293 | 0.004256 | 0.004034 | 5.5e-05 | 0.000905 | 0.30182 | 1.306 | 2.5479 | 99.51 | 20.89 | 37.359 | 75.2912 | 100.9528 | 70.4461 | 0.009284 | 1.0275 | 0.02467 | 24.0095 | 0.002167 | 0.14205 | -9.9e-05 | -0.007124 | -0.007124 | -0.81953
2019-03-21 12:06 | 1553184379.3771 | 1266 | 0.20961 | 0.000473 | -3e-06 | 101.2811 | 101.4432 | 123.7492 | 21.1274 | 0.002157 | 0.002048 | -152.1503 | 0.23972 | 8.4017 | 0.003316 | 0.004538 | 0.004398 | 5.8e-05 | 0.000906 | 0.32084 | 1.3105 | 2.5481 | 99.51 | 20.89 | 37.358 | 73.59 | 101.5021 | 70.9738 | 0.009284 | 1.0259 | 0.024647 | 24.0237 | -0.014304 | 0.14261 | -0.000117 | -0.008432 | -0.008432 | -0.69511
2019-03-21 12:06 | 1553184380.3772 | 1267 | 0.20959 | 0.000473 | -3e-06 | 101.1908 | 101.4922 | 123.3962 | 21.1279 | 0.002163 | 0.002054 | -152.1499 | 0.23809 | 8.4013 | 0.003276 | 0.004381 | 0.004228 | 5.5e-05 | 0.000909 | 0.29856 | 1.2763 | 2.5479 | 99.51 | 20.9 | 37.357 | 73.7293 | 100.5355 | 71.1084 | 0.00929 | 1.0255 | 0.024607 | 23.9941 | -0.026604 | 0.14164 | -1e-05 | -0.000541 | -0.000541 | -10.7705
2019-03-21 12:06 | 1553184381.3768 | 1268 | 0.20961 | 0.000469 | -3e-06 | 101.1637 | 101.4023 | 123.1859 | 21.1278 | 0.00211 | 0.002024 | -152.1477 | 0.2388 | 8.4025 | 0.003225 | 0.004169 | 0.004022 | 5.7e-05 | 0.000908 | 0.30934 | 1.324 | 2.5477 | 99.51 | 20.9 | 37.357 | 72.4609 | 99.3116 | 70.4042 | 0.00929 | 1.0276 | 0.024611 | 23.9491 | -0.021073 | 0.14206 | -8.9e-05 | -0.006359 | -0.006359 | -0.9182
2019-03-21 12:06 | 1553184382.3762 | 1269 | 0.20958 | 0.000465 | -3e-06 | 101.2784 | 101.528 | 123.0927 | 21.1278 | 0.002186 | 0.002009 | -152.1487 | 0.23963 | 8.4019 | 0.003214 | 0.004253 | 0.004049 | 5.8e-05 | 0.000909 | 0.31492 | 1.3458 | 2.548 | 99.51 | 20.9 | 37.354 | 74.288 | 99.0285 | 70.0415 | 0.009289 | 1.0287 | 0.024614 | 23.9274 | -0.022016 | 0.14255 | 8.3e-05 | 0.006346 | 0.006346 | 0.92325
2019-03-21 12:06 | 1553184383.3773 | 1270 | 0.20958 | 0.000469 | -3e-06 | 101.1968 | 101.2692 | 122.877 | 21.1269 | 0.002152 | 0.001994 | -152.157 | 0.23895 | 8.4021 | 0.003258 | 0.004237 | 0.003949 | 5.4e-05 | 0.000909 | 0.31598 | 1.3213 | 2.548 | 99.51 | 20.89 | 37.354 | 73.4624 | 100.0905 | 69.6751 | 0.009283 | 1.0298 | 0.024606 | 23.8937 | -0.006395 | 0.14215 | 5.1e-05 | 0.003995 | 0.003995 | 1.4623
2019-03-21 12:06 | 1553184384.377 | 1271 | 0.20961 | 0.000464 | -4e-06 | 101.2441 | 101.3124 | 122.634 | 21.1264 | 0.002155 | 0.002022 | -152.1496 | 0.23905 | 8.4013 | 0.003322 | 0.005548 | 0.005399 | 5.9e-05 | 0.000905 | 0.31277 | 1.3173 | 2.5479 | 99.51 | 20.9 | 37.351 | 73.5307 | 101.6371 | 70.3466 | 0.009288 | 1.0278 | 0.024558 | 23.8934 | -0.006029 | 0.14221 | -8e-05 | -0.005686 | -0.005686 | -1.028
2019-03-21 12:06 | 1553184385.3769 | 1272 | 0.20959 | 0.000469 | -3e-06 | 101.1905 | 101.4408 | 122.6927 | 21.1262 | 0.002159 | 0.002013 | -152.1432 | 0.23969 | 8.4033 | 0.003348 | 0.004222 | 0.003996 | 5.9e-05 | 0.000907 | 0.31564 | 1.3109 | 2.5479 | 99.51 | 20.9 | 37.349 | 73.6403 | 102.2644 | 70.1268 | 0.009288 | 1.0285 | 0.024571 | 23.8915 | -0.022097 | 0.14259 | 6e-06 | 0.000665 | 0.000665 | 8.8066
2019-03-21 12:06 | 1553184386.3764 | 1273 | 0.20957 | 0.000464 | -3e-06 | 101.1664 | 101.6309 | 122.7042 | 21.1265 | 0.002112 | 0.001944 | -152.1513 | 0.23902 | 8.4028 | 0.003337 | 0.004774 | 0.004514 | 5.5e-05 | 0.000907 | 0.33251 | 1.3351 | 2.548 | 99.51 | 20.89 | 37.347 | 72.5041 | 102.0047 | 68.4736 | 0.009282 | 1.0334 | 0.024632 | 23.835 | -0.040972 | 0.14219 | 0.000157 | 0.01178 | 0.01178 | 0.49611
2019-03-21 12:06 | 1553184387.3751 | 1274 | 0.20958 | 0.000461 | -3e-06 | 101.2242 | 101.513 | 122.9143 | 21.1276 | 0.002146 | 0.002034 | -152.1574 | 0.23956 | 8.4007 | 0.003384 | 0.003951 | 0.003854 | 5.8e-05 | 0.000905 | 0.33268 | 1.3173 | 2.548 | 99.51 | 20.9 | 37.344 | 73.3283 | 103.1432 | 70.6285 | 0.009286 | 1.027 | 0.024576 | 23.9305 | -0.025486 | 0.14251 | 5.9e-05 | 0.004558 | 0.004558 | 1.2851
2019-03-21 12:06 | 1553184388.3751 | 1275 | 0.20958 | 0.000462 | -4e-06 | 101.2173 | 101.3141 | 123.1237 | 21.1272 | 0.002135 | 0.001981 | -152.1223 | 0.23889 | 8.4009 | 0.003301 | 0.004473 | 0.004242 | 5.8e-05 | 0.000911 | 0.36658 | 1.3471 | 2.548 | 99.51 | 20.9 | 37.343 | 73.0497 | 101.1338 | 69.3711 | 0.009286 | 1.0307 | 0.024642 | 23.907 | -0.008548 | 0.14212 | 8.1e-05 | 0.006251 | 0.006251 | 0.93446
2019-03-21 12:06 | 1553184389.3765 | 1276 | 0.20959 | 0.000464 | -3e-06 | 101.1923 | 101.2024 | 122.9643 | 21.1269 | 0.002117 | 0.002053 | -152.1227 | 0.23999 | 8.4031 | 0.003285 | 0.004853 | 0.004647 | 6e-05 | 0.000909 | 0.32836 | 1.2851 | 2.548 | 99.51 | 20.9 | 37.342 | 72.6392 | 100.7426 | 71.0911 | 0.009286 | 1.0256 | 0.024564 | 23.9515 | -0.000892 | 0.14277 | 3.8e-05 | 0.002999 | 0.002999 | 1.9568
2019-03-21 12:06 | 1553184390.3752 | 1277 | 0.20958 | 0.000465 | -4e-06 | 101.2304 | 101.4568 | 123.1183 | 21.128 | 0.002107 | 0.001987 | -152.1189 | 0.23908 | 8.4023 | 0.003273 | 0.004664 | 0.004509 | 5.9e-05 | 0.000913 | 0.31437 | 1.2839 | 2.548 | 99.51 | 20.9 | 37.343 | 72.3792 | 100.4573 | 69.5046 | 0.009286 | 1.0303 | 0.024636 | 23.9112 | -0.01998 | 0.14223 | 7.9e-05 | 0.006055 | 0.006055 | 0.96542
2019-03-21 12:06 | 1553184391.3763 | 1278 | 0.20956 | 0.000471 | -4e-06 | 101.2155 | 101.401 | 123.2549 | 21.1264 | 0.002149 | 0.002031 | -152.118 | 0.23801 | 8.4032 | 0.003213 | 0.004505 | 0.004363 | 6.1e-05 | 0.000914 | 0.31421 | 1.304 | 2.5479 | 99.51 | 20.9 | 37.344 | 73.4011 | 99.0052 | 70.5612 | 0.009286 | 1.0272 | 0.024612 | 23.9613 | -0.01637 | 0.14159 | 0.000203 | 0.01526 | 0.01526 | 0.38136
2019-03-21 12:07 | 1553184392.3767 | 1279 | 0.20958 | 0.000465 | -4e-06 | 101.2555 | 101.3298 | 123.2098 | 21.1277 | 0.002134 | 0.002005 | -152.1239 | 0.23901 | 8.4007 | 0.003223 | 0.004956 | 0.004748 | 5.8e-05 | 0.000916 | 0.35706 | 1.26 | 2.5478 | 99.51 | 20.9 | 37.35 | 73.0322 | 99.2606 | 69.93 | 0.009288 | 1.029 | 0.02463 | 23.9349 | -0.006562 | 0.14219 | 5.1e-05 | 0.003978 | 0.003978 | 1.4691
2019-03-21 12:07 | 1553184393.3765 | 1280 | 0.20953 | 0.000467 | -4e-06 | 101.1732 | 101.4888 | 123.2579 | 21.1271 | 0.002184 | 0.002016 | -152.1244 | 0.23874 | 8.402 | 0.003303 | 0.004752 | 0.004621 | 5.3e-05 | 0.000913 | 0.3126 | 1.2843 | 2.5481 | 99.51 | 20.9 | 37.355 | 74.2422 | 101.1911 | 70.1992 | 0.009289 | 1.0282 | 0.024625 | 23.949 | -0.027858 | 0.14202 | 0.000394 | 0.029372 | 0.029372 | 0.19873
2019-03-21 12:07 | 1553184394.3754 | 1281 | 0.20959 | 0.000462 | -3e-06 | 101.2313 | 101.3202 | 123.4504 | 21.128 | 0.002131 | 0.002008 | -152.1518 | 0.23975 | 8.4016 | 0.003315 | 0.004988 | 0.004732 | 5.7e-05 | 0.000919 | 0.34307 | 1.3041 | 2.548 | 99.51 | 20.9 | 37.359 | 72.9741 | 101.4863 | 70.0067 | 0.00929 | 1.0288 | 0.024651 | 23.9609 | -0.007852 | 0.14263 | 1.6e-05 | 0.001405 | 0.001405 | 4.1728
2019-03-21 12:07 | 1553184395.3768 | 1282 | 0.20957 | 0.000466 | -3e-06 | 101.2239 | 101.3764 | 123.4182 | 21.1273 | 0.002074 | 0.001994 | -152.1628 | 0.23872 | 8.4018 | 0.003238 | 0.004804 | 0.004735 | 5.8e-05 | 0.000915 | 0.33826 | 1.3316 | 2.548 | 99.51 | 20.9 | 37.371 | 71.5904 | 99.6044 | 69.6849 | 0.009293 | 1.0298 | 0.02466 | 23.9466 | -0.01347 | 0.14202 | 0.000128 | 0.009681 | 0.009681 | 0.60293
2019-03-21 12:07 | 1553184396.3769 | 1283 | 0.20957 | 0.000465 | -3e-06 | 101.2009 | 101.405 | 123.7185 | 21.126 | 0.002098 | 0.001961 | -152.1653 | 0.23931 | 8.4018 | 0.003363 | 0.004195 | 0.004029 | 5.5e-05 | 0.000917 | 0.30287 | 1.2821 | 2.5478 | 99.51 | 20.9 | 37.379 | 72.1646 | 102.6285 | 68.8818 | 0.009295 | 1.0322 | 0.024719 | 23.9476 | -0.018021 | 0.14236 | 0.000151 | 0.011441 | 0.011441 | 0.51143
2019-03-21 12:07 | 1553184397.3766 | 1284 | 0.20956 | 0.000466 | -4e-06 | 101.1393 | 101.354 | 123.9426 | 21.1265 | 0.00213 | 0.002001 | -152.1651 | 0.23984 | 8.404 | 0.003343 | 0.004781 | 0.004559 | 5.6e-05 | 0.000914 | 0.29175 | 1.2923 | 2.5478 | 99.51 | 20.9 | 37.396 | 72.9289 | 102.1642 | 69.8523 | 0.009299 | 1.0293 | 0.024706 | 24.0033 | -0.018966 | 0.14268 | 0.000212 | 0.016009 | 0.016009 | 0.36631
2019-03-21 12:07 | 1553184398.3761 | 1285 | 0.20955 | 0.000471 | -4e-06 | 101.1393 | 101.388 | 124.0403 | 21.1266 | 0.002066 | 0.002031 | -152.1782 | 0.23844 | 8.4043 | 0.003285 | 0.004381 | 0.0043 | 5.7e-05 | 0.000914 | 0.323 | 1.2851 | 2.5479 | 99.51 | 20.9 | 37.406 | 71.3933 | 100.7479 | 70.5698 | 0.009302 | 1.0271 | 0.02469 | 24.0378 | -0.021972 | 0.14185 | 0.000255 | 0.019108 | 0.019108 | 0.3051
2019-03-21 12:07 | 1553184399.3762 | 1286 | 0.20958 | 0.000465 | -4e-06 | 101.1789 | 101.4289 | 124.3196 | 21.1259 | 0.002131 | 0.001975 | -152.169 | 0.23898 | 8.4038 | 0.00326 | 0.004533 | 0.004348 | 5.9e-05 | 0.000914 | 0.33373 | 1.3377 | 2.548 | 99.51 | 20.9 | 37.418 | 72.9557 | 100.1381 | 69.2148 | 0.009305 | 1.0312 | 0.024767 | 24.0174 | -0.022073 | 0.14217 | 6.2e-05 | 0.004852 | 0.004852 | 1.2043
2019-03-21 12:07 | 1553184400.3768 | 1287 | 0.20959 | 0.000458 | -4e-06 | 101.229 | 101.4166 | 124.37 | 21.1271 | 0.002104 | 0.001996 | -152.1827 | 0.23898 | 8.4032 | 0.003376 | 0.00441 | 0.004219 | 5.8e-05 | 0.000913 | 0.32947 | 1.3363 | 2.548 | 99.51 | 20.9 | 37.441 | 72.3098 | 102.9659 | 69.7181 | 0.009311 | 1.0297 | 0.024753 | 24.0399 | -0.016568 | 0.14217 | 4.1e-05 | 0.003294 | 0.003294 | 1.7738
2019-03-21 12:07 | 1553184401.3772 | 1288 | 0.20962 | 0.000464 | -4e-06 | 101.2123 | 101.3478 | 124.2726 | 21.1276 | 0.002112 | 0.001993 | -152.1845 | 0.23943 | 8.4023 | 0.003281 | 0.004425 | 0.00422 | 5.6e-05 | 0.000913 | 0.29856 | 1.3115 | 2.548 | 99.51 | 20.9 | 37.452 | 72.5177 | 100.6468 | 69.6554 | 0.009313 | 1.0299 | 0.024746 | 24.0283 | -0.011972 | 0.14244 | -0.000175 | -0.012808 | -0.012808 | -0.45706
2019-03-21 12:07 | 1553184402.375 | 1289 | 0.20957 | 0.000462 | -4e-06 | 101.1649 | 101.388 | 124.39 | 21.127 | 0.002069 | 0.001992 | -152.1736 | 0.23905 | 8.401 | 0.003332 | 0.003936 | 0.003756 | 5.8e-05 | 0.000911 | 0.29966 | 1.3218 | 2.548 | 99.51 | 20.9 | 37.464 | 71.4817 | 101.8787 | 69.6245 | 0.009316 | 1.03 | 0.024759 | 24.0385 | -0.019706 | 0.14221 | 0.000181 | 0.013684 | 0.013684 | 0.42715
2019-03-21 12:07 | 1553184403.375 | 1290 | 0.20958 | 0.000461 | -4e-06 | 101.1616 | 101.3151 | 124.6022 | 21.1271 | 0.002122 | 0.001938 | -152.1738 | 0.23916 | 8.4032 | 0.003303 | 0.005041 | 0.004937 | 5.5e-05 | 0.000915 | 0.32117 | 1.3068 | 2.5479 | 99.51 | 20.9 | 37.484 | 72.7603 | 101.187 | 68.3299 | 0.009321 | 1.0339 | 0.024827 | 24.0136 | -0.013563 | 0.14228 | 0.000103 | 0.00792 | 0.00792 | 0.73828
2019-03-21 12:07 | 1553184404.375 | 1291 | 0.20957 | 0.000462 | -3e-06 | 101.1977 | 101.3533 | 124.4247 | 21.1264 | 0.002073 | 0.001977 | -152.1545 | 0.23898 | 8.4031 | 0.003342 | 0.004115 | 0.00391 | 5.5e-05 | 0.000913 | 0.31022 | 1.2937 | 2.548 | 99.51 | 20.9 | 37.492 | 71.5623 | 102.1396 | 69.2713 | 0.009323 | 1.031 | 0.024775 | 24.0295 | -0.013743 | 0.14217 | 0.000166 | 0.012575 | 0.012575 | 0.46468
2019-03-21 12:07 | 1553184405.376 | 1292 | 0.2096 | 0.000459 | -3e-06 | 101.2048 | 101.3495 | 124.1393 | 21.1268 | 0.00206 | 0.002038 | -152.1444 | 0.23884 | 8.4004 | 0.003332 | 0.004612 | 0.004468 | 5.9e-05 | 0.000912 | 0.32532 | 1.3006 | 2.5478 | 99.51 | 20.9 | 37.508 | 71.2585 | 101.8919 | 70.7329 | 0.009327 | 1.0266 | 0.024694 | 24.0531 | -0.012781 | 0.14209 | -3.2e-05 | -0.002135 | -0.002135 | -2.7352
2019-03-21 12:07 | 1553184406.3751 | 1293 | 0.20957 | 0.000461 | -4e-06 | 101.1813 | 101.4807 | 123.7544 | 21.1271 | 0.002122 | 0.001986 | -152.1374 | 0.23904 | 8.4015 | 0.003318 | 0.004513 | 0.004335 | 5.7e-05 | 0.000917 | 0.32012 | 1.2775 | 2.5479 | 99.51 | 20.9 | 37.516 | 72.7557 | 101.5466 | 69.4763 | 0.009329 | 1.0304 | 0.024701 | 23.9719 | -0.026426 | 0.1422 | 0.000136 | 0.010295 | 0.010295 | 0.56771
2019-03-21 12:07 | 1553184407.3751 | 1294 | 0.20961 | 0.000467 | -4e-06 | 101.1527 | 101.4119 | 123.4455 | 21.1273 | 0.002117 | 0.001954 | -152.1503 | 0.23909 | 8.4021 | 0.003295 | 0.004107 | 0.003909 | 5.3e-05 | 0.000915 | 0.3115 | 1.3356 | 2.548 | 99.51 | 20.9 | 37.525 | 72.6351 | 100.9972 | 68.7054 | 0.009332 | 1.0327 | 0.024698 | 23.915 | -0.022889 | 0.14223 | -0.000117 | -0.008449 | -0.008449 | -0.69192
2019-03-21 12:07 | 1553184408.3752 | 1295 | 0.20961 | 0.00046 | -3e-06 | 101.1726 | 101.4793 | 123.0117 | 21.1276 | 0.002098 | 0.002016 | -152.1672 | 0.2388 | 8.4008 | 0.003249 | 0.004862 | 0.004652 | 5.7e-05 | 0.000918 | 0.33373 | 1.2812 | 2.5478 | 99.51 | 20.9 | 37.529 | 72.1731 | 99.8794 | 70.1997 | 0.009333 | 1.0282 | 0.024601 | 23.925 | -0.02707 | 0.14206 | -0.0001 | -0.007211 | -0.007211 | -0.80966
2019-03-21 12:07 | 1553184409.3762 | 1296 | 0.20958 | 0.000463 | -3e-06 | 101.2397 | 101.4255 | 123.0816 | 21.1266 | 0.002077 | 0.001976 | -152.1691 | 0.23878 | 8.4042 | 0.003239 | 0.004377 | 0.00423 | 5.7e-05 | 0.000917 | 0.29264 | 1.284 | 2.5478 | 99.51 | 20.9 | 37.531 | 71.6605 | 99.6438 | 69.2419 | 0.009333 | 1.0311 | 0.024642 | 23.8984 | -0.016401 | 0.14205 | 0.000109 | 0.008293 | 0.008293 | 0.70398
2019-03-21 12:07 | 1553184410.3757 | 1297 | 0.2096 | 0.000462 | -3e-06 | 101.2447 | 101.2341 | 123.0211 | 21.1259 | 0.002037 | 0.001986 | -152.1551 | 0.23932 | 8.4031 | 0.00333 | 0.003836 | 0.003791 | 5.4e-05 | 0.000919 | 0.30989 | 1.2775 | 2.5479 | 99.51 | 20.9 | 37.534 | 70.7095 | 101.8354 | 69.4807 | 0.009334 | 1.0304 | 0.024627 | 23.9009 | 0.000943 | 0.14237 | -4.4e-05 | -0.003063 | -0.003063 | -1.9101
2019-03-21 12:07 | 1553184411.3764 | 1298 | 0.2096 | 0.00046 | -3e-06 | 101.1953 | 101.5263 | 122.8472 | 21.127 | 0.00206 | 0.001944 | -152.144 | 0.23919 | 8.4025 | 0.003309 | 0.003984 | 0.003764 | 5.5e-05 | 0.000923 | 0.3194 | 1.3098 | 2.5479 | 99.51 | 20.9 | 37.534 | 71.2667 | 101.3324 | 68.4758 | 0.009334 | 1.0334 | 0.024646 | 23.8489 | -0.02921 | 0.14229 | -3.3e-05 | -0.002241 | -0.002241 | -2.6097
2019-03-21 12:07 | 1553184412.3764 | 1299 | 0.20958 | 0.000456 | -3e-06 | 101.1664 | 101.4429 | 122.7603 | 21.1269 | 0.002038 | 0.001951 | -152.1433 | 0.23857 | 8.4031 | 0.003279 | 0.004529 | 0.00438 | 5.4e-05 | 0.000925 | 0.31404 | 1.338 | 2.5478 | 99.51 | 20.9 | 37.533 | 70.7253 | 100.6108 | 68.6293 | 0.009333 | 1.033 | 0.024632 | 23.8459 | -0.02441 | 0.14192 | 0.000108 | 0.008211 | 0.008211 | 0.71043
2019-03-21 12:07 | 1553184413.3756 | 1300 | 0.20958 | 0.000464 | -4e-06 | 101.1682 | 101.405 | 122.3674 | 21.1274 | 0.002077 | 0.002 | -152.127 | 0.23969 | 8.401 | 0.003304 | 0.004567 | 0.004327 | 5.3e-05 | 0.000925 | 0.33931 | 1.3278 | 2.5479 | 99.51 | 20.9 | 37.528 | 71.67 | 101.2055 | 69.8285 | 0.009332 | 1.0294 | 0.024549 | 23.8494 | -0.020918 | 0.14259 | 6.3e-05 | 0.004839 | 0.004839 | 1.2112
2019-03-21 12:07 | 1553184414.3755 | 1301 | 0.20959 | 0.000456 | -3e-06 | 101.2325 | 101.3499 | 122.5481 | 21.1271 | 0.002092 | 0.001925 | -152.1383 | 0.24032 | 8.4012 | 0.003327 | 0.004618 | 0.004459 | 5.4e-05 | 0.000925 | 0.32012 | 1.3581 | 2.548 | 99.51 | 20.9 | 37.525 | 72.0393 | 101.776 | 68.0085 | 0.009332 | 1.0348 | 0.024633 | 23.8036 | -0.010362 | 0.14297 | 4.1e-05 | 0.003181 | 0.003181 | 1.8474
2019-03-21 12:07 | 1553184415.3756 | 1302 | 0.20959 | 0.000462 | -4e-06 | 101.1977 | 101.2566 | 122.653 | 21.1275 | 0.00209 | 0.001975 | -152.1393 | 0.2393 | 8.4001 | 0.003238 | 0.003894 | 0.003701 | 5.5e-05 | 0.000922 | 0.33395 | 1.3356 | 2.5481 | 99.51 | 20.9 | 37.517 | 71.9812 | 99.6242 | 69.207 | 0.00933 | 1.0312 | 0.0246 | 23.8556 | -0.005203 | 0.14236 | 1.1e-05 | 0.001027 | 0.001027 | 5.6983
2019-03-21 12:07 | 1553184416.3756 | 1303 | 0.20961 | 0.00046 | -3e-06 | 101.2501 | 101.3253 | 122.7697 | 21.1272 | 0.002123 | 0.001937 | -152.1264 | 0.23982 | 8.4023 | 0.003336 | 0.00463 | 0.00447 | 5.2e-05 | 0.000925 | 0.27273 | 1.3168 | 2.5479 | 99.51 | 20.9 | 37.512 | 72.7806 | 101.9966 | 68.3026 | 0.009328 | 1.034 | 0.024645 | 23.8354 | -0.006645 | 0.14267 | -9.6e-05 | -0.006889 | -0.006889 | -0.85111
2019-03-21 12:07 | 1553184417.3756 | 1304 | 0.20958 | 0.000461 | -3e-06 | 101.161 | 101.355 | 122.8476 | 21.1267 | 0.002017 | 0.001992 | -152.1061 | 0.2396 | 8.4031 | 0.003299 | 0.004767 | 0.004644 | 5.3e-05 | 0.000926 | 0.28456 | 1.3237 | 2.5479 | 99.51 | 20.9 | 37.506 | 70.219 | 101.0791 | 69.6179 | 0.009327 | 1.03 | 0.024605 | 23.8888 | -0.017133 | 0.14254 | 0.000101 | 0.007686 | 0.007686 | 0.7622
2019-03-21 12:07 | 1553184418.3749 | 1305 | 0.20957 | 0.000462 | -4e-06 | 101.1428 | 101.358 | 123.0591 | 21.1273 | 0.002 | 0.001975 | -152.0737 | 0.23975 | 8.4028 | 0.003291 | 0.004393 | 0.004195 | 5.3e-05 | 0.000928 | 0.33588 | 1.3549 | 2.5481 | 99.51 | 20.9 | 37.492 | 69.8089 | 100.9 | 69.2051 | 0.009323 | 1.0312 | 0.024641 | 23.895 | -0.019011 | 0.14262 | 0.000138 | 0.010474 | 0.010474 | 0.55968
2019-03-21 12:07 | 1553184419.3758 | 1306 | 0.20959 | 0.000456 | -4e-06 | 101.158 | 101.3897 | 123.2046 | 21.1282 | 0.002067 | 0.001987 | -152.0895 | 0.24029 | 8.4026 | 0.003363 | 0.004569 | 0.004383 | 5.4e-05 | 0.000925 | 0.30807 | 1.3666 | 2.5481 | 99.51 | 20.9 | 37.486 | 71.4252 | 102.6338 | 69.4961 | 0.009322 | 1.0304 | 0.024645 | 23.9193 | -0.020462 | 0.14295 | 5.7e-05 | 0.004446 | 0.004446 | 1.3213
2019-03-21 12:07 | 1553184420.3752 | 1307 | 0.20958 | 0.000467 | -4e-06 | 101.1571 | 101.2712 | 123.4168 | 21.1274 | 0.002055 | 0.00195 | -152.1032 | 0.23954 | 8.4025 | 0.00329 | 0.00488 | 0.004699 | 5.5e-05 | 0.000924 | 0.3449 | 1.3254 | 2.5479 | 99.51 | 20.9 | 37.472 | 71.1485 | 100.8756 | 68.6102 | 0.009318 | 1.033 | 0.024698 | 23.9089 | -0.010079 | 0.1425 | 8.3e-05 | 0.006384 | 0.006384 | 0.91741
2019-03-21 12:07 | 1553184421.3749 | 1308 | 0.20959 | 0.000462 | -4e-06 | 101.1357 | 101.1925 | 123.9935 | 21.1278 | 0.002043 | 0.001955 | -152.1163 | 0.24 | 8.4055 | 0.003352 | 0.004503 | 0.004465 | 5.5e-05 | 0.000925 | 0.30735 | 1.3275 | 2.5478 | 99.51 | 20.9 | 37.463 | 70.8517 | 102.3705 | 68.7347 | 0.009316 | 1.0326 | 0.024751 | 23.969 | -0.005024 | 0.14278 | 2.8e-05 | 0.002299 | 0.002299 | 2.5523
2019-03-21 12:07 | 1553184422.376 | 1309 | 0.20957 | 0.000461 | -4e-06 | 101.1679 | 101.4228 | 124.2635 | 21.1267 | 0.002093 | 0.001906 | -152.1333 | 0.23956 | 8.4057 | 0.003345 | 0.005021 | 0.004839 | 5.6e-05 | 0.000928 | 0.3126 | 1.315 | 2.548 | 99.51 | 20.9 | 37.455 | 72.0576 | 102.1988 | 67.5588 | 0.009314 | 1.0362 | 0.024821 | 23.9538 | -0.022507 | 0.14252 | 0.000154 | 0.011703 | 0.011703 | 0.5005
2019-03-21 12:07 | 1553184423.3753 | 1310 | 0.20962 | 0.000463 | -4e-06 | 101.1601 | 101.4211 | 124.5566 | 21.128 | 0.002114 | 0.00193 | -152.1445 | 0.23895 | 8.404 | 0.003414 | 0.004621 | 0.004476 | 5.4e-05 | 0.000926 | 0.31598 | 1.3253 | 2.5479 | 99.51 | 20.9 | 37.437 | 72.5694 | 103.8691 | 68.1341 | 0.00931 | 1.0345 | 0.024829 | 24.0023 | -0.023042 | 0.14215 | -0.000123 | -0.008895 | -0.008895 | -0.65685
2019-03-21 12:07 | 1553184424.376 | 1311 | 0.20957 | 0.00046 | -3e-06 | 101.1056 | 101.3342 | 124.2935 | 21.1274 | 0.002001 | 0.001951 | -152.1406 | 0.23961 | 8.4016 | 0.003334 | 0.004754 | 0.004631 | 5.4e-05 | 0.000923 | 0.30702 | 1.3616 | 2.5479 | 99.51 | 20.9 | 37.428 | 69.8455 | 101.9324 | 68.6437 | 0.009307 | 1.0329 | 0.024785 | 23.9948 | -0.020201 | 0.14254 | 0.000149 | 0.01128 | 0.01128 | 0.51936
2019-03-21 12:07 | 1553184425.3761 | 1312 | 0.20959 | 0.000463 | -4e-06 | 101.1819 | 101.5018 | 124.4074 | 21.1269 | 0.002029 | 0.001925 | -152.1389 | 0.23939 | 8.4015 | 0.003349 | 0.004669 | 0.004479 | 5.2e-05 | 0.000924 | 0.3178 | 1.3148 | 2.5479 | 99.51 | 20.9 | 37.419 | 70.5124 | 102.3066 | 68.0045 | 0.009305 | 1.0349 | 0.024819 | 23.9834 | -0.028234 | 0.14241 | 6.1e-05 | 0.004807 | 0.004807 | 1.2175
2019-03-21 12:07 | 1553184426.3766 | 1313 | 0.20958 | 0.000463 | -4e-06 | 101.1154 | 101.2906 | 124.5876 | 21.1261 | 0.002107 | 0.001933 | -152.1451 | 0.23999 | 8.4042 | 0.003398 | 0.004102 | 0.003937 | 5.4e-05 | 0.000926 | 0.31526 | 1.359 | 2.5479 | 99.51 | 20.9 | 37.404 | 72.3922 | 103.4898 | 68.1994 | 0.009301 | 1.0343 | 0.02483 | 24.0076 | -0.015482 | 0.14277 | 0.000104 | 0.007963 | 0.007963 | 0.73692
2019-03-21 12:07 | 1553184427.3766 | 1314 | 0.20961 | 0.000461 | -4e-06 | 101.1321 | 101.2664 | 124.7282 | 21.1266 | 0.002029 | 0.001984 | -152.1477 | 0.2397 | 8.4024 | 0.003355 | 0.004912 | 0.004764 | 5.5e-05 | 0.000928 | 0.32748 | 1.3539 | 2.5478 | 99.51 | 20.9 | 37.396 | 70.5254 | 102.4432 | 69.4213 | 0.009299 | 1.0306 | 0.0248 | 24.0641 | -0.011871 | 0.1426 | -6.9e-05 | -0.004871 | -0.004871 | -1.2031
2019-03-21 12:07 | 1553184428.3753 | 1315 | 0.20958 | 0.000461 | -4e-06 | 101.1408 | 101.2065 | 124.9529 | 21.1268 | 0.002048 | 0.001948 | -152.1701 | 0.23934 | 8.4026 | 0.003327 | 0.004457 | 0.004294 | 5.6e-05 | 0.000931 | 0.30254 | 1.3612 | 2.5478 | 99.51 | 20.9 | 37.384 | 70.9764 | 101.7593 | 68.5739 | 0.009296 | 1.0331 | 0.024853 | 24.0559 | -0.005811 | 0.14238 | 0.000115 | 0.008845 | 0.008845 | 0.66159
2019-03-21 12:07 | 1553184429.3767 | 1316 | 0.20958 | 0.000461 | -4e-06 | 101.1271 | 101.1643 | 124.8982 | 21.1271 | 0.002025 | 0.001914 | -152.168 | 0.2396 | 8.4028 | 0.003357 | 0.004283 | 0.004159 | 5.4e-05 | 0.000932 | 0.33019 | 1.3618 | 2.548 | 99.51 | 20.89 | 37.38 | 70.4299 | 102.5 | 67.744 | 0.00929 | 1.0356 | 0.024878 | 24.0214 | -0.003289 | 0.14253 | 0.000127 | 0.009762 | 0.009762 | 0.60008
2019-03-21 12:07 | 1553184430.3763 | 1317 | 0.20958 | 0.000458 | -3e-06 | 101.1652 | 101.3407 | 124.7294 | 21.1262 | 0.002045 | 0.001923 | -152.1609 | 0.2387 | 8.403 | 0.003303 | 0.004414 | 0.004191 | 5.2e-05 | 0.000933 | 0.32228 | 1.3369 | 2.5479 | 99.51 | 20.89 | 37.374 | 70.8876 | 101.1745 | 67.9631 | 0.009288 | 1.035 | 0.024853 | 24.0129 | -0.015502 | 0.142 | 0.000115 | 0.008788 | 0.008788 | 0.66414
2019-03-21 12:07 | 1553184431.3754 | 1318 | 0.20958 | 0.000457 | -4e-06 | 101.1145 | 101.3693 | 124.2337 | 21.1266 | 0.002031 | 0.001965 | -152.1779 | 0.23999 | 8.405 | 0.003395 | 0.004605 | 0.004428 | 5.1e-05 | 0.000931 | 0.30397 | 1.3266 | 2.5478 | 99.51 | 20.89 | 37.366 | 70.5527 | 103.4062 | 68.9838 | 0.009286 | 1.0319 | 0.024766 | 24.001 | -0.022506 | 0.14277 | 7.6e-05 | 0.005901 | 0.005901 | 0.99444
2019-03-21 12:07 | 1553184432.3759 | 1319 | 0.20956 | 0.000459 | -4e-06 | 101.0433 | 101.2327 | 123.8445 | 21.1266 | 0.002029 | 0.001933 | -152.1782 | 0.23909 | 8.4044 | 0.003391 | 0.004744 | 0.004527 | 5.2e-05 | 0.00093 | 0.28241 | 1.2963 | 2.548 | 99.51 | 20.89 | 37.361 | 70.5223 | 103.3145 | 68.2148 | 0.009285 | 1.0342 | 0.024755 | 23.9364 | -0.016748 | 0.14223 | 0.000226 | 0.017055 | 0.017055 | 0.34277
2019-03-21 12:07 | 1553184433.3746 | 1320 | 0.20959 | 0.000453 | -4e-06 | 101.0618 | 101.2882 | 123.9342 | 21.127 | 0.002016 | 0.001937 | -152.1962 | 0.23967 | 8.4036 | 0.003356 | 0.004651 | 0.004479 | 5.2e-05 | 0.000927 | 0.32565 | 1.3023 | 2.5479 | 99.51 | 20.89 | 37.358 | 70.1922 | 102.4653 | 68.3024 | 0.009284 | 1.034 | 0.024761 | 23.9481 | -0.020017 | 0.14258 | 5.5e-05 | 0.004313 | 0.004313 | 1.3588
2019-03-21 12:07 | 1553184434.3746 | 1321 | 0.20959 | 0.000459 | -4e-06 | 101.0558 | 101.2627 | 123.9931 | 21.1272 | 0.002025 | 0.001994 | -152.1824 | 0.2392 | 8.4015 | 0.003314 | 0.00474 | 0.004549 | 5.1e-05 | 0.000929 | 0.29574 | 1.3023 | 2.5479 | 99.51 | 20.89 | 37.357 | 70.4255 | 101.4518 | 69.6636 | 0.009284 | 1.0298 | 0.024718 | 24.0015 | -0.018288 | 0.1423 | 5.6e-05 | 0.004376 | 0.004376 | 1.3365
2019-03-21 12:07 | 1553184435.375 | 1322 | 0.20956 | 0.000456 | -3e-06 | 101.0624 | 101.3046 | 123.7802 | 21.1261 | 0.002079 | 0.001999 | -152.1931 | 0.23889 | 8.402 | 0.003222 | 0.004409 | 0.004216 | 5e-05 | 0.00093 | 0.31421 | 1.3094 | 2.548 | 99.51 | 20.89 | 37.356 | 71.7038 | 99.2323 | 69.7911 | 0.009284 | 1.0295 | 0.024692 | 23.9854 | -0.021408 | 0.14211 | 0.000214 | 0.016085 | 0.016085 | 0.36312
2019-03-21 12:07 | 1553184436.3761 | 1323 | 0.20958 | 0.000458 | -4e-06 | 100.9671 | 100.9957 | 123.4091 | 21.1271 | 0.00202 | 0.001942 | -152.2094 | 0.24091 | 8.4039 | 0.003212 | 0.005471 | 0.005316 | 5.5e-05 | 0.000932 | 0.29767 | 1.3017 | 2.548 | 99.51 | 20.89 | 37.353 | 70.2999 | 98.9783 | 68.4328 | 0.009283 | 1.0336 | 0.024704 | 23.9019 | -0.002533 | 0.14332 | 0.000135 | 0.010288 | 0.010288 | 0.57253
2019-03-21 12:07 | 1553184437.3758 | 1324 | 0.20958 | 0.000459 | -3e-06 | 101.0097 | 101.0757 | 123.2307 | 21.1256 | 0.002072 | 0.001981 | -152.2075 | 0.23976 | 8.3996 | 0.00333 | 0.00397 | 0.003844 | 5.2e-05 | 0.000934 | 0.29623 | 1.3132 | 2.5479 | 99.51 | 20.89 | 37.352 | 71.5544 | 101.8443 | 69.3665 | 0.009283 | 1.0307 | 0.024652 | 23.9173 | -0.005846 | 0.14263 | 8.9e-05 | 0.006777 | 0.006777 | 0.865
2019-03-21 12:07 | 1553184438.3751 | 1325 | 0.20961 | 0.000455 | -4e-06 | 101.0424 | 101.1329 | 123.4089 | 21.1275 | 0.00203 | 0.001942 | -152.2121 | 0.23964 | 8.401 | 0.003289 | 0.005315 | 0.005134 | 5.3e-05 | 0.000932 | 0.33843 | 1.3282 | 2.5479 | 99.51 | 20.89 | 37.351 | 70.5325 | 100.8467 | 68.4325 | 0.009283 | 1.0336 | 0.024704 | 23.9019 | -0.008006 | 0.14256 | -7.9e-05 | -0.00557 | -0.00557 | -1.052
2019-03-21 12:07 | 1553184439.3763 | 1326 | 0.2096 | 0.000458 | -3e-06 | 101.0415 | 101.1748 | 123.2823 | 21.1274 | 0.001993 | 0.001912 | -152.2193 | 0.23959 | 8.4015 | 0.0033 | 0.004924 | 0.004712 | 5e-05 | 0.000931 | 0.30021 | 1.2916 | 2.5479 | 99.51 | 20.89 | 37.35 | 69.6615 | 101.1191 | 67.7047 | 0.009282 | 1.0358 | 0.024718 | 23.8642 | -0.011789 | 0.14253 | -8e-06 | -0.000375 | -0.000375 | -15.6318
2019-03-21 12:07 | 1553184440.3759 | 1327 | 0.20961 | 0.000457 | -3e-06 | 100.976 | 101.123 | 123.3368 | 21.1269 | 0.002026 | 0.001935 | -152.2192 | 0.23894 | 8.4018 | 0.003344 | 0.004414 | 0.004309 | 5e-05 | 0.00093 | 0.32101 | 1.3146 | 2.548 | 99.51 | 20.89 | 37.349 | 70.4447 | 102.1904 | 68.2456 | 0.009282 | 1.0341 | 0.024704 | 23.8884 | -0.013016 | 0.14214 | -6.3e-05 | -0.004453 | -0.004453 | -1.3118
2019-03-21 12:07 | 1553184441.3764 | 1328 | 0.20958 | 0.000456 | -4e-06 | 100.962 | 101.2552 | 123.5157 | 21.1277 | 0.002055 | 0.001959 | -152.2198 | 0.23924 | 8.4035 | 0.003284 | 0.003956 | 0.003845 | 5.2e-05 | 0.000933 | 0.32964 | 1.3063 | 2.5478 | 99.51 | 20.89 | 37.346 | 71.1354 | 100.7269 | 68.8351 | 0.009281 | 1.0323 | 0.0247 | 23.9263 | -0.025937 | 0.14233 | 0.000119 | 0.009081 | 0.009081 | 0.64418
2019-03-21 12:07 | 1553184442.3764 | 1329 | 0.20956 | 0.000457 | -4e-06 | 100.9197 | 101.0604 | 123.4966 | 21.1272 | 0.00198 | 0.001947 | -152.186 | 0.23965 | 8.4006 | 0.003202 | 0.00473 | 0.004547 | 5.1e-05 | 0.00093 | 0.31061 | 1.28 | 2.548 | 99.51 | 20.89 | 37.345 | 69.3408 | 98.7488 | 68.5451 | 0.009281 | 1.0332 | 0.024709 | 23.9143 | -0.01246 | 0.14257 | 0.000207 | 0.015612 | 0.015612 | 0.37532
2019-03-21 12:07 | 1553184443.3751 | 1330 | 0.20958 | 0.000453 | -4e-06 | 100.9289 | 101.0771 | 123.3685 | 21.1269 | 0.00199 | 0.001939 | -152.1822 | 0.23917 | 8.4011 | 0.003262 | 0.004154 | 0.003986 | 5.2e-05 | 0.000927 | 0.28479 | 1.2989 | 2.5478 | 99.51 | 20.89 | 37.345 | 69.5845 | 100.2003 | 68.346 | 0.009281 | 1.0338 | 0.024703 | 23.895 | -0.013119 | 0.14228 | 0.000142 | 0.010775 | 0.010775 | 0.54273
2019-03-21 12:07 | 1553184444.3763 | 1331 | 0.20957 | 0.000458 | -4e-06 | 100.9 | 101.0359 | 123.6649 | 21.1278 | 0.002052 | 0.001947 | -152.1894 | 0.23903 | 8.406 | 0.003197 | 0.005316 | 0.005104 | 5.4e-05 | 0.000929 | 0.30215 | 1.2751 | 2.5479 | 99.51 | 20.89 | 37.344 | 71.059 | 98.6261 | 68.5381 | 0.009281 | 1.0332 | 0.024726 | 23.9304 | -0.012033 | 0.1422 | 0.000183 | 0.013845 | 0.013845 | 0.42212
2019-03-21 12:07 | 1553184445.3756 | 1332 | 0.20954 | 0.000456 | -4e-06 | 100.9218 | 101.0641 | 123.9494 | 21.1275 | 0.002082 | 0.001937 | -152.1713 | 0.23869 | 8.4013 | 0.003288 | 0.004598 | 0.00446 | 5.3e-05 | 0.00093 | 0.29551 | 1.3583 | 2.548 | 99.51 | 20.89 | 37.343 | 71.7941 | 100.8153 | 68.3007 | 0.009281 | 1.034 | 0.024763 | 23.9496 | -0.012607 | 0.142 | 0.000331 | 0.024844 | 0.024844 | 0.23491
2019-03-21 12:07 | 1553184446.3761 | 1333 | 0.20957 | 0.000453 | -4e-06 | 100.855 | 100.9923 | 123.9105 | 21.1279 | 0.00202 | 0.001914 | -152.1592 | 0.2398 | 8.401 | 0.003266 | 0.004736 | 0.004538 | 5.3e-05 | 0.00093 | 0.32925 | 1.331 | 2.5481 | 99.51 | 20.89 | 37.341 | 70.2952 | 100.2985 | 67.7628 | 0.00928 | 1.0356 | 0.024778 | 23.9269 | -0.012162 | 0.14266 | 0.00019 | 0.014361 | 0.014361 | 0.40829
2019-03-21 12:07 | 1553184447.3752 | 1334 | 0.20954 | 0.000454 | -4e-06 | 100.8741 | 100.8843 | 124.0081 | 21.128 | 0.002012 | 0.001944 | -152.1548 | 0.23924 | 8.4025 | 0.00326 | 0.004091 | 0.003921 | 5.3e-05 | 0.000932 | 0.30917 | 1.3254 | 2.5481 | 99.51 | 20.89 | 37.341 | 70.1018 | 100.1503 | 68.4781 | 0.00928 | 1.0334 | 0.024762 | 23.9615 | -0.000903 | 0.14233 | 0.000333 | 0.024994 | 0.024994 | 0.23404
2019-03-21 12:07 | 1553184448.3744 | 1335 | 0.20955 | 0.00046 | -4e-06 | 100.8065 | 100.98 | 123.668 | 21.1275 | 0.002019 | 0.001947 | -152.156 | 0.23976 | 8.4018 | 0.003223 | 0.004132 | 0.003956 | 5.3e-05 | 0.000935 | 0.32997 | 1.3282 | 2.5479 | 99.51 | 20.89 | 37.341 | 70.2754 | 99.2463 | 68.551 | 0.00928 | 1.0332 | 0.024726 | 23.9311 | -0.015384 | 0.14264 | 0.000312 | 0.023406 | 0.023406 | 0.25046
2019-03-21 12:07 | 1553184449.3744 | 1336 | 0.20956 | 0.000457 | -4e-06 | 100.8276 | 100.8502 | 123.6401 | 21.1281 | 0.002003 | 0.001909 | -152.1432 | 0.23848 | 8.4026 | 0.003241 | 0.005076 | 0.004951 | 5.3e-05 | 0.000933 | 0.29927 | 1.3155 | 2.548 | 99.51 | 20.89 | 37.339 | 69.8931 | 99.6971 | 67.6352 | 0.00928 | 1.036 | 0.024756 | 23.8963 | -0.002005 | 0.14187 | 0.000246 | 0.018574 | 0.018574 | 0.31393
2019-03-21 12:07 | 1553184450.3755 | 1337 | 0.20957 | 0.000457 | -4e-06 | 100.8193 | 100.8931 | 123.506 | 21.1266 | 0.002083 | 0.001921 | -152.1405 | 0.23996 | 8.4025 | 0.003284 | 0.004107 | 0.003891 | 5.1e-05 | 0.000935 | 0.33804 | 1.3178 | 2.5479 | 99.51 | 20.89 | 37.338 | 71.8033 | 100.7305 | 67.9128 | 0.009279 | 1.0351 | 0.024733 | 23.8931 | -0.006551 | 0.14275 | 0.000158 | 0.01197 | 0.01197 | 0.49015
2019-03-21 12:08 | 1553184451.3758 | 1338 | 0.20958 | 0.000452 | -4e-06 | 100.788 | 100.9708 | 123.3255 | 21.1275 | 0.001969 | 0.001944 | -152.1208 | 0.24025 | 8.4045 | 0.003237 | 0.004192 | 0.004114 | 5.2e-05 | 0.000932 | 0.31188 | 1.3083 | 2.5478 | 99.5 | 20.89 | 37.337 | 69.083 | 99.5945 | 68.4645 | 0.00928 | 1.0335 | 0.024694 | 23.8949 | -0.016208 | 0.14292 | 0.000141 | 0.010733 | 0.010733 | 0.54731
2019-03-21 12:08 | 1553184452.3756 | 1339 | 0.20954 | 0.000456 | -4e-06 | 100.7564 | 101.0335 | 123.3466 | 21.1278 | 0.001961 | 0.001937 | -152.1498 | 0.23975 | 8.402 | 0.003231 | 0.004747 | 0.00459 | 5.1e-05 | 0.000931 | 0.32117 | 1.3172 | 2.548 | 99.5 | 20.88 | 37.337 | 68.8914 | 99.4458 | 68.3126 | 0.009274 | 1.0339 | 0.024702 | 23.8917 | -0.024561 | 0.14263 | 0.000336 | 0.025152 | 0.025152 | 0.23307
2019-03-21 12:08 | 1553184453.3756 | 1340 | 0.20958 | 0.000454 | -3e-06 | 100.8503 | 100.8829 | 123.3157 | 21.1274 | 0.002037 | 0.001912 | -152.1565 | 0.2401 | 8.4002 | 0.003206 | 0.00476 | 0.004605 | 5.5e-05 | 0.00093 | 0.30038 | 1.3186 | 2.5481 | 99.5 | 20.88 | 37.336 | 70.7008 | 98.8312 | 67.703 | 0.009274 | 1.0358 | 0.024721 | 23.8673 | -0.002896 | 0.14284 | 0.00012 | 0.009108 | 0.009108 | 0.64452
2019-03-21 12:08 | 1553184454.3746 | 1341 | 0.20959 | 0.000454 | -4e-06 | 100.8151 | 100.9636 | 122.9774 | 21.1277 | 0.002054 | 0.001929 | -152.1418 | 0.23992 | 8.3995 | 0.003264 | 0.004401 | 0.004243 | 5.4e-05 | 0.000932 | 0.29789 | 1.3114 | 2.5479 | 99.5 | 20.88 | 37.335 | 71.1208 | 100.2473 | 68.1019 | 0.009274 | 1.0346 | 0.024673 | 23.8485 | -0.013169 | 0.14273 | 7.2e-05 | 0.005533 | 0.005533 | 1.0603
2019-03-21 12:08 | 1553184455.3747 | 1342 | 0.20961 | 0.000462 | -4e-06 | 100.8166 | 101.013 | 122.8082 | 21.1273 | 0.00198 | 0.001897 | -152.1628 | 0.24008 | 8.4013 | 0.003258 | 0.0039 | 0.003772 | 5.3e-05 | 0.000934 | 0.29878 | 1.3394 | 2.5481 | 99.5 | 20.88 | 37.333 | 69.3455 | 100.0909 | 67.3309 | 0.009273 | 1.0369 | 0.024684 | 23.8052 | -0.017411 | 0.14283 | -5.9e-05 | -0.004134 | -0.004134 | -1.4201
2019-03-21 12:08 | 1553184456.3762 | 1343 | 0.2096 | 0.000463 | -4e-06 | 100.8211 | 100.9906 | 123.0338 | 21.1273 | 0.002009 | 0.001931 | -152.1647 | 0.2395 | 8.4013 | 0.003254 | 0.004356 | 0.004117 | 5.3e-05 | 0.000937 | 0.28622 | 1.3188 | 2.548 | 99.5 | 20.88 | 37.331 | 70.0278 | 100.0029 | 68.1585 | 0.009273 | 1.0344 | 0.024676 | 23.856 | -0.015023 | 0.14248 | 1.2e-05 | 0.001129 | 0.001129 | 5.189
2019-03-21 12:08 | 1553184457.3758 | 1344 | 0.20954 | 0.000457 | -4e-06 | 100.8086 | 100.8431 | 123.2234 | 21.1279 | 0.002005 | 0.001925 | -152.1697 | 0.24068 | 8.4025 | 0.003257 | 0.00392 | 0.003733 | 5.2e-05 | 0.000937 | 0.33373 | 1.3222 | 2.548 | 99.5 | 20.88 | 37.331 | 69.9311 | 100.0859 | 68.0141 | 0.009273 | 1.0348 | 0.024701 | 23.8693 | -0.003063 | 0.14318 | 0.000366 | 0.027382 | 0.027382 | 0.21491
2019-03-21 12:08 | 1553184458.3749 | 1345 | 0.20958 | 0.000454 | -4e-06 | 100.7463 | 100.8799 | 123.13 | 21.1276 | 0.002 | 0.001957 | -152.1557 | 0.2402 | 8.4023 | 0.003288 | 0.004119 | 0.003924 | 5.3e-05 | 0.000935 | 0.33356 | 1.305 | 2.5478 | 99.5 | 20.88 | 37.332 | 69.8207 | 100.8146 | 68.7852 | 0.009273 | 1.0325 | 0.024663 | 23.8872 | -0.011852 | 0.14289 | 0.000102 | 0.007766 | 0.007766 | 0.75626
2019-03-21 12:08 | 1553184459.375 | 1346 | 0.20959 | 0.000459 | -4e-06 | 100.8071 | 101.0062 | 123.2126 | 21.127 | 0.002105 | 0.001937 | -152.1581 | 0.24089 | 8.4016 | 0.003341 | 0.003967 | 0.003846 | 5.2e-05 | 0.000933 | 0.32748 | 1.3317 | 2.5479 | 99.5 | 20.88 | 37.333 | 72.3518 | 102.1112 | 68.3017 | 0.009273 | 1.034 | 0.024689 | 23.8783 | -0.017653 | 0.1433 | 6.3e-05 | 0.004938 | 0.004938 | 1.1927
2019-03-21 12:08 | 1553184460.3759 | 1347 | 0.20963 | 0.000459 | -4e-06 | 100.8261 | 100.9858 | 123.2354 | 21.1278 | 0.001998 | 0.001939 | -152.1706 | 0.24007 | 8.4053 | 0.00328 | 0.004377 | 0.004249 | 5.4e-05 | 0.000932 | 0.3183 | 1.3337 | 2.5478 | 99.5 | 20.88 | 37.334 | 69.7695 | 100.6298 | 68.3477 | 0.009273 | 1.0338 | 0.02469 | 23.8821 | -0.014152 | 0.14282 | -0.000221 | -0.016124 | -0.016124 | -0.36405
2019-03-21 12:08 | 1553184461.3764 | 1348 | 0.2096 | 0.000456 | -4e-06 | 100.8199 | 101.0171 | 123.6217 | 21.1279 | 0.002043 | 0.001916 | -152.1697 | 0.23971 | 8.4017 | 0.00331 | 0.004383 | 0.004168 | 5.3e-05 | 0.000932 | 0.32189 | 1.313 | 2.5479 | 99.5 | 20.88 | 37.336 | 70.8611 | 101.3675 | 67.7934 | 0.009274 | 1.0355 | 0.024748 | 23.9001 | -0.017482 | 0.14261 | -1e-06 | 0.000199 | 0.000199 | 29.4833
2019-03-21 12:08 | 1553184462.376 | 1349 | 0.2096 | 0.000458 | -4e-06 | 100.8443 | 100.9255 | 123.691 | 21.1274 | 0.001997 | 0.001907 | -152.1682 | 0.23978 | 8.4022 | 0.003342 | 0.004793 | 0.004605 | 5e-05 | 0.000931 | 0.31725 | 1.3111 | 2.548 | 99.5 | 20.88 | 37.337 | 69.7566 | 102.1297 | 67.5931 | 0.009274 | 1.0361 | 0.024763 | 23.8998 | -0.007199 | 0.14264 | -2.1e-05 | -0.001309 | -0.001309 | -4.4772
2019-03-21 12:08 | 1553184463.3751 | 1350 | 0.20958 | 0.000455 | -4e-06 | 100.8148 | 101.0106 | 123.8222 | 21.1281 | 0.002022 | 0.001925 | -152.1488 | 0.24004 | 8.4039 | 0.003291 | 0.004171 | 0.004007 | 5.5e-05 | 0.000934 | 0.2943 | 1.3561 | 2.548 | 99.5 | 20.88 | 37.337 | 70.3499 | 100.899 | 68.0174 | 0.009274 | 1.0348 | 0.02476 | 23.9273 | -0.017358 | 0.1428 | 0.000144 | 0.010939 | 0.010939 | 0.53654
2019-03-21 12:08 | 1553184464.3742 | 1351 | 0.20957 | 0.00046 | -4e-06 | 100.8473 | 100.9415 | 124.2736 | 21.1273 | 0.00196 | 0.001931 | -152.146 | 0.23973 | 8.4027 | 0.003283 | 0.004697 | 0.004529 | 5.2e-05 | 0.000934 | 0.30143 | 1.3599 | 2.5479 | 99.51 | 20.88 | 37.339 | 68.8527 | 100.6958 | 68.15 | 0.009274 | 1.0344 | 0.024801 | 23.9756 | -0.008353 | 0.14262 | 0.000158 | 0.012008 | 0.012008 | 0.48814
2019-03-21 12:08 | 1553184465.3746 | 1352 | 0.20956 | 0.000457 | -4e-06 | 100.7823 | 101.0985 | 124.4949 | 21.1272 | 0.001977 | 0.001952 | -152.1318 | 0.23956 | 8.4034 | 0.003258 | 0.005006 | 0.004852 | 5.4e-05 | 0.000932 | 0.30613 | 1.3024 | 2.548 | 99.51 | 20.88 | 37.34 | 69.2612 | 100.0995 | 68.6615 | 0.009274 | 1.0329 | 0.024804 | 24.0149 | -0.028018 | 0.14252 | 0.000245 | 0.018452 | 0.018452 | 0.31745
2019-03-21 12:08 | 1553184466.3743 | 1353 | 0.2096 | 0.000458 | -4e-06 | 100.8544 | 100.8867 | 124.5562 | 21.1278 | 0.001966 | 0.001884 | -152.1346 | 0.23972 | 8.4013 | 0.00327 | 0.004607 | 0.00444 | 5.2e-05 | 0.000935 | 0.29623 | 1.2853 | 2.548 | 99.51 | 20.88 | 37.344 | 68.9898 | 100.3794 | 67.0311 | 0.009275 | 1.0378 | 0.02487 | 23.9634 | -0.002858 | 0.14261 | 2.7e-05 | 0.002238 | 0.002238 | 2.6194
2019-03-21 12:08 | 1553184467.3744 | 1354 | 0.2096 | 0.000457 | -4e-06 | 100.8115 | 100.8049 | 124.685 | 21.1276 | 0.002013 | 0.001923 | -152.1364 | 0.24052 | 8.4006 | 0.00332 | 0.004227 | 0.004054 | 5.6e-05 | 0.000928 | 0.32173 | 1.2776 | 2.5479 | 99.51 | 20.88 | 37.347 | 70.131 | 101.6044 | 67.96 | 0.009276 | 1.035 | 0.024849 | 24.0085 | 0.000586 | 0.14308 | 1e-05 | 0.000936 | 0.000936 | 6.2845
2019-03-21 12:08 | 1553184468.3748 | 1355 | 0.2096 | 0.000458 | -4e-06 | 100.8434 | 100.9183 | 124.6745 | 21.1274 | 0.001964 | 0.001938 | -152.1436 | 0.23923 | 8.404 | 0.003271 | 0.004845 | 0.004658 | 5.5e-05 | 0.000925 | 0.28407 | 1.2864 | 2.5479 | 99.51 | 20.88 | 37.351 | 68.9533 | 100.4213 | 68.3319 | 0.009277 | 1.0339 | 0.024834 | 24.0206 | -0.006644 | 0.14232 | -1.5e-05 | -0.000859 | -0.000859 | -6.8133
2019-03-21 12:08 | 1553184469.3745 | 1356 | 0.2096 | 0.000459 | -4e-06 | 100.8142 | 100.8298 | 124.2921 | 21.1262 | 0.001996 | 0.00193 | -152.1436 | 0.23965 | 8.4027 | 0.003313 | 0.004479 | 0.00432 | 5.5e-05 | 0.000927 | 0.30973 | 1.2927 | 2.5479 | 99.51 | 20.87 | 37.353 | 69.729 | 101.4359 | 68.1272 | 0.009272 | 1.0345 | 0.024803 | 23.9765 | -0.001382 | 0.14257 | -1.9e-05 | -0.001206 | -0.001206 | -4.858
2019-03-21 12:08 | 1553184470.3755 | 1357 | 0.20958 | 0.000453 | -4e-06 | 100.8297 | 100.8839 | 124.2407 | 21.1267 | 0.002042 | 0.00194 | -152.1278 | 0.23876 | 8.4021 | 0.003162 | 0.00479 | 0.004555 | 5.6e-05 | 0.000929 | 0.30685 | 1.3071 | 2.548 | 99.51 | 20.87 | 37.355 | 70.8196 | 97.7871 | 68.3802 | 0.009272 | 1.0337 | 0.024789 | 23.9805 | -0.00481 | 0.14204 | 0.000129 | 0.009855 | 0.009855 | 0.5924
2019-03-21 12:08 | 1553184471.3755 | 1358 | 0.20963 | 0.000454 | -4e-06 | 100.8118 | 101.0205 | 123.8884 | 21.1273 | 0.002032 | 0.001856 | -152.1266 | 0.24023 | 8.4037 | 0.00322 | 0.004498 | 0.004404 | 5.1e-05 | 0.00093 | 0.30093 | 1.3522 | 2.5478 | 99.51 | 20.87 | 37.359 | 70.5909 | 99.1738 | 66.3587 | 0.009273 | 1.0399 | 0.024827 | 23.8755 | -0.018497 | 0.14291 | -0.000211 | -0.015462 | -0.015462 | -0.37987
2019-03-21 12:08 | 1553184472.3755 | 1359 | 0.20963 | 0.000455 | -4e-06 | 100.8148 | 100.9211 | 123.891 | 21.127 | 0.002006 | 0.001903 | -152.1047 | 0.2399 | 8.4014 | 0.003222 | 0.004465 | 0.004288 | 5.1e-05 | 0.000927 | 0.3402 | 1.3419 | 2.548 | 99.51 | 20.87 | 37.361 | 69.9622 | 99.2354 | 67.4913 | 0.009274 | 1.0364 | 0.024786 | 23.9155 | -0.009423 | 0.14272 | -0.000159 | -0.011642 | -0.011642 | -0.50386
2019-03-21 12:08 | 1553184473.3757 | 1360 | 0.2096 | 0.000459 | -4e-06 | 100.7892 | 101.0454 | 123.7436 | 21.1275 | 0.00203 | 0.001937 | -152.1118 | 0.23874 | 8.4048 | 0.003234 | 0.004744 | 0.00456 | 5.8e-05 | 0.000929 | 0.31022 | 1.3336 | 2.548 | 99.51 | 20.87 | 37.364 | 70.5397 | 99.5098 | 68.3101 | 0.009274 | 1.0339 | 0.024742 | 23.93 | -0.022708 | 0.14203 | 2.9e-05 | 0.002379 | 0.002379 | 2.4536
2019-03-21 12:08 | 1553184474.3751 | 1361 | 0.20963 | 0.000458 | -4e-06 | 100.8452 | 100.8656 | 123.5333 | 21.1291 | 0.002044 | 0.001872 | -152.1261 | 0.23921 | 8.4012 | 0.003221 | 0.003934 | 0.003781 | 5.6e-05 | 0.000929 | 0.30486 | 1.3265 | 2.548 | 99.51 | 20.87 | 37.366 | 70.8709 | 99.2084 | 66.7351 | 0.009275 | 1.0387 | 0.024778 | 23.8544 | -0.001805 | 0.14231 | -0.000169 | -0.012327 | -0.012327 | -0.47446
2019-03-21 12:08 | 1553184475.3757 | 1362 | 0.20962 | 0.000458 | -4e-06 | 100.8106 | 100.791 | 123.8818 | 21.1273 | 0.002004 | 0.001931 | -152.1536 | 0.23991 | 8.4022 | 0.003231 | 0.004918 | 0.004781 | 5.4e-05 | 0.000931 | 0.30829 | 1.3239 | 2.5479 | 99.51 | 20.87 | 37.367 | 69.9243 | 99.4548 | 68.1616 | 0.009275 | 1.0344 | 0.024761 | 23.9382 | 0.001745 | 0.14272 | -0.000123 | -0.008876 | -0.008876 | -0.66088
2019-03-21 12:08 | 1553184476.3759 | 1363 | 0.20962 | 0.000459 | -3e-06 | 100.8214 | 100.9558 | 123.9022 | 21.1267 | 0.002001 | 0.001883 | -152.1666 | 0.23969 | 8.403 | 0.003278 | 0.004675 | 0.004503 | 5.7e-05 | 0.00093 | 0.2995 | 1.3212 | 2.5479 | 99.51 | 20.87 | 37.37 | 69.8434 | 100.5854 | 67.0081 | 0.009276 | 1.0379 | 0.024805 | 23.8996 | -0.01192 | 0.14259 | -0.000128 | -0.009335 | -0.009335 | -0.62777
2019-03-21 12:08 | 1553184477.3753 | 1364 | 0.20963 | 0.000459 | -4e-06 | 100.8032 | 100.9705 | 123.8682 | 21.1271 | 0.001962 | 0.001901 | -152.166 | 0.23922 | 8.4017 | 0.003246 | 0.004532 | 0.004313 | 5.5e-05 | 0.000929 | 0.28799 | 1.3651 | 2.548 | 99.51 | 20.87 | 37.372 | 68.8939 | 99.8184 | 67.4391 | 0.009276 | 1.0366 | 0.024786 | 23.9115 | -0.01483 | 0.14231 | -0.000168 | -0.012254 | -0.012254 | -0.47732
2019-03-21 12:08 | 1553184478.3759 | 1365 | 0.20963 | 0.000459 | -4e-06 | 100.7761 | 100.854 | 123.9543 | 21.1281 | 0.001974 | 0.001916 | -152.1671 | 0.24017 | 8.4018 | 0.003255 | 0.004511 | 0.004368 | 5.5e-05 | 0.000928 | 0.32189 | 1.3663 | 2.548 | 99.51 | 20.87 | 37.373 | 69.1912 | 100.0281 | 67.7873 | 0.009277 | 1.0355 | 0.024782 | 23.932 | -0.006911 | 0.14288 | -0.00017 | -0.012435 | -0.012435 | -0.47226
2019-03-21 12:08 | 1553184479.3744 | 1366 | 0.20962 | 0.000456 | -4e-06 | 100.7508 | 101.045 | 124.2729 | 21.1269 | 0.002022 | 0.001921 | -152.1553 | 0.23989 | 8.4027 | 0.003272 | 0.004602 | 0.004424 | 5.4e-05 | 0.000933 | 0.31526 | 1.328 | 2.548 | 99.51 | 20.87 | 37.373 | 70.3426 | 100.4272 | 67.9281 | 0.009277 | 1.0351 | 0.024809 | 23.9677 | -0.026089 | 0.14271 | -0.000132 | -0.009618 | -0.009618 | -0.60983
2019-03-21 12:08 | 1553184480.3754 | 1367 | 0.20962 | 0.000453 | -4e-06 | 100.8428 | 100.916 | 124.3318 | 21.1266 | 0.00193 | 0.001904 | -152.1728 | 0.23937 | 8.3997 | 0.0032 | 0.004489 | 0.004365 | 5.6e-05 | 0.000935 | 0.32997 | 1.2846 | 2.548 | 99.51 | 20.88 | 37.373 | 68.1283 | 98.684 | 67.5003 | 0.009282 | 1.0364 | 0.02483 | 23.9583 | -0.006486 | 0.1424 | -8.9e-05 | -0.006412 | -0.006412 | -0.91277
2019-03-21 12:08 | 1553184481.3752 | 1368 | 0.20962 | 0.000454 | -3e-06 | 100.796 | 100.9957 | 124.2743 | 21.1267 | 0.001922 | 0.001853 | -152.1801 | 0.23965 | 8.4013 | 0.003256 | 0.004814 | 0.004604 | 5.5e-05 | 0.000933 | 0.3366 | 1.3114 | 2.5478 | 99.51 | 20.88 | 37.374 | 67.9428 | 100.048 | 66.2865 | 0.009283 | 1.0401 | 0.024869 | 23.9101 | -0.017697 | 0.14257 | -0.000123 | -0.00896 | -0.00896 | -0.65398
2019-03-21 12:08 | 1553184482.3742 | 1369 | 0.20961 | 0.000456 | -4e-06 | 100.8422 | 100.7181 | 124.16 | 21.1264 | 0.002016 | 0.001817 | -152.2247 | 0.23953 | 8.4014 | 0.00326 | 0.005235 | 0.005069 | 5.6e-05 | 0.000936 | 0.32173 | 1.345 | 2.548 | 99.51 | 20.88 | 37.374 | 70.2017 | 100.1558 | 65.4183 | 0.009283 | 1.0428 | 0.024889 | 23.8685 | 0.011018 | 0.1425 | -8.2e-05 | -0.005884 | -0.005884 | -0.9954
2019-03-21 12:08 | 1553184483.3745 | 1370 | 0.20961 | 0.000457 | -4e-06 | 100.7338 | 100.9551 | 124.4862 | 21.1265 | 0.001996 | 0.001875 | -152.2206 | 0.23976 | 8.4045 | 0.003317 | 0.004876 | 0.004767 | 5.2e-05 | 0.000933 | 0.30182 | 1.3216 | 2.5479 | 99.51 | 20.88 | 37.373 | 69.7324 | 101.5327 | 66.8182 | 0.009282 | 1.0385 | 0.02487 | 23.9492 | -0.019634 | 0.14264 | -5.4e-05 | -0.003778 | -0.003778 | -1.5518
2019-03-21 12:08 | 1553184484.3742 | 1371 | 0.20961 | 0.000456 | -4e-06 | 100.7582 | 100.8281 | 124.6386 | 21.1273 | 0.002001 | 0.001938 | -152.2174 | 0.24011 | 8.4018 | 0.003336 | 0.004477 | 0.004337 | 5.7e-05 | 0.000936 | 0.30414 | 1.322 | 2.548 | 99.5 | 20.88 | 37.372 | 69.8353 | 101.9812 | 68.3193 | 0.009283 | 1.0339 | 0.024831 | 24.0167 | -0.006202 | 0.14284 | -7.9e-05 | -0.005688 | -0.005688 | -1.0321
2019-03-21 12:08 | 1553184485.3743 | 1372 | 0.20961 | 0.000457 | -4e-06 | 100.7227 | 100.9051 | 124.3069 | 21.1268 | 0.002026 | 0.001953 | -152.2117 | 0.24063 | 8.4015 | 0.003426 | 0.00485 | 0.004691 | 5.3e-05 | 0.000934 | 0.32477 | 1.32 | 2.548 | 99.5 | 20.88 | 37.371 | 70.4369 | 104.165 | 68.6792 | 0.009283 | 1.0328 | 0.024785 | 23.9974 | -0.016176 | 0.14315 | -6.4e-05 | -0.004506 | -0.004506 | -1.3057
2019-03-21 12:08 | 1553184486.3758 | 1373 | 0.20961 | 0.000461 | -4e-06 | 100.7665 | 100.9589 | 124.2426 | 21.1275 | 0.002007 | 0.001903 | -152.1987 | 0.23855 | 8.4021 | 0.003373 | 0.005241 | 0.005033 | 5.3e-05 | 0.000935 | 0.28799 | 1.2992 | 2.5479 | 99.5 | 20.88 | 37.371 | 69.9766 | 102.8838 | 67.4894 | 0.009283 | 1.0364 | 0.024822 | 23.9494 | -0.017056 | 0.14191 | -8.3e-05 | -0.005955 | -0.005955 | -0.9794
2019-03-21 12:08 | 1553184487.3754 | 1374 | 0.2096 | 0.000459 | -4e-06 | 100.7156 | 100.8243 | 123.9572 | 21.1277 | 0.002036 | 0.001913 | -152.2067 | 0.24023 | 8.4041 | 0.003363 | 0.00427 | 0.004163 | 5.5e-05 | 0.000934 | 0.32244 | 1.3145 | 2.5479 | 99.5 | 20.88 | 37.371 | 70.6949 | 102.6334 | 67.7384 | 0.009283 | 1.0357 | 0.024784 | 23.9306 | -0.009653 | 0.14291 | -5e-06 | -0.00016 | -0.00016 | -36.7891
2019-03-21 12:08 | 1553184488.3755 | 1375 | 0.20962 | 0.000454 | -4e-06 | 100.7227 | 100.7899 | 124.0386 | 21.1268 | 0.001996 | 0.001853 | -152.1993 | 0.24039 | 8.4018 | 0.003425 | 0.004951 | 0.004698 | 5.3e-05 | 0.000937 | 0.33644 | 1.2734 | 2.548 | 99.5 | 20.88 | 37.37 | 69.7164 | 104.1376 | 66.2918 | 0.009282 | 1.0401 | 0.024845 | 23.8876 | -0.005966 | 0.14301 | -0.000115 | -0.008335 | -0.008335 | -0.70518
2019-03-21 12:08 | 1553184489.3749 | 1376 | 0.20971 | 0.000457 | -4e-06 | 100.7624 | 100.8594 | 124.0811 | 21.1283 | 0.002005 | 0.00191 | -152.2109 | 0.23889 | 8.3985 | 0.003454 | 0.004095 | 0.003997 | 5.3e-05 | 0.000938 | 0.31758 | 1.2863 | 2.5478 | 99.5 | 20.87 | 37.368 | 69.943 | 104.8569 | 67.6461 | 0.009276 | 1.0359 | 0.0248 | 23.9393 | -0.008611 | 0.14211 | -0.000649 | -0.04812 | -0.04812 | -0.12138
2019-03-21 12:08 | 1553184490.3756 | 1377 | 0.20964 | 0.000455 | -4e-06 | 100.7254 | 100.8989 | 124.0773 | 21.1278 | 0.001978 | 0.001921 | -152.209 | 0.2384 | 8.4034 | 0.003426 | 0.004549 | 0.004398 | 5.3e-05 | 0.000938 | 0.33467 | 1.315 | 2.5478 | 99.5 | 20.87 | 37.367 | 69.2844 | 104.173 | 67.9212 | 0.009276 | 1.0351 | 0.024789 | 23.9486 | -0.015394 | 0.14182 | -0.000243 | -0.017847 | -0.017847 | -0.32661
2019-03-21 12:08 | 1553184491.3746 | 1378 | 0.20963 | 0.000458 | -4e-06 | 100.6399 | 100.8005 | 124.305 | 21.1276 | 0.001961 | 0.001915 | -152.1948 | 0.23899 | 8.4037 | 0.003479 | 0.005016 | 0.00493 | 5.4e-05 | 0.000939 | 0.31974 | 1.3568 | 2.548 | 99.5 | 20.87 | 37.367 | 68.8889 | 105.4518 | 67.7744 | 0.009276 | 1.0356 | 0.024817 | 23.9654 | -0.014261 | 0.14217 | -0.000157 | -0.011469 | -0.011469 | -0.50951
2019-03-21 12:08 | 1553184492.3761 | 1379 | 0.20963 | 0.00046 | -3e-06 | 100.7144 | 100.9068 | 124.4153 | 21.1275 | 0.001998 | 0.001861 | -152.1668 | 0.24026 | 8.4029 | 0.003473 | 0.004481 | 0.004318 | 5.3e-05 | 0.000943 | 0.33804 | 1.3207 | 2.5478 | 99.5 | 20.87 | 37.366 | 69.7584 | 105.3062 | 66.4837 | 0.009276 | 1.0395 | 0.024876 | 23.9306 | -0.017068 | 0.14293 | -0.000191 | -0.014084 | -0.014084 | -0.41711
2019-03-21 12:08 | 1553184493.3757 | 1380 | 0.2096 | 0.000454 | -4e-06 | 100.7239 | 100.8134 | 124.1508 | 21.1275 | 0.001988 | 0.001821 | -152.1585 | 0.24045 | 8.4017 | 0.003504 | 0.005184 | 0.005 | 5.4e-05 | 0.000939 | 0.3115 | 1.3366 | 2.548 | 99.5 | 20.88 | 37.366 | 69.5198 | 106.0655 | 65.506 | 0.009281 | 1.0425 | 0.024885 | 23.8707 | -0.007945 | 0.14305 | -1.2e-05 | -0.000686 | -0.000686 | -8.5649
2019-03-21 12:08 | 1553184494.3764 | 1381 | 0.20961 | 0.000458 | -4e-06 | 100.771 | 100.998 | 124.2393 | 21.1264 | 0.001942 | 0.001879 | -152.1755 | 0.2395 | 8.4015 | 0.003465 | 0.005248 | 0.00512 | 5.1e-05 | 0.000944 | 0.36481 | 1.3633 | 2.5479 | 99.5 | 20.87 | 37.364 | 68.4346 | 105.1005 | 66.904 | 0.009275 | 1.0382 | 0.024843 | 23.9285 | -0.020129 | 0.14248 | -5.3e-05 | -0.00375 | -0.00375 | -1.5617
2019-03-21 12:08 | 1553184495.3753 | 1382 | 0.20961 | 0.000458 | -4e-06 | 100.732 | 100.7177 | 124.474 | 21.127 | 0.001951 | 0.001903 | -152.1671 | 0.23967 | 8.4017 | 0.003397 | 0.004327 | 0.004095 | 4.9e-05 | 0.000944 | 0.34092 | 1.2877 | 2.5479 | 99.5 | 20.87 | 37.363 | 68.6402 | 103.4507 | 67.4804 | 0.009275 | 1.0364 | 0.024845 | 23.9713 | 0.001265 | 0.14258 | -5.6e-05 | -0.003886 | -0.003886 | -1.508
2019-03-21 12:08 | 1553184496.3739 | 1383 | 0.20958 | 0.000458 | -4e-06 | 100.7788 | 100.886 | 124.4775 | 21.1274 | 0.002001 | 0.001866 | -152.1833 | 0.23978 | 8.4023 | 0.003422 | 0.004993 | 0.004838 | 5e-05 | 0.00094 | 0.29031 | 1.2719 | 2.5478 | 99.5 | 20.87 | 37.363 | 69.84 | 104.0714 | 66.6024 | 0.009275 | 1.0391 | 0.024877 | 23.9408 | -0.009512 | 0.14264 | 0.000112 | 0.008617 | 0.008617 | 0.68037
2019-03-21 12:08 | 1553184497.3743 | 1384 | 0.20961 | 0.000458 | -4e-06 | 100.763 | 100.9606 | 124.8085 | 21.1267 | 0.001953 | 0.001891 | -152.18 | 0.23922 | 8.4008 | 0.003473 | 0.004576 | 0.004469 | 5.3e-05 | 0.000941 | 0.33019 | 1.2777 | 2.548 | 99.5 | 20.87 | 37.363 | 68.6985 | 105.3092 | 67.1906 | 0.009275 | 1.0373 | 0.024889 | 23.9933 | -0.017524 | 0.14231 | -4.8e-05 | -0.003301 | -0.003301 | -1.7717
2019-03-21 12:08 | 1553184498.375 | 1385 | 0.20958 | 0.000455 | -4e-06 | 100.7856 | 100.873 | 124.8813 | 21.1272 | 0.001947 | 0.001863 | -152.1941 | 0.23922 | 8.4012 | 0.003408 | 0.005368 | 0.005065 | 5.1e-05 | 0.000945 | 0.27815 | 1.293 | 2.548 | 99.5 | 20.88 | 37.364 | 68.5392 | 103.7259 | 66.5358 | 0.009281 | 1.0393 | 0.02492 | 23.9772 | -0.007757 | 0.14231 | 0.000161 | 0.012286 | 0.012286 | 0.47606
2019-03-21 12:08 | 1553184499.3741 | 1386 | 0.20959 | 0.000457 | -4e-06 | 100.7642 | 100.9657 | 124.7386 | 21.1271 | 0.001966 | 0.001893 | -152.1997 | 0.23947 | 8.4019 | 0.003426 | 0.004884 | 0.0047 | 5.1e-05 | 0.000946 | 0.31116 | 1.3124 | 2.5479 | 99.5 | 20.87 | 37.366 | 68.9953 | 104.1716 | 67.2341 | 0.009276 | 1.0372 | 0.02488 | 23.9881 | -0.017871 | 0.14246 | 6.5e-05 | 0.00508 | 0.00508 | 1.1526
2019-03-21 12:08 | 1553184500.3755 | 1387 | 0.20954 | 0.000457 | -4e-06 | 100.7695 | 100.7041 | 124.7681 | 21.1268 | 0.00197 | 0.001912 | -152.2034 | 0.2403 | 8.4045 | 0.003441 | 0.004978 | 0.004823 | 5.4e-05 | 0.000948 | 0.31614 | 1.301 | 2.548 | 99.5 | 20.87 | 37.368 | 69.0955 | 104.5188 | 67.7072 | 0.009276 | 1.0358 | 0.024866 | 24.0076 | 0.005808 | 0.14296 | 0.000375 | 0.028265 | 0.028265 | 0.20787
2019-03-21 12:08 | 1553184501.3752 | 1388 | 0.20957 | 0.000454 | -4e-06 | 100.7439 | 100.8989 | 124.7146 | 21.1276 | 0.001935 | 0.00188 | -152.2225 | 0.23952 | 8.4019 | 0.003418 | 0.004602 | 0.004409 | 5.2e-05 | 0.000952 | 0.37465 | 1.3181 | 2.5481 | 99.5 | 20.88 | 37.371 | 68.2672 | 103.9748 | 66.9443 | 0.009283 | 1.0381 | 0.024889 | 23.9756 | -0.013754 | 0.14249 | 0.000186 | 0.014157 | 0.014157 | 0.41369
2019-03-21 12:08 | 1553184502.3743 | 1389 | 0.20956 | 0.000458 | -3e-06 | 100.7779 | 101.0532 | 124.8707 | 21.1273 | 0.001966 | 0.001861 | -152.2104 | 0.23894 | 8.4031 | 0.003418 | 0.004619 | 0.004502 | 5.2e-05 | 0.000953 | 0.32908 | 1.318 | 2.5481 | 99.5 | 20.88 | 37.376 | 69.0056 | 103.9736 | 66.4664 | 0.009284 | 1.0395 | 0.024922 | 23.9737 | -0.024406 | 0.14215 | 0.000221 | 0.016777 | 0.016777 | 0.34824
2019-03-21 12:08 | 1553184503.3749 | 1390 | 0.20957 | 0.000459 | -4e-06 | 100.7484 | 100.885 | 124.877 | 21.1273 | 0.00202 | 0.001835 | -152.2382 | 0.23879 | 8.4047 | 0.003454 | 0.005075 | 0.004959 | 5.1e-05 | 0.000958 | 0.35275 | 1.3166 | 2.5479 | 99.5 | 20.88 | 37.378 | 70.3062 | 104.8332 | 65.8483 | 0.009284 | 1.0414 | 0.024945 | 23.9525 | -0.012119 | 0.14205 | 0.000209 | 0.015914 | 0.015914 | 0.36688
2019-03-21 12:08 | 1553184504.3743 | 1391 | 0.20955 | 0.00046 | -4e-06 | 100.7966 | 101.0706 | 124.5522 | 21.128 | 0.002028 | 0.001886 | -152.2334 | 0.23876 | 8.4022 | 0.003514 | 0.004995 | 0.004736 | 5.7e-05 | 0.000954 | 0.3194 | 1.3733 | 2.5481 | 99.5 | 20.88 | 37.383 | 70.501 | 106.3003 | 67.0726 | 0.009286 | 1.0377 | 0.024868 | 23.9645 | -0.024277 | 0.14204 | 0.000295 | 0.02228 | 0.02228 | 0.26202
2019-03-21 12:08 | 1553184505.3743 | 1392 | 0.20958 | 0.000462 | -3e-06 | 100.774 | 100.7951 | 124.2177 | 21.1285 | 0.002009 | 0.001882 | -152.239 | 0.23902 | 8.4023 | 0.00348 | 0.004621 | 0.004539 | 5e-05 | 0.000949 | 0.29822 | 1.3683 | 2.548 | 99.5 | 20.88 | 37.386 | 70.0401 | 105.4818 | 66.988 | 0.009286 | 1.0379 | 0.024837 | 23.9293 | -0.001869 | 0.14219 | 8.9e-05 | 0.006861 | 0.006861 | 0.85182
2019-03-21 12:08 | 1553184506.3748 | 1393 | 0.2096 | 0.000457 | -4e-06 | 100.7972 | 100.9527 | 124.0257 | 21.1271 | 0.00201 | 0.001855 | -152.2273 | 0.23973 | 8.4032 | 0.003435 | 0.004963 | 0.004721 | 5.1e-05 | 0.000946 | 0.29717 | 1.3466 | 2.5479 | 99.5 | 20.88 | 37.388 | 70.0581 | 104.38 | 66.3226 | 0.009287 | 1.04 | 0.024843 | 23.8874 | -0.01379 | 0.14262 | 2.7e-05 | 0.002269 | 0.002269 | 2.5838
2019-03-21 12:08 | 1553184507.3745 | 1394 | 0.20962 | 0.000456 | -3e-06 | 100.8386 | 100.9401 | 123.6785 | 21.1278 | 0.001972 | 0.001839 | -152.2444 | 0.23779 | 8.4012 | 0.003415 | 0.004406 | 0.00427 | 4.9e-05 | 0.000948 | 0.32516 | 1.3014 | 2.5479 | 99.5 | 20.88 | 37.392 | 69.1448 | 103.9064 | 65.9558 | 0.009288 | 1.0411 | 0.024821 | 23.8411 | -0.008999 | 0.14146 | -0.000118 | -0.0086 | -0.0086 | -0.6761
2019-03-21 12:08 | 1553184508.3749 | 1395 | 0.20962 | 0.000459 | -3e-06 | 100.7797 | 100.8945 | 123.6572 | 21.1287 | 0.001946 | 0.001893 | -152.2333 | 0.23988 | 8.4012 | 0.00343 | 0.004933 | 0.004609 | 4.9e-05 | 0.00095 | 0.32764 | 1.3057 | 2.5479 | 99.5 | 20.88 | 37.394 | 68.5303 | 104.2642 | 67.2471 | 0.009288 | 1.0372 | 0.024772 | 23.8844 | -0.010188 | 0.1427 | -9.9e-05 | -0.007185 | -0.007185 | -0.81629
2019-03-21 12:08 | 1553184509.3751 | 1396 | 0.20959 | 0.000458 | -4e-06 | 100.819 | 100.7753 | 123.8816 | 21.1285 | 0.001991 | 0.001904 | -152.2372 | 0.23974 | 8.4008 | 0.003344 | 0.005193 | 0.005106 | 5e-05 | 0.000949 | 0.31166 | 1.3258 | 2.5479 | 99.5 | 20.88 | 37.395 | 69.5932 | 102.1724 | 67.5201 | 0.009289 | 1.0363 | 0.024784 | 23.9156 | 0.003876 | 0.14262 | 9.3e-05 | 0.007161 | 0.007161 | 0.81851
2019-03-21 12:08 | 1553184510.3737 | 1397 | 0.20959 | 0.000459 | -4e-06 | 100.7558 | 101.0474 | 124.0806 | 21.1289 | 0.001979 | 0.00189 | -152.2359 | 0.23945 | 8.4038 | 0.003397 | 0.004433 | 0.00428 | 5.2e-05 | 0.000951 | 0.3183 | 1.2814 | 2.5479 | 99.5 | 20.88 | 37.396 | 69.3137 | 103.4698 | 67.1656 | 0.009289 | 1.0374 | 0.024817 | 23.9224 | -0.02585 | 0.14245 | 8.8e-05 | 0.006796 | 0.006796 | 0.86145
2019-03-21 12:08 | 1553184511.3748 | 1398 | 0.20958 | 0.000454 | -4e-06 | 100.9203 | 100.9882 | 124.0682 | 21.128 | 0.002036 | 0.001877 | -152.2464 | 0.23949 | 8.4003 | 0.003336 | 0.00483 | 0.004585 | 5.1e-05 | 0.000949 | 0.31868 | 1.289 | 2.5479 | 99.5 | 20.88 | 37.396 | 70.6937 | 101.992 | 66.8634 | 0.009289 | 1.0383 | 0.024827 | 23.9105 | -0.006015 | 0.14247 | 0.000162 | 0.012326 | 0.012326 | 0.47504
2019-03-21 12:09 | 1553184512.3741 | 1399 | 0.20957 | 0.000459 | -3e-06 | 100.8711 | 100.9698 | 124.1609 | 21.1279 | 0.001957 | 0.001876 | -152.2612 | 0.23927 | 8.4027 | 0.003485 | 0.003876 | 0.003725 | 5.2e-05 | 0.00095 | 0.31205 | 1.2818 | 2.5479 | 99.5 | 20.88 | 37.394 | 68.7783 | 105.5944 | 66.8486 | 0.009288 | 1.0384 | 0.024837 | 23.919 | -0.008744 | 0.14234 | 0.000187 | 0.014161 | 0.014161 | 0.41313
2019-03-21 12:09 | 1553184513.3742 | 1400 | 0.20959 | 0.000455 | -3e-06 | 100.8497 | 100.8666 | 124.0067 | 21.1281 | 0.001985 | 0.001858 | -152.2337 | 0.23939 | 8.4017 | 0.003504 | 0.004673 | 0.004482 | 5.3e-05 | 0.000956 | 0.33196 | 1.3495 | 2.548 | 99.5 | 20.88 | 37.391 | 69.4648 | 106.0581 | 66.3953 | 0.009288 | 1.0398 | 0.024838 | 23.8882 | -0.001499 | 0.14241 | 9.2e-05 | 0.007075 | 0.007075 | 0.82726
2019-03-21 12:09 | 1553184514.3739 | 1401 | 0.20961 | 0.000466 | -4e-06 | 100.8178 | 100.9006 | 124.1271 | 21.1284 | 0.001976 | 0.001895 | -152.2279 | 0.23997 | 8.4016 | 0.003506 | 0.004497 | 0.004364 | 5.4e-05 | 0.000955 | 0.30525 | 1.2948 | 2.5479 | 99.5 | 20.88 | 37.389 | 69.243 | 106.1039 | 67.2952 | 0.009287 | 1.037 | 0.024817 | 23.9314 | -0.007347 | 0.14276 | -4.4e-05 | -0.003061 | -0.003061 | -1.9169
2019-03-21 12:09 | 1553184515.3749 | 1402 | 0.20957 | 0.000458 | -4e-06 | 100.7963 | 100.9405 | 124.1564 | 21.1273 | 0.00198 | 0.00192 | -152.2349 | 0.23867 | 8.4037 | 0.003554 | 0.004352 | 0.004209 | 5.1e-05 | 0.000955 | 0.31437 | 1.2933 | 2.548 | 99.5 | 20.88 | 37.384 | 69.333 | 107.2806 | 67.8887 | 0.009286 | 1.0352 | 0.024798 | 23.9551 | -0.012782 | 0.14199 | 0.000174 | 0.013179 | 0.013179 | 0.44281
2019-03-21 12:09 | 1553184516.3754 | 1403 | 0.20961 | 0.000458 | -3e-06 | 100.844 | 101.0556 | 124.4578 | 21.1274 | 0.001985 | 0.00184 | -152.2386 | 0.24015 | 8.4024 | 0.003527 | 0.005291 | 0.005023 | 5.2e-05 | 0.000957 | 0.31404 | 1.3336 | 2.548 | 99.5 | 20.88 | 37.381 | 69.4456 | 106.6122 | 65.9648 | 0.009285 | 1.0411 | 0.024899 | 23.9164 | -0.018748 | 0.14286 | -7.2e-05 | -0.005188 | -0.005188 | -1.1318
2019-03-21 12:09 | 1553184517.375 | 1404 | 0.2096 | 0.000457 | -4e-06 | 100.8413 | 100.8199 | 124.712 | 21.1282 | 0.001931 | 0.001909 | -152.237 | 0.23937 | 8.4027 | 0.003455 | 0.004752 | 0.004536 | 5.1e-05 | 0.00096 | 0.33428 | 1.3534 | 2.5481 | 99.5 | 20.88 | 37.376 | 68.1664 | 104.8632 | 67.6196 | 0.009284 | 1.036 | 0.024864 | 23.9992 | 0.001899 | 0.1424 | -9e-06 | -0.00042 | -0.00042 | -13.9362
2019-03-21 12:09 | 1553184518.3751 | 1405 | 0.20961 | 0.000456 | -4e-06 | 100.8264 | 100.7181 | 124.7325 | 21.1277 | 0.001957 | 0.001852 | -152.2633 | 0.23801 | 8.4018 | 0.003417 | 0.005195 | 0.005095 | 5.3e-05 | 0.000961 | 0.30596 | 1.3284 | 2.548 | 99.5 | 20.88 | 37.373 | 68.7945 | 103.9466 | 66.2499 | 0.009283 | 1.0402 | 0.024916 | 23.9528 | 0.009617 | 0.14159 | -7.3e-05 | -0.005204 | -0.005204 | -1.1183
2019-03-21 12:09 | 1553184519.3754 | 1406 | 0.20961 | 0.000455 | -4e-06 | 100.7519 | 100.7361 | 124.4104 | 21.1277 | 0.00199 | 0.001854 | -152.2675 | 0.23913 | 8.3988 | 0.003466 | 0.004768 | 0.0045 | 5.5e-05 | 0.000962 | 0.31526 | 1.3441 | 2.5481 | 99.5 | 20.88 | 37.37 | 69.5706 | 105.1269 | 66.31 | 0.009282 | 1.04 | 0.024881 | 23.924 | 0.001404 | 0.14226 | -5.3e-05 | -0.003736 | -0.003736 | -1.5651
2019-03-21 12:09 | 1553184520.3752 | 1407 | 0.20963 | 0.000457 | -3e-06 | 100.8437 | 100.8812 | 124.3313 | 21.1279 | 0.001967 | 0.001864 | -152.2624 | 0.23942 | 8.4028 | 0.003478 | 0.005005 | 0.004898 | 5.8e-05 | 0.000962 | 0.31902 | 1.332 | 2.5478 | 99.5 | 20.88 | 37.368 | 69.0288 | 105.43 | 66.549 | 0.009282 | 1.0393 | 0.024865 | 23.9248 | -0.003327 | 0.14243 | -0.000163 | -0.011988 | -0.011988 | -0.48833
2019-03-21 12:09 | 1553184521.3742 | 1408 | 0.20959 | 0.000458 | -3e-06 | 100.8413 | 100.9878 | 124.3574 | 21.1273 | 0.001954 | 0.001892 | -152.246 | 0.23921 | 8.4033 | 0.003466 | 0.005629 | 0.005455 | 5.6e-05 | 0.00096 | 0.28567 | 1.3068 | 2.5479 | 99.5 | 20.88 | 37.368 | 68.7138 | 105.1241 | 67.2107 | 0.009282 | 1.0373 | 0.024843 | 23.9506 | -0.012985 | 0.14231 | 5.9e-05 | 0.004582 | 0.004582 | 1.2764
2019-03-21 12:09 | 1553184522.3742 | 1409 | 0.20962 | 0.00046 | -4e-06 | 100.8804 | 100.9568 | 124.3916 | 21.1278 | 0.0019 | 0.001899 | -152.2519 | 0.23924 | 8.4035 | 0.003394 | 0.005143 | 0.00495 | 5.4e-05 | 0.000959 | 0.32565 | 1.3447 | 2.5478 | 99.5 | 20.88 | 37.369 | 67.411 | 103.3878 | 67.3999 | 0.009282 | 1.0367 | 0.02484 | 23.9606 | -0.006778 | 0.14232 | -0.000108 | -0.007816 | -0.007816 | -0.74836
2019-03-21 12:09 | 1553184523.3744 | 1410 | 0.20962 | 0.000458 | -4e-06 | 100.855 | 100.9006 | 124.3668 | 21.1276 | 0.001942 | 0.001904 | -152.2797 | 0.23825 | 8.402 | 0.003453 | 0.004621 | 0.004443 | 5.3e-05 | 0.000962 | 0.32189 | 1.3311 | 2.548 | 99.5 | 20.89 | 37.372 | 68.4256 | 104.8184 | 67.5019 | 0.009289 | 1.0364 | 0.024834 | 23.9618 | -0.004044 | 0.14174 | -0.000102 | -0.007397 | -0.007397 | -0.78754
2019-03-21 12:09 | 1553184524.3754 | 1411 | 0.20959 | 0.000459 | -4e-06 | 100.7573 | 101.0011 | 124.2243 | 21.1278 | 0.001934 | 0.001869 | -152.2705 | 0.23911 | 8.4001 | 0.003516 | 0.00426 | 0.004072 | 5.3e-05 | 0.000958 | 0.31758 | 1.2976 | 2.5479 | 99.5 | 20.89 | 37.376 | 68.2208 | 106.3365 | 66.6636 | 0.00929 | 1.0389 | 0.02485 | 23.9186 | -0.021617 | 0.14225 | 7.7e-05 | 0.006008 | 0.006008 | 0.97309
2019-03-21 12:09 | 1553184525.3761 | 1412 | 0.20959 | 0.000455 | -4e-06 | 100.7993 | 100.8649 | 124.3212 | 21.1278 | 0.001949 | 0.001889 | -152.2557 | 0.23967 | 8.399 | 0.003519 | 0.004343 | 0.004214 | 5.6e-05 | 0.000956 | 0.32117 | 1.2621 | 2.548 | 99.5 | 20.89 | 37.386 | 68.5996 | 106.4246 | 67.1427 | 0.009292 | 1.0375 | 0.024842 | 23.9447 | -0.005815 | 0.14258 | 9e-05 | 0.006941 | 0.006941 | 0.84423
2019-03-21 12:09 | 1553184526.3739 | 1413 | 0.20957 | 0.000458 | -4e-06 | 100.8193 | 100.8788 | 124.1431 | 21.1275 | 0.001916 | 0.001893 | -152.246 | 0.23961 | 8.4016 | 0.00357 | 0.00482 | 0.0047 | 5.6e-05 | 0.000956 | 0.28998 | 1.2888 | 2.5479 | 99.5 | 20.88 | 37.392 | 67.7952 | 107.6566 | 67.2477 | 0.009288 | 1.0372 | 0.02482 | 23.9313 | -0.005283 | 0.14255 | 0.000204 | 0.015479 | 0.015479 | 0.37849
2019-03-21 12:09 | 1553184527.3746 | 1414 | 0.20957 | 0.000457 | -4e-06 | 100.7779 | 100.7218 | 123.8736 | 21.1278 | 0.001974 | 0.001832 | -152.238 | 0.23949 | 8.4002 | 0.00345 | 0.004524 | 0.004284 | 5.9e-05 | 0.000956 | 0.29048 | 1.3164 | 2.5479 | 99.5 | 20.89 | 37.401 | 69.1924 | 104.7384 | 65.7903 | 0.009296 | 1.0416 | 0.024847 | 23.8541 | 0.004975 | 0.14247 | 0.000183 | 0.01388 | 0.01388 | 0.42186
2019-03-21 12:09 | 1553184528.3752 | 1415 | 0.20962 | 0.000458 | -3e-06 | 100.8094 | 100.8104 | 123.818 | 21.1277 | 0.002012 | 0.001875 | -152.2179 | 0.2388 | 8.4007 | 0.003454 | 0.005288 | 0.005069 | 5.8e-05 | 0.000957 | 0.28766 | 1.2951 | 2.548 | 99.5 | 20.89 | 37.406 | 70.1147 | 104.8539 | 66.8071 | 0.009297 | 1.0385 | 0.024804 | 23.8844 | -8.3e-05 | 0.14206 | -7.9e-05 | -0.005716 | -0.005716 | -1.0215
2019-03-21 12:09 | 1553184529.374 | 1416 | 0.20952 | 0.000457 | -3e-06 | 100.6798 | 100.8417 | 123.7706 | 21.1282 | 0.001979 | 0.001979 | -152.2184 | 0.23929 | 8.4031 | 0.003522 | 0.005017 | 0.004853 | 5.6e-05 | 0.000959 | 0.33301 | 1.3395 | 2.5479 | 99.5 | 20.89 | 37.41 | 69.3199 | 106.4895 | 69.3105 | 0.009298 | 1.0309 | 0.024708 | 23.9676 | -0.01437 | 0.14235 | 0.000483 | 0.03608 | 0.03608 | 0.16216
2019-03-21 12:09 | 1553184530.3737 | 1417 | 0.20957 | 0.000458 | -4e-06 | 100.7612 | 100.8574 | 123.5617 | 21.1266 | 0.001908 | 0.001816 | -152.2099 | 0.23941 | 8.4011 | 0.003518 | 0.004285 | 0.004146 | 5.4e-05 | 0.000956 | 0.32964 | 1.348 | 2.548 | 99.5 | 20.89 | 37.42 | 67.6173 | 106.391 | 65.3888 | 0.009301 | 1.0429 | 0.02483 | 23.8099 | -0.008536 | 0.14242 | 0.000202 | 0.015279 | 0.015279 | 0.38312
2019-03-21 12:09 | 1553184531.3748 | 1418 | 0.20964 | 0.000457 | -3e-06 | 100.7788 | 100.8952 | 123.4387 | 21.1267 | 0.002026 | 0.001918 | -152.2179 | 0.2387 | 8.4024 | 0.003487 | 0.00474 | 0.004501 | 5.6e-05 | 0.000958 | 0.31996 | 1.3373 | 2.548 | 99.5 | 20.89 | 37.424 | 70.4514 | 105.6482 | 67.8521 | 0.009302 | 1.0353 | 0.024728 | 23.8845 | -0.010328 | 0.142 | -0.000222 | -0.016279 | -0.016279 | -0.35853
2019-03-21 12:09 | 1553184532.3752 | 1419 | 0.20954 | 0.000455 | -3e-06 | 100.6742 | 100.8404 | 123.3145 | 21.1278 | 0.001997 | 0.001889 | -152.2222 | 0.23913 | 8.4021 | 0.003441 | 0.004309 | 0.004109 | 5.7e-05 | 0.000957 | 0.32156 | 1.3608 | 2.5478 | 99.5 | 20.89 | 37.426 | 69.757 | 104.5237 | 67.1489 | 0.009302 | 1.0375 | 0.024741 | 23.8478 | -0.014752 | 0.14226 | 0.0004 | 0.029922 | 0.029922 | 0.1954
2019-03-21 12:09 | 1553184533.3752 | 1420 | 0.20958 | 0.000458 | -4e-06 | 100.6688 | 100.9943 | 123.5333 | 21.1277 | 0.001979 | 0.001894 | -152.1966 | 0.2401 | 8.4016 | 0.003516 | 0.005106 | 0.004909 | 5.5e-05 | 0.000958 | 0.31902 | 1.3346 | 2.5479 | 99.5 | 20.89 | 37.431 | 69.3062 | 106.348 | 67.2719 | 0.009303 | 1.0371 | 0.024759 | 23.8733 | -0.028873 | 0.14283 | 0.000121 | 0.009248 | 0.009248 | 0.63477
2019-03-21 12:09 | 1553184534.3749 | 1421 | 0.20961 | 0.000462 | -4e-06 | 100.7019 | 100.9817 | 123.5495 | 21.1274 | 0.002005 | 0.001905 | -152.2022 | 0.23981 | 8.4007 | 0.003433 | 0.004656 | 0.004465 | 5.6e-05 | 0.000955 | 0.30165 | 1.3214 | 2.5479 | 99.5 | 20.89 | 37.431 | 69.932 | 104.3306 | 67.5298 | 0.009303 | 1.0363 | 0.024751 | 23.8839 | -0.024819 | 0.14266 | -1.8e-05 | -0.001119 | -0.001119 | -5.2391
2019-03-21 12:09 | 1553184535.374 | 1422 | 0.20955 | 0.00046 | -4e-06 | 100.7591 | 100.668 | 123.5891 | 21.1274 | 0.001974 | 0.001851 | -152.2108 | 0.23993 | 8.4027 | 0.003515 | 0.005153 | 0.004972 | 5.6e-05 | 0.000958 | 0.32925 | 1.2713 | 2.548 | 99.5 | 20.89 | 37.428 | 69.1882 | 106.3123 | 66.2378 | 0.009303 | 1.0402 | 0.024802 | 23.8424 | 0.008089 | 0.14274 | 0.000321 | 0.024145 | 0.024145 | 0.24297
2019-03-21 12:09 | 1553184536.3753 | 1423 | 0.2096 | 0.000459 | -3e-06 | 100.718 | 100.7702 | 123.7774 | 21.1274 | 0.001963 | 0.001884 | -152.2166 | 0.23997 | 8.4019 | 0.00344 | 0.004926 | 0.004716 | 5.6e-05 | 0.000959 | 0.31205 | 1.2924 | 2.5479 | 99.5 | 20.89 | 37.426 | 68.9376 | 104.5169 | 67.0227 | 0.009302 | 1.0378 | 0.024792 | 23.8881 | -0.004636 | 0.14276 | 4e-06 | 0.000456 | 0.000456 | 12.8573
2019-03-21 12:09 | 1553184537.3742 | 1424 | 0.20957 | 0.000457 | -4e-06 | 100.645 | 100.7777 | 123.8703 | 21.128 | 0.001947 | 0.001808 | -152.2038 | 0.2396 | 8.4035 | 0.003524 | 0.004287 | 0.004139 | 5.5e-05 | 0.000958 | 0.28567 | 1.3105 | 2.548 | 99.5 | 20.89 | 37.426 | 68.5375 | 106.545 | 65.2052 | 0.009302 | 1.0434 | 0.024868 | 23.8331 | -0.011786 | 0.14254 | 0.000221 | 0.016737 | 0.016737 | 0.35003
2019-03-21 12:09 | 1553184538.3742 | 1425 | 0.20959 | 0.000458 | -3e-06 | 100.6661 | 100.9708 | 123.9583 | 21.1285 | 0.001946 | 0.001864 | -152.2142 | 0.23943 | 8.4036 | 0.003443 | 0.005014 | 0.004873 | 5.4e-05 | 0.000956 | 0.32156 | 1.3336 | 2.5479 | 99.49 | 20.89 | 37.432 | 68.5263 | 104.5709 | 66.5378 | 0.009305 | 1.0393 | 0.024828 | 23.8885 | -0.027029 | 0.14244 | 7.8e-05 | 0.006007 | 0.006007 | 0.9746
2019-03-21 12:09 | 1553184539.3752 | 1426 | 0.20958 | 0.000459 | -3e-06 | 100.6784 | 100.8547 | 124.4898 | 21.1276 | 0.001913 | 0.001805 | -152.2127 | 0.23946 | 8.4029 | 0.00345 | 0.00502 | 0.00483 | 5.7e-05 | 0.000955 | 0.31669 | 1.2949 | 2.548 | 99.49 | 20.89 | 37.438 | 67.718 | 104.7421 | 65.1391 | 0.009306 | 1.0436 | 0.024932 | 23.8903 | -0.01565 | 0.14245 | 0.000126 | 0.009627 | 0.009627 | 0.6082
2019-03-21 12:09 | 1553184540.3743 | 1427 | 0.20958 | 0.000459 | -3e-06 | 100.6125 | 100.8744 | 124.962 | 21.1278 | 0.001873 | 0.001851 | -152.2169 | 0.23922 | 8.4023 | 0.003433 | 0.004697 | 0.004545 | 5.6e-05 | 0.000953 | 0.30215 | 1.3465 | 2.548 | 99.49 | 20.89 | 37.453 | 66.7697 | 104.3477 | 66.2343 | 0.00931 | 1.0403 | 0.024939 | 23.9743 | -0.023253 | 0.14231 | 0.00011 | 0.008432 | 0.008432 | 0.69367
2019-03-21 12:09 | 1553184541.3743 | 1428 | 0.20956 | 0.000458 | -3e-06 | 100.6 | 100.5713 | 125.177 | 21.1274 | 0.001922 | 0.001858 | -152.2053 | 0.23947 | 8.402 | 0.003439 | 0.004441 | 0.00428 | 5.4e-05 | 0.00095 | 0.31349 | 1.2889 | 2.5479 | 99.49 | 20.89 | 37.462 | 67.9418 | 104.4715 | 66.404 | 0.009312 | 1.0397 | 0.024955 | 24.0009 | 0.002553 | 0.14246 | 0.000234 | 0.017781 | 0.017781 | 0.32929
2019-03-21 12:09 | 1553184542.3737 | 1429 | 0.20959 | 0.000456 | -4e-06 | 100.6092 | 100.7726 | 125.2869 | 21.1266 | 0.002016 | 0.001855 | -152.2028 | 0.23859 | 8.404 | 0.003363 | 0.004039 | 0.0039 | 5.5e-05 | 0.000953 | 0.30735 | 1.3076 | 2.5479 | 99.49 | 20.89 | 37.472 | 70.212 | 102.6458 | 66.3385 | 0.009314 | 1.0399 | 0.024968 | 24.0091 | -0.01451 | 0.14194 | 0.000107 | 0.008288 | 0.008288 | 0.70388
2019-03-21 12:09 | 1553184543.3734 | 1430 | 0.20959 | 0.000459 | -4e-06 | 100.5824 | 100.7235 | 125.5549 | 21.1266 | 0.001984 | 0.001856 | -152.2103 | 0.2396 | 8.4038 | 0.003461 | 0.004625 | 0.00445 | 5.7e-05 | 0.000947 | 0.30541 | 1.3156 | 2.548 | 99.49 | 20.89 | 37.488 | 69.4244 | 105.019 | 66.3619 | 0.009318 | 1.0399 | 0.024994 | 24.0356 | -0.012541 | 0.14253 | 5.3e-05 | 0.004251 | 0.004251 | 1.3782
2019-03-21 12:09 | 1553184544.3745 | 1431 | 0.20957 | 0.000458 | -3e-06 | 100.6208 | 100.6592 | 125.6392 | 21.1265 | 0.001954 | 0.00185 | -152.2127 | 0.23976 | 8.4032 | 0.003428 | 0.004734 | 0.004576 | 5.3e-05 | 0.00095 | 0.29446 | 1.3313 | 2.5479 | 99.49 | 20.89 | 37.495 | 68.7185 | 104.2144 | 66.215 | 0.00932 | 1.0403 | 0.025008 | 24.0385 | -0.003405 | 0.14263 | 0.000209 | 0.015898 | 0.015898 | 0.36873
2019-03-21 12:09 | 1553184545.3739 | 1432 | 0.20959 | 0.000459 | -4e-06 | 100.6051 | 100.7021 | 125.3444 | 21.1271 | 0.001935 | 0.001834 | -152.2105 | 0.23941 | 8.4029 | 0.003397 | 0.004601 | 0.004443 | 5.8e-05 | 0.000954 | 0.32604 | 1.298 | 2.5479 | 99.49 | 20.89 | 37.501 | 68.2534 | 103.4585 | 65.8192 | 0.009322 | 1.0415 | 0.024993 | 23.9963 | -0.008621 | 0.14243 | 6.7e-05 | 0.005271 | 0.005271 | 1.1106
2019-03-21 12:09 | 1553184546.3735 | 1433 | 0.20958 | 0.00046 | -4e-06 | 100.6977 | 100.7051 | 125.1195 | 21.127 | 0.001965 | 0.001867 | -152.2213 | 0.24 | 8.4008 | 0.003448 | 0.005222 | 0.005078 | 5.4e-05 | 0.000953 | 0.32278 | 1.2788 | 2.5479 | 99.49 | 20.89 | 37.512 | 68.971 | 104.7086 | 66.6222 | 0.009324 | 1.0391 | 0.024941 | 24.0031 | -0.000659 | 0.14278 | 0.000106 | 0.008206 | 0.008206 | 0.71509
2019-03-21 12:09 | 1553184547.3736 | 1434 | 0.20955 | 0.000455 | -4e-06 | 100.676 | 100.6476 | 125.1301 | 21.1261 | 0.001933 | 0.001882 | -152.2367 | 0.23968 | 8.4057 | 0.003335 | 0.005156 | 0.004995 | 5.2e-05 | 0.000956 | 0.32925 | 1.2901 | 2.5479 | 99.49 | 20.89 | 37.517 | 68.1987 | 101.9508 | 66.9811 | 0.009326 | 1.038 | 0.024929 | 24.0168 | 0.002523 | 0.14259 | 0.000293 | 0.022155 | 0.022155 | 0.26451
2019-03-21 12:09 | 1553184548.375 | 1435 | 0.20958 | 0.000456 | -4e-06 | 100.6506 | 100.6793 | 124.7578 | 21.127 | 0.001958 | 0.001856 | -152.2417 | 0.23928 | 8.4014 | 0.003353 | 0.004707 | 0.004525 | 5.4e-05 | 0.000958 | 0.3235 | 1.3311 | 2.548 | 99.49 | 20.89 | 37.522 | 68.8111 | 102.3956 | 66.3563 | 0.009327 | 1.0399 | 0.024914 | 23.959 | -0.002542 | 0.14235 | 0.000145 | 0.011071 | 0.011071 | 0.52846
2019-03-21 12:09 | 1553184549.3737 | 1436 | 0.20958 | 0.000459 | -4e-06 | 100.6426 | 100.8727 | 124.1776 | 21.1269 | 0.001951 | 0.001875 | -152.252 | 0.2397 | 8.4004 | 0.003434 | 0.004963 | 0.00475 | 5.2e-05 | 0.000961 | 0.3298 | 1.3285 | 2.5479 | 99.49 | 20.89 | 37.521 | 68.637 | 104.3687 | 66.821 | 0.009327 | 1.0385 | 0.024839 | 23.9196 | -0.020427 | 0.1426 | 0.00013 | 0.009926 | 0.009926 | 0.59044
2019-03-21 12:09 | 1553184550.3747 | 1437 | 0.20956 | 0.000458 | -4e-06 | 100.6447 | 100.7688 | 123.8412 | 21.1267 | 0.001923 | 0.001854 | -152.2534 | 0.24063 | 8.4023 | 0.003453 | 0.005089 | 0.004984 | 5.6e-05 | 0.000962 | 0.3361 | 1.3332 | 2.548 | 99.49 | 20.89 | 37.51 | 67.9615 | 104.8278 | 66.3146 | 0.009324 | 1.04 | 0.024824 | 23.8694 | -0.011026 | 0.14315 | 0.000256 | 0.019302 | 0.019302 | 0.30481
2019-03-21 12:09 | 1553184551.3752 | 1438 | 0.20959 | 0.000452 | -4e-06 | 100.6083 | 100.8604 | 123.5673 | 21.1264 | 0.001938 | 0.001905 | -152.2609 | 0.23983 | 8.4005 | 0.003449 | 0.005063 | 0.004864 | 5.1e-05 | 0.00096 | 0.29175 | 1.3061 | 2.548 | 99.49 | 20.89 | 37.502 | 68.3185 | 104.7104 | 67.5275 | 0.009322 | 1.0363 | 0.024753 | 23.8855 | -0.022386 | 0.14267 | 0.000108 | 0.008233 | 0.008233 | 0.71226
2019-03-21 12:09 | 1553184552.3749 | 1439 | 0.20958 | 0.000447 | -4e-06 | 100.6542 | 100.8032 | 123.4168 | 21.1271 | 0.001977 | 0.00187 | -152.2627 | 0.24022 | 8.4015 | 0.003452 | 0.004805 | 0.004635 | 5.5e-05 | 0.000967 | 0.31006 | 1.3859 | 2.5479 | 99.49 | 20.89 | 37.493 | 69.255 | 104.7843 | 66.6957 | 0.00932 | 1.0388 | 0.024768 | 23.8418 | -0.013232 | 0.14291 | 0.000149 | 0.011332 | 0.011332 | 0.51831
2019-03-21 12:09 | 1553184553.375 | 1440 | 0.20959 | 0.000451 | -4e-06 | 100.6706 | 100.7569 | 122.9781 | 21.127 | 0.00192 | 0.001817 | -152.2596 | 0.23997 | 8.4045 | 0.003436 | 0.005341 | 0.005144 | 5.5e-05 | 0.00097 | 0.31349 | 1.3289 | 2.5479 | 99.49 | 20.89 | 37.493 | 67.8978 | 104.4114 | 65.4234 | 0.00932 | 1.0427 | 0.02477 | 23.7548 | -0.007665 | 0.14276 | 8.8e-05 | 0.006757 | 0.006757 | 0.86837
2019-03-21 12:09 | 1553184554.3753 | 1441 | 0.20959 | 0.000451 | -4e-06 | 100.6992 | 100.7218 | 123.1023 | 21.1264 | 0.001965 | 0.001842 | -152.2623 | 0.24009 | 8.404 | 0.003462 | 0.004932 | 0.004706 | 5.3e-05 | 0.000974 | 0.33516 | 1.3191 | 2.548 | 99.49 | 20.89 | 37.464 | 68.9711 | 105.0367 | 66.0209 | 0.009312 | 1.0409 | 0.024761 | 23.7878 | -0.002009 | 0.14283 | 9.3e-05 | 0.007157 | 0.007157 | 0.82023
2019-03-21 12:09 | 1553184555.374 | 1442 | 0.20957 | 0.000457 | -4e-06 | 100.7177 | 100.7058 | 123.4354 | 21.1277 | 0.001952 | 0.001853 | -152.2718 | 0.23989 | 8.4027 | 0.003406 | 0.004798 | 0.004625 | 5.4e-05 | 0.000973 | 0.32676 | 1.2924 | 2.548 | 99.49 | 20.89 | 37.455 | 68.6728 | 103.6902 | 66.2873 | 0.00931 | 1.0401 | 0.024785 | 23.8293 | 0.001053 | 0.14271 | 0.00019 | 0.014386 | 0.014386 | 0.4077
2019-03-21 12:09 | 1553184556.3745 | 1443 | 0.2096 | 0.000452 | -4e-06 | 100.7227 | 100.7699 | 123.4945 | 21.1279 | 0.001967 | 0.001866 | -152.251 | 0.23869 | 8.4035 | 0.003445 | 0.005273 | 0.005147 | 5.6e-05 | 0.00097 | 0.31758 | 1.3276 | 2.548 | 99.49 | 20.89 | 37.447 | 69.0195 | 104.6205 | 66.6036 | 0.009308 | 1.0391 | 0.024779 | 23.8461 | -0.004182 | 0.142 | 9e-06 | 0.000933 | 0.000933 | 6.253
2019-03-21 12:09 | 1553184557.3745 | 1444 | 0.20958 | 0.000454 | -4e-06 | 100.7728 | 100.9824 | 123.5532 | 21.1275 | 0.00199 | 0.00183 | -152.2593 | 0.23865 | 8.401 | 0.003462 | 0.004963 | 0.00473 | 5.2e-05 | 0.000961 | 0.29822 | 1.3296 | 2.548 | 99.49 | 20.89 | 37.43 | 69.5695 | 105.0276 | 65.7236 | 0.009304 | 1.0418 | 0.024817 | 23.8208 | -0.018583 | 0.14197 | 0.000182 | 0.0138 | 0.0138 | 0.42284
2019-03-21 12:09 | 1553184558.3732 | 1445 | 0.20958 | 0.000456 | -4e-06 | 100.6986 | 100.8516 | 123.8389 | 21.127 | 0.001873 | 0.001821 | -152.2505 | 0.23988 | 8.4013 | 0.003454 | 0.00443 | 0.004238 | 5.5e-05 | 0.000962 | 0.32261 | 1.3155 | 2.5479 | 99.49 | 20.89 | 37.423 | 66.7595 | 104.8532 | 65.5023 | 0.009302 | 1.0425 | 0.024854 | 23.8406 | -0.013578 | 0.1427 | 0.00015 | 0.011465 | 0.011465 | 0.51158
2019-03-21 12:09 | 1553184559.3737 | 1446 | 0.20957 | 0.000455 | -4e-06 | 100.7758 | 100.8979 | 124.2663 | 21.1278 | 0.001972 | 0.001843 | -152.2348 | 0.24038 | 8.4003 | 0.003472 | 0.004462 | 0.004273 | 5.3e-05 | 0.000965 | 0.33019 | 1.3169 | 2.5478 | 99.49 | 20.89 | 37.415 | 69.1479 | 105.2897 | 66.0347 | 0.0093 | 1.0409 | 0.024877 | 23.9004 | -0.010834 | 0.143 | 0.00022 | 0.016701 | 0.016701 | 0.35192
2019-03-21 12:09 | 1553184560.3742 | 1447 | 0.20963 | 0.000452 | -4e-06 | 100.6924 | 100.9524 | 124.3423 | 21.1274 | 0.001968 | 0.001874 | -152.2532 | 0.24041 | 8.3997 | 0.003439 | 0.004735 | 0.004661 | 5.4e-05 | 0.000964 | 0.35153 | 1.3011 | 2.5478 | 99.49 | 20.89 | 37.404 | 69.0471 | 104.4802 | 66.7809 | 0.009298 | 1.0386 | 0.024857 | 23.934 | -0.023071 | 0.14302 | -0.000173 | -0.01271 | -0.01271 | -0.46249
2019-03-21 12:09 | 1553184561.3734 | 1448 | 0.20961 | 0.000451 | -4e-06 | 100.7135 | 100.8727 | 124.3262 | 21.128 | 0.001942 | 0.001807 | -152.2379 | 0.23991 | 8.4029 | 0.00351 | 0.005251 | 0.004964 | 5.5e-05 | 0.000963 | 0.29015 | 1.3186 | 2.5479 | 99.49 | 20.89 | 37.399 | 68.4135 | 106.2091 | 65.1827 | 0.009296 | 1.0435 | 0.024914 | 23.8762 | -0.014127 | 0.14272 | -1.6e-05 | -0.000981 | -0.000981 | -5.9806
2019-03-21 12:09 | 1553184562.3748 | 1449 | 0.20968 | 0.000452 | -4e-06 | 100.7624 | 100.9068 | 124.2407 | 21.1279 | 0.00198 | 0.001869 | -152.2287 | 0.23972 | 8.4025 | 0.003503 | 0.004305 | 0.004135 | 5.6e-05 | 0.000959 | 0.3402 | 1.3397 | 2.5479 | 99.49 | 20.89 | 37.39 | 69.3314 | 106.0295 | 66.6746 | 0.009294 | 1.0389 | 0.024851 | 23.9205 | -0.012809 | 0.14261 | -0.000476 | -0.035296 | -0.035296 | -0.16607
2019-03-21 12:09 | 1553184563.3735 | 1450 | 0.20958 | 0.000449 | -4e-06 | 100.7671 | 100.822 | 124.4325 | 21.1285 | 0.001959 | 0.001835 | -152.2355 | 0.24045 | 8.4009 | 0.003498 | 0.005113 | 0.004936 | 5.6e-05 | 0.00096 | 0.29391 | 1.3566 | 2.548 | 99.49 | 20.89 | 37.386 | 68.8452 | 105.9002 | 65.8463 | 0.009293 | 1.0414 | 0.024901 | 23.9098 | -0.004865 | 0.14305 | 0.000127 | 0.009727 | 0.009727 | 0.60444
2019-03-21 12:09 | 1553184564.3745 | 1451 | 0.20958 | 0.000453 | -3e-06 | 100.7335 | 100.9647 | 124.4682 | 21.1275 | 0.001973 | 0.001917 | -152.2338 | 0.2384 | 8.4014 | 0.003456 | 0.004103 | 0.003923 | 5.4e-05 | 0.000964 | 0.31188 | 1.3544 | 2.5481 | 99.49 | 20.89 | 37.381 | 69.1789 | 104.886 | 67.8209 | 0.009292 | 1.0354 | 0.024832 | 23.9828 | -0.020505 | 0.14182 | 0.00018 | 0.013626 | 0.013626 | 0.42778
2019-03-21 12:09 | 1553184565.374 | 1452 | 0.20959 | 0.00045 | -4e-06 | 100.6786 | 100.7423 | 124.4191 | 21.1277 | 0.001927 | 0.001829 | -152.2125 | 0.23945 | 8.4041 | 0.003458 | 0.004734 | 0.004622 | 5.4e-05 | 0.000967 | 0.32781 | 1.3363 | 2.5479 | 99.49 | 20.89 | 37.375 | 68.0661 | 104.9388 | 65.7099 | 0.00929 | 1.0419 | 0.024904 | 23.9037 | -0.00565 | 0.14245 | 0.000111 | 0.008533 | 0.008533 | 0.68609
2019-03-21 12:09 | 1553184566.3737 | 1453 | 0.20956 | 0.000452 | -4e-06 | 100.7365 | 101.0134 | 124.5038 | 21.1271 | 0.001947 | 0.001838 | -152.2179 | 0.23857 | 8.4026 | 0.003446 | 0.004645 | 0.004429 | 6e-05 | 0.000964 | 0.3194 | 1.3142 | 2.5479 | 99.49 | 20.89 | 37.373 | 68.5499 | 104.6575 | 65.9287 | 0.00929 | 1.0412 | 0.024905 | 23.9195 | -0.024555 | 0.14193 | 0.00029 | 0.021895 | 0.021895 | 0.26642
2019-03-21 12:09 | 1553184567.3737 | 1454 | 0.2096 | 0.000448 | -4e-06 | 100.7239 | 100.8448 | 124.3445 | 21.1277 | 0.001984 | 0.001839 | -152.2123 | 0.24056 | 8.4051 | 0.003523 | 0.004289 | 0.004153 | 5.6e-05 | 0.000966 | 0.30663 | 1.3448 | 2.5481 | 99.49 | 20.89 | 37.372 | 69.4411 | 106.5283 | 65.9377 | 0.00929 | 1.0412 | 0.024889 | 23.9045 | -0.010725 | 0.14311 | 6.3e-05 | 0.004962 | 0.004962 | 1.1855
2019-03-21 12:09 | 1553184568.3752 | 1455 | 0.2096 | 0.00045 | -4e-06 | 100.6921 | 100.761 | 124.3684 | 21.127 | 0.001956 | 0.001817 | -152.2166 | 0.23949 | 8.4023 | 0.003477 | 0.004899 | 0.004754 | 5.6e-05 | 0.000964 | 0.31581 | 1.3381 | 2.5478 | 99.49 | 20.89 | 37.371 | 68.7587 | 105.4118 | 65.4171 | 0.009289 | 1.0428 | 0.02491 | 23.8885 | -0.006122 | 0.14247 | 5e-05 | 0.004015 | 0.004015 | 1.4584
2019-03-21 12:09 | 1553184569.3737 | 1456 | 0.20958 | 0.000448 | -4e-06 | 100.6995 | 100.87 | 124.4783 | 21.1268 | 0.00198 | 0.001831 | -152.2159 | 0.23932 | 8.404 | 0.003472 | 0.004911 | 0.004652 | 5.5e-05 | 0.00097 | 0.32068 | 1.3412 | 2.5479 | 99.49 | 20.89 | 37.372 | 69.3467 | 105.2841 | 65.7636 | 0.00929 | 1.0417 | 0.024908 | 23.9113 | -0.01513 | 0.14237 | 0.000155 | 0.011841 | 0.011841 | 0.49418
2019-03-21 12:09 | 1553184570.3749 | 1457 | 0.20958 | 0.000451 | -4e-06 | 100.6542 | 100.7661 | 124.4773 | 21.1277 | 0.001965 | 0.001848 | -152.1948 | 0.23986 | 8.4042 | 0.003475 | 0.005214 | 0.005174 | 5.8e-05 | 0.00097 | 0.31725 | 1.3209 | 2.5479 | 99.49 | 20.89 | 37.374 | 68.9699 | 105.3502 | 66.1679 | 0.00929 | 1.0405 | 0.024893 | 23.9254 | -0.009937 | 0.14269 | 0.000143 | 0.010921 | 0.010921 | 0.53701
2019-03-21 12:10 | 1553184571.3742 | 1458 | 0.20965 | 0.000449 | -4e-06 | 100.6694 | 100.8891 | 124.5759 | 21.1275 | 0.00196 | 0.00183 | -152.2087 | 0.23952 | 8.402 | 0.003502 | 0.004903 | 0.004733 | 5.9e-05 | 0.000967 | 0.32516 | 1.3088 | 2.5479 | 99.49 | 20.89 | 37.373 | 68.8529 | 105.9968 | 65.7385 | 0.00929 | 1.0418 | 0.024919 | 23.9198 | -0.019494 | 0.14249 | -0.000251 | -0.018514 | -0.018514 | -0.31632
2019-03-21 12:10 | 1553184572.373 | 1459 | 0.20957 | 0.000451 | -3e-06 | 100.6685 | 100.6609 | 124.9716 | 21.1273 | 0.001931 | 0.001874 | -152.2156 | 0.24045 | 8.4009 | 0.003478 | 0.005058 | 0.004885 | 5.5e-05 | 0.00097 | 0.31653 | 1.2932 | 2.548 | 99.49 | 20.89 | 37.374 | 68.1664 | 105.431 | 66.7879 | 0.00929 | 1.0386 | 0.02492 | 23.9948 | 0.00068 | 0.14305 | 0.000195 | 0.01482 | 0.01482 | 0.3967
2019-03-21 12:10 | 1553184573.373 | 1460 | 0.20958 | 0.000449 | -4e-06 | 100.6286 | 100.8104 | 125.1003 | 21.127 | 0.001995 | 0.001849 | -152.2088 | 0.23983 | 8.4038 | 0.003462 | 0.004657 | 0.0045 | 5.5e-05 | 0.00097 | 0.32261 | 1.3101 | 2.548 | 99.49 | 20.89 | 37.374 | 69.7079 | 105.0428 | 66.1876 | 0.00929 | 1.0404 | 0.024955 | 23.9859 | -0.016144 | 0.14267 | 0.000192 | 0.014622 | 0.014622 | 0.40103
2019-03-21 12:10 | 1553184574.3734 | 1461 | 0.20956 | 0.000455 | -4e-06 | 100.5434 | 100.8819 | 125.2524 | 21.1264 | 0.001867 | 0.001859 | -152.2114 | 0.23989 | 8.4014 | 0.00346 | 0.004194 | 0.004004 | 5.7e-05 | 0.000969 | 0.31653 | 1.3334 | 2.5481 | 99.49 | 20.89 | 37.373 | 66.6115 | 104.995 | 66.4343 | 0.00929 | 1.0396 | 0.024961 | 24.0092 | -0.030065 | 0.14271 | 0.000253 | 0.019237 | 0.019237 | 0.3049
2019-03-21 12:10 | 1553184575.3732 | 1462 | 0.20956 | 0.000452 | -4e-06 | 100.5651 | 100.6871 | 125.2172 | 21.127 | 0.001952 | 0.001834 | -152.2236 | 0.24002 | 8.4037 | 0.003516 | 0.004607 | 0.004411 | 5.5e-05 | 0.000969 | 0.29142 | 1.3481 | 2.548 | 99.49 | 20.89 | 37.374 | 68.67 | 106.3483 | 65.8333 | 0.00929 | 1.0415 | 0.02498 | 23.9846 | -0.01084 | 0.14279 | 0.000283 | 0.021449 | 0.021449 | 0.27361
2019-03-21 12:10 | 1553184576.3742 | 1463 | 0.20956 | 0.000453 | -4e-06 | 100.6083 | 100.8778 | 125.3275 | 21.128 | 0.001972 | 0.001857 | -152.2335 | 0.2401 | 8.4047 | 0.00347 | 0.00507 | 0.004909 | 5.3e-05 | 0.000971 | 0.29015 | 1.3204 | 2.548 | 99.49 | 20.89 | 37.374 | 69.1498 | 105.2443 | 66.3694 | 0.00929 | 1.0398 | 0.024971 | 24.0141 | -0.023926 | 0.14283 | 0.000265 | 0.020105 | 0.020105 | 0.292
2019-03-21 12:10 | 1553184577.3736 | 1464 | 0.2096 | 0.000453 | -3e-06 | 100.5717 | 100.7324 | 125.5261 | 21.1279 | 0.001913 | 0.001843 | -152.2392 | 0.23967 | 8.4028 | 0.003435 | 0.004589 | 0.004415 | 5.4e-05 | 0.000972 | 0.33467 | 1.3312 | 2.548 | 99.49 | 20.89 | 37.373 | 67.7364 | 104.3728 | 66.0526 | 0.00929 | 1.0408 | 0.025002 | 24.0219 | -0.014281 | 0.14258 | 4.7e-05 | 0.003736 | 0.003736 | 1.5686
2019-03-21 12:10 | 1553184578.3743 | 1465 | 0.20958 | 0.000454 | -3e-06 | 100.5476 | 100.76 | 125.4359 | 21.1274 | 0.001947 | 0.00181 | -152.2223 | 0.23944 | 8.4042 | 0.003526 | 0.00441 | 0.004267 | 5.2e-05 | 0.000977 | 0.29623 | 1.3677 | 2.5481 | 99.49 | 20.89 | 37.368 | 68.5495 | 106.595 | 65.2475 | 0.009289 | 1.0433 | 0.025023 | 23.9848 | -0.018877 | 0.14244 | 0.000171 | 0.013111 | 0.013111 | 0.44654
2019-03-21 12:10 | 1553184579.3733 | 1466 | 0.20956 | 0.000449 | -4e-06 | 100.5312 | 100.7242 | 125.1071 | 21.1277 | 0.001929 | 0.001873 | -152.2102 | 0.24028 | 8.4044 | 0.003495 | 0.004774 | 0.004617 | 5.5e-05 | 0.000978 | 0.31277 | 1.3169 | 2.548 | 99.49 | 20.89 | 37.366 | 68.1144 | 105.8413 | 66.7595 | 0.009288 | 1.0386 | 0.024935 | 24.0068 | -0.017159 | 0.14294 | 0.00031 | 0.023499 | 0.023499 | 0.25001
2019-03-21 12:10 | 1553184580.3734 | 1467 | 0.20963 | 0.000453 | -3e-06 | 100.4722 | 100.685 | 124.8787 | 21.1273 | 0.001936 | 0.001909 | -152.218 | 0.24004 | 8.4036 | 0.003421 | 0.004662 | 0.004425 | 5.4e-05 | 0.000977 | 0.27887 | 1.3239 | 2.548 | 99.49 | 20.89 | 37.361 | 68.2801 | 104.0436 | 67.6344 | 0.009287 | 1.036 | 0.02488 | 24.0157 | -0.018931 | 0.1428 | -0.000141 | -0.010354 | -0.010354 | -0.56685
2019-03-21 12:10 | 1553184581.3734 | 1468 | 0.20963 | 0.000449 | -3e-06 | 100.4519 | 100.6132 | 124.8475 | 21.1277 | 0.001891 | 0.001882 | -152.2161 | 0.24011 | 8.4022 | 0.003499 | 0.004901 | 0.004674 | 5.5e-05 | 0.000981 | 0.2912 | 1.3491 | 2.5481 | 99.49 | 20.89 | 37.358 | 67.1878 | 105.9364 | 66.9841 | 0.009286 | 1.038 | 0.0249 | 23.9898 | -0.014348 | 0.14284 | -0.000168 | -0.012377 | -0.012377 | -0.47434
2019-03-21 12:10 | 1553184582.3748 | 1469 | 0.20957 | 0.000452 | -3e-06 | 100.4871 | 100.512 | 124.282 | 21.1278 | 0.002001 | 0.001845 | -152.2157 | 0.23994 | 8.402 | 0.00345 | 0.00458 | 0.004567 | 5.3e-05 | 0.000978 | 0.30215 | 1.36 | 2.548 | 99.49 | 20.89 | 37.355 | 69.8459 | 104.737 | 66.0845 | 0.009285 | 1.0407 | 0.024877 | 23.9037 | -0.002221 | 0.14274 | 0.000223 | 0.0169 | 0.0169 | 0.34713
2019-03-21 12:10 | 1553184583.3745 | 1470 | 0.20955 | 0.000454 | -4e-06 | 100.5055 | 100.5962 | 123.9135 | 21.1274 | 0.001979 | 0.001837 | -152.2099 | 0.24051 | 8.4033 | 0.003474 | 0.0043 | 0.004097 | 5.3e-05 | 0.000981 | 0.29839 | 1.3611 | 2.5478 | 99.49 | 20.89 | 37.352 | 69.3216 | 105.3264 | 65.9078 | 0.009285 | 1.0413 | 0.024846 | 23.862 | -0.00806 | 0.14308 | 0.000314 | 0.023725 | 0.023725 | 0.24786
2019-03-21 12:10 | 1553184584.3745 | 1471 | 0.20959 | 0.000453 | -4e-06 | 100.4898 | 100.728 | 123.3603 | 21.1279 | 0.001909 | 0.00183 | -152.2096 | 0.2393 | 8.402 | 0.003393 | 0.004743 | 0.004455 | 5.4e-05 | 0.00098 | 0.31725 | 1.3188 | 2.5479 | 99.49 | 20.89 | 37.35 | 67.638 | 103.3651 | 65.7194 | 0.009284 | 1.0418 | 0.024798 | 23.8021 | -0.021177 | 0.14236 | 0.000104 | 0.008012 | 0.008012 | 0.73028
2019-03-21 12:10 | 1553184585.3737 | 1472 | 0.20956 | 0.000452 | -3e-06 | 100.4805 | 100.715 | 123.0091 | 21.1273 | 0.001928 | 0.001837 | -152.2087 | 0.23956 | 8.4013 | 0.003389 | 0.004996 | 0.004731 | 5e-05 | 0.00098 | 0.3194 | 1.3096 | 2.548 | 99.49 | 20.89 | 37.345 | 68.0789 | 103.2788 | 65.9077 | 0.009283 | 1.0413 | 0.024756 | 23.7748 | -0.02085 | 0.14252 | 0.000273 | 0.020529 | 0.020529 | 0.28532
2019-03-21 12:10 | 1553184586.3737 | 1473 | 0.20956 | 0.000454 | -4e-06 | 100.4418 | 100.497 | 122.7601 | 21.1278 | 0.001956 | 0.001824 | -152.1983 | 0.24003 | 8.4024 | 0.003479 | 0.004647 | 0.004539 | 5.5e-05 | 0.00098 | 0.30198 | 1.3058 | 2.548 | 99.49 | 20.89 | 37.343 | 68.7681 | 105.4527 | 65.5811 | 0.009282 | 1.0423 | 0.024742 | 23.7393 | -0.004919 | 0.1428 | 0.000271 | 0.020398 | 0.020398 | 0.28772
2019-03-21 12:10 | 1553184587.3728 | 1474 | 0.20958 | 0.000453 | -4e-06 | 100.4379 | 100.557 | 122.7676 | 21.1284 | 0.001985 | 0.001808 | -152.1731 | 0.23943 | 8.4035 | 0.003461 | 0.004809 | 0.004662 | 5.6e-05 | 0.000982 | 0.30541 | 1.286 | 2.5478 | 99.49 | 20.89 | 37.341 | 69.448 | 105.0253 | 65.1932 | 0.009282 | 1.0435 | 0.024757 | 23.7264 | -0.010598 | 0.14244 | 0.000129 | 0.009865 | 0.009865 | 0.59341
2019-03-21 12:10 | 1553184588.3748 | 1475 | 0.20957 | 0.000452 | -4e-06 | 100.4498 | 100.3703 | 122.8584 | 21.1283 | 0.001948 | 0.001805 | -152.1841 | 0.24009 | 8.4042 | 0.003494 | 0.004544 | 0.004354 | 5.5e-05 | 0.000983 | 0.32139 | 1.2758 | 2.548 | 99.49 | 20.89 | 37.337 | 68.5619 | 105.8229 | 65.1196 | 0.009281 | 1.0437 | 0.024769 | 23.7326 | 0.007081 | 0.14283 | 0.000243 | 0.018301 | 0.018301 | 0.32077
2019-03-21 12:10 | 1553184589.3728 | 1476 | 0.20955 | 0.000452 | -4e-06 | 100.4534 | 100.6013 | 122.8699 | 21.1276 | 0.001932 | 0.001882 | -152.2005 | 0.23999 | 8.4025 | 0.003508 | 0.004947 | 0.004764 | 5.4e-05 | 0.000984 | 0.29502 | 1.3097 | 2.5478 | 99.49 | 20.89 | 37.336 | 68.1876 | 106.1512 | 66.9708 | 0.009281 | 1.038 | 0.024703 | 23.7986 | -0.013156 | 0.14277 | 0.00033 | 0.024733 | 0.024733 | 0.23725
2019-03-21 12:10 | 1553184590.3743 | 1477 | 0.20956 | 0.000451 | -4e-06 | 100.4123 | 100.5818 | 122.9333 | 21.1279 | 0.001928 | 0.001815 | -152.1988 | 0.23848 | 8.4043 | 0.003385 | 0.00511 | 0.004981 | 5.3e-05 | 0.000985 | 0.30901 | 1.3252 | 2.5478 | 99.49 | 20.89 | 37.334 | 68.0928 | 103.1671 | 65.3659 | 0.00928 | 1.0429 | 0.024768 | 23.7485 | -0.015091 | 0.14187 | 0.00028 | 0.021121 | 0.021121 | 0.27607
2019-03-21 12:10 | 1553184591.374 | 1478 | 0.20958 | 0.000453 | -4e-06 | 100.4543 | 100.5335 | 123.256 | 21.1283 | 0.00198 | 0.001806 | -152.1973 | 0.2388 | 8.4036 | 0.003482 | 0.005018 | 0.004925 | 5.3e-05 | 0.000986 | 0.29822 | 1.3099 | 2.5479 | 99.49 | 20.89 | 37.332 | 69.327 | 105.5232 | 65.1461 | 0.00928 | 1.0436 | 0.024808 | 23.7719 | -0.007048 | 0.14206 | 0.000186 | 0.014058 | 0.014058 | 0.41533
2019-03-21 12:10 | 1553184592.374 | 1479 | 0.20957 | 0.000451 | -4e-06 | 100.4665 | 100.4763 | 123.5948 | 21.1284 | 0.001939 | 0.00187 | -152.2017 | 0.24002 | 8.4038 | 0.003386 | 0.004822 | 0.004589 | 5.6e-05 | 0.000991 | 0.32244 | 1.3631 | 2.5479 | 99.49 | 20.89 | 37.33 | 68.3617 | 103.2092 | 66.7008 | 0.009279 | 1.0388 | 0.024786 | 23.8592 | -0.000868 | 0.14279 | 0.000219 | 0.016507 | 0.016507 | 0.35553
2019-03-21 12:10 | 1553184593.3733 | 1480 | 0.20955 | 0.000452 | -4e-06 | 100.4647 | 100.4606 | 123.9687 | 21.1286 | 0.001943 | 0.001793 | -152.2225 | 0.23964 | 8.4017 | 0.003387 | 0.004403 | 0.004246 | 5.4e-05 | 0.00099 | 0.3402 | 1.3329 | 2.548 | 99.49 | 20.89 | 37.327 | 68.4523 | 103.2338 | 64.8334 | 0.009278 | 1.0446 | 0.024891 | 23.8295 | 0.000368 | 0.14256 | 0.000325 | 0.024533 | 0.024533 | 0.23883
2019-03-21 12:10 | 1553184594.3743 | 1481 | 0.20958 | 0.000451 | -4e-06 | 100.4597 | 100.5396 | 124.5552 | 21.1273 | 0.00193 | 0.001831 | -152.2297 | 0.2395 | 8.4021 | 0.003431 | 0.004659 | 0.004441 | 5.8e-05 | 0.000989 | 0.30646 | 1.3826 | 2.5482 | 99.49 | 20.89 | 37.326 | 68.1296 | 104.2961 | 65.7603 | 0.009278 | 1.0417 | 0.024916 | 23.9185 | -0.007116 | 0.14248 | 0.000164 | 0.012538 | 0.012538 | 0.46705
2019-03-21 12:10 | 1553184595.3731 | 1482 | 0.20959 | 0.000455 | -4e-06 | 100.3947 | 100.6425 | 125.169 | 21.1283 | 0.001936 | 0.001831 | -152.2214 | 0.23959 | 8.4006 | 0.003395 | 0.004742 | 0.004539 | 5.8e-05 | 0.000991 | 0.2975 | 1.3468 | 2.5479 | 99.49 | 20.89 | 37.324 | 68.2727 | 103.419 | 65.7487 | 0.009278 | 1.0417 | 0.024978 | 23.977 | -0.022047 | 0.14253 | 0.000105 | 0.008127 | 0.008127 | 0.72081
2019-03-21 12:10 | 1553184596.3736 | 1483 | 0.20956 | 0.000451 | -4e-06 | 100.4057 | 100.4913 | 125.3639 | 21.1282 | 0.001977 | 0.001876 | -152.2408 | 0.24047 | 8.4011 | 0.003448 | 0.005494 | 0.005357 | 5.5e-05 | 0.00099 | 0.29806 | 1.2825 | 2.5479 | 99.49 | 20.89 | 37.32 | 69.2724 | 104.6947 | 66.8278 | 0.009277 | 1.0384 | 0.024958 | 24.0338 | -0.007615 | 0.14306 | 0.000287 | 0.021781 | 0.021781 | 0.26995
2019-03-21 12:10 | 1553184597.3733 | 1484 | 0.20954 | 0.000453 | -4e-06 | 100.3432 | 100.5941 | 125.3772 | 21.1291 | 0.001918 | 0.001836 | -152.2587 | 0.24027 | 8.4015 | 0.003471 | 0.004746 | 0.004576 | 5.6e-05 | 0.000987 | 0.3027 | 1.2889 | 2.548 | 99.49 | 20.89 | 37.319 | 67.8554 | 105.2604 | 65.8758 | 0.009276 | 1.0414 | 0.024994 | 24.0014 | -0.022341 | 0.14294 | 0.000433 | 0.032741 | 0.032741 | 0.17943
2019-03-21 12:10 | 1553184598.3732 | 1485 | 0.20956 | 0.000454 | -4e-06 | 100.4358 | 100.588 | 125.4049 | 21.128 | 0.001911 | 0.00179 | -152.2521 | 0.23988 | 8.4019 | 0.003462 | 0.004942 | 0.004787 | 5.5e-05 | 0.00099 | 0.32947 | 1.2993 | 2.5479 | 99.49 | 20.89 | 37.318 | 67.6678 | 105.0509 | 64.7633 | 0.009276 | 1.0448 | 0.025038 | 23.9647 | -0.01354 | 0.14271 | 0.000263 | 0.020026 | 0.020026 | 0.29288
2019-03-21 12:10 | 1553184599.3738 | 1486 | 0.20956 | 0.00045 | -4e-06 | 100.4182 | 100.5587 | 125.1402 | 21.1277 | 0.001942 | 0.001867 | -152.2461 | 0.24121 | 8.4011 | 0.003437 | 0.005288 | 0.005065 | 5.9e-05 | 0.000994 | 0.34036 | 1.2843 | 2.5478 | 99.49 | 20.89 | 37.317 | 68.4358 | 104.4371 | 66.6222 | 0.009276 | 1.0391 | 0.024943 | 24.0051 | -0.012501 | 0.1435 | 0.000289 | 0.021854 | 0.021854 | 0.26986
2019-03-21 12:10 | 1553184600.3744 | 1487 | 0.20952 | 0.000452 | -4e-06 | 100.4004 | 100.4116 | 125.0749 | 21.1285 | 0.00195 | 0.001897 | -152.2486 | 0.24026 | 8.4007 | 0.003488 | 0.004945 | 0.004786 | 5.9e-05 | 0.000993 | 0.32189 | 1.2898 | 2.5479 | 99.49 | 20.89 | 37.315 | 68.6083 | 105.6675 | 67.3485 | 0.009275 | 1.0368 | 0.02491 | 24.0245 | -0.000996 | 0.14293 | 0.000521 | 0.039219 | 0.039219 | 0.14978
2019-03-21 12:10 | 1553184601.3745 | 1488 | 0.2096 | 0.000455 | -4e-06 | 100.434 | 100.5672 | 124.8836 | 21.1281 | 0.001963 | 0.00185 | -152.2546 | 0.23954 | 8.4027 | 0.003459 | 0.00498 | 0.004677 | 5.5e-05 | 0.000993 | 0.31598 | 1.3012 | 2.5479 | 99.49 | 20.89 | 37.313 | 68.9199 | 104.9567 | 66.2183 | 0.009275 | 1.0403 | 0.024932 | 23.9662 | -0.011851 | 0.1425 | 6.7e-05 | 0.005249 | 0.005249 | 1.1159
2019-03-21 12:10 | 1553184602.3759 | 1489 | 0.20958 | 0.000451 | -4e-06 | 100.4409 | 100.6694 | 124.6261 | 21.1284 | 0.001918 | 0.001817 | -152.2355 | 0.24012 | 8.4024 | 0.003456 | 0.005656 | 0.005558 | 5.5e-05 | 0.000987 | 0.3095 | 1.3281 | 2.5479 | 99.49 | 20.89 | 37.312 | 67.8578 | 104.8822 | 65.4139 | 0.009275 | 1.0428 | 0.024936 | 23.9131 | -0.020324 | 0.14284 | 0.000169 | 0.012924 | 0.012924 | 0.45426
2019-03-21 12:10 | 1553184603.3729 | 1490 | 0.20959 | 0.000453 | -4e-06 | 100.4248 | 100.3823 | 124.2961 | 21.1286 | 0.001989 | 0.001838 | -152.2242 | 0.23989 | 8.4008 | 0.003533 | 0.00534 | 0.005201 | 5.2e-05 | 0.000984 | 0.30574 | 1.3105 | 2.5479 | 99.49 | 20.89 | 37.311 | 69.566 | 106.7662 | 65.9229 | 0.009274 | 1.0412 | 0.024884 | 23.8994 | 0.00379 | 0.14271 | 8.7e-05 | 0.006757 | 0.006757 | 0.86811
2019-03-21 12:10 | 1553184604.3726 | 1491 | 0.20958 | 0.000451 | -3e-06 | 100.3959 | 100.4388 | 124.0123 | 21.129 | 0.001938 | 0.001812 | -152.2312 | 0.24041 | 8.4032 | 0.003577 | 0.004306 | 0.004118 | 5.7e-05 | 0.000987 | 0.34451 | 1.3198 | 2.5479 | 99.49 | 20.89 | 37.31 | 68.3363 | 107.8241 | 65.2935 | 0.009274 | 1.0431 | 0.024879 | 23.8499 | -0.003821 | 0.14302 | 0.000137 | 0.010432 | 0.010432 | 0.56348
2019-03-21 12:10 | 1553184605.3726 | 1492 | 0.20956 | 0.000453 | -4e-06 | 100.5082 | 100.4742 | 123.7532 | 21.1283 | 0.001938 | 0.001835 | -152.2461 | 0.23978 | 8.4034 | 0.003474 | 0.004774 | 0.004644 | 5.9e-05 | 0.000985 | 0.30934 | 1.2779 | 2.548 | 99.49 | 20.89 | 37.31 | 68.3297 | 105.3329 | 65.8565 | 0.009274 | 1.0414 | 0.024832 | 23.8448 | 0.003027 | 0.14265 | 0.000276 | 0.020825 | 0.020825 | 0.28152
2019-03-21 12:10 | 1553184606.3741 | 1493 | 0.2096 | 0.000455 | -4e-06 | 100.4272 | 100.4684 | 123.9811 | 21.1288 | 0.001873 | 0.001856 | -152.2304 | 0.24013 | 8.4009 | 0.003557 | 0.00431 | 0.004154 | 5.6e-05 | 0.000986 | 0.33251 | 1.3089 | 2.548 | 99.49 | 20.89 | 37.308 | 66.7582 | 107.3397 | 66.3486 | 0.009274 | 1.0399 | 0.024837 | 23.8841 | -0.003673 | 0.14285 | 4.9e-05 | 0.003892 | 0.003892 | 1.5086
2019-03-21 12:10 | 1553184607.3738 | 1494 | 0.20963 | 0.000453 | -4e-06 | 100.4147 | 100.604 | 123.9947 | 21.1284 | 0.001957 | 0.001819 | -152.2189 | 0.24028 | 8.4023 | 0.003499 | 0.004504 | 0.004313 | 5.8e-05 | 0.000987 | 0.33395 | 1.332 | 2.548 | 99.49 | 20.89 | 37.308 | 68.7882 | 105.9475 | 65.463 | 0.009274 | 1.0426 | 0.024871 | 23.8542 | -0.016847 | 0.14294 | -0.000114 | -0.008249 | -0.008249 | -0.71221
2019-03-21 12:10 | 1553184608.3742 | 1495 | 0.20961 | 0.000453 | -4e-06 | 100.4641 | 100.4943 | 124.0832 | 21.1294 | 0.001982 | 0.001866 | -152.2281 | 0.23893 | 8.4004 | 0.003414 | 0.004382 | 0.004189 | 5.7e-05 | 0.000985 | 0.31061 | 1.3283 | 2.548 | 99.49 | 20.89 | 37.308 | 69.399 | 103.8656 | 66.5901 | 0.009274 | 1.0392 | 0.024838 | 23.9024 | -0.002687 | 0.14214 | 7e-06 | 0.000802 | 0.000802 | 7.2814
2019-03-21 12:10 | 1553184609.373 | 1496 | 0.20962 | 0.000456 | -4e-06 | 100.381 | 100.3554 | 124.3217 | 21.1279 | 0.001939 | 0.001836 | -152.2238 | 0.23956 | 8.3995 | 0.003392 | 0.005149 | 0.004963 | 6.1e-05 | 0.000988 | 0.33069 | 1.3041 | 2.548 | 99.49 | 20.89 | 37.306 | 68.3425 | 103.3479 | 65.8693 | 0.009273 | 1.0414 | 0.024889 | 23.8999 | 0.002286 | 0.14251 | -7e-05 | -0.004984 | -0.004984 | -1.1753
2019-03-21 12:10 | 1553184610.374 | 1497 | 0.2096 | 0.000449 | -4e-06 | 100.4829 | 100.4504 | 124.2623 | 21.1292 | 0.001999 | 0.001802 | -152.2303 | 0.23934 | 8.4021 | 0.003438 | 0.00448 | 0.004278 | 5.4e-05 | 0.000987 | 0.32421 | 1.2876 | 2.5479 | 99.49 | 20.89 | 37.306 | 69.7888 | 104.448 | 65.0572 | 0.009273 | 1.0439 | 0.024913 | 23.8656 | 0.002895 | 0.14238 | 5.8e-05 | 0.004593 | 0.004593 | 1.2742
2019-03-21 12:10 | 1553184611.3735 | 1498 | 0.20959 | 0.000449 | -4e-06 | 100.5303 | 100.5464 | 124.2926 | 21.1293 | 0.001951 | 0.001847 | -152.2213 | 0.24019 | 8.4036 | 0.003458 | 0.004358 | 0.004179 | 6e-05 | 0.000986 | 0.30917 | 1.3054 | 2.5478 | 99.49 | 20.9 | 37.305 | 68.6323 | 104.9491 | 66.1429 | 0.009279 | 1.0405 | 0.024876 | 23.9068 | -0.001437 | 0.14289 | 0.000121 | 0.009287 | 0.009287 | 0.63236
2019-03-21 12:10 | 1553184612.3745 | 1499 | 0.2096 | 0.000453 | -4e-06 | 100.5047 | 100.5757 | 124.5008 | 21.1295 | 0.001932 | 0.001839 | -152.2385 | 0.23914 | 8.4024 | 0.003423 | 0.004337 | 0.004185 | 5.7e-05 | 0.000989 | 0.28694 | 1.3309 | 2.548 | 99.49 | 20.9 | 37.304 | 68.177 | 104.0923 | 65.9474 | 0.009278 | 1.0411 | 0.024904 | 23.9199 | -0.006322 | 0.14226 | 3.1e-05 | 0.00257 | 0.00257 | 2.275
2019-03-21 12:10 | 1553184613.3742 | 1500 | 0.20961 | 0.000453 | -4e-06 | 100.4892 | 100.4248 | 124.827 | 21.1295 | 0.001968 | 0.001889 | -152.2466 | 0.23921 | 8.4022 | 0.003426 | 0.005444 | 0.005261 | 6e-05 | 0.000985 | 0.29247 | 1.3517 | 2.548 | 99.49 | 20.9 | 37.304 | 69.0513 | 104.1764 | 67.1551 | 0.009278 | 1.0374 | 0.024892 | 23.9938 | 0.005727 | 0.14231 | -4e-05 | -0.00274 | -0.00274 | -2.1348
2019-03-21 12:10 | 1553184614.3741 | 1501 | 0.2096 | 0.000449 | -4e-06 | 100.4915 | 100.525 | 124.9395 | 21.1293 | 0.001957 | 0.00181 | -152.2379 | 0.23955 | 8.4043 | 0.003448 | 0.005012 | 0.004816 | 6e-05 | 0.000988 | 0.30845 | 1.3088 | 2.5478 | 99.49 | 20.9 | 37.301 | 68.7789 | 104.7076 | 65.2392 | 0.009278 | 1.0433 | 0.024974 | 23.937 | -0.002975 | 0.14251 | 5.5e-05 | 0.004364 | 0.004364 | 1.3421
2019-03-21 12:10 | 1553184615.3737 | 1502 | 0.2096 | 0.000452 | -4e-06 | 100.457 | 100.6881 | 125.016 | 21.128 | 0.002003 | 0.001833 | -152.2353 | 0.23885 | 8.4052 | 0.003469 | 0.004519 | 0.004287 | 5.7e-05 | 0.000984 | 0.3131 | 1.3169 | 2.548 | 99.49 | 20.9 | 37.3 | 69.8867 | 105.2011 | 65.796 | 0.009277 | 1.0416 | 0.024961 | 23.964 | -0.020556 | 0.14209 | 3.5e-05 | 0.002888 | 0.002888 | 2.0224
2019-03-21 12:10 | 1553184616.3733 | 1503 | 0.20963 | 0.000449 | -4e-06 | 100.5109 | 100.6333 | 124.52 | 21.1284 | 0.001961 | 0.001819 | -152.2448 | 0.23905 | 8.4002 | 0.003503 | 0.006048 | 0.005861 | 5.7e-05 | 0.000983 | 0.30198 | 1.3182 | 2.5481 | 99.49 | 20.9 | 37.296 | 68.8779 | 106.0229 | 65.464 | 0.009276 | 1.0426 | 0.024923 | 23.9047 | -0.010883 | 0.14221 | -0.00015 | -0.011029 | -0.011029 | -0.52996
2019-03-21 12:10 | 1553184617.4361 | 1504 | 0.2096 | 0.000451 | -3e-06 | 100.5425 | 100.6592 | 123.9783 | 21.1287 | 0.001956 | 0.001881 | -152.2678 | 0.2395 | 8.4036 | 0.00348 | 0.004796 | 0.004567 | 5.6e-05 | 0.000982 | 0.31404 | 1.3479 | 2.548 | 99.49 | 20.9 | 37.295 | 68.7683 | 105.485 | 66.9593 | 0.009276 | 1.038 | 0.024814 | 23.9053 | -0.010373 | 0.14248 | 6.9e-05 | 0.005341 | 0.005341 | 1.0965
2019-03-21 12:10 | 1553184618.3821 | 1505 | 0.20957 | 0.000449 | -4e-06 | 100.5547 | 100.5468 | 123.9276 | 21.1298 | 0.001989 | 0.001829 | -152.2543 | 0.23941 | 8.4012 | 0.003459 | 0.005405 | 0.005305 | 5.5e-05 | 0.000988 | 0.30702 | 1.2896 | 2.5479 | 99.49 | 20.9 | 37.293 | 69.5632 | 104.9677 | 65.7025 | 0.009276 | 1.0419 | 0.024855 | 23.8562 | 0.000706 | 0.14242 | 0.000216 | 0.016403 | 0.016403 | 0.35687
2019-03-21 12:10 | 1553184619.3811 | 1506 | 0.20963 | 0.000455 | -4e-06 | 100.5404 | 100.6295 | 123.8518 | 21.1295 | 0.001941 | 0.001858 | -152.2607 | 0.23952 | 8.4046 | 0.003502 | 0.004836 | 0.00469 | 5.8e-05 | 0.000992 | 0.30574 | 1.2948 | 2.5479 | 99.49 | 20.9 | 37.289 | 68.4094 | 106.0099 | 66.4053 | 0.009275 | 1.0397 | 0.024822 | 23.8736 | -0.007926 | 0.14249 | -0.00011 | -0.007921 | -0.007921 | -0.73939
2019-03-21 12:10 | 1553184620.38 | 1507 | 0.20962 | 0.000454 | -3e-06 | 100.5419 | 100.5801 | 124.3409 | 21.1286 | 0.001883 | 0.001833 | -152.2479 | 0.24008 | 8.4016 | 0.00343 | 0.004835 | 0.004642 | 5.6e-05 | 0.000989 | 0.35513 | 1.3111 | 2.5481 | 99.49 | 20.9 | 37.288 | 67.0083 | 104.2722 | 65.7918 | 0.009274 | 1.0416 | 0.024894 | 23.8991 | -0.003402 | 0.14282 | -4.5e-05 | -0.003191 | -0.003191 | -1.8396
2019-03-21 12:10 | 1553184621.3802 | 1508 | 0.2096 | 0.00045 | -4e-06 | 100.5771 | 100.5767 | 124.332 | 21.1296 | 0.001917 | 0.001898 | -152.2688 | 0.23978 | 8.402 | 0.003473 | 0.005862 | 0.005588 | 5.7e-05 | 0.000987 | 0.29336 | 1.3206 | 2.5479 | 99.49 | 20.9 | 37.283 | 67.8152 | 105.2987 | 67.369 | 0.009273 | 1.0368 | 0.024835 | 23.9537 | 2.8e-05 | 0.14264 | 4e-05 | 0.003222 | 0.003222 | 1.8194
2019-03-21 12:10 | 1553184622.3801 | 1509 | 0.20962 | 0.000452 | -3e-06 | 100.5776 | 100.7109 | 124.4088 | 21.1282 | 0.001942 | 0.001779 | -152.2662 | 0.23919 | 8.4036 | 0.003518 | 0.00513 | 0.005046 | 5.7e-05 | 0.000986 | 0.30646 | 1.3504 | 2.5481 | 99.49 | 20.9 | 37.282 | 68.4308 | 106.4014 | 64.5013 | 0.009273 | 1.0456 | 0.024948 | 23.86 | -0.011845 | 0.1423 | -4.9e-05 | -0.003471 | -0.003471 | -1.6848
2019-03-21 12:10 | 1553184623.3812 | 1510 | 0.20959 | 0.000451 | -4e-06 | 100.5145 | 100.682 | 124.1398 | 21.1293 | 0.001967 | 0.001834 | -152.2719 | 0.23905 | 8.4037 | 0.003495 | 0.005081 | 0.004879 | 5.6e-05 | 0.000988 | 0.3131 | 1.3694 | 2.548 | 99.49 | 20.9 | 37.281 | 69.0378 | 105.8491 | 65.8355 | 0.009273 | 1.0415 | 0.024872 | 23.8813 | -0.014893 | 0.14221 | 9e-05 | 0.006961 | 0.006961 | 0.83961
2019-03-21 12:10 | 1553184624.3812 | 1511 | 0.2096 | 0.000458 | -4e-06 | 100.5896 | 100.6885 | 123.967 | 21.1283 | 0.001946 | 0.001813 | -152.2759 | 0.23974 | 8.4023 | 0.003574 | 0.004935 | 0.004763 | 5.4e-05 | 0.000987 | 0.30198 | 1.3224 | 2.548 | 99.49 | 20.9 | 37.277 | 68.5175 | 107.7597 | 65.3312 | 0.009272 | 1.043 | 0.024873 | 23.8469 | -0.008789 | 0.14262 | 4.1e-05 | 0.003281 | 0.003281 | 1.7864
2019-03-21 12:10 | 1553184625.3812 | 1512 | 0.20961 | 0.000455 | -4e-06 | 100.5759 | 100.6173 | 123.8842 | 21.129 | 0.001965 | 0.001838 | -152.2717 | 0.23984 | 8.4015 | 0.003574 | 0.004696 | 0.004597 | 5.6e-05 | 0.000986 | 0.31045 | 1.3129 | 2.548 | 99.49 | 20.9 | 37.275 | 68.9736 | 107.7652 | 65.9296 | 0.009271 | 1.0412 | 0.024843 | 23.86 | -0.003682 | 0.14268 | -3.2e-05 | -0.002134 | -0.002134 | -2.7485
2019-03-21 12:10 | 1553184626.3817 | 1513 | 0.2096 | 0.000454 | -3e-06 | 100.5666 | 100.604 | 124.2996 | 21.1287 | 0.001997 | 0.001831 | -152.2747 | 0.24 | 8.4023 | 0.003538 | 0.004734 | 0.004564 | 5.4e-05 | 0.000982 | 0.32261 | 1.3186 | 2.5479 | 99.49 | 20.9 | 37.272 | 69.7454 | 106.8899 | 65.7638 | 0.00927 | 1.0417 | 0.02489 | 23.8941 | -0.003322 | 0.14278 | 8.2e-05 | 0.006345 | 0.006345 | 0.92481
2019-03-21 12:10 | 1553184627.3804 | 1514 | 0.20957 | 0.000453 | -3e-06 | 100.5988 | 100.7232 | 124.7658 | 21.1302 | 0.001925 | 0.001808 | -152.2665 | 0.2402 | 8.4052 | 0.003527 | 0.004361 | 0.00426 | 5.8e-05 | 0.000983 | 0.28816 | 1.2787 | 2.5479 | 99.48 | 20.9 | 37.269 | 68.0171 | 106.6072 | 65.212 | 0.009271 | 1.0434 | 0.024957 | 23.9194 | -0.011053 | 0.14289 | 0.00024 | 0.018191 | 0.018191 | 0.32285
2019-03-21 12:10 | 1553184628.3808 | 1515 | 0.20962 | 0.000455 | -4e-06 | 100.6197 | 100.7879 | 124.4639 | 21.1292 | 0.001954 | 0.001774 | -152.2678 | 0.23993 | 8.4023 | 0.003598 | 0.004721 | 0.004453 | 5.7e-05 | 0.000984 | 0.29911 | 1.3275 | 2.5481 | 99.48 | 20.9 | 37.268 | 68.7048 | 108.3476 | 64.3766 | 0.00927 | 1.046 | 0.024958 | 23.8609 | -0.014944 | 0.14274 | -5e-05 | -0.003485 | -0.003485 | -1.6832
2019-03-21 12:10 | 1553184629.3809 | 1516 | 0.20958 | 0.000451 | -4e-06 | 100.6664 | 100.8111 | 124.3496 | 21.1308 | 0.001902 | 0.001857 | -152.2743 | 0.24013 | 8.4047 | 0.00354 | 0.004402 | 0.004233 | 5.9e-05 | 0.000982 | 0.29894 | 1.2953 | 2.5479 | 99.48 | 20.9 | 37.264 | 67.4591 | 106.9297 | 66.3933 | 0.009269 | 1.0398 | 0.024872 | 23.9211 | -0.012841 | 0.14285 | 0.000175 | 0.01331 | 0.01331 | 0.44112
2019-03-21 12:10 | 1553184630.3806 | 1517 | 0.20956 | 0.000453 | -4e-06 | 100.6819 | 100.6622 | 123.99 | 21.1303 | 0.001954 | 0.001833 | -152.2886 | 0.23905 | 8.4023 | 0.003549 | 0.004272 | 0.004127 | 5.6e-05 | 0.00099 | 0.3366 | 1.3185 | 2.5479 | 99.48 | 20.9 | 37.262 | 68.7136 | 107.1591 | 65.8115 | 0.009269 | 1.0415 | 0.024858 | 23.866 | 0.001751 | 0.14221 | 0.000293 | 0.022106 | 0.022106 | 0.2644
2019-03-21 12:10 | 1553184631.3818 | 1518 | 0.20959 | 0.000449 | -4e-06 | 100.6739 | 100.7269 | 124.1846 | 21.1294 | 0.001918 | 0.001838 | -152.2762 | 0.2398 | 8.4034 | 0.003466 | 0.004845 | 0.004626 | 5.5e-05 | 0.00099 | 0.30541 | 1.2967 | 2.5478 | 99.48 | 20.91 | 37.261 | 67.8595 | 105.1436 | 65.9331 | 0.009274 | 1.0412 | 0.024873 | 23.889 | -0.004713 | 0.14266 | 9.7e-05 | 0.007505 | 0.007505 | 0.78121
2019-03-21 12:11 | 1553184632.381 | 1519 | 0.20957 | 0.000452 | -4e-06 | 100.6816 | 100.8693 | 124.3121 | 21.1298 | 0.001908 | 0.001812 | -152.2746 | 0.23898 | 8.4018 | 0.003424 | 0.004561 | 0.004394 | 5.6e-05 | 0.000993 | 0.33771 | 1.311 | 2.548 | 99.48 | 20.91 | 37.259 | 67.6024 | 104.1192 | 65.3099 | 0.009274 | 1.0431 | 0.024908 | 23.8793 | -0.016657 | 0.14217 | 0.000244 | 0.01854 | 0.01854 | 0.31517
2019-03-21 12:11 | 1553184633.3807 | 1520 | 0.20957 | 0.000456 | -3e-06 | 100.7213 | 100.9718 | 124.3811 | 21.1298 | 0.001936 | 0.001858 | -152.2653 | 0.23833 | 8.4009 | 0.003539 | 0.004932 | 0.004769 | 5.8e-05 | 0.000993 | 0.30525 | 1.3546 | 2.5481 | 99.49 | 20.91 | 37.259 | 68.2874 | 106.9011 | 66.3994 | 0.009273 | 1.0397 | 0.024875 | 23.9243 | -0.022224 | 0.14178 | 0.000214 | 0.01618 | 0.01618 | 0.36016
2019-03-21 12:11 | 1553184634.3808 | 1521 | 0.20959 | 0.000454 | -4e-06 | 100.712 | 100.7811 | 124.5182 | 21.1296 | 0.00191 | 0.001779 | -152.262 | 0.23983 | 8.4018 | 0.00348 | 0.004624 | 0.004445 | 5.5e-05 | 0.000995 | 0.31686 | 1.3873 | 2.5479 | 99.49 | 20.91 | 37.256 | 67.6647 | 105.4741 | 64.5043 | 0.009272 | 1.0456 | 0.024959 | 23.8706 | -0.006132 | 0.14267 | 0.000134 | 0.010263 | 0.010263 | 0.57139
2019-03-21 12:11 | 1553184635.3799 | 1522 | 0.20958 | 0.000457 | -4e-06 | 100.7597 | 100.9333 | 124.566 | 21.1302 | 0.00196 | 0.001802 | -152.2495 | 0.23958 | 8.4017 | 0.003476 | 0.005295 | 0.005161 | 5.4e-05 | 0.000997 | 0.31006 | 1.3559 | 2.5479 | 99.49 | 20.91 | 37.255 | 68.8594 | 105.3898 | 65.0659 | 0.009272 | 1.0438 | 0.024943 | 23.8951 | -0.015401 | 0.14253 | 0.000201 | 0.01529 | 0.01529 | 0.38311
2019-03-21 12:11 | 1553184636.3803 | 1523 | 0.20958 | 0.000452 | -4e-06 | 100.7874 | 100.9057 | 124.1341 | 21.1295 | 0.001894 | 0.001876 | -152.234 | 0.23913 | 8.4026 | 0.003454 | 0.005144 | 0.004919 | 5.5e-05 | 0.000994 | 0.34916 | 1.34 | 2.5479 | 99.49 | 20.91 | 37.254 | 67.2785 | 104.847 | 66.8305 | 0.009272 | 1.0384 | 0.024835 | 23.9157 | -0.010497 | 0.14226 | 0.000197 | 0.014952 | 0.014952 | 0.39104
2019-03-21 12:11 | 1553184637.3809 | 1524 | 0.20961 | 0.000455 | -4e-06 | 100.8386 | 101.1288 | 123.7159 | 21.1299 | 0.001871 | 0.001797 | -152.2315 | 0.23937 | 8.4051 | 0.003459 | 0.004797 | 0.004594 | 5.3e-05 | 0.000993 | 0.30685 | 1.3178 | 2.548 | 99.49 | 20.91 | 37.252 | 66.721 | 104.963 | 64.9487 | 0.009271 | 1.0442 | 0.024862 | 23.8093 | -0.025703 | 0.1424 | 3e-06 | 0.000435 | 0.000435 | 13.4531
2019-03-21 12:11 | 1553184638.3817 | 1525 | 0.20962 | 0.000454 | -3e-06 | 100.8232 | 100.9963 | 124.0461 | 21.1303 | 0.001939 | 0.001858 | -152.2511 | 0.23966 | 8.4044 | 0.003503 | 0.004973 | 0.004833 | 5.6e-05 | 0.000994 | 0.29574 | 1.3127 | 2.548 | 99.49 | 20.91 | 37.252 | 68.3479 | 106.0328 | 66.4092 | 0.009271 | 1.0397 | 0.024841 | 23.8924 | -0.015351 | 0.14257 | -4e-05 | -0.002803 | -0.002803 | -2.0907
2019-03-21 12:11 | 1553184639.3806 | 1526 | 0.20959 | 0.000453 | -4e-06 | 100.8294 | 100.9504 | 124.5104 | 21.1298 | 0.001944 | 0.001843 | -152.2505 | 0.23982 | 8.4035 | 0.003498 | 0.004115 | 0.003981 | 5.3e-05 | 0.000998 | 0.30934 | 1.3106 | 2.5479 | 99.48 | 20.91 | 37.25 | 68.4636 | 105.9103 | 66.0374 | 0.009272 | 1.0409 | 0.024901 | 23.924 | -0.010723 | 0.14267 | 0.000107 | 0.008213 | 0.008213 | 0.71393
2019-03-21 12:11 | 1553184640.3805 | 1527 | 0.20956 | 0.000451 | -4e-06 | 100.8503 | 101.1677 | 124.6674 | 21.1301 | 0.001974 | 0.001818 | -152.2511 | 0.2395 | 8.4024 | 0.00354 | 0.004393 | 0.004189 | 5.4e-05 | 0.000997 | 0.31974 | 1.3136 | 2.548 | 99.48 | 20.91 | 37.248 | 69.204 | 106.9424 | 65.4469 | 0.009271 | 1.0427 | 0.024939 | 23.9182 | -0.028106 | 0.14248 | 0.000282 | 0.021333 | 0.021333 | 0.2745
2019-03-21 12:11 | 1553184641.3805 | 1528 | 0.20957 | 0.00045 | -4e-06 | 100.9674 | 101.1115 | 124.6853 | 21.1306 | 0.001956 | 0.001823 | -152.2551 | 0.23947 | 8.4045 | 0.003392 | 0.005515 | 0.005308 | 5.4e-05 | 0.000995 | 0.31166 | 1.3011 | 2.5479 | 99.48 | 20.91 | 37.246 | 68.7601 | 103.3363 | 65.5542 | 0.009271 | 1.0423 | 0.024937 | 23.9237 | -0.012757 | 0.14246 | 0.000218 | 0.016594 | 0.016594 | 0.35284
2019-03-21 12:11 | 1553184642.3811 | 1529 | 0.20955 | 0.00045 | -3e-06 | 100.9245 | 101.2317 | 124.5045 | 21.1306 | 0.00192 | 0.001826 | -152.2521 | 0.23942 | 8.4044 | 0.003552 | 0.003977 | 0.003853 | 5.3e-05 | 0.000997 | 0.30989 | 1.3278 | 2.5481 | 99.48 | 20.91 | 37.242 | 67.8996 | 107.2326 | 65.6463 | 0.00927 | 1.0421 | 0.024915 | 23.9096 | -0.027188 | 0.14243 | 0.00037 | 0.027921 | 0.027921 | 0.20966
2019-03-21 12:11 | 1553184643.3813 | 1530 | 0.20956 | 0.000454 | -4e-06 | 100.8753 | 101.1285 | 124.4538 | 21.1307 | 0.001903 | 0.001827 | -152.2588 | 0.24015 | 8.4011 | 0.003402 | 0.004869 | 0.00467 | 5.1e-05 | 0.000997 | 0.33157 | 1.298 | 2.548 | 99.48 | 20.91 | 37.242 | 67.4823 | 103.5973 | 65.6534 | 0.00927 | 1.042 | 0.02491 | 23.905 | -0.022423 | 0.14286 | 0.000302 | 0.022806 | 0.022806 | 0.25746
2019-03-21 12:11 | 1553184644.3806 | 1531 | 0.20957 | 0.000452 | -4e-06 | 100.9182 | 101.0314 | 124.1752 | 21.1304 | 0.001927 | 0.0018 | -152.2462 | 0.23913 | 8.4025 | 0.003359 | 0.005113 | 0.004868 | 5.2e-05 | 0.000996 | 0.2896 | 1.3111 | 2.5479 | 99.48 | 20.92 | 37.236 | 68.0617 | 102.534 | 65.0146 | 0.009274 | 1.044 | 0.024905 | 23.8557 | -0.01003 | 0.14226 | 0.000274 | 0.020788 | 0.020788 | 0.28127
2019-03-21 12:11 | 1553184645.3804 | 1532 | 0.20955 | 0.000454 | -3e-06 | 100.9775 | 101.0335 | 124.1757 | 21.1307 | 0.001955 | 0.001851 | -152.2628 | 0.24049 | 8.4028 | 0.00349 | 0.004996 | 0.004862 | 5.3e-05 | 0.000996 | 0.30558 | 1.3201 | 2.5479 | 99.48 | 20.92 | 37.232 | 68.7303 | 105.7237 | 66.2388 | 0.009273 | 1.0402 | 0.024861 | 23.8989 | -0.004957 | 0.14306 | 0.000354 | 0.026623 | 0.026623 | 0.22086
2019-03-21 12:11 | 1553184646.3808 | 1533 | 0.20962 | 0.000454 | -4e-06 | 101.0234 | 101.2566 | 124.5576 | 21.1306 | 0.001901 | 0.001862 | -152.2648 | 0.2401 | 8.4037 | 0.00349 | 0.004801 | 0.004607 | 5.2e-05 | 0.000998 | 0.30254 | 1.3362 | 2.5479 | 99.48 | 20.92 | 37.225 | 67.4489 | 105.7292 | 66.5134 | 0.009271 | 1.0394 | 0.024889 | 23.9453 | -0.020622 | 0.14283 | -5e-05 | -0.003475 | -0.003475 | -1.6894
2019-03-21 12:11 | 1553184647.3805 | 1534 | 0.20959 | 0.000451 | -4e-06 | 100.9921 | 101.1929 | 124.8029 | 21.1312 | 0.001903 | 0.00181 | -152.2646 | 0.23863 | 8.4034 | 0.003472 | 0.004934 | 0.004717 | 5e-05 | 0.000997 | 0.33069 | 1.4223 | 2.5479 | 99.48 | 20.92 | 37.221 | 67.4961 | 105.2901 | 65.263 | 0.00927 | 1.0432 | 0.024959 | 23.9247 | -0.017764 | 0.14196 | 0.000102 | 0.007891 | 0.007891 | 0.73937
2019-03-21 12:11 | 1553184648.3815 | 1535 | 0.20961 | 0.000452 | -4e-06 | 100.9435 | 101.2596 | 125.0735 | 21.1306 | 0.001922 | 0.001837 | -152.2544 | 0.24025 | 8.4014 | 0.003532 | 0.004722 | 0.004535 | 4.9e-05 | 0.000993 | 0.31332 | 1.3367 | 2.548 | 99.48 | 20.92 | 37.218 | 67.9378 | 106.7289 | 65.9015 | 0.009269 | 1.0413 | 0.024963 | 23.9732 | -0.027965 | 0.14293 | 2.6e-05 | 0.002161 | 0.002161 | 2.7186
2019-03-21 12:11 | 1553184649.3819 | 1536 | 0.20962 | 0.000454 | -3e-06 | 101.0618 | 101.371 | 125.1528 | 21.1306 | 0.001953 | 0.001836 | -152.2444 | 0.2391 | 8.4028 | 0.003446 | 0.004773 | 0.004686 | 5e-05 | 0.00099 | 0.32388 | 1.3001 | 2.5481 | 99.48 | 20.92 | 37.212 | 68.7023 | 104.6469 | 65.8703 | 0.009268 | 1.0414 | 0.024972 | 23.9797 | -0.027323 | 0.14224 | -7e-05 | -0.005051 | -0.005051 | -1.1575
2019-03-21 12:11 | 1553184650.3796 | 1537 | 0.2096 | 0.000451 | -3e-06 | 100.9912 | 101.3526 | 125.0536 | 21.1301 | 0.001937 | 0.001852 | -152.2398 | 0.23972 | 8.4026 | 0.003455 | 0.004603 | 0.004434 | 4.9e-05 | 0.000988 | 0.31349 | 1.2926 | 2.548 | 99.48 | 20.92 | 37.208 | 68.3024 | 104.8704 | 66.2631 | 0.009267 | 1.0402 | 0.024947 | 23.9841 | -0.031952 | 0.14261 | 5e-05 | 0.00396 | 0.00396 | 1.4802
2019-03-21 12:11 | 1553184651.3807 | 1538 | 0.2096 | 0.000451 | -4e-06 | 101.0713 | 101.1499 | 124.8745 | 21.1307 | 0.001957 | 0.001833 | -152.2377 | 0.23926 | 8.4018 | 0.003417 | 0.004717 | 0.004505 | 4.7e-05 | 0.000987 | 0.28854 | 1.2815 | 2.5479 | 99.48 | 20.93 | 37.201 | 68.7928 | 103.9471 | 65.8102 | 0.009271 | 1.0416 | 0.024946 | 23.9509 | -0.006954 | 0.14234 | 5.8e-05 | 0.00458 | 0.00458 | 1.2773
2019-03-21 12:11 | 1553184652.3802 | 1539 | 0.20962 | 0.000454 | -4e-06 | 101.1241 | 101.1796 | 124.8515 | 21.1299 | 0.001933 | 0.001824 | -152.2177 | 0.23931 | 8.4036 | 0.00344 | 0.004449 | 0.004288 | 5.1e-05 | 0.00099 | 0.3235 | 1.3453 | 2.548 | 99.48 | 20.93 | 37.198 | 68.203 | 104.5039 | 65.5809 | 0.00927 | 1.0423 | 0.024952 | 23.9406 | -0.004908 | 0.14236 | -3.4e-05 | -0.002326 | -0.002326 | -2.5155
2019-03-21 12:11 | 1553184653.3798 | 1540 | 0.20961 | 0.000456 | -3e-06 | 101.099 | 101.3165 | 125.1324 | 21.1313 | 0.001952 | 0.0018 | -152.2099 | 0.23963 | 8.4009 | 0.003464 | 0.00443 | 0.004247 | 4.9e-05 | 0.000994 | 0.29159 | 1.3587 | 2.5479 | 99.48 | 20.93 | 37.196 | 68.6593 | 105.0982 | 65.0085 | 0.00927 | 1.044 | 0.025002 | 23.9473 | -0.019218 | 0.14255 | -2.8e-05 | -0.001912 | -0.001912 | -3.064
2019-03-21 12:11 | 1553184654.3808 | 1541 | 0.20963 | 0.000452 | -3e-06 | 101.1521 | 101.4623 | 125.0594 | 21.1303 | 0.001894 | 0.001818 | -152.2368 | 0.23911 | 8.4011 | 0.00351 | 0.004896 | 0.004772 | 5.2e-05 | 0.000994 | 0.32964 | 1.3482 | 2.5481 | 99.48 | 20.93 | 37.191 | 67.2838 | 106.2004 | 65.4478 | 0.009268 | 1.0427 | 0.024978 | 23.9559 | -0.027387 | 0.14225 | -0.000149 | -0.010979 | -0.010979 | -0.53251
2019-03-21 12:11 | 1553184655.3813 | 1542 | 0.20963 | 0.000455 | -4e-06 | 101.1577 | 101.291 | 124.9087 | 21.1303 | 0.001906 | 0.001835 | -152.2505 | 0.24042 | 8.4028 | 0.003497 | 0.004804 | 0.004607 | 5.1e-05 | 0.000995 | 0.31061 | 1.3377 | 2.5479 | 99.48 | 20.92 | 37.19 | 67.5685 | 105.895 | 65.8418 | 0.009262 | 1.0415 | 0.024948 | 23.9553 | -0.011771 | 0.14303 | -9.5e-05 | -0.006913 | -0.006913 | -0.85039
2019-03-21 12:11 | 1553184656.3809 | 1543 | 0.20963 | 0.00045 | -4e-06 | 101.1986 | 101.4071 | 124.651 | 21.1308 | 0.001899 | 0.001853 | -152.2688 | 0.23974 | 8.4017 | 0.003544 | 0.005126 | 0.004931 | 5e-05 | 0.000994 | 0.31332 | 1.3492 | 2.548 | 99.48 | 20.93 | 37.188 | 67.3882 | 107.0305 | 66.2919 | 0.009268 | 1.0401 | 0.024906 | 23.9465 | -0.018412 | 0.14262 | -0.000106 | -0.007712 | -0.007712 | -0.76008
2019-03-21 12:11 | 1553184657.38 | 1544 | 0.20962 | 0.000453 | -4e-06 | 101.1834 | 101.3148 | 124.6266 | 21.131 | 0.001913 | 0.001832 | -152.2655 | 0.24017 | 8.4019 | 0.003477 | 0.004607 | 0.004509 | 4.8e-05 | 0.000997 | 0.30486 | 1.2934 | 2.5479 | 99.48 | 20.93 | 37.188 | 67.7276 | 105.3937 | 65.7848 | 0.009268 | 1.0416 | 0.024922 | 23.9263 | -0.01161 | 0.14288 | -6.4e-05 | -0.00456 | -0.00456 | -1.2878
2019-03-21 12:11 | 1553184658.3801 | 1545 | 0.20965 | 0.000459 | -4e-06 | 101.1616 | 101.4255 | 124.2219 | 21.1308 | 0.001952 | 0.001803 | -152.2615 | 0.24029 | 8.4016 | 0.003526 | 0.005044 | 0.004909 | 5e-05 | 0.000994 | 0.34811 | 1.2581 | 2.5478 | 99.48 | 20.93 | 37.186 | 68.6744 | 106.599 | 65.0838 | 0.009267 | 1.0438 | 0.024908 | 23.8627 | -0.0233 | 0.14295 | -0.000265 | -0.019569 | -0.019569 | -0.30023
2019-03-21 12:11 | 1553184659.3801 | 1546 | 0.20964 | 0.000452 | -4e-06 | 101.2281 | 101.3744 | 124.1029 | 21.1306 | 0.001988 | 0.001883 | -152.2792 | 0.24028 | 8.4022 | 0.003561 | 0.004609 | 0.004448 | 5e-05 | 0.000996 | 0.33716 | 1.2896 | 2.548 | 99.48 | 20.93 | 37.184 | 69.529 | 107.454 | 67.0056 | 0.009267 | 1.0379 | 0.024825 | 23.9189 | -0.012921 | 0.14294 | -0.000159 | -0.011599 | -0.011599 | -0.50649
2019-03-21 12:11 | 1553184660.3801 | 1547 | 0.20962 | 0.000454 | -4e-06 | 101.2719 | 101.451 | 124.1522 | 21.1317 | 0.001935 | 0.001826 | -152.2938 | 0.23989 | 8.4036 | 0.003531 | 0.00471 | 0.00451 | 5.1e-05 | 0.000995 | 0.28528 | 1.3352 | 2.548 | 99.48 | 20.93 | 37.183 | 68.2704 | 106.718 | 65.6349 | 0.009266 | 1.0421 | 0.02488 | 23.8754 | -0.015811 | 0.14271 | -3.7e-05 | -0.00258 | -0.00258 | -2.2736
2019-03-21 12:11 | 1553184661.3802 | 1548 | 0.20962 | 0.000454 | -4e-06 | 101.2334 | 101.4813 | 123.9466 | 21.1312 | 0.001956 | 0.001876 | -152.2916 | 0.23972 | 8.4025 | 0.003545 | 0.00494 | 0.004726 | 4.9e-05 | 0.000998 | 0.31022 | 1.3615 | 2.5478 | 99.48 | 20.93 | 37.182 | 68.7558 | 107.0522 | 66.8487 | 0.009266 | 1.0384 | 0.024815 | 23.8983 | -0.021878 | 0.14261 | -7.2e-05 | -0.005169 | -0.005169 | -1.134
2019-03-21 12:11 | 1553184662.3816 | 1549 | 0.20962 | 0.000456 | -4e-06 | 101.2668 | 101.6094 | 123.921 | 21.1309 | 0.001947 | 0.001805 | -152.327 | 0.2396 | 8.4021 | 0.003529 | 0.004671 | 0.00456 | 5.3e-05 | 0.000994 | 0.335 | 1.3504 | 2.5479 | 99.48 | 20.93 | 37.179 | 68.5504 | 106.6692 | 65.1347 | 0.009265 | 1.0436 | 0.024876 | 23.8355 | -0.03021 | 0.14253 | -5.1e-05 | -0.003604 | -0.003604 | -1.6254
2019-03-21 12:11 | 1553184663.3804 | 1550 | 0.20963 | 0.000453 | -4e-06 | 101.2868 | 101.2838 | 123.8802 | 21.131 | 0.001952 | 0.001847 | -152.3233 | 0.23958 | 8.4026 | 0.003518 | 0.005559 | 0.005334 | 5.1e-05 | 0.000995 | 0.27759 | 1.3289 | 2.548 | 99.48 | 20.93 | 37.179 | 68.6659 | 106.3908 | 66.1516 | 0.009265 | 1.0405 | 0.024834 | 23.8674 | 0.000261 | 0.14252 | -0.000104 | -0.007479 | -0.007479 | -0.78324
2019-03-21 12:11 | 1553184664.3804 | 1551 | 0.20963 | 0.000452 | -4e-06 | 101.3002 | 101.4432 | 124.4358 | 21.1317 | 0.001939 | 0.00183 | -152.3103 | 0.2386 | 8.4016 | 0.003494 | 0.005074 | 0.00492 | 5.2e-05 | 0.000996 | 0.32709 | 1.3461 | 2.548 | 99.48 | 20.93 | 37.178 | 68.3515 | 105.8217 | 65.7434 | 0.009265 | 1.0418 | 0.024905 | 23.9065 | -0.01262 | 0.14194 | -9.8e-05 | -0.007124 | -0.007124 | -0.81891
2019-03-21 12:11 | 1553184665.3808 | 1552 | 0.20962 | 0.000454 | -4e-06 | 101.2436 | 101.5106 | 124.9087 | 21.1313 | 0.001931 | 0.00183 | -152.2977 | 0.23961 | 8.4036 | 0.00349 | 0.004774 | 0.004595 | 5e-05 | 0.000992 | 0.31509 | 1.3001 | 2.5479 | 99.48 | 20.93 | 37.178 | 68.1679 | 105.7172 | 65.7242 | 0.009265 | 1.0418 | 0.024953 | 23.9512 | -0.023564 | 0.14254 | -4.6e-05 | -0.003196 | -0.003196 | -1.8329
2019-03-21 12:11 | 1553184666.3804 | 1553 | 0.20963 | 0.000451 | -3e-06 | 101.2531 | 101.2525 | 124.9552 | 21.1308 | 0.001944 | 0.001867 | -152.2851 | 0.23937 | 8.4024 | 0.003512 | 0.004722 | 0.00456 | 5.5e-05 | 0.000994 | 0.30663 | 1.2919 | 2.5479 | 99.48 | 20.93 | 37.177 | 68.4707 | 106.2605 | 66.6165 | 0.009265 | 1.0391 | 0.024925 | 23.9872 | 5.5e-05 | 0.1424 | -0.000137 | -0.010029 | -0.010029 | -0.58357
2019-03-21 12:11 | 1553184667.3799 | 1554 | 0.20963 | 0.000451 | -4e-06 | 101.3064 | 101.4115 | 125.1683 | 21.131 | 0.001913 | 0.001813 | -152.2801 | 0.24118 | 8.4028 | 0.003563 | 0.005278 | 0.005049 | 5.4e-05 | 0.000995 | 0.29878 | 1.3426 | 2.548 | 99.48 | 20.93 | 37.174 | 67.7189 | 107.4954 | 65.3149 | 0.009264 | 1.0431 | 0.024994 | 23.9616 | -0.009275 | 0.14348 | -0.000133 | -0.009778 | -0.009778 | -0.60309
2019-03-21 12:11 | 1553184668.381 | 1555 | 0.2096 | 0.000455 | -4e-06 | 101.2951 | 101.2838 | 125.3899 | 21.1314 | 0.001928 | 0.001799 | -152.3007 | 0.23984 | 8.4014 | 0.003476 | 0.00476 | 0.004583 | 5.3e-05 | 0.000995 | 0.33754 | 1.2932 | 2.5479 | 99.48 | 20.93 | 37.172 | 68.0891 | 105.3815 | 64.9846 | 0.009264 | 1.0441 | 0.025028 | 23.9711 | 0.000998 | 0.14268 | 8.1e-05 | 0.006357 | 0.006357 | 0.92252
2019-03-21 12:11 | 1553184669.3807 | 1556 | 0.20959 | 0.000454 | -4e-06 | 101.3011 | 101.6571 | 125.3174 | 21.1303 | 0.001972 | 0.001861 | -152.3023 | 0.24006 | 8.4052 | 0.003541 | 0.004662 | 0.004532 | 5.3e-05 | 0.000997 | 0.28528 | 1.2919 | 2.548 | 99.48 | 20.92 | 37.169 | 69.1489 | 106.9496 | 66.47 | 0.009257 | 1.0395 | 0.024966 | 24.0167 | -0.031381 | 0.14281 | 0.000147 | 0.011298 | 0.011298 | 0.51952
2019-03-21 12:11 | 1553184670.3807 | 1557 | 0.2096 | 0.000454 | -4e-06 | 101.3171 | 101.3911 | 125.2888 | 21.1314 | 0.001954 | 0.001803 | -152.3004 | 0.23889 | 8.4033 | 0.00337 | 0.00526 | 0.005087 | 5.2e-05 | 0.000994 | 0.31614 | 1.2681 | 2.5479 | 99.48 | 20.92 | 37.168 | 68.7139 | 102.8217 | 65.0774 | 0.009257 | 1.0438 | 0.025015 | 23.9647 | -0.006525 | 0.14212 | 7.6e-05 | 0.00596 | 0.00596 | 0.98006
2019-03-21 12:11 | 1553184671.3801 | 1558 | 0.20961 | 0.000453 | -4e-06 | 101.1563 | 101.3669 | 125.6279 | 21.1301 | 0.001966 | 0.001794 | -152.3195 | 0.24086 | 8.402 | 0.003447 | 0.00502 | 0.004839 | 4.9e-05 | 0.000993 | 0.31061 | 1.3178 | 2.5479 | 99.48 | 20.92 | 37.166 | 69.0122 | 104.6716 | 64.8606 | 0.009256 | 1.0445 | 0.025056 | 23.9894 | -0.018607 | 0.14329 | 1e-05 | 0.001007 | 0.001007 | 5.8482
2019-03-21 12:11 | 1553184672.3808 | 1559 | 0.20959 | 0.000453 | -4e-06 | 101.2477 | 101.3809 | 125.6786 | 21.131 | 0.001908 | 0.001827 | -152.3101 | 0.24043 | 8.4015 | 0.003445 | 0.005001 | 0.004909 | 5.4e-05 | 0.000997 | 0.29662 | 1.321 | 2.5478 | 99.48 | 20.92 | 37.164 | 67.6074 | 104.633 | 65.6623 | 0.009256 | 1.042 | 0.025032 | 24.0227 | -0.011755 | 0.14303 | 0.000104 | 0.008114 | 0.008114 | 0.72454
2019-03-21 12:11 | 1553184673.3798 | 1560 | 0.20958 | 0.000452 | -4e-06 | 101.2587 | 101.5474 | 125.4228 | 21.1312 | 0.001931 | 0.001818 | -152.3039 | 0.2398 | 8.4039 | 0.003454 | 0.004636 | 0.00449 | 5.2e-05 | 0.000994 | 0.34379 | 1.3225 | 2.5479 | 99.48 | 20.92 | 37.163 | 68.1538 | 104.836 | 65.4431 | 0.009256 | 1.0427 | 0.025014 | 23.9905 | -0.025463 | 0.14266 | 0.000169 | 0.012954 | 0.012954 | 0.45262
2019-03-21 12:11 | 1553184674.3813 | 1561 | 0.2096 | 0.000456 | -4e-06 | 101.2602 | 101.4176 | 125.2007 | 21.1303 | 0.001868 | 0.001809 | -152.3077 | 0.23954 | 8.4021 | 0.003442 | 0.004531 | 0.00433 | 5.1e-05 | 0.000995 | 0.3095 | 1.308 | 2.548 | 99.48 | 20.92 | 37.162 | 66.6456 | 104.5608 | 65.2183 | 0.009255 | 1.0434 | 0.025001 | 23.9613 | -0.013894 | 0.1425 | 6.2e-05 | 0.004873 | 0.004873 | 1.2019
2019-03-21 12:11 | 1553184675.38 | 1562 | 0.2096 | 0.000454 | -3e-06 | 101.1548 | 101.3526 | 124.7794 | 21.1313 | 0.001931 | 0.001837 | -152.3187 | 0.2401 | 8.4028 | 0.00348 | 0.005165 | 0.005031 | 5.4e-05 | 0.000994 | 0.3131 | 1.3389 | 2.548 | 99.48 | 20.92 | 37.158 | 68.1523 | 105.4813 | 65.8998 | 0.009254 | 1.0413 | 0.024933 | 23.945 | -0.017476 | 0.14284 | 7.4e-05 | 0.005738 | 0.005738 | 1.0231
2019-03-21 12:11 | 1553184676.38 | 1563 | 0.20959 | 0.000456 | -4e-06 | 101.2745 | 101.3185 | 124.7078 | 21.1314 | 0.001923 | 0.001806 | -152.3227 | 0.23975 | 8.402 | 0.003406 | 0.004619 | 0.004512 | 5.6e-05 | 0.000996 | 0.33229 | 1.3221 | 2.5479 | 99.48 | 20.93 | 37.157 | 67.9629 | 103.6822 | 65.1539 | 0.00926 | 1.0436 | 0.024954 | 23.9118 | -0.003885 | 0.14262 | 0.000145 | 0.011081 | 0.011081 | 0.52902
2019-03-21 12:11 | 1553184677.3801 | 1564 | 0.2096 | 0.000451 | -3e-06 | 101.223 | 101.4388 | 124.3837 | 21.1318 | 0.00195 | 0.001795 | -152.319 | 0.24009 | 8.4021 | 0.003428 | 0.004464 | 0.004185 | 5e-05 | 0.000992 | 0.30414 | 1.3417 | 2.5479 | 99.48 | 20.93 | 37.153 | 68.6302 | 104.2196 | 64.9013 | 0.009259 | 1.0444 | 0.024931 | 23.8717 | -0.019046 | 0.14283 | 5.5e-05 | 0.004323 | 0.004323 | 1.3581
2019-03-21 12:11 | 1553184678.3805 | 1565 | 0.20961 | 0.000453 | -4e-06 | 101.2897 | 101.4956 | 124.1675 | 21.1324 | 0.001946 | 0.001843 | -152.3159 | 0.23994 | 8.4012 | 0.003471 | 0.004608 | 0.004426 | 4.9e-05 | 0.000994 | 0.29479 | 1.3244 | 2.5481 | 99.48 | 20.93 | 37.152 | 68.5206 | 105.2651 | 66.05 | 0.009259 | 1.0408 | 0.024867 | 23.8915 | -0.018165 | 0.14274 | 1.6e-05 | 0.00141 | 0.00141 | 4.1596
2019-03-21 12:11 | 1553184679.3802 | 1566 | 0.20959 | 0.000452 | -3e-06 | 101.2546 | 101.3621 | 124.109 | 21.1315 | 0.001942 | 0.001789 | -152.2907 | 0.2397 | 8.4016 | 0.003487 | 0.004677 | 0.004521 | 5e-05 | 0.000991 | 0.29966 | 1.376 | 2.5481 | 99.48 | 20.93 | 37.15 | 68.4318 | 105.6484 | 64.7583 | 0.009258 | 1.0448 | 0.024908 | 23.8403 | -0.009496 | 0.1426 | 0.000111 | 0.008467 | 0.008467 | 0.69217
2019-03-21 12:11 | 1553184680.3808 | 1567 | 0.20961 | 0.000449 | -3e-06 | 101.2933 | 101.5396 | 123.9717 | 21.1318 | 0.00198 | 0.001812 | -152.286 | 0.23933 | 8.4015 | 0.003415 | 0.004751 | 0.00456 | 5.1e-05 | 0.000994 | 0.31022 | 1.3579 | 2.5481 | 99.48 | 20.93 | 37.146 | 69.3484 | 103.905 | 65.2931 | 0.009257 | 1.0431 | 0.024875 | 23.846 | -0.02172 | 0.14238 | 3.5e-05 | 0.002864 | 0.002864 | 2.0431
2019-03-21 12:11 | 1553184681.3806 | 1568 | 0.20958 | 0.000456 | -4e-06 | 101.2924 | 101.4749 | 123.4577 | 21.1333 | 0.001953 | 0.001833 | -152.2626 | 0.24017 | 8.406 | 0.003452 | 0.004328 | 0.004216 | 5.4e-05 | 0.000988 | 0.30879 | 1.3484 | 2.548 | 99.48 | 20.93 | 37.144 | 68.7019 | 104.7943 | 65.8152 | 0.009257 | 1.0415 | 0.024804 | 23.8149 | -0.016097 | 0.14288 | 0.000185 | 0.014023 | 0.014023 | 0.41876
2019-03-21 12:11 | 1553184682.3803 | 1569 | 0.20962 | 0.000456 | -3e-06 | 101.2444 | 101.5089 | 123.2758 | 21.1318 | 0.001929 | 0.001765 | -152.2626 | 0.23922 | 8.4012 | 0.003429 | 0.005007 | 0.004822 | 5.1e-05 | 0.000984 | 0.3095 | 1.3328 | 2.5481 | 99.48 | 20.93 | 37.142 | 68.1068 | 104.2312 | 64.1729 | 0.009256 | 1.0466 | 0.024846 | 23.7396 | -0.023335 | 0.14231 | -3.5e-05 | -0.002381 | -0.002381 | -2.4567
2019-03-21 12:11 | 1553184683.3794 | 1570 | 0.20956 | 0.000451 | -4e-06 | 101.2218 | 101.3008 | 123.1661 | 21.1319 | 0.001901 | 0.001835 | -152.2863 | 0.23965 | 8.4018 | 0.003444 | 0.005201 | 0.005017 | 5.3e-05 | 0.000977 | 0.32803 | 1.3421 | 2.548 | 99.48 | 20.93 | 37.14 | 67.45 | 104.611 | 65.8652 | 0.009256 | 1.0414 | 0.024773 | 23.7885 | -0.006981 | 0.14257 | 0.000297 | 0.022358 | 0.022358 | 0.26208
2019-03-21 12:11 | 1553184684.3798 | 1571 | 0.20959 | 0.000455 | -4e-06 | 101.2796 | 101.5011 | 123.0119 | 21.1318 | 0.001959 | 0.001818 | -152.2747 | 0.24053 | 8.4013 | 0.003484 | 0.005355 | 0.005198 | 5.3e-05 | 0.000974 | 0.31686 | 1.3297 | 2.5479 | 99.48 | 20.93 | 37.137 | 68.8258 | 105.588 | 65.4444 | 0.009255 | 1.0427 | 0.024773 | 23.7588 | -0.01954 | 0.14309 | 0.00011 | 0.008452 | 0.008452 | 0.6958
2019-03-21 12:11 | 1553184685.3799 | 1572 | 0.20958 | 0.000455 | -4e-06 | 101.251 | 101.4858 | 122.8429 | 21.1321 | 0.001928 | 0.00179 | -152.2673 | 0.23965 | 8.4031 | 0.003449 | 0.005199 | 0.004903 | 5.2e-05 | 0.000977 | 0.29247 | 1.2923 | 2.5481 | 99.48 | 20.93 | 37.132 | 68.0821 | 104.7167 | 64.7651 | 0.009254 | 1.0448 | 0.024781 | 23.7187 | -0.020715 | 0.14257 | 0.000158 | 0.011964 | 0.011964 | 0.48979
2019-03-21 12:11 | 1553184686.3805 | 1573 | 0.20955 | 0.000451 | -4e-06 | 101.217 | 101.4173 | 123.0816 | 21.1327 | 0.001958 | 0.001803 | -152.2639 | 0.23898 | 8.4007 | 0.003434 | 0.004355 | 0.004198 | 5.3e-05 | 0.000974 | 0.27959 | 1.3169 | 2.5479 | 99.48 | 20.93 | 37.131 | 68.8198 | 104.3533 | 65.0833 | 0.009253 | 1.0438 | 0.024793 | 23.7529 | -0.017681 | 0.14217 | 0.00036 | 0.02701 | 0.02701 | 0.21633
2019-03-21 12:11 | 1553184687.3805 | 1574 | 0.20959 | 0.000454 | -3e-06 | 101.2659 | 101.4149 | 123.3962 | 21.1315 | 0.001944 | 0.001801 | -152.2834 | 0.23984 | 8.4012 | 0.003395 | 0.004951 | 0.004783 | 5.1e-05 | 0.000975 | 0.31332 | 1.323 | 2.5481 | 99.48 | 20.93 | 37.13 | 68.4659 | 103.4127 | 65.0462 | 0.009253 | 1.0439 | 0.024826 | 23.7819 | -0.013154 | 0.14268 | 0.00011 | 0.008436 | 0.008436 | 0.69516
2019-03-21 12:11 | 1553184688.3811 | 1575 | 0.20959 | 0.000452 | -3e-06 | 101.2692 | 101.2256 | 123.5527 | 21.1312 | 0.001978 | 0.001783 | -152.2733 | 0.23953 | 8.401 | 0.003443 | 0.005052 | 0.0049 | 5.4e-05 | 0.000977 | 0.30414 | 1.3503 | 2.5478 | 99.48 | 20.93 | 37.129 | 69.2979 | 104.5806 | 64.6134 | 0.009253 | 1.0452 | 0.024858 | 23.7817 | 0.003853 | 0.1425 | 0.00015 | 0.011439 | 0.011439 | 0.51197
2019-03-21 12:11 | 1553184689.3807 | 1576 | 0.20956 | 0.000456 | -4e-06 | 101.2373 | 101.4139 | 123.3307 | 21.1321 | 0.00188 | 0.001836 | -152.2872 | 0.23916 | 8.401 | 0.003446 | 0.004604 | 0.00443 | 5.5e-05 | 0.000979 | 0.30757 | 1.3561 | 2.5481 | 99.48 | 20.93 | 37.129 | 66.929 | 104.6491 | 65.8849 | 0.009253 | 1.0413 | 0.024789 | 23.8051 | -0.01559 | 0.14228 | 0.000292 | 0.021963 | 0.021963 | 0.26624
2019-03-21 12:11 | 1553184690.3807 | 1577 | 0.20955 | 0.000454 | -3e-06 | 101.2072 | 101.3781 | 123.3107 | 21.1317 | 0.001957 | 0.001779 | -152.2657 | 0.23933 | 8.4024 | 0.003424 | 0.003817 | 0.003689 | 5.3e-05 | 0.000976 | 0.2756 | 1.3561 | 2.548 | 99.48 | 20.93 | 37.129 | 68.7976 | 104.1249 | 64.5053 | 0.009253 | 1.0456 | 0.024837 | 23.7546 | -0.015095 | 0.14238 | 0.000363 | 0.027302 | 0.027302 | 0.21434
2019-03-21 12:11 | 1553184691.3798 | 1578 | 0.20958 | 0.000451 | -3e-06 | 101.2164 | 101.3226 | 123.5659 | 21.1321 | 0.00187 | 0.001801 | -152.2887 | 0.23984 | 8.4064 | 0.003455 | 0.005166 | 0.004972 | 5.7e-05 | 0.000979 | 0.31492 | 1.3478 | 2.5477 | 99.48 | 20.93 | 37.13 | 66.6981 | 104.8702 | 65.0437 | 0.009253 | 1.0439 | 0.024843 | 23.7982 | -0.009379 | 0.14268 | 0.000221 | 0.016727 | 0.016727 | 0.35058
2019-03-21 12:12 | 1553184692.3803 | 1579 | 0.20958 | 0.000456 | -4e-06 | 101.2173 | 101.4742 | 123.5093 | 21.1302 | 0.00188 | 0.001799 | -152.2917 | 0.23911 | 8.4019 | 0.003391 | 0.004316 | 0.004232 | 5.3e-05 | 0.000981 | 0.30254 | 1.3188 | 2.5477 | 99.48 | 20.93 | 37.13 | 66.9358 | 103.3187 | 64.9887 | 0.009253 | 1.0441 | 0.02484 | 23.7908 | -0.022669 | 0.14225 | 0.00016 | 0.012176 | 0.012176 | 0.48014
2019-03-21 12:12 | 1553184693.38 | 1580 | 0.20957 | 0.000452 | -3e-06 | 101.2248 | 101.448 | 123.982 | 21.1314 | 0.001954 | 0.001827 | -152.2739 | 0.23989 | 8.4025 | 0.003411 | 0.00426 | 0.004045 | 5.3e-05 | 0.000981 | 0.31061 | 1.3309 | 2.5479 | 99.48 | 20.93 | 37.13 | 68.7115 | 103.8063 | 65.6611 | 0.009253 | 1.042 | 0.024862 | 23.8599 | -0.019699 | 0.14271 | 0.000256 | 0.019303 | 0.019303 | 0.30386
2019-03-21 12:12 | 1553184694.3799 | 1581 | 0.20959 | 0.000456 | -4e-06 | 101.259 | 101.5127 | 124.0222 | 21.1306 | 0.001929 | 0.001816 | -152.2822 | 0.23944 | 8.402 | 0.003342 | 0.004591 | 0.004462 | 5.4e-05 | 0.000979 | 0.31885 | 1.2822 | 2.5479 | 99.48 | 20.93 | 37.13 | 68.1071 | 102.128 | 65.4067 | 0.009253 | 1.0428 | 0.024876 | 23.8548 | -0.022376 | 0.14244 | 0.000129 | 0.009843 | 0.009843 | 0.59479
2019-03-21 12:12 | 1553184695.3797 | 1582 | 0.20959 | 0.000455 | -3e-06 | 101.2108 | 101.4759 | 124.3423 | 21.1316 | 0.001884 | 0.0018 | -152.2602 | 0.23934 | 8.4008 | 0.003365 | 0.005028 | 0.004897 | 5.7e-05 | 0.000977 | 0.31868 | 1.2861 | 2.548 | 99.48 | 20.93 | 37.132 | 67.0261 | 102.6937 | 65.0073 | 0.009254 | 1.044 | 0.024923 | 23.8715 | -0.023399 | 0.14238 | 0.000107 | 0.008227 | 0.008227 | 0.71135
2019-03-21 12:12 | 1553184696.3805 | 1583 | 0.20964 | 0.000456 | -4e-06 | 101.2608 | 101.3621 | 124.209 | 21.13 | 0.00192 | 0.001863 | -152.2703 | 0.23907 | 8.4027 | 0.003339 | 0.004583 | 0.004324 | 5.8e-05 | 0.000974 | 0.33356 | 1.3026 | 2.548 | 99.48 | 20.93 | 37.133 | 67.8959 | 102.0679 | 66.5222 | 0.009254 | 1.0394 | 0.024854 | 23.9121 | -0.008944 | 0.14222 | -0.00017 | -0.012444 | -0.012444 | -0.46974
2019-03-21 12:12 | 1553184697.38 | 1584 | 0.20961 | 0.000453 | -4e-06 | 101.2626 | 101.3969 | 124.1933 | 21.1317 | 0.001885 | 0.001798 | -152.283 | 0.23896 | 8.4035 | 0.003392 | 0.004579 | 0.004417 | 5.6e-05 | 0.000978 | 0.3183 | 1.2988 | 2.5479 | 99.48 | 20.93 | 37.134 | 67.063 | 103.3531 | 64.9558 | 0.009254 | 1.0442 | 0.024909 | 23.8554 | -0.011851 | 0.14216 | 1.9e-05 | 0.001635 | 0.001635 | 3.5734
2019-03-21 12:12 | 1553184698.3791 | 1585 | 0.20963 | 0.000453 | -4e-06 | 101.2158 | 101.5106 | 124.2231 | 21.1306 | 0.001854 | 0.001817 | -152.2718 | 0.23971 | 8.4021 | 0.003368 | 0.004301 | 0.004184 | 5.4e-05 | 0.000977 | 0.30613 | 1.3363 | 2.5479 | 99.48 | 20.93 | 37.134 | 66.316 | 102.7652 | 65.42 | 0.009254 | 1.0428 | 0.024895 | 23.8746 | -0.026013 | 0.1426 | -0.000108 | -0.007837 | -0.007837 | -0.74784
2019-03-21 12:12 | 1553184699.3791 | 1586 | 0.20959 | 0.000456 | -4e-06 | 101.2132 | 101.3345 | 124.5189 | 21.1311 | 0.001916 | 0.001797 | -152.2885 | 0.23937 | 8.4046 | 0.003375 | 0.005232 | 0.005043 | 5.5e-05 | 0.000978 | 0.28152 | 1.3174 | 2.548 | 99.48 | 20.93 | 37.135 | 67.8049 | 102.9274 | 64.946 | 0.009254 | 1.0442 | 0.024942 | 23.8863 | -0.010721 | 0.1424 | 0.000107 | 0.008274 | 0.008274 | 0.70733
2019-03-21 12:12 | 1553184700.3793 | 1587 | 0.20965 | 0.000451 | -3e-06 | 101.2346 | 101.4122 | 124.524 | 21.1307 | 0.001907 | 0.001857 | -152.2996 | 0.23876 | 8.4029 | 0.003424 | 0.004282 | 0.004135 | 5.6e-05 | 0.000977 | 0.27129 | 1.3434 | 2.548 | 99.48 | 20.93 | 37.136 | 67.5955 | 104.1177 | 66.3831 | 0.009255 | 1.0398 | 0.02489 | 23.9375 | -0.015677 | 0.14204 | -0.000215 | -0.015847 | -0.015847 | -0.36839
2019-03-21 12:12 | 1553184701.3793 | 1588 | 0.20959 | 0.00045 | -4e-06 | 101.1998 | 101.4837 | 124.7409 | 21.1309 | 0.001926 | 0.001774 | -152.2685 | 0.23853 | 8.4018 | 0.003326 | 0.004726 | 0.004501 | 5.4e-05 | 0.000981 | 0.29479 | 1.3795 | 2.5481 | 99.48 | 20.93 | 37.136 | 68.0514 | 101.7475 | 64.3929 | 0.009255 | 1.0459 | 0.024985 | 23.888 | -0.025064 | 0.14191 | 0.000158 | 0.012122 | 0.012122 | 0.48111
2019-03-21 12:12 | 1553184702.3804 | 1589 | 0.20963 | 0.000455 | -4e-06 | 101.2662 | 101.2218 | 125.0252 | 21.1314 | 0.001966 | 0.001835 | -152.2634 | 0.2402 | 8.4017 | 0.003385 | 0.004773 | 0.004607 | 5.1e-05 | 0.00098 | 0.34772 | 1.3339 | 2.5479 | 99.48 | 20.93 | 37.137 | 69.0055 | 103.1689 | 65.8571 | 0.009255 | 1.0414 | 0.02496 | 23.967 | 0.003921 | 0.14289 | -9.6e-05 | -0.006972 | -0.006972 | -0.84242
2019-03-21 12:12 | 1553184703.3804 | 1590 | 0.20961 | 0.00045 | -4e-06 | 101.1309 | 101.5682 | 124.7836 | 21.1323 | 0.001891 | 0.001769 | -152.2601 | 0.24074 | 8.4053 | 0.003389 | 0.004897 | 0.004768 | 5.5e-05 | 0.000979 | 0.29286 | 1.3233 | 2.548 | 99.48 | 20.93 | 37.141 | 67.2141 | 103.268 | 64.2593 | 0.009256 | 1.0463 | 0.024994 | 23.8874 | -0.03859 | 0.14322 | 2e-06 | 0.000389 | 0.000389 | 15.132
2019-03-21 12:12 | 1553184704.3795 | 1591 | 0.20961 | 0.000451 | -4e-06 | 101.2066 | 101.514 | 124.5165 | 21.1317 | 0.00193 | 0.001832 | -152.2625 | 0.23958 | 8.405 | 0.003334 | 0.004885 | 0.004687 | 5.4e-05 | 0.000982 | 0.30663 | 1.3803 | 2.5479 | 99.48 | 20.93 | 37.142 | 68.1464 | 101.9492 | 65.7747 | 0.009256 | 1.0417 | 0.024912 | 23.9153 | -0.027129 | 0.14252 | 2.3e-05 | 0.001985 | 0.001985 | 2.9517
2019-03-21 12:12 | 1553184705.3799 | 1592 | 0.20957 | 0.000456 | -4e-06 | 101.17 | 101.1745 | 124.2463 | 21.1316 | 0.001964 | 0.001833 | -152.291 | 0.23989 | 8.4018 | 0.003386 | 0.004419 | 0.004274 | 5.3e-05 | 0.000979 | 0.30182 | 1.3475 | 2.5479 | 99.48 | 20.93 | 37.144 | 68.9646 | 103.207 | 65.8089 | 0.009257 | 1.0416 | 0.024883 | 23.8906 | -0.000399 | 0.14271 | 0.000243 | 0.018405 | 0.018405 | 0.31869
2019-03-21 12:12 | 1553184706.3796 | 1593 | 0.2096 | 0.000453 | -3e-06 | 101.1548 | 101.5171 | 124.0543 | 21.1317 | 0.001934 | 0.001807 | -152.2868 | 0.23972 | 8.4043 | 0.003365 | 0.004599 | 0.00435 | 5.5e-05 | 0.000984 | 0.30486 | 1.3019 | 2.548 | 99.48 | 20.94 | 37.145 | 68.2246 | 102.6869 | 65.1714 | 0.009263 | 1.0435 | 0.024888 | 23.8496 | -0.031982 | 0.14261 | 7.8e-05 | 0.006074 | 0.006074 | 0.96501
2019-03-21 12:12 | 1553184707.3797 | 1594 | 0.20959 | 0.000456 | -4e-06 | 101.2346 | 101.3219 | 123.8513 | 21.1322 | 0.00195 | 0.001792 | -152.2773 | 0.23952 | 8.4021 | 0.003397 | 0.003576 | 0.003475 | 5.3e-05 | 0.000986 | 0.28407 | 1.3307 | 2.5479 | 99.48 | 20.94 | 37.147 | 68.6094 | 103.473 | 64.8111 | 0.009263 | 1.0446 | 0.02488 | 23.8174 | -0.007713 | 0.14249 | 0.000115 | 0.008804 | 0.008804 | 0.66518
2019-03-21 12:12 | 1553184708.3807 | 1595 | 0.2096 | 0.000454 | -3e-06 | 101.2114 | 101.401 | 123.4875 | 21.1317 | 0.001909 | 0.001802 | -152.2782 | 0.23968 | 8.4037 | 0.003285 | 0.00488 | 0.004633 | 5.5e-05 | 0.000986 | 0.30934 | 1.3245 | 2.5479 | 99.48 | 20.94 | 37.15 | 67.6338 | 100.7458 | 65.0659 | 0.009264 | 1.0438 | 0.024835 | 23.7914 | -0.016739 | 0.14259 | 5.5e-05 | 0.004311 | 0.004311 | 1.3592
2019-03-21 12:12 | 1553184709.3808 | 1596 | 0.20957 | 0.000459 | -4e-06 | 101.2051 | 101.3666 | 123.2849 | 21.1324 | 0.001873 | 0.001811 | -152.2739 | 0.23962 | 8.4028 | 0.003356 | 0.004274 | 0.004056 | 5.6e-05 | 0.00099 | 0.34827 | 1.3379 | 2.548 | 99.48 | 20.94 | 37.153 | 66.7671 | 102.4643 | 65.2882 | 0.009265 | 1.0432 | 0.024806 | 23.7797 | -0.014257 | 0.14255 | 0.000228 | 0.017226 | 0.017226 | 0.34012
2019-03-21 12:12 | 1553184710.3804 | 1597 | 0.20963 | 0.000455 | -4e-06 | 101.2772 | 101.3737 | 123.4725 | 21.1319 | 0.001955 | 0.001853 | -152.2796 | 0.23807 | 8.402 | 0.003286 | 0.005109 | 0.004966 | 5.5e-05 | 0.000989 | 0.30309 | 1.3045 | 2.5479 | 99.48 | 20.94 | 37.157 | 68.7346 | 100.7874 | 66.2922 | 0.009266 | 1.0401 | 0.024788 | 23.833 | -0.008518 | 0.14163 | -9.1e-05 | -0.006553 | -0.006553 | -0.88826
2019-03-21 12:12 | 1553184711.3809 | 1598 | 0.20963 | 0.00045 | -4e-06 | 101.3061 | 101.5164 | 123.2276 | 21.1322 | 0.001901 | 0.001797 | -152.2844 | 0.23955 | 8.4052 | 0.003333 | 0.005098 | 0.004903 | 5.4e-05 | 0.000993 | 0.30071 | 1.2738 | 2.5479 | 99.48 | 20.94 | 37.16 | 67.448 | 101.9073 | 64.9478 | 0.009266 | 1.0442 | 0.024813 | 23.7622 | -0.018549 | 0.14251 | -9.3e-05 | -0.006678 | -0.006678 | -0.87713
2019-03-21 12:12 | 1553184712.379 | 1599 | 0.20959 | 0.000456 | -4e-06 | 101.2507 | 101.5617 | 122.9669 | 21.1319 | 0.001899 | 0.001842 | -152.2911 | 0.23936 | 8.4029 | 0.003285 | 0.004548 | 0.004332 | 5.5e-05 | 0.000998 | 0.3454 | 1.2708 | 2.5479 | 99.48 | 20.94 | 37.161 | 67.4018 | 100.757 | 66.02 | 0.009267 | 1.0409 | 0.024747 | 23.7746 | -0.027433 | 0.14239 | 0.000113 | 0.008613 | 0.008613 | 0.67951
2019-03-21 12:12 | 1553184713.38 | 1600 | 0.20958 | 0.00045 | -4e-06 | 101.276 | 101.4139 | 122.8324 | 21.1329 | 0.001916 | 0.001803 | -152.297 | 0.24008 | 8.4006 | 0.003327 | 0.004508 | 0.004422 | 5.4e-05 | 0.000998 | 0.3282 | 1.295 | 2.5479 | 99.48 | 20.94 | 37.163 | 67.8083 | 101.7737 | 65.0873 | 0.009267 | 1.0438 | 0.024768 | 23.7289 | -0.012169 | 0.14282 | 0.000187 | 0.014156 | 0.014156 | 0.41465
2019-03-21 12:12 | 1553184714.3799 | 1601 | 0.20961 | 0.000452 | -4e-06 | 101.2829 | 101.5682 | 123.3499 | 21.1322 | 0.001893 | 0.001793 | -152.2944 | 0.23958 | 8.4039 | 0.003298 | 0.00485 | 0.004609 | 5e-05 | 0.00099 | 0.31404 | 1.3229 | 2.5479 | 99.48 | 20.93 | 37.163 | 67.2429 | 101.0734 | 64.8564 | 0.009261 | 1.0445 | 0.024828 | 23.7708 | -0.025161 | 0.14253 | 0 | 0.000292 | 0.000292 | 20.0381
2019-03-21 12:12 | 1553184715.3791 | 1602 | 0.2096 | 0.000454 | -4e-06 | 101.316 | 101.4977 | 123.4441 | 21.1331 | 0.001896 | 0.001878 | -152.2777 | 0.23866 | 8.4038 | 0.003355 | 0.004619 | 0.004487 | 5.3e-05 | 0.00099 | 0.31061 | 1.2956 | 2.5479 | 99.48 | 20.93 | 37.165 | 67.3123 | 102.4549 | 66.8896 | 0.009262 | 1.0382 | 0.024764 | 23.8512 | -0.01603 | 0.14198 | 9.5e-05 | 0.007334 | 0.007334 | 0.79565
2019-03-21 12:12 | 1553184716.3792 | 1603 | 0.20958 | 0.000453 | -4e-06 | 101.2227 | 101.5426 | 123.4626 | 21.1311 | 0.001931 | 0.001838 | -152.2668 | 0.23981 | 8.4014 | 0.003374 | 0.004957 | 0.004736 | 5.3e-05 | 0.000993 | 0.34108 | 1.2678 | 2.548 | 99.48 | 20.93 | 37.168 | 68.1677 | 102.9205 | 65.9208 | 0.009263 | 1.0412 | 0.024801 | 23.819 | -0.028228 | 0.14266 | 0.00017 | 0.012861 | 0.012861 | 0.45592
2019-03-21 12:12 | 1553184717.3802 | 1604 | 0.20958 | 0.000453 | -4e-06 | 101.2871 | 101.6663 | 123.4286 | 21.1328 | 0.001935 | 0.001833 | -152.2482 | 0.23981 | 8.4022 | 0.003392 | 0.004198 | 0.003986 | 5.3e-05 | 0.000997 | 0.32139 | 1.2911 | 2.5478 | 99.48 | 20.93 | 37.168 | 68.2638 | 103.3384 | 65.8021 | 0.009263 | 1.0416 | 0.024802 | 23.8116 | -0.033427 | 0.14266 | 0.000226 | 0.017135 | 0.017135 | 0.3422
2019-03-21 12:12 | 1553184718.3795 | 1605 | 0.20961 | 0.000458 | -4e-06 | 101.3049 | 101.3788 | 123.5898 | 21.132 | 0.001878 | 0.001812 | -152.261 | 0.23944 | 8.4036 | 0.00341 | 0.004929 | 0.004813 | 5.6e-05 | 0.000999 | 0.32189 | 1.2857 | 2.548 | 99.48 | 20.93 | 37.168 | 66.8861 | 103.7936 | 65.3121 | 0.009263 | 1.0431 | 0.024836 | 23.8099 | -0.006522 | 0.14244 | -1.3e-05 | -0.000709 | -0.000709 | -8.2515
2019-03-21 12:12 | 1553184719.3793 | 1606 | 0.20961 | 0.000456 | -4e-06 | 101.3183 | 101.5195 | 123.6807 | 21.1336 | 0.001981 | 0.001814 | -152.2649 | 0.23993 | 8.4033 | 0.003385 | 0.004737 | 0.004545 | 5.4e-05 | 0.001001 | 0.28766 | 1.2873 | 2.5479 | 99.48 | 20.93 | 37.169 | 69.3679 | 103.1861 | 65.3401 | 0.009263 | 1.043 | 0.024844 | 23.8196 | -0.01774 | 0.14274 | -1.2e-05 | -0.000636 | -0.000636 | -9.224
2019-03-21 12:12 | 1553184720.3804 | 1607 | 0.20958 | 0.000455 | -4e-06 | 101.2504 | 101.4687 | 123.602 | 21.1326 | 0.00195 | 0.001795 | -152.2635 | 0.23915 | 8.4024 | 0.003352 | 0.0046 | 0.004386 | 5.4e-05 | 0.001002 | 0.31404 | 1.3016 | 2.5481 | 99.48 | 20.93 | 37.169 | 68.6286 | 102.39 | 64.8975 | 0.009263 | 1.0444 | 0.024852 | 23.7965 | -0.019267 | 0.14227 | 0.000179 | 0.013601 | 0.013601 | 0.42991
2019-03-21 12:12 | 1553184721.3808 | 1608 | 0.2096 | 0.000449 | -4e-06 | 101.3303 | 101.5764 | 123.6321 | 21.1325 | 0.001905 | 0.001829 | -152.2788 | 0.2395 | 8.4054 | 0.003304 | 0.005142 | 0.004953 | 5.6e-05 | 0.001002 | 0.33787 | 1.343 | 2.548 | 99.48 | 20.93 | 37.166 | 67.5271 | 101.2223 | 65.7059 | 0.009262 | 1.0419 | 0.024826 | 23.8278 | -0.021698 | 0.14248 | 7e-05 | 0.005424 | 0.005424 | 1.0797
2019-03-21 12:12 | 1553184722.3804 | 1609 | 0.20957 | 0.000453 | -4e-06 | 101.3892 | 101.466 | 123.791 | 21.1334 | 0.001913 | 0.001906 | -152.2743 | 0.24022 | 8.4014 | 0.003288 | 0.00424 | 0.004113 | 5.4e-05 | 0.001003 | 0.30735 | 1.3526 | 2.5478 | 99.48 | 20.93 | 37.166 | 67.7401 | 100.8204 | 67.5646 | 0.009262 | 1.0362 | 0.024774 | 23.9084 | -0.00677 | 0.14291 | 0.00029 | 0.021782 | 0.021782 | 0.26964
2019-03-21 12:12 | 1553184723.3796 | 1610 | 0.20959 | 0.000453 | -4e-06 | 101.3243 | 101.5587 | 123.7924 | 21.1341 | 0.001947 | 0.001829 | -152.2876 | 0.23926 | 8.4038 | 0.00337 | 0.004124 | 0.003956 | 5.5e-05 | 0.001003 | 0.30901 | 1.3004 | 2.5478 | 99.48 | 20.93 | 37.165 | 68.5559 | 102.8212 | 65.7098 | 0.009262 | 1.0419 | 0.024842 | 23.8434 | -0.020664 | 0.14234 | 0.000162 | 0.012332 | 0.012332 | 0.47437
2019-03-21 12:12 | 1553184724.3806 | 1611 | 0.20959 | 0.000455 | -4e-06 | 101.3228 | 101.5501 | 123.6994 | 21.1342 | 0.001915 | 0.001803 | -152.2878 | 0.23957 | 8.402 | 0.003276 | 0.004894 | 0.004714 | 5.7e-05 | 0.001005 | 0.31454 | 1.3156 | 2.5479 | 99.48 | 20.93 | 37.165 | 67.7776 | 100.5361 | 65.0903 | 0.009262 | 1.0438 | 0.024855 | 23.8127 | -0.020046 | 0.14252 | 0.000144 | 0.011053 | 0.011053 | 0.52996
2019-03-21 12:12 | 1553184725.381 | 1612 | 0.20958 | 0.000458 | -4e-06 | 101.3454 | 101.5174 | 123.9619 | 21.1333 | 0.00193 | 0.001848 | -152.2928 | 0.23928 | 8.4024 | 0.003285 | 0.00476 | 0.004628 | 5.4e-05 | 0.001 | 0.33826 | 1.3364 | 2.548 | 99.48 | 20.93 | 37.165 | 68.1323 | 100.759 | 66.1701 | 0.009262 | 1.0404 | 0.024842 | 23.8759 | -0.015167 | 0.14235 | 0.000226 | 0.017082 | 0.017082 | 0.34249
2019-03-21 12:12 | 1553184726.3797 | 1613 | 0.20959 | 0.000453 | -4e-06 | 101.2683 | 101.4357 | 124.2407 | 21.1332 | 0.001906 | 0.001845 | -152.3028 | 0.24005 | 8.4034 | 0.003268 | 0.005099 | 0.00489 | 5.5e-05 | 0.001001 | 0.30685 | 1.3144 | 2.548 | 99.48 | 20.93 | 37.163 | 67.5597 | 100.3525 | 66.089 | 0.009261 | 1.0407 | 0.024873 | 23.8999 | -0.014775 | 0.14281 | 0.00015 | 0.011449 | 0.011449 | 0.51267
2019-03-21 12:12 | 1553184727.3811 | 1614 | 0.20961 | 0.000456 | -4e-06 | 101.3297 | 101.5488 | 124.6444 | 21.1335 | 0.001891 | 0.001802 | -152.3218 | 0.24017 | 8.4025 | 0.003237 | 0.004598 | 0.004394 | 5.4e-05 | 0.001004 | 0.31332 | 1.3318 | 2.548 | 99.48 | 20.93 | 37.163 | 67.1941 | 99.6041 | 65.0684 | 0.009261 | 1.0438 | 0.024951 | 23.9027 | -0.019321 | 0.14288 | 4.9e-05 | 0.003881 | 0.003881 | 1.5129
2019-03-21 12:12 | 1553184728.3788 | 1615 | 0.20957 | 0.000456 | -4e-06 | 101.3219 | 101.5413 | 124.5984 | 21.1339 | 0.00192 | 0.001768 | -152.3158 | 0.23956 | 8.403 | 0.003243 | 0.004991 | 0.004905 | 5.5e-05 | 0.001002 | 0.29662 | 1.2929 | 2.5477 | 99.48 | 20.93 | 37.16 | 67.8935 | 99.7545 | 64.243 | 0.009261 | 1.0464 | 0.024976 | 23.8691 | -0.019345 | 0.14251 | 0.00027 | 0.020541 | 0.020541 | 0.28515
2019-03-21 12:12 | 1553184729.3805 | 1616 | 0.20961 | 0.000452 | -4e-06 | 101.3154 | 101.7289 | 124.8299 | 21.1343 | 0.001937 | 0.001805 | -152.3068 | 0.23979 | 8.4046 | 0.003259 | 0.004433 | 0.004247 | 5.3e-05 | 0.001002 | 0.31006 | 1.3012 | 2.5479 | 99.48 | 20.93 | 37.16 | 68.3167 | 100.133 | 65.1394 | 0.009261 | 1.0436 | 0.024966 | 23.923 | -0.036439 | 0.14265 | 3.9e-05 | 0.00322 | 0.00322 | 1.8209
2019-03-21 12:12 | 1553184730.3793 | 1617 | 0.20957 | 0.000456 | -4e-06 | 101.2477 | 101.6418 | 124.9986 | 21.1349 | 0.001826 | 0.001778 | -152.308 | 0.23991 | 8.4009 | 0.003287 | 0.004621 | 0.004425 | 5.5e-05 | 0.001002 | 0.2912 | 1.2747 | 2.548 | 99.48 | 20.93 | 37.158 | 65.6334 | 100.8092 | 64.4851 | 0.00926 | 1.0456 | 0.025008 | 23.916 | -0.034743 | 0.14272 | 0.000286 | 0.021796 | 0.021796 | 0.26913
2019-03-21 12:12 | 1553184731.3789 | 1618 | 0.20958 | 0.000451 | -4e-06 | 101.2778 | 101.5048 | 125.3505 | 21.1345 | 0.001933 | 0.001796 | -152.3136 | 0.23924 | 8.4039 | 0.003251 | 0.005281 | 0.005158 | 5.2e-05 | 0.001009 | 0.32621 | 1.2838 | 2.5478 | 99.48 | 20.93 | 37.157 | 68.221 | 99.948 | 64.9226 | 0.00926 | 1.0443 | 0.025026 | 23.9651 | -0.020028 | 0.14232 | 0.000241 | 0.018353 | 0.018353 | 0.31872
2019-03-21 12:12 | 1553184732.3806 | 1619 | 0.20961 | 0.000449 | -4e-06 | 101.2918 | 101.5372 | 125.3998 | 21.1336 | 0.001892 | 0.001848 | -152.3088 | 0.23845 | 8.4049 | 0.003251 | 0.004272 | 0.004113 | 5.4e-05 | 0.001005 | 0.29391 | 1.2653 | 2.5479 | 99.48 | 20.93 | 37.156 | 67.2223 | 99.9299 | 66.1603 | 0.00926 | 1.0405 | 0.024986 | 24.0137 | -0.021642 | 0.14185 | 2e-06 | 0.000405 | 0.000405 | 14.3865
2019-03-21 12:12 | 1553184733.3795 | 1620 | 0.20956 | 0.000454 | -4e-06 | 101.2927 | 101.5157 | 125.5387 | 21.1341 | 0.001961 | 0.001816 | -152.3155 | 0.23963 | 8.3995 | 0.003302 | 0.005506 | 0.005394 | 5.1e-05 | 0.00101 | 0.323 | 1.2896 | 2.5481 | 99.48 | 20.93 | 37.156 | 68.8764 | 101.1633 | 65.3882 | 0.00926 | 1.0429 | 0.025028 | 23.9996 | -0.019673 | 0.14256 | 0.000316 | 0.02405 | 0.02405 | 0.24362
2019-03-21 12:12 | 1553184734.3801 | 1621 | 0.20958 | 0.000454 | -4e-06 | 101.2465 | 101.4153 | 125.2646 | 21.1335 | 0.001891 | 0.001838 | -152.3316 | 0.23933 | 8.4019 | 0.00334 | 0.004785 | 0.004644 | 5.5e-05 | 0.00101 | 0.30718 | 1.3221 | 2.5479 | 99.48 | 20.93 | 37.155 | 67.2128 | 102.0901 | 65.9255 | 0.009259 | 1.0412 | 0.024981 | 23.9924 | -0.014894 | 0.14238 | 0.000213 | 0.01622 | 0.01622 | 0.36078
2019-03-21 12:12 | 1553184735.3792 | 1622 | 0.20957 | 0.000455 | -4e-06 | 101.2635 | 101.4565 | 125.0723 | 21.1349 | 0.001935 | 0.0018 | -152.3328 | 0.24038 | 8.4 | 0.003313 | 0.004787 | 0.00457 | 5.3e-05 | 0.001009 | 0.32659 | 1.3248 | 2.548 | 99.48 | 20.93 | 37.154 | 68.2516 | 101.4261 | 65.0089 | 0.009259 | 1.044 | 0.024996 | 23.9416 | -0.017028 | 0.143 | 0.000243 | 0.018523 | 0.018523 | 0.31731
2019-03-21 12:12 | 1553184736.3795 | 1623 | 0.20957 | 0.000451 | -5e-06 | 101.2698 | 101.4905 | 124.9207 | 21.1337 | 0.001927 | 0.001795 | -152.3244 | 0.23978 | 8.4001 | 0.003308 | 0.004892 | 0.004732 | 5.5e-05 | 0.001012 | 0.35098 | 1.3487 | 2.548 | 99.48 | 20.93 | 37.154 | 68.0805 | 101.312 | 64.8898 | 0.009259 | 1.0444 | 0.024985 | 23.9228 | -0.019478 | 0.14264 | 0.000249 | 0.018987 | 0.018987 | 0.30877
2019-03-21 12:12 | 1553184737.3807 | 1624 | 0.20961 | 0.000452 | -5e-06 | 101.2447 | 101.6053 | 124.6829 | 21.1338 | 0.001923 | 0.001847 | -152.3103 | 0.23929 | 8.4054 | 0.003385 | 0.004647 | 0.004446 | 5.1e-05 | 0.001011 | 0.30879 | 1.3256 | 2.5478 | 99.48 | 20.93 | 37.154 | 67.9753 | 103.18 | 66.1372 | 0.009259 | 1.0406 | 0.024915 | 23.9441 | -0.031798 | 0.14235 | 5.5e-05 | 0.004387 | 0.004387 | 1.3336
2019-03-21 12:12 | 1553184738.3794 | 1625 | 0.20956 | 0.000451 | -4e-06 | 101.2075 | 101.4534 | 124.7705 | 21.1352 | 0.001934 | 0.001786 | -152.3141 | 0.24027 | 8.4047 | 0.003331 | 0.004849 | 0.004674 | 5.3e-05 | 0.001011 | 0.32045 | 1.3228 | 2.548 | 99.48 | 20.93 | 37.155 | 68.2411 | 101.8829 | 64.6728 | 0.009259 | 1.0451 | 0.024978 | 23.9008 | -0.021707 | 0.14293 | 0.00032 | 0.02425 | 0.02425 | 0.24225
2019-03-21 12:12 | 1553184739.3793 | 1626 | 0.20958 | 0.000452 | -5e-06 | 101.2343 | 101.5975 | 124.4203 | 21.1345 | 0.001901 | 0.001834 | -152.2911 | 0.23958 | 8.4025 | 0.003366 | 0.004824 | 0.004576 | 5e-05 | 0.001006 | 0.29247 | 1.2898 | 2.5479 | 99.48 | 20.93 | 37.154 | 67.4526 | 102.7155 | 65.8305 | 0.009259 | 1.0415 | 0.0249 | 23.9081 | -0.03203 | 0.14253 | 0.000191 | 0.01455 | 0.01455 | 0.4026
2019-03-21 12:12 | 1553184740.3796 | 1627 | 0.20959 | 0.000455 | -4e-06 | 101.1929 | 101.354 | 124.4729 | 21.1344 | 0.001916 | 0.001801 | -152.294 | 0.23985 | 8.4018 | 0.003332 | 0.005615 | 0.005438 | 5.6e-05 | 0.001007 | 0.28208 | 1.2894 | 2.5479 | 99.48 | 20.93 | 37.154 | 67.7988 | 101.8969 | 65.0484 | 0.009259 | 1.0439 | 0.024934 | 23.8855 | -0.014225 | 0.14269 | 0.000132 | 0.01013 | 0.01013 | 0.5789
2019-03-21 12:12 | 1553184741.3794 | 1628 | 0.20962 | 0.000449 | -4e-06 | 101.1863 | 101.4452 | 124.5663 | 21.1348 | 0.001903 | 0.001842 | -152.2804 | 0.24008 | 8.4023 | 0.00332 | 0.004671 | 0.004533 | 5.4e-05 | 0.001006 | 0.33035 | 1.2993 | 2.548 | 99.48 | 20.93 | 37.156 | 67.4981 | 101.6153 | 66.0204 | 0.00926 | 1.0409 | 0.024908 | 23.9288 | -0.022856 | 0.14282 | -5.2e-05 | -0.003649 | -0.003649 | -1.6087
2019-03-21 12:12 | 1553184742.3805 | 1629 | 0.20961 | 0.000453 | -4e-06 | 101.2087 | 101.3076 | 124.7334 | 21.1348 | 0.00191 | 0.001853 | -152.2794 | 0.23997 | 8.4021 | 0.003371 | 0.005428 | 0.005204 | 5.2e-05 | 0.001008 | 0.31797 | 1.2804 | 2.548 | 99.48 | 20.93 | 37.157 | 67.6628 | 102.836 | 66.2811 | 0.00926 | 1.0401 | 0.024915 | 23.954 | -0.00874 | 0.14276 | 5.4e-05 | 0.004276 | 0.004276 | 1.3722
2019-03-21 12:12 | 1553184743.379 | 1630 | 0.20948 | 0.000451 | -4e-06 | 101.2522 | 101.3482 | 125.0007 | 21.1349 | 0.001923 | 0.001776 | -152.2591 | 0.23941 | 8.4015 | 0.003408 | 0.005147 | 0.005018 | 5.3e-05 | 0.001008 | 0.32764 | 1.2642 | 2.5479 | 99.48 | 20.93 | 37.158 | 67.9663 | 103.732 | 64.4267 | 0.00926 | 1.0458 | 0.02501 | 23.9141 | -0.008475 | 0.14243 | 0.000815 | 0.061528 | 0.061528 | 0.09514
2019-03-21 12:12 | 1553184744.3786 | 1631 | 0.20958 | 0.000451 | -4e-06 | 101.1613 | 101.3274 | 125.2268 | 21.1348 | 0.001974 | 0.001834 | -152.2478 | 0.23963 | 8.4019 | 0.003431 | 0.005017 | 0.004911 | 5e-05 | 0.001006 | 0.3147 | 1.2814 | 2.5478 | 99.48 | 20.93 | 37.16 | 69.2047 | 104.2947 | 65.8267 | 0.009261 | 1.0415 | 0.024981 | 23.9853 | -0.014672 | 0.14255 | 0.000238 | 0.018107 | 0.018107 | 0.32357
2019-03-21 12:12 | 1553184745.3797 | 1632 | 0.20961 | 0.00045 | -4e-06 | 101.2027 | 101.2644 | 125.2336 | 21.1348 | 0.001918 | 0.001799 | -152.263 | 0.23861 | 8.4037 | 0.003342 | 0.004857 | 0.004703 | 5.3e-05 | 0.001004 | 0.32892 | 1.2913 | 2.5479 | 99.48 | 20.93 | 37.162 | 67.8453 | 102.1352 | 64.9823 | 0.009261 | 1.0441 | 0.025013 | 23.956 | -0.005447 | 0.14195 | 3.3e-05 | 0.002746 | 0.002746 | 2.1246
2019-03-21 12:12 | 1553184746.3791 | 1633 | 0.20953 | 0.000453 | -3e-06 | 101.1217 | 101.2596 | 125.281 | 21.1354 | 0.001956 | 0.001865 | -152.2636 | 0.24013 | 8.404 | 0.003374 | 0.004806 | 0.004646 | 5.5e-05 | 0.001 | 0.3095 | 1.296 | 2.5479 | 99.48 | 20.93 | 37.164 | 68.7724 | 102.9174 | 66.5835 | 0.009262 | 1.0392 | 0.024958 | 24.0172 | -0.012192 | 0.14286 | 0.000487 | 0.036747 | 0.036747 | 0.15978
2019-03-21 12:12 | 1553184747.3787 | 1634 | 0.20957 | 0.000455 | -4e-06 | 101.0749 | 101.2416 | 125.1869 | 21.1348 | 0.001901 | 0.001775 | -152.2692 | 0.24 | 8.4031 | 0.003381 | 0.005043 | 0.004799 | 5e-05 | 0.001004 | 0.2844 | 1.2911 | 2.5479 | 99.48 | 20.93 | 37.165 | 67.4418 | 103.0834 | 64.4217 | 0.009262 | 1.0458 | 0.025029 | 23.9317 | -0.014735 | 0.14278 | 0.00025 | 0.019105 | 0.019105 | 0.30714
2019-03-21 12:12 | 1553184748.3788 | 1635 | 0.20972 | 0.000452 | -4e-06 | 101.1133 | 101.2222 | 125.0418 | 21.1348 | 0.001889 | 0.001809 | -152.2863 | 0.24059 | 8.4016 | 0.003303 | 0.00536 | 0.00524 | 5.2e-05 | 0.001006 | 0.31636 | 1.3014 | 2.5478 | 99.48 | 20.94 | 37.166 | 67.1564 | 101.2032 | 65.222 | 0.009268 | 1.0434 | 0.024985 | 23.9462 | -0.00962 | 0.14313 | -0.000646 | -0.048219 | -0.048219 | -0.122
2019-03-21 12:12 | 1553184749.3793 | 1636 | 0.20957 | 0.000455 | -4e-06 | 101.094 | 101.2136 | 124.982 | 21.1351 | 0.001856 | 0.001788 | -152.2867 | 0.24012 | 8.4023 | 0.003408 | 0.004838 | 0.004685 | 5.1e-05 | 0.001006 | 0.29911 | 1.3299 | 2.5481 | 99.48 | 20.94 | 37.166 | 66.3631 | 103.7314 | 64.7226 | 0.009268 | 1.0449 | 0.024997 | 23.9228 | -0.01058 | 0.14284 | 0.000302 | 0.022944 | 0.022944 | 0.25588
2019-03-21 12:12 | 1553184750.3788 | 1637 | 0.2098 | 0.000449 | -4e-06 | 101.1405 | 101.2004 | 124.8705 | 21.1349 | 0.001948 | 0.001817 | -152.3013 | 0.23976 | 8.4012 | 0.003384 | 0.004284 | 0.004101 | 5.2e-05 | 0.001007 | 0.32421 | 1.3 | 2.548 | 99.48 | 20.94 | 37.166 | 68.5661 | 103.1566 | 65.414 | 0.009268 | 1.0428 | 0.02496 | 23.9366 | -0.005295 | 0.14264 | -0.001102 | -0.082376 | -0.082376 | -0.071166
2019-03-21 12:12 | 1553184751.3804 | 1638 | 0.20963 | 0.000453 | -4e-06 | 101.0973 | 101.2361 | 124.5076 | 21.1351 | 0.001933 | 0.001793 | -152.3227 | 0.23946 | 8.4034 | 0.003391 | 0.004404 | 0.004227 | 5.4e-05 | 0.001003 | 0.31636 | 1.3229 | 2.548 | 99.48 | 20.94 | 37.168 | 68.2154 | 103.3141 | 64.8467 | 0.009268 | 1.0445 | 0.024945 | 23.8817 | -0.012277 | 0.14245 | -9.4e-05 | -0.006796 | -0.006796 | -0.86145
2019-03-21 12:13 | 1553184752.3793 | 1639 | 0.20966 | 0.000451 | -4e-06 | 101.0538 | 101.2603 | 124.2245 | 21.1343 | 0.001903 | 0.001792 | -152.3374 | 0.23943 | 8.4032 | 0.003348 | 0.004346 | 0.004164 | 5.4e-05 | 0.001007 | 0.30702 | 1.3521 | 2.548 | 99.48 | 20.94 | 37.17 | 67.4875 | 102.2816 | 64.8319 | 0.009269 | 1.0446 | 0.024917 | 23.854 | -0.018262 | 0.14244 | -0.000252 | -0.018639 | -0.018639 | -0.31407
2019-03-21 12:13 | 1553184753.3801 | 1640 | 0.20957 | 0.000452 | -5e-06 | 101.0821 | 101.1636 | 124.032 | 21.1346 | 0.001907 | 0.001824 | -152.3355 | 0.23963 | 8.4014 | 0.003443 | 0.00449 | 0.004288 | 5.2e-05 | 0.001008 | 0.2959 | 1.3691 | 2.548 | 99.48 | 20.94 | 37.169 | 67.5782 | 104.5761 | 65.5924 | 0.009269 | 1.0422 | 0.02487 | 23.8623 | -0.007209 | 0.14255 | 0.000284 | 0.02152 | 0.02152 | 0.27226
2019-03-21 12:13 | 1553184754.3792 | 1641 | 0.20962 | 0.00045 | -4e-06 | 101.0383 | 101.2933 | 124.0442 | 21.1349 | 0.001887 | 0.001821 | -152.3201 | 0.24028 | 8.4024 | 0.003487 | 0.004715 | 0.004619 | 5.1e-05 | 0.001008 | 0.30093 | 1.3777 | 2.5478 | 99.48 | 20.94 | 37.17 | 67.1033 | 105.6499 | 65.5208 | 0.009269 | 1.0424 | 0.024874 | 23.861 | -0.022552 | 0.14294 | -1.5e-05 | -0.000879 | -0.000879 | -6.6799
2019-03-21 12:13 | 1553184755.3803 | 1642 | 0.20966 | 0.000453 | -4e-06 | 101.071 | 101.2903 | 124.1804 | 21.1347 | 0.001869 | 0.001862 | -152.3047 | 0.23938 | 8.4036 | 0.003408 | 0.00451 | 0.004375 | 5.5e-05 | 0.001008 | 0.30525 | 1.3631 | 2.548 | 99.48 | 20.94 | 37.17 | 66.6826 | 103.7444 | 66.508 | 0.009269 | 1.0394 | 0.024851 | 23.9089 | -0.01938 | 0.14241 | -0.000288 | -0.021269 | -0.021269 | -0.27519
2019-03-21 12:13 | 1553184756.3792 | 1643 | 0.20957 | 0.000453 | -4e-06 | 101.0514 | 100.9912 | 124.0837 | 21.1344 | 0.001952 | 0.001813 | -152.3037 | 0.23938 | 8.4031 | 0.00346 | 0.004572 | 0.004361 | 5e-05 | 0.00101 | 0.30773 | 1.3727 | 2.548 | 99.48 | 20.94 | 37.17 | 68.6787 | 105.007 | 65.3169 | 0.009269 | 1.0431 | 0.024885 | 23.8576 | 0.005325 | 0.14241 | 0.000275 | 0.020797 | 0.020797 | 0.28143
2019-03-21 12:13 | 1553184757.3793 | 1644 | 0.20966 | 0.000452 | -3e-06 | 101.06 | 101.1728 | 123.9384 | 21.1329 | 0.001884 | 0.001755 | -152.3082 | 0.23997 | 8.4003 | 0.003403 | 0.004756 | 0.004632 | 5e-05 | 0.001012 | 0.30596 | 1.3329 | 2.548 | 99.48 | 20.94 | 37.171 | 67.0291 | 103.6215 | 63.9422 | 0.009269 | 1.0473 | 0.024921 | 23.7952 | -0.009974 | 0.14276 | -0.000283 | -0.020975 | -0.020975 | -0.27973
2019-03-21 12:13 | 1553184758.3807 | 1645 | 0.20968 | 0.000455 | -4e-06 | 100.9548 | 101.198 | 124.0935 | 21.134 | 0.001888 | 0.001809 | -152.2942 | 0.23956 | 8.403 | 0.00342 | 0.004628 | 0.004441 | 5.4e-05 | 0.001012 | 0.2855 | 1.2696 | 2.548 | 99.48 | 20.93 | 37.171 | 67.1345 | 104.033 | 65.2313 | 0.009263 | 1.0433 | 0.024889 | 23.8555 | -0.021515 | 0.14251 | -0.000392 | -0.029044 | -0.029044 | -0.20167
2019-03-21 12:13 | 1553184759.3797 | 1646 | 0.20953 | 0.000453 | -4e-06 | 101.0302 | 101.0713 | 124.3912 | 21.1344 | 0.001888 | 0.001796 | -152.3153 | 0.23968 | 8.4012 | 0.003419 | 0.005042 | 0.004888 | 5.4e-05 | 0.00101 | 0.27992 | 1.2649 | 2.548 | 99.48 | 20.93 | 37.171 | 67.1341 | 104.0004 | 64.9249 | 0.009263 | 1.0443 | 0.02493 | 23.8733 | -0.003633 | 0.14259 | 0.000514 | 0.038745 | 0.038745 | 0.15125
2019-03-21 12:13 | 1553184760.3795 | 1647 | 0.2097 | 0.000454 | -4e-06 | 101.0144 | 101.2089 | 124.4538 | 21.134 | 0.001893 | 0.001818 | -152.3033 | 0.24004 | 8.4052 | 0.003433 | 0.004664 | 0.004434 | 5.3e-05 | 0.001011 | 0.27704 | 1.2908 | 2.5478 | 99.48 | 20.93 | 37.172 | 67.2516 | 104.3465 | 65.449 | 0.009264 | 1.0427 | 0.024917 | 23.8978 | -0.0172 | 0.1428 | -0.000509 | -0.037876 | -0.037876 | -0.15496
2019-03-21 12:13 | 1553184761.3786 | 1648 | 0.20961 | 0.000449 | -4e-06 | 100.9977 | 101.0038 | 124.44 | 21.1343 | 0.001923 | 0.001833 | -152.2949 | 0.23942 | 8.4012 | 0.003408 | 0.004689 | 0.00446 | 4.9e-05 | 0.001009 | 0.30502 | 1.3435 | 2.5479 | 99.48 | 20.93 | 37.173 | 67.964 | 103.7401 | 65.8005 | 0.009264 | 1.0416 | 0.024903 | 23.9089 | -0.000539 | 0.14243 | 5.8e-05 | 0.004586 | 0.004586 | 1.2766
2019-03-21 12:13 | 1553184762.379 | 1649 | 0.20957 | 0.000453 | -4e-06 | 101.0478 | 101.0883 | 124.3405 | 21.1343 | 0.001962 | 0.001853 | -152.2838 | 0.23916 | 8.4016 | 0.003374 | 0.004939 | 0.00477 | 5e-05 | 0.001007 | 0.3027 | 1.3259 | 2.5478 | 99.48 | 20.93 | 37.173 | 68.9071 | 102.9141 | 66.2868 | 0.009264 | 1.0401 | 0.024875 | 23.9165 | -0.003584 | 0.14228 | 0.0003 | 0.022665 | 0.022665 | 0.258
2019-03-21 12:13 | 1553184763.3786 | 1650 | 0.20965 | 0.000453 | -4e-06 | 101.0049 | 101.1353 | 124.428 | 21.1353 | 0.001931 | 0.001822 | -152.2869 | 0.23905 | 8.4003 | 0.00338 | 0.004869 | 0.004718 | 4.9e-05 | 0.001006 | 0.33588 | 1.3125 | 2.5479 | 99.48 | 20.93 | 37.174 | 68.1609 | 103.0502 | 65.5417 | 0.009264 | 1.0424 | 0.024911 | 23.8986 | -0.01154 | 0.14221 | -0.000178 | -0.01309 | -0.01309 | -0.44651
2019-03-21 12:13 | 1553184764.3786 | 1651 | 0.20955 | 0.000456 | -4e-06 | 101.0094 | 101.1871 | 124.3686 | 21.1345 | 0.00192 | 0.001809 | -152.2926 | 0.239 | 8.399 | 0.003397 | 0.005228 | 0.005074 | 4.9e-05 | 0.001002 | 0.31166 | 1.3339 | 2.548 | 99.48 | 20.93 | 37.173 | 67.9114 | 103.4605 | 65.2405 | 0.009264 | 1.0433 | 0.024917 | 23.8823 | -0.015722 | 0.14218 | 0.000406 | 0.030689 | 0.030689 | 0.19042
2019-03-21 12:13 | 1553184765.3801 | 1652 | 0.20963 | 0.000447 | -5e-06 | 100.9569 | 101.1646 | 124.1738 | 21.1341 | 0.001938 | 0.001868 | -152.3061 | 0.24005 | 8.4018 | 0.003465 | 0.004809 | 0.004677 | 4.9e-05 | 0.000999 | 0.3027 | 1.3384 | 2.5479 | 99.48 | 20.93 | 37.173 | 68.3242 | 105.1167 | 66.6423 | 0.009264 | 1.039 | 0.024846 | 23.9129 | -0.01838 | 0.14281 | -8.9e-05 | -0.006334 | -0.006334 | -0.9266
2019-03-21 12:13 | 1553184766.3801 | 1653 | 0.20961 | 0.000455 | -4e-06 | 101.0106 | 101.1976 | 124.1212 | 21.1343 | 0.001954 | 0.001837 | -152.3252 | 0.24033 | 8.4022 | 0.003437 | 0.004534 | 0.004368 | 5.2e-05 | 0.000998 | 0.30807 | 1.3606 | 2.548 | 99.48 | 20.93 | 37.175 | 68.73 | 104.4484 | 65.9037 | 0.009264 | 1.0413 | 0.024867 | 23.8819 | -0.01655 | 0.14297 | 3.9e-05 | 0.003173 | 0.003173 | 1.8519
2019-03-21 12:13 | 1553184767.3799 | 1654 | 0.20963 | 0.00045 | -4e-06 | 100.9829 | 101.1155 | 124.4656 | 21.1349 | 0.00197 | 0.001775 | -152.3144 | 0.23993 | 8.4027 | 0.003412 | 0.00472 | 0.00456 | 5.2e-05 | 0.001003 | 0.35153 | 1.3039 | 2.5478 | 99.48 | 20.93 | 37.175 | 69.1097 | 103.842 | 64.4094 | 0.009264 | 1.0459 | 0.024957 | 23.8622 | -0.011746 | 0.14274 | -7.2e-05 | -0.005182 | -0.005182 | -1.132
2019-03-21 12:13 | 1553184768.3799 | 1655 | 0.20959 | 0.000453 | -4e-06 | 101.0153 | 101.1169 | 124.417 | 21.1344 | 0.001941 | 0.001809 | -152.3164 | 0.23923 | 8.4016 | 0.003482 | 0.005299 | 0.005134 | 5.5e-05 | 0.001003 | 0.31813 | 1.3035 | 2.5479 | 99.48 | 20.93 | 37.173 | 68.3979 | 105.5198 | 65.2411 | 0.009264 | 1.0433 | 0.024921 | 23.8869 | -0.00899 | 0.14232 | 0.000174 | 0.013331 | 0.013331 | 0.43879
2019-03-21 12:13 | 1553184769.3799 | 1656 | 0.20967 | 0.00045 | -4e-06 | 100.9691 | 101.0328 | 124.3215 | 21.1345 | 0.001901 | 0.001793 | -152.3029 | 0.23957 | 8.4029 | 0.003395 | 0.005135 | 0.004961 | 5.6e-05 | 0.001006 | 0.33987 | 1.3327 | 2.5479 | 99.48 | 20.93 | 37.173 | 67.437 | 103.4203 | 64.8579 | 0.009264 | 1.0445 | 0.024926 | 23.8642 | -0.005636 | 0.14252 | -0.000308 | -0.022819 | -0.022819 | -0.2567
2019-03-21 12:13 | 1553184770.38 | 1657 | 0.2096 | 0.00045 | -5e-06 | 100.9805 | 101.1384 | 124.5304 | 21.1341 | 0.001954 | 0.001767 | -152.2806 | 0.23965 | 8.4024 | 0.003395 | 0.005055 | 0.004981 | 5.2e-05 | 0.001004 | 0.30773 | 1.2866 | 2.5478 | 99.48 | 20.93 | 37.174 | 68.7134 | 103.43 | 64.2141 | 0.009264 | 1.0465 | 0.024971 | 23.8615 | -0.013975 | 0.14257 | 0.000108 | 0.008435 | 0.008435 | 0.69472
2019-03-21 12:13 | 1553184771.3801 | 1658 | 0.20962 | 0.000458 | -4e-06 | 101.0165 | 101.0311 | 124.7881 | 21.1333 | 0.001925 | 0.001827 | -152.2754 | 0.23946 | 8.4024 | 0.003456 | 0.004347 | 0.004165 | 5.4e-05 | 0.001009 | 0.3183 | 1.3096 | 2.5478 | 99.48 | 20.93 | 37.175 | 68.0162 | 104.8997 | 65.667 | 0.009264 | 1.042 | 0.024943 | 23.9376 | -0.00129 | 0.14245 | -5.2e-05 | -0.003661 | -0.003661 | -1.5992
2019-03-21 12:13 | 1553184772.3801 | 1659 | 0.20964 | 0.000449 | -4e-06 | 100.9596 | 101.1087 | 125.0918 | 21.1338 | 0.001912 | 0.001804 | -152.2735 | 0.23921 | 8.4032 | 0.003354 | 0.004267 | 0.004125 | 5.3e-05 | 0.001009 | 0.31564 | 1.2915 | 2.5479 | 99.48 | 20.93 | 37.178 | 67.7171 | 102.4366 | 65.1038 | 0.009265 | 1.0437 | 0.024994 | 23.9468 | -0.013202 | 0.1423 | -0.000156 | -0.011487 | -0.011487 | -0.50916
2019-03-21 12:13 | 1553184773.3802 | 1660 | 0.20961 | 0.000453 | -4e-06 | 100.945 | 101.119 | 124.97 | 21.1348 | 0.001903 | 0.001818 | -152.278 | 0.24032 | 8.4026 | 0.003365 | 0.004739 | 0.004605 | 5.5e-05 | 0.001009 | 0.32477 | 1.3509 | 2.548 | 99.48 | 20.93 | 37.181 | 67.4825 | 102.6875 | 65.4592 | 0.009266 | 1.0426 | 0.024969 | 23.9477 | -0.015399 | 0.14297 | 6.7e-05 | 0.005277 | 0.005277 | 1.1136
2019-03-21 12:13 | 1553184774.3793 | 1661 | 0.20964 | 0.000449 | -4e-06 | 100.9995 | 101.0798 | 124.8871 | 21.1344 | 0.001936 | 0.001769 | -152.2903 | 0.23991 | 8.4016 | 0.003401 | 0.00468 | 0.00448 | 5e-05 | 0.001003 | 0.32925 | 1.3266 | 2.5481 | 99.48 | 20.93 | 37.186 | 68.2838 | 103.5633 | 64.2769 | 0.009267 | 1.0463 | 0.025004 | 23.8979 | -0.007104 | 0.14272 | -0.00016 | -0.011732 | -0.011732 | -0.50001
2019-03-21 12:13 | 1553184775.3792 | 1662 | 0.20966 | 0.000453 | -4e-06 | 100.9811 | 101.2654 | 124.9493 | 21.1347 | 0.001899 | 0.001814 | -152.281 | 0.23952 | 8.4022 | 0.003435 | 0.004686 | 0.004411 | 5e-05 | 0.001005 | 0.33804 | 1.3106 | 2.5478 | 99.48 | 20.93 | 37.189 | 67.3857 | 104.4008 | 65.36 | 0.009268 | 1.0429 | 0.02497 | 23.9422 | -0.025151 | 0.14249 | -0.000242 | -0.017898 | -0.017898 | -0.3272
2019-03-21 12:13 | 1553184776.3786 | 1663 | 0.2096 | 0.000454 | -4e-06 | 100.9823 | 101.2201 | 125.0963 | 21.1341 | 0.001934 | 0.001855 | -152.2711 | 0.23911 | 8.4031 | 0.003381 | 0.004907 | 0.004774 | 5.2e-05 | 0.001007 | 0.30093 | 1.3078 | 2.548 | 99.48 | 20.93 | 37.196 | 68.23 | 103.0852 | 66.3427 | 0.00927 | 1.0399 | 0.024949 | 23.991 | -0.021043 | 0.14225 | 0.000131 | 0.01007 | 0.01007 | 0.58055
2019-03-21 12:13 | 1553184777.3784 | 1664 | 0.2096 | 0.000455 | -4e-06 | 100.9441 | 101.0355 | 125.189 | 21.1339 | 0.001931 | 0.001764 | -152.2654 | 0.23908 | 8.4053 | 0.003383 | 0.005292 | 0.005254 | 5.2e-05 | 0.001004 | 0.29662 | 1.3355 | 2.5478 | 99.48 | 20.93 | 37.199 | 68.1649 | 103.1368 | 64.1492 | 0.00927 | 1.0467 | 0.025039 | 23.9222 | -0.008094 | 0.14223 | 7.4e-05 | 0.005839 | 0.005839 | 1.0011
2019-03-21 12:13 | 1553184778.3784 | 1665 | 0.20962 | 0.000457 | -4e-06 | 100.9534 | 101.0631 | 125.082 | 21.1354 | 0.001915 | 0.001827 | -152.2621 | 0.23938 | 8.4025 | 0.003459 | 0.004854 | 0.004557 | 5.1e-05 | 0.001006 | 0.30845 | 1.3093 | 2.5479 | 99.48 | 20.93 | 37.202 | 67.7858 | 104.9677 | 65.6659 | 0.009271 | 1.042 | 0.024972 | 23.9657 | -0.009718 | 0.14241 | -9e-06 | -0.000442 | -0.000442 | -13.2399
2019-03-21 12:13 | 1553184779.3788 | 1666 | 0.20958 | 0.000454 | -4e-06 | 100.9465 | 101.119 | 124.8909 | 21.135 | 0.001927 | 0.001769 | -152.2742 | 0.23943 | 8.4013 | 0.003486 | 0.004515 | 0.004362 | 5.3e-05 | 0.001003 | 0.30093 | 1.3477 | 2.548 | 99.48 | 20.93 | 37.211 | 68.0786 | 105.6217 | 64.2587 | 0.009273 | 1.0463 | 0.025005 | 23.8976 | -0.015267 | 0.14244 | 0.000189 | 0.014498 | 0.014498 | 0.40379
2019-03-21 12:13 | 1553184780.3796 | 1667 | 0.20966 | 0.000453 | -4e-06 | 100.9474 | 101.1959 | 124.4644 | 21.1358 | 0.001895 | 0.001838 | -152.2933 | 0.23882 | 8.4018 | 0.003454 | 0.004724 | 0.00458 | 5.2e-05 | 0.001006 | 0.33284 | 1.3339 | 2.548 | 99.48 | 20.93 | 37.215 | 67.2971 | 104.8587 | 65.934 | 0.009274 | 1.0412 | 0.024901 | 23.9159 | -0.021994 | 0.14208 | -0.000254 | -0.018733 | -0.018733 | -0.31171
2019-03-21 12:13 | 1553184781.3786 | 1668 | 0.20962 | 0.000454 | -4e-06 | 100.8479 | 101.1564 | 124.297 | 21.1352 | 0.001859 | 0.001804 | -152.2969 | 0.2396 | 8.4024 | 0.003407 | 0.004811 | 0.004659 | 5.6e-05 | 0.001007 | 0.32228 | 1.2944 | 2.5477 | 99.48 | 20.93 | 37.219 | 66.4366 | 103.7087 | 65.1153 | 0.009275 | 1.0437 | 0.024914 | 23.871 | -0.027323 | 0.14254 | -3.2e-05 | -0.002161 | -0.002161 | -2.7111
2019-03-21 12:13 | 1553184782.38 | 1669 | 0.2096 | 0.000449 | -4e-06 | 100.9358 | 101.1728 | 124.2064 | 21.1358 | 0.001911 | 0.001808 | -152.3084 | 0.23971 | 8.403 | 0.003447 | 0.005286 | 0.005072 | 5e-05 | 0.001007 | 0.33483 | 1.2725 | 2.5479 | 99.48 | 20.93 | 37.226 | 67.6894 | 104.6817 | 65.2111 | 0.009277 | 1.0434 | 0.024901 | 23.8657 | -0.020977 | 0.1426 | 0.000139 | 0.010627 | 0.010627 | 0.55152
2019-03-21 12:13 | 1553184783.3801 | 1670 | 0.20961 | 0.000453 | -4e-06 | 100.9027 | 100.9337 | 124.4691 | 21.1354 | 0.001935 | 0.001777 | -152.3111 | 0.23946 | 8.4002 | 0.003468 | 0.004795 | 0.004559 | 5.6e-05 | 0.001004 | 0.30359 | 1.2973 | 2.548 | 99.48 | 20.93 | 37.228 | 68.2664 | 105.1973 | 64.4748 | 0.009278 | 1.0457 | 0.024955 | 23.8649 | -0.002745 | 0.14245 | 2.6e-05 | 0.002182 | 0.002182 | 2.6837
2019-03-21 12:13 | 1553184784.3798 | 1671 | 0.20967 | 0.000451 | -4e-06 | 100.8205 | 101.1554 | 124.6174 | 21.1368 | 0.001883 | 0.001817 | -152.3053 | 0.24005 | 8.403 | 0.003536 | 0.00476 | 0.004572 | 5.5e-05 | 0.001005 | 0.26521 | 1.331 | 2.5478 | 99.48 | 20.93 | 37.232 | 67.0203 | 106.8339 | 65.419 | 0.009279 | 1.0428 | 0.024935 | 23.9125 | -0.029665 | 0.14281 | -0.000297 | -0.021994 | -0.021994 | -0.26686
2019-03-21 12:13 | 1553184785.3791 | 1672 | 0.20964 | 0.000453 | -4e-06 | 100.8324 | 100.9759 | 124.5081 | 21.1354 | 0.001889 | 0.001814 | -152.3054 | 0.23992 | 8.4003 | 0.003465 | 0.004846 | 0.004577 | 5.6e-05 | 0.001005 | 0.29789 | 1.3199 | 2.548 | 99.48 | 20.93 | 37.234 | 67.157 | 105.1257 | 65.3435 | 0.009279 | 1.043 | 0.024927 | 23.8993 | -0.012722 | 0.14273 | -0.000114 | -0.008276 | -0.008276 | -0.70879
2019-03-21 12:13 | 1553184786.3789 | 1673 | 0.20964 | 0.00045 | -4e-06 | 100.8434 | 100.9183 | 124.497 | 21.1364 | 0.001883 | 0.001768 | -152.3179 | 0.2395 | 8.404 | 0.003412 | 0.004557 | 0.004455 | 5.7e-05 | 0.001007 | 0.33301 | 1.3097 | 2.5478 | 99.48 | 20.93 | 37.234 | 67.0183 | 103.8318 | 64.2577 | 0.009279 | 1.0463 | 0.024966 | 23.8599 | -0.006644 | 0.14248 | -0.000143 | -0.010476 | -0.010476 | -0.55898
2019-03-21 12:13 | 1553184787.3788 | 1674 | 0.20961 | 0.000452 | -4e-06 | 100.861 | 100.9647 | 124.237 | 21.1369 | 0.001914 | 0.001772 | -152.3241 | 0.23902 | 8.4029 | 0.003454 | 0.004444 | 0.004378 | 5.5e-05 | 0.001009 | 0.33931 | 1.3244 | 2.5479 | 99.48 | 20.93 | 37.236 | 67.7514 | 104.8433 | 64.3395 | 0.00928 | 1.0461 | 0.024937 | 23.8378 | -0.00919 | 0.14219 | 7.5e-05 | 0.005874 | 0.005874 | 0.99496
2019-03-21 12:13 | 1553184788.3793 | 1675 | 0.20955 | 0.00045 | -4e-06 | 100.8339 | 100.9258 | 124.306 | 21.1362 | 0.001888 | 0.00182 | -152.3258 | 0.24037 | 8.4024 | 0.003484 | 0.004373 | 0.004164 | 5.7e-05 | 0.001004 | 0.3246 | 1.306 | 2.5479 | 99.48 | 20.93 | 37.237 | 67.1428 | 105.5821 | 65.4925 | 0.00928 | 1.0425 | 0.024901 | 23.8851 | -0.008154 | 0.143 | 0.000409 | 0.030849 | 0.030849 | 0.19051
2019-03-21 12:13 | 1553184789.38 | 1676 | 0.20967 | 0.000455 | -4e-06 | 100.8276 | 100.8547 | 124.5806 | 21.1365 | 0.001883 | 0.001823 | -152.3356 | 0.2401 | 8.4029 | 0.003419 | 0.005038 | 0.004829 | 5.6e-05 | 0.001009 | 0.33301 | 1.3432 | 2.5479 | 99.48 | 20.93 | 37.238 | 67.0059 | 103.9983 | 65.5748 | 0.00928 | 1.0423 | 0.024925 | 23.9144 | -0.002398 | 0.14283 | -0.000298 | -0.022005 | -0.022005 | -0.26677
2019-03-21 12:13 | 1553184790.38 | 1677 | 0.2096 | 0.000451 | -4e-06 | 100.8294 | 100.9575 | 124.9749 | 21.1357 | 0.001938 | 0.001834 | -152.3342 | 0.23927 | 8.4019 | 0.003416 | 0.004307 | 0.004116 | 5.2e-05 | 0.001008 | 0.2995 | 1.2972 | 2.5479 | 99.48 | 20.93 | 37.239 | 68.332 | 103.9236 | 65.8237 | 0.00928 | 1.0415 | 0.024956 | 23.961 | -0.011357 | 0.14234 | 0.000141 | 0.010808 | 0.010808 | 0.54127
2019-03-21 12:13 | 1553184791.3785 | 1678 | 0.20961 | 0.000451 | -4e-06 | 100.7582 | 100.7927 | 125.001 | 21.1366 | 0.001883 | 0.001778 | -152.3364 | 0.24017 | 8.4061 | 0.003398 | 0.004457 | 0.004264 | 5.6e-05 | 0.001009 | 0.3131 | 1.3602 | 2.548 | 99.48 | 20.93 | 37.239 | 67.0184 | 103.4846 | 64.4798 | 0.00928 | 1.0457 | 0.025008 | 23.916 | -0.003059 | 0.14288 | 7.2e-05 | 0.005634 | 0.005634 | 1.0424
2019-03-21 12:13 | 1553184792.3792 | 1679 | 0.20959 | 0.000452 | -4e-06 | 100.7293 | 100.8577 | 124.9397 | 21.1361 | 0.001936 | 0.001805 | -152.3332 | 0.23948 | 8.4029 | 0.003444 | 0.004442 | 0.004422 | 5.4e-05 | 0.001012 | 0.30381 | 1.316 | 2.5479 | 99.48 | 20.93 | 37.237 | 68.2959 | 104.6145 | 65.1475 | 0.00928 | 1.0436 | 0.024977 | 23.9338 | -0.011396 | 0.14247 | 0.000155 | 0.011878 | 0.011878 | 0.49294
2019-03-21 12:13 | 1553184793.3782 | 1680 | 0.20964 | 0.000455 | -4e-06 | 100.751 | 100.7579 | 124.6947 | 21.1364 | 0.001903 | 0.00176 | -152.3287 | 0.23938 | 8.4031 | 0.00342 | 0.004803 | 0.004594 | 5.3e-05 | 0.001009 | 0.29717 | 1.3259 | 2.5479 | 99.48 | 20.93 | 37.237 | 67.4878 | 104.0384 | 64.0625 | 0.00928 | 1.047 | 0.024993 | 23.8719 | -0.000611 | 0.14241 | -0.000154 | -0.01131 | -0.01131 | -0.51751
2019-03-21 12:13 | 1553184794.3796 | 1681 | 0.20962 | 0.000455 | -4e-06 | 100.7224 | 100.7596 | 124.5876 | 21.1361 | 0.001879 | 0.001753 | -152.3357 | 0.23905 | 8.4015 | 0.003405 | 0.004687 | 0.004502 | 5.6e-05 | 0.001012 | 0.32173 | 1.322 | 2.5479 | 99.48 | 20.93 | 37.235 | 66.9283 | 103.66 | 63.8897 | 0.009279 | 1.0475 | 0.024988 | 23.8555 | -0.003302 | 0.14221 | 2e-06 | 0.000391 | 0.000391 | 14.9457
2019-03-21 12:13 | 1553184795.3784 | 1682 | 0.20963 | 0.000449 | -4e-06 | 100.7242 | 100.9092 | 124.789 | 21.1369 | 0.001917 | 0.001801 | -152.3335 | 0.23947 | 8.4026 | 0.003371 | 0.004617 | 0.004434 | 5.9e-05 | 0.00101 | 0.28888 | 1.287 | 2.5479 | 99.48 | 20.93 | 37.232 | 67.8293 | 102.8505 | 65.0404 | 0.009279 | 1.0439 | 0.024966 | 23.9155 | -0.016406 | 0.14246 | -3.6e-05 | -0.00246 | -0.00246 | -2.3797
2019-03-21 12:13 | 1553184796.3784 | 1683 | 0.20959 | 0.000453 | -4e-06 | 100.7904 | 100.9858 | 124.8287 | 21.1373 | 0.001947 | 0.001822 | -152.3164 | 0.23881 | 8.4024 | 0.003352 | 0.004558 | 0.004397 | 5.3e-05 | 0.001005 | 0.29192 | 1.2849 | 2.5479 | 99.48 | 20.93 | 37.23 | 68.5591 | 102.39 | 65.5405 | 0.009278 | 1.0424 | 0.024952 | 23.937 | -0.017324 | 0.14207 | 0.000175 | 0.013367 | 0.013367 | 0.43683
2019-03-21 12:13 | 1553184797.3799 | 1684 | 0.20961 | 0.000448 | -4e-06 | 100.7466 | 100.8703 | 124.8191 | 21.1369 | 0.001928 | 0.001761 | -152.326 | 0.24019 | 8.4016 | 0.003382 | 0.005237 | 0.005049 | 5.5e-05 | 0.001006 | 0.3011 | 1.275 | 2.548 | 99.48 | 20.93 | 37.228 | 68.0892 | 103.0998 | 64.0835 | 0.009278 | 1.0469 | 0.025004 | 23.8846 | -0.01098 | 0.14289 | 3.2e-05 | 0.00266 | 0.00266 | 2.2074
2019-03-21 12:13 | 1553184798.3795 | 1685 | 0.20965 | 0.000452 | -4e-06 | 100.7156 | 100.7678 | 124.9517 | 21.1364 | 0.001905 | 0.001813 | -152.3406 | 0.23922 | 8.402 | 0.003334 | 0.004925 | 0.0047 | 5.5e-05 | 0.001009 | 0.30917 | 1.2808 | 2.5479 | 99.48 | 20.93 | 37.227 | 67.5528 | 101.9516 | 65.3195 | 0.009277 | 1.0431 | 0.024972 | 23.941 | -0.004636 | 0.14231 | -0.000195 | -0.014393 | -0.014393 | -0.40638
2019-03-21 12:13 | 1553184799.3799 | 1686 | 0.2096 | 0.000454 | -4e-06 | 100.7385 | 100.7004 | 125.0627 | 21.1368 | 0.001933 | 0.001794 | -152.3306 | 0.23989 | 8.4033 | 0.003397 | 0.004751 | 0.004577 | 5.8e-05 | 0.001007 | 0.32278 | 1.302 | 2.5479 | 99.48 | 20.92 | 37.225 | 68.2107 | 103.4806 | 64.8772 | 0.009271 | 1.0444 | 0.024999 | 23.936 | 0.003389 | 0.14271 | 0.000123 | 0.009524 | 0.009524 | 0.6159
2019-03-21 12:13 | 1553184800.3789 | 1687 | 0.2096 | 0.000455 | -4e-06 | 100.7656 | 100.6687 | 124.8393 | 21.1363 | 0.001908 | 0.001782 | -152.3488 | 0.24004 | 8.4043 | 0.003282 | 0.004897 | 0.004684 | 5.7e-05 | 0.001006 | 0.29856 | 1.2943 | 2.5479 | 99.48 | 20.92 | 37.22 | 67.6244 | 100.6993 | 64.5849 | 0.00927 | 1.0453 | 0.024988 | 23.9042 | 0.00861 | 0.1428 | 9.4e-05 | 0.007301 | 0.007301 | 0.8039
2019-03-21 12:13 | 1553184801.38 | 1688 | 0.20958 | 0.000457 | -4e-06 | 100.7219 | 100.7886 | 124.904 | 21.1367 | 0.001942 | 0.001829 | -152.3484 | 0.23925 | 8.402 | 0.003351 | 0.004908 | 0.004744 | 5.7e-05 | 0.001008 | 0.31924 | 1.3226 | 2.5478 | 99.48 | 20.92 | 37.22 | 68.4275 | 102.3485 | 65.7139 | 0.00927 | 1.0418 | 0.024953 | 23.9504 | -0.005924 | 0.14233 | 0.000256 | 0.019471 | 0.019471 | 0.30044
2019-03-21 12:13 | 1553184802.38 | 1689 | 0.20964 | 0.000449 | -5e-06 | 100.7153 | 100.8114 | 124.9892 | 21.1358 | 0.001935 | 0.001798 | -152.3479 | 0.23964 | 8.4022 | 0.003335 | 0.004614 | 0.004443 | 5.6e-05 | 0.001008 | 0.33411 | 1.3313 | 2.548 | 99.48 | 20.92 | 37.218 | 68.2574 | 101.9808 | 64.959 | 0.009269 | 1.0442 | 0.024989 | 23.9318 | -0.008531 | 0.14256 | -0.000141 | -0.010257 | -0.010257 | -0.57123
2019-03-21 12:13 | 1553184803.3798 | 1690 | 0.20965 | 0.000456 | -5e-06 | 100.6843 | 100.7566 | 125.0268 | 21.1352 | 0.00191 | 0.001742 | -152.3475 | 0.23985 | 8.4006 | 0.003319 | 0.004497 | 0.004282 | 5.3e-05 | 0.001008 | 0.3027 | 1.3031 | 2.5479 | 99.48 | 20.92 | 37.218 | 67.6663 | 101.5821 | 63.6164 | 0.009269 | 1.0483 | 0.025043 | 23.8879 | -0.006417 | 0.14269 | -0.000194 | -0.014305 | -0.014305 | -0.40996
2019-03-21 12:13 | 1553184804.3798 | 1691 | 0.20964 | 0.000455 | -4e-06 | 100.7558 | 100.904 | 125.2073 | 21.135 | 0.001906 | 0.001761 | -152.3351 | 0.24025 | 8.4028 | 0.003292 | 0.004917 | 0.004771 | 5.7e-05 | 0.001013 | 0.33124 | 1.318 | 2.548 | 99.48 | 20.92 | 37.22 | 67.5771 | 100.9358 | 64.0799 | 0.00927 | 1.0469 | 0.025043 | 23.9215 | -0.013149 | 0.14293 | -0.000135 | -0.009865 | -0.009865 | -0.59544
2019-03-21 12:13 | 1553184805.3793 | 1692 | 0.20963 | 0.000448 | -5e-06 | 100.718 | 100.8254 | 125.458 | 21.1359 | 0.001891 | 0.001841 | -152.3389 | 0.2393 | 8.4019 | 0.003283 | 0.004728 | 0.004515 | 5.3e-05 | 0.001009 | 0.30773 | 1.3239 | 2.5479 | 99.48 | 20.92 | 37.224 | 67.2176 | 100.7231 | 66.0144 | 0.009271 | 1.0409 | 0.024997 | 24.0141 | -0.009532 | 0.14236 | -7.3e-05 | -0.005221 | -0.005221 | -1.1208
2019-03-21 12:13 | 1553184806.379 | 1693 | 0.20954 | 0.00045 | -4e-06 | 100.7272 | 100.9749 | 125.4655 | 21.1361 | 0.001857 | 0.001834 | -152.3336 | 0.2389 | 8.401 | 0.003293 | 0.004965 | 0.004786 | 5.8e-05 | 0.00101 | 0.32084 | 1.2935 | 2.548 | 99.48 | 20.92 | 37.227 | 66.3867 | 100.9635 | 65.8328 | 0.009272 | 1.0415 | 0.025004 | 24.0083 | -0.021966 | 0.14212 | 0.000472 | 0.035676 | 0.035676 | 0.16373
2019-03-21 12:13 | 1553184807.379 | 1694 | 0.2096 | 0.000453 | -3e-06 | 100.8133 | 100.8114 | 125.3387 | 21.1353 | 0.001927 | 0.001785 | -152.3444 | 0.23956 | 8.4018 | 0.003279 | 0.003986 | 0.00381 | 5.7e-05 | 0.00101 | 0.33588 | 1.2679 | 2.5479 | 99.48 | 20.92 | 37.229 | 68.0672 | 100.6071 | 64.6486 | 0.009272 | 1.0451 | 0.025035 | 23.9543 | 0.00017 | 0.14251 | 0.000129 | 0.009937 | 0.009937 | 0.58946
2019-03-21 12:13 | 1553184808.3784 | 1695 | 0.2096 | 0.000458 | -4e-06 | 100.7445 | 100.7947 | 125.258 | 21.1356 | 0.001895 | 0.001794 | -152.3332 | 0.23963 | 8.4037 | 0.003263 | 0.004294 | 0.004107 | 5.2e-05 | 0.00101 | 0.29679 | 1.3183 | 2.5479 | 99.48 | 20.92 | 37.235 | 67.2969 | 100.2228 | 64.8801 | 0.009274 | 1.0444 | 0.025019 | 23.9548 | -0.004457 | 0.14255 | 9.3e-05 | 0.007299 | 0.007299 | 0.80267
2019-03-21 12:13 | 1553184809.3782 | 1696 | 0.20961 | 0.000456 | -4e-06 | 100.8479 | 100.9054 | 125.3392 | 21.1358 | 0.001913 | 0.001793 | -152.3353 | 0.23936 | 8.399 | 0.00327 | 0.004263 | 0.004067 | 5.5e-05 | 0.001013 | 0.29662 | 1.3199 | 2.548 | 99.48 | 20.92 | 37.237 | 67.7423 | 100.4037 | 64.8442 | 0.009274 | 1.0445 | 0.025028 | 23.9612 | -0.005101 | 0.14239 | 3.7e-05 | 0.003016 | 0.003016 | 1.9406
2019-03-21 12:13 | 1553184810.3791 | 1697 | 0.20958 | 0.000455 | -4e-06 | 100.7603 | 100.9201 | 125.3688 | 21.1355 | 0.001934 | 0.0018 | -152.3279 | 0.23909 | 8.402 | 0.003268 | 0.004625 | 0.004475 | 5.6e-05 | 0.001011 | 0.29264 | 1.3232 | 2.548 | 99.48 | 20.92 | 37.244 | 68.2304 | 100.3493 | 65.0112 | 0.009276 | 1.044 | 0.025025 | 23.97 | -0.014171 | 0.14223 | 0.000247 | 0.018816 | 0.018816 | 0.31069
2019-03-21 12:14 | 1553184811.3796 | 1698 | 0.20939 | 0.000455 | -4e-06 | 100.7636 | 100.8182 | 125.3784 | 21.1361 | 0.001925 | 0.001772 | -152.3149 | 0.24004 | 8.4029 | 0.003272 | 0.004621 | 0.004486 | 5.6e-05 | 0.001013 | 0.32565 | 1.3226 | 2.548 | 99.48 | 20.92 | 37.247 | 68.0189 | 100.4478 | 64.3349 | 0.009277 | 1.0461 | 0.025051 | 23.9469 | -0.00485 | 0.1428 | 0.001378 | 0.10397 | 0.10397 | 0.056453
2019-03-21 12:14 | 1553184812.3791 | 1699 | 0.20946 | 0.000452 | -4e-06 | 100.7141 | 100.9306 | 125.519 | 21.136 | 0.001909 | 0.001798 | -152.3256 | 0.23992 | 8.4033 | 0.003228 | 0.004739 | 0.004593 | 5.7e-05 | 0.001016 | 0.32045 | 1.3215 | 2.5478 | 99.48 | 20.92 | 37.253 | 67.6459 | 99.3776 | 64.9599 | 0.009278 | 1.0442 | 0.025042 | 23.9825 | -0.019207 | 0.14273 | 0.000992 | 0.074878 | 0.074878 | 0.078343
2019-03-21 12:14 | 1553184813.3783 | 1700 | 0.2096 | 0.000454 | -4e-06 | 100.737 | 100.8475 | 125.347 | 21.1356 | 0.001913 | 0.001765 | -152.338 | 0.23951 | 8.4001 | 0.003304 | 0.004392 | 0.004213 | 5.4e-05 | 0.001011 | 0.30934 | 1.3212 | 2.5479 | 99.48 | 20.92 | 37.256 | 67.7328 | 101.2197 | 64.1804 | 0.009279 | 1.0466 | 0.025054 | 23.9384 | -0.009802 | 0.14249 | 0.000129 | 0.009977 | 0.009977 | 0.58699
2019-03-21 12:14 | 1553184814.3787 | 1701 | 0.20956 | 0.000455 | -4e-06 | 100.7907 | 100.8996 | 125.0503 | 21.1362 | 0.001889 | 0.001861 | -152.3435 | 0.23909 | 8.4029 | 0.003301 | 0.005092 | 0.004879 | 5.2e-05 | 0.001014 | 0.33301 | 1.323 | 2.5479 | 99.48 | 20.92 | 37.258 | 67.1554 | 101.1584 | 66.4931 | 0.009279 | 1.0395 | 0.024939 | 23.9919 | -0.009662 | 0.14224 | 0.000339 | 0.025631 | 0.025631 | 0.22808
2019-03-21 12:14 | 1553184815.3784 | 1702 | 0.20963 | 0.000455 | -4e-06 | 100.7275 | 100.9899 | 124.6827 | 21.1355 | 0.001887 | 0.001769 | -152.3494 | 0.23974 | 8.4019 | 0.003208 | 0.004874 | 0.004722 | 5.2e-05 | 0.001017 | 0.33915 | 1.3134 | 2.5479 | 99.48 | 20.92 | 37.261 | 67.1049 | 98.8964 | 64.2593 | 0.00928 | 1.0463 | 0.024984 | 23.8777 | -0.023267 | 0.14262 | -7.4e-05 | -0.005313 | -0.005313 | -1.1032
2019-03-21 12:14 | 1553184816.3794 | 1703 | 0.20966 | 0.000456 | -4e-06 | 100.7299 | 100.9136 | 124.8381 | 21.1365 | 0.001897 | 0.001816 | -152.3613 | 0.23972 | 8.4019 | 0.00336 | 0.004376 | 0.004194 | 5.5e-05 | 0.001015 | 0.3314 | 1.3205 | 2.5479 | 99.48 | 20.92 | 37.261 | 67.3593 | 102.5848 | 65.4 | 0.00928 | 1.0428 | 0.024958 | 23.9329 | -0.016296 | 0.14261 | -0.000251 | -0.018584 | -0.018584 | -0.31539
2019-03-21 12:14 | 1553184817.3785 | 1704 | 0.20963 | 0.000455 | -4e-06 | 100.7779 | 100.763 | 124.9606 | 21.1365 | 0.001919 | 0.001815 | -152.35 | 0.23977 | 8.4032 | 0.003259 | 0.004329 | 0.004151 | 5.3e-05 | 0.00102 | 0.32532 | 1.3215 | 2.5478 | 99.48 | 20.92 | 37.261 | 67.8873 | 100.1429 | 65.3727 | 0.00928 | 1.0429 | 0.024971 | 23.9437 | 0.001316 | 0.14264 | -3.7e-05 | -0.00249 | -0.00249 | -2.3541
2019-03-21 12:14 | 1553184818.3796 | 1705 | 0.20957 | 0.000459 | -4e-06 | 100.7838 | 100.888 | 124.8963 | 21.1362 | 0.00189 | 0.001823 | -152.3315 | 0.23941 | 8.404 | 0.003298 | 0.004155 | 0.004004 | 5.6e-05 | 0.001018 | 0.33268 | 1.3244 | 2.548 | 99.48 | 20.92 | 37.263 | 67.1703 | 101.0727 | 65.5589 | 0.009281 | 1.0423 | 0.024958 | 23.9441 | -0.009244 | 0.14243 | 0.000312 | 0.02366 | 0.02366 | 0.24741
2019-03-21 12:14 | 1553184819.3786 | 1706 | 0.20959 | 0.000458 | -4e-06 | 100.7808 | 101.1053 | 124.4125 | 21.1363 | 0.001909 | 0.001807 | -152.3375 | 0.2384 | 8.4039 | 0.003242 | 0.004162 | 0.003998 | 5.5e-05 | 0.001016 | 0.34252 | 1.3162 | 2.5481 | 99.48 | 20.92 | 37.264 | 67.633 | 99.7169 | 65.1922 | 0.009281 | 1.0435 | 0.024923 | 23.8848 | -0.028753 | 0.14182 | 0.000183 | 0.014013 | 0.014013 | 0.41597
2019-03-21 12:14 | 1553184820.3786 | 1707 | 0.20958 | 0.000454 | -4e-06 | 100.7993 | 100.9493 | 124.033 | 21.1358 | 0.001908 | 0.001807 | -152.3292 | 0.23938 | 8.4002 | 0.003254 | 0.004092 | 0.003944 | 5.5e-05 | 0.001011 | 0.33212 | 1.2881 | 2.5479 | 99.48 | 20.92 | 37.264 | 67.6114 | 100.0185 | 65.1902 | 0.009281 | 1.0435 | 0.024885 | 23.8483 | -0.013303 | 0.14241 | 0.000213 | 0.016219 | 0.016219 | 0.36088
2019-03-21 12:14 | 1553184821.3781 | 1708 | 0.20964 | 0.000458 | -4e-06 | 100.8297 | 100.7971 | 123.8217 | 21.1385 | 0.001902 | 0.001763 | -152.3335 | 0.23972 | 8.4032 | 0.003307 | 0.004592 | 0.004404 | 5.6e-05 | 0.001009 | 0.31653 | 1.2639 | 2.5479 | 99.48 | 20.92 | 37.265 | 67.4658 | 101.3041 | 64.1341 | 0.009281 | 1.0467 | 0.024902 | 23.7907 | 0.002893 | 0.14261 | -0.000122 | -0.008906 | -0.008906 | -0.65817
2019-03-21 12:14 | 1553184822.3787 | 1709 | 0.20956 | 0.00046 | -4e-06 | 100.844 | 101.0835 | 123.6929 | 21.1365 | 0.001892 | 0.001797 | -152.319 | 0.23936 | 8.4036 | 0.003311 | 0.004528 | 0.004209 | 5.5e-05 | 0.001005 | 0.33876 | 1.3193 | 2.5479 | 99.48 | 20.92 | 37.266 | 67.2251 | 101.3987 | 64.9495 | 0.009281 | 1.0442 | 0.024859 | 23.8071 | -0.02122 | 0.14239 | 0.000339 | 0.025547 | 0.025547 | 0.22908
2019-03-21 12:14 | 1553184823.379 | 1710 |  | 0.00046 |  | 100.8214 | 101.0621 | 123.8612 | 21.1366 | 0.001936 | 0.001803 | -152.3116 | 0.23923 | 8.4022 | 0.00335 | 0.004482 | 0.0043 | 5.2e-05 | 0.001 | 0.33268 | 1.3152 | 2.548 | 99.48 | 20.92 | 37.266 | 68.2846 | 102.328 | 65.0871 | 0.009281 | 1.0438 | 0.024871 | 23.8281 | -0.02133 | 0.14232
2019-03-21 12:14 | 1553184824.379 | 1711 |  | 0.000456 |  | 100.7838 | 100.9143 | 124.1212 | 21.1367 | 0.00195 | 0.001813 | -152.3233 | 0.23986 | 8.4032 | 0.00339 | 0.004823 | 0.004696 | 5e-05 | 0.000997 | 0.34202 | 1.3199 | 2.5479 | 99.48 | 20.92 | 37.265 | 68.6224 | 103.2911 | 65.3361 | 0.009281 | 1.043 | 0.024888 | 23.8619 | -0.011569 | 0.14269
2019-03-21 12:14 | 1553184825.379 | 1712 |  |  |  | 100.883 | 100.9888 | 124.1856 | 21.1366 | 0.001867 | 0.001787 | -152.3185 | 0.23869 | 8.4022 | 0.003372 | 0.004906 | 0.004734 | 5.4e-05 | 0.000998 | 0.31006 | 1.3219 | 2.5479 | 99.48 | 20.91 | 37.264 | 66.6236 | 102.8692 | 64.7077 | 0.009275 | 1.045 | 0.024918 | 23.8459 | -0.009376 | 0.142
2019-03-21 12:14 | 1553184826.379 | 1713 |  |  |  | 100.8815 | 100.966 | 123.9135 | 21.1363 | 0.001886 | 0.001789 | -152.3102 | 0.239 | 8.4041 | 0.003316 | 0.004701 | 0.004476 | 5.6e-05 | 0.001002 | 0.34346 | 1.331 | 2.5481 | 99.48 | 20.91 | 37.264 | 67.079 | 101.5157 | 64.7558 | 0.009275 | 1.0448 | 0.024889 | 23.8215 | -0.007487 | 0.14218
2019-03-21 12:14 | 1553184827.379 | 1714 |  |  |  | 100.872 | 100.9715 | 123.9621 | 21.1364 | 0.001882 | 0.001782 | -152.3055 | 0.23978 | 8.4029 | 0.003369 | 0.005728 | 0.005587 | 5.5e-05 | 0.001004 | 0.32477 | 1.3241 | 2.5479 | 99.48 | 20.91 | 37.261 | 66.987 | 102.7944 | 64.5826 | 0.009274 | 1.0453 | 0.0249 | 23.82 | -0.008815 | 0.14264
2019-03-21 12:14 | 1553184828.379 | 1715 |  |  |  | 100.8395 | 100.933 |  | 21.1369 | 0.00188 | 0.001834 | -152.3168 | 0.23937 | 8.4011 | 0.003304 | 0.004737 | 0.004584 | 5.5e-05 | 0.001002 | 0.33085 | 1.3537 | 2.5479 | 99.47 | 20.91 | 37.261 | 66.9533 | 101.2311 | 65.8468 | 0.009275 | 1.0414 |  |  | -0.008286 | 0.1424
2019-03-21 12:14 | 1553184829.379 | 1716 |  |  |  | 100.8884 | 101.182 |  | 21.1362 | 0.0019 |  | -152.3044 | 0.23889 | 8.4019 | 0.003333 | 0.004806 | 0.004666 | 5.1e-05 | 0.001004 | 0.3126 | 1.3083 | 2.5479 | 99.47 | 20.91 | 37.26 | 67.4206 | 101.92 |  | 0.009275 |  |  |  | -0.025989 | 0.14212
2019-03-21 12:14 | 1553184830.379 | 1717 |  |  |  | 100.8339 | 101.0696 |  | 21.137 | 0.001926 |  |  | 0.23898 | 8.4018 | 0.003311 | 0.005055 | 0.004791 | 5.6e-05 | 0.000998 | 0.29302 | 1.3349 | 2.5482 | 99.47 | 20.91 | 37.258 | 68.0424 | 101.3978 |  | 0.009275 |  |  |  | -0.020883 | 0.14217
2019-03-21 12:14 | 1553184831.379 | 1718 |  |  |  | 100.8807 | 100.9739 |  | 21.1371 |  |  |  | 0.23908 | 8.3999 | 0.003316 | 0.004124 | 0.003969 | 5.6e-05 | 0.000993 | 0.32388 | 1.2984 | 2.5478 | 99.48 | 20.91 | 37.257 |  | 101.5094 |  | 0.009273 |  |  |  | -0.008261 | 0.14223
2019-03-21 12:14 | 1553184832.379 | 1719 |  |  |  | 100.8783 | 101.1435 |  | 21.1362 |  |  |  | 0.23885 | 8.4005 | 0.003342 | 0.004696 | 0.004479 | 5.6e-05 | 0.000995 | 0.29336 | 1.3379 | 2.5479 | 99.48 | 20.91 | 37.254 |  | 102.1373 |  | 0.009273 |  |  |  | -0.023484 | 0.14209
2019-03-21 12:14 | 1553184833.379 | 1720 |  |  |  | 100.903 | 101.0716 |  | 21.1378 |  |  |  | 0.23961 | 8.4029 | 0.003289 | 0.004865 | 0.004698 | 5.1e-05 | 0.001 | 0.32549 | 1.3532 | 2.548 | 99.47 | 20.91 | 37.253 |  | 100.8529 |  | 0.009273 |  |  |  | -0.014934 | 0.14255
2019-03-21 12:14 | 1553184834.379 | 1721 |  |  |  | 100.8556 | 100.9408 |  | 21.1377 |  |  |  | 0.24034 | 8.4018 | 0.003366 | 0.004326 | 0.004203 | 5.4e-05 | 0.001002 | 0.31741 | 1.325 | 2.5481 | 99.47 | 20.91 | 37.251 |  | 102.717 |  | 0.009273 |  |  |  | -0.007553 | 0.14298
2019-03-21 12:14 | 1553184835.379

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-21 11:45 | 0 | 0 | 0 | 0 | 0 | 52.9 | -0.000845
2019-03-21 11:45 | 1 | 0 | 0 | 0 | 0 | 53.3012 | -0.01175
2019-03-21 11:45 | 2 | 0 | 0 | 0 | 0 | 53.2337 | 0.002132
2019-03-21 11:45 | 3 | 0 | 0 | 0 | 0 | 53.2289 | 0.000571
2019-03-21 11:45 | 4 | 0 | 0 | 0 | 0 | 53.2778 | 0.018479
2019-03-21 11:45 | 5 | 0 | 0 | 0 | 0 | 53.3433 | 0.023306
2019-03-21 11:45 | 6 | 0 | 0 | 0 | 0 | 53.0493 | 0.20379
2019-03-21 11:45 | 7 | 0 | 0 | 0 | 0 | 52.9741 | 0.59054
2019-03-21 11:45 | 8 | 0 | 0 | 0 | 0 | 52.8713 | 1.1759
2019-03-21 11:45 | 9 | 0 | 0 | 0 | 0 | 52.7616 | 1.6911
2019-03-21 11:45 | 10 | 0 | 0 | 0 | 0 | 52.6545 | 2.2274
2019-03-21 11:45 | 11 | 0 | 0 | 0 | 0 | 52.5108 | 2.6499
2019-03-21 11:45 | 12 | 0 | 0 | 0 | 0 | 52.3598 | 2.7508
2019-03-21 11:45 | 13 | 0 | 0 | 0 | 0 | 52.2426 | 2.9481
2019-03-21 11:45 | 14 | 0 | 0 | 0 | 0 | 52.1221 | 3.1041
2019-03-21 11:45 | 15 | 0.12099 | 0 | 0 | 0 | 51.9885 | 3.2235
2019-03-21 11:45 | 16 | 0.11834 | 0 | 0 | 0 | 51.8792 | 3.2005
2019-03-21 11:45 | 17 | 0.11582 | 0 | 0 | 0 | 51.7515 | 3.1945
2019-03-21 11:45 | 18 | 0.11598 | 0 | 0 | 0 | 51.6461 | 3.2993
2019-03-21 11:46 | 19 | 0.14381 | 0 | 0 | 0 | 51.5136 | 3.3233
2019-03-21 11:46 | 20 | 0.13109 | 0 | 0 | 0 | 51.3679 | 3.3975
2019-03-21 11:46 | 21 | 0.11654 | 0 | 0 | 0 | 51.2514 | 3.451
2019-03-21 11:46 | 22 | 0.1585 | 0 | 0 | 0 | 51.1216 | 3.6191
2019-03-21 11:46 | 23 | 0.17128 | 0 | 0 | 0 | 50.9414 | 3.9227
2019-03-21 11:46 | 24 | 0.14046 | 0 | 0 | 0 | 50.79 | 4.137
2019-03-21 11:46 | 25 | 0.13924 | 0 | 0 | 0 | 50.654 | 4.2764
2019-03-21 11:46 | 26 | 0.15035 | 0 | 0 | 0 | 50.5078 | 4.3392
2019-03-21 11:46 | 27 | 0.15517 | 0 | 0 | 0 | 50.3548 | 4.3984
2019-03-21 11:46 | 28 | 0.14779 | 0 | 0 | 0 | 50.2006 | 4.384
2019-03-21 11:46 | 29 | 0.14344 | 0 | 0 | 0 | 50.0586 | 4.3575
2019-03-21 11:46 | 30 | 0.15244 | 0 | 0 | 0 | 49.9111 | 4.2702
2019-03-21 11:46 | 31 | 0.1329 | 0 | 0 | 0 | 49.7604 | 4.1208
2019-03-21 11:46 | 32 | 0.11518 | 0 | 0 | 0 | 49.6449 | 3.966
2019-03-21 11:46 | 33 | 0.12576 | 0 | 0 | 0 | 49.5231 | 3.815
2019-03-21 11:46 | 34 | 0.13082 | 0 | 0 | 0 | 49.395 | 3.5941
2019-03-21 11:46 | 35 | 0.097243 | 0 | 0 | 0 | 49.2704 | 3.3476
2019-03-21 11:46 | 36 | 0.09394 | 0 | 0 | 0 | 49.1931 | 3.302
2019-03-21 11:46 | 37 | 0.12439 | 0 | 0 | 0 | 49.0747 | 3.0169
2019-03-21 11:46 | 38 | 0.10391 | 0 | 0 | 0 | 48.9559 | 2.7366
2019-03-21 11:46 | 39 | 0.088912 | 0 | 0 | 0 | 48.866 | 2.8231
2019-03-21 11:46 | 40 | 0.10487 | 0 | 0 | 0 | 48.7709 | 2.6294
2019-03-21 11:46 | 41 | 0.099927 | 0 | 0 | 0 | 48.6611 | 2.6305
2019-03-21 11:46 | 42 | 0.078213 | 0 | 0 | 0 | 48.5752 | 2.4668
2019-03-21 11:46 | 43 | 0.078266 | 0 | 0 | 0 | 48.4994 | 2.3556
2019-03-21 11:46 | 44 | 0.087275 | 0 | 0 | 0 | 48.4157 | 2.318
2019-03-21 11:46 | 45 | 0.097149 | 0.018332 | 2.5027 | 23.3984 | 48.3249 | 2.2731
2019-03-21 11:46 | 46 | 0.089808 | 0.019144 | 2.6628 | 24.926 | 48.226 | 2.2385
2019-03-21 11:46 | 47 | 0.071969 | 0.023019 | 3.2658 | 30.799 | 48.1475 | 2.2166
2019-03-21 11:46 | 48 | 0.087629 | 0.018172 | 2.6344 | 25.0973 | 48.0743 | 2.1993
2019-03-21 11:46 | 49 | 0.092776 | 0.01671 | 2.4471 | 23.4462 | 47.9747 | 2.1752
2019-03-21 11:46 | 50 | 0.069973 | 0.021576 | 3.1939 | 30.7308 | 47.8952 | 2.1503
2019-03-21 11:46 | 51 | 0.077385 | 0.01925 | 2.8466 | 26.6922 | 47.8283 | 2.0656
2019-03-21 11:46 | 52 | 0.071374 | 0.020248 | 3.0379 | 28.8062 | 47.7429 | 2.056
2019-03-21 11:46 | 53 | 0.066879 | 0.021341 | 3.2021 | 30.3684 | 47.6849 | 2.031
2019-03-21 11:46 | 54 | 0.083004 | 0.01694 | 2.5425 | 24.4496 | 47.6043 | 2.0294
2019-03-21 11:46 | 55 | 0.086291 | 0.016077 | 2.4142 | 23.8045 | 47.5217 | 2.0541
2019-03-21 11:46 | 56 | 0.081588 | 0.016516 | 2.5225 | 24.4245 | 47.4341 | 1.9927
2019-03-21 11:46 | 57 | 0.076916 | 0.017223 | 2.6408 | 25.4373 | 47.3586 | 1.9565
2019-03-21 11:46 | 58 | 0.083841 | 0.015348 | 2.4016 | 22.8862 | 47.2778 | 1.9188
2019-03-21 11:46 | 59 | 0.074131 | 0.016795 | 2.6864 | 24.437 | 47.1949 | 1.8115
2019-03-21 11:46 | 60 | 0.05798 | 0.021485 | 3.3941 | 31.1185 | 47.1308 | 1.8043
2019-03-21 11:46 | 61 | 0.069735 | 0.017558 | 2.7818 | 24.9321 | 47.0719 | 1.7386
2019-03-21 11:46 | 62 | 0.089468 | 0.013482 | 2.1437 | 18.6935 | 46.9895 | 1.6725
2019-03-21 11:46 | 63 | 0.076995 | 0.015393 | 2.4525 | 20.9155 | 46.9012 | 1.6104
2019-03-21 11:46 | 64 | 0.05384 | 0.021688 | 3.4589 | 28.4175 | 46.8386 | 1.53
2019-03-21 11:46 | 65 | 0.050032 | 0.022957 | 3.6666 | 29.3944 | 46.789 | 1.4707
2019-03-21 11:46 | 66 | 0.043831 | 0.025261 | 4.1132 | 30.8582 | 46.7387 | 1.3525
2019-03-21 11:46 | 67 | 0.034068 | 0.030653 | 5.2033 | 37.579 | 46.702 | 1.2803
2019-03-21 11:46 | 68 | 0.029118 | 0.035067 | 5.9525 | 41.2779 | 46.6686 | 1.2019
2019-03-21 11:46 | 69 | 0.051632 | 0.019401 | 3.2912 | 22.0213 | 46.6371 | 1.137
2019-03-21 11:46 | 70 | 0.060249 | 0.01591 | 2.759 | 17.895 | 46.5686 | 1.0782
2019-03-21 11:46 | 71 | 0.049507 | 0.018458 | 3.2762 | 21.2775 | 46.5217 | 1.0534
2019-03-21 11:46 | 72 | 0.049435 | 0.018072 | 3.2041 | 20.9055 | 46.468 | 1.0335
2019-03-21 11:46 | 73 | 0.031864 | 0.027342 | 4.8507 | 32.0984 | 46.4266 | 1.0228
2019-03-21 11:46 | 74 | 0.024013 | 0.034459 | 6.2814 | 40.2123 | 46.4016 | 0.96563
2019-03-21 11:46 | 75 | 0.039795 | 0.020308 | 3.7108 | 24.5304 | 46.3724 | 0.9762
2019-03-21 11:46 | 76 | 0.051586 | 0.015646 | 2.7905 | 18.8488 | 46.3224 | 0.97232
2019-03-21 11:46 | 77 | 0.0591 | 0.013649 | 2.3777 | 16.4822 | 46.2707 | 0.9741
2019-03-21 11:46 | 78 | 0.054806 | 0.014716 | 2.5021 | 17.4118 | 46.2073 | 0.95426
2019-03-21 11:46 | 79 | 0.053071 | 0.014792 | 2.5251 | 17.4823 | 46.1608 | 0.9278
2019-03-21 11:47 | 80 | 0.056302 | 0.013521 | 2.3253 | 16.4074 | 46.1007 | 0.92376
2019-03-21 11:47 | 81 | 0.032163 | 0.023733 | 4.0003 | 28.4244 | 46.0546 | 0.91421
2019-03-21 11:47 | 82 | 0.020852 | 0.036585 | 6.05 | 43.7765 | 46.033 | 0.91285
2019-03-21 11:47 | 83 | 0.038286 | 0.01935 | 3.2305 | 23.7076 | 46.0052 | 0.90767
2019-03-21 11:47 | 84 | 0.057389 | 0.012896 | 2.1125 | 15.8054 | 45.9558 | 0.90705
2019-03-21 11:47 | 85 | 0.055117 | 0.013414 | 2.1584 | 16.3848 | 45.8962 | 0.90308
2019-03-21 11:47 | 86 | 0.038823 | 0.019685 | 3.0251 | 24.0306 | 45.8499 | 0.93294
2019-03-21 11:47 | 87 | 0.016446 | 0.047705 | 7.006 | 56.6378 | 45.8199 | 0.93145
2019-03-21 11:47 | 88 | 0.018867 | 0.042849 | 6.0246 | 48.5293 | 45.8107 | 0.91558
2019-03-21 11:47 | 89 | 0.025758 | 0.031323 | 4.3411 | 35.8933 | 45.7801 | 0.92453
2019-03-21 11:47 | 90 | 0.033403 | 0.024857 | 3.3096 | 28.057 | 45.7586 | 0.93719
2019-03-21 11:47 | 91 | 0.046953 | 0.018149 | 2.3259 | 19.6195 | 45.7119 | 0.9212
2019-03-21 11:47 | 92 | 0.05676 | 0.015421 | 1.9051 | 16.1275 | 45.6654 | 0.91541
2019-03-21 11:47 | 93 | 0.07386 | 0.012485 | 1.4508 | 12.7491 | 45.5979 | 0.94165
2019-03-21 11:47 | 94 | 0.048139 | 0.019179 | 2.2045 | 19.2475 | 45.528 | 0.92655
2019-03-21 11:47 | 95 | 0.033227 | 0.027738 | 3.1575 | 28.3551 | 45.4989 | 0.94214
2019-03-21 11:47 | 96 | 0.05976 | 0.015425 | 1.7435 | 15.7474 | 45.4512 | 0.94106
2019-03-21 11:47 | 97 | 0.072388 | 0.012733 | 1.4294 | 12.8868 | 45.3825 | 0.93285
2019-03-21 11:47 | 98 | 0.057476 | 0.015659 | 1.7877 | 16.3987 | 45.3137 | 0.94254
2019-03-21 11:47 | 99 | 0.044381 | 0.019862 | 2.3049 | 21.4407 | 45.2677 | 0.95157
2019-03-21 11:47 | 100 | 0.04204 | 0.021509 | 2.4226 | 22.3964 | 45.2224 | 0.94154
2019-03-21 11:47 | 101 | 0.033059 | 0.026652 | 3.0605 | 28.8653 | 45.1847 | 0.95425
2019-03-21 11:47 | 102 | 0.030706 | 0.028675 | 3.2756 | 30.722 | 45.1546 | 0.94334
2019-03-21 11:47 | 103 | 0.034116 | 0.025101 | 2.9282 | 28.0505 | 45.1214 | 0.95698
2019-03-21 11:47 | 104 | 0.046738 | 0.018804 | 2.1283 | 20.2068 | 45.0844 | 0.94442
2019-03-21 11:47 | 105 | 0.039229 | 0.022457 | 2.5245 | 24.0977 | 45.0331 | 0.94532
2019-03-21 11:47 | 106 | 0.020685 | 0.043687 | 4.7867 | 46.0539 | 45.0085 | 0.95262
2019-03-21 11:47 | 107 | 0.03061 | 0.028788 | 3.2166 | 31.169 | 44.985 | 0.95408
2019-03-21 11:47 | 108 | 0.034474 | 0.025623 | 2.8557 | 28.1064 | 44.9478 | 0.96893
2019-03-21 11:47 | 109 | 0.045358 | 0.019452 | 2.1593 | 20.5726 | 44.9143 | 0.93313
2019-03-21 11:47 | 110 | 0.05692 | 0.015065 | 1.7101 | 16.5532 | 44.857 | 0.94221
2019-03-21 11:47 | 111 | 0.059826 | 0.013984 | 1.6226 | 15.5692 | 44.8032 | 0.93144
2019-03-21 11:47 | 112 | 0.044805 | 0.018678 | 2.1657 | 21.1037 | 44.7419 | 0.94556
2019-03-21 11:47 | 113 | 0.038807 | 0.021558 | 2.4886 | 23.5901 | 44.7114 | 0.91546
2019-03-21 11:47 | 114 | 0.048964 | 0.01708 | 1.9622 | 18.3805 | 44.6605 | 0.89999
2019-03-21 11:47 | 115 | 0.039829 | 0.020454 | 2.4089 | 22.0668 | 44.6175 | 0.8789
2019-03-21 11:47 | 116 | 0.042669 | 0.019108 | 2.2437 | 20.94 | 44.5781 | 0.89349
2019-03-21 11:47 | 117 | 0.051213 | 0.015468 | 1.8591 | 17.0554 | 44.5305 | 0.87346
2019-03-21 11:47 | 118 | 0.060164 | 0.0128 | 1.579 | 14.7231 | 44.4764 | 0.8858
2019-03-21 11:47 | 119 | 0.036427 | 0.021073 | 2.5839 | 23.6231 | 44.418 | 0.86051
2019-03-21 11:47 | 120 | 0.026821 | 0.027794 | 3.4981 | 32.1333 | 44.4002 | 0.86186
2019-03-21 11:47 | 121 | 0.043499 | 0.016658 | 2.1474 | 19.9942 | 44.3579 | 0.86972
2019-03-21 11:47 | 122 | 0.038858 | 0.018041 | 2.3888 | 22.1439 | 44.3178 | 0.86047
2019-03-21 11:47 | 123 | 0.033277 | 0.020396 | 2.7735 | 25.3425 | 44.2812 | 0.84333
2019-03-21 11:47 | 124 | 0.0194 | 0.035083 | 4.7504 | 43.9323 | 44.2528 | 0.85229
2019-03-21 11:47 | 125 | 0.02543 | 0.025858 | 3.6024 | 33.0875 | 44.2371 | 0.84141
2019-03-21 11:47 | 126 | 0.047142 | 0.013906 | 1.9321 | 17.8365 | 44.1979 | 0.84085
2019-03-21 11:47 | 127 | 0.051259 | 0.013248 | 1.7674 | 16.4117 | 44.1475 | 0.84125
2019-03-21 11:47 | 128 | 0.031314 | 0.021695 | 2.8785 | 26.4563 | 44.1017 | 0.82845
2019-03-21 11:47 | 129 | 0.024916 | 0.028141 | 3.6033 | 32.5203 | 44.0819 | 0.81026
2019-03-21 11:47 | 130 | 0.029818 | 0.023549 | 2.9963 | 27.6213 | 44.0485 | 0.82361
2019-03-21 11:47 | 131 | 0.035988 | 0.019554 | 2.4751 | 23.1786 | 44.0216 | 0.83416
2019-03-21 11:47 | 132 | 0.039826 | 0.018289 | 2.2318 | 20.7097 | 43.9774 | 0.82478
2019-03-21 11:47 | 133 | 0.031022 | 0.023441 | 2.8537 | 25.9056 | 43.9444 | 0.80364
2019-03-21 11:47 | 134 | 0.051762 | 0.014035 | 1.6997 | 15.4675 | 43.9087 | 0.80064
2019-03-21 11:47 | 135 | 0.05925 | 0.011863 | 1.4775 | 13.6577 | 43.8442 | 0.80922
2019-03-21 11:47 | 136 | 0.041462 | 0.016957 | 2.1016 | 19.3994 | 43.7967 | 0.80433
2019-03-21 11:47 | 137 | 0.028561 | 0.023882 | 3.0448 | 27.7481 | 43.7599 | 0.79251
2019-03-21 11:47 | 138 | 0.045265 | 0.015588 | 1.9193 | 17.1057 | 43.7321 | 0.77429
2019-03-21 11:48 | 139 | 0.057407 | 0.012288 | 1.5036 | 13.3997 | 43.6702 | 0.76924
2019-03-21 11:48 | 140 | 0.052102 | 0.013959 | 1.6486 | 14.6661 | 43.6218 | 0.76413
2019-03-21 11:48 | 141 | 0.047359 | 0.015394 | 1.8075 | 16.1375 | 43.5664 | 0.76425
2019-03-21 11:48 | 142 | 0.042851 | 0.016432 | 1.9875 | 17.7056 | 43.5269 | 0.7587
2019-03-21 11:48 | 143 | 0.045666 | 0.015434 | 1.8557 | 16.1759 | 43.4788 | 0.73869
2019-03-21 11:48 | 144 | 0.047898 | 0.014267 | 1.7629 | 15.416 | 43.4359 | 0.73839
2019-03-21 11:48 | 145 | 0.038012 | 0.018035 | 2.22 | 19.5281 | 43.3866 | 0.7423
2019-03-21 11:48 | 146 | 0.0090321 | 0.073287 | 9.2847 | 81.9357 | 43.3638 | 0.74005
2019-03-21 11:48 | 147 | 0.023115 | 0.028646 | 3.6105 | 31.6421 | 43.358 | 0.73142
2019-03-21 11:48 | 148 | 0.045468 | 0.014498 | 1.8194 | 16.0418 | 43.3152 | 0.72939
2019-03-21 11:48 | 149 | 0.032661 | 0.01951 | 2.5271 | 22.3817 | 43.2754 | 0.73101
2019-03-21 11:48 | 150 | 0.025313 | 0.025182 | 3.2406 | 28.0204 | 43.2491 | 0.70929
2019-03-21 11:48 | 151 | 0.031024 | 0.020567 | 2.6328 | 23.1177 | 43.2212 | 0.7172
2019-03-21 11:48 | 152 | 0.046369 | 0.014783 | 1.7577 | 15.3924 | 43.185 | 0.71372
2019-03-21 11:48 | 153 | 0.035988 | 0.019042 | 2.2562 | 19.6557 | 43.135 | 0.70737
2019-03-21 11:48 | 154 | 0.013277 | 0.053416 | 6.0977 | 54.0304 | 43.1157 | 0.71733
2019-03-21 11:48 | 155 | 0.031407 | 0.022518 | 2.5546 | 22.2763 | 43.0983 | 0.69963
2019-03-21 11:48 | 156 | 0.057182 | 0.012779 | 1.4001 | 12.3162 | 43.0509 | 0.70427
2019-03-21 11:48 | 157 | 0.047282 | 0.014959 | 1.688 | 14.8275 | 42.9933 | 0.70108
2019-03-21 11:48 | 158 | 0.015533 | 0.047068 | 5.1174 | 45.1062 | 42.9617 | 0.70065
2019-03-21 11:48 | 159 | 0.024459 | 0.028964 | 3.2315 | 28.1639 | 42.9526 | 0.68886
2019-03-21 11:48 | 160 | 0.031764 | 0.022993 | 2.4784 | 21.3314 | 42.9123 | 0.67756
2019-03-21 11:48 | 161 | 0.030003 | 0.023531 | 2.6039 | 22.8245 | 42.8916 | 0.6848
2019-03-21 11:48 | 162 | 0.020889 | 0.033779 | 3.7311 | 32.3986 | 42.8547 | 0.67678
2019-03-21 11:48 | 163 | 0.002647 | 0.2582 | 29.3351 | 255.7798 | 42.8513 | 0.67705
2019-03-21 11:48 | 164 | 0.021253 | 0.032122 | 3.6326 | 32.0032 | 42.8404 | 0.68015
2019-03-21 11:48 | 165 | 0.036553 | 0.018703 | 2.1073 | 18.238 | 42.8091 | 0.66666
2019-03-21 11:48 | 166 | 0.030558 | 0.020918 | 2.5123 | 22.0702 | 42.7725 | 0.67442
2019-03-21 11:48 | 167 | 0.029947 | 0.021267 | 2.5459 | 21.5121 | 42.747 | 0.64422
2019-03-21 11:48 | 168 | 0.041706 | 0.01417 | 1.8165 | 15.6709 | 42.7096 | 0.65358
2019-03-21 11:48 | 169 | 0.043649 | 0.013568 | 1.7344 | 14.8636 | 42.6656 | 0.64878
2019-03-21 11:48 | 170 | 0.051372 | 0.011519 | 1.469 | 12.4831 | 42.6218 | 0.64129
2019-03-21 11:48 | 171 | 0.037043 | 0.015994 | 2.0262 | 17.5266 | 42.5677 | 0.64925
2019-03-21 11:48 | 172 | 0.041309 | 0.014908 | 1.809 | 15.4879 | 42.5432 | 0.63979
2019-03-21 11:48 | 173 | 0.061337 | 0.010057 | 1.2151 | 10.68 | 42.4814 | 0.65508
2019-03-21 11:48 | 174 | 0.052856 | 0.012122 | 1.4089 | 12.1744 | 42.4271 | 0.6435
2019-03-21 11:48 | 175 | 0.048512 | 0.013696 | 1.5288 | 13.3423 | 42.3754 | 0.64726
2019-03-21 11:48 | 176 | 0.047932 | 0.013414 | 1.5463 | 13.2548 | 42.3296 | 0.63533
2019-03-21 11:48 | 177 | 0.02858 | 0.021716 | 2.5872 | 22.6308 | 42.2837 | 0.64678
2019-03-21 11:48 | 178 | 0.027192 | 0.021953 | 2.7067 | 23.4309 | 42.2683 | 0.63714
2019-03-21 11:48 | 179 | 0.032391 | 0.018472 | 2.2753 | 20.0013 | 42.2275 | 0.64786
2019-03-21 11:48 | 180 | 0.021082 | 0.027267 | 3.4827 | 29.9826 | 42.207 | 0.63209
2019-03-21 11:48 | 181 | 0.027515 | 0.021755 | 2.6586 | 22.4584 | 42.1818 | 0.61794
2019-03-21 11:48 | 182 | 0.015463 | 0.038682 | 4.7179 | 40.2324 | 42.1556 | 0.62213
2019-03-21 11:48 | 183 | 0.022799 | 0.027243 | 3.1902 | 26.8615 | 42.1465 | 0.61241
2019-03-21 11:48 | 184 | 0.031891 | 0.020877 | 2.2625 | 19.3773 | 42.1096 | 0.61796
2019-03-21 11:48 | 185 | 0.018575 | 0.035857 | 3.8782 | 33.4469 | 42.0879 | 0.62126
2019-03-21 11:48 | 186 | 0.017493 | 0.039315 | 4.0981 | 35.0732 | 42.0695 | 0.61354
2019-03-21 11:48 | 187 | 0.032923 | 0.020925 | 2.1746 | 18.9896 | 42.0484 | 0.6252
2019-03-21 11:48 | 188 | 0.04698 | 0.015139 | 1.5179 | 13.1863 | 42.004 | 0.6195
2019-03-21 11:48 | 189 | 0.044803 | 0.015342 | 1.5864 | 13.7773 | 41.9587 | 0.61727
2019-03-21 11:48 | 190 | 0.03935 | 0.017446 | 1.8012 | 15.5091 | 41.9153 | 0.61029
2019-03-21 11:48 | 191 | 0.045332 | 0.015622 | 1.5553 | 13.4632 | 41.877 | 0.61032
2019-03-21 11:48 | 192 | 0.059978 | 0.01182 | 1.1762 | 10.2383 | 41.823 | 0.61407
2019-03-21 11:48 | 193 | 0.047699 | 0.014384 | 1.4775 | 12.9182 | 41.7636 | 0.61618
2019-03-21 11:48 | 194 | 0.03698 | 0.018537 | 1.8993 | 16.5068 | 41.7275 | 0.61042
2019-03-21 11:48 | 195 | 0.043977 | 0.016101 | 1.5931 | 14.0237 | 41.6855 | 0.61673
2019-03-21 11:48 | 196 | 0.040075 | 0.017085 | 1.7433 | 15.3726 | 41.642 | 0.61606
2019-03-21 11:48 | 197 | 0.027411 | 0.024952 | 2.5382 | 22.2012 | 41.6077 | 0.60855
2019-03-21 11:48 | 198 | 0.016757 | 0.04218 | 4.1544 | 36.1869 | 41.5863 | 0.60638
2019-03-21 11:49 | 199 | 0.028727 | 0.024655 | 2.4183 | 21.4874 | 41.5685 | 0.61726
2019-03-21 11:49 | 200 | 0.036871 | 0.019234 | 1.8915 | 16.6065 | 41.5294 | 0.6123
2019-03-21 11:49 | 201 | 0.033705 | 0.021044 | 2.0631 | 18.0642 | 41.4972 | 0.60886
2019-03-21 11:49 | 202 | 0.048677 | 0.014109 | 1.4248 | 12.5113 | 41.4582 | 0.60901
2019-03-21 11:49 | 203 | 0.050298 | 0.013183 | 1.378 | 11.9119 | 41.4033 | 0.59915
2019-03-21 11:49 | 204 | 0.030493 | 0.021728 | 2.2641 | 18.951 | 41.3626 | 0.57788
2019-03-21 11:49 | 205 | 0.035934 | 0.01784 | 1.9163 | 15.8909 | 41.3362 | 0.57102
2019-03-21 11:49 | 206 | 0.053569 | 0.01197 | 1.2851 | 10.7442 | 41.2873 | 0.57556
2019-03-21 11:49 | 207 | 0.060488 | 0.01062 | 1.1348 | 9.4728 | 41.2318 | 0.57299
2019-03-21 11:49 | 208 | 0.054708 | 0.011325 | 1.2544 | 10.4271 | 41.1697 | 0.57045
2019-03-21 11:49 | 209 | 0.051854 | 0.01194 | 1.3169 | 10.9757 | 41.122 | 0.56913
2019-03-21 11:49 | 210 | 0.053042 | 0.012106 | 1.2814 | 10.6911 | 41.0654 | 0.56707
2019-03-21 11:49 | 211 | 0.038986 | 0.016494 | 1.7403 | 14.3522 | 41.0196 | 0.55954
2019-03-21 11:49 | 212 | 0.028807 | 0.022287 | 2.3402 | 19.2239 | 40.9868 | 0.55377
2019-03-21 11:49 | 213 | 0.02071 | 0.031048 | 3.2497 | 26.6579 | 40.9608 | 0.55209
2019-03-21 11:49 | 214 | 0.028087 | 0.022909 | 2.3927 | 19.1983 | 40.9413 | 0.53922
2019-03-21 11:49 | 215 | 0.037518 | 0.017181 | 1.7916 | 14.702 | 40.9039 | 0.55159
2019-03-21 11:49 | 216 | 0.039478 | 0.016297 | 1.6902 | 14.2141 | 40.8684 | 0.56114
2019-03-21 11:49 | 217 | 0.027551 | 0.023347 | 2.4097 | 19.7646 | 40.8282 | 0.54454
2019-03-21 11:49 | 218 | 0.031604 | 0.020366 | 2.0986 | 16.936 | 40.8095 | 0.53524
2019-03-21 11:49 | 219 | 0.045255 | 0.014731 | 1.4592 | 11.7364 | 40.763 | 0.53113
2019-03-21 11:49 | 220 | 0.028324 | 0.024321 | 2.3161 | 18.2529 | 40.726 | 0.517
2019-03-21 11:49 | 221 | 0.024484 | 0.028168 | 2.6839 | 20.2557 | 40.7035 | 0.49595
2019-03-21 11:49 | 222 | 0.034736 | 0.02053 | 1.8774 | 15.2105 | 40.6733 | 0.52835
2019-03-21 11:49 | 223 | 0.041176 | 0.01731 | 1.5785 | 12.4895 | 40.6348 | 0.51427
2019-03-21 11:49 | 224 | 0.037341 | 0.019136 | 1.7421 | 13.4016 | 40.5936 | 0.50043
2019-03-21 11:49 | 225 | 0.033999 | 0.021656 | 1.9068 | 14.8867 | 40.5602 | 0.50614
2019-03-21 11:49 | 226 | 0.035454 | 0.020834 | 1.8252 | 14.3927 | 40.5241 | 0.51028
2019-03-21 11:49 | 227 | 0.047758 | 0.015469 | 1.3476 | 10.5274 | 40.487 | 0.50277
2019-03-21 11:49 | 228 | 0.040799 | 0.018681 | 1.5745 | 12.8833 | 40.4335 | 0.52562
2019-03-21 11:49 | 229 | 0.023116 | 0.031992 | 2.7713 | 23.1245 | 40.408 | 0.53455
2019-03-21 11:49 | 230 | 0.029171 | 0.026127 | 2.1834 | 17.3883 | 40.3818 | 0.50723
2019-03-21 11:49 | 231 | 0.025497 | 0.029874 | 2.4881 | 20.1114 | 40.3511 | 0.51278
2019-03-21 11:49 | 232 | 0.027919 | 0.027278 | 2.2692 | 18.1077 | 40.3286 | 0.50555
2019-03-21 11:49 | 233 | 0.063197 | 0.012046 | 0.99813 | 8.0563 | 40.288 | 0.50914
2019-03-21 11:49 | 234 | 0.059375 | 0.012822 | 1.0569 | 8.5392 | 40.2117 | 0.50702
2019-03-21 11:49 | 235 | 0.035862 | 0.021869 | 1.7476 | 14.0977 | 40.1749 | 0.50557
2019-03-21 11:49 | 236 | 0.027951 | 0.02887 | 2.2321 | 17.5683 | 40.1364 | 0.49104
2019-03-21 11:49 | 237 | 0.030997 | 0.026051 | 2.0036 | 15.9966 | 40.1154 | 0.49584
2019-03-21 11:49 | 238 | 0.048099 | 0.017215 | 1.2812 | 10.3491 | 40.0712 | 0.49778
2019-03-21 11:49 | 239 | 0.037893 | 0.021256 | 1.6262 | 13.3432 | 40.0253 | 0.50562
2019-03-21 11:49 | 240 | 0.038105 | 0.02172 | 1.6118 | 13.152 | 39.9932 | 0.50116
2019-03-21 11:49 | 241 | 0.053815 | 0.015411 | 1.1369 | 9.0707 | 39.946 | 0.48814
2019-03-21 11:49 | 242 | 0.034008 | 0.024415 | 1.7979 | 14.992 | 39.8939 | 0.50984
2019-03-21 11:49 | 243 | 0.017287 | 0.047982 | 3.5312 | 28.7411 | 39.8773 | 0.49685
2019-03-21 11:49 | 244 | 0.032373 | 0.024927 | 1.8824 | 15.0471 | 39.8514 | 0.48712
2019-03-21 11:49 | 245 | 0.042096 | 0.019121 | 1.4362 | 12.0979 | 39.8131 | 0.50927
2019-03-21 11:49 | 246 | 0.050236 | 0.016042 | 1.2007 | 9.8599 | 39.7686 | 0.49532
2019-03-21 11:49 | 247 | 0.028393 | 0.027513 | 2.1134 | 17.1245 | 39.72 | 0.48622
2019-03-21 11:49 | 248 | 0.013287 | 0.058891 | 4.5101 | 36.0318 | 39.7099 | 0.47876
2019-03-21 11:49 | 249 | 0.035358 | 0.022129 | 1.6875 | 13.8525 | 39.6844 | 0.4898
2019-03-21 11:49 | 250 | 0.039401 | 0.019932 | 1.5191 | 12.21 | 39.6431 | 0.48109
2019-03-21 11:49 | 251 | 0.023338 | 0.033586 | 2.5519 | 20.3139 | 39.6101 | 0.47408
2019-03-21 11:49 | 252 | 0.044139 | 0.017776 | 1.3466 | 10.8078 | 39.5882 | 0.47705
2019-03-21 11:49 | 253 | 0.049071 | 0.015497 | 1.2036 | 9.8102 | 39.5253 | 0.4814
2019-03-21 11:49 | 254 | 0.038572 | 0.019142 | 1.5275 | 12.8047 | 39.4944 | 0.49391
2019-03-21 11:49 | 255 | 0.029218 | 0.024509 | 2.0176 | 16.1567 | 39.4485 | 0.47206
2019-03-21 11:49 | 256 | 0.012189 | 0.05881 | 4.8295 | 39.7665 | 39.4368 | 0.4847
2019-03-21 11:49 | 257 | 0.031592 | 0.022713 | 1.8583 | 14.7721 | 39.4151 | 0.46668
2019-03-21 11:49 | 258 | 0.049227 | 0.014579 | 1.1891 | 9.7562 | 39.3736 | 0.48027
2019-03-21 11:50 | 259 | 0.048367 | 0.015288 | 1.211 | 9.4618 | 39.3216 | 0.45764
2019-03-21 11:50 | 260 | 0.037325 | 0.020421 | 1.5608 | 12.7887 | 39.2794 | 0.47734
2019-03-21 11:50 | 261 | 0.047677 | 0.016492 | 1.2211 | 10.06 | 39.2424 | 0.47963
2019-03-21 11:50 | 262 | 0.044425 | 0.017661 | 1.3068 | 11.2258 | 39.1873 | 0.4987
2019-03-21 11:50 | 263 | 0.028478 | 0.027587 | 2.026 | 17.5685 | 39.1565 | 0.50031
2019-03-21 11:50 | 264 | 0.03265 | 0.024797 | 1.7769 | 14.9421 | 39.1257 | 0.48785
2019-03-21 11:50 | 265 | 0.03133 | 0.025798 | 1.8498 | 15.5728 | 39.0917 | 0.4879
2019-03-21 11:50 | 266 | 0.039954 | 0.020262 | 1.4474 | 12.0268 | 39.0608 | 0.48052
2019-03-21 11:50 | 267 | 0.047147 | 0.01666 | 1.2274 | 9.9633 | 39.013 | 0.46974
2019-03-21 11:50 | 268 | 0.017897 | 0.04393 | 3.2385 | 26.9123 | 38.9751 | 0.48164
2019-03-21 11:50 | 269 | 0.0079524 | 0.098807 | 7.2793 | 62.9951 | 38.9721 | 0.50096
2019-03-21 11:50 | 270 | 0.041898 | 0.018185 | 1.3742 | 11.5533 | 38.9483 | 0.48407
2019-03-21 11:50 | 271 | 0.050464 | 0.015545 | 1.138 | 9.7985 | 38.8942 | 0.49447
2019-03-21 11:50 | 272 | 0.030338 | 0.025853 | 1.8927 | 15.8504 | 38.8549 | 0.48088
2019-03-21 11:50 | 273 | 0.028386 | 0.028447 | 2.0205 | 17.4352 | 38.83 | 0.49492
2019-03-21 11:50 | 274 | 0.021691 | 0.036237 | 2.6429 | 22.122 | 38.7986 | 0.47984
2019-03-21 11:50 | 275 | 0.022363 | 0.036054 | 2.5495 | 21.3516 | 38.7843 | 0.47749
2019-03-21 11:50 | 276 | 0.039833 | 0.020291 | 1.4349 | 11.2664 | 38.7501 | 0.44877
2019-03-21 11:50 | 277 | 0.03218 | 0.024458 | 1.777 | 14.0731 | 38.7104 | 0.45288
2019-03-21 11:50 | 278 | 0.028863 | 0.027286 | 1.9863 | 16.4606 | 38.685 | 0.4751
2019-03-21 11:50 | 279 | 0.038151 | 0.020648 | 1.4993 | 12.2945 | 38.6497 | 0.46905
2019-03-21 11:50 | 280 | 0.035964 | 0.021216 | 1.5801 | 12.9205 | 38.6114 | 0.46467
2019-03-21 11:50 | 281 | 0.027615 | 0.027691 | 2.0636 | 16.8773 | 38.5797 | 0.46607
2019-03-21 11:50 | 282 | 0.018845 | 0.040448 | 3.0143 | 25.1197 | 38.556 | 0.47337
2019-03-21 11:50 | 283 | 0.028148 | 0.027191 | 2.0251 | 16.8126 | 38.5374 | 0.47325
2019-03-21 11:50 | 284 | 0.038167 | 0.020006 | 1.4881 | 12.3233 | 38.4991 | 0.47034
2019-03-21 11:50 | 285 | 0.042556 | 0.017976 | 1.3345 | 11.2708 | 38.463 | 0.47964
2019-03-21 11:50 | 286 | 0.028797 | 0.027298 | 1.9594 | 16.2809 | 38.4187 | 0.46884
2019-03-21 11:50 | 287 | 0.016932 | 0.046439 | 3.3333 | 27.894 | 38.4045 | 0.4723
2019-03-21 11:50 | 288 | 0.040021 | 0.020217 | 1.4051 | 11.7444 | 38.3769 | 0.47003
2019-03-21 11:50 | 289 | 0.033884 | 0.02387 | 1.658 | 13.5785 | 38.3316 | 0.4601
2019-03-21 11:50 | 290 | 0.00048375 | 1.6282 | 116.2684 | 964.5705 | 38.3157 | 0.46661
2019-03-21 11:50 | 291 | 0.0038146 | 0.20602 | 14.7115 | 120.0742 | 38.3213 | 0.45803
2019-03-21 11:50 | 292 | 0.035039 | 0.022449 | 1.6041 | 13.1115 | 38.3008 | 0.45942
2019-03-21 11:50 | 293 | 0.035081 | 0.022392 | 1.5938 | 13.0401 | 38.2585 | 0.45746
2019-03-21 11:50 | 294 | 0.02834 | 0.02688 | 1.9712 | 16.6261 | 38.2335 | 0.47118
2019-03-21 11:50 | 295 | 0.014045 | 0.055907 | 3.9637 | 33.9654 | 38.2038 | 0.47703
2019-03-21 11:50 | 296 | 0.0072455 | 0.10857 | 7.7028 | 64.2357 | 38.2031 | 0.46542
2019-03-21 11:50 | 297 | 0.020597 | 0.038184 | 2.702 | 22.7037 | 38.184 | 0.46762
2019-03-21 11:50 | 298 | 0.031474 | 0.025004 | 1.767 | 14.7403 | 38.1622 | 0.46394
2019-03-21 11:50 | 299 | 0.033306 | 0.023606 | 1.6714 | 13.2374 | 38.1236 | 0.44089
2019-03-21 11:50 | 300 | 0.027365 | 0.028722 | 2.0284 | 17.4496 | 38.0973 | 0.4775
2019-03-21 11:50 | 301 | 0.039895 | 0.01914 | 1.3945 | 11.781 | 38.0648 | 0.47
2019-03-21 11:50 | 302 | 0.040323 | 0.01893 | 1.3783 | 11.1804 | 38.0203 | 0.45082
2019-03-21 11:50 | 303 | 0.034138 | 0.021722 | 1.6331 | 13.5351 | 37.9859 | 0.46206
2019-03-21 11:50 | 304 | 0.034837 | 0.021242 | 1.5908 | 13.0313 | 37.9501 | 0.45397
2019-03-21 11:50 | 305 | 0.050166 | 0.01431 | 1.1026 | 9.3016 | 37.9128 | 0.46662
2019-03-21 11:50 | 306 | 0.05113 | 0.013991 | 1.0801 | 8.886 | 37.8533 | 0.45434
2019-03-21 11:50 | 307 | 0.039242 | 0.018286 | 1.4062 | 11.7009 | 37.814 | 0.45916
2019-03-21 11:50 | 308 | 0.034889 | 0.019893 | 1.5808 | 13.1135 | 37.7733 | 0.45752
2019-03-21 11:50 | 309 | 0.030586 | 0.022642 | 1.7992 | 15.1647 | 37.7436 | 0.46383
2019-03-21 11:50 | 310 | 0.041002 | 0.016359 | 1.3403 | 11.2804 | 37.7092 | 0.46252
2019-03-21 11:50 | 311 | 0.020993 | 0.031996 | 2.6249 | 21.3933 | 37.6684 | 0.44911
2019-03-21 11:50 | 312 | 0.017971 | 0.037394 | 3.0617 | 25.6743 | 37.6635 | 0.46139
2019-03-21 11:50 | 313 | 0.026607 | 0.024354 | 2.0611 | 16.9936 | 37.6295 | 0.45215
2019-03-21 11:50 | 314 | 0.019003 | 0.034121 | 2.8858 | 24.6129 | 37.6139 | 0.46772
2019-03-21 11:50 | 315 | 0.012995 | 0.049891 | 4.2168 | 34.7795 | 37.5912 | 0.45197
2019-03-21 11:50 | 316 | 0.019216 | 0.033803 | 2.8571 | 23.9532 | 37.5845 | 0.4603
2019-03-21 11:50 | 317 | 0.036276 | 0.017878 | 1.513 | 12.2221 | 37.5504 | 0.44337
2019-03-21 11:50 | 318 | 0.024304 | 0.026631 | 2.2434 | 18.4757 | 37.5185 | 0.44903
2019-03-21 11:50 | 319 | 0.029108 | 0.022264 | 1.868 | 15.6626 | 37.4984 | 0.45591
2019-03-21 11:51 | 320 | 0.034509 | 0.018689 | 1.5702 | 13.1604 | 37.4599 | 0.45415
2019-03-21 11:51 | 321 | 0.035487 | 0.018205 | 1.5315 | 12.5921 | 37.4305 | 0.44685
2019-03-21 11:51 | 322 | 0.034231 | 0.018872 | 1.5813 | 13.2788 | 37.3897 | 0.45455
2019-03-21 11:51 | 323 | 0.030629 | 0.021092 | 1.7697 | 14.8162 | 37.3625 | 0.4538
2019-03-21 11:51 | 324 | 0.032538 | 0.019906 | 1.6713 | 13.8234 | 37.3269 | 0.44978
2019-03-21 11:51 | 325 | 0.040387 | 0.016028 | 1.343 | 11.2774 | 37.296 | 0.45546
2019-03-21 11:51 | 326 | 0.042228 | 0.015324 | 1.2866 | 10.9286 | 37.248 | 0.46149
2019-03-21 11:51 | 327 | 0.026072 | 0.024827 | 2.0775 | 17.2165 | 37.2162 | 0.44888
2019-03-21 11:51 | 328 | 0.015104 | 0.042928 | 3.5946 | 30.6187 | 37.195 | 0.46247
2019-03-21 11:51 | 329 | 0.0063797 | 0.10156 | 8.4982 | 70.4758 | 37.1846 | 0.44961
2019-03-21 11:51 | 330 | 0.019571 | 0.033112 | 2.7652 | 23.0944 | 37.1766 | 0.45199
2019-03-21 11:51 | 331 | 0.033702 | 0.018564 | 1.612 | 13.7903 | 37.1454 | 0.46476
2019-03-21 11:51 | 332 | 0.020686 | 0.030158 | 2.6136 | 21.465 | 37.1161 | 0.44402
2019-03-21 11:51 | 333 | 0.0018302 | 0.34041 | 29.5008 | 246.3825 | 37.1054 | 0.45092
2019-03-21 11:51 | 334 | 0.013157 | 0.049141 | 4.101 | 34.5104 | 37.105 | 0.45406
2019-03-21 11:51 | 335 | 0.027636 | 0.023417 | 1.9529 | 16.2144 | 37.078 | 0.4481
2019-03-21 11:51 | 336 | 0.022945 | 0.028176 | 2.3435 | 19.9573 | 37.0544 | 0.45792
2019-03-21 11:51 | 337 | 0.014628 | 0.045685 | 3.6737 | 30.8726 | 37.0328 | 0.45159
2019-03-21 11:51 | 338 | 0.030215 | 0.022135 | 1.7799 | 15.4866 | 37.0193 | 0.46793
2019-03-21 11:51 | 339 | 0.046212 | 0.0145 | 1.1652 | 9.8854 | 36.9727 | 0.45683
2019-03-21 11:51 | 340 | 0.023793 | 0.028172 | 2.2669 | 18.7511 | 36.9362 | 0.44614
2019-03-21 11:51 | 341 | 0.017298 | 0.038787 | 3.1148 | 25.9351 | 36.9219 | 0.44864
2019-03-21 11:51 | 342 | 0.021202 | 0.031645 | 2.5361 | 21.2871 | 36.8982 | 0.45133
2019-03-21 11:51 | 343 | 0.03655 | 0.019009 | 1.4717 | 11.8595 | 36.8767 | 0.43346
2019-03-21 11:51 | 344 | 0.036523 | 0.018388 | 1.4687 | 12.1828 | 36.829 | 0.44496
2019-03-21 11:51 | 345 | 0.024507 | 0.027308 | 2.1827 | 17.7808 | 36.807 | 0.43576
2019-03-21 11:51 | 346 | 0.01887 | 0.035557 | 2.8535 | 22.7293 | 36.7785 | 0.4289
2019-03-21 11:51 | 347 | 0.021598 | 0.031056 | 2.4788 | 19.444 | 36.7666 | 0.41996
2019-03-21 11:51 | 348 | 0.037596 | 0.017867 | 1.4212 | 11.1999 | 36.7325 | 0.42107
2019-03-21 11:51 | 349 | 0.023378 | 0.027775 | 2.2899 | 19.0374 | 36.6984 | 0.44505
2019-03-21 11:51 | 350 | 0.015783 | 0.041054 | 3.38 | 27.2589 | 36.6844 | 0.43023
2019-03-21 11:51 | 351 | 0.027131 | 0.024717 | 1.9595 | 15.4336 | 36.6619 | 0.41873
2019-03-21 11:51 | 352 | 0.038129 | 0.017584 | 1.3901 | 10.8516 | 36.63 | 0.41377
2019-03-21 11:51 | 353 | 0.037247 | 0.018004 | 1.4263 | 11.0175 | 36.5887 | 0.41037
2019-03-21 11:51 | 354 | 0.03717 | 0.018028 | 1.4243 | 11.167 | 36.5559 | 0.41508
2019-03-21 11:51 | 355 | 0.028165 | 0.023015 | 1.8781 | 14.7178 | 36.5162 | 0.41452
2019-03-21 11:51 | 356 | 0.031 | 0.021663 | 1.6986 | 13.1201 | 36.4965 | 0.40672
2019-03-21 11:51 | 357 | 0.045634 | 0.014701 | 1.1535 | 8.9525 | 36.4519 | 0.40854
2019-03-21 11:51 | 358 | 0.025129 | 0.026717 | 2.0948 | 16.1174 | 36.4133 | 0.40501
2019-03-21 11:51 | 359 | 0.021611 | 0.03108 | 2.4302 | 18.0996 | 36.3978 | 0.39116
2019-03-21 11:51 | 360 | 0.040826 | 0.016399 | 1.2805 | 9.9257 | 36.3647 | 0.40523
2019-03-21 11:51 | 361 | 0.030904 | 0.022452 | 1.6923 | 13.0564 | 36.323 | 0.40349
2019-03-21 11:51 | 362 | 0.013639 | 0.052532 | 3.8266 | 29.7712 | 36.3052 | 0.40605
2019-03-21 11:51 | 363 | 0.011468 | 0.06258 | 4.5617 | 34.6309 | 36.292 | 0.39716
2019-03-21 11:51 | 364 | 0.013072 | 0.053168 | 3.9798 | 30.9213 | 36.2806 | 0.40421
2019-03-21 11:51 | 365 | 0.022237 | 0.031332 | 2.3469 | 18.4204 | 36.2642 | 0.40962
2019-03-21 11:51 | 366 | 0.016691 | 0.043023 | 3.1034 | 23.8572 | 36.2394 | 0.3982
2019-03-21 11:51 | 367 | 0.020309 | 0.034274 | 2.5493 | 20.3156 | 36.229 | 0.41259
2019-03-21 11:51 | 368 | 0.020814 | 0.03446 | 2.477 | 19.6103 | 36.1997 | 0.40817
2019-03-21 11:51 | 369 | 0.0065267 | 0.11383 | 7.882 | 61.2064 | 36.1903 | 0.39948
2019-03-21 11:51 | 370 | 0.022971 | 0.032373 | 2.2415 | 17.6796 | 36.1796 | 0.40613
2019-03-21 11:51 | 371 | 0.030843 | 0.024838 | 1.6654 | 13.3785 | 36.1464 | 0.41263
2019-03-21 11:51 | 372 | 0.015735 | 0.048775 | 3.2656 | 26.3286 | 36.1234 | 0.41427
2019-03-21 11:51 | 373 | 0.015292 | 0.051768 | 3.3569 | 27.2838 | 36.1115 | 0.41722
2019-03-21 11:51 | 374 | 0.029318 | 0.027791 | 1.7406 | 13.8979 | 36.0895 | 0.40746
2019-03-21 11:51 | 375 | 0.018904 | 0.043103 | 2.7055 | 21.222 | 36.058 | 0.40118
2019-03-21 11:51 | 376 | 0.028495 | 0.0286 | 1.7861 | 14.1255 | 36.0473 | 0.40251
2019-03-21 11:51 | 377 | 0.037873 | 0.022067 | 1.338 | 10.8487 | 36.001 | 0.41087
2019-03-21 11:51 | 378 | 0.031171 | 0.026745 | 1.6239 | 12.7879 | 35.9754 | 0.39861
2019-03-21 11:52 | 379 | 0.033819 | 0.024748 | 1.4929 | 12.3589 | 35.9371 | 0.41796
2019-03-21 11:52 | 380 | 0.021753 | 0.039583 | 2.3301 | 18.5708 | 35.9106 | 0.40397
2019-03-21 11:52 | 381 | 0.031772 | 0.027121 | 1.5873 | 12.8511 | 35.8884 | 0.40831
2019-03-21 11:52 | 382 | 0.044131 | 0.020046 | 1.1432 | 8.9042 | 35.8463 | 0.39296
2019-03-21 11:52 | 383 | 0.038795 | 0.022807 | 1.2968 | 10.4676 | 35.8044 | 0.40609
2019-03-21 11:52 | 384 | 0.03606 | 0.025106 | 1.3869 | 11.0951 | 35.7685 | 0.40009
2019-03-21 11:52 | 385 | 0.037937 | 0.024412 | 1.3081 | 10.2306 | 35.7315 | 0.38812
2019-03-21 11:52 | 386 | 0.027175 | 0.03332 | 1.8286 | 14.3637 | 35.6958 | 0.39034
2019-03-21 11:52 | 387 | 0.017336 | 0.053492 | 2.8622 | 23.4241 | 35.6774 | 0.40609
2019-03-21 11:52 | 388 | -0.00029208 | -3.2566 | -169.3795 | -1339.5424 | 35.6621 | 0.39126
2019-03-21 11:52 | 389 | 0.013444 | 0.06906 | 3.6735 | 28.8161 | 35.6701 | 0.38742
2019-03-21 11:52 | 390 | 0.033337 | 0.028451 | 1.4743 | 11.7622 | 35.6338 | 0.39211
2019-03-21 11:52 | 391 | 0.012382 | 0.076677 | 3.9733 | 32.044 | 35.6128 | 0.39676
2019-03-21 11:52 | 392 | 0.019518 | 0.048733 | 2.529 | 17.9827 | 35.603 | 0.35099
2019-03-21 11:52 | 393 | 0.033902 | 0.028695 | 1.4515 | 11.1985 | 35.5723 | 0.37966
2019-03-21 11:52 | 394 | 0.0087146 | 0.11143 | 5.6159 | 43.348 | 35.5445 | 0.37776
2019-03-21 11:52 | 395 | -0.00089767 | -1.0832 | -54.5904 | -408.4812 | 35.5514 | 0.36668
2019-03-21 11:52 | 396 | 0.010071 | 0.098846 | 4.8333 | 39.3514 | 35.5406 | 0.39631
2019-03-21 11:52 | 397 | 0.019732 | 0.050447 | 2.4686 | 19.0908 | 35.5304 | 0.37671
2019-03-21 11:52 | 398 | 0.051032 | 0.019987 | 0.9544 | 7.4512 | 35.4971 | 0.38025
2019-03-21 11:52 | 399 | 0.032418 | 0.031465 | 1.5036 | 12.191 | 35.4405 | 0.39521
2019-03-21 11:52 | 400 | 0.011301 | 0.090206 | 4.2816 | 33.1791 | 35.4334 | 0.37495
2019-03-21 11:52 | 401 | 0.026846 | 0.03801 | 1.8041 | 13.7681 | 35.4094 | 0.36962
2019-03-21 11:52 | 402 | 0.019536 | 0.053376 | 2.4715 | 18.9268 | 35.3845 | 0.36976
2019-03-21 11:52 | 403 | 0.0066253 | 0.1607 | 7.2632 | 56.4533 | 35.3715 | 0.37402
2019-03-21 11:52 | 404 | 0.022834 | 0.045746 | 2.1087 | 16.0851 | 35.3641 | 0.36729
2019-03-21 11:52 | 405 | 0.039552 | 0.027023 | 1.2186 | 9.6406 | 35.3255 | 0.38131
2019-03-21 11:52 | 406 | 0.034444 | 0.03108 | 1.3941 | 10.3848 | 35.2907 | 0.35769
2019-03-21 11:52 | 407 | 0.022556 | 0.047355 | 2.1176 | 16.3598 | 35.2588 | 0.36901
2019-03-21 11:52 | 408 | 0.016549 | 0.065961 | 2.8934 | 23.3122 | 35.2439 | 0.3858
2019-03-21 11:52 | 409 | 0.024177 | 0.045106 | 1.9786 | 15.1977 | 35.2223 | 0.36743
2019-03-21 11:52 | 410 | 0.023261 | 0.047 | 2.0444 | 16.3622 | 35.1974 | 0.3806
2019-03-21 11:52 | 411 | 0.017463 | 0.062789 | 2.7438 | 20.6683 | 35.1771 | 0.36092
2019-03-21 11:52 | 412 | 0.016833 | 0.064996 | 2.8272 | 22.0347 | 35.161 | 0.37091
2019-03-21 11:52 | 413 | 0.029575 | 0.037763 | 1.5998 | 12.7425 | 35.1405 | 0.37686
2019-03-21 11:52 | 414 | 0.018516 | 0.059068 | 2.5634 | 19.7376 | 35.1072 | 0.36545
2019-03-21 11:52 | 415 | 0.019055 | 0.057414 | 2.4782 | 19.6694 | 35.1012 | 0.37481
2019-03-21 11:52 | 416 | 0.020819 | 0.053795 | 2.2701 | 17.4739 | 35.0689 | 0.36378
2019-03-21 11:52 | 417 | 0.0079689 | 0.14047 | 5.9141 | 46.0542 | 35.0627 | 0.367
2019-03-21 11:52 | 418 | 0.013406 | 0.08533 | 3.5219 | 26.7533 | 35.0485 | 0.35866
2019-03-21 11:52 | 419 | 0.026134 | 0.04454 | 1.7848 | 13.9527 | 35.0339 | 0.36464
2019-03-21 11:52 | 420 | 0.02543 | 0.045689 | 1.8366 | 14.4999 | 34.9993 | 0.36873
2019-03-21 11:52 | 421 | 0.027695 | 0.042056 | 1.6879 | 13.0744 | 34.9825 | 0.3621
2019-03-21 11:52 | 422 | 0.036503 | 0.031919 | 1.278 | 9.6798 | 34.9429 | 0.35334
2019-03-21 11:52 | 423 | 0.021013 | 0.056455 | 2.2127 | 17.3899 | 34.9149 | 0.36542
2019-03-21 11:52 | 424 | 0.025473 | 0.046614 | 1.8284 | 14.2102 | 34.8966 | 0.36198
2019-03-21 11:52 | 425 | 0.024819 | 0.047863 | 1.8656 | 14.5668 | 34.8648 | 0.36153
2019-03-21 11:52 | 426 | 0.021079 | 0.057382 | 2.1953 | 17.4289 | 34.8477 | 0.36738
2019-03-21 11:52 | 427 | 0.022044 | 0.053837 | 2.1017 | 16.1392 | 34.8214 | 0.35577
2019-03-21 11:52 | 428 | 0.027208 | 0.044567 | 1.6983 | 12.8954 | 34.8024 | 0.35086
2019-03-21 11:52 | 429 | 0.035309 | 0.034958 | 1.308 | 9.6679 | 34.767 | 0.34136
2019-03-21 11:52 | 430 | 0.01297 | 0.095084 | 3.5466 | 26.5335 | 34.7389 | 0.34413
2019-03-21 11:52 | 431 | 0.0085152 | 0.14777 | 5.4011 | 40.0491 | 34.7368 | 0.34103
2019-03-21 11:52 | 432 | 0.026401 | 0.047623 | 1.7365 | 13.1663 | 34.7166 | 0.34761
2019-03-21 11:52 | 433 | 0.014364 | 0.088975 | 3.1828 | 24.0557 | 34.6906 | 0.34555
2019-03-21 11:52 | 434 | 0.013419 | 0.093519 | 3.41 | 25.4673 | 34.6855 | 0.34174
2019-03-21 11:52 | 435 | 0.023206 | 0.054153 | 1.9667 | 14.8104 | 34.6611 | 0.34369
2019-03-21 11:52 | 436 | 0.019676 | 0.063936 | 2.3071 | 17.5249 | 34.6419 | 0.34483
2019-03-21 11:52 | 437 | 0.026171 | 0.048881 | 1.7304 | 12.7324 | 34.6199 | 0.33322
2019-03-21 11:52 | 438 | 0.023732 | 0.054788 | 1.9141 | 14.1653 | 34.592 | 0.33617
2019-03-21 11:52 | 439 | 0.0043718 | 0.30317 | 10.3472 | 77.2399 | 34.5759 | 0.33768
2019-03-21 11:53 | 440 | 0.022843 | 0.058006 | 1.9734 | 14.8672 | 34.5745 | 0.3396
2019-03-21 11:53 | 441 | 0.032087 | 0.041946 | 1.4036 | 10.1391 | 34.5322 | 0.32534
2019-03-21 11:53 | 442 | 0.012539 | 0.10751 | 3.5914 | 23.9992 | 34.5168 | 0.30091
2019-03-21 11:53 | 443 | 0.027776 | 0.049538 | 1.6242 | 12.184 | 34.4995 | 0.33843
2019-03-21 11:53 | 444 | 0.03056 | 0.044871 | 1.4688 | 10.1405 | 34.4635 | 0.3099
2019-03-21 11:53 | 445 | 0.034016 | 0.040301 | 1.3117 | 9.8299 | 34.4391 | 0.33437
2019-03-21 11:53 | 446 | 0.017652 | 0.078923 | 2.5383 | 18.7658 | 34.4001 | 0.33125
2019-03-21 11:53 | 447 | 0.016806 | 0.084377 | 2.6518 | 19.5527 | 34.4002 | 0.3286
2019-03-21 11:53 | 448 | 0.02532 | 0.055912 | 1.7572 | 13.2537 | 34.3639 | 0.33558
2019-03-21 11:53 | 449 | 0.0243 | 0.059338 | 1.8318 | 13.1823 | 34.3514 | 0.32032
2019-03-21 11:53 | 450 | 0.032729 | 0.043404 | 1.3641 | 10.1509 | 34.314 | 0.33223
2019-03-21 11:53 | 451 | 0.014383 | 0.10032 | 3.0892 | 21.8179 | 34.2923 | 0.31381
2019-03-21 11:53 | 452 | 0.0032091 | 0.44328 | 13.8289 | 103.4364 | 34.2837 | 0.33194
2019-03-21 11:53 | 453 | 0.0039709 | 0.36361 | 11.1421 | 82.3704 | 34.2824 | 0.32709
2019-03-21 11:53 | 454 | 0.017089 | 0.084315 | 2.5922 | 18.6042 | 34.2729 | 0.31793
2019-03-21 11:53 | 455 | 0.0067036 | 0.21581 | 6.5637 | 46.7899 | 34.253 | 0.31366
2019-03-21 11:53 | 456 | 0.022303 | 0.065754 | 1.9714 | 14.3028 | 34.2541 | 0.31899
2019-03-21 11:53 | 457 | 0.021675 | 0.066639 | 2.0302 | 15.1131 | 34.2122 | 0.32758
2019-03-21 11:53 | 458 | 0.005429 | 0.26641 | 8.0757 | 57.3476 | 34.2142 | 0.31134
2019-03-21 11:53 | 459 | 0.020216 | 0.07276 | 2.1688 | 15.8946 | 34.1943 | 0.32133
2019-03-21 11:53 | 460 | 0.022153 | 0.066246 | 1.9713 | 14.8831 | 34.176 | 0.3297
2019-03-21 11:53 | 461 | 0.030561 | 0.047299 | 1.4302 | 10.7188 | 34.1491 | 0.32757
2019-03-21 11:53 | 462 | 0.025489 | 0.05763 | 1.7063 | 12.854 | 34.118 | 0.32764
2019-03-21 11:53 | 463 | 0.025556 | 0.057403 | 1.6996 | 12.3018 | 34.0972 | 0.31438
2019-03-21 11:53 | 464 | 0.026408 | 0.056459 | 1.6469 | 12.4204 | 34.0669 | 0.32799
2019-03-21 11:53 | 465 | 0.01485 | 0.10058 | 2.9264 | 22.167 | 34.0471 | 0.32918
2019-03-21 11:53 | 466 | 0.016811 | 0.090286 | 2.5866 | 18.5964 | 34.0341 | 0.31263
2019-03-21 11:53 | 467 | 0.021678 | 0.069963 | 1.9992 | 14.6738 | 34.0126 | 0.31809
2019-03-21 11:53 | 468 | 0.021325 | 0.071234 | 2.0407 | 14.6228 | 33.9925 | 0.31183
2019-03-21 11:53 | 469 | 0.00055958 | 2.7121 | 77.1718 | 597.4535 | 33.9747 | 0.33432
2019-03-21 11:53 | 470 | -0.000797 | -1.9308 | -54.153 | -406.2095 | 33.9865 | 0.32375
2019-03-21 11:53 | 471 | 0.016258 | 0.094688 | 2.6513 | 19.0911 | 33.9715 | 0.31039
2019-03-21 11:53 | 472 | 0.017819 | 0.08648 | 2.4276 | 17.8218 | 33.9573 | 0.31757
2019-03-21 11:53 | 473 | 0.021691 | 0.071071 | 1.99 | 14.511 | 33.9362 | 0.31476
2019-03-21 11:53 | 474 | 0.0057056 | 0.27449 | 7.584 | 52.8817 | 33.9185 | 0.30172
2019-03-21 11:53 | 475 | 0.00322 | 0.48678 | 13.4264 | 91.1183 | 33.9212 | 0.2934
2019-03-21 11:53 | 476 | 0.026398 | 0.060194 | 1.6247 | 12.024 | 33.9059 | 0.31741
2019-03-21 11:53 | 477 | 0.026179 | 0.060557 | 1.6261 | 11.1389 | 33.8738 | 0.2916
2019-03-21 11:53 | 478 | 0.01651 | 0.098528 | 2.579 | 19.8868 | 33.8559 | 0.32832
2019-03-21 11:53 | 479 | 0.028535 | 0.057067 | 1.4912 | 11.4028 | 33.8358 | 0.32538
2019-03-21 11:53 | 480 | 0.037689 | 0.043325 | 1.135 | 8.1114 | 33.7998 | 0.30571
2019-03-21 11:53 | 481 | 0.019213 | 0.086197 | 2.2226 | 18.0027 | 33.767 | 0.34589
2019-03-21 11:53 | 482 | 0.018001 | 0.092047 | 2.3775 | 18.2769 | 33.7576 | 0.32901
2019-03-21 11:53 | 483 | 0.024519 | 0.068416 | 1.7336 | 13.2187 | 33.7287 | 0.3241
2019-03-21 11:53 | 484 | 0.024623 | 0.067198 | 1.7223 | 12.5093 | 33.7097 | 0.30802
2019-03-21 11:53 | 485 | 0.031313 | 0.052746 | 1.3577 | 10.1032 | 33.6784 | 0.31636
2019-03-21 11:53 | 486 | 0.019131 | 0.086292 | 2.2155 | 16.469 | 33.651 | 0.31507
2019-03-21 11:53 | 487 | 0.032797 | 0.050487 | 1.2951 | 9.9459 | 33.6346 | 0.32619
2019-03-21 11:53 | 488 | 0.032753 | 0.049918 | 1.2921 | 9.4747 | 33.5888 | 0.31032
2019-03-21 11:53 | 489 | 0.011401 | 0.14303 | 3.7118 | 26.4473 | 33.5741 | 0.30153
2019-03-21 11:53 | 490 | 0.015744 | 0.10361 | 2.6863 | 20.9346 | 33.56 | 0.3296
2019-03-21 11:53 | 491 | 0.024111 | 0.067531 | 1.748 | 13.1352 | 33.541 | 0.3167
2019-03-21 11:53 | 492 | 0.030058 | 0.054084 | 1.3975 | 10.1627 | 33.5126 | 0.30547
2019-03-21 11:53 | 493 | 0.021614 | 0.076369 | 1.9421 | 13.4559 | 33.4844 | 0.29084
2019-03-21 11:53 | 494 | 0.019087 | 0.08644 | 2.1982 | 16.3603 | 33.468 | 0.31227
2019-03-21 11:53 | 495 | 0.028013 | 0.059854 | 1.4997 | 10.8565 | 33.4437 | 0.30412
2019-03-21 11:53 | 496 | 0.020073 | 0.083376 | 2.0799 | 16.8257 | 33.4159 | 0.33775
2019-03-21 11:53 | 497 | 0.010414 | 0.16113 | 4.0407 | 29.3246 | 33.4041 | 0.30539
2019-03-21 11:53 | 498 | 0.011858 | 0.14139 | 3.5488 | 26.9238 | 33.3921 | 0.31925
2019-03-21 11:54 | 499 | 0.024378 | 0.069526 | 1.7107 | 12.585 | 33.3778 | 0.30679
2019-03-21 11:54 | 500 | 0.019766 | 0.085789 | 2.109 | 16.4487 | 33.3475 | 0.32513
2019-03-21 11:54 | 501 | 0.0071285 | 0.23816 | 5.8598 | 43.9069 | 33.3404 | 0.31299
2019-03-21 11:54 | 502 | 0.0043253 | 0.39848 | 9.6465 | 72.0467 | 33.3306 | 0.31163
2019-03-21 11:54 | 503 | 0.01725 | 0.10011 | 2.434 | 19.281 | 33.3274 | 0.3326
2019-03-21 11:54 | 504 | 0.034124 | 0.051342 | 1.2253 | 9.1602 | 33.2956 | 0.31258
2019-03-21 11:54 | 505 | 0.013106 | 0.1332 | 3.1816 | 25.1753 | 33.268 | 0.32995
2019-03-21 11:54 | 506 | 0.013524 | 0.12917 | 3.0933 | 23.6059 | 33.2649 | 0.31924
2019-03-21 11:54 | 507 | 0.022959 | 0.075058 | 1.8202 | 13.3306 | 33.2384 | 0.30606
2019-03-21 11:54 | 508 | 0.022898 | 0.076213 | 1.8284 | 14.9111 | 33.2212 | 0.34144
2019-03-21 11:54 | 509 | 0.018108 | 0.096491 | 2.3027 | 16.7842 | 33.1941 | 0.30393
2019-03-21 11:54 | 510 | 0.014125 | 0.12393 | 2.9603 | 22.6799 | 33.1848 | 0.32035
2019-03-21 11:54 | 511 | 0.014129 | 0.12389 | 2.9697 | 21.0399 | 33.1649 | 0.29728
2019-03-21 11:54 | 512 | 0.012564 | 0.14119 | 3.331 | 25.3671 | 33.1568 | 0.31872
2019-03-21 11:54 | 513 | 0.010711 | 0.16764 | 3.8933 | 28.2889 | 33.1398 | 0.30302
2019-03-21 11:54 | 514 | 0.007659 | 0.23734 | 5.4222 | 39.2154 | 33.1349 | 0.30035
2019-03-21 11:54 | 515 | 0.035809 | 0.051443 | 1.1614 | 8.3462 | 33.1177 | 0.29887
2019-03-21 11:54 | 516 | 0.025459 | 0.073222 | 1.631 | 11.9657 | 33.0719 | 0.30463
2019-03-21 11:54 | 517 | 0.020865 | 0.090518 | 1.9861 | 14.8084 | 33.0661 | 0.30897
2019-03-21 11:54 | 518 | 0.027969 | 0.068313 | 1.4754 | 10.661 | 33.0276 | 0.29818
2019-03-21 11:54 | 519 | 0.020643 | 0.092515 | 1.9928 | 15.2295 | 33.0132 | 0.31438
2019-03-21 11:54 | 520 | 0.019815 | 0.097713 | 2.0757 | 16.4815 | 32.9848 | 0.32657
2019-03-21 11:54 | 521 | 0.026141 | 0.074022 | 1.5668 | 11.5563 | 32.9717 | 0.3021
2019-03-21 11:54 | 522 | 0.030236 | 0.063972 | 1.3565 | 9.4743 | 32.9331 | 0.28647
2019-03-21 11:54 | 523 | 0.029031 | 0.066785 | 1.4136 | 10.9029 | 32.9129 | 0.31653
2019-03-21 11:54 | 524 | 0.018846 | 0.10286 | 2.1752 | 15.7454 | 32.8774 | 0.29674
2019-03-21 11:54 | 525 | 0.0088127 | 0.22024 | 4.6411 | 30.5737 | 32.8745 | 0.26944
2019-03-21 11:54 | 526 | 0.035349 | 0.054776 | 1.1532 | 9.0108 | 32.8518 | 0.31853
2019-03-21 11:54 | 527 | 0.015932 | 0.12024 | 2.5623 | 17.7228 | 32.8139 | 0.28236
2019-03-21 11:54 | 528 | 0.0095181 | 0.20446 | 4.2969 | 34.2661 | 32.8173 | 0.32615
2019-03-21 11:54 | 529 | 0.023305 | 0.083179 | 1.7418 | 11.6678 | 32.79 | 0.27192
2019-03-21 11:54 | 530 | 0.020897 | 0.093962 | 1.9424 | 15.0431 | 32.7742 | 0.31435
2019-03-21 11:54 | 531 | 0.017963 | 0.10915 | 2.2564 | 17.9803 | 32.7485 | 0.32299
2019-03-21 11:54 | 532 | 0.023995 | 0.082656 | 1.6809 | 11.6203 | 32.736 | 0.27883
2019-03-21 11:54 | 533 | 0.033855 | 0.058595 | 1.1862 | 8.9629 | 32.6998 | 0.30344
2019-03-21 11:54 | 534 | 0.02667 | 0.075039 | 1.5042 | 12.2168 | 32.6726 | 0.32582
2019-03-21 11:54 | 535 | 0.013309 | 0.15046 | 3.0159 | 22.3028 | 32.6482 | 0.29682
2019-03-21 11:54 | 536 | 0.0075852 | 0.26362 | 5.2746 | 35.3077 | 32.6438 | 0.26782
2019-03-21 11:54 | 537 | 0.017046 | 0.1172 | 2.3363 | 15.9789 | 32.6289 | 0.27237
2019-03-21 11:54 | 538 | 0.029826 | 0.067789 | 1.3334 | 9.4728 | 32.6085 | 0.28254
2019-03-21 11:54 | 539 | 0.041136 | 0.04976 | 0.96499 | 6.642 | 32.5702 | 0.27322
2019-03-21 11:54 | 540 | 0.028206 | 0.072712 | 1.4088 | 9.1334 | 32.5325 | 0.25761
2019-03-21 11:54 | 541 | 0.013411 | 0.1532 | 2.9738 | 20.1812 | 32.5147 | 0.27064
2019-03-21 11:54 | 542 | 0.011547 | 0.17776 | 3.4315 | 26.1214 | 32.5026 | 0.30162
2019-03-21 11:54 | 543 | 0.0064518 | 0.32273 | 6.1599 | 35.4896 | 32.4918 | 0.22897
2019-03-21 11:54 | 544 | 0.0072865 | 0.28569 | 5.4529 | 36.734 | 32.488 | 0.26766
2019-03-21 11:54 | 545 | 0.017041 | 0.12346 | 2.3174 | 18.5135 | 32.4752 | 0.31549
2019-03-21 11:54 | 546 | 0.010495 | 0.20023 | 3.7688 | 22.7294 | 32.4577 | 0.23856
2019-03-21 11:54 | 547 | 0.0062995 | 0.33332 | 6.2622 | 49.345 | 32.454 | 0.31085
2019-03-21 11:54 | 548 | 0.0061128 | 0.34464 | 6.4749 | 46.6373 | 32.444 | 0.28508
2019-03-21 11:54 | 549 | 0.0048733 | 0.43213 | 8.1034 | 57.5002 | 32.4415 | 0.28021
2019-03-21 11:54 | 550 | 0.015824 | 0.13293 | 2.4788 | 17.2967 | 32.4311 | 0.2737
2019-03-21 11:54 | 551 | 0.031656 | 0.067284 | 1.2421 | 9.0239 | 32.4088 | 0.28566
2019-03-21 11:54 | 552 | 0.028485 | 0.073796 | 1.38 | 7.2656 | 32.3725 | 0.20696
2019-03-21 11:54 | 553 | 0.022938 | 0.092614 | 1.7049 | 11.9554 | 32.3526 | 0.27424
2019-03-21 11:54 | 554 | 0.033346 | 0.064504 | 1.169 | 10.9046 | 32.3233 | 0.36363
2019-03-21 11:54 | 555 | 0.021292 | 0.10128 | 1.8494 | 8.593 | 32.2908 | 0.18296
2019-03-21 11:54 | 556 | 0.023596 | 0.092471 | 1.6625 | 13.802 | 32.2778 | 0.32568
2019-03-21 11:54 | 557 | 0.022985 | 0.095637 | 1.7107 | 13.2152 | 32.2441 | 0.30375
2019-03-21 11:54 | 558 | 0.016789 | 0.13112 | 2.33 | 9.1389 | 32.2331 | 0.15343
2019-03-21 11:54 | 559 | 0.019728 | 0.11127 | 1.9886 | 20.0833 | 32.2088 | 0.39619
2019-03-21 11:55 | 560 | 0.0053467 | 0.41082 | 7.3349 | 49.8935 | 32.1973 | 0.26676
2019-03-21 11:55 | 561 | 0.0059916 | 0.37106 | 6.5245 | 34.4997 | 32.1944 | 0.20671
2019-03-21 11:55 | 562 | 0.022604 | 0.098525 | 1.7389 | 18.4878 | 32.1816 | 0.4179
2019-03-21 11:55 | 563 | 0.012516 | 0.17972 | 3.142 | 7.9933 | 32.1555 | 0.10005
2019-03-21 11:55 | 564 | -0.00049358 | -4.5462 | -79.254 | -628.8462 | 32.1571 | 0.31039
2019-03-21 11:55 | 565 | 0.020212 | 0.11232 | 1.9416 | 20.5644 | 32.1485 | 0.41564
2019-03-21 11:55 | 566 | 0.026467 | 0.086634 | 1.4711 | 7.3184 | 32.1198 | 0.1937
2019-03-21 11:55 | 567 | 0.019109 | 0.12145 | 2.0376 | 10.1673 | 32.099 | 0.19429
2019-03-21 11:55 | 568 | 0.02342 | 0.099164 | 1.6764 | 11.3826 | 32.079 | 0.26659
2019-03-21 11:55 | 569 | 0.030443 | 0.076202 | 1.2858 | 11.2382 | 32.0515 | 0.34212
2019-03-21 11:55 | 570 | 0.027751 | 0.083492 | 1.3995 | 18.4247 | 32.0207 | 0.51131
2019-03-21 11:55 | 571 | 0.017648 | 0.13283 | 2.2021 | 4.913 | 31.998 | 0.086703
2019-03-21 11:55 | 572 | 0.012665 | 0.18302 | 3.0591 | 17.8168 | 31.9842 | 0.22564
2019-03-21 11:55 | 573 | 0.010085 | 0.23024 | 3.8447 | 38.5291 | 31.9711 | 0.38858
2019-03-21 11:55 | 574 | 0.03246 | 0.07149 | 1.1813 | 5.2715 | 31.9578 | 0.17111
2019-03-21 11:55 | 575 | 0.04791 | 0.048451 | 0.80599 | 6.549 | 31.9083 | 0.31376
2019-03-21 11:55 | 576 | 0.024677 | 0.093881 | 1.5662 | 12.9187 | 31.8712 | 0.3188
2019-03-21 11:55 | 577 | 0.029488 | 0.078539 | 1.3065 | 9.1349 | 31.853 | 0.26937
2019-03-21 11:55 | 578 | 0.038241 | 0.059807 | 1.006 | 5.9254 | 31.8111 | 0.22659
2019-03-21 11:55 | 579 | 0.024116 | 0.09571 | 1.5921 | 9.187 | 31.7816 | 0.22155
2019-03-21 11:55 | 580 | 0.025371 | 0.090904 | 1.5078 | 11.5731 | 31.7594 | 0.29362
2019-03-21 11:55 | 581 | 0.037625 | 0.06188 | 1.0152 | 8.5651 | 31.7282 | 0.32227
2019-03-21 11:55 | 582 | 0.032213 | 0.072189 | 1.1832 | 8.268 | 31.6881 | 0.26634
2019-03-21 11:55 | 583 | 0.033047 | 0.071042 | 1.1517 | 8.2461 | 31.6629 | 0.27251
2019-03-21 11:55 | 584 | 0.030326 | 0.078267 | 1.254 | 8.8347 | 31.6233 | 0.26792
2019-03-21 11:55 | 585 | 0.0059114 | 0.3956 | 6.3889 | 46.2253 | 31.607 | 0.27326
2019-03-21 11:55 | 586 | 0.01151 | 0.2039 | 3.3057 | 23.7038 | 31.6045 | 0.27284
2019-03-21 11:55 | 587 | 0.016741 | 0.13989 | 2.2592 | 12.7553 | 31.5834 | 0.21354
2019-03-21 11:55 | 588 | 0.012501 | 0.18971 | 3.025 | 21.3438 | 31.573 | 0.26683
2019-03-21 11:55 | 589 | 0.021403 | 0.11092 | 1.7704 | 12.9874 | 31.5556 | 0.27798
2019-03-21 11:55 | 590 | 0.02195 | 0.10806 | 1.7314 | 12.251 | 31.5322 | 0.26891
2019-03-21 11:55 | 591 | 0.014101 | 0.16811 | 2.6831 | 19.3576 | 31.5139 | 0.27296
2019-03-21 11:55 | 592 | 0.0094508 | 0.25347 | 4.0024 | 33.0078 | 31.503 | 0.31195
2019-03-21 11:55 | 593 | 0.020323 | 0.11683 | 1.8611 | 13.127 | 31.4914 | 0.26678
2019-03-21 11:55 | 594 | 0.016126 | 0.14737 | 2.3522 | 13.1286 | 31.4656 | 0.21171
2019-03-21 11:55 | 595 | 0.016442 | 0.14567 | 2.3091 | 18.1553 | 31.4584 | 0.29851
2019-03-21 11:55 | 596 | 0.018955 | 0.12624 | 1.98 | 11.4429 | 31.4319 | 0.2169
2019-03-21 11:55 | 597 | 0.027546 | 0.08713 | 1.3745 | 11.0117 | 31.419 | 0.30333
2019-03-21 11:55 | 598 | 0.03147 | 0.076307 | 1.1991 | 7.481 | 31.3784 | 0.23542
2019-03-21 11:55 | 599 | 0.014122 | 0.16987 | 2.6617 | 19.0638 | 31.3611 | 0.26921
2019-03-21 11:55 | 600 | 0.013639 | 0.17759 | 2.7478 | 21.8993 | 31.3461 | 0.29868
2019-03-21 11:55 | 601 | 0.025692 | 0.095235 | 1.4652 | 8.9354 | 31.3303 | 0.22957
2019-03-21 11:55 | 602 | 0.034954 | 0.070635 | 1.0744 | 7.1184 | 31.2954 | 0.24882
2019-03-21 11:55 | 603 | 0.033567 | 0.074304 | 1.1185 | 8.0621 | 31.2632 | 0.27062
2019-03-21 11:55 | 604 | 0.02985 | 0.083607 | 1.2598 | 9.6168 | 31.2292 | 0.28706
2019-03-21 11:55 | 605 | 0.025281 | 0.099717 | 1.4828 | 10.3017 | 31.2039 | 0.26044
2019-03-21 11:55 | 606 | 0.02034 | 0.12376 | 1.844 | 13.1114 | 31.179 | 0.26668
2019-03-21 11:55 | 607 | 0.0092817 | 0.27091 | 4.0403 | 29.3559 | 31.1647 | 0.27247
2019-03-21 11:55 | 608 | 0.00038442 | 6.5359 | 97.2868 | 666.8493 | 31.1597 | 0.25635
2019-03-21 11:55 | 609 | 0.00024133 | 10.3219 | 154.4699 | 1088.6602 | 31.1612 | 0.26273
2019-03-21 11:55 | 610 | 0.01582 | 0.15763 | 2.3636 | 15.1512 | 31.1554 | 0.23969
2019-03-21 11:55 | 611 | 0.01763 | 0.1413 | 2.1104 | 15.7485 | 31.1329 | 0.27765
2019-03-21 11:55 | 612 | 0.0066139 | 0.37361 | 5.666 | 33.9703 | 31.1237 | 0.22468
2019-03-21 11:55 | 613 | -0.0038027 | -0.65066 | -9.8677 | -70.8656 | 31.1191 | 0.26948
2019-03-21 11:55 | 614 | 0.016538 | 0.14965 | 2.2629 | 15.8053 | 31.1242 | 0.26139
2019-03-21 11:55 | 615 | 0.01151 | 0.21249 | 3.2406 | 21.7355 | 31.0914 | 0.25018
2019-03-21 11:55 | 616 | 0.0082848 | 0.29546 | 4.5191 | 34.8948 | 31.1007 | 0.2891
2019-03-21 11:55 | 617 | 0.031369 | 0.078724 | 1.1846 | 7.5796 | 31.0694 | 0.23776
2019-03-21 11:55 | 618 | 0.012639 | 0.19393 | 2.9546 | 22.7265 | 31.0473 | 0.28724
2019-03-21 11:55 | 619 | 0.011772 | 0.20802 | 3.1569 | 21.7629 | 31.0404 | 0.25619
2019-03-21 11:56 | 620 | 0.026244 | 0.092419 | 1.4258 | 10.2474 | 31.0201 | 0.26894
2019-03-21 11:56 | 621 | 0.016377 | 0.14812 | 2.2963 | 18.3365 | 30.9931 | 0.30029
2019-03-21 11:56 | 622 | 0.021293 | 0.1148 | 1.7628 | 11.6244 | 30.9839 | 0.24752
2019-03-21 11:56 | 623 | 0.040117 | 0.061169 | 0.93467 | 7.4176 | 30.9477 | 0.29757
2019-03-21 11:56 | 624 | 0.020434 | 0.12097 | 1.8382 | 14.0477 | 30.9125 | 0.28705
2019-03-21 11:56 | 625 | 0.018924 | 0.13197 | 1.9905 | 13.4765 | 30.903 | 0.25503
2019-03-21 11:56 | 626 | 0.028858 | 0.087225 | 1.2971 | 11.1625 | 30.8716 | 0.32212
2019-03-21 11:56 | 627 | 0.029308 | 0.087438 | 1.2791 | 9.0997 | 30.8472 | 0.2667
2019-03-21 11:56 | 628 | 0.033605 | 0.076861 | 1.1229 | 8.7047 | 30.8129 | 0.29252
2019-03-21 11:56 | 629 | 0.016557 | 0.15762 | 2.2624 | 16.729 | 30.7853 | 0.27698
2019-03-21 11:56 | 630 | -0.00042217 | -6.2327 | -89.0151 | -615.2854 | 30.7794 | 0.25975
2019-03-21 11:56 | 631 | 0.01678 | 0.15849 | 2.2458 | 16.4381 | 30.7779 | 0.27582
2019-03-21 11:56 | 632 | 0.020327 | 0.13173 | 1.8594 | 14.6246 | 30.7485 | 0.29727
2019-03-21 11:56 | 633 | 0.014321 | 0.18734 | 2.652 | 19.15 | 30.7395 | 0.27425
2019-03-21 11:56 | 634 | 0.02376 | 0.11264 | 1.5992 | 11.2802 | 30.716 | 0.26801
2019-03-21 11:56 | 635 | 0.041276 | 0.064743 | 0.91389 | 7.3594 | 30.69 | 0.30377
2019-03-21 11:56 | 636 | 0.045405 | 0.058963 | 0.83391 | 6.1681 | 30.6377 | 0.28007
2019-03-21 11:56 | 637 | 0.011376 | 0.23164 | 3.3501 | 25.1551 | 30.6081 | 0.28616
2019-03-21 11:56 | 638 | 0.012302 | 0.2145 | 3.1112 | 21.9376 | 30.6058 | 0.26988
2019-03-21 11:56 | 639 | 0.061643 | 0.043258 | 0.61775 | 4.7275 | 30.5717 | 0.29142
2019-03-21 11:56 | 640 | 0.066207 | 0.040554 | 0.57188 | 4.6329 | 30.4937 | 0.30673
2019-03-21 11:56 | 641 | 0.026813 | 0.10025 | 1.4165 | 10.6352 | 30.4505 | 0.28517
2019-03-21 11:56 | 642 | 0.028278 | 0.095006 | 1.345 | 10.1075 | 30.4311 | 0.28582
2019-03-21 11:56 | 643 | 0.032479 | 0.082762 | 1.175 | 8.3095 | 30.3926 | 0.26989
2019-03-21 11:56 | 644 | 0.01819 | 0.14769 | 2.0867 | 15.8027 | 30.3699 | 0.28744
2019-03-21 11:56 | 645 | 0.026826 | 0.10104 | 1.4222 | 10.9765 | 30.3511 | 0.29446
2019-03-21 11:56 | 646 | 0.030874 | 0.087921 | 1.2387 | 9.6558 | 30.3172 | 0.29811
2019-03-21 11:56 | 647 | 0.018959 | 0.14146 | 2.0026 | 15.3895 | 30.2929 | 0.29177
2019-03-21 11:56 | 648 | 0.030945 | 0.087423 | 1.2305 | 9.095 | 30.2744 | 0.28145
2019-03-21 11:56 | 649 | 0.021048 | 0.12835 | 1.8013 | 13.8287 | 30.2362 | 0.29107
2019-03-21 11:56 | 650 | -0.0044083 | -0.60815 | -8.6342 | -66.2278 | 30.2354 | 0.29195
2019-03-21 11:56 | 651 | 0.01654 | 0.1624 | 2.3057 | 17.6535 | 30.2337 | 0.29199
2019-03-21 11:56 | 652 | 0.03952 | 0.068621 | 0.96956 | 6.782 | 30.2014 | 0.26802
2019-03-21 11:56 | 653 | 0.029417 | 0.092917 | 1.2954 | 9.9275 | 30.1625 | 0.29204
2019-03-21 11:56 | 654 | 0.035965 | 0.07534 | 1.0574 | 8.1302 | 30.1402 | 0.29241
2019-03-21 11:56 | 655 | 0.015056 | 0.18174 | 2.5361 | 21.5122 | 30.0968 | 0.32389
2019-03-21 11:56 | 656 | 0.0043395 | 0.62389 | 8.7644 | 67.7023 | 30.1072 | 0.29379
2019-03-21 11:56 | 657 | 0.028971 | 0.092671 | 1.3208 | 9.3009 | 30.0794 | 0.26946
2019-03-21 11:56 | 658 | 0.034287 | 0.077404 | 1.1118 | 8.7277 | 30.0535 | 0.29925
2019-03-21 11:56 | 659 | 0.024599 | 0.10783 | 1.5532 | 12.7079 | 30.0146 | 0.3126
2019-03-21 11:56 | 660 | 0.031424 | 0.083681 | 1.2114 | 9.5449 | 30.0002 | 0.29994
2019-03-21 11:56 | 661 | 0.054464 | 0.048326 | 0.70362 | 5.1043 | 29.9485 | 0.278
2019-03-21 11:56 | 662 | 0.033433 | 0.07811 | 1.1463 | 8.786 | 29.9016 | 0.29374
2019-03-21 11:56 | 663 | 0.029226 | 0.089429 | 1.3125 | 10.7478 | 29.879 | 0.31412
2019-03-21 11:56 | 664 | 0.013469 | 0.19232 | 2.8452 | 20.8551 | 29.8451 | 0.28091
2019-03-21 11:56 | 665 | 0.016725 | 0.15504 | 2.2958 | 17.6429 | 29.8475 | 0.29508
2019-03-21 11:56 | 666 | 0.0183 | 0.14184 | 2.0944 | 17.7277 | 29.8114 | 0.32442
2019-03-21 11:56 | 667 | 0.021841 | 0.11904 | 1.7713 | 13.1239 | 29.8103 | 0.28663
2019-03-21 11:56 | 668 | 0.030669 | 0.083913 | 1.2587 | 10.8249 | 29.7668 | 0.33199
2019-03-21 11:56 | 669 | 0.022857 | 0.11439 | 1.69 | 13.5944 | 29.7523 | 0.31073
2019-03-21 11:56 | 670 | 0.042261 | 0.061898 | 0.91623 | 7.0638 | 29.7159 | 0.29852
2019-03-21 11:56 | 671 | 0.019432 | 0.1345 | 2.0022 | 17.2577 | 29.6777 | 0.33536
2019-03-21 11:56 | 672 | -0.0030126 | -0.86612 | -12.8935 | -107.9439 | 29.6773 | 0.32519
2019-03-21 11:56 | 673 | 0.0019744 | 1.3189 | 19.6893 | 167.3254 | 29.6762 | 0.33037
2019-03-21 11:56 | 674 | 0.026519 | 0.098387 | 1.4743 | 11.7856 | 29.6677 | 0.31255
2019-03-21 11:56 | 675 | 0.04495 | 0.058014 | 0.86688 | 7.8037 | 29.6246 | 0.35078
2019-03-21 11:56 | 676 | 0.04567 | 0.057088 | 0.85545 | 7.2718 | 29.5826 | 0.3321
2019-03-21 11:56 | 677 | 0.044067 | 0.059149 | 0.88966 | 7.6881 | 29.5348 | 0.33879
2019-03-21 11:56 | 678 | 0.030761 | 0.084702 | 1.2752 | 10.1346 | 29.4973 | 0.31176
2019-03-21 11:56 | 679 | 0.026075 | 0.10001 | 1.5112 | 12.0749 | 29.4714 | 0.31485
2019-03-21 11:57 | 680 | 0.023254 | 0.1133 | 1.6907 | 13.818 | 29.4441 | 0.32133
2019-03-21 11:57 | 681 | 0.028304 | 0.09381 | 1.3953 | 11.9721 | 29.4225 | 0.33886
2019-03-21 11:57 | 682 | 0.041505 | 0.064569 | 0.95289 | 7.8899 | 29.3855 | 0.32747
2019-03-21 11:57 | 683 | 0.043987 | 0.061272 | 0.89644 | 7.4383 | 29.3418 | 0.32719
2019-03-21 11:57 | 684 | 0.055189 | 0.048965 | 0.72236 | 6.3376 | 29.2965 | 0.34977
2019-03-21 11:57 | 685 | 0.035308 | 0.076895 | 1.1206 | 9.3956 | 29.2384 | 0.33174
2019-03-21 11:57 | 686 | 0.0381 | 0.071657 | 1.0382 | 8.4882 | 29.2207 | 0.3234
2019-03-21 11:57 | 687 | 0.055583 | 0.049517 | 0.7134 | 5.7104 | 29.1587 | 0.3174
2019-03-21 11:57 | 688 | 0.040625 | 0.069513 | 0.97684 | 8.6882 | 29.1163 | 0.35296
2019-03-21 11:57 | 689 | 0.059854 | 0.048646 | 0.66372 | 7.1237 | 29.0703 | 0.42639
2019-03-21 11:57 | 690 | 0.058948 | 0.05058 | 0.67468 | 10.7706 | 29.001 | 0.63491
2019-03-21 11:57 | 691 | 0.047956 | 0.06354 | 0.83434 | 16.2843 | 28.9545 | 0.78093
2019-03-21 11:57 | 692 | 0.070016 | 0.044411 | 0.58725 | 11.4866 | 28.897 | 0.80425
2019-03-21 11:57 | 693 | 0.09459 | 0.033234 | 0.44715 | 9.5434 | 28.8144 | 0.90272
2019-03-21 11:57 | 694 | 0.065657 | 0.048241 | 0.67677 | 14.3139 | 28.7204 | 0.93981
2019-03-21 11:57 | 695 | 0.066101 | 0.047917 | 0.70987 | 14.8691 | 28.6766 | 0.98287
2019-03-21 11:57 | 696 | 0.089394 | 0.03534 | 0.55374 | 12.486 | 28.5823 | 1.1162
2019-03-21 11:57 | 697 | 0.090905 | 0.034407 | 0.58133 | 13.355 | 28.5025 | 1.214
2019-03-21 11:57 | 698 | 0.095445 | 0.032432 | 0.58871 | 13.8918 | 28.4007 | 1.3259
2019-03-21 11:57 | 699 | 0.085237 | 0.0362 | 0.7002 | 16.5609 | 28.3147 | 1.4116
2019-03-21 11:57 | 700 | 0.096782 | 0.032066 | 0.65579 | 15.0302 | 28.2252 | 1.4547
2019-03-21 11:57 | 701 | 0.10463 | 0.029758 | 0.64452 | 14.6075 | 28.1216 | 1.5283
2019-03-21 11:57 | 702 | 0.11169 | 0.028036 | 0.64437 | 14.4393 | 28.0167 | 1.6128
2019-03-21 11:57 | 703 | 0.10101 | 0.031494 | 0.75315 | 16.7941 | 27.9024 | 1.6964
2019-03-21 11:57 | 704 | 0.10063 | 0.031865 | 0.80013 | 18.3812 | 27.8121 | 1.8498
2019-03-21 11:57 | 705 | 0.1187 | 0.027442 | 0.71812 | 16.9343 | 27.6968 | 2.0102
2019-03-21 11:57 | 706 | 0.11506 | 0.028781 | 0.78485 | 18.7828 | 27.5798 | 2.1611
2019-03-21 11:57 | 707 | 0.10451 | 0.032163 | 0.91701 | 21.8635 | 27.4682 | 2.2849
2019-03-21 11:57 | 708 | 0.12195 | 0.027933 | 0.82977 | 19.911 | 27.3641 | 2.4282
2019-03-21 11:57 | 709 | 0.1333 | 0.025893 | 0.80123 | 19.2271 | 27.2252 | 2.5629
2019-03-21 11:57 | 710 | 0.1349 | 0.026239 | 0.83445 | 20.0991 | 27.0994 | 2.7113
2019-03-21 11:57 | 711 | 0.16085 | 0.022354 | 0.75481 | 18.3206 | 26.9502 | 2.9469
2019-03-21 11:57 | 712 | 0.16789 | 0.021688 | 0.77272 | 19.5257 | 26.7825 | 3.2781
2019-03-21 11:57 | 713 | 0.15304 | 0.024071 | 0.90998 | 22.2023 | 26.62 | 3.3979
2019-03-21 11:57 | 714 | 0.15327 | 0.024263 | 0.98549 | 23.1069 | 26.4728 | 3.5415
2019-03-21 11:57 | 715 | 0.16916 | 0.022191 | 0.96744 | 22.058 | 26.3092 | 3.7312
2019-03-21 11:57 | 716 | 0.18009 | 0.020742 | 0.98021 | 21.4961 | 26.1353 | 3.8712
2019-03-21 11:57 | 717 | 0.18864 | 0.019651 | 1.0022 | 21.6562 | 25.9499 | 4.0853
2019-03-21 11:57 | 718 | 0.19216 | 0.019078 | 1.0433 | 22.2742 | 25.7594 | 4.2802
2019-03-21 11:57 | 719 | 0.19849 | 0.018387 | 1.0697 | 22.3344 | 25.5656 | 4.4331
2019-03-21 11:57 | 720 | 0.18978 | 0.018983 | 1.1782 | 23.8043 | 25.3661 | 4.5175
2019-03-21 11:57 | 721 | 0.18497 | 0.019459 | 1.2741 | 24.7142 | 25.1861 | 4.5714
2019-03-21 11:57 | 722 | 0.15726 | 0.022867 | 1.548 | 29.5259 | 25.0009 | 4.6431
2019-03-21 11:57 | 723 | 0.16175 | 0.022622 | 1.5833 | 28.8027 | 24.8632 | 4.6589
2019-03-21 11:57 | 724 | 0.19623 | 0.019102 | 1.3504 | 23.6698 | 24.6702 | 4.6448
2019-03-21 11:57 | 725 | 0.18545 | 0.020812 | 1.4733 | 24.9999 | 24.4819 | 4.6363
2019-03-21 11:57 | 726 | 0.15639 | 0.025652 | 1.799 | 29.3709 | 24.3052 | 4.5934
2019-03-21 11:57 | 727 | 0.13021 | 0.031968 | 2.2307 | 34.4363 | 24.1682 | 4.4841
2019-03-21 11:57 | 728 | 0.13494 | 0.031998 | 2.2042 | 32.2097 | 24.0373 | 4.3463
2019-03-21 11:57 | 729 | 0.12317 | 0.035937 | 2.4709 | 33.9827 | 23.9013 | 4.1855
2019-03-21 11:57 | 730 | 0.12145 | 0.036964 | 2.5503 | 32.5188 | 23.789 | 3.9495
2019-03-21 11:57 | 731 | 0.11166 | 0.040237 | 2.8113 | 32.7317 | 23.6603 | 3.6549
2019-03-21 11:57 | 732 | 0.095854 | 0.04638 | 3.2992 | 34.8 | 23.567 | 3.3357
2019-03-21 11:57 | 733 | 0.098522 | 0.043631 | 3.2166 | 30.5843 | 23.4646 | 3.0132
2019-03-21 11:57 | 734 | 0.085044 | 0.048665 | 3.7072 | 31.8813 | 23.374 | 2.7113
2019-03-21 11:57 | 735 | 0.052817 | 0.074878 | 5.9114 | 47.5925 | 23.2989 | 2.5137
2019-03-21 11:57 | 736 | 0.045452 | 0.082544 | 6.7741 | 52.434 | 23.2618 | 2.3832
2019-03-21 11:57 | 737 | 0.068941 | 0.051728 | 4.3816 | 32.7614 | 23.2003 | 2.2586
2019-03-21 11:57 | 738 | 0.060776 | 0.056289 | 4.8543 | 34.3554 | 23.1308 | 2.088
2019-03-21 11:58 | 739 | 0.054555 | 0.059623 | 5.2686 | 35.6684 | 23.0788 | 1.9459
2019-03-21 11:58 | 740 | 0.061654 | 0.050392 | 4.5319 | 30.1532 | 23.0184 | 1.8591
2019-03-21 11:58 | 741 | 0.064098 | 0.046544 | 4.232 | 27.2234 | 22.956 | 1.745
2019-03-21 11:58 | 742 | 0.075693 | 0.03778 | 3.476 | 21.9867 | 22.8884 | 1.6642
2019-03-21 11:58 | 743 | 0.076135 | 0.036108 | 3.3383 | 21.6139 | 22.8076 | 1.6456
2019-03-21 11:58 | 744 | 0.062926 | 0.042443 | 3.9182 | 25.5856 | 22.7393 | 1.61
2019-03-21 11:58 | 745 | 0.069719 | 0.03744 | 3.4277 | 22.3796 | 22.6769 | 1.5603
2019-03-21 11:58 | 746 | 0.080167 | 0.031751 | 2.8844 | 18.7404 | 22.5987 | 1.5024
2019-03-21 11:58 | 747 | 0.082832 | 0.030002 | 2.7086 | 17.5208 | 22.519 | 1.4513
2019-03-21 11:58 | 748 | 0.066039 | 0.036711 | 3.2998 | 21.2697 | 22.4377 | 1.4046
2019-03-21 11:58 | 749 | 0.067038 | 0.035592 | 3.1621 | 21.621 | 22.3829 | 1.4494
2019-03-21 11:58 | 750 | 0.073351 | 0.031957 | 2.807 | 19.2988 | 22.3024 | 1.4156
2019-03-21 11:58 | 751 | 0.055478 | 0.041892 | 3.6075 | 25.6536 | 22.2416 | 1.4232
2019-03-21 11:58 | 752 | 0.052649 | 0.043696 | 3.698 | 26.4703 | 22.1878 | 1.3936
2019-03-21 11:58 | 753 | 0.068827 | 0.033067 | 2.7525 | 20.3391 | 22.1314 | 1.3999
2019-03-21 11:58 | 754 | 0.077191 | 0.028892 | 2.3932 | 18.2019 | 22.0521 | 1.405
2019-03-21 11:58 | 755 | 0.064971 | 0.034362 | 2.7822 | 21.2896 | 21.9819 | 1.3832
2019-03-21 11:58 | 756 | 0.070688 | 0.031217 | 2.4964 | 19.2239 | 21.9187 | 1.3589
2019-03-21 11:58 | 757 | 0.061896 | 0.03523 | 2.7904 | 22.0326 | 21.8436 | 1.3637
2019-03-21 11:58 | 758 | 0.056072 | 0.038617 | 3.0336 | 23.9333 | 21.7939 | 1.342
2019-03-21 11:58 | 759 | 0.073166 | 0.029191 | 2.2742 | 18.1074 | 21.7263 | 1.3248
2019-03-21 11:58 | 760 | 0.06424 | 0.032894 | 2.5441 | 21.0569 | 21.6529 | 1.3527
2019-03-21 11:58 | 761 | 0.070543 | 0.029654 | 2.2819 | 18.6262 | 21.5943 | 1.3139
2019-03-21 11:58 | 762 | 0.094327 | 0.022216 | 1.6831 | 14.3422 | 21.5086 | 1.3529
2019-03-21 11:58 | 763 | 0.067068 | 0.031301 | 2.3361 | 19.8417 | 21.4177 | 1.3307
2019-03-21 11:58 | 764 | 0.049421 | 0.0425 | 3.1312 | 26.2417 | 21.3725 | 1.2969
2019-03-21 11:58 | 765 | 0.066941 | 0.031401 | 2.2807 | 19.7539 | 21.3104 | 1.3224
2019-03-21 11:58 | 766 | 0.062015 | 0.033908 | 2.4326 | 21.3894 | 21.2429 | 1.3265
2019-03-21 11:58 | 767 | 0.065183 | 0.031867 | 2.2865 | 20.3324 | 21.1846 | 1.3253
2019-03-21 11:58 | 768 | 0.079808 | 0.026029 | 1.8459 | 17.079 | 21.1105 | 1.363
2019-03-21 11:58 | 769 | 0.061349 | 0.033533 | 2.3832 | 21.4874 | 21.0329 | 1.3182
2019-03-21 11:58 | 770 | 0.049467 | 0.041225 | 2.9363 | 26.1749 | 20.9866 | 1.2948
2019-03-21 11:58 | 771 | 0.050814 | 0.040068 | 2.8265 | 26.2383 | 20.9303 | 1.3333
2019-03-21 11:58 | 772 | 0.0583 | 0.034619 | 2.4486 | 23.0303 | 20.8831 | 1.3427
2019-03-21 11:58 | 773 | 0.06359 | 0.03136 | 2.2258 | 20.1379 | 20.8143 | 1.2806
2019-03-21 11:58 | 774 | 0.055195 | 0.036066 | 2.5347 | 24.3528 | 20.7587 | 1.3442
2019-03-21 11:58 | 775 | 0.074339 | 0.026523 | 1.8709 | 17.6255 | 20.6982 | 1.3103
2019-03-21 11:58 | 776 | 0.067087 | 0.029072 | 2.0558 | 19.6036 | 20.6164 | 1.3151
2019-03-21 11:58 | 777 | 0.043034 | 0.044877 | 3.1889 | 29.6832 | 20.5677 | 1.2774
2019-03-21 11:58 | 778 | 0.062791 | 0.03077 | 2.172 | 20.9217 | 20.5198 | 1.3137
2019-03-21 11:58 | 779 | 0.080548 | 0.023681 | 1.6819 | 16.3936 | 20.4422 | 1.3205
2019-03-21 11:58 | 780 | 0.066918 | 0.028544 | 2.0169 | 20.0203 | 20.3661 | 1.3397
2019-03-21 11:58 | 781 | 0.074575 | 0.02563 | 1.8002 | 18.3222 | 20.3047 | 1.3664
2019-03-21 11:58 | 782 | 0.057869 | 0.032641 | 2.3106 | 23.0858 | 20.2223 | 1.3359
2019-03-21 11:58 | 783 | 0.047642 | 0.039728 | 2.7976 | 28.7539 | 20.1868 | 1.3699
2019-03-21 11:58 | 784 | 0.080809 | 0.023147 | 1.6446 | 16.783 | 20.1175 | 1.3562
2019-03-21 11:58 | 785 | 0.056522 | 0.033443 | 2.3402 | 23.3376 | 20.0379 | 1.3191
2019-03-21 11:58 | 786 | 0.047167 | 0.039651 | 2.7994 | 27.1461 | 20.0011 | 1.2804
2019-03-21 11:58 | 787 | 0.081929 | 0.023101 | 1.6046 | 16.8356 | 19.9334 | 1.3793
2019-03-21 11:58 | 788 | 0.073539 | 0.025689 | 1.7819 | 18.2395 | 19.8472 | 1.3413
2019-03-21 11:58 | 789 | 0.050419 | 0.037483 | 2.5882 | 25.2361 | 19.7904 | 1.2724
2019-03-21 11:58 | 790 | 0.05753 | 0.032777 | 2.2578 | 24.9531 | 19.7395 | 1.4356
2019-03-21 11:58 | 791 | 0.057003 | 0.033104 | 2.276 | 21.7607 | 19.6765 | 1.2404
2019-03-21 11:58 | 792 | 0.05126 | 0.03679 | 2.5227 | 25.4293 | 19.6266 | 1.3035
2019-03-21 11:58 | 793 | 0.057915 | 0.032837 | 2.2219 | 24.453 | 19.5714 | 1.4162
2019-03-21 11:58 | 794 | 0.050451 | 0.038107 | 2.5412 | 23.0718 | 19.5138 | 1.164
2019-03-21 11:58 | 795 | 0.04741 | 0.041043 | 2.6986 | 28.6483 | 19.4694 | 1.3582
2019-03-21 11:58 | 796 | 0.057056 | 0.033662 | 2.2357 | 30.035 | 19.416 | 1.7137
2019-03-21 11:58 | 797 | 0.059418 | 0.032723 | 2.1416 | 23.6844 | 19.3572 | 1.4073
2019-03-21 11:58 | 798 | 0.044489 | 0.043652 | 2.8522 | 36.4971 | 19.3012 | 1.6237
2019-03-21 11:58 | 799 | 0.045892 | 0.042373 | 2.7648 | 29.8928 | 19.2642 | 1.3718
2019-03-21 11:59 | 800 | 0.061001 | 0.031885 | 2.0759 | 21.131 | 19.2058 | 1.289
2019-03-21 11:59 | 801 | 0.061432 | 0.032016 | 2.0599 | 18.78 | 19.1451 | 1.1537
2019-03-21 11:59 | 802 | 0.076276 | 0.025893 | 1.6619 | 15.1133 | 19.0801 | 1.1528
2019-03-21 11:59 | 803 | 0.081943 | 0.024083 | 1.5457 | 12.9862 | 18.9952 | 1.0641
2019-03-21 11:59 | 804 | 0.061318 | 0.03259 | 2.0696 | 22.7129 | 18.9222 | 1.3927
2019-03-21 11:59 | 805 | 0.070227 | 0.028178 | 1.807 | 19.8878 | 18.8655 | 1.3967
2019-03-21 11:59 | 806 | 0.095131 | 0.020785 | 1.3304 | 15.8889 | 18.7782 | 1.5115
2019-03-21 11:59 | 807 | 0.073417 | 0.02662 | 1.7236 | 21.8634 | 18.6862 | 1.6051
2019-03-21 11:59 | 808 | 0.062045 | 0.031845 | 2.0393 | 19.3165 | 18.6295 | 1.1985
2019-03-21 11:59 | 809 | 0.07475 | 0.02609 | 1.6888 | 19.1156 | 18.5563 | 1.4289
2019-03-21 11:59 | 810 | 0.071739 | 0.027201 | 1.7627 | 21.7869 | 18.4834 | 1.563
2019-03-21 11:59 | 811 | 0.062158 | 0.03103 | 2.0357 | 22.9198 | 18.4151 | 1.4246
2019-03-21 11:59 | 812 | 0.045931 | 0.041472 | 2.7501 | 28.7371 | 18.3607 | 1.3199
2019-03-21 11:59 | 813 | 0.038829 | 0.049617 | 3.2604 | 40.1791 | 18.3202 | 1.5601
2019-03-21 11:59 | 814 | 0.062676 | 0.030679 | 2.0176 | 21.1623 | 18.2756 | 1.3264
2019-03-21 11:59 | 815 | 0.06564 | 0.029283 | 1.9285 | 23.8032 | 18.2002 | 1.5624
2019-03-21 11:59 | 816 | 0.033129 | 0.057977 | 3.8202 | 44.9177 | 18.1529 | 1.4881
2019-03-21 11:59 | 817 | 0.030316 | 0.063317 | 4.1835 | 47.5568 | 18.1265 | 1.4417
2019-03-21 11:59 | 818 | 0.060852 | 0.031922 | 2.0898 | 19.8886 | 18.0836 | 1.2103
2019-03-21 11:59 | 819 | 0.072168 | 0.027188 | 1.7647 | 19.8338 | 18.0091 | 1.4314
2019-03-21 11:59 | 820 | 0.066298 | 0.02968 | 1.928 | 22.9944 | 17.9437 | 1.5245
2019-03-21 11:59 | 821 | 0.075281 | 0.026433 | 1.7049 | 20.834 | 17.8733 | 1.5684
2019-03-21 11:59 | 822 | 0.084611 | 0.023482 | 1.5203 | 20.4474 | 17.793 | 1.7301
2019-03-21 11:59 | 823 | 0.081243 | 0.024771 | 1.5904 | 19.0999 | 17.7075 | 1.5517
2019-03-21 11:59 | 824 | 0.067937 | 0.029573 | 1.9059 | 22.563 | 17.6326 | 1.5329
2019-03-21 11:59 | 825 | 0.084869 | 0.023416 | 1.5324 | 17.3879 | 17.5648 | 1.4757
2019-03-21 11:59 | 826 | 0.08562 | 0.023217 | 1.5238 | 16.5721 | 17.4665 | 1.4189
2019-03-21 11:59 | 827 | 0.072019 | 0.02732 | 1.8214 | 18.723 | 17.3973 | 1.3484
2019-03-21 11:59 | 828 | 0.070768 | 0.027815 | 1.8628 | 22.5337 | 17.3199 | 1.5947
2019-03-21 11:59 | 829 | 0.067788 | 0.028923 | 1.9452 | 21.0755 | 17.2559 | 1.4287
2019-03-21 11:59 | 830 | 0.066472 | 0.029456 | 1.9862 | 22.2081 | 17.1843 | 1.4762
2019-03-21 11:59 | 831 | 0.051595 | 0.037996 | 2.5728 | 29.6242 | 17.1259 | 1.5285
2019-03-21 11:59 | 832 | 0.051429 | 0.038015 | 2.5875 | 30.1157 | 17.0778 | 1.5488
2019-03-21 11:59 | 833 | 0.048582 | 0.040193 | 2.7485 | 31.8922 | 17.0229 | 1.5494
2019-03-21 11:59 | 834 | 0.058182 | 0.033549 | 2.3079 | 26.4037 | 16.9775 | 1.5362
2019-03-21 11:59 | 835 | 0.064089 | 0.030495 | 2.1026 | 24.147 | 16.9071 | 1.5475
2019-03-21 11:59 | 836 | 0.068802 | 0.028693 | 1.9654 | 19.2059 | 16.8493 | 1.3214
2019-03-21 11:59 | 837 | 0.090531 | 0.022018 | 1.4999 | 11.5546 | 16.7663 | 1.0461
2019-03-21 11:59 | 838 | 0.076182 | 0.026219 | 1.7957 | 15.6383 | 16.6769 | 1.1913
2019-03-21 11:59 | 839 | 0.06232 | 0.032037 | 2.2051 | 26.2539 | 16.614 | 1.6362
2019-03-21 11:59 | 840 | 0.076969 | 0.02596 | 1.7967 | 22.4409 | 16.5454 | 1.7273
2019-03-21 11:59 | 841 | 0.082162 | 0.024587 | 1.6924 | 21.2054 | 16.4621 | 1.7423
2019-03-21 11:59 | 842 | 0.069225 | 0.029189 | 2.0202 | 30.0357 | 16.385 | 2.0792
2019-03-21 11:59 | 843 | 0.087318 | 0.023311 | 1.6055 | 21.8326 | 16.317 | 1.9064
2019-03-21 11:59 | 844 | 0.086698 | 0.023721 | 1.6231 | 20.3083 | 16.2153 | 1.7607
2019-03-21 11:59 | 845 | 0.04917 | 0.041771 | 2.8749 | 35.2839 | 16.1524 | 1.7349
2019-03-21 11:59 | 846 | 0.04621 | 0.044944 | 3.0733 | 39.8687 | 16.1087 | 1.8423
2019-03-21 11:59 | 847 | 0.064202 | 0.032426 | 2.2291 | 29.5309 | 16.0539 | 1.8959
2019-03-21 11:59 | 848 | 0.077325 | 0.027249 | 1.8596 | 23.1206 | 15.9815 | 1.7878
2019-03-21 11:59 | 849 | 0.063022 | 0.033403 | 2.2883 | 28.8434 | 15.9058 | 1.8178
2019-03-21 11:59 | 850 | 0.065614 | 0.031979 | 2.2017 | 24.7148 | 15.8516 | 1.6216
2019-03-21 11:59 | 851 | 0.084383 | 0.024892 | 1.7215 | 20.0135 | 15.7705 | 1.6888
2019-03-21 11:59 | 852 | 0.085011 | 0.024719 | 1.7173 | 21.4538 | 15.687 | 1.8238
2019-03-21 11:59 | 853 | 0.0862 | 0.024317 | 1.6977 | 23.043 | 15.6013 | 1.9863
2019-03-21 11:59 | 854 | 0.070678 | 0.030037 | 2.0856 | 30.1281 | 15.5184 | 2.1294
2019-03-21 11:59 | 855 | 0.068507 | 0.030928 | 2.157 | 26.5012 | 15.4562 | 1.8155
2019-03-21 11:59 | 856 | 0.099851 | 0.021665 | 1.4883 | 18.1601 | 15.3736 | 1.8133
2019-03-21 11:59 | 857 | 0.096344 | 0.022378 | 1.5444 | 16.9368 | 15.265 | 1.6318
2019-03-21 11:59 | 858 | 0.06817 | 0.031926 | 2.1879 | 20.8318 | 15.187 | 1.4201
2019-03-21 12:00 | 859 | 0.0785 | 0.027693 | 1.9039 | 20.7006 | 15.1202 | 1.625
2019-03-21 12:00 | 860 | 0.074202 | 0.029272 | 2.0194 | 22.5839 | 15.0331 | 1.6758
2019-03-21 12:00 | 861 | 0.053084 | 0.040985 | 2.8345 | 32.411 | 14.9757 | 1.7205
2019-03-21 12:00 | 862 | 0.05439 | 0.039951 | 2.7698 | 32.6047 | 14.9216 | 1.7734
2019-03-21 12:00 | 863 | 0.05965 | 0.036391 | 2.5327 | 24.7356 | 14.865 | 1.4755
2019-03-21 12:00 | 864 | 0.070133 | 0.030955 | 2.1622 | 24.2998 | 14.8008 | 1.7042
2019-03-21 12:00 | 865 | 0.083716 | 0.025958 | 1.8201 | 22.0047 | 14.7245 | 1.8421
2019-03-21 12:00 | 866 | 0.074963 | 0.028715 | 2.0392 | 28.1382 | 14.6391 | 2.1093
2019-03-21 12:00 | 867 | 0.059326 | 0.036316 | 2.5876 | 35.7409 | 14.5764 | 2.1204
2019-03-21 12:00 | 868 | 0.06521 | 0.033317 | 2.356 | 28.2381 | 14.5152 | 1.8414
2019-03-21 12:00 | 869 | 0.073926 | 0.029191 | 2.0925 | 22.4521 | 14.4449 | 1.6598
2019-03-21 12:00 | 870 | 0.070751 | 0.030501 | 2.1913 | 26.2553 | 14.3702 | 1.8576
2019-03-21 12:00 | 871 | 0.066497 | 0.032093 | 2.3312 | 31.7033 | 14.3036 | 2.1082
2019-03-21 12:00 | 872 | 0.077167 | 0.027656 | 2.0142 | 24.213 | 14.2338 | 1.8684
2019-03-21 12:00 | 873 | 0.077332 | 0.027619 | 2.0159 | 21.8738 | 14.1516 | 1.6915
2019-03-21 12:00 | 874 | 0.074546 | 0.028326 | 2.0935 | 27.1937 | 14.0799 | 2.0272
2019-03-21 12:00 | 875 | 0.080232 | 0.026296 | 1.9485 | 21.9158 | 14.0008 | 1.7584
2019-03-21 12:00 | 876 | 0.077466 | 0.026986 | 2.0258 | 25.1548 | 13.9217 | 1.9486
2019-03-21 12:00 | 877 | 0.061314 | 0.03408 | 2.5617 | 28.6636 | 13.8488 | 1.7575
2019-03-21 12:00 | 878 | 0.07094 | 0.029366 | 2.2117 | 27.7328 | 13.7932 | 1.9674
2019-03-21 12:00 | 879 | 0.070267 | 0.029694 | 2.2378 | 25.5549 | 13.7087 | 1.7957
2019-03-21 12:00 | 880 | 0.067299 | 0.031305 | 2.3385 | 29.3244 | 13.6528 | 1.9735
2019-03-21 12:00 | 881 | 0.089078 | 0.02365 | 1.7697 | 19.9531 | 13.5689 | 1.7774
2019-03-21 12:00 | 882 | 0.079641 | 0.026786 | 1.9869 | 25.3693 | 13.4822 | 2.0204
2019-03-21 12:00 | 883 | 0.058637 | 0.036837 | 2.7089 | 30.8604 | 13.4127 | 1.8096
2019-03-21 12:00 | 884 | 0.064772 | 0.033271 | 2.4531 | 30.363 | 13.3585 | 1.9667
2019-03-21 12:00 | 885 | 0.069805 | 0.030846 | 2.2767 | 27.0381 | 13.2831 | 1.8874
2019-03-21 12:00 | 886 | 0.054814 | 0.039322 | 2.9074 | 34.4179 | 13.2228 | 1.8866
2019-03-21 12:00 | 887 | 0.079969 | 0.026985 | 1.9982 | 23.7174 | 13.1642 | 1.8966
2019-03-21 12:00 | 888 | 0.091933 | 0.023198 | 1.7364 | 20.4487 | 13.0656 | 1.8799
2019-03-21 12:00 | 889 | 0.087365 | 0.024384 | 1.8275 | 21.5345 | 12.9853 | 1.8814
2019-03-21 12:00 | 890 | 0.068424 | 0.030814 | 2.3347 | 26.6845 | 12.8947 | 1.8259
2019-03-21 12:00 | 891 | 0.059888 | 0.034855 | 2.6707 | 31.1223 | 12.8456 | 1.8639
2019-03-21 12:00 | 892 | 0.07602 | 0.027471 | 2.1049 | 24.7157 | 12.769 | 1.8789
2019-03-21 12:00 | 893 | 0.069176 | 0.03011 | 2.3096 | 26.1316 | 12.6982 | 1.8077
2019-03-21 12:00 | 894 | 0.077693 | 0.026511 | 2.0518 | 23.8939 | 12.628 | 1.8564
2019-03-21 12:00 | 895 | 0.066956 | 0.030752 | 2.3776 | 27.7076 | 12.5473 | 1.8552
2019-03-21 12:00 | 896 | 0.056883 | 0.036209 | 2.7977 | 32.1632 | 12.4939 | 1.8296
2019-03-21 12:00 | 897 | 0.062396 | 0.032638 | 2.548 | 28.983 | 12.4296 | 1.8084
2019-03-21 12:00 | 898 | 0.07061 | 0.028903 | 2.2555 | 26.0066 | 12.3686 | 1.8363
2019-03-21 12:00 | 899 | 0.054607 | 0.036998 | 2.9164 | 34.1485 | 12.2934 | 1.8648
2019-03-21 12:00 | 900 | 0.07214 | 0.028025 | 2.2109 | 25.4074 | 12.2518 | 1.8329
2019-03-21 12:00 | 901 | 0.085757 | 0.023234 | 1.8514 | 21.0747 | 12.1504 | 1.8073
2019-03-21 12:00 | 902 | 0.053917 | 0.037054 | 2.9544 | 33.6136 | 12.0906 | 1.8124
2019-03-21 12:00 | 903 | 0.069942 | 0.028563 | 2.2765 | 25.6315 | 12.0325 | 1.7927
2019-03-21 12:00 | 904 | 0.061578 | 0.032388 | 2.5807 | 29.0158 | 11.955 | 1.7867
2019-03-21 12:00 | 905 | 0.069329 | 0.028889 | 2.2984 | 26.1205 | 11.9055 | 1.8109
2019-03-21 12:00 | 906 | 0.090415 | 0.022103 | 1.7541 | 19.8099 | 11.8136 | 1.7911
2019-03-21 12:00 | 907 | 0.078053 | 0.025629 | 2.0321 | 22.8221 | 11.7332 | 1.7813
2019-03-21 12:00 | 908 | 0.046663 | 0.042738 | 3.3836 | 37.7713 | 11.6621 | 1.7625
2019-03-21 12:00 | 909 | 0.055375 | 0.035639 | 2.8476 | 32.4877 | 11.63 | 1.799
2019-03-21 12:00 | 910 | 0.078569 | 0 | 0 | 0 | 11.5475 | 1.7633
2019-03-21 12:00 | 911 | 0.064835 | 0 | 0 | 0 | 11.4812 | 1.7436
2019-03-21 12:00 | 912 | 0.067504 | 0 | 0 | 0 | 11.4149 | 1.7726
2019-03-21 12:00 | 913 | 0.066954 | 0 | 0 | 0 | 11.3467 | 1.725
2019-03-21 12:00 | 914 | 0.064918 | 0 | 0 | 0 | 11.2818 | 1.7342
2019-03-21 12:00 | 915 | 0.049034 | 0 | 0 | 0 | 11.2197 | 1.7465
2019-03-21 12:00 | 916 | 0.058592 | 0 | 0 | 0 | 11.1779 | 1.7092
2019-03-21 12:00 | 917 | 0.058125 | 0 | 0 | 0 | 11.1041 | 1.7034
2019-03-21 12:00 | 918 | 0.065728 | 0 | 0 | 0 | 11.0595 | 1.6966
2019-03-21 12:00 | 919 | 0.087105 | 0 | 0 | 0 | 10.9701 | 1.6665
2019-03-21 12:01 | 920 | 0.070236 | 0 | 0 | 0 | 10.8946 | 1.6701
2019-03-21 12:01 | 921 | 0.051302 | 0 | 0 | 0 | 10.8305 | 1.6386
2019-03-21 12:01 | 922 | 0.057949 | 0 | 0 | 0 | 10.7856 | 1.6272
2019-03-21 12:01 | 923 | 0.068889 | 0 | 0 | 0 | 10.7132 | 1.592
2019-03-21 12:01 | 924 | 0.058913 | 0 | 0 | 0 | 10.6519 | 1.6127
2019-03-21 12:01 | 925 | 0.084278 | 0 | 0 | 0 | 10.5878 | 1.6347
2019-03-21 12:01 | 926 | 0.086484 | 0 | 0 | 0 | 10.4889 | 1.6159
2019-03-21 12:01 | 927 | 0.066435 | 0 | 0 | 0 | 10.4209 | 1.6026
2019-03-21 12:01 | 928 | 0.05341 | 0 | 0 | 0 | 10.3546 | 1.6248
2019-03-21 12:01 | 929 | 0.056201 | 0 | 0 | 0 | 10.3102 | 1.6376
2019-03-21 12:01 | 930 | 0.044757 | 0 | 0 | 0 | 10.2448 | 1.6183
2019-03-21 12:01 | 931 | 0.047888 | 0 | 0 | 0 | 10.217 | 1.5853
2019-03-21 12:01 | 932 | 0.066106 | 0 | 0 | 0 | 10.1455 | 1.5991
2019-03-21 12:01 | 933 | 0.05616 | 0 | 0 | 0 | 10.0908 | 1.5354
2019-03-21 12:01 | 934 | 0.073322 | 0 | 0 | 0 | 10.0284 | 1.5309
2019-03-21 12:01 | 935 | 0.050651 | 0 | 0 | 0 | 9.9534 | 1.5192
2019-03-21 12:01 | 936 | 0.02994 | 0 | 0 | 0 | 9.9264 | 1.5006
2019-03-21 12:01 | 937 | 0.05006 | 0 | 0 | 0 | 9.8833 | 1.4849
2019-03-21 12:01 | 938 | 0.066202 | 0 | 0 | 0 | 9.8267 | 1.4807
2019-03-21 12:01 | 939 | 0.060731 | 0 | 0 | 0 | 9.7564 | 1.4585
2019-03-21 12:01 | 940 | 0.05104 | 0 | 0 | 0 | 9.7062 | 1.4459
2019-03-21 12:01 | 941 | 0.070318 | 0 | 0 | 0 | 9.648 | 1.4498
2019-03-21 12:01 | 942 | 0.063312 | 0 | 0 | 0 | 9.5718 | 1.4098
2019-03-21 12:01 | 943 | 0.044934 | 0 | 0 | 0 | 9.5248 | 1.4036
2019-03-21 12:01 | 944 | 0.034845 | 0 | 0 | 0 | 9.4799 | 1.4233
2019-03-21 12:01 | 945 | 0.036386 | 0 | 0 | 0 | 9.4518 | 1.3947
2019-03-21 12:01 | 946 | 0.037826 | 0 | 0 | 0 | 9.4065 | 1.3818
2019-03-21 12:01 | 947 | 0.046764 | 0 | 0 | 0 | 9.3743 | 1.3782
2019-03-21 12:01 | 948 | 0.05593 | 0 | 0 | 0 | 9.3132 | 1.3743
2019-03-21 12:01 | 949 | 0.043666 | 0 | 0 | 0 | 9.267 | 1.341
2019-03-21 12:01 | 950 | 0.060805 | 0 | 0 | 0 | 9.2196 | 1.3463
2019-03-21 12:01 | 951 | 0.062854 | 0 | 0 | 0 | 9.1489 | 1.3401
2019-03-21 12:01 | 952 | 0.047396 | 0 | 0 | 0 | 9.098 | 1.3307
2019-03-21 12:01 | 953 | 0.053932 | 0 | 0 | 0 | 9.0489 | 1.3217
2019-03-21 12:01 | 954 | 0.065655 | 0 | 0 | 0 | 8.9886 | 1.3102
2019-03-21 12:01 | 955 | 0.063507 | 0 | 0 | 0 | 8.9211 | 1.3022
2019-03-21 12:01 | 956 | 0.05609 | 0 | 0 | 0 | 8.8634 | 1.2872
2019-03-21 12:01 | 957 | 0.04851 | 0 | 0 | 0 | 8.8093 | 1.2554
2019-03-21 12:01 | 958 | 0.038914 | 0 | 0 | 0 | 8.7668 | 1.2467
2019-03-21 12:01 | 959 | 0.030551 | 0 | 0 | 0 | 8.7306 | 1.2198
2019-03-21 12:01 | 960 | 0.042733 | 0 | 0 | 0 | 8.7009 | 1.2153
2019-03-21 12:01 | 961 | 0.039558 | 0 | 0 | 0 | 8.6487 | 1.2208
2019-03-21 12:01 | 962 | 0.02417 | 0 | 0 | 0 | 8.6246 | 1.1785
2019-03-21 12:01 | 963 | 0.034301 | 0 | 0 | 0 | 8.5944 | 1.1686
2019-03-21 12:01 | 964 | 0.041062 | 0 | 0 | 0 | 8.5564 | 1.1902
2019-03-21 12:01 | 965 | 0.035663 | 0 | 0 | 0 | 8.515 | 1.1847
2019-03-21 12:01 | 966 | 0.041232 | 0 | 0 | 0 | 8.4827 | 1.192
2019-03-21 12:01 | 967 | 0.0497 | 0 | 0 | 0 | 8.4323 | 1.1579
2019-03-21 12:01 | 968 | 0.030657 | 0 | 0 | 0 | 8.3891 | 1.1616
2019-03-21 12:01 | 969 | 0.048296 | 0 | 0 | 0 | 8.3629 | 1.1459
2019-03-21 12:01 | 970 | 0.054532 | 0 | 0 | 0 | 8.2952 | 1.1475
2019-03-21 12:01 | 971 | 0.03181 | 0 | 0 | 0 | 8.2603 | 1.1002
2019-03-21 12:01 | 972 | 0.04669 | 0 | 0 | 0 | 8.2233 | 1.0848
2019-03-21 12:01 | 973 | 0.049717 | 0 | 0 | 0 | 8.1696 | 1.0594
2019-03-21 12:01 | 974 | 0.029651 | 0 | 0 | 0 | 8.1292 | 1.0732
2019-03-21 12:01 | 975 | 0.035724 | 0 | 0 | 0 | 8.1045 | 1.0512
2019-03-21 12:01 | 976 | 0.037916 | 0 | 0 | 0 | 8.0587 | 1.042
2019-03-21 12:01 | 977 | 0.013987 | 0 | 0 | 0 | 8.0338 | 1.0366
2019-03-21 12:01 | 978 | 0.044209 | 0 | 0 | 0 | 8.0187 | 1.027
2019-03-21 12:02 | 979 | 0.043137 | 0 | 0 | 0 | 7.9526 | 1.0107
2019-03-21 12:02 | 980 | 0.011061 | 0 | 0 | 0 | 7.9398 | 0.99597
2019-03-21 12:02 | 981 | 0.023221 | 0 | 0 | 0 | 7.9202 | 0.9779
2019-03-21 12:02 | 982 | 0.038802 | 0 | 0 | 0 | 7.8917 | 0.96724
2019-03-21 12:02 | 983 | 0.040628 | 0 | 0 | 0 | 7.8464 | 0.95623
2019-03-21 12:02 | 984 | 0.021825 | 0 | 0 | 0 | 7.8154 | 0.94581
2019-03-21 12:02 | 985 | 0.026465 | 0 | 0 | 0 | 7.7973 | 0.94148
2019-03-21 12:02 | 986 | 0.036766 | 0 | 0 | 0 | 7.7604 | 0.90645
2019-03-21 12:02 | 987 | 0.045966 | 0 | 0 | 0 | 7.7244 | 0.92718
2019-03-21 12:02 | 988 | 0.035226 | 0 | 0 | 0 | 7.6734 | 0.90134
2019-03-21 12:02 | 989 | 0.030536 | 0 | 0 | 0 | 7.6529 | 0.91718
2019-03-21 12:02 | 990 | 0.030406 | 0 | 0 | 0 | 7.611 | 0.88302
2019-03-21 12:02 | 991 | 0.031362 | 0 | 0 | 0 | 7.5917 | 0.88631
2019-03-21 12:02 | 992 | 0.031742 | 0 | 0 | 0 | 7.549 | 0.8804
2019-03-21 12:02 | 993 | 0.0068176 | 0 | 0 | 0 | 7.5329 | 0.84493
2019-03-21 12:02 | 994 | 0.043335 | 0 | 0 | 0 | 7.5212 | 0.87054
2019-03-21 12:02 | 995 | 0.058002 | 0 | 0 | 0 | 7.4512 | 0.82913
2019-03-21 12:02 | 996 | 0.031829 | 0 | 0 | 0 | 7.4155 | 0.84503
2019-03-21 12:02 | 997 | 0.02699 | 0 | 0 | 0 | 7.3832 | 0.8137
2019-03-21 12:02 | 998 | 0.028617 | 0 | 0 | 0 | 7.3596 | 0.82047
2019-03-21 12:02 | 999 | 0.020421 | 0 | 0 | 0 | 7.3276 | 0.80016
2019-03-21 12:02 | 1000 | 0.032105 | 0 | 0 | 0 | 7.3141 | 0.78853
2019-03-21 12:02 | 1001 | 0.04144 | 0 | 0 | 0 | 7.2643 | 0.78965
2019-03-21 12:02 | 1002 | 0.020745 | 0 | 0 | 0 | 7.2379 | 0.78116
2019-03-21 12:02 | 1003 | 0.034592 | 0 | 0 | 0 | 7.2155 | 0.77426
2019-03-21 12:02 | 1004 | 0.027561 | 0 | 0 | 0 | 7.1732 | 0.76864
2019-03-21 12:02 | 1005 | 0.013381 | 0 | 0 | 0 | 7.162 | 0.75293
2019-03-21 12:02 | 1006 | 0.025179 | 0 | 0 | 0 | 7.1404 | 0.72091
2019-03-21 12:02 | 1007 | 0.029158 | 0 | 0 | 0 | 7.1131 | 0.71346
2019-03-21 12:02 | 1008 | 0.022852 | 0 | 0 | 0 | 7.0844 | 0.71236
2019-03-21 12:02 | 1009 | 0.032497 | 0 | 0 | 0 | 7.064 | 0.72844
2019-03-21 12:02 | 1010 | 0.034026 | 0 | 0 | 0 | 7.0216 | 0.69627
2019-03-21 12:02 | 1011 | 0.013626 | 0 | 0 | 0 | 7.0007 | 0.70043

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062478 | 0.062678 | 0 | 0.2095 | 2019-03-21 09:44 | Group A | 0.1 | 0.22 | Do Nothing | 0.20951 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025146 | 0.025271 | 0 | 0.0261 | 2019-03-21 09:44 | Group A | -0.01 | 0.1 | Do Nothing | 0.00046117 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002726 | 0.002739 | 0 | 0.0011 | 2019-03-21 09:44 | Group A | -0.01 | 0.1 | Do Nothing | -3.0915e-06 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.6721 | 183.6734 | 0 | 100 | 2019-03-21 09:44 | None | 0 | 1000 | Do Nothing | 100.6152 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.4138 | 209.9622 | 0 | 100 | 2019-03-21 09:44 | None | 0 | 1000 | Do Nothing | 100.7692 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -16.237 | 1447 | 0 |  | 2019-03-21 09:27 | None | -10 | 1000 | Do Nothing | 124.1744 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 20.9733 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0018925 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017707 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -151.6735 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23713 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4026 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0035818 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0049747 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0048017 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 6.1047e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0010148 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 75.1771 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.3184 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5479 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 99.6136 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 20.9226 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 37.2486 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 67.0822 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 107.7951 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 64.1521 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.009266 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0467 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.024928 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 23.8164 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN