## Cone Summary
Date: | 2019-03-28
Time: | 14:56
Program Version: | MIDAS.exe Ver. 2.124.0.842
File Name: | C0S-Test2
Test Series information: | Cube Test - Control - 75 kW/m2; Frame @ 30mm below bottom face of Cone Assembly New foam- Insulator= low Density Ceramic Board (6mm thick)
Test Series Folder: | E:\ConeLab\DATA\2019\03_Mar\Cube Test
Test Director: | Zammarano
Qualified Operator: | Shields
Post Processing Version | NFRL_Report58.m
Lab Name | NIST Cone Lab
Lab Location | Gaithersburg, MD
Test Parameters: | Value | Unit | Event Description | Time (s) | System Time
Calibration Factor, Cf | 0.0413 |  | Start Test | 0 | 2019-03-28 15:01
Fuel Energy Value, Ef | 13100 | KJ/kg | Ignition | 5 | 2019-03-28 15:01
Sample Area | 1 | m^2 | Dripping | 86 | 2019-03-28 15:02
Cone Heat Flux | 75 | kW/m^2 | Wetting | 87 | 2019-03-28 15:02
Cone Orientation |  |  | Flame Out | 286 | 2019-03-28 15:05
Sample Grid |  |  | Bottom Ignition | 104
Sample Thickness | 106 | mm
Sample Initial Mass | 44.2 | g
Cone Separation Distance | 30 | mm
Exhaust Flow Rate | 24.1739 | L/s
Test Result Summary: (Ignition to Flame Out) | Value | Unit | Test Result Summary: (10% to 90% Mass Loss) | Value | Unit
Peak Heat Release Rate | 12.084 | kW/m2 | Average Heat Release Rate | 5.8741 | kW/m2
Total Heat Release | 1.1373 | MJ/m2 | Average Mass Loss Rate | 0.23039 | g/m2s
Peak Mass Loss Rate | 0.47381 | g/m2s | Effective Heat of Combustion | 25.5838 | MJ/kg
Total Mass Lost | 45.0593 | g | Average Specific Extinction Area | 175.7031 | m2/kg
Time to Ignition | 5 | s | Peak CO Yield | 0.038059 | kg/kg
Time to Peak Heat Release | 125 | s | Peak CO2 Yield | 3.0251 | kg/kg
Time to Flame Out | 286 | s | Average CO Yield | 0.019106 | kg/kg
Average Heat Release Rate | 4.033 | kW/m2 | Average CO2 Yield | 2.1737 | kg/kg
Average Mass Loss Rate | 0.16012 | g/m2s
Effective Heat of Combustion | 25.2401 | MJ/kg
Dripping | 87 | s | Time to dripping
HRRD | 2.8447 | kW | HRR at dripping time
THRD | 0.23098 | MJ | Heat released from Ignition to Dripping
AHRRD | 2.8168 | kW | AHRR from Ignition to Dripping
Wetting | 86 | s | Time to wetting
HRRW | 2.8398 | kW | HRR at wetting time
THRW | 0.22814 | MJ | Heat released from Ignition to Wetting
AHRRW | 2.7833 | kW | AHRR from Ignition to wetting

## Scaled Data
Time Stamp | NTP Epoch Time (s) | Time (s) | 0: O2 (Vol fr) | 1: CO2 (Vol fr) | 2: CO (Vol fr) | 3: Io (%) | 4: I (%) | 5: Pe (Pa) | 6: CJC0 (°C) | 7: TCraw1 (V) | 8: TCraw2 (V) | 9: SampMass (g) | 10: BurnerFlow (SLPM) | 11: BurnerFlowSet (SLPM) | 12: TCraw3 (V) | 13: TCraw4 (V) | 14: TCraw5 (V) | 15: Hfspare (V) | 16: TCraw7 (V) | 17: HeatFlux (kW/m^2) | 18: CanopyFlow (V) | 19: MarkSig (V) | 20: Pamb (kPa) | 21: Tamb (°C) | 22: RHamb (%) | 23: Tsmk (°C) | 24: Tcone (°C) | 25: Te (°C) | 26: H2Oa (Vol fr) | 27: Rhoe (kg/m3) | 28: MFRe (kg/s) | 29: VFRe (L/s) | 30: Ksm (1/m) | 31: HRRmethane (kW) | 32: Phi (none) | 33: RHR (kW) | 34: RHRA (kW/m2) | 35: newCf (none) | 36: EvMark (none)
2019-03-28 14:56 | 1553799377.7712 | -261 | 0.20951 | 0.000445 | -1e-06 | 100.6681 | 100.6224 | 126.3671 | 21.2612 | 0.001875 | 0.001772 | -151.6362 | 0.23059 | 8.406 | 0.002527 | 0.003448 | 0.003424 | 4.9e-05 | 0.001306 | 0.3147 | 1.3515 | 2.5471 | 100.71 | 21.16 | 25.692 | 66.9521 | 82.6114 | 64.4534 | 0.006415 | 1.0457 | 0.025268 | 24.1624 | 0.004066 | 0.13718 | -8e-06 | -0.000561 | -0.000561 | -10.0979
2019-03-28 14:56 | 1553799378.7713 | -260 | 0.20951 | 0.000445 | -1e-06 | 100.6681 | 100.6224 | 126.3671 | 21.2612 | 0.001875 | 0.001772 | -151.6362 | 0.23059 | 8.406 | 0.002527 | 0.003448 | 0.003424 | 4.9e-05 | 0.001306 | 0.3147 | 1.3515 | 2.5471 | 100.71 | 21.16 | 25.691 | 66.9521 | 82.6114 | 64.4534 | 0.006414 | 1.0457 | 0.025268 | 24.1624 | 0.004066 | 0.13718 | -8e-06 | -0.00056 | -0.00056 | -10.1197
2019-03-28 14:56 | 1553799379.771 | -259 | 0.20949 | 0.000445 | 0 | 100.6669 | 100.744 | 126.3671 | 21.2612 | 0.00185 | 0.001748 | -151.6362 | 0.23007 | 8.404 | 0.002526 | 0.004033 | 0.003878 | 5.2e-05 | 0.001309 | 0.32211 | 1.3458 | 2.5474 | 100.7 | 21.16 | 25.692 | 66.3476 | 82.5825 | 63.8763 | 0.006415 | 1.0475 | 0.025289 | 24.1417 | -0.006845 | 0.13687 | -8e-06 | -0.000571 | -0.000571 | -9.8982
2019-03-28 14:56 | 1553799380.7718 | -258 | 0.20949 | 0.000444 | -1e-06 | 100.7217 | 100.7623 | 126.3671 | 21.2612 | 0.001824 | 0.001732 | -151.6362 | 0.22928 | 8.3985 | 0.002461 | 0.003515 | 0.003386 | 4.8e-05 | 0.001312 | 0.33301 | 1.3212 | 2.5474 | 100.7 | 21.16 | 25.691 | 65.7085 | 81.0259 | 63.4929 | 0.006415 | 1.0487 | 0.025304 | 24.128 | -0.003605 | 0.1364 | -6e-06 | -0.000407 | -0.000407 | -13.8534
2019-03-28 14:56 | 1553799381.7702 | -257 | 0.20949 | 0.000446 | -1e-06 | 100.6806 | 100.7684 | 126.3671 | 21.2612 | 0.001882 | 0.001768 | -151.6362 | 0.23001 | 8.4023 | 0.002489 | 0.003262 | 0.00319 | 4.8e-05 | 0.00132 | 0.32189 | 1.2971 | 2.5473 | 100.7 | 21.16 | 25.691 | 67.1082 | 81.6936 | 64.3667 | 0.006415 | 1.046 | 0.025271 | 24.1593 | -0.007799 | 0.13683 | -8e-06 | -0.00057 | -0.00057 | -9.9064
2019-03-28 14:56 | 1553799382.7704 | -256 | 0.20947 | 0.000447 | -1e-06 | 100.7065 | 100.8048 | 126.3671 | 21.2612 | 0.001857 | 0.0017 | -151.6362 | 0.23052 | 8.4011 | 0.002505 | 0.003868 | 0.003612 | 4.9e-05 | 0.00133 | 0.32908 | 1.2948 | 2.5472 | 100.7 | 21.16 | 25.693 | 66.5169 | 82.0717 | 62.7399 | 0.006416 | 1.0511 | 0.025332 | 24.101 | -0.008723 | 0.13714 | -1e-05 | -0.000685 | -0.000685 | -8.2642
2019-03-28 14:56 | 1553799383.7702 | -255 | 0.2095 | 0.000445 | 0 | 100.677 | 100.8327 | 126.3671 | 21.2612 | 0.001859 | 0.001812 | -151.6362 | 0.23049 | 8.402 | 0.002504 | 0.004072 | 0.003887 | 4.7e-05 | 0.001334 | 0.31526 | 1.3194 | 2.5474 | 100.7 | 21.16 | 25.693 | 66.5489 | 82.0693 | 65.4189 | 0.006416 | 1.0428 | 0.025232 | 24.1969 | -0.013814 | 0.13712 | -8e-06 | -0.000583 | -0.000583 | -9.7212
2019-03-28 14:56 | 1553799384.7703 | -254 | 0.20947 | 0.000443 | -1e-06 | 100.7012 | 100.6319 | 126.3671 | 21.2612 | 0.001789 | 0.001752 | -151.6362 | 0.2296 | 8.4053 | 0.002475 | 0.003866 | 0.003791 | 4.9e-05 | 0.001331 | 0.32947 | 1.3475 | 2.5473 | 100.7 | 21.16 | 25.697 | 64.8771 | 81.3715 | 63.9856 | 0.006417 | 1.0472 | 0.025285 | 24.1456 | 0.006157 | 0.13659 | -5e-06 | -0.000354 | -0.000354 | -15.9253
2019-03-28 14:56 | 1553799385.7703 | -253 | 0.20946 | 0.000449 | 0 | 100.6976 | 100.7701 | 126.3671 | 21.2612 | 0.001908 | 0.001772 | -151.6362 | 0.23004 | 8.4036 | 0.002506 | 0.003429 | 0.003354 | 4.9e-05 | 0.001331 | 0.34506 | 1.3526 | 2.5475 | 100.7 | 21.16 | 25.697 | 67.7293 | 82.0967 | 64.4618 | 0.006417 | 1.0457 | 0.025267 | 24.1627 | -0.006442 | 0.13685 | -1.3e-05 | -0.00098 | -0.00098 | -5.7691
2019-03-28 14:56 | 1553799386.7714 | -252 | 0.20951 | 0.000446 | -1e-06 | 100.7289 | 100.8415 | 126.3671 | 21.2612 | 0.001909 | 0.001734 | -151.6362 | 0.23029 | 8.4013 | 0.002453 | 0.00408 | 0.003878 | 5e-05 | 0.00133 | 0.35513 | 1.3135 | 2.5474 | 100.7 | 21.15 | 25.708 | 67.7536 | 80.8208 | 63.5398 | 0.006415 | 1.0486 | 0.025302 | 24.1297 | -0.009994 | 0.137 | -9e-06 | -0.000614 | -0.000614 | -9.2174
2019-03-28 14:56 | 1553799387.7717 | -251 | 0.20954 | 0.000443 | 0 | 100.6934 | 100.6641 | 126.3671 | 21.2612 | 0.001903 | 0.001766 | -151.6362 | 0.22945 | 8.4035 | 0.002547 | 0.003611 | 0.003461 | 5.3e-05 | 0.001324 | 0.34324 | 1.3298 | 2.5473 | 100.7 | 21.15 | 25.708 | 67.6201 | 83.0958 | 64.3244 | 0.006415 | 1.0461 | 0.025272 | 24.1578 | 0.002601 | 0.1365 | -6e-06 | -0.000464 | -0.000464 | -12.1454
2019-03-28 14:56 | 1553799388.7705 | -250 | 0.20934 | 0.000449 | 0 | 100.731 | 100.799 | 126.454 | 21.2612 | 0.001852 | 0.001816 | -151.6458 | 0.23074 | 8.4031 | 0.002487 | 0.003754 | 0.00362 | 5.1e-05 | 0.00132 | 0.33124 | 1.3062 | 2.5473 | 100.7 | 21.15 | 25.725 | 66.4014 | 81.6404 | 65.5197 | 0.00642 | 1.0424 | 0.025236 | 24.2088 | -0.006042 | 0.13727 | -1.2e-05 | -0.00092 | -0.00092 | -6.1632
2019-03-28 14:56 | 1553799389.7706 | -249 | 0.20948 | 0.000447 | 0 | 100.6791 | 100.7667 | 126.2624 | 21.2612 | 0.001841 | 0.001746 | -151.6556 | 0.23043 | 8.4059 | 0.002539 | 0.003545 | 0.003342 | 5.2e-05 | 0.001322 | 0.34164 | 1.3578 | 2.5475 | 100.7 | 21.15 | 25.731 | 66.1251 | 82.9071 | 63.8388 | 0.006421 | 1.0476 | 0.02528 | 24.1304 | -0.00778 | 0.13708 | -1e-05 | -0.000695 | -0.000695 | -8.1415
2019-03-28 14:56 | 1553799390.771 | -248 | 0.20954 | 0.000448 | 0 | 100.6952 | 100.5847 | 126.0068 | 21.2612 | 0.001914 | 0.001779 | -151.6528 | 0.22997 | 8.4012 | 0.002542 | 0.003956 | 0.003761 | 5.1e-05 | 0.001316 | 0.31813 | 1.3297 | 2.5474 | 100.7 | 21.15 | 25.745 | 67.8704 | 82.9758 | 64.6336 | 0.006425 | 1.0452 | 0.025225 | 24.1344 | 0.009826 | 0.13681 | -1.1e-05 | -0.000818 | -0.000818 | -6.9082
2019-03-28 14:56 | 1553799391.7705 | -247 | 0.20954 | 0.000446 | 0 | 100.6919 | 100.5959 | 125.8827 | 21.2612 | 0.001842 | 0.00172 | -151.6629 | 0.22893 | 8.4025 | 0.002545 | 0.003975 | 0.003773 | 4.9e-05 | 0.001311 | 0.32925 | 1.2893 | 2.5474 | 100.7 | 21.15 | 25.752 | 66.1556 | 83.0444 | 63.2132 | 0.006426 | 1.0496 | 0.025266 | 24.0717 | 0.008539 | 0.13619 | -9e-06 | -0.000731 | -0.000731 | -7.6952
2019-03-28 14:57 | 1553799392.7706 | -246 | 0.20945 | 0.000445 | 0 | 100.683 | 100.7827 | 125.8728 | 21.2612 | 0.001875 | 0.001789 | -151.6576 | 0.23009 | 8.4017 | 0.002513 | 0.003552 | 0.003397 | 5.2e-05 | 0.001311 | 0.31902 | 1.3011 | 2.5473 | 100.7 | 21.15 | 25.759 | 66.9521 | 82.283 | 64.8749 | 0.006428 | 1.0444 | 0.025202 | 24.1301 | -0.008854 | 0.13688 | -8e-06 | -0.000575 | -0.000575 | -9.8327
2019-03-28 14:57 | 1553799393.771 | -245 | 0.20973 | 0.00045 | 0 | 100.6916 | 100.727 | 125.7007 | 21.2612 | 0.00185 | 0.001781 | -151.6539 | 0.22975 | 8.4028 | 0.002493 | 0.003448 | 0.003332 | 4.9e-05 | 0.001306 | 0.33019 | 1.3035 | 2.5473 | 100.7 | 21.15 | 25.774 | 66.3327 | 81.7909 | 64.6812 | 0.006432 | 1.045 | 0.025192 | 24.1067 | -0.00314 | 0.13668 | -1.4e-05 | -0.00105 | -0.00105 | -5.3772
2019-03-28 14:57 | 1553799394.7706 | -244 | 0.20935 | 0.000448 | 0 | 100.6094 | 100.7443 | 125.5432 | 21.2612 | 0.001872 | 0.001787 | -151.656 | 0.22967 | 8.4027 | 0.002464 | 0.003903 | 0.00376 | 5.4e-05 | 0.001305 | 0.33107 | 1.3086 | 2.5473 | 100.7 | 21.15 | 25.781 | 66.8624 | 81.096 | 64.8141 | 0.006434 | 1.0446 | 0.025172 | 24.0964 | -0.011984 | 0.13663 | -1.2e-05 | -0.000869 | -0.000869 | -6.4906
2019-03-28 14:57 | 1553799395.7716 | -243 | 0.2093 | 0.000445 | 0 | 100.6681 | 100.655 | 125.5669 | 21.2612 | 0.001883 | 0.001714 | -151.6594 | 0.23075 | 8.4005 | 0.002512 | 0.00353 | 0.003334 | 4.9e-05 | 0.001305 | 0.32316 | 1.3261 | 2.5473 | 100.71 | 21.14 | 25.794 | 67.1411 | 82.2565 | 63.0594 | 0.006432 | 1.0501 | 0.02524 | 24.036 | 0.001167 | 0.13727 | -7e-06 | -0.00054 | -0.00054 | -10.4972
2019-03-28 14:57 | 1553799396.7716 | -242 | 0.20957 | 0.000446 | -1e-06 | 100.6997 | 100.6726 | 125.8024 | 21.2612 | 0.001848 | 0.0018 | -151.6451 | 0.22913 | 8.4029 | 0.002491 | 0.003924 | 0.003768 | 5e-05 | 0.001309 | 0.33843 | 1.4064 | 2.5475 | 100.71 | 21.14 | 25.8 | 66.2944 | 81.7412 | 65.1434 | 0.006434 | 1.0436 | 0.025185 | 24.133 | 0.002402 | 0.13631 | -9e-06 | -0.000652 | -0.000652 | -8.6389
2019-03-28 14:57 | 1553799397.7708 | -241 | 0.20942 | 0.000449 | 0 | 100.7232 | 100.5632 | 125.6695 | 21.2612 | 0.001912 | 0.001757 | -151.6332 | 0.23003 | 8.4026 | 0.002501 | 0.003927 | 0.003797 | 5e-05 | 0.00131 | 0.32244 | 1.3497 | 2.5474 | 100.71 | 21.14 | 25.805 | 67.8353 | 81.9976 | 64.1114 | 0.006435 | 1.0468 | 0.025211 | 24.0834 | 0.014217 | 0.13685 | -1.3e-05 | -0.000944 | -0.000944 | -5.9844
2019-03-28 14:57 | 1553799398.7712 | -240 | 0.20952 | 0.000447 | 0 | 100.6702 | 100.6281 | 125.9157 | 21.2612 | 0.001888 | 0.001791 | -151.651 | 0.22965 | 8.403 | 0.002501 | 0.00404 | 0.00382 | 5e-05 | 0.00131 | 0.35187 | 1.3786 | 2.5473 | 100.71 | 21.14 | 25.805 | 67.2597 | 81.9837 | 64.9108 | 0.006435 | 1.0443 | 0.025205 | 24.1355 | 0.003738 | 0.13662 | -1.1e-05 | -0.000808 | -0.000808 | -6.9836
2019-03-28 14:57 | 1553799399.7702 | -239 | 0.20948 | 0.000447 | 0 | 100.6836 | 100.7239 | 125.935 | 21.2612 | 0.001901 | 0.001766 | -151.6393 | 0.2302 | 8.4037 | 0.002534 | 0.003849 | 0.003668 | 4.9e-05 | 0.001308 | 0.30757 | 1.3288 | 2.5473 | 100.71 | 21.14 | 25.817 | 67.5677 | 82.7775 | 64.3292 | 0.006438 | 1.0461 | 0.025229 | 24.1166 | -0.003583 | 0.13695 | -1e-05 | -0.000764 | -0.000764 | -7.4039
2019-03-28 14:57 | 1553799400.7709 | -238 | 0.20951 | 0.000446 | 0 | 100.6583 | 100.9539 | 126.2324 | 21.2612 | 0.001855 | 0.00177 | -151.639 | 0.23042 | 8.4013 | 0.002485 | 0.003727 | 0.003558 | 4.6e-05 | 0.001302 | 0.3314 | 1.341 | 2.5474 | 100.71 | 21.14 | 25.824 | 66.4515 | 81.6123 | 64.4127 | 0.00644 | 1.0459 | 0.025256 | 24.1481 | -0.026233 | 0.13708 | -9e-06 | -0.000632 | -0.000632 | -8.9579
2019-03-28 14:57 | 1553799401.7712 | -237 | 0.20944 | 0.000447 | 0 | 100.6455 | 100.5503 | 126.2568 | 21.2612 | 0.001861 | 0.001724 | -151.6417 | 0.23032 | 8.4031 | 0.002494 | 0.003936 | 0.003749 | 5e-05 | 0.0013 | 0.32875 | 1.3425 | 2.5474 | 100.71 | 21.14 | 25.827 | 66.609 | 81.826 | 63.3064 | 0.00644 | 1.0493 | 0.0253 | 24.1108 | 0.008459 | 0.13702 | -1.1e-05 | -0.000832 | -0.000832 | -6.801
2019-03-28 14:57 | 1553799402.771 | -236 | 0.20954 | 0.000445 | 0 | 100.6294 | 100.6326 | 126.1505 | 21.2612 | 0.001883 | 0.001788 | -151.6509 | 0.22897 | 8.4037 | 0.002474 | 0.003629 | 0.003503 | 4.7e-05 | 0.0013 | 0.33732 | 1.3314 | 2.5474 | 100.71 | 21.14 | 25.827 | 67.1349 | 81.3489 | 64.8499 | 0.00644 | 1.0445 | 0.025231 | 24.1558 | -0.000281 | 0.13621 | -7e-06 | -0.000526 | -0.000526 | -10.6892
2019-03-28 14:57 | 1553799403.7703 | -235 | 0.20949 | 0.00045 | 0 | 100.5502 | 100.7518 | 126.2059 | 21.2612 | 0.001902 | 0.001789 | -151.6529 | 0.23003 | 8.4022 | 0.002487 | 0.003775 | 0.003637 | 4.7e-05 | 0.001297 | 0.32477 | 1.2997 | 2.5473 | 100.7 | 21.14 | 25.837 | 67.5958 | 81.6429 | 64.8826 | 0.006444 | 1.0444 | 0.025235 | 24.1623 | -0.017919 | 0.13685 | -1.4e-05 | -0.001016 | -0.001016 | -5.5644
2019-03-28 14:57 | 1553799404.771 | -234 | 0.2095 | 0.000445 | -1e-06 | 100.616 | 100.4352 | 126.2486 | 21.2612 | 0.001891 | 0.00179 | -151.6564 | 0.22928 | 8.4026 | 0.002483 | 0.003586 | 0.003384 | 4.7e-05 | 0.001295 | 0.32731 | 1.296 | 2.5473 | 100.7 | 21.14 | 25.837 | 67.3213 | 81.5515 | 64.8882 | 0.006444 | 1.0444 | 0.025239 | 24.1666 | 0.016088 | 0.1364 | -6e-06 | -0.000425 | -0.000425 | -13.2642
2019-03-28 14:57 | 1553799405.7709 | -233 | 0.20952 | 0.000446 | 0 | 100.5633 | 100.6244 | 126.1094 | 21.2612 | 0.001875 | 0.001787 | -151.6463 | 0.22954 | 8.4036 | 0.002518 | 0.003677 | 0.003572 | 4.9e-05 | 0.001292 | 0.32764 | 1.2919 | 2.5472 | 100.7 | 21.14 | 25.846 | 66.9368 | 82.3844 | 64.8256 | 0.006446 | 1.0446 | 0.025228 | 24.151 | -0.005436 | 0.13656 | -9e-06 | -0.000685 | -0.000685 | -8.2352
2019-03-28 14:57 | 1553799406.7713 | -232 | 0.20946 | 0.000444 | 0 | 100.6175 | 100.7943 | 125.9953 | 21.2612 | 0.0019 | 0.001709 | -151.6549 | 0.2305 | 8.4035 | 0.002517 | 0.003872 | 0.003638 | 4.9e-05 | 0.001292 | 0.35059 | 1.3125 | 2.5473 | 100.7 | 21.14 | 25.849 | 67.5398 | 82.3823 | 62.9462 | 0.006447 | 1.0504 | 0.025287 | 24.0729 | -0.015702 | 0.13712 | -6e-06 | -0.000465 | -0.000465 | -12.189
2019-03-28 14:57 | 1553799407.7714 | -231 | 0.20948 | 0.000447 | -1e-06 | 100.6005 | 100.6655 | 126.1791 | 21.2612 | 0.001853 | 0.001721 | -151.6725 | 0.22962 | 8.4015 | 0.002446 | 0.004425 | 0.00432 | 5e-05 | 0.001289 | 0.30845 | 1.283 | 2.5473 | 100.7 | 21.14 | 25.853 | 66.4077 | 80.6586 | 63.2445 | 0.006448 | 1.0495 | 0.025294 | 24.1011 | -0.005778 | 0.1366 | -9e-06 | -0.000669 | -0.000669 | -8.4277
2019-03-28 14:57 | 1553799408.771 | -230 | 0.20949 | 0.000447 | 0 | 100.5898 | 100.654 | 126.4124 | 21.2612 | 0.001856 | 0.001725 | -151.6434 | 0.22931 | 8.4026 | 0.002561 | 0.004179 | 0.003986 | 4.6e-05 | 0.001277 | 0.3235 | 1.3152 | 2.5473 | 100.7 | 21.14 | 25.859 | 66.4834 | 83.4329 | 63.3336 | 0.006449 | 1.0492 | 0.025314 | 24.1266 | -0.005705 | 0.13642 | -1e-05 | -0.000733 | -0.000733 | -7.6827
2019-03-28 14:57 | 1553799409.7701 | -229 | 0.2095 | 0.000443 | 0 | 100.5925 | 100.6268 | 126.7434 | 21.2612 | 0.001852 | 0.001751 | -151.6408 | 0.22939 | 8.4007 | 0.002542 | 0.003704 | 0.003574 | 4.8e-05 | 0.001275 | 0.31116 | 1.3496 | 2.5474 | 100.7 | 21.14 | 25.859 | 66.3872 | 82.9727 | 63.9622 | 0.006449 | 1.0473 | 0.025324 | 24.1807 | -0.003051 | 0.13646 | -4e-06 | -0.000272 | -0.000272 | -20.7274
2019-03-28 14:57 | 1553799410.7711 | -228 | 0.20948 | 0.000448 | 0 | 100.6318 | 100.6767 | 126.7413 | 21.2612 | 0.001891 | 0.00176 | -151.6458 | 0.22939 | 8.4025 | 0.002582 | 0.003759 | 0.003556 | 4.3e-05 | 0.001271 | 0.33052 | 1.3137 | 2.5474 | 100.7 | 21.14 | 25.866 | 67.3205 | 83.9345 | 64.1793 | 0.006451 | 1.0466 | 0.025315 | 24.1883 | -0.003994 | 0.13646 | -1.1e-05 | -0.000825 | -0.000825 | -6.8285
2019-03-28 14:57 | 1553799411.7703 | -227 | 0.20949 | 0.000443 | -1e-06 | 100.5668 | 100.5931 | 126.523 | 21.2612 | 0.001869 | 0.001789 | -151.6562 | 0.2291 | 8.4047 | 0.002523 | 0.003604 | 0.003481 | 4.7e-05 | 0.00126 | 0.30397 | 1.3271 | 2.5474 | 100.7 | 21.14 | 25.871 | 66.7994 | 82.5227 | 64.8781 | 0.006452 | 1.0444 | 0.025267 | 24.1925 | -0.002339 | 0.13629 | -5e-06 | -0.000333 | -0.000333 | -16.9102
2019-03-28 14:57 | 1553799412.7705 | -226 | 0.20948 | 0.000446 | 0 | 100.5954 | 100.6067 | 126.6277 | 21.2612 | 0.001871 | 0.001809 | -151.6511 | 0.22959 | 8.4038 | 0.00248 | 0.004018 | 0.003798 | 4.7e-05 | 0.001267 | 0.30901 | 1.3211 | 2.5474 | 100.7 | 21.14 | 25.873 | 66.8516 | 81.4861 | 65.3619 | 0.006453 | 1.0429 | 0.02526 | 24.2198 | -0.001005 | 0.13658 | -8e-06 | -0.000623 | -0.000623 | -9.0616
2019-03-28 14:57 | 1553799413.7703 | -225 | 0.20948 | 0.00045 | 0 | 100.5954 | 100.7127 | 126.5781 | 21.2612 | 0.001892 | 0.001779 | -151.6528 | 0.22924 | 8.4007 | 0.002495 | 0.004017 | 0.003846 | 4.8e-05 | 0.00127 | 0.29999 | 1.3141 | 2.5474 | 100.7 | 21.14 | 25.876 | 67.3439 | 81.8527 | 64.6322 | 0.006453 | 1.0452 | 0.025282 | 24.189 | -0.010423 | 0.13637 | -1.3e-05 | -0.000985 | -0.000985 | -5.7184
2019-03-28 14:57 | 1553799414.7696 | -224 | 0.20949 | 0.000453 | 0 | 100.5692 | 100.7212 | 126.3401 | 21.2612 | 0.001884 | 0.001725 | -151.6483 | 0.22934 | 8.4023 | 0.002519 | 0.002928 | 0.002725 | 4.6e-05 | 0.001266 | 0.31636 | 1.3253 | 2.5473 | 100.7 | 21.14 | 25.876 | 67.1575 | 82.4195 | 63.3437 | 0.006453 | 1.0492 | 0.025307 | 24.1201 | -0.013508 | 0.13643 | -1.8e-05 | -0.001345 | -0.001345 | -4.1889
2019-03-28 14:57 | 1553799415.7703 | -223 | 0.20952 | 0.000444 | 0 | 100.599 | 100.7552 | 126.787 | 21.2612 | 0.001891 | 0.001811 | -151.6494 | 0.23005 | 8.4015 | 0.002495 | 0.003874 | 0.003663 | 4.6e-05 | 0.001269 | 0.30735 | 1.3092 | 2.5474 | 100.7 | 21.14 | 25.872 | 67.3338 | 81.8464 | 65.404 | 0.006452 | 1.0428 | 0.025274 | 24.2366 | -0.013876 | 0.13685 | -7e-06 | -0.000518 | -0.000518 | -10.9109
2019-03-28 14:57 | 1553799416.7696 | -222 | 0.20948 | 0.000448 | 0 | 100.6029 | 100.6129 | 127.0584 | 21.2612 | 0.001817 | 0.001775 | -151.6355 | 0.22939 | 8.4026 | 0.002449 | 0.003594 | 0.003418 | 4.9e-05 | 0.001272 | 0.3194 | 1.3335 | 2.5475 | 100.7 | 21.14 | 25.872 | 65.536 | 80.7475 | 64.5376 | 0.006452 | 1.0455 | 0.025333 | 24.2314 | -0.000886 | 0.13646 | -1.1e-05 | -0.000813 | -0.000813 | -6.9298
2019-03-28 14:57 | 1553799417.7704 | -221 | 0.20951 | 0.000448 | 0 | 100.6291 | 100.5965 | 126.7516 | 21.2612 | 0.001899 | 0.001788 | -151.6262 | 0.22887 | 8.4035 | 0.002416 | 0.003754 | 0.003567 | 5e-05 | 0.001273 | 0.34595 | 1.3434 | 2.5473 | 100.7 | 21.15 | 25.863 | 67.5249 | 79.936 | 64.8545 | 0.006454 | 1.0445 | 0.025291 | 24.2135 | 0.002894 | 0.13615 | -1.1e-05 | -0.000839 | -0.000839 | -6.7054
2019-03-28 14:57 | 1553799418.7694 | -220 | 0.20951 | 0.000449 | 0 | 100.6032 | 100.7389 | 126.2347 | 21.2612 | 0.001874 | 0.001817 | -151.6472 | 0.22942 | 8.402 | 0.002492 | 0.003436 | 0.003309 | 4.6e-05 | 0.001274 | 0.33411 | 1.3154 | 2.5474 | 100.7 | 21.15 | 25.852 | 66.9242 | 81.7597 | 65.5594 | 0.006451 | 1.0423 | 0.025213 | 24.1893 | -0.012058 | 0.13648 | -1.3e-05 | -0.000953 | -0.000953 | -5.9178
2019-03-28 14:57 | 1553799419.7708 | -219 | 0.20949 | 0.000448 | -1e-06 | 100.6017 | 100.6699 | 126.1756 | 21.2612 | 0.001955 | 0.001758 | -151.6698 | 0.22908 | 8.402 | 0.002456 | 0.00383 | 0.003645 | 4.8e-05 | 0.001275 | 0.33948 | 1.303 | 2.5475 | 100.7 | 21.14 | 25.847 | 68.8556 | 80.9059 | 64.1167 | 0.006446 | 1.0468 | 0.025261 | 24.132 | -0.006065 | 0.13628 | -1.1e-05 | -0.000766 | -0.000766 | -7.3524
2019-03-28 14:57 | 1553799420.7695 | -218 | 0.20946 | 0.000446 | 0 | 100.6607 | 100.8867 | 126.2911 | 21.2612 | 0.00192 | 0.001754 | -151.659 | 0.2291 | 8.4012 | 0.002497 | 0.004489 | 0.004282 | 4.8e-05 | 0.001277 | 0.33212 | 1.3002 | 2.5473 | 100.7 | 21.14 | 25.84 | 68.0288 | 81.8969 | 64.02 | 0.006444 | 1.0471 | 0.025276 | 24.1396 | -0.02006 | 0.13629 | -8e-06 | -0.000596 | -0.000596 | -9.4473
2019-03-28 14:57 | 1553799421.7695 | -217 | 0.20951 | 0.000445 | -1e-06 | 100.6493 | 100.5622 | 126.7877 | 21.2612 | 0.001913 | 0.001729 | -151.6579 | 0.23105 | 8.4043 | 0.002492 | 0.003604 | 0.003465 | 4.7e-05 | 0.001276 | 0.32764 | 1.2947 | 2.5474 | 100.7 | 21.14 | 25.824 | 67.8557 | 81.7777 | 63.4359 | 0.00644 | 1.0489 | 0.025348 | 24.1661 | 0.007745 | 0.13745 | -7e-06 | -0.000487 | -0.000487 | -11.6549
2019-03-28 14:57 | 1553799422.7699 | -216 | 0.2095 | 0.000446 | -1e-06 | 100.6103 | 100.6662 | 126.8338 | 21.2612 | 0.00189 | 0.00179 | -151.6662 | 0.23067 | 8.4003 | 0.002517 | 0.003867 | 0.003624 | 4.8e-05 | 0.001279 | 0.30973 | 1.3011 | 2.5473 | 100.7 | 21.14 | 25.816 | 67.3112 | 82.3611 | 64.8865 | 0.006438 | 1.0444 | 0.025298 | 24.2225 | -0.004965 | 0.13723 | -8e-06 | -0.000574 | -0.000574 | -9.8762
2019-03-28 14:57 | 1553799423.7705 | -215 | 0.20951 | 0.000446 | 0 | 100.6115 | 100.655 | 126.9758 | 21.2612 | 0.001869 | 0.001807 | -151.6608 | 0.22995 | 8.4028 | 0.002459 | 0.003402 | 0.003239 | 5.1e-05 | 0.001285 | 0.3194 | 1.3511 | 2.5473 | 100.7 | 21.14 | 25.807 | 66.7915 | 80.9876 | 65.3034 | 0.006436 | 1.0431 | 0.025297 | 24.251 | -0.003863 | 0.1368 | -9e-06 | -0.000715 | -0.000715 | -7.9047
2019-03-28 14:57 | 1553799424.7706 | -214 | 0.20947 | 0.000446 | 0 | 100.6648 | 100.8194 | 127.1377 | 21.2612 | 0.00193 | 0.001767 | -151.6628 | 0.23027 | 8.4029 | 0.002461 | 0.003853 | 0.003703 | 4.8e-05 | 0.001288 | 0.32278 | 1.3357 | 2.5473 | 100.7 | 21.14 | 25.807 | 68.2721 | 81.0283 | 64.3369 | 0.006436 | 1.0461 | 0.025349 | 24.2318 | -0.013724 | 0.13699 | -9e-06 | -0.000722 | -0.000722 | -7.8409
2019-03-28 14:57 | 1553799425.7696 | -213 | 0.2095 | 0.000449 | 0 | 100.7164 | 100.7141 | 126.999 | 21.2612 | 0.001897 | 0.001786 | -151.6585 | 0.23014 | 8.4007 | 0.002466 | 0.003242 | 0.003049 | 4.4e-05 | 0.001293 | 0.31564 | 1.3117 | 2.5474 | 100.7 | 21.15 | 25.783 | 67.4679 | 81.1468 | 64.8023 | 0.006434 | 1.0447 | 0.025318 | 24.2352 | 0.000202 | 0.13691 | -1.2e-05 | -0.00086 | -0.00086 | -6.5721
2019-03-28 14:57 | 1553799426.7706 | -212 | 0.20949 | 0.000448 | 0 | 100.6833 | 100.7175 | 127.0178 | 21.2612 | 0.001901 | 0.001816 | -151.6801 | 0.22974 | 8.4012 | 0.002422 | 0.003117 | 0.002957 | 4.4e-05 | 0.001292 | 0.30486 | 1.3348 | 2.5474 | 100.7 | 21.15 | 25.783 | 67.5725 | 80.0878 | 65.5124 | 0.006434 | 1.0425 | 0.025293 | 24.2625 | -0.003036 | 0.13667 | -1.1e-05 | -0.000795 | -0.000795 | -7.1013
2019-03-28 14:57 | 1553799427.7697 | -211 | 0.20948 | 0.000446 | 0 | 100.6377 | 100.9128 | 127.0542 | 21.2612 | 0.001901 | 0.001737 | -151.6689 | 0.22922 | 8.4023 | 0.002408 | 0.003089 | 0.002882 | 4.7e-05 | 0.001294 | 0.33069 | 1.333 | 2.5473 | 100.7 | 21.15 | 25.763 | 67.5732 | 79.7495 | 63.6265 | 0.006429 | 1.0483 | 0.025367 | 24.1983 | -0.024417 | 0.13636 | -9e-06 | -0.000728 | -0.000728 | -7.7316
2019-03-28 14:57 | 1553799428.7688 | -210 | 0.20949 | 0.000448 | 0 | 100.7092 | 100.7698 | 127.1664 | 21.2612 | 0.001858 | 0.00177 | -151.6615 | 0.22987 | 8.4029 | 0.00239 | 0.003844 | 0.003611 | 4.9e-05 | 0.001295 | 0.3449 | 1.3321 | 2.5473 | 100.7 | 21.15 | 25.763 | 66.545 | 79.3145 | 64.4283 | 0.006429 | 1.0458 | 0.025348 | 24.2378 | -0.00538 | 0.13675 | -1e-05 | -0.000762 | -0.000762 | -7.4158
2019-03-28 14:57 | 1553799429.7689 | -209 | 0.2095 | 0.00045 | 0 | 100.6493 | 100.5639 | 126.9298 | 21.2612 | 0.001846 | 0.001766 | -151.6366 | 0.23036 | 8.4034 | 0.002428 | 0.003931 | 0.003752 | 4.7e-05 | 0.001299 | 0.31332 | 1.3538 | 2.5474 | 100.7 | 21.15 | 25.749 | 66.2345 | 80.2275 | 64.3191 | 0.006426 | 1.0462 | 0.025329 | 24.2113 | 0.007594 | 0.13704 | -1.3e-05 | -0.000983 | -0.000983 | -5.7565
2019-03-28 14:57 | 1553799430.7698 | -208 | 0.20948 | 0.000446 | -1e-06 | 100.6237 | 100.7932 | 126.7018 | 21.2612 | 0.001847 | 0.001743 | -151.6281 | 0.22981 | 8.4023 | 0.002509 | 0.003263 | 0.003099 | 4.6e-05 | 0.001305 | 0.31094 | 1.3296 | 2.5474 | 100.71 | 21.15 | 25.746 | 66.2728 | 82.1887 | 63.7608 | 0.006424 | 1.0479 | 0.025327 | 24.1695 | -0.015055 | 0.13672 | -8e-06 | -0.000585 | -0.000585 | -9.6578
2019-03-28 14:57 | 1553799431.7691 | -207 | 0.20948 | 0.000448 | 0 | 100.6547 | 100.6523 | 126.4195 | 21.2612 | 0.001869 | 0.00172 | -151.6091 | 0.23036 | 8.4003 | 0.002525 | 0.003776 | 0.003614 | 4.9e-05 | 0.001313 | 0.33588 | 1.3066 | 2.5474 | 100.71 | 21.15 | 25.742 | 66.7985 | 82.5662 | 63.2243 | 0.006423 | 1.0496 | 0.025319 | 24.1234 | 0.000218 | 0.13704 | -1e-05 | -0.000751 | -0.000751 | -7.5397
2019-03-28 14:57 | 1553799432.7689 | -206 | 0.20949 | 0.000443 | -1e-06 | 100.6577 | 100.7022 | 126.5711 | 21.2612 | 0.001855 | 0.001676 | -151.6117 | 0.22934 | 8.4039 | 0.00245 | 0.003132 | 0.00292 | 5e-05 | 0.001313 | 0.31094 | 1.3072 | 2.5474 | 100.71 | 21.15 | 25.743 | 66.4543 | 80.7555 | 62.1517 | 0.006423 | 1.0529 | 0.025375 | 24.0993 | -0.003955 | 0.13644 | -4e-06 | -0.000238 | -0.000238 | -23.6548
2019-03-28 14:57 | 1553799433.7699 | -205 | 0.20951 | 0.000447 | 0 | 100.6863 | 100.6499 | 126.411 | 21.2612 | 0.001859 | 0.001788 | -151.6144 | 0.23039 | 8.4012 | 0.002501 | 0.004126 | 0.003924 | 4.9e-05 | 0.001314 | 0.31332 | 1.2995 | 2.5474 | 100.71 | 21.15 | 25.743 | 66.5481 | 81.9789 | 64.8398 | 0.006423 | 1.0445 | 0.025258 | 24.1804 | 0.003234 | 0.13706 | -1e-05 | -0.000715 | -0.000715 | -7.9182
2019-03-28 14:57 | 1553799434.77 | -204 | 0.2095 | 0.000444 | -1e-06 | 100.6827 | 100.8177 | 126.2425 | 21.2612 | 0.001816 | 0.001798 | -151.6195 | 0.23055 | 8.4013 | 0.002501 | 0.00376 | 0.003611 | 5.2e-05 | 0.001317 | 0.33196 | 1.3169 | 2.5474 | 100.71 | 21.15 | 25.747 | 65.525 | 81.9806 | 65.084 | 0.006424 | 1.0438 | 0.025232 | 24.173 | -0.011985 | 0.13715 | -5e-06 | -0.000346 | -0.000346 | -16.3852
2019-03-28 14:57 | 1553799435.7699 | -203 | 0.20947 | 0.000448 | 0 | 100.7065 | 100.8034 | 126.0704 | 21.2612 | 0.001873 | 0.001796 | -151.6361 | 0.23022 | 8.4056 | 0.002527 | 0.003363 | 0.003233 | 5.1e-05 | 0.001313 | 0.31598 | 1.3194 | 2.5475 | 100.71 | 21.15 | 25.749 | 66.8944 | 82.6224 | 65.0395 | 0.006425 | 1.0439 | 0.025216 | 24.155 | -0.008603 | 0.13696 | -1.1e-05 | -0.000872 | -0.000872 | -6.4904
2019-03-28 14:57 | 1553799436.7701 | -202 | 0.20951 | 0.000443 | 0 | 100.6955 | 100.797 | 125.908 | 21.2612 | 0.001838 | 0.001757 | -151.6393 | 0.23005 | 8.403 | 0.002578 | 0.00326 | 0.003135 | 5.2e-05 | 0.001317 | 0.33356 | 1.3195 | 2.5474 | 100.71 | 21.15 | 25.754 | 66.0472 | 83.8361 | 64.0965 | 0.006426 | 1.0468 | 0.025235 | 24.1057 | -0.009008 | 0.13686 | -4e-06 | -0.000245 | -0.000245 | -23.0442
2019-03-28 14:57 | 1553799437.7691 | -201 | 0.20949 | 0.000446 | -1e-06 | 100.764 | 100.9638 | 126.0139 | 21.2612 | 0.001876 | 0.001816 | -151.6558 | 0.23066 | 8.4031 | 0.002502 | 0.003924 | 0.00373 | 5.1e-05 | 0.001312 | 0.3282 | 1.3338 | 2.5474 | 100.71 | 21.15 | 25.754 | 66.9562 | 82.007 | 65.5197 | 0.006426 | 1.0424 | 0.025192 | 24.1667 | -0.017715 | 0.13722 | -8e-06 | -0.000572 | -0.000572 | -9.9068
2019-03-28 14:57 | 1553799438.77 | -200 | 0.20949 | 0.000448 | -1e-06 | 100.7268 | 100.9441 | 126.2256 | 21.2612 | 0.001855 | 0.00179 | -151.6735 | 0.23122 | 8.4017 | 0.002486 | 0.003984 | 0.003801 | 5.1e-05 | 0.001307 | 0.30596 | 1.3054 | 2.5474 | 100.71 | 21.15 | 25.759 | 66.4724 | 81.6154 | 64.8975 | 0.006427 | 1.0444 | 0.025237 | 24.1647 | -0.019275 | 0.13756 | -1e-05 | -0.000708 | -0.000708 | -8.0233
2019-03-28 14:57 | 1553799439.7701 | -199 | 0.20949 | 0.000444 | 0 | 100.6717 | 100.9913 | 126.2279 | 21.2612 | 0.001926 | 0.001768 | -151.6864 | 0.2305 | 8.404 | 0.002479 | 0.004089 | 0.003977 | 5.2e-05 | 0.001304 | 0.31277 | 1.2541 | 2.5472 | 100.71 | 21.15 | 25.759 | 68.1801 | 81.4518 | 64.3682 | 0.006427 | 1.046 | 0.025257 | 24.146 | -0.028353 | 0.13712 | -5e-06 | -0.000384 | -0.000384 | -14.742
2019-03-28 14:57 | 1553799440.7691 | -198 | 0.20949 | 0.000447 | 0 | 100.7444 | 100.8717 | 126.2326 | 21.2612 | 0.001912 | 0.001774 | -151.6928 | 0.23089 | 8.4012 | 0.002597 | 0.004177 | 0.004 | 4.6e-05 | 0.001298 | 0.34346 | 1.2806 | 2.5473 | 100.71 | 21.15 | 25.764 | 67.8221 | 84.2894 | 64.5032 | 0.006429 | 1.0456 | 0.025252 | 24.1513 | -0.0113 | 0.13735 | -9e-06 | -0.000624 | -0.000624 | -9.0963
2019-03-28 14:57 | 1553799441.7701 | -197 | 0.20951 | 0.000444 | -1e-06 | 100.7375 | 100.9597 | 126.1613 | 21.2612 | 0.001851 | 0.001758 | -151.7118 | 0.23053 | 8.4017 | 0.002578 | 0.004631 | 0.004433 | 5e-05 | 0.001296 | 0.31564 | 1.2941 | 2.5474 | 100.71 | 21.15 | 25.764 | 66.3632 | 83.8455 | 64.13 | 0.006429 | 1.0467 | 0.025259 | 24.1311 | -0.019707 | 0.13714 | -5e-06 | -0.000304 | -0.000304 | -18.6039
2019-03-28 14:57 | 1553799442.7702 | -196 | 0.20947 | 0.000442 | -1e-06 | 100.7908 | 100.8751 | 125.8829 | 21.2612 | 0.001856 | 0.001796 | -151.6892 | 0.23023 | 8.401 | 0.002542 | 0.004165 | 0.003971 | 4.9e-05 | 0.001297 | 0.32693 | 1.2808 | 2.5474 | 100.71 | 21.16 | 25.764 | 66.4803 | 82.9773 | 65.0513 | 0.006433 | 1.0439 | 0.025197 | 24.1374 | -0.007477 | 0.13696 | -3e-06 | -0.000154 | -0.000154 | -36.7205
2019-03-28 14:57 | 1553799443.7692 | -195 | 0.20951 | 0.000441 | -1e-06 | 100.8191 | 100.9702 | 125.7932 | 21.2612 | 0.001895 | 0.001778 | -151.6962 | 0.23015 | 8.4035 | 0.002495 | 0.004024 | 0.003897 | 5.3e-05 | 0.001301 | 0.33892 | 1.2872 | 2.5474 | 100.71 | 21.16 | 25.762 | 67.4234 | 81.8371 | 64.6211 | 0.006432 | 1.0452 | 0.025204 | 24.1135 | -0.013397 | 0.13692 | -1e-06 | -5.3e-05 | -5.3e-05 | -106.8036
2019-03-28 14:57 | 1553799444.7692 | -194 | 0.20949 | 0.000446 | -1e-06 | 100.7896 | 100.8344 | 125.9782 | 21.2612 | 0.001894 | 0.001751 | -151.7087 | 0.23013 | 8.4009 | 0.002559 | 0.004258 | 0.004128 | 5.4e-05 | 0.001299 | 0.32676 | 1.2985 | 2.5474 | 100.71 | 21.16 | 25.759 | 67.3891 | 83.3846 | 63.9637 | 0.006431 | 1.0473 | 0.025247 | 24.1077 | -0.003967 | 0.1369 | -7e-06 | -0.000474 | -0.000474 | -11.9208
2019-03-28 14:57 | 1553799445.7686 | -193 | 0.20949 | 0.000444 | 0 | 100.7068 | 100.7885 | 126.161 | 21.2612 | 0.001894 | 0.001794 | -151.7031 | 0.23101 | 8.4045 | 0.002578 | 0.003897 | 0.003653 | 5e-05 | 0.001298 | 0.33124 | 1.3492 | 2.5474 | 100.71 | 21.16 | 25.757 | 67.4078 | 83.8337 | 64.9974 | 0.006431 | 1.0441 | 0.025227 | 24.1621 | -0.007249 | 0.13743 | -5e-06 | -0.000386 | -0.000386 | -14.6964
2019-03-28 14:57 | 1553799446.7693 | -192 | 0.20952 | 0.000442 | 0 | 100.8033 | 100.7549 | 126.1514 | 21.2612 | 0.001861 | 0.001779 | -151.6784 | 0.22999 | 8.4017 | 0.002558 | 0.004251 | 0.004095 | 5.5e-05 | 0.001295 | 0.33356 | 1.3509 | 2.5475 | 100.71 | 21.16 | 25.757 | 66.6073 | 83.3682 | 64.6406 | 0.006431 | 1.0452 | 0.025239 | 24.1485 | 0.004303 | 0.13682 | -3e-06 | -0.000173 | -0.000173 | -32.611
2019-03-28 14:57 | 1553799447.7697 | -191 | 0.20952 | 0.00044 | 0 | 100.8105 | 100.9172 | 126.2063 | 21.2612 | 0.001877 | 0.001723 | -151.6732 | 0.23031 | 8.4032 | 0.002594 | 0.00373 | 0.003574 | 5.6e-05 | 0.0013 | 0.32405 | 1.3453 | 2.5474 | 100.71 | 21.16 | 25.753 | 66.9897 | 84.223 | 63.2859 | 0.00643 | 1.0494 | 0.025295 | 24.1052 | -0.009467 | 0.13701 | 0 | 1e-06 | 1e-06 | 7903.0574
2019-03-28 14:57 | 1553799448.7694 | -190 | 0.20952 | 0.000445 | 0 | 100.8087 | 100.9692 | 126.1153 | 21.2612 | 0.001893 | 0.001729 | -151.6447 | 0.23023 | 8.4027 | 0.002582 | 0.004801 | 0.00469 | 5.3e-05 | 0.001297 | 0.32997 | 1.3516 | 2.5474 | 100.71 | 21.16 | 25.737 | 67.3689 | 83.9321 | 63.4233 | 0.006426 | 1.0489 | 0.025281 | 24.1015 | -0.014232 | 0.13696 | -6e-06 | -0.00044 | -0.00044 | -12.8501
2019-03-28 14:57 | 1553799449.7683 | -189 | 0.20953 | 0.00044 | -1e-06 | 100.7914 | 100.9451 | 126.2387 | 21.2612 | 0.001889 | 0.001769 | -151.6388 | 0.23057 | 8.4027 | 0.0025 | 0.004379 | 0.004222 | 5e-05 | 0.001296 | 0.31686 | 1.3499 | 2.5474 | 100.7 | 21.16 | 25.732 | 67.2854 | 81.95 | 64.3877 | 0.006425 | 1.0459 | 0.025257 | 24.1478 | -0.013628 | 0.13717 | 0 | 2.7e-05 | 2.7e-05 | 206.4624
2019-03-28 14:57 | 1553799450.7684 | -188 | 0.20952 | 0.000437 | -1e-06 | 100.8248 | 101.0181 | 126.0031 | 21.2612 | 0.001905 | 0.001725 | -151.6457 | 0.23068 | 8.4029 | 0.002535 | 0.004402 | 0.004247 | 5.3e-05 | 0.001296 | 0.33516 | 1.3475 | 2.5474 | 100.7 | 21.16 | 25.726 | 67.6754 | 82.7924 | 63.3328 | 0.006424 | 1.0492 | 0.025273 | 24.0875 | -0.017137 | 0.13723 | 3e-06 | 0.000304 | 0.000304 | 18.6396
2019-03-28 14:58 | 1553799451.7695 | -187 | 0.20952 | 0.000441 | 0 | 100.7524 | 100.9339 | 125.9106 | 21.2612 | 0.001861 | 0.001758 | -151.6642 | 0.23028 | 8.4049 | 0.002555 | 0.003863 | 0.003685 | 5.7e-05 | 0.001295 | 0.3298 | 1.356 | 2.5476 | 100.7 | 21.16 | 25.726 | 66.6136 | 83.2838 | 64.1386 | 0.006424 | 1.0467 | 0.025234 | 24.1075 | -0.016098 | 0.13699 | -1e-06 | -6.8e-05 | -6.8e-05 | -83.088
2019-03-28 14:58 | 1553799452.7699 | -186 | 0.20952 | 0.000439 | 0 | 100.8024 | 100.9101 | 125.8153 | 21.2612 | 0.00186 | 0.001785 | -151.6695 | 0.23036 | 8.4028 | 0.00255 | 0.004066 | 0.003907 | 5.3e-05 | 0.001294 | 0.32012 | 1.3413 | 2.5475 | 100.7 | 21.16 | 25.713 | 66.5715 | 83.166 | 64.7703 | 0.006421 | 1.0448 | 0.025201 | 24.1209 | -0.009549 | 0.13704 | 1e-06 | 6.6e-05 | 6.6e-05 | 86.0941
2019-03-28 14:58 | 1553799453.7696 | -185 | 0.2095 | 0.000448 | 0 | 100.7965 | 100.9529 | 125.7756 | 21.2612 | 0.00188 | 0.001752 | -151.6948 | 0.23048 | 8.403 | 0.002612 | 0.003945 | 0.00377 | 5.5e-05 | 0.001296 | 0.33356 | 1.3251 | 2.5474 | 100.7 | 21.16 | 25.713 | 67.0592 | 84.6569 | 63.9848 | 0.006421 | 1.0472 | 0.025226 | 24.089 | -0.013871 | 0.13711 | -1e-05 | -0.000698 | -0.000698 | -8.115
2019-03-28 14:58 | 1553799454.7692 | -184 | 0.2095 | 0.000439 | -1e-06 | 100.7881 | 100.9516 | 125.7465 | 21.2612 | 0.001844 | 0.001731 | -151.687 | 0.23072 | 8.4039 | 0.002591 | 0.004165 | 0.004054 | 5.7e-05 | 0.001302 | 0.323 | 1.3299 | 2.5474 | 100.7 | 21.16 | 25.704 | 66.2025 | 84.1538 | 63.4835 | 0.006418 | 1.0488 | 0.025242 | 24.0683 | -0.01449 | 0.13725 | 1e-06 | 0.000138 | 0.000138 | 40.9947
2019-03-28 14:58 | 1553799455.77 | -183 | 0.20951 | 0.000439 | -1e-06 | 100.7831 | 100.994 | 125.9127 | 21.2612 | 0.001821 | 0.00176 | -151.678 | 0.23078 | 8.4022 | 0.002533 | 0.004296 | 0.004145 | 5.8e-05 | 0.001307 | 0.31957 | 1.3171 | 2.5474 | 100.7 | 21.16 | 25.704 | 65.6334 | 82.7486 | 64.1723 | 0.006418 | 1.0466 | 0.025233 | 24.1089 | -0.018702 | 0.13729 | 2e-06 | 0.000221 | 0.000221 | 25.6912
2019-03-28 14:58 | 1553799456.7698 | -182 | 0.20947 | 0.000443 | -1e-06 | 100.8072 | 101.0898 | 125.8773 | 21.2612 | 0.001855 | 0.001768 | -151.7056 | 0.23083 | 8.4014 | 0.002564 | 0.004012 | 0.003786 | 5.4e-05 | 0.00131 | 0.3413 | 1.3297 | 2.5473 | 100.7 | 21.16 | 25.702 | 66.4668 | 83.5055 | 64.3699 | 0.006418 | 1.046 | 0.025222 | 24.1125 | -0.025041 | 0.13732 | -3e-06 | -0.000169 | -0.000169 | -33.4663
2019-03-28 14:58 | 1553799457.7697 | -181 | 0.20951 | 0.00044 | -1e-06 | 100.8665 | 101.065 | 126.0535 | 21.2612 | 0.001906 | 0.001765 | -151.6971 | 0.23056 | 8.4018 | 0.002534 | 0.003868 | 0.003699 | 5.4e-05 | 0.001308 | 0.3449 | 1.289 | 2.5473 | 100.7 | 21.16 | 25.7 | 67.6966 | 82.7698 | 64.2871 | 0.006417 | 1.0463 | 0.025242 | 24.1265 | -0.017589 | 0.13716 | 1e-06 | 0.000143 | 0.000143 | 39.4915
2019-03-28 14:58 | 1553799458.7692 | -180 | 0.20951 | 0.000437 | -1e-06 | 100.8727 | 101.1948 | 126.1035 | 21.2612 | 0.001855 | 0.001773 | -151.6875 | 0.22914 | 8.4002 | 0.002563 | 0.003799 | 0.003609 | 5.2e-05 | 0.001297 | 0.35187 | 1.2977 | 2.5473 | 100.7 | 21.16 | 25.703 | 66.4668 | 83.4819 | 64.4837 | 0.006418 | 1.0456 | 0.02524 | 24.1383 | -0.028512 | 0.13632 | 4e-06 | 0.000318 | 0.000318 | 17.6945
2019-03-28 14:58 | 1553799459.768 | -179 | 0.2095 | 0.000438 | -1e-06 | 100.8799 | 100.9753 | 126.1115 | 21.2612 | 0.001907 | 0.001749 | -151.6766 | 0.23015 | 8.4011 | 0.002557 | 0.003877 | 0.00373 | 5.4e-05 | 0.001293 | 0.31564 | 1.3062 | 2.5473 | 100.7 | 21.16 | 25.706 | 67.706 | 83.3245 | 63.9076 | 0.006419 | 1.0474 | 0.025262 | 24.1184 | -0.008459 | 0.13692 | 3e-06 | 0.000288 | 0.000288 | 19.6546
2019-03-28 14:58 | 1553799460.7692 | -178 | 0.20951 | 0.000441 | -1e-06 | 100.8778 | 101.0453 | 126.2462 | 21.2612 | 0.001854 | 0.001739 | -151.6691 | 0.23059 | 8.4035 | 0.00261 | 0.003641 | 0.003458 | 5.5e-05 | 0.001282 | 0.30237 | 1.3025 | 2.5474 | 100.7 | 21.16 | 25.71 | 66.4459 | 84.6165 | 63.671 | 0.00642 | 1.0482 | 0.025285 | 24.1228 | -0.014841 | 0.13718 | -1e-06 | 1.6e-05 | 1.6e-05 | 346.7495
2019-03-28 14:58 | 1553799461.7689 | -177 | 0.20951 | 0.000439 | -1e-06 | 100.8522 | 101.0334 | 126.2803 | 21.2612 | 0.001845 | 0.001762 | -151.6419 | 0.23005 | 8.404 | 0.002588 | 0.003607 | 0.003443 | 5.6e-05 | 0.001277 | 0.32477 | 1.331 | 2.5475 | 100.7 | 21.16 | 25.717 | 66.2189 | 84.0687 | 64.2269 | 0.006422 | 1.0464 | 0.025267 | 24.146 | -0.01606 | 0.13685 | 1e-06 | 0.000135 | 0.000135 | 41.8988
2019-03-28 14:58 | 1553799462.7688 | -176 | 0.2095 | 0.000441 | -1e-06 | 100.879 | 100.9594 | 126.6119 | 21.2612 | 0.001867 | 0.001777 | -151.6405 | 0.23027 | 8.4017 | 0.00264 | 0.003482 | 0.00332 | 5.3e-05 | 0.001271 | 0.32173 | 1.3367 | 2.5473 | 100.7 | 21.16 | 25.72 | 66.7571 | 85.3328 | 64.5828 | 0.006422 | 1.0453 | 0.025287 | 24.1904 | -0.007124 | 0.13699 | -5.9e-05 | -0.004485 | -0.004485 | -1.2614
2019-03-28 14:58 | 1553799463.7689 | -175 | 0.20952 | 0.000441 | -1e-06 | 100.8254 | 101.1924 | 126.7 | 21.2612 | 0.001905 | 0.001772 | -151.5985 | 0.23126 | 8.4027 | 0.002575 | 0.003724 | 0.003515 | 5.5e-05 | 0.001271 | 0.36641 | 1.2931 | 2.5473 | 100.7 | 21.16 | 25.723 | 67.6528 | 83.7575 | 64.4767 | 0.006423 | 1.0457 | 0.0253 | 24.195 | -0.032502 | 0.13757 | -0.000152 | -0.011508 | -0.011508 | -0.49375
2019-03-28 14:58 | 1553799464.769 | -174 | 0.20948 | 0.00044 | 0 | 100.9383 | 100.9855 | 126.622 | 21.2612 | 0.001907 | 0.001815 | -151.6048 | 0.22928 | 8.4054 | 0.002635 | 0.003615 | 0.003525 | 5.5e-05 | 0.001275 | 0.34147 | 1.3356 | 2.5474 | 100.7 | 21.16 | 25.73 | 67.7011 | 85.2149 | 65.4978 | 0.006425 | 1.0425 | 0.025254 | 24.2241 | -0.004187 | 0.1364 | 3.4e-05 | 0.00261 | 0.00261 | 2.1585
2019-03-28 14:58 | 1553799465.7684 | -173 | 0.20951 | 0.000439 | -1e-06 | 100.9844 | 101.0083 | 126.2239 | 21.2612 | 0.001883 | 0.001759 | -151.6091 | 0.23045 | 8.4002 | 0.00255 | 0.00413 | 0.003929 | 5.3e-05 | 0.001274 | 0.33644 | 1.3223 | 2.5474 | 100.7 | 21.16 | 25.732 | 67.1325 | 83.1629 | 64.1528 | 0.006425 | 1.0467 | 0.025265 | 24.1379 | -0.002113 | 0.13709 | -0.000135 | -0.010219 | -0.010219 | -0.55407
2019-03-28 14:58 | 1553799466.7681 | -172 | 0.20947 | 0.000445 | -2e-06 | 100.9052 | 100.9584 | 126.289 | 21.2612 | 0.001915 | 0.001785 | -151.6052 | 0.23003 | 8.4037 | 0.002651 | 0.003166 | 0.002886 | 5.5e-05 | 0.001274 | 0.30807 | 1.3269 | 2.5474 | 100.7 | 21.16 | 25.736 | 67.9024 | 85.5887 | 64.7789 | 0.006426 | 1.0447 | 0.025248 | 24.1666 | -0.00471 | 0.13685 | 0.000118 | 0.009109 | 0.009109 | 0.62043
2019-03-28 14:58 | 1553799467.7681 | -171 | 0.20952 | 0.000441 | -1e-06 | 100.9457 | 101.1048 | 126.2298 | 21.2612 | 0.001914 | 0.001748 | -151.5904 | 0.2304 | 8.4013 | 0.002623 | 0.003694 | 0.003553 | 5.7e-05 | 0.001278 | 0.31564 | 1.3605 | 2.5475 | 100.7 | 21.16 | 25.739 | 67.8805 | 84.9193 | 63.8975 | 0.006427 | 1.0475 | 0.025275 | 24.1294 | -0.014083 | 0.13706 | -0.000161 | -0.012168 | -0.012168 | -0.4652
2019-03-28 14:58 | 1553799468.769 | -170 | 0.20949 | 0.000438 | -1e-06 | 100.9109 | 101.0871 | 126.219 | 21.2612 | 0.001926 | 0.001754 | -151.6146 | 0.23062 | 8.4039 | 0.002627 | 0.003902 | 0.003689 | 5.4e-05 | 0.00128 | 0.32565 | 1.3117 | 2.5473 | 100.7 | 21.16 | 25.743 | 68.1668 | 85.0059 | 64.0412 | 0.006428 | 1.047 | 0.025268 | 24.1335 | -0.015608 | 0.1372 | 2.2e-05 | 0.001756 | 0.001756 | 3.2276
2019-03-28 14:58 | 1553799469.7691 | -169 | 0.20949 | 0.000438 | -1e-06 | 100.974 | 101.1364 | 126.3528 | 21.2612 | 0.001878 | 0.001772 | -151.6266 | 0.22929 | 8.4035 | 0.002548 | 0.003646 | 0.003538 | 5.6e-05 | 0.001283 | 0.35673 | 1.311 | 2.5474 | 100.7 | 21.17 | 25.752 | 67.0029 | 83.1201 | 64.4565 | 0.006434 | 1.0457 | 0.025266 | 24.1611 | -0.014371 | 0.1364 | 2e-06 | 0.000255 | 0.000255 | 22.0706
2019-03-28 14:58 | 1553799470.7682 | -168 | 0.2095 | 0.000436 | -1e-06 | 100.9171 | 101.0637 | 126.3812 | 21.2612 | 0.001841 | 0.001769 | -151.6292 | 0.22994 | 8.4045 | 0.002566 | 0.003383 | 0.003167 | 5.4e-05 | 0.001278 | 0.32748 | 1.3266 | 2.5472 | 100.7 | 21.17 | 25.756 | 66.1345 | 83.5553 | 64.3963 | 0.006435 | 1.0459 | 0.025271 | 24.1617 | -0.01298 | 0.13679 | -4.8e-05 | -0.003599 | -0.003599 | -1.5696
2019-03-28 14:58 | 1553799471.7686 | -167 | 0.20949 | 0.000439 | -1e-06 | 100.9624 | 101.0915 | 126.7047 | 21.2612 | 0.00186 | 0.001753 | -151.6169 | 0.23011 | 8.4042 | 0.002569 | 0.004255 | 0.004073 | 5.7e-05 | 0.001287 | 0.32244 | 1.3065 | 2.5474 | 100.7 | 21.17 | 25.756 | 66.5816 | 83.6224 | 64.006 | 0.006435 | 1.0471 | 0.025318 | 24.1786 | -0.011433 | 0.13689 | 1.3e-05 | 0.001064 | 0.001064 | 5.3112
2019-03-28 14:58 | 1553799472.7682 | -166 | 0.2095 | 0.00044 | -1e-06 | 100.9981 | 101.1255 | 126.4392 | 21.2612 | 0.001877 | 0.00178 | -151.6301 | 0.22988 | 8.4023 | 0.0026 | 0.003331 | 0.003176 | 5.3e-05 | 0.00129 | 0.31492 | 1.299 | 2.5474 | 100.7 | 21.17 | 25.773 | 66.9865 | 84.3613 | 64.6492 | 0.00644 | 1.0451 | 0.025267 | 24.1763 | -0.011272 | 0.13675 | -3.8e-05 | -0.002807 | -0.002807 | -2.0121
2019-03-28 14:58 | 1553799473.7682 | -165 | 0.20951 | 0.000439 | -1e-06 | 100.9931 | 101.1911 | 126.4157 | 21.2612 | 0.001854 | 0.001768 | -151.6444 | 0.22999 | 8.4022 | 0.002496 | 0.00396 | 0.003773 | 5.7e-05 | 0.001292 | 0.29894 | 1.3399 | 2.5476 | 100.7 | 21.17 | 25.773 | 66.4327 | 81.8611 | 64.3595 | 0.00644 | 1.046 | 0.025276 | 24.1637 | -0.017519 | 0.13682 | -0.000114 | -0.008583 | -0.008583 | -0.65836
2019-03-28 14:58 | 1553799474.7692 | -164 | 0.2095 | 0.000438 | -1e-06 | 101.0106 | 101.2111 | 126.2871 | 21.2612 | 0.001885 | 0.001748 | -151.627 | 0.22963 | 8.4033 | 0.002509 | 0.003608 | 0.003418 | 5.4e-05 | 0.001288 | 0.3126 | 1.3516 | 2.5473 | 100.7 | 21.17 | 25.79 | 67.1794 | 82.1755 | 63.884 | 0.006444 | 1.0475 | 0.025281 | 24.1344 | -0.017734 | 0.13661 | -3.2e-05 | -0.002383 | -0.002383 | -2.3677
2019-03-28 14:58 | 1553799475.7683 | -163 | 0.20949 | 0.000439 | -1e-06 | 101.0374 | 101.1201 | 126.5455 | 21.2612 | 0.001893 | 0.001794 | -151.6423 | 0.23023 | 8.4016 | 0.002551 | 0.003587 | 0.003399 | 5.2e-05 | 0.001289 | 0.33787 | 1.3494 | 2.5476 | 100.7 | 21.17 | 25.795 | 67.3665 | 83.1949 | 64.9967 | 0.006445 | 1.0441 | 0.025265 | 24.1989 | -0.00731 | 0.13696 | 2.9e-05 | 0.002267 | 0.002267 | 2.495
2019-03-28 14:58 | 1553799476.7688 | -162 | 0.20954 | 0.000444 | -1e-06 | 101.0544 | 101.1686 | 126.5833 | 21.2612 | 0.001871 | 0.00176 | -151.6353 | 0.22986 | 8.3995 | 0.002521 | 0.003627 | 0.003408 | 5.5e-05 | 0.001294 | 0.33301 | 1.3249 | 2.5475 | 100.7 | 21.17 | 25.806 | 66.8353 | 82.4563 | 64.1855 | 0.006448 | 1.0466 | 0.025299 | 24.1735 | -0.010104 | 0.13674 | -0.000279 | -0.021176 | -0.021176 | -0.26669
2019-03-28 14:58 | 1553799477.7684 | -161 | 0.20948 | 0.000438 | -1e-06 | 100.9377 | 101.1863 | 126.4885 | 21.2612 | 0.001841 | 0.001735 | -151.6577 | 0.23084 | 8.4032 | 0.002472 | 0.003967 | 0.003869 | 5.4e-05 | 0.001301 | 0.29806 | 1.3256 | 2.5474 | 100.7 | 21.17 | 25.806 | 66.122 | 81.2787 | 63.5757 | 0.006448 | 1.0485 | 0.025313 | 24.1425 | -0.022006 | 0.13733 | 8.5e-05 | 0.006562 | 0.006562 | 0.8643
2019-03-28 14:58 | 1553799478.7675 | -160 | 0.20964 | 0.000438 | -1e-06 | 101.0249 | 101.0246 | 126.5131 | 21.2612 | 0.001862 | 0.001766 | -151.6526 | 0.23006 | 8.4014 | 0.002542 | 0.00389 | 0.00366 | 5.3e-05 | 0.001308 | 0.31421 | 1.3154 | 2.5473 | 100.7 | 21.17 | 25.817 | 66.6246 | 82.9648 | 64.3167 | 0.006451 | 1.0462 | 0.025287 | 24.1715 | 3e-05 | 0.13687 | -0.000875 | -0.066561 | -0.066561 | -0.084923
2019-03-28 14:58 | 1553799479.7679 | -159 | 0.20932 | 0.00044 | -1e-06 | 101.0419 | 101.2512 | 126.4713 | 21.2612 | 0.001828 | 0.001795 | -151.6544 | 0.22966 | 8.4022 | 0.002564 | 0.003639 | 0.00349 | 5.2e-05 | 0.00132 | 0.31669 | 1.3042 | 2.5474 | 100.7 | 21.17 | 25.817 | 65.8052 | 83.4906 | 65.0154 | 0.006451 | 1.044 | 0.025257 | 24.1925 | -0.018507 | 0.13662 | 0.001047 | 0.079682 | 0.079682 | 0.070813
2019-03-28 14:58 | 1553799480.7677 | -158 | 0.20952 | 0.00044 | -1e-06 | 100.9594 | 101.0739 | 126.376 | 21.2612 | 0.001946 | 0.00174 | -151.671 | 0.23048 | 8.4031 | 0.00253 | 0.003765 | 0.00364 | 5.3e-05 | 0.001331 | 0.29446 | 1.3145 | 2.5474 | 100.7 | 21.17 | 25.832 | 68.6443 | 82.6722 | 63.6867 | 0.006454 | 1.0481 | 0.025297 | 24.1358 | -0.010134 | 0.13711 | -0.000174 | -0.013193 | -0.013193 | -0.42923
2019-03-28 14:58 | 1553799481.7689 | -157 | 0.2095 | 0.00044 | -1e-06 | 101.0732 | 101.0993 | 126.1951 | 21.2612 | 0.001893 | 0.001757 | -151.69 | 0.22956 | 8.4025 | 0.002579 | 0.003274 | 0.003103 | 5.4e-05 | 0.001339 | 0.3361 | 1.31 | 2.5474 | 100.7 | 21.17 | 25.832 | 67.3806 | 83.8657 | 64.1011 | 0.006454 | 1.0468 | 0.025264 | 24.1333 | -0.002314 | 0.13657 | -7.2e-05 | -0.005459 | -0.005459 | -1.0332
2019-03-28 14:58 | 1553799482.7675 | -156 | 0.20938 | 0.00044 | -1e-06 | 101.0089 | 101.2047 | 126.4969 | 21.2612 | 0.001891 | 0.001782 | -151.682 | 0.23039 | 8.4038 | 0.002573 | 0.004105 | 0.003879 | 5.5e-05 | 0.001334 | 0.30773 | 1.3303 | 2.5473 | 100.7 | 21.17 | 25.838 | 67.3188 | 83.7169 | 64.7108 | 0.006456 | 1.0449 | 0.025271 | 24.184 | -0.017321 | 0.13706 | 0.000693 | 0.052829 | 0.052829 | 0.10715
2019-03-28 14:58 | 1553799483.7686 | -155 | 0.2095 | 0.000443 | -1e-06 | 101.0196 | 101.2064 | 126.623 | 21.2612 | 0.00191 | 0.001749 | -151.6905 | 0.23079 | 8.4012 | 0.002439 | 0.003716 | 0.003547 | 5.7e-05 | 0.001338 | 0.33428 | 1.3724 | 2.5475 | 100.7 | 21.17 | 25.838 | 67.7877 | 80.5012 | 63.9129 | 0.006456 | 1.0474 | 0.025313 | 24.1675 | -0.016522 | 0.1373 | -8.5e-05 | -0.006461 | -0.006461 | -0.87762
2019-03-28 14:58 | 1553799484.768 | -154 | 0.20949 | 0.000439 | -1e-06 | 100.9615 | 101.1241 | 126.9492 | 21.2612 | 0.001888 | 0.001762 | -151.7018 | 0.23021 | 8.4011 | 0.002478 | 0.00374 | 0.003567 | 5.4e-05 | 0.001332 | 0.31061 | 1.3543 | 2.5475 | 100.7 | 21.17 | 25.843 | 67.2465 | 81.4253 | 64.2221 | 0.006457 | 1.0465 | 0.025334 | 24.2097 | -0.014397 | 0.13695 | -2.3e-05 | -0.001695 | -0.001695 | -3.3377
2019-03-28 14:58 | 1553799485.7687 | -153 | 0.20947 | 0.000442 | -1e-06 | 101.0669 | 101.2247 | 127.0112 | 21.2612 | 0.001918 | 0.001768 | -151.6975 | 0.23033 | 8.4037 | 0.002448 | 0.004674 | 0.004467 | 5.3e-05 | 0.001331 | 0.34003 | 1.3715 | 2.5474 | 100.7 | 21.17 | 25.843 | 67.968 | 80.7079 | 64.3689 | 0.006457 | 1.046 | 0.025335 | 24.2209 | -0.013951 | 0.13703 | 0.000129 | 0.009933 | 0.009933 | 0.56973
2019-03-28 14:58 | 1553799486.769 | -152 | 0.20948 | 0.000439 | -1e-06 | 101.0544 | 101.1965 | 127.1941 | 21.2612 | 0.001847 | 0.001722 | -151.6995 | 0.23014 | 8.4006 | 0.002565 | 0.003716 | 0.003573 | 5.3e-05 | 0.00133 | 0.32908 | 1.334 | 2.5473 | 100.7 | 21.17 | 25.846 | 66.2788 | 83.5163 | 63.257 | 0.006458 | 1.0495 | 0.025395 | 24.1983 | -0.012566 | 0.13691 | 4.5e-05 | 0.003506 | 0.003506 | 1.6127
2019-03-28 14:58 | 1553799487.7677 | -151 | 0.20949 | 0.000438 | -1e-06 | 101.0092 | 101.084 | 127.464 | 21.2612 | 0.0019 | 0.001802 | -151.7022 | 0.23058 | 8.4028 | 0.002576 | 0.003381 | 0.0032 | 5.5e-05 | 0.001331 | 0.34363 | 1.368 | 2.5475 | 100.7 | 21.17 | 25.846 | 67.5405 | 83.7955 | 65.1956 | 0.006458 | 1.0434 | 0.025349 | 24.2937 | -0.00663 | 0.13717 | -1.9e-05 | -0.001405 | -0.001405 | -4.0332
2019-03-28 14:58 | 1553799488.7691 | -150 | 0.20948 | 0.000442 | 0 | 100.9374 | 101.0953 | 127.4267 | 21.2612 | 0.001836 | 0.001776 | -151.7126 | 0.2304 | 8.4021 | 0.002583 | 0.004012 | 0.003857 | 5.4e-05 | 0.001317 | 0.3366 | 1.2955 | 2.5474 | 100.7 | 21.17 | 25.854 | 66.012 | 83.9515 | 64.5595 | 0.00646 | 1.0454 | 0.025369 | 24.2673 | -0.013981 | 0.13707 | 6.6e-05 | 0.00508 | 0.00508 | 1.1143
2019-03-28 14:58 | 1553799489.7688 | -149 | 0.20945 | 0.000441 | 0 | 100.9701 | 101.1928 | 127.1924 | 21.2612 | 0.001887 | 0.001789 | -151.7048 | 0.23072 | 8.4025 | 0.002505 | 0.004301 | 0.004145 | 5.5e-05 | 0.001311 | 0.31598 | 1.2655 | 2.5474 | 100.7 | 21.17 | 25.859 | 67.2323 | 82.0921 | 64.8718 | 0.006461 | 1.0444 | 0.025334 | 24.2562 | -0.0197 | 0.13725 | 0.000217 | 0.01658 | 0.01658 | 0.34188
2019-03-28 14:58 | 1553799490.7678 | -148 | 0.20949 | 0.000443 | -1e-06 | 101.0059 | 100.9611 | 127.2886 | 21.2612 | 0.00189 | 0.001743 | -151.7059 | 0.22978 | 8.4039 | 0.002545 | 0.004066 | 0.003898 | 5.1e-05 | 0.001306 | 0.33107 | 1.316 | 2.5473 | 100.7 | 21.17 | 25.861 | 67.2955 | 83.0513 | 63.7591 | 0.006462 | 1.0479 | 0.025386 | 24.2254 | 0.003969 | 0.1367 | -1.8e-05 | -0.00129 | -0.00129 | -4.3766
2019-03-28 14:58 | 1553799491.7678 | -147 | 0.20943 | 0.000441 | 0 | 100.9383 | 100.8367 | 127.3689 | 21.2612 | 0.001852 | 0.001779 | -151.7134 | 0.23055 | 8.4035 | 0.002529 | 0.003543 | 0.003321 | 5.8e-05 | 0.001299 | 0.3361 | 1.3181 | 2.5475 | 100.7 | 21.17 | 25.863 | 66.3913 | 82.6496 | 64.6454 | 0.006462 | 1.0451 | 0.02536 | 24.2649 | 0.009002 | 0.13715 | 0.000381 | 0.029108 | 0.029108 | 0.1946
2019-03-28 14:58 | 1553799492.7679 | -146 | 0.20947 | 0.00044 | -1e-06 | 100.9969 | 101.2322 | 127.3546 | 21.2612 | 0.001862 | 0.001826 | -151.7196 | 0.22982 | 8.402 | 0.002529 | 0.003987 | 0.003854 | 5.7e-05 | 0.001296 | 0.30807 | 1.3462 | 2.5474 | 100.7 | 21.17 | 25.863 | 66.6191 | 82.6582 | 65.7568 | 0.006462 | 1.0417 | 0.025317 | 24.3034 | -0.020808 | 0.13672 | 9.6e-05 | 0.007378 | 0.007378 | 0.76531
2019-03-28 14:58 | 1553799493.7669 | -145 | 0.20947 | 0.000444 | -1e-06 | 100.9263 | 101.189 | 127.272 | 21.2612 | 0.001872 | 0.001792 | -151.7028 | 0.23074 | 8.3998 | 0.002548 | 0.004331 | 0.004166 | 5.6e-05 | 0.001293 | 0.35153 | 1.3333 | 2.5473 | 100.7 | 21.17 | 25.869 | 66.8603 | 83.1271 | 64.9497 | 0.006464 | 1.0442 | 0.025339 | 24.2666 | -0.023249 | 0.13727 | 0.000142 | 0.010889 | 0.010889 | 0.52063
2019-03-28 14:58 | 1553799494.7669 | -144 | 0.20951 | 0.00044 | 0 | 100.9773 | 101.1557 | 127.2884 | 21.2612 | 0.001894 | 0.001767 | -151.6992 | 0.23009 | 8.4041 | 0.002618 | 0.004485 | 0.004265 | 5.9e-05 | 0.001293 | 0.30165 | 1.3082 | 2.5474 | 100.7 | 21.17 | 25.869 | 67.3994 | 84.799 | 64.3331 | 0.006464 | 1.0461 | 0.025364 | 24.246 | -0.015793 | 0.13688 | -9.5e-05 | -0.007231 | -0.007231 | -0.78182
2019-03-28 14:58 | 1553799495.7669 | -143 | 0.20949 | 0.000438 | 0 | 101.0035 | 101.1649 | 127.2072 | 21.2612 | 0.001874 | 0.001765 | -151.6945 | 0.23061 | 8.4027 | 0.00264 | 0.003796 | 0.003698 | 5.6e-05 | 0.001287 | 0.31045 | 1.3167 | 2.5474 | 100.7 | 21.17 | 25.868 | 66.9187 | 85.3265 | 64.2902 | 0.006463 | 1.0462 | 0.025358 | 24.2367 | -0.014283 | 0.13719 | -1.5e-05 | -0.001122 | -0.001122 | -5.0488
2019-03-28 14:58 | 1553799496.7684 | -142 | 0.2095 | 0.000439 | 0 | 101.0011 | 101.1819 | 127.202 | 21.2612 | 0.001883 | 0.001787 | -151.6912 | 0.23072 | 8.404 | 0.002612 | 0.004035 | 0.003821 | 5.7e-05 | 0.001278 | 0.29645 | 1.3663 | 2.5473 | 100.7 | 21.17 | 25.868 | 67.1332 | 84.6562 | 64.835 | 0.006463 | 1.0446 | 0.025337 | 24.2558 | -0.015995 | 0.13726 | -5.8e-05 | -0.004428 | -0.004428 | -1.2803
2019-03-28 14:58 | 1553799497.7671 | -141 | 0.20949 | 0.000441 | 0 | 100.991 | 101.0361 | 127.0084 | 21.2612 | 0.00191 | 0.001786 | -151.7012 | 0.22985 | 8.4036 | 0.002585 | 0.004103 | 0.003894 | 5.4e-05 | 0.001276 | 0.3449 | 1.3848 | 2.5475 | 100.7 | 21.17 | 25.859 | 67.7894 | 84.0155 | 64.8061 | 0.006461 | 1.0446 | 0.025318 | 24.2363 | -0.003999 | 0.13674 | -6e-06 | -0.000411 | -0.000411 | -13.742
2019-03-28 14:58 | 1553799498.767 | -140 | 0.20949 | 0.000443 | 0 | 100.9645 | 101.1483 | 126.8075 | 21.2612 | 0.001891 | 0.001826 | -151.671 | 0.22914 | 8.4031 | 0.002612 | 0.004172 | 0.00396 | 5.8e-05 | 0.001276 | 0.34954 | 1.3493 | 2.5474 | 100.7 | 21.17 | 25.859 | 67.3323 | 84.653 | 65.7537 | 0.006461 | 1.0417 | 0.025263 | 24.251 | -0.016267 | 0.13632 | 2.1e-05 | 0.001572 | 0.001572 | 3.5812
2019-03-28 14:58 | 1553799499.7685 | -139 | 0.20949 | 0.000444 | 0 | 100.9734 | 101.1377 | 126.654 | 21.2612 | 0.001872 | 0.001771 | -151.6869 | 0.22964 | 8.4012 | 0.00255 | 0.004223 | 0.004026 | 5.8e-05 | 0.001272 | 0.29142 | 1.3272 | 2.5474 | 100.7 | 21.17 | 25.85 | 66.8757 | 83.1605 | 64.4517 | 0.006459 | 1.0457 | 0.025296 | 24.1897 | -0.014544 | 0.13661 | -7e-06 | -0.000477 | -0.000477 | -11.8247
2019-03-28 14:58 | 1553799500.7681 | -138 | 0.20949 | 0.00044 | 0 | 100.9383 | 101.116 | 126.5523 | 21.2612 | 0.001888 | 0.001773 | -151.6971 | 0.22998 | 8.4044 | 0.002607 | 0.003607 | 0.003453 | 5.4e-05 | 0.00127 | 0.31188 | 1.3114 | 2.5474 | 100.7 | 21.17 | 25.843 | 67.2643 | 84.5431 | 64.4938 | 0.006457 | 1.0456 | 0.025285 | 24.1815 | -0.015735 | 0.13682 | 2.8e-05 | 0.002138 | 0.002138 | 2.6425
2019-03-28 14:58 | 1553799501.7676 | -137 | 0.20943 | 0.000439 | -1e-06 | 100.9329 | 101.1245 | 126.284 | 21.2612 | 0.001905 | 0.001751 | -151.6932 | 0.22981 | 8.4016 | 0.002593 | 0.003624 | 0.003477 | 5.5e-05 | 0.001266 | 0.30702 | 1.349 | 2.5473 | 100.7 | 21.17 | 25.84 | 67.6723 | 84.2028 | 63.9615 | 0.006456 | 1.0473 | 0.025278 | 24.1368 | -0.016961 | 0.13672 | 0.000383 | 0.02916 | 0.02916 | 0.19363
2019-03-28 14:58 | 1553799502.7675 | -136 | 0.20945 | 0.000439 | 0 | 100.9683 | 101.1204 | 126.1218 | 21.2612 | 0.001873 | 0.001725 | -151.6944 | 0.22938 | 8.4044 | 0.002522 | 0.00399 | 0.003821 | 5.6e-05 | 0.001268 | 0.32947 | 1.3399 | 2.5472 | 100.7 | 21.17 | 25.837 | 66.8923 | 82.4977 | 63.3446 | 0.006456 | 1.0492 | 0.025285 | 24.0993 | -0.01346 | 0.13646 | 0.000261 | 0.019919 | 0.019919 | 0.28294
2019-03-28 14:58 | 1553799503.7674 | -135 | 0.20945 | 0.000441 | 0 | 100.9088 | 101.0446 | 126.0143 | 21.2612 | 0.001935 | 0.001754 | -151.7125 | 0.22971 | 8.4032 | 0.002584 | 0.002831 | 0.002685 | 5.3e-05 | 0.001273 | 0.31188 | 1.3449 | 2.5475 | 100.7 | 21.17 | 25.835 | 68.3767 | 83.9814 | 64.0214 | 0.006455 | 1.0471 | 0.025248 | 24.1132 | -0.012035 | 0.13665 | 0.000267 | 0.020271 | 0.020271 | 0.27842
2019-03-28 14:58 | 1553799504.7685 | -134 | 0.20946 | 0.00044 | -1e-06 | 100.9022 | 100.8707 | 125.797 | 21.2612 | 0.001871 | 0.001746 | -151.7226 | 0.22994 | 8.4027 | 0.002556 | 0.004089 | 0.003916 | 5.4e-05 | 0.001274 | 0.34219 | 1.2893 | 2.5474 | 100.7 | 21.17 | 25.835 | 66.8485 | 83.3189 | 63.8287 | 0.006455 | 1.0477 | 0.025234 | 24.0855 | 0.002795 | 0.13679 | 0.000169 | 0.012855 | 0.012855 | 0.43947
2019-03-28 14:58 | 1553799505.8058 | -133 | 0.20949 | 0.000442 | -1e-06 | 100.8844 | 100.8296 | 125.4381 | 21.2612 | 0.001872 | 0.001797 | -151.7172 | 0.22987 | 8.4036 | 0.002574 | 0.00349 | 0.003315 | 5.4e-05 | 0.001277 | 0.29535 | 1.3058 | 2.5473 | 100.7 | 21.17 | 25.835 | 66.8617 | 83.7419 | 65.0638 | 0.006455 | 1.0439 | 0.025152 | 24.0952 | 0.004856 | 0.13675 | 2.3e-05 | 0.001831 | 0.001831 | 3.0846
2019-03-28 14:58 | 1553799506.7536 | -132 | 0.20948 | 0.000438 | -1e-06 | 100.9144 | 101.0874 | 125.097 | 21.2612 | 0.00192 | 0.001744 | -151.7368 | 0.22928 | 8.3999 | 0.002585 | 0.004152 | 0.003972 | 5.2e-05 | 0.001269 | 0.30309 | 1.3007 | 2.5473 | 100.7 | 21.17 | 25.838 | 68.024 | 84.0001 | 63.7803 | 0.006456 | 1.0478 | 0.025165 | 24.0167 | -0.015322 | 0.1364 | 6.2e-05 | 0.004742 | 0.004742 | 1.1879
2019-03-28 14:58 | 1553799507.7525 | -131 | 0.20949 | 0.000443 | 0 | 100.8704 | 101.0773 | 125.1095 | 21.2612 | 0.001919 | 0.001776 | -151.7163 | 0.23003 | 8.4024 | 0.002505 | 0.00404 | 0.003898 | 5.3e-05 | 0.001269 | 0.32997 | 1.3267 | 2.5473 | 100.7 | 21.17 | 25.842 | 68.0101 | 82.0724 | 64.565 | 0.006457 | 1.0454 | 0.025137 | 24.0458 | -0.018328 | 0.13685 | -1.1e-05 | -0.000811 | -0.000811 | -6.9717
2019-03-28 14:58 | 1553799508.7535 | -130 | 0.20949 | 0.000439 | -1e-06 | 100.8709 | 101.1021 | 125.2975 | 21.2612 | 0.001905 | 0.001774 | -151.7413 | 0.23007 | 8.4036 | 0.002537 | 0.003962 | 0.003675 | 5.1e-05 | 0.001266 | 0.31205 | 1.3315 | 2.5474 | 100.7 | 21.17 | 25.845 | 67.6605 | 82.8417 | 64.5087 | 0.006458 | 1.0456 | 0.025158 | 24.0619 | -0.020469 | 0.13687 | -1e-05 | -0.000749 | -0.000749 | -7.5517
2019-03-28 14:58 | 1553799509.7536 | -129 | 0.20951 | 0.000439 | 0 | 100.8391 | 101.0949 | 125.5735 | 21.2612 | 0.001867 | 0.001751 | -151.7271 | 0.23031 | 8.402 | 0.002473 | 0.004187 | 0.004088 | 5.5e-05 | 0.001267 | 0.33539 | 1.3061 | 2.5473 | 100.7 | 21.17 | 25.845 | 66.7446 | 81.309 | 63.949 | 0.006458 | 1.0473 | 0.025207 | 24.0684 | -0.022665 | 0.13701 | -0.000116 | -0.008762 | -0.008762 | -0.64582
2019-03-28 14:58 | 1553799510.7529 | -128 | 0.20951 | 0.00044 | 0 | 100.9007 | 100.9556 | 125.7165 | 21.2612 | 0.001891 | 0.001779 | -151.7048 | 0.22903 | 8.4015 | 0.002553 | 0.004125 | 0.003909 | 5.2e-05 | 0.001263 | 0.32892 | 1.3195 | 2.5474 | 100.7 | 21.16 | 25.857 | 67.3174 | 83.2324 | 64.6235 | 0.006457 | 1.0452 | 0.025196 | 24.1062 | -0.004865 | 0.13625 | -0.000116 | -0.008814 | -0.008814 | -0.63845
2019-03-28 14:58 | 1553799511.7526 | -127 | 0.2095 | 0.000445 | -1e-06 | 100.8701 | 100.8568 | 125.9099 | 21.2612 | 0.001901 | 0.001765 | -151.6907 | 0.22948 | 8.4001 | 0.002541 | 0.003594 | 0.003391 | 5.2e-05 | 0.001268 | 0.32947 | 1.3524 | 2.5475 | 100.7 | 21.16 | 25.859 | 67.56 | 82.9398 | 64.3049 | 0.006457 | 1.0462 | 0.025227 | 24.1133 | 0.001178 | 0.13652 | -6.7e-05 | -0.005087 | -0.005087 | -1.1083
2019-03-28 14:59 | 1553799512.7527 | -126 | 0.2095 | 0.000438 | -1e-06 | 100.8665 | 101.0589 | 126.2343 | 21.2612 | 0.001872 | 0.001796 | -151.6919 | 0.22942 | 8.4007 | 0.002503 | 0.003932 | 0.003757 | 5.1e-05 | 0.001281 | 0.3418 | 1.3147 | 2.5474 | 100.7 | 21.16 | 25.863 | 66.8624 | 82.0436 | 65.038 | 0.006458 | 1.0439 | 0.025232 | 24.1706 | -0.017048 | 0.13648 | -7.1e-05 | -0.005326 | -0.005326 | -1.0584
2019-03-28 14:59 | 1553799513.753 | -125 | 0.2095 | 0.000436 | -1e-06 | 100.8879 | 100.9696 | 126.5756 | 21.2612 | 0.001853 | 0.001751 | -151.6901 | 0.22981 | 8.4018 | 0.002506 | 0.004077 | 0.00392 | 5.3e-05 | 0.00128 | 0.3621 | 1.3264 | 2.5472 | 100.7 | 21.16 | 25.863 | 66.4154 | 82.096 | 63.9661 | 0.006458 | 1.0473 | 0.025307 | 24.1649 | -0.007235 | 0.13672 | -4e-05 | -0.003018 | -0.003018 | -1.871
2019-03-28 14:59 | 1553799514.7531 | -124 | 0.2095 | 0.000441 | 0 | 100.862 | 100.9407 | 126.6152 | 21.2612 | 0.001907 | 0.00177 | -151.6797 | 0.22995 | 8.4044 | 0.002542 | 0.003516 | 0.003359 | 5.3e-05 | 0.001282 | 0.32659 | 1.3475 | 2.5473 | 100.7 | 21.16 | 25.865 | 67.7081 | 82.9672 | 64.412 | 0.006459 | 1.0459 | 0.025294 | 24.1846 | -0.006974 | 0.1368 | -3.3e-05 | -0.002466 | -0.002466 | -2.291
2019-03-28 14:59 | 1553799515.7527 | -123 | 0.20951 | 0.000441 | 0 | 100.8418 | 101.0103 | 127.1335 | 21.2612 | 0.001929 | 0.001719 | -151.6745 | 0.23027 | 8.403 | 0.002553 | 0.004262 | 0.004111 | 5.5e-05 | 0.001283 | 0.3115 | 1.3313 | 2.5472 | 100.7 | 21.16 | 25.865 | 68.2308 | 83.2355 | 63.1954 | 0.006459 | 1.0497 | 0.025391 | 24.1904 | -0.014939 | 0.13699 | -0.000113 | -0.008644 | -0.008644 | -0.65448
2019-03-28 14:59 | 1553799516.7528 | -122 | 0.20951 | 0.000443 | -1e-06 | 100.8432 | 100.9638 | 127.3565 | 21.2612 | 0.001916 | 0.001763 | -151.6656 | 0.22969 | 8.3998 | 0.002524 | 0.003794 | 0.003668 | 5.5e-05 | 0.001284 | 0.31166 | 1.3491 | 2.5474 | 100.7 | 21.16 | 25.866 | 67.9322 | 82.5443 | 64.2582 | 0.006459 | 1.0463 | 0.025374 | 24.2498 | -0.010685 | 0.13664 | -0.000117 | -0.008862 | -0.008862 | -0.63679
2019-03-28 14:59 | 1553799517.7541 | -121 | 0.20951 | 0.000444 | 0 | 100.8581 | 100.7385 | 127.3907 | 21.2612 | 0.001877 | 0.001727 | -151.6807 | 0.22936 | 8.4041 | 0.002568 | 0.00342 | 0.003263 | 5.3e-05 | 0.001281 | 0.31365 | 1.3464 | 2.5473 | 100.7 | 21.16 | 25.868 | 66.9858 | 83.6005 | 63.3788 | 0.006459 | 1.0491 | 0.02541 | 24.2214 | 0.010612 | 0.13645 | -0.000141 | -0.010797 | -0.010797 | -0.52193
2019-03-28 14:59 | 1553799518.7528 | -120 | 0.20951 | 0.000438 | 0 | 100.8793 | 100.9747 | 127.2006 | 21.2612 | 0.001902 | 0.001814 | -151.6829 | 0.22998 | 8.4017 | 0.002559 | 0.00383 | 0.003659 | 5.4e-05 | 0.001279 | 0.31542 | 1.3386 | 2.5473 | 100.7 | 21.16 | 25.871 | 67.5944 | 83.3877 | 65.4867 | 0.00646 | 1.0425 | 0.025312 | 24.279 | -0.008452 | 0.13682 | -0.0001 | -0.007612 | -0.007612 | -0.74234
2019-03-28 14:59 | 1553799519.7542 | -119 | 0.20951 | 0.000441 | 0 | 100.7631 | 100.9509 | 127.1384 | 21.2612 | 0.001869 | 0.001772 | -151.7008 | 0.2301 | 8.4043 | 0.002553 | 0.004267 | 0.004037 | 5.4e-05 | 0.001274 | 0.32731 | 1.2845 | 2.5473 | 100.7 | 21.16 | 25.874 | 66.7939 | 83.2285 | 64.4712 | 0.006461 | 1.0457 | 0.025344 | 24.2367 | -0.016651 | 0.13689 | -0.000114 | -0.008665 | -0.008665 | -0.65242
2019-03-28 14:59 | 1553799520.7539 | -118 | 0.20949 | 0.000443 | 0 | 100.8108 | 100.9186 | 127.0654 | 21.2612 | 0.001883 | 0.001815 | -151.6923 | 0.22974 | 8.4022 | 0.002518 | 0.00376 | 0.003565 | 5.2e-05 | 0.001273 | 0.3063 | 1.2959 | 2.5474 | 100.7 | 21.16 | 25.88 | 67.127 | 82.4056 | 65.4992 | 0.006462 | 1.0425 | 0.025298 | 24.2666 | -0.009561 | 0.13667 | -2e-06 | -0.000167 | -0.000167 | -33.7506
2019-03-28 14:59 | 1553799521.7529 | -117 | 0.20951 | 0.000442 | 0 | 100.8376 | 100.9747 | 126.9626 | 21.2612 | 0.001884 | 0.001749 | -151.7025 | 0.23023 | 8.4031 | 0.002591 | 0.003637 | 0.003449 | 5.2e-05 | 0.001273 | 0.32604 | 1.292 | 2.5472 | 100.7 | 21.16 | 25.884 | 67.1551 | 84.1466 | 63.9021 | 0.006463 | 1.0475 | 0.025348 | 24.1995 | -0.01215 | 0.13696 | -0.000126 | -0.009579 | -0.009579 | -0.59051
2019-03-28 14:59 | 1553799522.7521 | -116 | 0.20952 | 0.000445 | 0 | 100.7875 | 100.9702 | 126.6985 | 21.2612 | 0.001907 | 0.001813 | -151.7158 | 0.22989 | 8.4019 | 0.002566 | 0.004341 | 0.004105 | 5.2e-05 | 0.001271 | 0.31509 | 1.3297 | 2.5474 | 100.7 | 21.16 | 25.886 | 67.7091 | 83.5497 | 65.4446 | 0.006464 | 1.0427 | 0.025264 | 24.2295 | -0.016199 | 0.13676 | -0.000162 | -0.012284 | -0.012284 | -0.45981
2019-03-28 14:59 | 1553799523.7521 | -115 | 0.20952 | 0.000444 | -1e-06 | 100.7762 | 101.0385 | 126.4803 | 21.2612 | 0.001891 | 0.001801 | -151.7106 | 0.22994 | 8.402 | 0.002614 | 0.003516 | 0.00346 | 5.2e-05 | 0.001271 | 0.32532 | 1.2929 | 2.5473 | 100.7 | 21.16 | 25.891 | 67.3167 | 84.6999 | 65.1542 | 0.006465 | 1.0436 | 0.025253 | 24.1983 | -0.02325 | 0.13679 | -0.000178 | -0.013527 | -0.013527 | -0.41763
2019-03-28 14:59 | 1553799524.7524 | -114 | 0.20946 | 0.000438 | 0 | 100.8376 | 100.8598 | 126.4631 | 21.2612 | 0.00185 | 0.001789 | -151.7092 | 0.23021 | 8.4034 | 0.002696 | 0.003688 | 0.003537 | 4.9e-05 | 0.001275 | 0.33356 | 1.3199 | 2.5474 | 100.7 | 21.16 | 25.892 | 66.3367 | 86.6699 | 64.8732 | 0.006465 | 1.0444 | 0.025261 | 24.1866 | -0.001973 | 0.13695 | 0.000171 | 0.012995 | 0.012995 | 0.43525
2019-03-28 14:59 | 1553799525.7524 | -113 | 0.20951 | 0.000441 | -1e-06 | 100.8188 | 100.9865 | 126.3063 | 21.2612 | 0.001924 | 0.001716 | -151.7216 | 0.22942 | 8.4004 | 0.002612 | 0.004283 | 0.00404 | 5.1e-05 | 0.001276 | 0.29518 | 1.2985 | 2.5473 | 100.7 | 21.16 | 25.893 | 68.1137 | 84.6499 | 63.1102 | 0.006466 | 1.0499 | 0.025312 | 24.1085 | -0.014868 | 0.13648 | -0.000113 | -0.008549 | -0.008549 | -0.65935
2019-03-28 14:59 | 1553799526.7531 | -112 | 0.2095 | 0.000442 | 0 | 100.8662 | 100.9862 | 126.311 | 21.2612 | 0.001912 | 0.001781 | -151.6961 | 0.23034 | 8.4042 | 0.002662 | 0.004178 | 0.004007 | 5.3e-05 | 0.001277 | 0.3314 | 1.3496 | 2.5474 | 100.7 | 21.16 | 25.893 | 67.8206 | 85.8502 | 64.686 | 0.006466 | 1.045 | 0.025253 | 24.1654 | -0.010637 | 0.13703 | -5.8e-05 | -0.00439 | -0.00439 | -1.2892
2019-03-28 14:59 | 1553799527.7531 | -111 | 0.20949 | 0.000441 | 0 | 100.8852 | 101.0728 | 126.0343 | 21.2612 | 0.001823 | 0.001791 | -151.689 | 0.23073 | 8.4031 | 0.002638 | 0.003293 | 0.00314 | 5.4e-05 | 0.001278 | 0.32693 | 1.3015 | 2.5473 | 100.7 | 21.16 | 25.893 | 65.7037 | 85.2899 | 64.9185 | 0.006466 | 1.0443 | 0.025217 | 24.1472 | -0.016617 | 0.13726 | -1.8e-05 | -0.00139 | -0.00139 | -4.0775
2019-03-28 14:59 | 1553799528.7531 | -110 | 0.2095 | 0.000445 | -1e-06 | 100.868 | 100.8693 | 126.3774 | 21.2612 | 0.001849 | 0.001789 | -151.6695 | 0.23085 | 8.4036 | 0.002594 | 0.003102 | 0.002902 | 4.9e-05 | 0.001281 | 0.32278 | 1.3388 | 2.5473 | 100.7 | 21.16 | 25.891 | 66.3055 | 84.2192 | 64.8631 | 0.006465 | 1.0445 | 0.025253 | 24.178 | -0.000122 | 0.13733 | -7.2e-05 | -0.005469 | -0.005469 | -1.0371
2019-03-28 14:59 | 1553799529.7522 | -109 | 0.20952 | 0.000442 | 0 | 100.8614 | 100.9747 | 126.2706 | 21.2612 | 0.001889 | 0.001784 | -151.6682 | 0.23033 | 8.4033 | 0.0026 | 0.003563 | 0.003404 | 5.1e-05 | 0.001288 | 0.32947 | 1.3197 | 2.5475 | 100.7 | 21.16 | 25.89 | 67.276 | 84.3721 | 64.7609 | 0.006465 | 1.0448 | 0.025246 | 24.1642 | -0.010037 | 0.13702 | -0.000145 | -0.011044 | -0.011044 | -0.51239
2019-03-28 14:59 | 1553799530.7536 | -108 | 0.20948 | 0.000441 | 0 | 100.8632 | 100.9835 | 126.1455 | 21.2612 | 0.001909 | 0.001769 | -151.676 | 0.23049 | 8.4035 | 0.002654 | 0.003513 | 0.003355 | 5e-05 | 0.001293 | 0.31133 | 1.3338 | 2.5473 | 100.7 | 21.16 | 25.89 | 67.7588 | 85.6652 | 64.3894 | 0.006465 | 1.0459 | 0.025248 | 24.1389 | -0.010661 | 0.13712 | 8.6e-05 | 0.006543 | 0.006543 | 0.86554
2019-03-28 14:59 | 1553799531.7532 | -107 | 0.20951 | 0.000442 | -1e-06 | 100.8969 | 101.0239 | 126.2556 | 21.2612 | 0.001909 | 0.001792 | -151.6643 | 0.23023 | 8.4031 | 0.002551 | 0.003713 | 0.003507 | 5.2e-05 | 0.001296 | 0.31598 | 1.3379 | 2.5474 | 100.7 | 21.16 | 25.887 | 67.7526 | 83.191 | 64.9396 | 0.006464 | 1.0442 | 0.025238 | 24.1691 | -0.011256 | 0.13697 | -0.000129 | -0.009781 | -0.009781 | -0.57832
2019-03-28 14:59 | 1553799532.7535 | -106 | 0.20943 | 0.000442 | 0 | 100.868 | 101.0311 | 126.1418 | 21.2612 | 0.001882 | 0.001782 | -151.6383 | 0.23027 | 8.4026 | 0.002582 | 0.003691 | 0.003559 | 5.3e-05 | 0.001297 | 0.30702 | 1.3379 | 2.5473 | 100.7 | 21.16 | 25.885 | 67.1005 | 83.9227 | 64.7055 | 0.006464 | 1.045 | 0.025236 | 24.1499 | -0.01445 | 0.13699 | 0.0004 | 0.030409 | 0.030409 | 0.18606
2019-03-28 14:59 | 1553799533.7532 | -105 | 0.20951 | 0.000441 | 0 | 100.8838 | 100.9845 | 125.8618 | 21.2612 | 0.001853 | 0.001803 | -151.6525 | 0.23075 | 8.402 | 0.002548 | 0.004502 | 0.004311 | 5.1e-05 | 0.001299 | 0.34562 | 1.3131 | 2.5473 | 100.7 | 21.16 | 25.882 | 66.4233 | 83.1045 | 65.2081 | 0.006463 | 1.0434 | 0.025189 | 24.141 | -0.008929 | 0.13727 | -8.6e-05 | -0.006499 | -0.006499 | -0.87238
2019-03-28 14:59 | 1553799534.7533 | -104 | 0.20954 | 0.000443 | 0 | 100.8844 | 100.8075 | 125.8303 | 21.2612 | 0.001913 | 0.001772 | -151.6364 | 0.2303 | 8.4043 | 0.0026 | 0.004103 | 0.003899 | 4.9e-05 | 0.001301 | 0.31006 | 1.3028 | 2.5474 | 100.7 | 21.16 | 25.882 | 67.8658 | 84.3613 | 64.4579 | 0.006463 | 1.0457 | 0.025214 | 24.1112 | 0.006815 | 0.137 | -0.000298 | -0.022656 | -0.022656 | -0.24975
2019-03-28 14:59 | 1553799535.7523 | -103 | 0.20936 | 0.000437 | 0 | 100.9326 | 101.115 | 125.6993 | 21.2612 | 0.001857 | 0.001784 | -151.6457 | 0.22998 | 8.4014 | 0.002551 | 0.004521 | 0.004369 | 5.2e-05 | 0.001306 | 0.33892 | 1.3065 | 2.5474 | 100.7 | 21.15 | 25.882 | 66.5068 | 83.1896 | 64.7515 | 0.006459 | 1.0448 | 0.02519 | 24.1091 | -0.016146 | 0.13682 | 0.000822 | 0.062327 | 0.062327 | 0.090659
2019-03-28 14:59 | 1553799536.7534 | -102 | 0.20954 | 0.000444 | -1e-06 | 100.9445 | 101.0593 | 125.5996 | 21.2612 | 0.001927 | 0.001734 | -151.6585 | 0.22904 | 8.4017 | 0.002617 | 0.003243 | 0.003079 | 5e-05 | 0.001305 | 0.33716 | 1.3067 | 2.5474 | 100.7 | 21.15 | 25.882 | 68.1942 | 84.7788 | 63.5601 | 0.006459 | 1.0485 | 0.025224 | 24.057 | -0.010161 | 0.13625 | -0.00028 | -0.021229 | -0.021229 | -0.26508
2019-03-28 14:59 | 1553799537.7524 | -101 | 0.20951 | 0.000444 | -1e-06 | 100.9028 | 101.0698 | 125.7261 | 21.2612 | 0.001862 | 0.001765 | -151.655 | 0.22967 | 8.401 | 0.002597 | 0.004097 | 0.003914 | 5.1e-05 | 0.001308 | 0.3361 | 1.3648 | 2.5475 | 100.7 | 21.15 | 25.887 | 66.6285 | 84.2824 | 64.2902 | 0.00646 | 1.0462 | 0.02521 | 24.0952 | -0.014788 | 0.13663 | -0.000109 | -0.008249 | -0.008249 | -0.68408
2019-03-28 14:59 | 1553799538.7518 | -100 | 0.20951 | 0.000438 | 0 | 100.8966 | 101.01 | 125.4571 | 21.2612 | 0.001875 | 0.001815 | -151.6781 | 0.2292 | 8.4038 | 0.002573 | 0.003546 | 0.00334 | 5e-05 | 0.00131 | 0.33699 | 1.3276 | 2.5475 | 100.7 | 21.15 | 25.89 | 66.9469 | 83.7089 | 65.4922 | 0.006461 | 1.0425 | 0.025138 | 24.1122 | -0.01005 | 0.13635 | -0.000122 | -0.009212 | -0.009212 | -0.61132
2019-03-28 14:59 | 1553799539.7525 | -99 | 0.20952 | 0.000441 | 0 | 100.9418 | 100.9132 | 125.6597 | 21.2612 | 0.001863 | 0.001792 | -151.6878 | 0.22942 | 8.4016 | 0.002589 | 0.004177 | 0.004002 | 5.3e-05 | 0.001306 | 0.31708 | 1.3242 | 2.5473 | 100.7 | 21.15 | 25.892 | 66.6518 | 84.0968 | 64.9497 | 0.006461 | 1.0442 | 0.025178 | 24.1124 | 0.002541 | 0.13648 | -0.000202 | -0.01535 | -0.01535 | -0.36721
2019-03-28 14:59 | 1553799540.7525 | -98 | 0.20949 | 0.000443 | 0 | 100.8909 | 101.0219 | 125.7892 | 21.2612 | 0.001872 | 0.001749 | -151.6874 | 0.22984 | 8.4023 | 0.002612 | 0.003742 | 0.003592 | 5.4e-05 | 0.001299 | 0.32947 | 1.3525 | 2.5475 | 100.7 | 21.15 | 25.892 | 66.868 | 84.6579 | 63.9232 | 0.006461 | 1.0474 | 0.02523 | 24.0881 | -0.011604 | 0.13673 | -9e-06 | -0.00062 | -0.00062 | -9.114
2019-03-28 14:59 | 1553799541.7515 | -97 | 0.20949 | 0.000443 | 0 | 100.8894 | 101.1235 | 125.9923 | 21.2612 | 0.001865 | 0.001786 | -151.7047 | 0.22983 | 8.4043 | 0.002553 | 0.003753 | 0.003567 | 5.3e-05 | 0.001293 | 0.31349 | 1.3096 | 2.5473 | 100.7 | 21.15 | 25.899 | 66.7073 | 83.23 | 64.7936 | 0.006463 | 1.0447 | 0.025217 | 24.1387 | -0.020725 | 0.13673 | -1.9e-05 | -0.001402 | -0.001402 | -4.028
2019-03-28 14:59 | 1553799542.7525 | -96 | 0.20949 | 0.000441 | 0 | 100.9025 | 101.0993 | 126.473 | 21.2612 | 0.001895 | 0.0018 | -151.7189 | 0.23057 | 8.401 | 0.002666 | 0.003822 | 0.003707 | 5.2e-05 | 0.001292 | 0.30613 | 1.31 | 2.5473 | 100.7 | 21.15 | 25.899 | 67.4297 | 85.9548 | 65.1316 | 0.006463 | 1.0436 | 0.025253 | 24.1968 | -0.01743 | 0.13716 | 2.5e-05 | 0.001885 | 0.001885 | 3.0047
2019-03-28 14:59 | 1553799543.7518 | -95 | 0.20951 | 0.000441 | 0 | 100.9701 | 101.1347 | 127.1232 | 21.2612 | 0.001891 | 0.001801 | -151.7119 | 0.23059 | 8.4002 | 0.002617 | 0.004054 | 0.003788 | 4.9e-05 | 0.001293 | 0.32421 | 1.3109 | 2.5474 | 100.7 | 21.15 | 25.908 | 67.3376 | 84.7726 | 65.1737 | 0.006465 | 1.0435 | 0.025316 | 24.2604 | -0.014564 | 0.13718 | -8.5e-05 | -0.006502 | -0.006502 | -0.8714
2019-03-28 14:59 | 1553799544.7516 | -94 | 0.20953 | 0.000443 | 0 | 100.8745 | 101.0093 | 127.2861 | 21.2612 | 0.001831 | 0.001722 | -151.739 | 0.22995 | 8.4013 | 0.002552 | 0.003895 | 0.0038 | 4.8e-05 | 0.001297 | 0.31045 | 1.3193 | 2.5473 | 100.7 | 21.15 | 25.908 | 65.8949 | 83.216 | 63.265 | 0.006465 | 1.0494 | 0.025404 | 24.2074 | -0.011944 | 0.1368 | -0.000211 | -0.016098 | -0.016098 | -0.35096
2019-03-28 14:59 | 1553799545.7527 | -93 | 0.20952 | 0.000442 | 0 | 100.9558 | 101.0467 | 127.3929 | 21.2612 | 0.001946 | 0.001734 | -151.734 | 0.22993 | 8.4052 | 0.002621 | 0.003588 | 0.003399 | 5.2e-05 | 0.001296 | 0.33196 | 1.3323 | 2.5474 | 100.7 | 21.15 | 25.917 | 68.6419 | 84.8613 | 63.5413 | 0.006468 | 1.0486 | 0.025404 | 24.2275 | -0.008045 | 0.13678 | -0.000162 | -0.012394 | -0.012394 | -0.45581
2019-03-28 14:59 | 1553799546.7517 | -92 | 0.20952 | 0.000439 | 0 | 100.9362 | 100.9115 | 127.3569 | 21.2612 | 0.001907 | 0.001778 | -151.743 | 0.22986 | 8.4049 | 0.002597 | 0.003882 | 0.003701 | 5.1e-05 | 0.001296 | 0.32781 | 1.3067 | 2.5475 | 100.7 | 21.15 | 25.917 | 67.7223 | 84.2841 | 64.5985 | 0.006468 | 1.0453 | 0.025361 | 24.2621 | 0.00219 | 0.13674 | -0.000186 | -0.014219 | -0.014219 | -0.39717
2019-03-28 14:59 | 1553799547.7517 | -91 | 0.20952 | 0.000443 | 0 | 100.9314 | 100.9682 | 127.272 | 21.2612 | 0.001901 | 0.001772 | -151.7427 | 0.23023 | 8.4039 | 0.002611 | 0.003833 | 0.003655 | 4.9e-05 | 0.001298 | 0.31636 | 1.3038 | 2.5473 | 100.7 | 21.15 | 25.921 | 67.5615 | 84.6405 | 64.4579 | 0.006469 | 1.0457 | 0.025358 | 24.2489 | -0.00326 | 0.13697 | -0.00014 | -0.01069 | -0.01069 | -0.52918
2019-03-28 14:59 | 1553799548.752 | -90 | 0.20953 | 0.00044 | 0 | 100.946 | 101.0796 | 127.3612 | 21.2612 | 0.001881 | 0.001775 | -151.7286 | 0.23024 | 8.3996 | 0.002579 | 0.004064 | 0.003874 | 4.8e-05 | 0.001295 | 0.32333 | 1.3194 | 2.5473 | 100.7 | 21.15 | 25.924 | 67.0834 | 83.8657 | 64.5337 | 0.006469 | 1.0455 | 0.025364 | 24.2601 | -0.011832 | 0.13697 | -0.000205 | -0.015642 | -0.015642 | -0.36166
2019-03-28 14:59 | 1553799549.7528 | -89 | 0.20954 | 0.000444 | 0 | 100.8516 | 101.1476 | 127.1042 | 21.2612 | 0.001874 | 0.001782 | -151.7165 | 0.23003 | 8.4065 | 0.002621 | 0.004303 | 0.00406 | 5e-05 | 0.001294 | 0.32012 | 1.3264 | 2.5473 | 100.7 | 21.15 | 25.926 | 66.9257 | 84.877 | 64.7132 | 0.00647 | 1.0449 | 0.025331 | 24.2421 | -0.026213 | 0.13685 | -0.0003 | -0.022886 | -0.022886 | -0.24696
2019-03-28 14:59 | 1553799550.752 | -88 | 0.20953 | 0.000442 | 0 | 100.9409 | 101.1306 | 127.06 | 21.2612 | 0.001911 | 0.001751 | -151.7205 | 0.23039 | 8.4016 | 0.002608 | 0.00429 | 0.004107 | 5.2e-05 | 0.001291 | 0.31653 | 1.3456 | 2.5475 | 100.7 | 21.15 | 25.929 | 67.8041 | 84.5689 | 63.9591 | 0.006471 | 1.0473 | 0.025355 | 24.2108 | -0.016789 | 0.13706 | -0.000217 | -0.01659 | -0.01659 | -0.34119
2019-03-28 14:59 | 1553799551.7522 | -87 | 0.20951 | 0.000444 | 0 | 100.8757 | 101.1618 | 126.7788 | 21.2612 | 0.00191 | 0.001768 | -151.7139 | 0.23085 | 8.4039 | 0.002573 | 0.00373 | 0.003583 | 4.7e-05 | 0.001295 | 0.30646 | 1.368 | 2.5475 | 100.7 | 21.15 | 25.929 | 67.7783 | 83.7193 | 64.3706 | 0.006471 | 1.046 | 0.025312 | 24.1988 | -0.025335 | 0.13733 | -0.0001 | -0.007628 | -0.007628 | -0.74351
2019-03-28 14:59 | 1553799552.752 | -86 | 0.20953 | 0.000442 | 0 | 100.991 | 101.0399 | 126.2793 | 21.2612 | 0.001831 | 0.001758 | -151.697 | 0.23009 | 8.4032 | 0.002602 | 0.003548 | 0.003405 | 4.9e-05 | 0.001294 | 0.29479 | 1.3351 | 2.5475 | 100.7 | 21.14 | 25.933 | 65.8761 | 84.4128 | 64.1254 | 0.006468 | 1.0468 | 0.025271 | 24.1423 | -0.00433 | 0.13688 | -0.000208 | -0.015817 | -0.015817 | -0.35742
2019-03-28 14:59 | 1553799553.7509 | -85 | 0.20952 | 0.000442 | -1e-06 | 100.9919 | 101.0518 | 125.8305 | 21.2612 | 0.001915 | 0.001743 | -151.6844 | 0.23038 | 8.4009 | 0.002603 | 0.003806 | 0.003677 | 5.1e-05 | 0.001286 | 0.2995 | 1.3053 | 2.5474 | 100.7 | 21.14 | 25.933 | 67.8993 | 84.4267 | 63.7569 | 0.006468 | 1.0479 | 0.02524 | 24.0862 | -0.005304 | 0.13705 | -0.000183 | -0.013866 | -0.013866 | -0.4082
2019-03-28 14:59 | 1553799554.752 | -84 | 0.20953 | 0.000443 | 0 | 100.9919 | 101.0698 | 125.9202 | 21.2612 | 0.001869 | 0.001759 | -151.6808 | 0.23057 | 8.4018 | 0.002571 | 0.004066 | 0.003842 | 5.1e-05 | 0.001288 | 0.30702 | 1.3211 | 2.5475 | 100.7 | 21.15 | 25.933 | 66.7891 | 83.6762 | 64.1511 | 0.006472 | 1.0467 | 0.025234 | 24.1088 | -0.006897 | 0.13716 | -0.000241 | -0.018341 | -0.018341 | -0.30887
2019-03-28 14:59 | 1553799555.752 | -83 | 0.20951 | 0.000444 | 0 | 100.9016 | 101.0671 | 126.0514 | 21.2612 | 0.001889 | 0.001731 | -151.6909 | 0.23101 | 8.4042 | 0.002584 | 0.003806 | 0.00372 | 4.8e-05 | 0.001298 | 0.34418 | 1.3065 | 2.5474 | 100.7 | 21.15 | 25.933 | 67.2816 | 83.9898 | 63.4821 | 0.006472 | 1.0488 | 0.025272 | 24.0975 | -0.014653 | 0.13743 | -0.000125 | -0.009488 | -0.009488 | -0.59822
2019-03-28 14:59 | 1553799556.7521 | -82 | 0.20954 | 0.00044 | 0 | 100.9228 | 100.9923 | 126.1469 | 21.2612 | 0.001892 | 0.001785 | -151.6778 | 0.22994 | 8.4019 | 0.002561 | 0.003987 | 0.003729 | 4.8e-05 | 0.001306 | 0.3314 | 1.3379 | 2.5475 | 100.7 | 21.15 | 25.932 | 67.3407 | 83.4375 | 64.7898 | 0.006471 | 1.0447 | 0.025233 | 24.1534 | -0.006162 | 0.13679 | -0.000263 | -0.019999 | -0.019999 | -0.28248
2019-03-28 14:59 | 1553799557.7511 | -81 | 0.20955 | 0.000441 | 0 | 100.9624 | 101.1401 | 126.3828 | 21.2612 | 0.001867 | 0.001812 | -151.6906 | 0.23051 | 8.4043 | 0.002609 | 0.004196 | 0.004062 | 4.9e-05 | 0.00131 | 0.32068 | 1.3025 | 2.5475 | 100.7 | 21.15 | 25.932 | 66.7439 | 84.5929 | 65.4251 | 0.006471 | 1.0427 | 0.025233 | 24.1986 | -0.01573 | 0.13713 | -0.000356 | -0.027063 | -0.027063 | -0.20927
2019-03-28 14:59 | 1553799558.7511 | -80 | 0.20952 | 0.000449 | -1e-06 | 100.9728 | 101.013 | 126.4157 | 21.2612 | 0.001867 | 0.001726 | -151.6948 | 0.22932 | 8.4038 | 0.002636 | 0.003772 | 0.003597 | 4.9e-05 | 0.00131 | 0.30525 | 1.2815 | 2.5474 | 100.7 | 21.15 | 25.933 | 66.7564 | 85.225 | 63.3624 | 0.006472 | 1.0491 | 0.025313 | 24.128 | -0.003563 | 0.13642 | -0.000178 | -0.013501 | -0.013501 | -0.41733
2019-03-28 14:59 | 1553799559.7511 | -79 | 0.20955 | 0.000449 | 0 | 101.0035 | 101.0728 | 126.396 | 21.2612 | 0.001941 | 0.001812 | -151.6844 | 0.22938 | 8.4014 | 0.002552 | 0.003505 | 0.003321 | 4.9e-05 | 0.001311 | 0.32372 | 1.3137 | 2.5474 | 100.7 | 21.15 | 25.933 | 68.531 | 83.2098 | 65.4369 | 0.006472 | 1.0427 | 0.025234 | 24.2003 | -0.006139 | 0.13646 | -0.000348 | -0.026484 | -0.026484 | -0.2128
2019-03-28 14:59 | 1553799560.7512 | -78 | 0.20952 | 0.00044 | 0 | 101.0252 | 101.1017 | 126.2845 | 21.2612 | 0.001875 | 0.001739 | -151.6746 | 0.22939 | 8.4034 | 0.002574 | 0.003683 | 0.003482 | 4.7e-05 | 0.001311 | 0.3063 | 1.3009 | 2.5474 | 100.7 | 21.15 | 25.93 | 66.9368 | 83.7332 | 63.6756 | 0.006471 | 1.0482 | 0.025288 | 24.1267 | -0.006769 | 0.13646 | -0.000147 | -0.01115 | -0.01115 | -0.50545
2019-03-28 14:59 | 1553799561.7522 | -77 | 0.20953 | 0.000441 | 0 | 100.9859 | 101.2913 | 126.239 | 21.2612 | 0.001877 | 0.001765 | -151.6779 | 0.2298 | 8.4033 | 0.002566 | 0.004077 | 0.003861 | 4.7e-05 | 0.001324 | 0.3246 | 1.3485 | 2.5474 | 100.7 | 21.14 | 25.929 | 67.0014 | 83.5546 | 64.2955 | 0.006467 | 1.0462 | 0.025261 | 24.1445 | -0.027006 | 0.13671 | -0.000255 | -0.01939 | -0.01939 | -0.29118
2019-03-28 14:59 | 1553799562.7513 | -76 | 0.20952 | 0.000443 | 0 | 100.9317 | 101.1771 | 126.0549 | 21.2612 | 0.001879 | 0.001785 | -151.6979 | 0.23053 | 8.4029 | 0.002561 | 0.003973 | 0.003797 | 5.1e-05 | 0.00133 | 0.31708 | 1.329 | 2.5474 | 100.7 | 21.15 | 25.929 | 67.0397 | 83.4343 | 64.7804 | 0.006471 | 1.0447 | 0.025224 | 24.1442 | -0.021723 | 0.13714 | -0.000171 | -0.012965 | -0.012965 | -0.43685
2019-03-28 14:59 | 1553799563.7524 | -75 | 0.20953 | 0.000446 | -1e-06 | 100.974 | 101.0525 | 125.6803 | 21.2612 | 0.00191 | 0.001788 | -151.7023 | 0.22928 | 8.4051 | 0.002611 | 0.003535 | 0.003389 | 5.2e-05 | 0.001331 | 0.29192 | 1.3822 | 2.5475 | 100.7 | 21.15 | 25.927 | 67.7815 | 84.6422 | 64.8538 | 0.00647 | 1.0445 | 0.025184 | 24.1109 | -0.006946 | 0.1364 | -0.000215 | -0.016242 | -0.016242 | -0.34685
2019-03-28 14:59 | 1553799564.7523 | -74 | 0.20954 | 0.000443 | 0 | 101.0115 | 101.0942 | 125.4092 | 21.2612 | 0.001918 | 0.001814 | -151.6916 | 0.22936 | 8.4041 | 0.002593 | 0.003511 | 0.003388 | 5e-05 | 0.001333 | 0.31349 | 1.3802 | 2.5473 | 100.7 | 21.15 | 25.927 | 67.9649 | 84.1911 | 65.4689 | 0.00647 | 1.0426 | 0.025134 | 24.1068 | -0.007321 | 0.13644 | -0.0003 | -0.022656 | -0.022656 | -0.24872
2019-03-28 14:59 | 1553799565.7517 | -73 | 0.20956 | 0.000441 | 0 | 101.0649 | 101.3052 | 125.1238 | 21.2612 | 0.001948 | 0.001801 | -151.689 | 0.2297 | 8.4011 | 0.002555 | 0.004098 | 0.003806 | 5.2e-05 | 0.001338 | 0.30397 | 1.3708 | 2.5473 | 100.7 | 21.15 | 25.929 | 68.6972 | 83.2776 | 65.1737 | 0.006471 | 1.0435 | 0.025116 | 24.0689 | -0.021247 | 0.13665 | -0.000383 | -0.028935 | -0.028935 | -0.19504
2019-03-28 14:59 | 1553799566.7514 | -72 | 0.20952 | 0.000444 | 0 | 101.0217 | 101.1024 | 124.8625 | 21.2612 | 0.001873 | 0.001777 | -151.6884 | 0.22968 | 8.4016 | 0.002637 | 0.004008 | 0.003911 | 5.7e-05 | 0.001341 | 0.32173 | 1.3187 | 2.5475 | 100.7 | 21.15 | 25.931 | 66.8843 | 85.2445 | 64.5891 | 0.006471 | 1.0453 | 0.025112 | 24.0229 | -0.007145 | 0.13664 | -0.000188 | -0.014161 | -0.014161 | -0.39851
2019-03-28 14:59 | 1553799567.7514 | -71 | 0.20949 | 0.00044 | 0 | 101.0729 | 101.1072 | 124.6931 | 21.2612 | 0.001855 | 0.001779 | -151.693 | 0.22945 | 8.402 | 0.0026 | 0.003891 | 0.003671 | 5.2e-05 | 0.001343 | 0.31382 | 1.3053 | 2.5474 | 100.7 | 21.15 | 25.933 | 66.4515 | 84.3666 | 64.6454 | 0.006472 | 1.0451 | 0.025093 | 24.0086 | -0.003031 | 0.1365 | 5e-05 | 0.003823 | 0.003823 | 1.4745
2019-03-28 14:59 | 1553799568.7504 | -70 | 0.20953 | 0.00044 | 0 | 101.0586 | 101.259 | 124.64 | 21.2612 | 0.001867 | 0.001755 | -151.7075 | 0.23002 | 8.4041 | 0.002585 | 0.003698 | 0.003521 | 5.3e-05 | 0.001341 | 0.3402 | 1.2923 | 2.5475 | 100.7 | 21.15 | 25.935 | 66.7564 | 84.0001 | 64.0638 | 0.006472 | 1.0469 | 0.025109 | 23.9829 | -0.017721 | 0.13684 | -0.000207 | -0.015673 | -0.015673 | -0.36059
2019-03-28 14:59 | 1553799569.7515 | -69 | 0.20952 | 0.000438 | 0 | 101.01 | 101.2169 | 124.8604 | 21.2612 | 0.001895 | 0.001779 | -151.7096 | 0.23092 | 8.4031 | 0.002616 | 0.0043 | 0.004091 | 5.4e-05 | 0.001333 | 0.32997 | 1.2921 | 2.5473 | 100.7 | 21.15 | 25.937 | 67.4126 | 84.7483 | 64.6322 | 0.006473 | 1.0452 | 0.02511 | 24.0243 | -0.018297 | 0.13737 | -0.000191 | -0.014491 | -0.014491 | -0.39152
2019-03-28 14:59 | 1553799570.7507 | -68 | 0.20954 | 0.000445 | -1e-06 | 101.0312 | 101.1727 | 124.7912 | 21.2612 | 0.001857 | 0.00181 | -151.7178 | 0.23093 | 8.4042 | 0.002618 | 0.003586 | 0.003376 | 5.4e-05 | 0.001324 | 0.31902 | 1.2955 | 2.5474 | 100.7 | 21.15 | 25.94 | 66.5162 | 84.7889 | 65.383 | 0.006473 | 1.0429 | 0.025075 | 24.0443 | -0.012521 | 0.13738 | -0.000261 | -0.019706 | -0.019706 | -0.28792
2019-03-28 15:00 | 1553799571.7515 | -67 | 0.20953 | 0.000445 | -1e-06 | 100.9949 | 101.1119 | 124.744 | 21.2612 | 0.001886 | 0.00177 | -151.7333 | 0.22992 | 8.4046 | 0.002642 | 0.004232 | 0.003981 | 5.6e-05 | 0.001315 | 0.3194 | 1.3088 | 2.5473 | 100.7 | 21.15 | 25.948 | 67.2066 | 85.3773 | 64.4204 | 0.006475 | 1.0458 | 0.025106 | 24.0055 | -0.010361 | 0.13678 | -0.000213 | -0.016095 | -0.016095 | -0.35098
2019-03-28 15:00 | 1553799572.7516 | -66 | 0.20956 | 0.000444 | 0 | 101.061 | 101.186 | 125.0107 | 21.2612 | 0.001913 | 0.001778 | -151.7227 | 0.23135 | 8.4024 | 0.002627 | 0.003938 | 0.003788 | 5.5e-05 | 0.001313 | 0.30718 | 1.3216 | 2.5474 | 100.7 | 21.15 | 25.948 | 67.8541 | 85.0153 | 64.6134 | 0.006475 | 1.0452 | 0.025126 | 24.0381 | -0.011055 | 0.13763 | -0.000406 | -0.030687 | -0.030687 | -0.18523
2019-03-28 15:00 | 1553799573.7516 | -65 | 0.20951 | 0.000443 | 0 | 101.0803 | 101.3242 | 125.5069 | 21.2612 | 0.001954 | 0.001809 | -151.7169 | 0.23062 | 8.402 | 0.002624 | 0.004121 | 0.003949 | 5.6e-05 | 0.001307 | 0.30397 | 1.3734 | 2.5475 | 100.7 | 21.15 | 25.964 | 68.833 | 84.9409 | 65.3448 | 0.006479 | 1.043 | 0.025148 | 24.1118 | -0.021556 | 0.1372 | -0.000118 | -0.008909 | -0.008909 | -0.63604
2019-03-28 15:00 | 1553799574.751 | -64 | 0.20953 | 0.000445 | 0 | 101.0952 | 101.1849 | 125.5749 | 21.2612 | 0.001878 | 0.001764 | -151.7195 | 0.22989 | 8.4018 | 0.002661 | 0.00376 | 0.003485 | 5.5e-05 | 0.001291 | 0.31454 | 1.332 | 2.5475 | 100.7 | 21.15 | 25.964 | 67.0216 | 85.8439 | 64.2683 | 0.006479 | 1.0463 | 0.025195 | 24.0799 | -0.007934 | 0.13676 | -0.000244 | -0.01851 | -0.01851 | -0.30515
2019-03-28 15:00 | 1553799575.7506 | -63 | 0.20951 | 0.000443 | 0 | 101.1021 | 101.1591 | 125.4773 | 21.2612 | 0.001879 | 0.001766 | -151.7078 | 0.2305 | 8.4025 | 0.002641 | 0.003591 | 0.003439 | 5.7e-05 | 0.001285 | 0.33627 | 1.3467 | 2.5475 | 100.7 | 21.15 | 25.973 | 67.0341 | 85.3508 | 64.3143 | 0.006482 | 1.0462 | 0.025184 | 24.0722 | -0.005046 | 0.13712 | -0.000128 | -0.009683 | -0.009683 | -0.58483
2019-03-28 15:00 | 1553799576.7517 | -62 | 0.20954 | 0.000442 | 0 | 101.1125 | 101.1306 | 125.5519 | 21.2612 | 0.001913 | 0.001794 | -151.7204 | 0.23033 | 8.4045 | 0.002621 | 0.0039 | 0.003729 | 5.8e-05 | 0.001277 | 0.33069 | 1.3319 | 2.5474 | 100.7 | 21.15 | 25.973 | 67.8526 | 84.8724 | 65.006 | 0.006482 | 1.044 | 0.025165 | 24.104 | -0.0016 | 0.13702 | -0.000278 | -0.021059 | -0.021059 | -0.26872
2019-03-28 15:00 | 1553799577.7518 | -61 | 0.20951 | 0.000447 | 0 | 101.1104 | 101.3749 | 125.3627 | 21.2612 | 0.001921 | 0.00178 | -151.6888 | 0.22966 | 8.4041 | 0.002638 | 0.003958 | 0.003837 | 5.8e-05 | 0.001278 | 0.31957 | 1.3014 | 2.5475 | 100.7 | 21.15 | 25.989 | 68.0498 | 85.2743 | 64.6478 | 0.006486 | 1.0451 | 0.02516 | 24.0731 | -0.023362 | 0.13662 | -0.000121 | -0.0092 | -0.0092 | -0.61332
2019-03-28 15:00 | 1553799578.7507 | -60 | 0.20953 | 0.000446 | 0 | 101.195 | 101.1849 | 125.2285 | 21.2612 | 0.00192 | 0.001822 | -151.6847 | 0.2304 | 8.4036 | 0.002632 | 0.003656 | 0.003523 | 5.8e-05 | 0.001274 | 0.31758 | 1.2932 | 2.5472 | 100.7 | 21.15 | 25.989 | 68.0365 | 85.137 | 65.6601 | 0.006486 | 1.042 | 0.025109 | 24.0962 | 0.000889 | 0.13706 | -0.000221 | -0.016713 | -0.016713 | -0.3387
2019-03-28 15:00 | 1553799579.7512 | -59 | 0.20955 | 0.000445 | 0 | 101.1846 | 101.2787 | 124.9149 | 21.2612 | 0.001914 | 0.001821 | -151.6784 | 0.22997 | 8.4007 | 0.002596 | 0.00406 | 0.003859 | 5.9e-05 | 0.00127 | 0.32853 | 1.3081 | 2.5473 | 100.7 | 21.14 | 26.004 | 67.8704 | 84.2762 | 65.6397 | 0.006485 | 1.0421 | 0.025078 | 24.0653 | -0.008317 | 0.13681 | -0.000356 | -0.026858 | -0.026858 | -0.21037
2019-03-28 15:00 | 1553799580.7509 | -58 | 0.2095 | 0.000447 | 0 | 101.1057 | 101.1921 | 125.1513 | 21.2612 | 0.001943 | 0.001763 | -151.6806 | 0.22976 | 8.4009 | 0.00258 | 0.004294 | 0.004084 | 6e-05 | 0.001276 | 0.29214 | 1.2924 | 2.5474 | 100.7 | 21.14 | 26.004 | 68.5741 | 83.8837 | 64.2433 | 0.006485 | 1.0464 | 0.025154 | 24.0384 | -0.007643 | 0.13668 | -3.4e-05 | -0.00253 | -0.00253 | -2.2308
2019-03-28 15:00 | 1553799581.7509 | -57 | 0.20952 | 0.000451 | 0 | 101.1155 | 101.3585 | 125.1228 | 21.2612 | 0.001936 | 0.001773 | -151.6687 | 0.22975 | 8.4016 | 0.00257 | 0.003604 | 0.003505 | 6.1e-05 | 0.001277 | 0.32244 | 1.2974 | 2.5473 | 100.7 | 21.15 | 26.013 | 68.4188 | 83.645 | 64.4806 | 0.006491 | 1.0457 | 0.025142 | 24.0441 | -0.021475 | 0.13668 | -0.000179 | -0.013506 | -0.013506 | -0.41793
2019-03-28 15:00 | 1553799582.7521 | -56 | 0.20952 | 0.000445 | 0 | 101.1542 | 101.4119 | 124.9308 | 21.2612 | 0.001911 | 0.001757 | -151.6749 | 0.23044 | 8.4044 | 0.002654 | 0.004375 | 0.004079 | 6.1e-05 | 0.001274 | 0.31902 | 1.341 | 2.5474 | 100.7 | 21.14 | 26.015 | 67.8096 | 85.6613 | 64.1136 | 0.006488 | 1.0468 | 0.025136 | 24.0126 | -0.022756 | 0.13709 | -0.000153 | -0.011574 | -0.011574 | -0.48916
2019-03-28 15:00 | 1553799583.7509 | -55 | 0.2096 | 0.000448 | 0 | 101.1777 | 101.3477 | 125.086 | 21.2612 | 0.001924 | 0.001756 | -151.6783 | 0.23006 | 8.4012 | 0.002617 | 0.003481 | 0.003332 | 6.2e-05 | 0.00127 | 0.31349 | 1.3222 | 2.5474 | 100.7 | 21.14 | 26.016 | 68.1209 | 84.774 | 64.0753 | 0.006488 | 1.0469 | 0.025153 | 24.0261 | -0.015011 | 0.13686 | -0.000655 | -0.049639 | -0.049639 | -0.11387
2019-03-28 15:00 | 1553799584.751 | -54 | 0.20952 | 0.000444 | 0 | 101.1873 | 101.2661 | 125.1829 | 21.2612 | 0.001897 | 0.001785 | -151.6891 | 0.23075 | 8.4035 | 0.00265 | 0.003857 | 0.00372 | 6.3e-05 | 0.001268 | 0.2975 | 1.3314 | 2.5474 | 100.7 | 21.14 | 26.016 | 67.4665 | 85.5577 | 64.7717 | 0.006488 | 1.0448 | 0.025137 | 24.0602 | -0.006969 | 0.13728 | -0.000188 | -0.014217 | -0.014217 | -0.39878
2019-03-28 15:00 | 1553799585.751 | -53 | 0.20959 | 0.000446 | 0 | 101.1107 | 101.2152 | 125.413 | 21.2612 | 0.001924 | 0.00183 | -151.6811 | 0.23086 | 8.4035 | 0.002631 | 0.003352 | 0.003197 | 6.2e-05 | 0.001263 | 0.34219 | 1.3082 | 2.5474 | 100.7 | 21.14 | 26.016 | 68.1168 | 85.1228 | 65.8629 | 0.006488 | 1.0414 | 0.02512 | 24.1212 | -0.009236 | 0.13734 | -0.000611 | -0.046184 | -0.046184 | -0.12281
2019-03-28 15:00 | 1553799586.751 | -52 | 0.20955 | 0.000444 | 0 | 101.178 | 101.4482 | 125.5608 | 21.2612 | 0.001953 | 0.001737 | -151.6721 | 0.23089 | 8.3998 | 0.002583 | 0.004123 | 0.003909 | 6.3e-05 | 0.001259 | 0.32875 | 1.314 | 2.5473 | 100.7 | 21.15 | 26.014 | 68.8219 | 83.9571 | 63.6342 | 0.006492 | 1.0483 | 0.025217 | 24.0559 | -0.023855 | 0.13735 | -0.000316 | -0.024021 | -0.024021 | -0.23616
2019-03-28 15:00 | 1553799587.751 | -51 | 0.20955 | 0.000447 | 0 | 101.1554 | 101.2988 | 125.8017 | 21.2612 | 0.00192 | 0.001787 | -151.7082 | 0.23066 | 8.3998 | 0.002563 | 0.004704 | 0.004456 | 6.4e-05 | 0.001254 | 0.33428 | 1.2764 | 2.5474 | 100.7 | 21.15 | 26.014 | 68.0351 | 83.475 | 64.8179 | 0.006492 | 1.0446 | 0.025197 | 24.1213 | -0.012667 | 0.13722 | -0.00032 | -0.024252 | -0.024252 | -0.23368
2019-03-28 15:00 | 1553799588.7505 | -50 | 0.20952 | 0.000448 | -1e-06 | 101.1229 | 101.451 | 126.1495 | 21.2612 | 0.00183 | 0.001786 | -151.6923 | 0.2293 | 8.4016 | 0.002685 | 0.003826 | 0.003693 | 6.4e-05 | 0.001246 | 0.31974 | 1.3214 | 2.5476 | 100.7 | 21.15 | 26.009 | 65.8699 | 86.4177 | 64.7953 | 0.00649 | 1.0447 | 0.025233 | 24.1538 | -0.028967 | 0.13641 | -0.000167 | -0.012637 | -0.012637 | -0.44583
2019-03-28 15:00 | 1553799589.7502 | -49 | 0.20955 | 0.000444 | 0 | 101.1944 | 101.3171 | 126.3209 | 21.2612 | 0.001909 | 0.001757 | -151.6978 | 0.23014 | 8.4011 | 0.002682 | 0.003835 | 0.003667 | 6.2e-05 | 0.001243 | 0.33019 | 1.2907 | 2.5474 | 100.7 | 21.15 | 26.009 | 67.7605 | 86.3506 | 64.1035 | 0.00649 | 1.0468 | 0.025276 | 24.1455 | -0.010838 | 0.13691 | -0.000313 | -0.023801 | -0.023801 | -0.23757
2019-03-28 15:00 | 1553799590.7501 | -48 | 0.20951 | 0.000437 | -1e-06 | 101.1742 | 101.205 | 126.2885 | 21.2612 | 0.001882 | 0.001784 | -151.6895 | 0.22991 | 8.4034 | 0.002697 | 0.004218 | 0.003955 | 6.4e-05 | 0.001245 | 0.32731 | 1.3213 | 2.5475 | 100.7 | 21.15 | 26.001 | 67.1113 | 86.7096 | 64.7428 | 0.006489 | 1.0448 | 0.025249 | 24.1652 | -0.002725 | 0.13677 | -8e-05 | -0.006055 | -0.006055 | -0.93294
2019-03-28 15:00 | 1553799591.7515 | -47 | 0.20953 | 0.00044 | 0 | 101.1694 | 101.223 | 126.4138 | 21.2612 | 0.001833 | 0.00176 | -151.6844 | 0.22997 | 8.4037 | 0.002705 | 0.004036 | 0.003915 | 6.3e-05 | 0.001239 | 0.31598 | 1.2853 | 2.5474 | 100.7 | 21.15 | 26.001 | 65.9425 | 86.9026 | 64.1668 | 0.006489 | 1.0466 | 0.025283 | 24.1566 | -0.004738 | 0.13681 | -0.000205 | -0.015568 | -0.015568 | -0.36294
2019-03-28 15:00 | 1553799592.7502 | -46 | 0.2095 | 0.000438 | 0 | 101.1399 | 101.2318 | 126.581 | 21.2612 | 0.001891 | 0.001797 | -151.6837 | 0.23118 | 8.4035 | 0.002725 | 0.004362 | 0.00415 | 6e-05 | 0.00124 | 0.32764 | 1.2641 | 2.5473 | 100.7 | 21.14 | 25.993 | 67.3167 | 87.3677 | 65.059 | 0.006483 | 1.0439 | 0.025266 | 24.2045 | -0.008125 | 0.13753 | -6.5e-05 | -0.004918 | -0.004918 | -1.1549
2019-03-28 15:00 | 1553799593.7516 | -45 | 0.2095 | 0.000437 | -1e-06 | 101.1745 | 101.4003 | 126.3413 | 21.2612 | 0.001923 | 0.001762 | -151.6827 | 0.23007 | 8.404 | 0.002693 | 0.004055 | 0.003904 | 6.7e-05 | 0.001245 | 0.30165 | 1.3192 | 2.5475 | 100.7 | 21.14 | 25.993 | 68.0921 | 86.6073 | 64.2291 | 0.006483 | 1.0464 | 0.025273 | 24.1519 | -0.019947 | 0.13687 | -5e-05 | -0.003761 | -0.003761 | -1.5032
2019-03-28 15:00 | 1553799594.7503 | -44 | 0.20954 | 0.000442 | 0 | 101.1754 | 101.2464 | 126.1387 | 21.2612 | 0.001928 | 0.001772 | -151.6849 | 0.23023 | 8.4032 | 0.002641 | 0.004549 | 0.004369 | 6.5e-05 | 0.001245 | 0.31492 | 1.3308 | 2.5473 | 100.7 | 21.15 | 25.99 | 68.2168 | 85.3554 | 64.4618 | 0.006486 | 1.0457 | 0.025244 | 24.1409 | -0.006282 | 0.13697 | -0.000258 | -0.019566 | -0.019566 | -0.28911
2019-03-28 15:00 | 1553799595.7502 | -43 | 0.20952 | 0.000441 | -1e-06 | 101.1009 | 101.1466 | 125.943 | 21.2612 | 0.001905 | 0.001794 | -151.6976 | 0.23092 | 8.4053 | 0.002657 | 0.003917 | 0.003737 | 6.4e-05 | 0.001248 | 0.33229 | 1.3265 | 2.5474 | 100.7 | 21.15 | 25.99 | 67.6583 | 85.7333 | 65.0022 | 0.006486 | 1.044 | 0.025205 | 24.1414 | -0.004039 | 0.13738 | -0.000187 | -0.014177 | -0.014177 | -0.40021
2019-03-28 15:00 | 1553799596.7506 | -42 | 0.2095 | 0.000444 | 0 | 101.0744 | 101.1676 | 125.7904 | 21.2612 | 0.001891 | 0.001775 | -151.7104 | 0.23018 | 8.402 | 0.00267 | 0.003418 | 0.003213 | 6.3e-05 | 0.00125 | 0.31509 | 1.2965 | 2.5473 | 100.7 | 21.14 | 25.987 | 67.3313 | 86.0515 | 64.5352 | 0.006481 | 1.0455 | 0.025207 | 24.1101 | -0.008247 | 0.13693 | -2.1e-05 | -0.001612 | -0.001612 | -3.508
2019-03-28 15:00 | 1553799597.7513 | -41 | 0.20954 | 0.000441 | 0 | 101.1456 | 101.3246 | 125.7913 | 21.2612 | 0.001829 | 0.001827 | -151.6935 | 0.23072 | 8.4013 | 0.002606 | 0.003586 | 0.003405 | 6.1e-05 | 0.001247 | 0.29175 | 1.2836 | 2.5474 | 100.7 | 21.15 | 25.986 | 65.8403 | 84.5039 | 65.7833 | 0.006485 | 1.0416 | 0.02516 | 24.1547 | -0.015816 | 0.13726 | -0.000285 | -0.021566 | -0.021566 | -0.26286
2019-03-28 15:00 | 1553799598.7511 | -40 | 0.20953 | 0.000442 | -1e-06 | 101.0467 | 101.3375 | 125.5242 | 21.2612 | 0.001862 | 0.001786 | -151.7186 | 0.23079 | 8.405 | 0.002625 | 0.003492 | 0.003258 | 6e-05 | 0.001245 | 0.34899 | 1.291 | 2.5474 | 100.7 | 21.15 | 25.988 | 66.6184 | 84.9597 | 64.803 | 0.006485 | 1.0447 | 0.02517 | 24.0941 | -0.025705 | 0.1373 | -0.000224 | -0.016956 | -0.016956 | -0.33441
2019-03-28 15:00 | 1553799599.7495 | -39 | 0.20952 | 0.000441 | -1e-06 | 101.0893 | 101.0535 | 125.5871 | 21.2612 | 0.001904 | 0.001764 | -151.72 | 0.22915 | 8.4024 | 0.002595 | 0.003978 | 0.003763 | 6.4e-05 | 0.00124 | 0.29551 | 1.3149 | 2.5474 | 100.7 | 21.15 | 25.99 | 67.6302 | 84.2435 | 64.2635 | 0.006486 | 1.0463 | 0.025197 | 24.0809 | 0.003168 | 0.13632 | -0.000155 | -0.01172 | -0.01172 | -0.48038
2019-03-28 15:00 | 1553799600.7504 | -38 | 0.20947 | 0.000442 | 0 | 101.0833 | 101.1754 | 125.7263 | 21.2612 | 0.001928 | 0.001819 | -151.7361 | 0.23056 | 8.4018 | 0.002606 | 0.003775 | 0.003669 | 6e-05 | 0.001238 | 0.33229 | 1.3407 | 2.5474 | 100.7 | 21.15 | 25.99 | 68.213 | 84.5143 | 65.5938 | 0.006486 | 1.0422 | 0.025161 | 24.1417 | -0.008148 | 0.13716 | 0.000151 | 0.011407 | 0.011407 | 0.49661
2019-03-28 15:00 | 1553799601.7499 | -37 | 0.20948 | 0.000436 | 0 | 101.0038 | 101.1523 | 125.7085 | 21.2612 | 0.001878 | 0.001806 | -151.7315 | 0.22993 | 8.4024 | 0.002681 | 0.003837 | 0.003616 | 6.1e-05 | 0.00124 | 0.35419 | 1.3654 | 2.5474 | 100.7 | 21.15 | 25.995 | 67.006 | 86.3248 | 65.2885 | 0.006487 | 1.0432 | 0.025171 | 24.1291 | -0.013145 | 0.13678 | 7.8e-05 | 0.005907 | 0.005907 | 0.95631
2019-03-28 15:00 | 1553799602.7505 | -36 | 0.20952 | 0.00044 | 0 | 100.9433 | 101.2413 | 126.0427 | 21.2612 | 0.001971 | 0.001791 | -151.746 | 0.23014 | 8.4005 | 0.002602 | 0.00381 | 0.00366 | 6.1e-05 | 0.001241 | 0.31382 | 1.3507 | 2.5475 | 100.7 | 21.15 | 25.995 | 69.2618 | 84.4229 | 64.9295 | 0.006487 | 1.0443 | 0.025217 | 24.1484 | -0.026368 | 0.13691 | -0.000128 | -0.009702 | -0.009702 | -0.5828
2019-03-28 15:00 | 1553799603.7506 | -35 | 0.20951 | 0.00044 | -1e-06 | 100.929 | 101.0626 | 126.1873 | 21.2612 | 0.001921 | 0.001799 | -151.7488 | 0.23079 | 8.401 | 0.0026 | 0.00412 | 0.003991 | 6.4e-05 | 0.00124 | 0.31094 | 1.3241 | 2.5474 | 100.7 | 21.15 | 26.004 | 68.0389 | 84.3697 | 65.1145 | 0.006489 | 1.0437 | 0.025225 | 24.1688 | -0.011834 | 0.1373 | -7.8e-05 | -0.005861 | -0.005861 | -0.96757
2019-03-28 15:00 | 1553799604.7496 | -34 | 0.20953 | 0.000445 | 0 | 100.8656 | 100.8517 | 126.3436 | 21.2612 | 0.001976 | 0.001825 | -151.7351 | 0.22921 | 8.4029 | 0.00259 | 0.004051 | 0.003907 | 6.2e-05 | 0.001243 | 0.28279 | 1.345 | 2.5475 | 100.7 | 21.15 | 26.012 | 69.3594 | 84.124 | 65.7405 | 0.006491 | 1.0418 | 0.025217 | 24.2062 | 0.001234 | 0.13636 | -0.000215 | -0.016336 | -0.016336 | -0.34473
2019-03-28 15:00 | 1553799605.7497 | -33 | 0.20944 | 0.000446 | -1e-06 | 100.9359 | 100.991 | 126.8467 | 21.2612 | 0.001866 | 0.001794 | -151.7285 | 0.22963 | 8.4009 | 0.00263 | 0.003453 | 0.003287 | 6e-05 | 0.001251 | 0.3402 | 1.3312 | 2.5473 | 100.7 | 21.15 | 26.016 | 66.7167 | 85.0886 | 64.9974 | 0.006492 | 1.0441 | 0.025295 | 24.2277 | -0.00488 | 0.13661 | 0.000295 | 0.022572 | 0.022572 | 0.24995
2019-03-28 15:00 | 1553799606.7496 | -32 | 0.2095 | 0.000448 | -1e-06 | 100.8295 | 100.8853 | 127.0788 | 21.2612 | 0.001926 | 0.001852 | -151.7223 | 0.22881 | 8.4031 | 0.002672 | 0.004038 | 0.00387 | 6.3e-05 | 0.001251 | 0.29463 | 1.3041 | 2.5475 | 100.7 | 21.15 | 26.016 | 68.1724 | 86.1016 | 66.3781 | 0.006492 | 1.0398 | 0.025267 | 24.2993 | -0.004945 | 0.13612 | -4.8e-05 | -0.003617 | -0.003617 | -1.5544
2019-03-28 15:00 | 1553799607.7509 | -31 | 0.20952 | 0.000442 | -1e-06 | 100.845 | 100.7983 | 127.1687 | 21.2612 | 0.00192 | 0.001782 | -151.7269 | 0.22951 | 8.3998 | 0.002704 | 0.004102 | 0.00388 | 6.3e-05 | 0.00125 | 0.29878 | 1.3131 | 2.5475 | 100.7 | 21.15 | 26.034 | 68.0296 | 86.8737 | 64.7132 | 0.006497 | 1.0449 | 0.025338 | 24.2482 | 0.004143 | 0.13653 | -0.000161 | -0.012195 | -0.012195 | -0.46238
2019-03-28 15:00 | 1553799608.7497 | -30 | 0.20951 | 0.000447 | -1e-06 | 100.7128 | 100.849 | 127.2079 | 21.2612 | 0.001983 | 0.001814 | -151.7138 | 0.22891 | 8.4062 | 0.002727 | 0.004251 | 0.004059 | 6.2e-05 | 0.001256 | 0.34219 | 1.3136 | 2.5475 | 100.7 | 21.15 | 26.034 | 69.5363 | 87.4272 | 65.4679 | 0.006497 | 1.0426 | 0.025313 | 24.279 | -0.012085 | 0.13618 | -8.4e-05 | -0.006347 | -0.006347 | -0.88615
2019-03-28 15:00 | 1553799609.7497 | -29 | 0.20952 | 0.000443 | -1e-06 | 100.7259 | 100.9787 | 127.0415 | 21.2612 | 0.001869 | 0.00179 | -151.7151 | 0.22992 | 8.4022 | 0.002714 | 0.004437 | 0.004246 | 6.1e-05 | 0.001252 | 0.30381 | 1.2947 | 2.5474 | 100.7 | 21.15 | 26.052 | 66.8002 | 87.1124 | 64.8997 | 0.006501 | 1.0444 | 0.025318 | 24.2428 | -0.022424 | 0.13678 | -0.000126 | -0.009512 | -0.009512 | -0.59391
2019-03-28 15:00 | 1553799610.7508 | -28 | 0.20947 | 0.000445 | -1e-06 | 100.68 | 100.6835 | 127.0535 | 21.2612 | 0.001903 | 0.001779 | -151.7129 | 0.22878 | 8.4025 | 0.002709 | 0.004084 | 0.003869 | 6e-05 | 0.001256 | 0.33107 | 1.3257 | 2.5473 | 100.7 | 21.15 | 26.052 | 67.6177 | 86.9962 | 64.6367 | 0.006501 | 1.0452 | 0.025329 | 24.2345 | -0.00031 | 0.1361 | 0.000124 | 0.009522 | 0.009522 | 0.59031
2019-03-28 15:00 | 1553799611.7511 | -27 | 0.20947 | 0.000445 | -1e-06 | 100.6112 | 100.588 | 126.7664 | 21.2612 | 0.001949 | 0.001809 | -151.7191 | 0.22971 | 8.404 | 0.002777 | 0.003983 | 0.003909 | 5.9e-05 | 0.001254 | 0.2907 | 1.311 | 2.5474 | 100.7 | 21.15 | 26.065 | 68.719 | 88.637 | 65.3566 | 0.006504 | 1.043 | 0.025274 | 24.2329 | 0.00206 | 0.13665 | 0.000131 | 0.009986 | 0.009986 | 0.56514
2019-03-28 15:00 | 1553799612.7512 | -26 | 0.20952 | 0.000445 | -1e-06 | 100.5094 | 100.6516 | 126.6319 | 21.2612 | 0.001947 | 0.001777 | -151.7092 | 0.22959 | 8.4027 | 0.002827 | 0.003673 | 0.003461 | 6.5e-05 | 0.001263 | 0.32587 | 1.3648 | 2.5474 | 100.7 | 21.15 | 26.072 | 68.6815 | 89.8174 | 64.5807 | 0.006506 | 1.0453 | 0.025289 | 24.1923 | -0.012648 | 0.13658 | -0.000136 | -0.010326 | -0.010326 | -0.54629
2019-03-28 15:00 | 1553799613.7499 | -25 | 0.20952 | 0.000447 | -1e-06 | 100.5731 | 100.7453 | 126.5232 | 21.2612 | 0.00192 | 0.001775 | -151.7142 | 0.22915 | 8.4019 | 0.002775 | 0.003806 | 0.003691 | 6.4e-05 | 0.001268 | 0.29502 | 1.3653 | 2.5475 | 100.7 | 21.15 | 26.073 | 68.025 | 88.5657 | 64.5414 | 0.006506 | 1.0455 | 0.02528 | 24.1805 | -0.015306 | 0.13632 | -0.000141 | -0.010686 | -0.010686 | -0.5269
2019-03-28 15:00 | 1553799614.7503 | -24 | 0.20948 | 0.000441 | 0 | 100.5007 | 100.4715 | 126.5298 | 21.2612 | 0.001966 | 0.001844 | -151.7109 | 0.22924 | 8.4037 | 0.002722 | 0.004051 | 0.003953 | 6.3e-05 | 0.001278 | 0.31741 | 1.3587 | 2.5474 | 100.7 | 21.14 | 26.073 | 69.1339 | 87.3006 | 66.197 | 0.006502 | 1.0404 | 0.025219 | 24.2403 | 0.002599 | 0.13637 | 6.6e-05 | 0.005061 | 0.005061 | 1.1128
2019-03-28 15:00 | 1553799615.7503 | -23 | 0.20953 | 0.000447 | 0 | 100.3997 | 100.6614 | 126.227 | 21.2612 | 0.001974 | 0.001822 | -151.7133 | 0.22988 | 8.4031 | 0.002756 | 0.004076 | 0.003922 | 6.5e-05 | 0.001285 | 0.30486 | 1.2969 | 2.5473 | 100.7 | 21.15 | 26.07 | 69.3337 | 88.1105 | 65.6765 | 0.006506 | 1.042 | 0.025208 | 24.1927 | -0.023283 | 0.13676 | -0.000245 | -0.018577 | -0.018577 | -0.30404
2019-03-28 15:00 | 1553799616.749 | -22 | 0.20956 | 0.000445 | -1e-06 | 100.3831 | 100.5656 | 126.2138 | 21.2612 | 0.001927 | 0.001799 | -151.7159 | 0.2295 | 8.4012 | 0.002703 | 0.003947 | 0.003704 | 6.7e-05 | 0.001282 | 0.33069 | 1.3161 | 2.5474 | 100.7 | 21.15 | 26.066 | 68.2043 | 86.8386 | 65.1044 | 0.006505 | 1.0437 | 0.025228 | 24.171 | -0.016252 | 0.13653 | -0.000414 | -0.031394 | -0.031394 | -0.17961
2019-03-28 15:00 | 1553799617.7504 | -21 | 0.20954 | 0.00045 | -1e-06 | 100.3074 | 100.3913 | 126.2373 | 21.2612 | 0.001947 | 0.001787 | -151.7084 | 0.22964 | 8.402 | 0.002733 | 0.004201 | 0.004029 | 6.4e-05 | 0.00128 | 0.30215 | 1.3056 | 2.5474 | 100.7 | 21.15 | 26.052 | 68.6762 | 87.5583 | 64.8193 | 0.006501 | 1.0446 | 0.025241 | 24.1631 | -0.007482 | 0.13661 | -0.000271 | -0.02057 | -0.02057 | -0.27428
2019-03-28 15:00 | 1553799618.7489 | -20 | 0.20952 | 0.000439 | 0 | 100.2958 | 100.4569 | 126.3098 | 21.2612 | 0.001892 | 0.00177 | -151.7042 | 0.22953 | 8.4042 | 0.002631 | 0.003599 | 0.003377 | 6.2e-05 | 0.001283 | 0.3131 | 1.3728 | 2.5474 | 100.7 | 21.15 | 26.044 | 67.3633 | 85.1158 | 64.4214 | 0.006499 | 1.0458 | 0.025263 | 24.1558 | -0.014357 | 0.13654 | -0.00018 | -0.013634 | -0.013634 | -0.41363
2019-03-28 15:00 | 1553799619.7494 | -19 | 0.20953 | 0.000447 | -1e-06 | 100.2097 | 100.3716 | 126.6622 | 21.2612 | 0.001898 | 0.001784 | -151.7119 | 0.23003 | 8.403 | 0.002671 | 0.004087 | 0.003854 | 6.1e-05 | 0.001282 | 0.31293 | 1.3209 | 2.5473 | 100.7 | 21.15 | 26.035 | 67.503 | 86.0734 | 64.7633 | 0.006497 | 1.0448 | 0.025285 | 24.2017 | -0.014442 | 0.13685 | -0.000208 | -0.015769 | -0.015769 | -0.35841
2019-03-28 15:00 | 1553799620.7491 | -18 | 0.20951 | 0.000448 | 0 | 100.1549 | 100.3299 | 126.6457 | 21.2612 | 0.00192 | 0.001806 | -151.7186 | 0.22943 | 8.4033 | 0.002655 | 0.00412 | 0.0039 | 6.2e-05 | 0.001284 | 0.32388 | 1.3321 | 2.5473 | 100.7 | 21.15 | 26.015 | 68.024 | 85.6794 | 65.2846 | 0.006492 | 1.0432 | 0.025264 | 24.2188 | -0.01561 | 0.13649 | -0.000119 | -0.009019 | -0.009019 | -0.62501
2019-03-28 15:00 | 1553799621.7491 | -17 | 0.20953 | 0.000445 | 0 | 100.1501 | 100.0384 | 127.1567 | 21.2612 | 0.001973 | 0.00181 | -151.7362 | 0.22931 | 8.4018 | 0.00262 | 0.003673 | 0.003512 | 6e-05 | 0.001282 | 0.31349 | 1.2841 | 2.5474 | 100.7 | 21.15 | 26.015 | 69.2938 | 84.845 | 65.3869 | 0.006492 | 1.0429 | 0.025311 | 24.2713 | 0.009988 | 0.13642 | -0.000223 | -0.017005 | -0.017005 | -0.33131
2019-03-28 15:00 | 1553799622.7501 | -16 | 0.20953 | 0.000442 | 0 | 100.0334 | 100.0591 | 127.0955 | 21.2612 | 0.001841 | 0.001762 | -151.7314 | 0.22994 | 8.4005 | 0.002678 | 0.003637 | 0.00341 | 6.2e-05 | 0.001277 | 0.32709 | 1.307 | 2.5474 | 100.7 | 21.14 | 25.982 | 66.1352 | 86.2522 | 64.2332 | 0.00648 | 1.0464 | 0.025349 | 24.224 | -0.002299 | 0.13679 | -0.000202 | -0.015389 | -0.015389 | -0.36712
2019-03-28 15:00 | 1553799623.7501 | -15 | 0.20951 | 0.000446 | 0 | 99.9673 | 99.8546 | 127.6816 | 21.2612 | 0.00188 | 0.001821 | -151.7527 | 0.23103 | 8.4016 | 0.00266 | 0.003695 | 0.003506 | 6.1e-05 | 0.001273 | 0.33445 | 1.2991 | 2.5475 | 100.7 | 21.14 | 25.982 | 67.056 | 85.8088 | 65.6484 | 0.00648 | 1.0421 | 0.025354 | 24.3307 | 0.010088 | 0.13744 | -7.4e-05 | -0.005593 | -0.005593 | -1.0148
2019-03-28 15:00 | 1553799624.7505 | -14 | 0.20952 | 0.000444 | 0 | 99.8577 | 99.8077 | 127.871 | 21.2612 | 0.001873 | 0.001706 | -151.7571 | 0.22975 | 8.4023 | 0.002625 | 0.004197 | 0.00396 | 6.3e-05 | 0.001276 | 0.32587 | 1.3078 | 2.5473 | 100.7 | 21.14 | 25.962 | 66.8923 | 84.9652 | 62.8846 | 0.006475 | 1.0506 | 0.025477 | 24.2492 | 0.004476 | 0.13668 | -0.000151 | -0.011622 | -0.011622 | -0.48571
2019-03-28 15:00 | 1553799625.7502 | -13 | 0.20954 | 0.000447 | 0 | 99.8157 | 99.6776 | 127.7799 | 21.2612 | 0.001914 | 0.001765 | -151.7861 | 0.2298 | 8.402 | 0.00268 | 0.003909 | 0.003781 | 6.3e-05 | 0.001277 | 0.31133 | 1.315 | 2.5474 | 100.7 | 21.14 | 25.962 | 67.8923 | 86.288 | 64.2871 | 0.006475 | 1.0463 | 0.025415 | 24.2911 | 0.012381 | 0.13671 | -0.000267 | -0.020403 | -0.020403 | -0.27672
2019-03-28 15:00 | 1553799626.7507 | -12 | 0.20952 | 0.000448 | 0 | 99.6632 | 99.644 | 127.6217 | 21.2612 | 0.001852 | 0.001731 | -151.7879 | 0.2294 | 8.4019 | 0.002767 | 0.00369 | 0.003469 | 6.3e-05 | 0.001274 | 0.31636 | 1.2841 | 2.5473 | 100.7 | 21.14 | 25.946 | 66.4007 | 88.3885 | 63.4765 | 0.006471 | 1.0488 | 0.02543 | 24.2469 | 0.001724 | 0.13647 | -0.000145 | -0.011117 | -0.011117 | -0.50699
2019-03-28 15:00 | 1553799627.7504 | -11 | 0.20954 | 0.000446 | 0 | 99.6346 | 99.6796 | 127.5806 | 21.2612 | 0.001879 | 0.001718 | -131.1774 | 0.22929 | 8.4046 | 0.00278 | 0.003801 | 0.003695 | 6e-05 | 0.001279 | 0.3079 | 1.317 | 2.5474 | 100.7 | 21.14 | 25.93 | 67.0303 | 88.7073 | 63.1679 | 0.006467 | 1.0497 | 0.025437 | 24.2319 | -0.004044 | 0.1364 | -0.000289 | -0.022144 | -0.022144 | -0.25441
2019-03-28 15:00 | 1553799628.7494 | -10 | 0.20953 | 0.000443 | -1e-06 | 99.5401 | 99.4143 | 127.7665 | 21.2612 | 0.001854 | 0.001683 | -89.9162 | 0.22989 | 8.4006 | 0.002783 | 0.00402 | 0.003761 | 5.9e-05 | 0.001278 | 0.36105 | 1.3115 | 2.5474 | 100.7 | 21.14 | 25.923 | 66.439 | 88.7651 | 62.3282 | 0.006465 | 1.0524 | 0.025487 | 24.2192 | 0.011316 | 0.13676 | -0.000196 | -0.015027 | -0.015027 | -0.37586
2019-03-28 15:00 | 1553799629.7494 | -9 | 0.20952 | 0.000444 | -1e-06 | 99.4687 | 99.4374 | 127.8346 | 21.2612 | 0.00189 | 0.001726 | -48.6223 | 0.22996 | 8.4045 | 0.002841 | 0.004103 | 0.003936 | 6.1e-05 | 0.001275 | 0.33035 | 1.3193 | 2.5475 | 100.7 | 21.14 | 25.91 | 67.315 | 90.162 | 63.3516 | 0.006462 | 1.0492 | 0.025455 | 24.2626 | 0.002812 | 0.1368 | -0.000178 | -0.013591 | -0.013591 | -0.4157
2019-03-28 15:00 | 1553799630.7484 | -8 | 0.20952 | 0.000447 | 0 | 99.4228 | 99.272 | 127.9905 | 21.2612 | 0.00188 | 0.001788 | -7.3425 | 0.22962 | 8.4015 | 0.002794 | 0.004167 | 0.004017 | 6.2e-05 | 0.001274 | 0.31454 | 1.3714 | 2.5473 | 100.7 | 21.14 | 25.904 | 67.0654 | 89.0432 | 64.8499 | 0.00646 | 1.0445 | 0.025414 | 24.3314 | 0.01358 | 0.1366 | -0.00013 | -0.009956 | -0.009956 | -0.56663
2019-03-28 15:00 | 1553799631.7497 | -7 | 0.20956 | 0.000443 | 0 | 99.2399 | 99.134 | 128.2212 | 21.2612 | 0.001871 | 0.001724 | 33.9728 | 0.22983 | 8.4031 | 0.002808 | 0.004343 | 0.004162 | 6e-05 | 0.001281 | 0.34274 | 1.3492 | 2.5475 | 100.7 | 21.14 | 25.899 | 66.8572 | 89.362 | 63.3165 | 0.006459 | 1.0493 | 0.025495 | 24.298 | 0.009549 | 0.13673 | -0.000386 | -0.029655 | -0.029655 | -0.19042
2019-03-28 15:01 | 1553799632.7484 | -6 | 0.20951 | 0.000446 | -1e-06 | 99.2539 | 99.098 | 128.2048 | 21.2612 | 0.001859 | 0.001725 | 75.2717 | 0.24608 | 8.4038 | 0.002825 | 0.004683 | 0.004491 | 5.8e-05 | 0.001286 | 0.32659 | 1.3217 | 2.5474 | 100.7 | 21.14 | 25.889 | 66.5691 | 89.7767 | 63.3343 | 0.006457 | 1.0492 | 0.025493 | 24.2971 | 0.01406 | 0.14639 | -0.000101 | -0.007686 | -0.007686 | -0.78665
2019-03-28 15:01 | 1553799633.7498 | -5 | 0.20952 | 0.000442 | -1e-06 | 99.122 | 99.1283 | 127.6398 | 21.2612 | 0.00187 | 0.001683 | 260.9227 | 0.26194 | 8.4025 | 0.002804 | 0.004615 | 0.00458 | 5.8e-05 | 0.001277 | 0.31133 | 1.32 | 2.5475 | 100.7 | 21.14 | 25.884 | 66.8165 | 89.2813 | 62.3205 | 0.006455 | 1.0524 | 0.025475 | 24.2069 | -0.000567 | 0.15583 | -0.000145 | -0.011051 | -0.011051 | -0.58238
2019-03-28 15:01 | 1553799634.7499 | -4 | 0.20953 | 0.000443 | -1e-06 | 99.0969 | 98.9108 | 127.4217 | 21.2612 | 0.00182 | 0.002001 | 261.1849 | 0.2621 | 8.403 | 0.002844 | 0.004925 | 0.004916 | 5.8e-05 | 0.001275 | 0.30071 | 1.3505 | 2.5473 | 100.7 | 21.14 | 25.872 | 65.6289 | 90.2292 | 69.9605 | 0.006452 | 1.029 | 0.025168 | 24.4601 | 0.016815 | 0.15592 | -0.000213 | -0.016119 | -0.016119 | -0.39949
2019-03-28 15:01 | 1553799635.7486 | -3 | 0.20952 | 0.000444 | -1e-06 | 99.0621 | 98.9366 | 127.1879 | 21.2612 | 0.001861 | 0.001693 | 261.0061 | 0.26184 | 8.4044 | 0.002876 | 0.004614 | 0.004519 | 5.4e-05 | 0.001281 | 0.36265 | 1.3584 | 2.5475 | 100.7 | 21.14 | 25.866 | 66.6121 | 91.0179 | 62.569 | 0.006451 | 1.0516 | 0.025421 | 24.173 | 0.011334 | 0.15577 | -0.000153 | -0.011675 | -0.011675 | -0.55103
2019-03-28 15:01 | 1553799636.7497 | -2 | 0.20953 | 0.000444 | -1e-06 | 99.0064 | 98.8585 | 127.2689 | 21.2612 | 0.001861 | 0.001689 | 261.2755 | 0.26158 | 8.403 | 0.002801 | 0.004208 | 0.004146 | 5.9e-05 | 0.001285 | 0.35098 | 1.2933 | 2.5474 | 100.7 | 21.14 | 25.866 | 66.5979 | 89.1985 | 62.4628 | 0.006451 | 1.0519 | 0.025433 | 24.1769 | 0.013371 | 0.15561 | -0.00023 | -0.017542 | -0.017542 | -0.36636
2019-03-28 15:01 | 1553799637.7496 | -1 | 0.20952 | 0.000444 | -1e-06 | 98.9236 | 98.7794 | 126.2061 | 21.2612 | 0.001795 | 0.001743 | 261.2602 | 0.26192 | 8.3999 | 0.00281 | 0.004533 | 0.004426 | 5.8e-05 | 0.001285 | 0.35977 | 1.2739 | 2.5473 | 100.7 | 21.13 | 25.855 | 65.0068 | 89.4135 | 63.7663 | 0.006444 | 1.0479 | 0.025277 | 24.1224 | 0.013052 | 0.15582 | -0.000151 | -0.01144 | -0.01144 | -0.56251
2019-03-28 15:01 | 1553799638.7487 | 0 | 0.20951 | 0.00044 | -1e-06 | 98.8304 | 98.515 | 124.8391 | 21.2612 | 0.001871 | 0.001772 | 261.554 | 0.26088 | 8.4042 | 0.002846 | 0.004368 | 0.004257 | 6e-05 | 0.001285 | 0.31349 | 1.2661 | 2.8079 | 100.7 | 21.13 | 25.855 | 66.8478 | 90.2964 | 64.4579 | 0.006444 | 1.0457 | 0.025114 | 24.016 | 0.028583 | 0.1552 | -8e-05 | -0.005966 | -0.005966 | -1.0744
2019-03-28 15:01 | 1553799639.7497 | 1 | 0.2095 | 0.000441 | 1e-06 | 98.7339 | 98.4654 | 122.9758 | 21.2609 | 0.001846 | 0.001726 | 261.3882 | 0.2618 | 8.4019 | 0.002705 | 0.004121 | 0.004009 | 5.7e-05 | 0.001288 | 0.35441 | 1.2645 | 2.5813 | 100.7 | 21.13 | 25.858 | 66.2506 | 86.8967 | 63.3684 | 0.006445 | 1.0491 | 0.024966 | 23.7976 | 0.024349 | 0.15574 | -1.5e-05 | -0.001126 | -0.001126 | -5.7135
2019-03-28 15:01 | 1553799640.7497 | 2 | 0.20946 | 0.000439 | 3e-06 | 98.5024 | 96.1268 | 121.2213 | 21.2607 | 0.001853 | 0.001784 | 261.3411 | 0.26193 | 8.4033 | 0.002841 | 0.016659 | 0.01771 | 5.8e-05 | 0.001287 | -1.1594 | 1.276 | 2.5475 | 100.7 | 21.13 | 25.858 | 66.4096 | 90.1615 | 64.7424 | 0.006445 | 1.0448 | 0.024737 | 23.6754 | 0.21837 | 0.15582 | 0.000236 | 0.0174 | 0.0174 | 0.36986
2019-03-28 15:01 | 1553799641.7493 | 3 | 0.20942 | 0.000444 | 8e-06 | 98.4062 | 94.4451 | 119.0074 | 21.26 | 0.00191 | 0.001878 | 261.3012 | 0.26193 | 8.4009 | 0.002737 | 0.004098 | 0.004283 | 5.8e-05 | 0.001284 | 0.31078 | 1.3031 | 2.5473 | 100.7 | 21.13 | 25.87 | 67.7858 | 87.6643 | 67.0142 | 0.006448 | 1.0379 | 0.024428 | 23.537 | 0.36749 | 0.15582 | 0.000444 | 0.032124 | 0.032124 | 0.20033
2019-03-28 15:01 | 1553799642.7498 | 4 | 0.20884 | 0.000446 | 1.5e-05 | 98.0402 | 95.432 | 117.396 | 21.2611 | 0.001888 | 0.001916 | 261.1749 | 0.26188 | 8.402 | 0.002768 | 0.005359 | 0.005493 | 5.8e-05 | 0.001281 | 0.35187 | 1.3057 | 2.5473 | 100.7 | 21.13 | 25.886 | 67.2547 | 88.4141 | 67.9359 | 0.006452 | 1.0351 | 0.02423 | 23.4087 | 0.24117 | 0.15579 | 0.003936 | 0.28619 | 0.28619 | 0.022482
2019-03-28 15:01 | 1553799643.7488 | 5 | 0.20779 | 0.000477 | 2.4e-05 | 98.1182 | 95.3668 | 115.8139 | 21.2611 | 0.001992 | 0.00195 | 261.3179 | 0.26136 | 8.4035 | 0.002783 | 0.003998 | 0.003967 | 5.4e-05 | 0.001275 | 0.28456 | 1.3306 | 1.7996 | 100.7 | 21.13 | 25.896 | 69.7662 | 88.7657 | 68.747 | 0.006454 | 1.0326 | 0.024037 | 23.2781 | 0.2544 | 0.15548 | 0.01022 | 0.73765 | 0.73765 | 0.008705
2019-03-28 15:01 | 1553799644.7492 | 6 | 0.20684 | 0.000581 | 3.6e-05 | 98.2326 | 91.6332 | 113.9113 | 21.2612 | 0.002027 | 0.002049 | 261.0224 | 0.26216 | 8.4021 | 0.00283 | 0.00417 | 0.004145 | 5.3e-05 | 0.001269 | 0.29142 | 1.3328 | 2.5472 | 100.7 | 21.13 | 25.908 | 70.6031 | 89.8956 | 71.1349 | 0.006457 | 1.0254 | 0.023756 | 23.1666 | 0.62204 | 0.15596 | 0.01576 | 1.1236 | 1.1236 | 0.005733
2019-03-28 15:01 | 1553799645.7499 | 7 | 0.20577 | 0.001074 | 4.9e-05 | 98.2746 | 87.6309 | 112.2267 | 21.2612 | 0.002169 | 0.002099 | 260.9265 | 0.26192 | 8.402 | 0.002798 | 0.004046 | 0.003961 | 5.1e-05 | 0.001268 | 0.29679 | 1.3027 | 2.5473 | 100.69 | 21.13 | 25.934 | 74.0115 | 89.1424 | 72.3277 | 0.006464 | 1.0219 | 0.023539 | 23.0344 | 1.0253 | 0.15582 | 0.02155 | 1.5214 | 1.5214 | 0.00423
2019-03-28 15:01 | 1553799646.7489 | 8 | 0.2043 | 0.001764 | 6.4e-05 | 98.2317 | 83.4958 | 110.5625 | 21.2612 | 0.002163 | 0.002141 | 260.8442 | 0.26175 | 8.4035 | 0.002794 | 0.004005 | 0.003892 | 5.3e-05 | 0.001269 | 0.31293 | 1.3395 | 2.5474 | 100.69 | 21.13 | 25.95 | 73.8539 | 89.0383 | 73.345 | 0.006468 | 1.0189 | 0.02333 | 22.8966 | 1.4538 | 0.15571 | 0.029473 | 2.0608 | 2.0608 | 0.003121
2019-03-28 15:01 | 1553799647.7479 | 9 | 0.20328 | 0.002551 | 8e-05 | 98.1983 | 78.3272 | 109.8001 | 21.2612 | 0.002214 | 0.002126 | 260.7546 | 0.26228 | 8.4039 | 0.002835 | 0.004071 | 0.004029 | 5e-05 | 0.001267 | 0.30309 | 1.3656 | 2.5474 | 100.69 | 21.13 | 25.95 | 75.0878 | 90.0322 | 72.97 | 0.006468 | 1.02 | 0.023262 | 22.8052 | 2.0223 | 0.15603 | 0.034578 | 2.4091 | 2.4091 | 0.002675
2019-03-28 15:01 | 1553799648.749 | 10 | 0.20258 | 0.00339 | 9.4e-05 | 98.1474 | 73.7158 | 109.0389 | 21.2612 | 0.002236 | 0.002291 | 260.6236 | 0.26203 | 8.4005 | 0.002811 | 0.003437 | 0.003397 | 5.3e-05 | 0.001256 | 0.30359 | 1.3142 | 2.5473 | 100.69 | 21.12 | 26.003 | 75.6188 | 89.4518 | 76.9454 | 0.006478 | 1.0084 | 0.023049 | 22.8562 | 2.5604 | 0.15588 | 0.037716 | 2.6024 | 2.6024 | 0.002474
2019-03-28 15:01 | 1553799649.749 | 11 | 0.20179 | 0.004198 | 0.000109 | 98.1146 | 70.5214 | 108.8647 | 21.2612 | 0.002261 | 0.002247 | 260.5357 | 0.26129 | 8.4047 | 0.002848 | 0.004262 | 0.00421 | 5.7e-05 | 0.001246 | 0.29717 | 1.3063 | 2.5472 | 100.69 | 21.12 | 26.003 | 76.2081 | 90.3316 | 75.8758 | 0.006478 | 1.0115 | 0.023066 | 22.803 | 2.9537 | 0.15544 | 0.041432 | 2.8595 | 2.8595 | 0.002245
2019-03-28 15:01 | 1553799650.748 | 12 | 0.20135 | 0.0048 | 0.00012 | 97.9958 | 68.7823 | 108.7598 | 21.2612 | 0.002392 | 0.002323 | 260.4222 | 0.26234 | 8.4019 | 0.002952 | 0.004343 | 0.004269 | 5.3e-05 | 0.001236 | 0.30237 | 1.3141 | 2.5474 | 100.69 | 21.12 | 26.058 | 79.3559 | 92.8397 | 77.6967 | 0.006491 | 1.0063 | 0.022995 | 22.8514 | 3.1662 | 0.15607 | 0.043335 | 2.9806 | 2.9806 | 0.002163
2019-03-28 15:01 | 1553799651.749 | 13 | 0.20097 | 0.005283 | 0.00013 | 97.9913 | 66.8717 | 108.8567 | 21.2612 | 0.002427 | 0.002374 | 260.3248 | 0.26199 | 8.4057 | 0.002951 | 0.004278 | 0.004265 | 5.5e-05 | 0.001225 | 0.29878 | 1.3121 | 2.5475 | 100.69 | 21.12 | 26.058 | 80.201 | 92.8282 | 78.9387 | 0.006491 | 1.0027 | 0.022964 | 22.902 | 3.4177 | 0.15585 | 0.044985 | 3.0892 | 3.0892 | 0.002084
2019-03-28 15:01 | 1553799652.7491 | 14 | 0.20075 | 0.005684 | 0.000139 | 97.9362 | 66.481 | 108.8713 | 21.2612 | 0.002462 | 0.002452 | 260.1902 | 0.26149 | 8.4018 | 0.003 | 0.004336 | 0.00417 | 5.3e-05 | 0.001217 | 0.30702 | 1.3036 | 2.5474 | 100.69 | 21.12 | 26.106 | 81.0525 | 94.0085 | 80.817 | 0.006503 | 0.9974 | 0.022905 | 22.9645 | 3.4651 | 0.15556 | 0.045813 | 3.1372 | 3.1372 | 0.002048
2019-03-28 15:01 | 1553799653.7481 | 15 | 0.2006 | 0.005903 | 0.000147 | 97.809 | 67.0147 | 108.6032 | 21.2612 | 0.002519 | 0.002453 | 260.1083 | 0.26136 | 8.4022 | 0.002991 | 0.00427 | 0.004236 | 5.6e-05 | 0.001211 | 0.29623 | 1.3083 | 2.5474 | 100.69 | 21.12 | 26.106 | 82.4282 | 93.77 | 80.8225 | 0.006503 | 0.99738 | 0.022876 | 22.9364 | 3.382 | 0.15548 | 0.046449 | 3.1763 | 3.1763 | 0.002022
2019-03-28 15:01 | 1553799654.7492 | 16 | 0.20056 | 0.006124 | 0.000153 | 97.7107 | 67.0714 | 108.7795 | 21.2612 | 0.002497 | 0.002572 | 259.94 | 0.26157 | 8.3995 | 0.003019 | 0.004631 | 0.004545 | 5.4e-05 | 0.001209 | 0.32421 | 1.3115 | 2.5474 | 100.69 | 21.12 | 26.134 | 81.9 | 94.4485 | 83.6873 | 0.00651 | 0.98937 | 0.022803 | 23.0477 | 3.3654 | 0.15561 | 0.046417 | 3.1636 | 3.1636 | 0.002031
2019-03-28 15:01 | 1553799655.7492 | 17 | 0.20053 | 0.006321 | 0.000157 | 97.6568 | 66.5843 | 109.1654 | 21.2612 | 0.002555 | 0.002532 | 259.8435 | 0.2626 | 8.4025 | 0.003093 | 0.004108 | 0.003983 | 5.7e-05 | 0.001214 | 0.29679 | 1.3301 | 2.5474 | 100.69 | 21.12 | 26.134 | 83.2877 | 96.2329 | 82.7323 | 0.00651 | 0.99203 | 0.022874 | 23.0577 | 3.4257 | 0.15622 | 0.04634 | 3.1679 | 3.1679 | 0.002037
2019-03-28 15:01 | 1553799656.7483 | 18 | 0.20053 | 0.006498 | 0.000161 | 97.5931 | 66.0268 | 109.2976 | 21.2612 | 0.002676 | 0.002583 | 259.7211 | 0.26233 | 8.4019 | 0.003063 | 0.004779 | 0.004709 | 5.6e-05 | 0.001212 | 0.32084 | 1.3297 | 2.5473 | 100.69 | 21.12 | 26.164 | 86.1851 | 95.5155 | 83.9595 | 0.006518 | 0.98862 | 0.022848 | 23.1114 | 3.495 | 0.15606 | 0.046101 | 3.1478 | 3.1478 | 0.002048
2019-03-28 15:01 | 1553799657.7493 | 19 | 0.20055 | 0.006642 | 0.000162 | 97.4915 | 65.7299 | 109.1379 | 21.2612 | 0.002696 | 0.002557 | 259.5815 | 0.26189 | 8.4055 | 0.003111 | 0.005433 | 0.005258 | 5.2e-05 | 0.001205 | 0.31454 | 1.3623 | 2.5474 | 100.69 | 21.12 | 26.177 | 86.6848 | 96.6762 | 83.3206 | 0.006521 | 0.99039 | 0.022852 | 23.0738 | 3.526 | 0.1558 | 0.045846 | 3.1309 | 3.1309 | 0.002055
2019-03-28 15:01 | 1553799658.7487 | 20 | 0.20056 | 0.006763 | 0.000164 | 97.4048 | 66.306 | 109.6079 | 21.2612 | 0.002701 | 0.002633 | 259.4439 | 0.26188 | 8.405 | 0.003154 | 0.005476 | 0.005384 | 5.5e-05 | 0.0012 | 0.3079 | 1.3611 | 2.5475 | 100.69 | 21.12 | 26.18 | 86.7878 | 97.7067 | 85.162 | 0.006522 | 0.9853 | 0.022842 | 23.1831 | 3.44 | 0.15579 | 0.04561 | 3.1134 | 3.1134 | 0.002067
2019-03-28 15:01 | 1553799659.7494 | 21 | 0.20057 | 0.006865 | 0.000164 | 97.364 | 67.279 | 109.8104 | 21.2612 | 0.002763 | 0.002687 | 259.2742 | 0.26141 | 8.4016 | 0.003233 | 0.0042 | 0.004198 | 5.7e-05 | 0.001204 | 0.31166 | 1.361 | 2.5474 | 100.69 | 21.12 | 26.183 | 88.2783 | 99.6355 | 86.4646 | 0.006523 | 0.98173 | 0.022822 | 23.2466 | 3.306 | 0.15551 | 0.045437 | 3.0989 | 3.0989 | 0.002073
2019-03-28 15:01 | 1553799660.7483 | 22 | 0.20061 | 0.006945 | 0.000163 | 97.2288 | 67.0626 | 109.6478 | 21.2612 | 0.002814 | 0.002757 | 259.1278 | 0.26163 | 8.4041 | 0.0032 | 0.004278 | 0.004186 | 5.9e-05 | 0.001205 | 0.32117 | 1.3114 | 2.5473 | 100.69 | 21.12 | 26.184 | 89.526 | 98.838 | 88.1363 | 0.006523 | 0.97719 | 0.022752 | 23.2833 | 3.3224 | 0.15564 | 0.045092 | 3.066 | 3.066 | 0.002097
2019-03-28 15:01 | 1553799661.7484 | 23 | 0.20072 | 0.007011 | 0.00016 | 97.2827 | 67.6422 | 110.0491 | 21.2612 | 0.002864 | 0.002748 | 258.9738 | 0.2606 | 8.4027 | 0.003145 | 0.004407 | 0.004269 | 5.5e-05 | 0.001203 | 0.3115 | 1.2889 | 2.5472 | 100.69 | 21.12 | 26.183 | 90.7091 | 97.5028 | 87.9191 | 0.006523 | 0.97778 | 0.022801 | 23.3189 | 3.2504 | 0.15503 | 0.044373 | 3.0238 | 3.0238 | 0.002117
2019-03-28 15:01 | 1553799662.7485 | 24 | 0.20084 | 0.007062 | 0.000157 | 97.2413 | 68.9355 | 110.193 | 21.2612 | 0.002884 | 0.002746 | 258.8667 | 0.26266 | 8.4034 | 0.003199 | 0.004533 | 0.004437 | 6e-05 | 0.001201 | 0.31332 | 1.2733 | 2.5472 | 100.69 | 21.12 | 26.178 | 91.2047 | 98.7941 | 87.8895 | 0.006521 | 0.97786 | 0.022817 | 23.3332 | 3.0771 | 0.15625 | 0.043569 | 2.9713 | 2.9713 | 0.002172
2019-03-28 15:01 | 1553799663.7485 | 25 | 0.20097 | 0.007095 | 0.000154 | 97.255 | 69.2406 | 110.5355 | 21.2612 | 0.002867 | 0.002735 | 258.6733 | 0.26154 | 8.4029 | 0.003188 | 0.004609 | 0.00451 | 6e-05 | 0.001207 | 0.30917 | 1.2812 | 2.5472 | 100.69 | 21.12 | 26.174 | 90.7857 | 98.5478 | 87.6067 | 0.00652 | 0.97862 | 0.022861 | 23.3603 | 3.0389 | 0.15559 | 0.042781 | 2.9235 | 2.9235 | 0.002198
2019-03-28 15:01 | 1553799664.7475 | 26 | 0.20106 | 0.007113 | 0.00015 | 97.2124 | 69.6697 | 110.6796 | 21.2612 | 0.002919 | 0.002836 | 258.5781 | 0.26128 | 8.4007 | 0.003147 | 0.003826 | 0.003728 | 5.6e-05 | 0.001205 | 0.29717 | 1.2968 | 2.5474 | 100.69 | 21.13 | 26.17 | 92.0341 | 97.5507 | 90.0454 | 0.006523 | 0.97205 | 0.022799 | 23.4544 | 2.9797 | 0.15543 | 0.042236 | 2.8786 | 2.8786 | 0.00223
2019-03-28 15:01 | 1553799665.7486 | 27 | 0.20114 | 0.007114 | 0.000147 | 97.0894 | 70.1864 | 110.5317 | 21.2612 | 0.002926 | 0.002841 | 258.4171 | 0.26129 | 8.4024 | 0.003131 | 0.004895 | 0.00477 | 5.4e-05 | 0.001203 | 0.33019 | 1.2838 | 2.5472 | 100.69 | 21.13 | 26.161 | 92.2132 | 97.1666 | 90.1697 | 0.006521 | 0.97172 | 0.02278 | 23.4427 | 2.9023 | 0.15544 | 0.041732 | 2.8422 | 2.8422 | 0.002259
2019-03-28 15:01 | 1553799666.7486 | 28 | 0.20125 | 0.0071 | 0.000145 | 97.0757 | 70.5597 | 110.3041 | 21.2612 | 0.002966 | 0.002804 | 258.2905 | 0.26092 | 8.4045 | 0.003225 | 0.005168 | 0.005093 | 5.4e-05 | 0.001198 | 0.29374 | 1.3095 | 2.5475 | 100.69 | 21.13 | 26.161 | 93.1739 | 99.4254 | 89.2719 | 0.006521 | 0.97413 | 0.022784 | 23.3896 | 2.8536 | 0.15522 | 0.041098 | 2.7997 | 2.7997 | 0.00229
2019-03-28 15:01 | 1553799667.7486 | 29 | 0.20132 | 0.007087 | 0.000142 | 97.1344 | 70.5767 | 110.635 | 21.2612 | 0.00296 | 0.002898 | 258.1751 | 0.26121 | 8.4035 | 0.003213 | 0.004654 | 0.004604 | 5.3e-05 | 0.001185 | 0.32565 | 1.3132 | 2.5473 | 100.69 | 21.13 | 26.158 | 93.0439 | 99.1477 | 91.5377 | 0.00652 | 0.96808 | 0.022748 | 23.4978 | 2.8568 | 0.1554 | 0.040723 | 2.7699 | 2.7699 | 0.002317
2019-03-28 15:01 | 1553799668.7476 | 30 | 0.20137 | 0.007063 | 0.000141 | 97.0617 | 70.4823 | 110.5841 | 21.2612 | 0.003007 | 0.00291 | 258.0682 | 0.26228 | 8.4017 | 0.003184 | 0.00532 | 0.005205 | 5.5e-05 | 0.001177 | 0.31868 | 1.276 | 2.5473 | 100.69 | 21.13 | 26.151 | 94.1793 | 98.4475 | 91.8396 | 0.006519 | 0.96727 | 0.022733 | 23.5021 | 2.8621 | 0.15603 | 0.040418 | 2.7475 | 2.7475 | 0.002346
2019-03-28 15:01 | 1553799669.7477 | 31 | 0.20135 | 0.007025 | 0.00014 | 97.1186 | 70.1932 | 110.3862 | 21.2612 | 0.003075 | 0.002854 | 257.9579 | 0.26078 | 8.3996 | 0.00319 | 0.00535 | 0.005332 | 5.3e-05 | 0.001178 | 0.36226 | 1.2882 | 2.5473 | 100.69 | 21.13 | 26.149 | 95.8013 | 98.5785 | 90.4855 | 0.006518 | 0.97088 | 0.022755 | 23.4374 | 2.9041 | 0.15514 | 0.040618 | 2.7638 | 2.7638 | 0.002318
2019-03-28 15:01 | 1553799670.749 | 32 | 0.20131 | 0.006992 | 0.000139 | 97.0861 | 70.2302 | 110.4904 | 21.2612 | 0.003117 | 0.002882 | 257.8465 | 0.26093 | 8.4027 | 0.003261 | 0.004846 | 0.004704 | 5.2e-05 | 0.001176 | 0.33157 | 1.2891 | 2.5472 | 100.69 | 21.13 | 26.149 | 96.8283 | 100.3032 | 91.1545 | 0.006518 | 0.96909 | 0.022745 | 23.4701 | 2.8964 | 0.15523 | 0.040884 | 2.7807 | 2.7807 | 0.002306
2019-03-28 15:01 | 1553799671.7491 | 33 | 0.20129 | 0.006964 | 0.000139 | 97.1022 | 69.7532 | 110.4308 | 21.2612 | 0.003041 | 0.002964 | 257.7334 | 0.26173 | 8.4061 | 0.003282 | 0.00475 | 0.004698 | 5.7e-05 | 0.001183 | 0.30862 | 1.3128 | 2.5474 | 100.69 | 21.13 | 26.143 | 94.9824 | 100.8039 | 93.137 | 0.006517 | 0.96385 | 0.022677 | 23.5275 | 2.9589 | 0.15571 | 0.041048 | 2.7835 | 2.7835 | 0.00231
2019-03-28 15:01 | 1553799672.7477 | 34 | 0.2013 | 0.006947 | 0.000138 | 97.0418 | 69.5487 | 110.239 | 21.2612 | 0.00312 | 0.002956 | 257.5983 | 0.26168 | 8.4035 | 0.003281 | 0.005103 | 0.005052 | 5.8e-05 | 0.00118 | 0.31653 | 1.2917 | 2.5473 | 100.69 | 21.13 | 26.143 | 96.9012 | 100.7906 | 92.9289 | 0.006517 | 0.9644 | 0.022664 | 23.5004 | 2.9796 | 0.15567 | 0.041009 | 2.7793 | 2.7793 | 0.002313
2019-03-28 15:01 | 1553799673.748 | 35 | 0.20132 | 0.00694 | 0.000138 | 96.9613 | 70.4989 | 110.1961 | 21.2612 | 0.003199 | 0.00294 | 257.5093 | 0.26195 | 8.4027 | 0.003312 | 0.00466 | 0.004611 | 5.7e-05 | 0.001188 | 0.33411 | 1.3005 | 2.5474 | 100.69 | 21.13 | 26.134 | 98.7965 | 101.5262 | 92.5589 | 0.006514 | 0.96537 | 0.022671 | 23.4839 | 2.8508 | 0.15583 | 0.040909 | 2.7734 | 2.7734 | 0.002321
2019-03-28 15:01 | 1553799674.7478 | 36 | 0.20139 | 0.006931 | 0.000137 | 96.9962 | 70.9518 | 110.4578 | 21.2612 | 0.003159 | 0.003034 | 257.3709 | 0.26112 | 8.4022 | 0.003234 | 0.004609 | 0.004486 | 5.2e-05 | 0.001182 | 0.32836 | 1.3113 | 2.5474 | 100.69 | 21.13 | 26.131 | 97.8352 | 99.6442 | 94.8179 | 0.006514 | 0.95945 | 0.022628 | 23.5843 | 2.7967 | 0.15534 | 0.040504 | 2.7409 | 2.7409 | 0.002341
2019-03-28 15:01 | 1553799675.7482 | 37 | 0.20139 | 0.006917 | 0.000135 | 97.0689 | 71.9771 | 110.7014 | 21.2612 | 0.003136 | 0.002998 | 257.2703 | 0.26101 | 8.4016 | 0.003238 | 0.004498 | 0.004418 | 5.5e-05 | 0.001186 | 0.33179 | 1.3246 | 2.5474 | 100.69 | 21.13 | 26.126 | 97.2828 | 99.7376 | 93.9555 | 0.006512 | 0.9617 | 0.022679 | 23.5826 | 2.6751 | 0.15527 | 0.040527 | 2.7487 | 2.7487 | 0.002333
2019-03-28 15:01 | 1553799676.748 | 38 | 0.20139 | 0.006907 | 0.000134 | 96.959 | 72.4248 | 110.9772 | 21.2612 | 0.003196 | 0.003003 | 257.1061 | 0.26177 | 8.4019 | 0.003218 | 0.004395 | 0.00437 | 5.4e-05 | 0.001188 | 0.32261 | 1.3547 | 2.5474 | 100.69 | 21.13 | 26.124 | 98.7385 | 99.2574 | 94.0643 | 0.006512 | 0.96142 | 0.022704 | 23.6155 | 2.6095 | 0.15573 | 0.040505 | 2.7504 | 2.7504 | 0.002338
2019-03-28 15:01 | 1553799677.7479 | 39 | 0.20137 | 0.006893 | 0.000133 | 96.8961 | 71.5568 | 111.0345 | 21.2612 | 0.003219 | 0.00302 | 256.9905 | 0.26137 | 8.4021 | 0.003223 | 0.004434 | 0.004322 | 5.5e-05 | 0.001192 | 0.3178 | 1.3084 | 2.5473 | 100.69 | 21.13 | 26.122 | 99.2905 | 99.3884 | 94.4695 | 0.006511 | 0.96035 | 0.022698 | 23.6346 | 2.7115 | 0.15549 | 0.040636 | 2.7585 | 2.7585 | 0.002328
2019-03-28 15:01 | 1553799678.7475 | 40 | 0.20142 | 0.00688 | 0.000132 | 96.9089 | 71.1916 | 111.0268 | 21.2612 | 0.003254 | 0.003013 | 256.8805 | 0.26201 | 8.402 | 0.003243 | 0.004067 | 0.00399 | 5.6e-05 | 0.001198 | 0.33157 | 1.2903 | 2.5474 | 100.69 | 21.13 | 26.116 | 100.1315 | 99.8648 | 94.3014 | 0.00651 | 0.96079 | 0.022702 | 23.6283 | 2.7585 | 0.15587 | 0.040408 | 2.7436 | 2.7436 | 0.002346
2019-03-28 15:01 | 1553799679.747 | 41 | 0.20153 | 0.006859 | 0.000129 | 96.9459 | 71.5898 | 111.3354 | 21.2612 | 0.003215 | 0.002985 | 256.7731 | 0.26155 | 8.4016 | 0.003197 | 0.004906 | 0.004721 | 5.5e-05 | 0.001199 | 0.34147 | 1.2774 | 2.5473 | 100.69 | 21.13 | 26.112 | 99.1878 | 98.7643 | 93.6432 | 0.006509 | 0.96252 | 0.022754 | 23.64 | 2.712 | 0.1556 | 0.039764 | 2.7063 | 2.7063 | 0.002375
2019-03-28 15:01 | 1553799680.7471 | 42 | 0.20154 | 0.006847 | 0.000126 | 96.95 | 72.3769 | 111.2922 | 21.2612 | 0.003206 | 0.003073 | 256.6576 | 0.26055 | 8.4028 | 0.003271 | 0.004954 | 0.00484 | 5.5e-05 | 0.001208 | 0.32516 | 1.3448 | 2.5474 | 100.69 | 21.13 | 26.108 | 98.9681 | 100.5451 | 95.7615 | 0.006508 | 0.95699 | 0.022684 | 23.7035 | 2.6146 | 0.155 | 0.039723 | 2.6954 | 2.6954 | 0.002375
2019-03-28 15:01 | 1553799681.748 | 43 | 0.20158 | 0.00683 | 0.000122 | 96.942 | 73.374 | 111.4173 | 21.2612 | 0.003254 | 0.003014 | 256.5491 | 0.2608 | 8.4045 | 0.003252 | 0.004589 | 0.004482 | 5.5e-05 | 0.001215 | 0.31188 | 1.2997 | 2.5473 | 100.69 | 21.13 | 26.108 | 100.1298 | 100.0836 | 94.3335 | 0.006508 | 0.96071 | 0.022741 | 23.6709 | 2.4914 | 0.15515 | 0.039493 | 2.6867 | 2.6867 | 0.002385
2019-03-28 15:01 | 1553799682.748 | 44 | 0.20175 | 0.006808 | 0.000119 | 96.9009 | 73.1923 | 111.6342 | 21.2612 | 0.003248 | 0.003042 | 256.4477 | 0.2613 | 8.4044 | 0.00324 | 0.004028 | 0.004033 | 5.4e-05 | 0.001212 | 0.29839 | 1.3026 | 2.5474 | 100.69 | 21.13 | 26.091 | 99.9997 | 99.798 | 95.001 | 0.006504 | 0.95897 | 0.022742 | 23.7154 | 2.5098 | 0.15545 | 0.038488 | 2.6188 | 2.6188 | 0.002452
2019-03-28 15:01 | 1553799683.747 | 45 | 0.2019 | 0.006774 | 0.000115 | 96.8642 | 74.5617 | 111.403 | 21.2612 | 0.003288 | 0.003106 | 256.337 | 0.26188 | 8.404 | 0.003223 | 0.005 | 0.004891 | 5.6e-05 | 0.001215 | 0.30845 | 1.3005 | 2.5474 | 100.69 | 21.13 | 26.091 | 100.9532 | 99.3853 | 96.5487 | 0.006504 | 0.95495 | 0.022671 | 23.7406 | 2.3406 | 0.15579 | 0.037654 | 2.5543 | 2.5543 | 0.002519
2019-03-28 15:01 | 1553799684.7472 | 46 | 0.20201 | 0.00673 | 0.000112 | 96.808 | 75.178 | 111.0549 | 21.2612 | 0.003287 | 0.003083 | 256.2211 | 0.26163 | 8.4019 | 0.003157 | 0.004541 | 0.004469 | 5.5e-05 | 0.001216 | 0.31277 | 1.3267 | 2.5474 | 100.69 | 21.13 | 26.07 | 100.943 | 97.7951 | 95.99 | 0.006498 | 0.9564 | 0.022653 | 23.6856 | 2.2618 | 0.15564 | 0.037053 | 2.5118 | 2.5118 | 0.002559
2019-03-28 15:01 | 1553799685.7476 | 47 | 0.20202 | 0.006671 | 0.00011 | 96.8312 | 76.1689 | 111.0059 | 21.2612 | 0.003307 | 0.003082 | 256.0954 | 0.26217 | 8.4017 | 0.003214 | 0.004748 | 0.004581 | 5.5e-05 | 0.001216 | 0.3131 | 1.3552 | 2.5475 | 100.69 | 21.13 | 26.07 | 101.4107 | 99.1673 | 95.9736 | 0.006498 | 0.95644 | 0.022648 | 23.6798 | 2.1468 | 0.15596 | 0.037094 | 2.5142 | 2.5142 | 0.002562
2019-03-28 15:01 | 1553799686.7472 | 48 | 0.20203 | 0.00662 | 0.000108 | 96.8386 | 76.0337 | 110.8139 | 21.2612 | 0.003337 | 0.003103 | 256.0355 | 0.26108 | 8.404 | 0.003225 | 0.004222 | 0.004328 | 5.4e-05 | 0.001217 | 0.31205 | 1.3502 | 2.5474 | 100.69 | 21.13 | 26.055 | 102.1419 | 99.4199 | 96.4758 | 0.006495 | 0.95514 | 0.022613 | 23.6754 | 2.1634 | 0.15531 | 0.037061 | 2.5082 | 2.5082 | 0.002557
2019-03-28 15:01 | 1553799687.7482 | 49 | 0.20206 | 0.006563 | 0.000106 | 96.819 | 75.4056 | 110.4867 | 21.2612 | 0.003287 | 0.003133 | 255.9095 | 0.26202 | 8.404 | 0.003191 | 0.005113 | 0.004993 | 5.5e-05 | 0.001222 | 0.33964 | 1.3649 | 2.5473 | 100.69 | 21.13 | 26.055 | 100.9232 | 98.6191 | 97.1988 | 0.006495 | 0.95328 | 0.022558 | 23.6636 | 2.2358 | 0.15588 | 0.03699 | 2.4973 | 2.4973 | 0.002578
2019-03-28 15:01 | 1553799688.7473 | 50 | 0.20208 | 0.006526 | 0.000106 | 96.8383 | 75.9485 | 110.2017 | 21.2612 | 0.003371 | 0.003112 | 255.7886 | 0.26139 | 8.4029 | 0.003226 | 0.003776 | 0.003712 | 5.8e-05 | 0.001221 | 0.30845 | 1.3573 | 2.5473 | 100.69 | 21.13 | 26.053 | 102.969 | 99.4677 | 96.7061 | 0.006494 | 0.95455 | 0.022544 | 23.6173 | 2.1734 | 0.1555 | 0.036918 | 2.4909 | 2.4909 | 0.002578
2019-03-28 15:01 | 1553799689.7473 | 51 | 0.20209 | 0.006493 | 0.000105 | 96.8637 | 76.2315 | 110.0897 | 21.2612 | 0.003351 | 0.003162 | 255.6653 | 0.2618 | 8.402 | 0.003192 | 0.004728 | 0.004679 | 5.6e-05 | 0.001224 | 0.30453 | 1.3142 | 2.5474 | 100.69 | 21.13 | 26.056 | 102.482 | 98.6333 | 97.9089 | 0.006495 | 0.95145 | 0.022496 | 23.6436 | 2.1425 | 0.15575 | 0.03689 | 2.4838 | 2.4838 | 0.00259
2019-03-28 15:01 | 1553799690.7474 | 52 | 0.20208 | 0.006471 | 0.000105 | 96.9155 | 76.6979 | 110.0954 | 21.2612 | 0.003369 | 0.003181 | 255.5621 | 0.26104 | 8.4022 | 0.003291 | 0.004475 | 0.004373 | 5.1e-05 | 0.001219 | 0.32676 | 1.3133 | 2.5473 | 100.69 | 21.13 | 26.062 | 102.9164 | 101.0323 | 98.3644 | 0.006496 | 0.95029 | 0.022483 | 23.6588 | 2.0927 | 0.15529 | 0.036952 | 2.4865 | 2.4865 | 0.002579
2019-03-28 15:02 | 1553799691.7474 | 53 | 0.202 | 0.006443 | 0.000105 | 96.917 | 76.9432 | 109.3929 | 21.2612 | 0.003396 | 0.003161 | 255.4207 | 0.26121 | 8.4032 | 0.003163 | 0.003774 | 0.003752 | 5.4e-05 | 0.001218 | 0.32084 | 1.3092 | 2.5473 | 100.69 | 21.13 | 26.087 | 103.5694 | 97.9372 | 97.8947 | 0.006503 | 0.95149 | 0.022425 | 23.5682 | 2.0643 | 0.15539 | 0.037506 | 2.5173 | 2.5173 | 0.002549
2019-03-28 15:02 | 1553799692.7478 | 54 | 0.20192 | 0.006425 | 0.000104 | 96.9417 | 75.941 | 109.3252 | 21.2612 | 0.003439 | 0.003179 | 255.3041 | 0.26087 | 8.4037 | 0.003225 | 0.005249 | 0.005153 | 5.6e-05 | 0.001221 | 0.30215 | 1.2824 | 2.5472 | 100.69 | 21.13 | 26.104 | 104.6199 | 99.4356 | 98.3102 | 0.006507 | 0.95043 | 0.022405 | 23.5741 | 2.1838 | 0.15519 | 0.037963 | 2.5457 | 2.5457 | 0.002518
2019-03-28 15:02 | 1553799693.7466 | 55 | 0.20175 | 0.006416 | 0.000104 | 96.9691 | 76.9038 | 109.2515 | 21.2612 | 0.003384 | 0.003221 | 255.1811 | 0.26105 | 8.4038 | 0.003208 | 0.004915 | 0.00491 | 5.2e-05 | 0.001224 | 0.30757 | 1.3387 | 2.5474 | 100.69 | 21.14 | 26.144 | 103.2693 | 99.0113 | 99.3374 | 0.006521 | 0.94781 | 0.022367 | 23.5988 | 2.0737 | 0.1553 | 0.038994 | 2.6102 | 2.6102 | 0.002457
2019-03-28 15:02 | 1553799694.7464 | 56 | 0.20168 | 0.00642 | 0.000104 | 96.945 | 77.0155 | 108.8741 | 21.2612 | 0.003423 | 0.003232 | 255.0665 | 0.26137 | 8.4014 | 0.003252 | 0.005222 | 0.00513 | 5.2e-05 | 0.001219 | 0.33323 | 1.3505 | 2.5475 | 100.69 | 21.14 | 26.166 | 104.2202 | 100.0867 | 99.5932 | 0.006526 | 0.94715 | 0.022321 | 23.566 | 2.0585 | 0.15549 | 0.03945 | 2.6352 | 2.6352 | 0.002437
2019-03-28 15:02 | 1553799695.7477 | 57 | 0.20162 | 0.006449 | 0.000104 | 97.0218 | 77.095 | 108.4457 | 21.2612 | 0.003431 | 0.003202 | 254.9447 | 0.26037 | 8.4033 | 0.003298 | 0.004559 | 0.004519 | 5.4e-05 | 0.001218 | 0.32244 | 1.3508 | 2.5474 | 100.69 | 21.14 | 26.189 | 104.424 | 101.2026 | 98.8835 | 0.006532 | 0.94896 | 0.022298 | 23.4972 | 2.0563 | 0.15489 | 0.039761 | 2.6532 | 2.6532 | 0.002411
2019-03-28 15:02 | 1553799696.7475 | 58 | 0.20133 | 0.00648 | 0.000104 | 97.0194 | 78.2216 | 108.5922 | 21.2612 | 0.003502 | 0.003215 | 254.7743 | 0.26059 | 8.4022 | 0.003255 | 0.004202 | 0.004124 | 5.4e-05 | 0.001212 | 0.33483 | 1.3242 | 2.5473 | 100.69 | 21.14 | 26.189 | 106.1309 | 100.1668 | 99.1955 | 0.006532 | 0.94817 | 0.022304 | 23.523 | 1.9263 | 0.15502 | 0.041428 | 2.7649 | 2.7649 | 0.002316
2019-03-28 15:02 | 1553799697.7476 | 59 | 0.2012 | 0.00652 | 0.000105 | 97.0465 | 77.764 | 108.6469 | 21.2612 | 0.003478 | 0.003267 | 254.6501 | 0.26084 | 8.4029 | 0.00328 | 0.004464 | 0.004368 | 5.2e-05 | 0.001211 | 0.33035 | 1.2988 | 2.5475 | 100.69 | 21.14 | 26.256 | 105.5683 | 100.7679 | 100.4382 | 0.006549 | 0.94501 | 0.022272 | 23.5681 | 1.9813 | 0.15517 | 0.042143 | 2.8085 | 2.8085 | 0.002282
2019-03-28 15:02 | 1553799698.7466 | 60 | 0.20106 | 0.006576 | 0.000105 | 97.0784 | 77.4881 | 108.5748 | 21.2612 | 0.003537 | 0.003282 | 254.5057 | 0.26117 | 8.4011 | 0.00324 | 0.004415 | 0.004312 | 5.5e-05 | 0.001207 | 0.32012 | 1.3127 | 2.5474 | 100.69 | 21.14 | 26.256 | 106.9986 | 99.8058 | 100.8189 | 0.006549 | 0.94405 | 0.022253 | 23.5723 | 2.0161 | 0.15537 | 0.042953 | 2.8599 | 2.8599 | 0.002244
2019-03-28 15:02 | 1553799699.7476 | 61 | 0.20097 | 0.006643 | 0.000104 | 97.0254 | 77.3478 | 108.6086 | 21.2612 | 0.003563 | 0.003259 | 254.3894 | 0.26129 | 8.4019 | 0.003271 | 0.004904 | 0.004827 | 5.7e-05 | 0.001213 | 0.31116 | 1.3327 | 2.5474 | 100.69 | 21.14 | 26.312 | 107.6277 | 100.5528 | 100.2539 | 0.006563 | 0.94548 | 0.022274 | 23.5581 | 2.0274 | 0.15544 | 0.043396 | 2.8919 | 2.8919 | 0.00222
2019-03-28 15:02 | 1553799700.748 | 62 | 0.20076 | 0.006707 | 0.000104 | 97.0733 | 77.2972 | 108.5131 | 21.2612 | 0.003565 | 0.003339 | 254.2589 | 0.26117 | 8.4065 | 0.003287 | 0.004537 | 0.004571 | 5.5e-05 | 0.001217 | 0.32372 | 1.342 | 2.5475 | 100.69 | 21.14 | 26.327 | 107.6656 | 100.943 | 102.1961 | 0.006566 | 0.94059 | 0.022206 | 23.6089 | 2.0377 | 0.15537 | 0.044551 | 2.9597 | 2.9597 | 0.002168
2019-03-28 15:02 | 1553799701.7477 | 63 | 0.2006 | 0.006788 | 0.000104 | 97.0516 | 77.2832 | 108.6582 | 21.2612 | 0.003573 | 0.003282 | 254.1573 | 0.26105 | 8.4044 | 0.003295 | 0.004736 | 0.004669 | 5.7e-05 | 0.001224 | 0.31061 | 1.3138 | 2.5475 | 100.69 | 21.14 | 26.353 | 107.8576 | 101.1305 | 100.8102 | 0.006573 | 0.94407 | 0.022262 | 23.5811 | 2.0373 | 0.1553 | 0.045402 | 3.0237 | 3.0237 | 0.002121
2019-03-28 15:02 | 1553799702.748 | 64 | 0.20055 | 0.006882 | 0.000104 | 96.9938 | 77.8411 | 107.8444 | 21.2612 | 0.003629 | 0.003351 | 254.0191 | 0.2611 | 8.404 | 0.00329 | 0.00508 | 0.004962 | 5.7e-05 | 0.00123 | 0.29479 | 1.3134 | 2.5474 | 100.69 | 21.14 | 26.353 | 109.2124 | 101.0137 | 102.4811 | 0.006573 | 0.93987 | 0.022129 | 23.545 | 1.9676 | 0.15533 | 0.045638 | 3.0212 | 3.0212 | 0.002123
2019-03-28 15:02 | 1553799703.7477 | 65 | 0.20055 | 0.00698 | 0.000104 | 97.0879 | 78.6523 | 107.8052 | 21.2612 | 0.00361 | 0.003332 | 253.8925 | 0.26011 | 8.4009 | 0.003324 | 0.004429 | 0.004386 | 6e-05 | 0.001234 | 0.29407 | 1.2791 | 2.5473 | 100.69 | 21.14 | 26.365 | 108.7679 | 101.8277 | 102.032 | 0.006576 | 0.941 | 0.022139 | 23.5267 | 1.8835 | 0.15474 | 0.045513 | 3.0142 | 3.0142 | 0.00212
2019-03-28 15:02 | 1553799704.7468 | 66 | 0.20053 | 0.007075 | 0.000104 | 97.1395 | 78.2739 | 108.0425 | 21.2612 | 0.003636 | 0.003361 | 253.7835 | 0.26083 | 8.4031 | 0.003347 | 0.004889 | 0.004812 | 5.6e-05 | 0.001229 | 0.31957 | 1.2691 | 2.5473 | 100.69 | 21.14 | 26.365 | 109.3881 | 102.3908 | 102.7317 | 0.006576 | 0.93925 | 0.022142 | 23.5745 | 1.9314 | 0.15517 | 0.045496 | 3.0136 | 3.0136 | 0.002127
2019-03-28 15:02 | 1553799705.7472 | 67 | 0.20049 | 0.007154 | 0.000104 | 97.1788 | 78.0945 | 108.016 | 21.2612 | 0.003671 | 0.003416 | 253.6448 | 0.26072 | 8.4029 | 0.003249 | 0.005221 | 0.005194 | 5.5e-05 | 0.001231 | 0.29679 | 1.2726 | 2.5472 | 100.69 | 21.14 | 26.374 | 110.2405 | 100.0011 | 104.0569 | 0.006578 | 0.93595 | 0.022101 | 23.6131 | 1.9556 | 0.1551 | 0.04564 | 3.0174 | 3.0174 | 0.002123
2019-03-28 15:02 | 1553799706.7478 | 68 | 0.20051 | 0.007225 | 0.000105 | 97.1448 | 76.9167 | 107.3951 | 21.2612 | 0.003676 | 0.003425 | 253.5419 | 0.26033 | 8.4018 | 0.003329 | 0.004401 | 0.004354 | 5.9e-05 | 0.001228 | 0.33339 | 1.2651 | 2.5473 | 100.69 | 21.14 | 26.374 | 110.3606 | 101.9369 | 104.2706 | 0.006578 | 0.93542 | 0.022031 | 23.5518 | 2.0884 | 0.15487 | 0.04542 | 2.9934 | 2.9934 | 0.002137
2019-03-28 15:02 | 1553799707.7479 | 69 | 0.20054 | 0.007281 | 0.000105 | 97.2357 | 76.8688 | 107.2677 | 21.2612 | 0.003672 | 0.003472 | 253.4027 | 0.26145 | 8.4021 | 0.003283 | 0.004417 | 0.004273 | 5.7e-05 | 0.001223 | 0.28622 | 1.2745 | 2.5473 | 100.69 | 21.14 | 26.376 | 110.2746 | 100.8339 | 105.4007 | 0.006579 | 0.93262 | 0.021985 | 23.5731 | 2.1023 | 0.15553 | 0.045215 | 2.9736 | 2.9736 | 0.00216
2019-03-28 15:02 | 1553799708.747 | 70 | 0.20048 | 0.007334 | 0.000104 | 97.2613 | 77.7918 | 106.9146 | 21.2612 | 0.003801 | 0.00342 | 253.2897 | 0.26154 | 8.4033 | 0.003352 | 0.005426 | 0.005318 | 5.6e-05 | 0.001224 | 0.32084 | 1.2769 | 2.5473 | 100.69 | 21.14 | 26.376 | 113.4111 | 102.5111 | 104.1424 | 0.006579 | 0.93573 | 0.021985 | 23.4951 | 1.9979 | 0.15559 | 0.045465 | 2.9902 | 2.9902 | 0.002149
2019-03-28 15:02 | 1553799709.7459 | 71 | 0.20023 | 0.007386 | 0.000104 | 97.2452 | 78.2759 | 106.6264 | 21.2612 | 0.00377 | 0.003471 | 253.1579 | 0.26191 | 8.4031 | 0.003329 | 0.00532 | 0.005315 | 5.8e-05 | 0.001228 | 0.32372 | 1.3318 | 2.5474 | 100.69 | 21.14 | 26.376 | 112.6473 | 101.9582 | 105.3772 | 0.006579 | 0.93268 | 0.02192 | 23.5018 | 1.9409 | 0.15581 | 0.046942 | 3.0778 | 3.0778 | 0.002091
2019-03-28 15:02 | 1553799710.7463 | 72 | 0.2001 | 0.007442 | 0.000104 | 97.3456 | 77.9804 | 106.31 | 21.2612 | 0.003777 | 0.003508 | 253.0225 | 0.26123 | 8.4033 | 0.003358 | 0.004796 | 0.004661 | 5.7e-05 | 0.001227 | 0.30845 | 1.3196 | 2.5474 | 100.69 | 21.14 | 26.375 | 112.8106 | 102.6392 | 106.2775 | 0.006578 | 0.93047 | 0.021861 | 23.4948 | 1.984 | 0.15541 | 0.047598 | 3.1124 | 3.1124 | 0.002062
2019-03-28 15:02 | 1553799711.746 | 73 | 0.2001 | 0.007505 | 0.000103 | 97.3128 | 78.9313 | 106.4546 | 21.2612 | 0.003819 | 0.003521 | 252.8753 | 0.26108 | 8.4021 | 0.003355 | 0.004494 | 0.00443 | 5.9e-05 | 0.001222 | 0.30613 | 1.291 | 2.5473 | 100.69 | 21.14 | 26.375 | 113.8491 | 102.5888 | 106.597 | 0.006578 | 0.92969 | 0.021867 | 23.5206 | 1.8726 | 0.15531 | 0.047545 | 3.1097 | 3.1097 | 0.002063
2019-03-28 15:02 | 1553799712.7463 | 74 | 0.20007 | 0.007567 | 0.000103 | 97.3387 | 80.0999 | 106.6616 | 21.2612 | 0.003823 | 0.003499 | 252.7306 | 0.2615 | 8.4037 | 0.003308 | 0.004503 | 0.004512 | 5.7e-05 | 0.001226 | 0.34202 | 1.3177 | 2.5474 | 100.69 | 21.14 | 26.37 | 113.9502 | 101.439 | 106.0562 | 0.006577 | 0.93101 | 0.021904 | 23.5267 | 1.7435 | 0.15557 | 0.047666 | 3.1229 | 3.1229 | 0.002057
2019-03-28 15:02 | 1553799713.7461 | 75 | 0.20008 | 0.007623 | 0.000102 | 97.3274 | 80.3112 | 106.7717 | 21.2612 | 0.003804 | 0.003562 | 252.5867 | 0.261 | 8.4029 | 0.003384 | 0.004863 | 0.004745 | 5.8e-05 | 0.001228 | 0.32244 | 1.2964 | 2.5475 | 100.69 | 21.14 | 26.37 | 113.4759 | 103.2795 | 107.5962 | 0.006577 | 0.92725 | 0.021871 | 23.5866 | 1.7189 | 0.15527 | 0.047546 | 3.1104 | 3.1104 | 0.002062
2019-03-28 15:02 | 1553799714.747 | 76 | 0.20009 | 0.007669 | 0.000101 | 97.4126 | 80.0921 | 106.541 | 21.2612 | 0.003826 | 0.003536 | 252.4493 | 0.26106 | 8.3979 | 0.003311 | 0.004601 | 0.004492 | 5.8e-05 | 0.001228 | 0.34562 | 1.3087 | 2.5475 | 100.69 | 21.13 | 26.358 | 114.0143 | 101.5184 | 106.9678 | 0.00657 | 0.92878 | 0.021865 | 23.5417 | 1.7511 | 0.15531 | 0.047408 | 3.1007 | 3.1007 | 0.002069
2019-03-28 15:02 | 1553799715.7465 | 77 | 0.20018 | 0.007711 | 0.0001 | 97.4203 | 79.2818 | 106.5442 | 21.2612 | 0.003846 | 0.003555 | 252.3446 | 0.26088 | 8.4017 | 0.003397 | 0.005332 | 0.005319 | 5.8e-05 | 0.001222 | 0.3126 | 1.3387 | 2.5475 | 100.69 | 21.13 | 26.353 | 114.4958 | 103.6016 | 107.4316 | 0.006569 | 0.92765 | 0.021852 | 23.5564 | 1.8428 | 0.1552 | 0.046835 | 3.0616 | 3.0616 | 0.002094
2019-03-28 15:02 | 1553799716.7461 | 78 | 0.20044 | 0.007742 | 9.9e-05 | 97.3486 | 79.4456 | 106.7524 | 21.2612 | 0.003851 | 0.003504 | 252.1904 | 0.26048 | 8.4034 | 0.003382 | 0.004914 | 0.004795 | 5.4e-05 | 0.001216 | 0.34324 | 1.3103 | 2.5475 | 100.69 | 21.13 | 26.353 | 114.6205 | 103.2284 | 106.187 | 0.006569 | 0.93069 | 0.021909 | 23.5408 | 1.8178 | 0.15496 | 0.045267 | 2.9671 | 2.9671 | 0.002157
2019-03-28 15:02 | 1553799717.7471 | 79 | 0.20048 | 0.007745 | 9.8e-05 | 97.3882 | 80.6363 | 106.8794 | 21.2612 | 0.003918 | 0.003589 | 252.0831 | 0.26167 | 8.4022 | 0.003406 | 0.004688 | 0.004649 | 5.7e-05 | 0.001216 | 0.34506 | 1.3245 | 2.5476 | 100.69 | 21.14 | 26.342 | 116.2511 | 103.802 | 108.2619 | 0.00657 | 0.92563 | 0.021863 | 23.6191 | 1.6883 | 0.15567 | 0.044977 | 2.942 | 2.942 | 0.002185
2019-03-28 15:02 | 1553799718.7472 | 80 | 0.20049 | 0.007734 | 9.7e-05 | 97.406 | 81.3103 | 107.3817 | 21.2612 | 0.00393 | 0.003625 | 251.9394 | 0.26249 | 8.402 | 0.003375 | 0.005045 | 0.004863 | 5.8e-05 | 0.001208 | 0.34324 | 1.3143 | 2.5475 | 100.69 | 21.14 | 26.33 | 116.5621 | 103.0491 | 109.1289 | 0.006567 | 0.92353 | 0.021889 | 23.7014 | 1.6155 | 0.15616 | 0.044964 | 2.9448 | 2.9448 | 0.00219
2019-03-28 15:02 | 1553799719.7466 | 81 | 0.20054 | 0.007708 | 9.7e-05 | 97.3903 | 80.4831 | 107.0327 | 21.2612 | 0.003948 | 0.003603 | 251.799 | 0.26081 | 8.4014 | 0.0034 | 0.004681 | 0.004695 | 5.7e-05 | 0.001205 | 0.35994 | 1.2659 | 2.5473 | 100.69 | 21.14 | 26.324 | 116.985 | 103.6741 | 108.5835 | 0.006566 | 0.92485 | 0.021869 | 23.646 | 1.7055 | 0.15515 | 0.044702 | 2.925 | 2.925 | 0.002191
2019-03-28 15:02 | 1553799720.7466 | 82 | 0.20051 | 0.007682 | 9.5e-05 | 97.325 | 80.4661 | 107.4137 | 21.2612 | 0.003944 | 0.003607 | 251.6731 | 0.26156 | 8.4016 | 0.003436 | 0.005167 | 0.005057 | 5.6e-05 | 0.001202 | 0.30215 | 1.2771 | 2.5472 | 100.69 | 21.14 | 26.318 | 116.8979 | 104.5277 | 108.6882 | 0.006564 | 0.92459 | 0.021905 | 23.6913 | 1.7014 | 0.1556 | 0.044917 | 2.944 | 2.944 | 0.002183
2019-03-28 15:02 | 1553799721.7463 | 83 | 0.2006 | 0.007656 | 9.5e-05 | 97.3947 | 80.7437 | 107.1534 | 21.2612 | 0.003984 | 0.003627 | 251.556 | 0.26098 | 8.4044 | 0.003454 | 0.005066 | 0.005001 | 5.7e-05 | 0.001202 | 0.32316 | 1.348 | 2.5474 | 100.69 | 21.14 | 26.308 | 117.868 | 104.9643 | 109.1658 | 0.006562 | 0.92344 | 0.021865 | 23.6774 | 1.677 | 0.15526 | 0.044388 | 2.9041 | 2.9041 | 0.002208
2019-03-28 15:02 | 1553799722.7467 | 84 | 0.20062 | 0.007625 | 9.5e-05 | 97.3137 | 80.6734 | 107.2454 | 21.2612 | 0.003965 | 0.003648 | 251.4536 | 0.26099 | 8.4014 | 0.003448 | 0.005174 | 0.005063 | 5.5e-05 | 0.001193 | 0.31758 | 1.3227 | 2.5473 | 100.69 | 21.14 | 26.303 | 117.4007 | 104.821 | 109.6853 | 0.00656 | 0.92219 | 0.021859 | 23.7036 | 1.6774 | 0.15526 | 0.044336 | 2.9 | 2.9 | 0.002211
2019-03-28 15:02 | 1553799723.7474 | 85 | 0.20073 | 0.007604 | 9.5e-05 | 97.325 | 81.13 | 107.1487 | 21.2612 | 0.003963 | 0.00361 | 251.3384 | 0.2618 | 8.4039 | 0.003477 | 0.004909 | 0.004831 | 5.5e-05 | 0.001189 | 0.32565 | 1.2936 | 2.5474 | 100.69 | 21.14 | 26.303 | 117.3541 | 105.5236 | 108.7513 | 0.00656 | 0.92444 | 0.021876 | 23.664 | 1.6279 | 0.15574 | 0.043694 | 2.8603 | 2.8603 | 0.002249
2019-03-28 15:02 | 1553799724.7454 | 86 | 0.2008 | 0.007575 | 9.6e-05 | 97.3387 | 80.5711 | 107.6491 | 21.2612 | 0.004017 | 0.003642 | 251.2094 | 0.26143 | 8.4047 | 0.003339 | 0.004122 | 0.004019 | 5.9e-05 | 0.001189 | 0.30541 | 1.2921 | 2.5474 | 100.69 | 21.14 | 26.289 | 118.6723 | 102.192 | 109.5339 | 0.006557 | 0.92255 | 0.021905 | 23.7435 | 1.691 | 0.15553 | 0.043325 | 2.8398 | 2.8398 | 0.002262
2019-03-28 15:02 | 1553799725.7454 | 87 | 0.20079 | 0.00754 | 9.8e-05 | 97.4617 | 81.1711 | 107.9343 | 21.2612 | 0.00395 | 0.003671 | 251.0774 | 0.26096 | 8.4026 | 0.003442 | 0.004378 | 0.004334 | 5.6e-05 | 0.001188 | 0.31133 | 1.3109 | 2.5474 | 100.69 | 21.14 | 26.286 | 117.0285 | 104.6928 | 110.243 | 0.006556 | 0.92085 | 0.021913 | 23.7969 | 1.636 | 0.15524 | 0.043384 | 2.8447 | 2.8447 | 0.002254
2019-03-28 15:02 | 1553799726.7455 | 88 | 0.2008 | 0.0075 | 9.9e-05 | 97.3992 | 80.9924 | 107.8479 | 21.2612 | 0.003986 | 0.003655 | 250.9739 | 0.26091 | 8.4042 | 0.003424 | 0.005004 | 0.004907 | 5.5e-05 | 0.001187 | 0.30309 | 1.2755 | 2.5474 | 100.69 | 21.15 | 26.283 | 117.9124 | 104.2454 | 109.8628 | 0.00656 | 0.92176 | 0.021915 | 23.7756 | 1.6499 | 0.15522 | 0.043379 | 2.8446 | 2.8446 | 0.002254
2019-03-28 15:02 | 1553799727.7455 | 89 | 0.20078 | 0.007456 | 0.000101 | 97.4671 | 80.8459 | 107.8667 | 21.2612 | 0.004043 | 0.003683 | 250.8263 | 0.26191 | 8.4044 | 0.003446 | 0.005137 | 0.005031 | 5.5e-05 | 0.00119 | 0.33644 | 1.3407 | 2.5475 | 100.69 | 21.15 | 26.283 | 119.308 | 104.774 | 110.5333 | 0.00656 | 0.92015 | 0.021898 | 23.7985 | 1.6724 | 0.15581 | 0.043536 | 2.8525 | 2.8525 | 0.002256
2019-03-28 15:02 | 1553799728.7465 | 90 | 0.20075 | 0.007415 | 0.000101 | 97.4474 | 80.5983 | 107.4054 | 21.2612 | 0.004076 | 0.003666 | 250.7128 | 0.26126 | 8.4015 | 0.00344 | 0.00507 | 0.004946 | 5.4e-05 | 0.001191 | 0.28926 | 1.3269 | 2.5475 | 100.69 | 21.15 | 26.277 | 120.1211 | 104.6378 | 110.1287 | 0.006558 | 0.92112 | 0.021863 | 23.735 | 1.698 | 0.15542 | 0.043755 | 2.8622 | 2.8622 | 0.002243
2019-03-28 15:02 | 1553799729.7466 | 91 | 0.20076 | 0.007398 | 0.000102 | 97.485 | 79.5122 | 107.5291 | 21.2612 | 0.004054 | 0.003679 | 250.6027 | 0.2613 | 8.4011 | 0.003459 | 0.005017 | 0.004978 | 5.5e-05 | 0.001194 | 0.32836 | 1.3034 | 2.5474 | 100.69 | 21.15 | 26.277 | 119.5825 | 105.0925 | 110.4275 | 0.006558 | 0.9204 | 0.021867 | 23.758 | 1.8228 | 0.15545 | 0.043768 | 2.8635 | 2.8635 | 0.002242
2019-03-28 15:02 | 1553799730.747 | 92 | 0.20075 | 0.007388 | 0.000104 | 97.5234 | 79.5733 | 106.5572 | 21.2612 | 0.004108 | 0.003683 | 250.4741 | 0.26141 | 8.403 | 0.003438 | 0.005069 | 0.005061 | 5.3e-05 | 0.00119 | 0.34219 | 1.3083 | 2.5475 | 100.69 | 21.15 | 26.271 | 120.9027 | 104.5883 | 110.5444 | 0.006557 | 0.92012 | 0.021764 | 23.6539 | 1.8194 | 0.15551 | 0.043787 | 2.8513 | 2.8513 | 0.002253
2019-03-28 15:02 | 1553799731.7467 | 93 | 0.20082 | 0.00738 | 0.000106 | 97.5198 | 80.5093 | 104.5928 | 21.2612 | 0.004102 | 0.003729 | 250.332 | 0.2602 | 8.4017 | 0.003498 | 0.005 | 0.00495 | 5.3e-05 | 0.001187 | 0.32101 | 1.3034 | 2.5475 | 100.69 | 21.15 | 26.271 | 120.7422 | 106.031 | 111.6546 | 0.006557 | 0.91747 | 0.021532 | 23.4688 | 1.7145 | 0.15479 | 0.043381 | 2.7946 | 2.7946 | 0.002288
2019-03-28 15:02 | 1553799732.7458 | 94 | 0.20084 | 0.007377 | 0.000109 | 97.5499 | 80.4243 | 102.1061 | 21.2612 | 0.004073 | 0.003768 | 250.2101 | 0.26152 | 8.4034 | 0.003431 | 0.004748 | 0.004621 | 5.2e-05 | 0.00118 | 0.3178 | 1.2875 | 2.5475 | 100.69 | 21.15 | 26.262 | 120.037 | 104.416 | 112.6117 | 0.006554 | 0.91519 | 0.021248 | 23.2169 | 1.7267 | 0.15558 | 0.043305 | 2.7527 | 2.7527 | 0.002334
2019-03-28 15:02 | 1553799733.7467 | 95 | 0.20082 | 0.007376 | 0.000115 | 97.5165 | 80.5446 | 99.3964 | 21.2612 | 0.004145 | 0.003945 | 250.0497 | 0.26161 | 8.4019 | 0.003488 | 0.005052 | 0.005043 | 5.5e-05 | 0.001175 | 0.32261 | 1.3265 | 2.5474 | 100.69 | 21.15 | 26.262 | 121.8036 | 105.8028 | 116.9074 | 0.006554 | 0.90511 | 0.020848 | 23.034 | 1.7103 | 0.15563 | 0.043423 | 2.708 | 2.708 | 0.002374
2019-03-28 15:02 | 1553799734.746 | 96 | 0.20054 | 0.007379 | 0.000125 | 97.5845 | 80.4386 | 95.9777 | 21.2612 | 0.00417 | 0.004035 | 249.8893 | 0.26077 | 8.4021 | 0.003433 | 0.004949 | 0.004839 | 5.8e-05 | 0.001172 | 0.31725 | 1.35 | 2.5474 | 100.69 | 21.15 | 26.249 | 122.4143 | 104.4587 | 119.1267 | 0.006551 | 0.89999 | 0.020429 | 22.6987 | 1.7283 | 0.15513 | 0.045053 | 2.7524 | 2.7524 | 0.002328
2019-03-28 15:02 | 1553799735.7474 | 97 | 0.20004 | 0.007375 | 0.000138 | 97.5445 | 80.4936 | 93.0293 | 21.2612 | 0.004387 | 0.004267 | 249.7133 | 0.26103 | 8.4034 | 0.003481 | 0.004659 | 0.004562 | 5.5e-05 | 0.001175 | 0.32012 | 1.289 | 2.5473 | 100.69 | 21.15 | 26.249 | 127.7328 | 105.6226 | 124.8023 | 0.006551 | 0.88716 | 0.019968 | 22.5084 | 1.7185 | 0.15529 | 0.048046 | 2.868 | 2.868 | 0.002236
2019-03-28 15:02 | 1553799736.7461 | 98 | 0.19808 | 0.007413 | 0.00015 | 97.5916 | 80.3635 | 89.5552 | 21.2612 | 0.004473 | 0.004359 | 249.4829 | 0.2611 | 8.406 | 0.003421 | 0.004339 | 0.004258 | 5.6e-05 | 0.001178 | 0.3501 | 1.2798 | 2.5474 | 100.69 | 21.15 | 26.238 | 129.8486 | 104.1652 | 127.0421 | 0.006548 | 0.88219 | 0.019537 | 22.1462 | 1.7373 | 0.15532 | 0.059645 | 3.4803 | 3.4803 | 0.001843
2019-03-28 15:02 | 1553799737.7458 | 99 | 0.19512 | 0.007648 | 0.000162 | 97.4662 | 77.786 | 86.1982 | 21.2612 | 0.004704 | 0.004556 | 249.2681 | 0.26022 | 8.4032 | 0.003487 | 0.004873 | 0.004759 | 5.6e-05 | 0.001174 | 0.29966 | 1.3206 | 2.5474 | 100.69 | 21.15 | 26.238 | 135.5504 | 105.7831 | 131.8984 | 0.006548 | 0.87161 | 0.019052 | 21.8586 | 2.0174 | 0.1548 | 0.076815 | 4.3646 | 4.3646 | 0.001465
2019-03-28 15:02 | 1553799738.7458 | 100 | 0.19158 | 0.008735 | 0.000175 | 97.4677 | 72.9392 | 83.9847 | 21.2612 | 0.004831 | 0.004712 | 249.0124 | 0.26039 | 8.4042 | 0.003514 | 0.005103 | 0.005087 | 5.7e-05 | 0.001174 | 0.31404 | 1.2959 | 2.5473 | 100.69 | 21.15 | 26.234 | 138.6834 | 106.4269 | 135.7415 | 0.006547 | 0.86342 | 0.018717 | 21.6782 | 2.593 | 0.15491 | 0.096324 | 5.3674 | 5.3674 | 0.001192
2019-03-28 15:02 | 1553799739.746 | 101 | 0.1858 | 0.010465 | 0.000194 | 97.3804 | 71.6785 | 81.8309 | 21.2612 | 0.005078 | 0.004848 | 248.7638 | 0.26087 | 8.4023 | 0.003534 | 0.005164 | 0.005106 | 6e-05 | 0.001172 | 0.31061 | 1.2905 | 2.5472 | 100.69 | 21.15 | 26.228 | 144.7965 | 106.9253 | 139.097 | 0.006546 | 0.85639 | 0.018401 | 21.4861 | 2.7409 | 0.15519 | 0.12801 | 6.9909 | 6.9909 | 0.000917
2019-03-28 15:02 | 1553799740.7459 | 102 | 0.18166 | 0.012734 | 0.000242 | 97.3992 | 74.862 | 80.0447 | 21.2612 | 0.005223 | 0.00488 | 248.4546 | 0.26145 | 8.3994 | 0.003555 | 0.005344 | 0.005257 | 5.2e-05 | 0.001172 | 0.33069 | 1.2984 | 2.5474 | 100.69 | 21.15 | 26.228 | 148.3881 | 107.4277 | 139.8973 | 0.006546 | 0.85473 | 0.018181 | 21.2709 | 2.3539 | 0.15553 | 0.14938 | 8.0422 | 8.0422 | 0.000799
2019-03-28 15:02 | 1553799741.745 | 103 | 0.17927 | 0.015415 | 0.000321 | 97.2759 | 79.4724 | 79.4522 | 21.2612 | 0.005402 | 0.004971 | 248.1468 | 0.2598 | 8.402 | 0.003624 | 0.005096 | 0.005042 | 5.4e-05 | 0.001173 | 0.33196 | 1.2972 | 2.5473 | 100.69 | 21.15 | 26.227 | 152.838 | 109.1131 | 142.1483 | 0.006546 | 0.8501 | 0.018064 | 21.2497 | 1.8081 | 0.15455 | 0.16016 | 8.5551 | 8.5551 | 0.000746
2019-03-28 15:02 | 1553799742.746 | 104 | 0.17574 | 0.018305 | 0.000416 | 97.2684 | 79.8196 | 80.0332 | 21.2612 | 0.005432 | 0.005078 | 247.79 | 0.26056 | 8.4048 | 0.003688 | 0.005001 | 0.004859 | 5.6e-05 | 0.001167 | 0.31006 | 1.3308 | 2.5475 | 100.69 | 21.15 | 26.227 | 153.5698 | 110.6478 | 144.7836 | 0.006546 | 0.84474 | 0.018073 | 21.3948 | 1.7684 | 0.15501 | 0.17724 | 9.4524 | 9.4524 | 0.000677
2019-03-28 15:02 | 1553799743.746 | 105 | 0.17445 | 0.021274 | 0.000514 | 97.2532 | 79.8322 | 80.8167 | 21.2612 | 0.005557 | 0.005153 | 247.4494 | 0.26044 | 8.405 | 0.003682 | 0.004997 | 0.004963 | 5.5e-05 | 0.00117 | 0.31166 | 1.3031 | 2.5474 | 100.69 | 21.16 | 26.226 | 156.6878 | 110.5017 | 146.6416 | 0.006549 | 0.841 | 0.018121 | 21.547 | 1.7656 | 0.15494 | 0.18149 | 9.697 | 9.697 | 0.00066
2019-03-28 15:02 | 1553799744.7461 | 106 | 0.1737 | 0.024051 | 0.000604 | 97.2139 | 78.7227 | 81.9272 | 21.2612 | 0.005669 | 0.005195 | 247.0671 | 0.26161 | 8.4009 | 0.00372 | 0.005714 | 0.005595 | 5.4e-05 | 0.001173 | 0.31852 | 1.3466 | 2.5474 | 100.69 | 21.16 | 26.226 | 159.4752 | 111.4336 | 147.6933 | 0.006549 | 0.8389 | 0.018222 | 21.7217 | 1.8871 | 0.15563 | 0.18288 | 9.8205 | 9.8205 | 0.000655
2019-03-28 15:02 | 1553799745.7461 | 107 | 0.17297 | 0.026317 | 0.000676 | 97.2523 | 75.5425 | 82.7006 | 21.2612 | 0.005735 | 0.005273 | 246.6815 | 0.26073 | 8.4036 | 0.003831 | 0.005017 | 0.004955 | 5.4e-05 | 0.001173 | 0.3183 | 1.341 | 2.5473 | 100.69 | 21.16 | 26.23 | 161.1316 | 114.1288 | 149.6311 | 0.00655 | 0.83505 | 0.018266 | 21.8742 | 2.2595 | 0.15511 | 0.18468 | 9.9358 | 9.9358 | 0.000645
2019-03-28 15:02 | 1553799746.745 | 108 | 0.17264 | 0.027395 | 0.000723 | 97.1579 | 72.6905 | 83.8984 | 21.2612 | 0.005846 | 0.005323 | 246.3263 | 0.26086 | 8.4016 | 0.003915 | 0.005567 | 0.005503 | 5.3e-05 | 0.001174 | 0.30685 | 1.3347 | 2.5472 | 100.69 | 21.16 | 26.234 | 163.9 | 116.195 | 150.8663 | 0.006551 | 0.83262 | 0.018371 | 22.0642 | 2.5951 | 0.15519 | 0.18541 | 10.0299 | 10.0299 | 0.000639
2019-03-28 15:02 | 1553799747.7461 | 109 | 0.17278 | 0.027851 | 0.000752 | 97.0579 | 72.7313 | 84.6168 | 21.2612 | 0.005888 | 0.005365 | 245.9162 | 0.26044 | 8.4043 | 0.004077 | 0.005486 | 0.005478 | 5.2e-05 | 0.001175 | 0.32316 | 1.3191 | 2.5474 | 100.69 | 21.16 | 26.235 | 164.9421 | 120.137 | 151.8983 | 0.006552 | 0.8306 | 0.018427 | 22.1854 | 2.5808 | 0.15494 | 0.18409 | 9.9885 | 9.9885 | 0.000641
2019-03-28 15:02 | 1553799748.7451 | 110 | 0.17289 | 0.028264 | 0.000779 | 97.0358 | 71.6061 | 84.5018 | 21.2612 | 0.005952 | 0.005396 | 245.5514 | 0.2612 | 8.3995 | 0.004198 | 0.005803 | 0.005735 | 5e-05 | 0.00118 | 0.33373 | 1.313 | 2.5475 | 100.69 | 21.16 | 26.235 | 166.5288 | 123.1095 | 152.6719 | 0.006552 | 0.82909 | 0.018398 | 22.1905 | 2.7182 | 0.15539 | 0.18304 | 9.9156 | 9.9156 | 0.000647
2019-03-28 15:02 | 1553799749.7466 | 111 | 0.17315 | 0.02858 | 0.000799 | 97.0141 | 70.2309 | 84.007 | 21.2612 | 0.006091 | 0.005493 | 245.1006 | 0.26087 | 8.4034 | 0.004271 | 0.006067 | 0.005973 | 5.2e-05 | 0.001176 | 0.34562 | 1.332 | 2.5475 | 100.69 | 21.16 | 26.237 | 169.9972 | 124.8952 | 155.1017 | 0.006552 | 0.82439 | 0.018292 | 22.1884 | 2.8897 | 0.15519 | 0.18118 | 9.7589 | 9.7589 | 0.000657
2019-03-28 15:02 | 1553799750.7453 | 112 | 0.17317 | 0.028786 | 0.000811 | 97.0248 | 67.8447 | 83.639 | 21.2612 | 0.006037 | 0.005608 | 244.6971 | 0.26016 | 8.4021 | 0.004294 | 0.005577 | 0.005455 | 5.6e-05 | 0.001174 | 0.32084 | 1.3719 | 2.5474 | 100.69 | 21.16 | 26.237 | 168.6639 | 125.4435 | 157.9514 | 0.006552 | 0.81894 | 0.018191 | 22.2133 | 3.1999 | 0.15477 | 0.18082 | 9.6857 | 9.6857 | 0.00066
2019-03-28 15:02 | 1553799751.7453 | 113 | 0.17315 | 0.028871 | 0.00081 | 96.9131 | 67.3059 | 82.9233 | 21.2612 | 0.006151 | 0.005663 | 244.2574 | 0.26133 | 8.4036 | 0.004315 | 0.005995 | 0.005845 | 5.5e-05 | 0.001183 | 0.32532 | 1.328 | 2.5474 | 100.69 | 21.16 | 26.238 | 171.5101 | 125.9758 | 159.3249 | 0.006552 | 0.81634 | 0.018085 | 22.1533 | 3.2609 | 0.15547 | 0.18086 | 9.6306 | 9.6306 | 0.000667
2019-03-28 15:03 | 1553799752.7467 | 114 | 0.17284 | 0.028888 | 0.0008 | 96.8422 | 68.022 | 81.7145 | 21.2612 | 0.006298 | 0.005772 | 243.8253 | 0.25973 | 8.4022 | 0.004361 | 0.005745 | 0.005709 | 5.7e-05 | 0.001187 | 0.33771 | 1.3053 | 2.5474 | 100.69 | 21.16 | 26.238 | 175.1699 | 127.089 | 162.04 | 0.006552 | 0.81124 | 0.017896 | 22.0602 | 3.1597 | 0.15451 | 0.18266 | 9.6242 | 9.6242 | 0.000663
2019-03-28 15:03 | 1553799753.7463 | 115 | 0.17245 | 0.028885 | 0.000783 | 96.7859 | 68.4487 | 80.4287 | 21.2612 | 0.006473 | 0.005917 | 243.3766 | 0.2603 | 8.4043 | 0.004484 | 0.005804 | 0.005727 | 5.9e-05 | 0.00119 | 0.33411 | 1.3381 | 2.5474 | 100.69 | 21.17 | 26.237 | 179.5691 | 130.1264 | 165.654 | 0.006556 | 0.80456 | 0.017682 | 21.9766 | 3.0985 | 0.15485 | 0.18492 | 9.6254 | 9.6254 | 0.000664
2019-03-28 15:03 | 1553799754.7453 | 116 | 0.17121 | 0.028933 | 0.000762 | 96.7192 | 69.8392 | 78.8581 | 21.2612 | 0.006529 | 0.005989 | 242.9375 | 0.26115 | 8.4026 | 0.004509 | 0.006319 | 0.006213 | 5.7e-05 | 0.001194 | 0.34075 | 1.3132 | 2.5474 | 100.69 | 21.17 | 26.231 | 180.9647 | 130.7288 | 167.4523 | 0.006555 | 0.80128 | 0.017472 | 21.8055 | 2.9125 | 0.15536 | 0.19201 | 9.8712 | 9.8712 | 0.00065
2019-03-28 15:03 | 1553799755.7453 | 117 | 0.16983 | 0.029061 | 0.000737 | 96.6632 | 71.2606 | 77.1735 | 21.2612 | 0.006694 | 0.006132 | 242.4773 | 0.26083 | 8.4012 | 0.004525 | 0.00582 | 0.00578 | 5.7e-05 | 0.001193 | 0.30215 | 1.3343 | 2.5474 | 100.69 | 21.17 | 26.227 | 185.0954 | 131.1192 | 171.0278 | 0.006554 | 0.79483 | 0.017215 | 21.6587 | 2.7271 | 0.15517 | 0.19982 | 10.1157 | 10.1157 | 0.000634
2019-03-28 15:03 | 1553799756.7454 | 118 | 0.16773 | 0.029317 | 0.000706 | 96.5485 | 71.7199 | 76.1706 | 21.2612 | 0.006822 | 0.006222 | 242.0142 | 0.26031 | 8.4013 | 0.004566 | 0.005919 | 0.005892 | 5.7e-05 | 0.001194 | 0.32997 | 1.3026 | 2.5474 | 100.69 | 21.17 | 26.223 | 188.2941 | 132.1293 | 173.2814 | 0.006553 | 0.79082 | 0.017059 | 21.572 | 2.659 | 0.15486 | 0.21155 | 10.6036 | 10.6036 | 0.000603
2019-03-28 15:03 | 1553799757.7455 | 119 | 0.16677 | 0.029781 | 0.000663 | 96.4118 | 71.6132 | 75.0368 | 21.2612 | 0.006949 | 0.00639 | 241.5947 | 0.26005 | 8.4047 | 0.00461 | 0.006053 | 0.005899 | 5.6e-05 | 0.001194 | 0.32068 | 1.3264 | 2.5475 | 100.69 | 21.17 | 26.223 | 191.4641 | 133.235 | 177.4931 | 0.006553 | 0.78343 | 0.016853 | 21.5116 | 2.6597 | 0.15471 | 0.21659 | 10.7211 | 10.7211 | 0.000596
2019-03-28 15:03 | 1553799758.7455 | 120 | 0.16474 | 0.030973 | 0.000618 | 96.2864 | 72.8434 | 73.9674 | 21.2612 | 0.007068 | 0.006538 | 241.0952 | 0.26043 | 8.4029 | 0.004753 | 0.006189 | 0.006069 | 5.4e-05 | 0.001191 | 0.32731 | 1.3195 | 2.5475 | 100.69 | 21.17 | 26.206 | 194.4421 | 136.7519 | 181.1796 | 0.006548 | 0.77707 | 0.016664 | 21.445 | 2.4957 | 0.15493 | 0.22695 | 11.0997 | 11.0997 | 0.000576
2019-03-28 15:03 | 1553799759.7459 | 121 | 0.1627 | 0.032319 | 0.000576 | 96.2715 | 73.8609 | 73.2783 | 21.2612 | 0.007262 | 0.006563 | 240.6541 | 0.26035 | 8.4028 | 0.004764 | 0.005682 | 0.005613 | 5.7e-05 | 0.001191 | 0.29358 | 1.34 | 2.5474 | 100.69 | 21.17 | 26.206 | 199.2867 | 137.0153 | 181.8071 | 0.006548 | 0.776 | 0.016575 | 21.3596 | 2.3702 | 0.15488 | 0.23721 | 11.5301 | 11.5301 | 0.000555
2019-03-28 15:03 | 1553799760.7456 | 122 | 0.1615 | 0.033633 | 0.000543 | 96.1244 | 74.1591 | 72.8053 | 21.2612 | 0.00744 | 0.006681 | 240.1749 | 0.25985 | 8.4034 | 0.004828 | 0.006127 | 0.006045 | 5.8e-05 | 0.001188 | 0.3314 | 1.3137 | 2.5474 | 100.69 | 21.17 | 26.191 | 203.7536 | 138.5898 | 184.7656 | 0.006545 | 0.77098 | 0.016468 | 21.3596 | 2.3205 | 0.15458 | 0.24279 | 11.7202 | 11.7202 | 0.000545
2019-03-28 15:03 | 1553799761.7456 | 123 | 0.16026 | 0.034825 | 0.000502 | 96.094 | 74.7176 | 72.5945 | 21.2612 | 0.007539 | 0.006722 | 239.7305 | 0.26059 | 8.4023 | 0.004884 | 0.0059 | 0.005812 | 5.5e-05 | 0.001184 | 0.30845 | 1.2959 | 2.5473 | 100.69 | 21.17 | 26.191 | 206.2299 | 139.9727 | 185.7765 | 0.006545 | 0.76928 | 0.016426 | 21.3522 | 2.2505 | 0.15502 | 0.24868 | 11.9689 | 11.9689 | 0.000535
2019-03-28 15:03 | 1553799762.7458 | 124 | 0.15957 | 0.035646 | 0.000475 | 96.1017 | 75.8972 | 72.5941 | 21.2612 | 0.007735 | 0.006832 | 239.2704 | 0.26141 | 8.4021 | 0.004939 | 0.006386 | 0.006266 | 5.3e-05 | 0.001182 | 0.3418 | 1.3389 | 2.5474 | 100.69 | 21.18 | 26.179 | 211.1074 | 141.3521 | 188.5351 | 0.006546 | 0.76469 | 0.016377 | 21.4162 | 2.1112 | 0.15551 | 0.25183 | 12.0817 | 12.0817 | 0.000532
2019-03-28 15:03 | 1553799763.7457 | 125 | 0.15947 | 0.036571 | 0.000449 | 95.8921 | 75.9943 | 72.9898 | 21.2612 | 0.007799 | 0.006886 | 238.8323 | 0.26003 | 8.4012 | 0.005009 | 0.006044 | 0.006017 | 5.7e-05 | 0.00118 | 0.34899 | 1.2763 | 2.5474 | 100.69 | 21.18 | 26.176 | 212.7115 | 143.0825 | 189.8893 | 0.006545 | 0.76245 | 0.016397 | 21.506 | 2.0802 | 0.15469 | 0.25153 | 12.084 | 12.084 | 0.000529
2019-03-28 15:03 | 1553799764.7457 | 126 | 0.15947 | 0.037158 | 0.000422 | 95.803 | 77.4368 | 73.2086 | 21.2612 | 0.00786 | 0.006999 | 238.3866 | 0.26093 | 8.4022 | 0.005065 | 0.005875 | 0.005622 | 5.7e-05 | 0.001178 | 0.3178 | 1.3266 | 2.5474 | 100.69 | 21.18 | 26.174 | 214.2367 | 144.4663 | 192.7118 | 0.006544 | 0.75783 | 0.016372 | 21.6037 | 1.9037 | 0.15523 | 0.25103 | 12.0431 | 12.0431 | 0.000532
2019-03-28 15:03 | 1553799765.7447 | 127 | 0.15948 | 0.037684 | 0.000398 | 95.789 | 78.3228 | 73.6919 | 21.2612 | 0.008011 | 0.007026 | 237.945 | 0.2607 | 8.4037 | 0.005163 | 0.007287 | 0.007209 | 5.7e-05 | 0.001183 | 0.33771 | 1.2837 | 2.5473 | 100.69 | 21.18 | 26.177 | 217.9792 | 146.9008 | 193.3937 | 0.006545 | 0.75672 | 0.016414 | 21.6908 | 1.8006 | 0.15509 | 0.25052 | 12.0509 | 12.0509 | 0.000532
2019-03-28 15:03 | 1553799766.7448 | 128 | 0.15979 | 0.038157 | 0.000378 | 95.7768 | 78.776 | 74.3364 | 21.2612 | 0.008082 | 0.007069 | 237.5234 | 0.2605 | 8.401 | 0.005096 | 0.00684 | 0.006743 | 5.5e-05 | 0.001187 | 0.32532 | 1.2957 | 2.5474 | 100.69 | 21.18 | 26.18 | 219.7434 | 145.2386 | 194.4739 | 0.006546 | 0.75498 | 0.016467 | 21.8107 | 1.7479 | 0.15497 | 0.2483 | 11.9858 | 11.9858 | 0.000534
2019-03-28 15:03 | 1553799767.7457 | 129 | 0.16042 | 0.038543 | 0.000361 | 95.7166 | 79.1426 | 75.1488 | 21.2612 | 0.008146 | 0.007124 | 237.0722 | 0.25959 | 8.4033 | 0.005141 | 0.006722 | 0.006577 | 5.7e-05 | 0.001187 | 0.32659 | 1.2699 | 2.5472 | 100.69 | 21.18 | 26.19 | 221.3387 | 146.3527 | 195.8428 | 0.006548 | 0.75277 | 0.016532 | 21.9616 | 1.7007 | 0.15443 | 0.24442 | 11.8505 | 11.8505 | 0.000538
2019-03-28 15:03 | 1553799768.7448 | 130 | 0.16167 | 0.03882 | 0.000345 | 95.6088 | 78.8997 | 75.9302 | 21.2612 | 0.008214 | 0.007137 | 236.6781 | 0.26052 | 8.4064 | 0.005235 | 0.006406 | 0.00634 | 5.7e-05 | 0.001182 | 0.32587 | 1.2959 | 2.5475 | 100.69 | 21.18 | 26.196 | 223.0238 | 148.6855 | 196.1724 | 0.00655 | 0.75225 | 0.016612 | 22.0832 | 1.7181 | 0.15499 | 0.23711 | 11.5602 | 11.5602 | 0.000554
2019-03-28 15:03 | 1553799769.7449 | 131 | 0.16272 | 0.038951 | 0.00033 | 95.5808 | 77.3559 | 76.6595 | 21.2612 | 0.008253 | 0.007126 | 236.2462 | 0.25976 | 8.4046 | 0.005275 | 0.006805 | 0.006669 | 5.7e-05 | 0.001178 | 0.31022 | 1.3073 | 2.5475 | 100.69 | 21.18 | 26.212 | 223.9944 | 149.6818 | 195.8948 | 0.006554 | 0.75269 | 0.016697 | 22.1825 | 1.8923 | 0.15453 | 0.23102 | 11.3281 | 11.3281 | 0.000563
2019-03-28 15:03 | 1553799770.7459 | 132 | 0.16362 | 0.03891 | 0.000315 | 95.512 | 75.4915 | 77.3054 | 21.2612 | 0.008275 | 0.007217 | 235.8579 | 0.25983 | 8.402 | 0.005231 | 0.007177 | 0.007106 | 5.7e-05 | 0.001185 | 0.3282 | 1.3162 | 2.5474 | 100.69 | 21.18 | 26.212 | 224.5595 | 148.579 | 198.1762 | 0.006554 | 0.74905 | 0.016726 | 22.3298 | 2.104 | 0.15457 | 0.22593 | 11.1042 | 11.1042 | 0.000575
2019-03-28 15:03 | 1553799771.7439 | 133 | 0.16459 | 0.038715 | 0.000302 | 95.481 | 76.0164 | 78.1089 | 21.2612 | 0.008347 | 0.007186 | 235.4485 | 0.26034 | 8.4038 | 0.005283 | 0.006863 | 0.006729 | 5.2e-05 | 0.001176 | 0.3402 | 1.2922 | 2.5473 | 100.69 | 21.18 | 26.227 | 226.3303 | 149.8839 | 197.408 | 0.006558 | 0.75027 | 0.016826 | 22.4273 | 2.0392 | 0.15488 | 0.22064 | 10.9154 | 10.9154 | 0.000586
2019-03-28 15:03 | 1553799772.745 | 134 | 0.16594 | 0.038374 | 0.000289 | 95.3729 | 76.6337 | 78.8659 | 21.2612 | 0.008336 | 0.007206 | 235.0704 | 0.25975 | 8.4019 | 0.00531 | 0.007032 | 0.006872 | 5.5e-05 | 0.001178 | 0.29806 | 1.3066 | 2.5474 | 100.69 | 21.18 | 26.227 | 226.0716 | 150.5359 | 197.8924 | 0.006558 | 0.7495 | 0.016899 | 22.5473 | 1.9567 | 0.15452 | 0.21321 | 10.6017 | 10.6017 | 0.000602
2019-03-28 15:03 | 1553799773.744 | 135 | 0.16754 | 0.037912 | 0.000276 | 95.3634 | 76.4621 | 79.2938 | 21.2612 | 0.008376 | 0.007198 | 234.676 | 0.25987 | 8.4031 | 0.005292 | 0.006603 | 0.00659 | 5.7e-05 | 0.001171 | 0.29734 | 1.3193 | 2.5474 | 100.69 | 21.18 | 26.249 | 227.0655 | 150.0918 | 197.6984 | 0.006563 | 0.74981 | 0.016948 | 22.6037 | 1.9758 | 0.1546 | 0.20453 | 10.2089 | 10.2089 | 0.000625
2019-03-28 15:03 | 1553799774.7441 | 136 | 0.16866 | 0.037282 | 0.000262 | 95.2791 | 76.8871 | 80.3123 | 21.2612 | 0.008425 | 0.007224 | 234.3222 | 0.25916 | 8.4026 | 0.005325 | 0.007039 | 0.006901 | 5.7e-05 | 0.001171 | 0.347 | 1.2899 | 2.5474 | 100.69 | 21.18 | 26.249 | 228.2682 | 150.9098 | 198.358 | 0.006563 | 0.74876 | 0.017045 | 22.7643 | 1.9184 | 0.15417 | 0.19873 | 9.9816 | 9.9816 | 0.000638
2019-03-28 15:03 | 1553799775.745 | 137 | 0.16961 | 0.036178 | 0.000248 | 95.2615 | 76.9252 | 80.8362 | 21.2612 | 0.008431 | 0.007165 | 233.957 | 0.26087 | 8.4019 | 0.005365 | 0.006879 | 0.006694 | 5.7e-05 | 0.001175 | 0.32421 | 1.3154 | 2.5474 | 100.69 | 21.18 | 26.265 | 228.4195 | 151.8968 | 196.8782 | 0.006567 | 0.75111 | 0.017127 | 22.8026 | 1.9123 | 0.15519 | 0.1944 | 9.8164 | 9.8164 | 0.000653
2019-03-28 15:03 | 1553799776.7451 | 138 | 0.17071 | 0.035212 | 0.000234 | 95.2177 | 78.3663 | 81.2342 | 21.2612 | 0.008457 | 0.007214 | 233.5965 | 0.2599 | 8.4053 | 0.005402 | 0.006866 | 0.006768 | 5.7e-05 | 0.001175 | 0.32029 | 1.3317 | 2.5473 | 100.69 | 21.18 | 26.265 | 229.0584 | 152.8283 | 198.1097 | 0.006567 | 0.74915 | 0.017147 | 22.8886 | 1.7421 | 0.15461 | 0.18902 | 9.5612 | 9.5612 | 0.000668
2019-03-28 15:03 | 1553799777.7441 | 139 | 0.17135 | 0.034348 | 0.000222 | 95.218 | 79.5693 | 81.8696 | 21.2612 | 0.008421 | 0.007238 | 233.2531 | 0.26057 | 8.4028 | 0.00537 | 0.006596 | 0.006505 | 5.7e-05 | 0.001178 | 0.35723 | 1.2999 | 2.5473 | 100.69 | 21.18 | 26.277 | 228.1871 | 152.0251 | 198.7012 | 0.00657 | 0.74821 | 0.017203 | 22.9924 | 1.6059 | 0.15502 | 0.18627 | 9.4559 | 9.4559 | 0.000677
2019-03-28 15:03 | 1553799778.7452 | 140 | 0.17243 | 0.033425 | 0.00021 | 95.106 | 80.8167 | 82.2389 | 21.2612 | 0.00844 | 0.007102 | 232.8786 | 0.25968 | 8.4003 | 0.005358 | 0.006764 | 0.006659 | 5.9e-05 | 0.001175 | 0.30574 | 1.3317 | 2.5476 | 100.69 | 21.18 | 26.28 | 228.6408 | 151.7323 | 195.3016 | 0.006571 | 0.75364 | 0.017304 | 22.961 | 1.4562 | 0.15449 | 0.18096 | 9.2455 | 9.2455 | 0.00069
2019-03-28 15:03 | 1553799779.7445 | 141 | 0.17349 | 0.032512 | 0.000201 | 95.0128 | 81.3698 | 82.6156 | 21.2612 | 0.008449 | 0.007153 | 232.576 | 0.25995 | 8.4016 | 0.005286 | 0.007154 | 0.007055 | 5.8e-05 | 0.001178 | 0.32781 | 1.3178 | 2.5474 | 100.69 | 21.18 | 26.284 | 228.859 | 149.9557 | 196.5658 | 0.006572 | 0.75162 | 0.017321 | 23.0445 | 1.3865 | 0.15464 | 0.17581 | 8.9961 | 8.9961 | 0.00071
2019-03-28 15:03 | 1553799780.7454 | 142 | 0.17417 | 0.031625 | 0.000192 | 95.0256 | 80.9784 | 83.293 | 21.2612 | 0.008424 | 0.007173 | 232.2638 | 0.25926 | 8.4019 | 0.005336 | 0.00656 | 0.006524 | 6e-05 | 0.00118 | 0.29806 | 1.3018 | 2.5473 | 100.69 | 21.18 | 26.294 | 228.2417 | 151.192 | 197.0722 | 0.006574 | 0.7508 | 0.017382 | 23.1513 | 1.4308 | 0.15423 | 0.17278 | 8.8751 | 8.8751 | 0.000718
2019-03-28 15:03 | 1553799781.7442 | 143 | 0.17507 | 0.030752 | 0.000184 | 94.8972 | 82.0557 | 83.7984 | 21.2612 | 0.008454 | 0.007183 | 231.9669 | 0.25988 | 8.4017 | 0.005407 | 0.006367 | 0.006346 | 5.7e-05 | 0.001175 | 0.31725 | 1.3297 | 2.5474 | 100.69 | 21.18 | 26.294 | 228.9957 | 152.9605 | 197.3245 | 0.006574 | 0.7504 | 0.01743 | 23.2277 | 1.3005 | 0.1546 | 0.16847 | 8.6818 | 8.6818 | 0.000735
2019-03-28 15:03 | 1553799782.7446 | 144 | 0.17625 | 0.029978 | 0.000178 | 94.7656 | 82.1484 | 83.9178 | 21.2612 | 0.008444 | 0.007092 | 231.6362 | 0.25964 | 8.4024 | 0.005253 | 0.006579 | 0.006419 | 6e-05 | 0.001176 | 0.34435 | 1.3226 | 2.5475 | 100.69 | 21.18 | 26.31 | 228.7569 | 149.1212 | 195.0499 | 0.006578 | 0.75405 | 0.017485 | 23.1879 | 1.278 | 0.15446 | 0.16246 | 8.4036 | 8.4036 | 0.000759
2019-03-28 15:03 | 1553799783.7443 | 145 | 0.1776 | 0.02922 | 0.000174 | 94.7054 | 81.8125 | 83.9268 | 21.2612 | 0.008484 | 0.007153 | 231.3522 | 0.25906 | 8.4035 | 0.005294 | 0.00706 | 0.00706 | 5.6e-05 | 0.001166 | 0.28926 | 1.2996 | 2.5474 | 100.69 | 21.18 | 26.31 | 229.7287 | 150.1378 | 196.5683 | 0.006578 | 0.75161 | 0.017457 | 23.2268 | 1.309 | 0.15412 | 0.15541 | 8.0318 | 8.0318 | 0.000792
2019-03-28 15:03 | 1553799784.7444 | 146 | 0.17826 | 0.02848 | 0.000172 | 94.6071 | 81.6796 | 84.0401 | 21.2612 | 0.008414 | 0.007087 | 231.0266 | 0.2597 | 8.4016 | 0.005307 | 0.006808 | 0.006606 | 5.5e-05 | 0.001164 | 0.3621 | 1.2794 | 2.5473 | 100.69 | 21.18 | 26.323 | 227.9963 | 150.4634 | 194.9299 | 0.006582 | 0.75424 | 0.0175 | 23.2019 | 1.3142 | 0.1545 | 0.15238 | 7.897 | 7.897 | 0.000808
2019-03-28 15:03 | 1553799785.7453 | 147 | 0.17872 | 0.027955 | 0.00017 | 94.5082 | 81.6035 | 84.3516 | 21.2612 | 0.008414 | 0.007061 | 230.7397 | 0.25945 | 8.4033 | 0.005264 | 0.006619 | 0.006524 | 5.6e-05 | 0.001166 | 0.3298 | 1.3066 | 2.5474 | 100.69 | 21.18 | 26.323 | 227.9913 | 149.4062 | 194.2758 | 0.006582 | 0.7553 | 0.017544 | 23.2286 | 1.3132 | 0.15435 | 0.15026 | 7.8082 | 7.8082 | 0.000816
2019-03-28 15:03 | 1553799786.7434 | 148 | 0.17933 | 0.027495 | 0.00017 | 94.4293 | 82.2786 | 84.4903 | 21.2612 | 0.008399 | 0.007111 | 230.4485 | 0.25933 | 8.403 | 0.005225 | 0.007285 | 0.007212 | 5.5e-05 | 0.001166 | 0.2995 | 1.3088 | 2.5473 | 100.69 | 21.19 | 26.341 | 227.6179 | 148.4373 | 195.5142 | 0.00659 | 0.7533 | 0.017536 | 23.2784 | 1.232 | 0.15428 | 0.14717 | 7.6464 | 7.6464 | 0.000833
2019-03-28 15:03 | 1553799787.7444 | 149 | 0.17977 | 0.027029 | 0.000169 | 94.3382 | 82.8214 | 84.7938 | 21.2612 | 0.008382 | 0.007122 | 230.1781 | 0.25948 | 8.4024 | 0.005265 | 0.006368 | 0.006249 | 6.1e-05 | 0.001164 | 0.31542 | 1.3312 | 2.5474 | 100.69 | 21.19 | 26.341 | 227.2137 | 149.4248 | 195.7893 | 0.00659 | 0.75286 | 0.017562 | 23.3271 | 1.1646 | 0.15436 | 0.14511 | 7.552 | 7.552 | 0.000844
2019-03-28 15:03 | 1553799788.7444 | 150 | 0.17996 | 0.026569 | 0.000168 | 94.2443 | 81.9361 | 85.1508 | 21.2612 | 0.008432 | 0.007121 | 229.9336 | 0.25912 | 8.4014 | 0.00523 | 0.006933 | 0.006892 | 6e-05 | 0.00117 | 0.31797 | 1.2979 | 2.5475 | 100.69 | 21.19 | 26.349 | 228.4421 | 148.5532 | 195.782 | 0.006592 | 0.75287 | 0.017599 | 23.3759 | 1.2518 | 0.15415 | 0.14451 | 7.537 | 7.537 | 0.000845
2019-03-28 15:03 | 1553799789.7438 | 151 | 0.18011 | 0.026117 | 0.000166 | 94.1791 | 80.8439 | 85.4285 | 21.2612 | 0.008424 | 0.007028 | 229.6368 | 0.25877 | 8.4024 | 0.005188 | 0.006098 | 0.005953 | 6.1e-05 | 0.00117 | 0.33483 | 1.3214 | 2.5474 | 100.69 | 21.19 | 26.349 | 228.2627 | 147.5204 | 193.452 | 0.006592 | 0.75663 | 0.017672 | 23.3558 | 1.3656 | 0.15394 | 0.14413 | 7.5489 | 7.5489 | 0.000842
2019-03-28 15:03 | 1553799790.7445 | 152 | 0.18005 | 0.025702 | 0.000164 | 94.0793 | 79.8471 | 85.4569 | 21.2612 | 0.008411 | 0.007089 | 229.381 | 0.25929 | 8.4019 | 0.005192 | 0.006634 | 0.006526 | 6e-05 | 0.001167 | 0.33035 | 1.2685 | 2.5474 | 100.69 | 21.19 | 26.352 | 227.9196 | 147.6194 | 194.9674 | 0.006593 | 0.75418 | 0.017646 | 23.3976 | 1.4671 | 0.15425 | 0.14496 | 7.5808 | 7.5808 | 0.00084
2019-03-28 15:03 | 1553799791.7436 | 153 | 0.18051 | 0.02534 | 0.000162 | 94.0037 | 79.3776 | 85.7031 | 21.2612 | 0.008349 | 0.007032 | 229.1101 | 0.25944 | 8.4027 | 0.005169 | 0.006972 | 0.00687 | 5.7e-05 | 0.001163 | 0.29734 | 1.3106 | 2.5474 | 100.69 | 21.19 | 26.352 | 226.3882 | 147.0327 | 193.5438 | 0.006593 | 0.75648 | 0.017698 | 23.3956 | 1.5127 | 0.15434 | 0.14263 | 7.4827 | 7.4827 | 0.000852
2019-03-28 15:03 | 1553799792.7436 | 154 | 0.181 | 0.025009 | 0.000161 | 93.8842 | 79.6355 | 85.8898 | 21.2612 | 0.008397 | 0.007022 | 228.8851 | 0.2597 | 8.4032 | 0.005261 | 0.007113 | 0.007009 | 5.7e-05 | 0.001162 | 0.33859 | 1.2817 | 2.5474 | 100.69 | 21.19 | 26.343 | 227.5842 | 149.325 | 193.3019 | 0.006591 | 0.75687 | 0.017722 | 23.415 | 1.4723 | 0.15449 | 0.14012 | 7.3628 | 7.3628 | 0.000867
2019-03-28 15:03 | 1553799793.7436 | 155 | 0.18133 | 0.024697 | 0.000161 | 93.8416 | 80.389 | 86.2705 | 21.2612 | 0.008343 | 0.007004 | 228.6275 | 0.2593 | 8.4018 | 0.00525 | 0.007338 | 0.007319 | 5.4e-05 | 0.001156 | 0.347 | 1.3162 | 2.5474 | 100.69 | 21.19 | 26.332 | 226.2449 | 149.0437 | 192.8466 | 0.006588 | 0.75761 | 0.01777 | 23.4554 | 1.384 | 0.15426 | 0.13857 | 7.302 | 7.302 | 0.000872
2019-03-28 15:03 | 1553799794.744 | 156 | 0.18228 | 0.024369 | 0.000163 | 93.7365 | 81.0977 | 86.8124 | 21.2612 | 0.008315 | 0.006951 | 228.4061 | 0.25876 | 8.4029 | 0.005209 | 0.006685 | 0.006588 | 6e-05 | 0.001163 | 0.34683 | 1.291 | 2.5474 | 100.69 | 21.19 | 26.332 | 225.5425 | 148.0355 | 191.5114 | 0.006588 | 0.75979 | 0.017851 | 23.4952 | 1.2955 | 0.15394 | 0.13336 | 7.0633 | 7.0633 | 0.0009
2019-03-28 15:03 | 1553799795.7437 | 157 | 0.18305 | 0.024015 | 0.000165 | 93.59 | 81.2438 | 87.3173 | 21.2612 | 0.008325 | 0.006917 | 228.1635 | 0.25842 | 8.3993 | 0.005221 | 0.006722 | 0.006658 | 6e-05 | 0.001171 | 0.3131 | 1.289 | 2.5475 | 100.69 | 21.19 | 26.311 | 225.801 | 148.3415 | 190.6597 | 0.006583 | 0.76119 | 0.01792 | 23.5418 | 1.2654 | 0.15373 | 0.12928 | 6.8758 | 6.8758 | 0.000923
2019-03-28 15:03 | 1553799796.7448 | 158 | 0.18382 | 0.023641 | 0.000169 | 93.402 | 81.7425 | 87.722 | 21.2612 | 0.008271 | 0.006926 | 227.9296 | 0.25869 | 8.4045 | 0.005228 | 0.006601 | 0.006561 | 5.8e-05 | 0.001171 | 0.3314 | 1.3377 | 2.5474 | 100.69 | 21.19 | 26.304 | 224.4515 | 148.5058 | 190.8927 | 0.006581 | 0.7608 | 0.017957 | 23.6022 | 1.1927 | 0.15389 | 0.12514 | 6.6721 | 6.6721 | 0.000953
2019-03-28 15:03 | 1553799797.7437 | 159 | 0.18436 | 0.023243 | 0.000174 | 93.2182 | 81.9979 | 87.7079 | 21.2612 | 0.008292 | 0.006888 | 227.6614 | 0.25811 | 8.4016 | 0.005175 | 0.00676 | 0.006716 | 5.9e-05 | 0.001176 | 0.33826 | 1.3057 | 2.5472 | 100.69 | 21.19 | 26.297 | 224.9808 | 147.1791 | 189.9396 | 0.006579 | 0.76237 | 0.017974 | 23.5761 | 1.1471 | 0.15355 | 0.1224 | 6.5338 | 6.5338 | 0.000971
2019-03-28 15:03 | 1553799798.7448 | 160 | 0.1849 | 0.022841 | 0.000179 | 93.0553 | 82.3088 | 88.2217 | 21.2612 | 0.008278 | 0.006913 | 227.4123 | 0.25873 | 8.4031 | 0.005192 | 0.007449 | 0.007327 | 5.7e-05 | 0.001179 | 0.34722 | 1.3045 | 2.5473 | 100.69 | 21.19 | 26.297 | 224.6159 | 147.6177 | 190.5639 | 0.006579 | 0.76134 | 0.018014 | 23.661 | 1.0976 | 0.15392 | 0.11967 | 6.4038 | 6.4038 | 0.000993
2019-03-28 15:03 | 1553799799.7438 | 161 | 0.1854 | 0.022432 | 0.000182 | 92.8569 | 82.5843 | 88.321 | 21.2612 | 0.008215 | 0.00685 | 227.2023 | 0.25815 | 8.4033 | 0.005152 | 0.006746 | 0.006638 | 5.8e-05 | 0.001178 | 0.30718 | 1.2823 | 2.5475 | 100.69 | 21.19 | 26.288 | 223.064 | 146.6101 | 188.9898 | 0.006577 | 0.76394 | 0.018055 | 23.6341 | 1.0487 | 0.15357 | 0.11721 | 6.2873 | 6.2873 | 0.001009
2019-03-28 15:03 | 1553799800.7453 | 162 | 0.18579 | 0.022027 | 0.000185 | 92.7512 | 82.779 | 88.5424 | 21.2612 | 0.008217 | 0.006864 | 226.9259 | 0.25797 | 8.403 | 0.005116 | 0.006514 | 0.006432 | 5.6e-05 | 0.001179 | 0.33268 | 1.2962 | 2.5474 | 100.69 | 21.19 | 26.287 | 223.1003 | 145.7309 | 189.3478 | 0.006577 | 0.76335 | 0.018071 | 23.6728 | 1.0174 | 0.15346 | 0.11532 | 6.1925 | 6.1925 | 0.001024
2019-03-28 15:03 | 1553799801.7439 | 163 | 0.18609 | 0.02143 | 0.000187 | 92.4084 | 82.6312 | 88.6407 | 21.2612 | 0.008178 | 0.006811 | 226.7065 | 0.25903 | 8.3991 | 0.005102 | 0.006571 | 0.006534 | 5.5e-05 | 0.001182 | 0.29087 | 1.2653 | 2.5474 | 100.69 | 21.19 | 26.288 | 222.1413 | 145.3722 | 188.0279 | 0.006577 | 0.76553 | 0.018106 | 23.6521 | 1.0003 | 0.1541 | 0.11424 | 6.1473 | 6.1473 | 0.001035
2019-03-28 15:03 | 1553799802.7443 | 164 | 0.18652 | 0.020984 | 0.000188 | 92.2321 | 81.9786 | 89.3538 | 21.2612 | 0.008137 | 0.006776 | 226.4724 | 0.25971 | 8.4008 | 0.005028 | 0.006634 | 0.006577 | 5.6e-05 | 0.001178 | 0.29966 | 1.2853 | 2.5473 | 100.69 | 21.19 | 26.293 | 221.1144 | 143.5483 | 187.1446 | 0.006578 | 0.767 | 0.018197 | 23.7243 | 1.0541 | 0.1545 | 0.11221 | 6.069 | 6.069 | 0.001051
2019-03-28 15:03 | 1553799803.7429 | 165 | 0.18696 | 0.020598 | 0.00019 | 92.0197 | 81.5397 | 89.7047 | 21.2612 | 0.008132 | 0.006767 | 226.238 | 0.26018 | 8.4008 | 0.004995 | 0.006231 | 0.006128 | 5.6e-05 | 0.00118 | 0.32709 | 1.317 | 2.5475 | 100.69 | 21.19 | 26.304 | 220.9893 | 142.7379 | 186.9033 | 0.006581 | 0.7674 | 0.018237 | 23.7646 | 1.0815 | 0.15478 | 0.11004 | 5.9658 | 5.9658 | 0.001072
2019-03-28 15:03 | 1553799804.743 | 166 | 0.1874 | 0.020206 | 0.00019 | 91.8043 | 81.1972 | 89.8056 | 21.2612 | 0.008115 | 0.006758 | 226.0068 | 0.25879 | 8.4034 | 0.004979 | 0.006291 | 0.006237 | 5.4e-05 | 0.001183 | 0.32892 | 1.3083 | 2.5473 | 100.69 | 21.19 | 26.31 | 220.564 | 142.333 | 186.6931 | 0.006582 | 0.76775 | 0.018251 | 23.7726 | 1.0982 | 0.15396 | 0.10793 | 5.8571 | 5.8571 | 0.001086
2019-03-28 15:03 | 1553799805.7433 | 167 | 0.1876 | 0.019838 | 0.000192 | 91.6429 | 81.1996 | 90.1936 | 21.2612 | 0.0081 | 0.00674 | 225.8013 | 0.25894 | 8.4042 | 0.00495 | 0.006376 | 0.006318 | 5.6e-05 | 0.001185 | 0.33052 | 1.3128 | 2.5474 | 100.69 | 21.19 | 26.31 | 220.2034 | 141.6282 | 186.2425 | 0.006582 | 0.7685 | 0.0183 | 23.8122 | 1.0822 | 0.15404 | 0.10716 | 5.831 | 5.831 | 0.001091
2019-03-28 15:03 | 1553799806.8107 | 168 | 0.18771 | 0.019482 | 0.000195 | 91.4758 | 80.8616 | 90.4056 | 21.2612 | 0.008024 | 0.006637 | 225.5874 | 0.25884 | 8.4032 | 0.004961 | 0.006183 | 0.00602 | 5.6e-05 | 0.00119 | 0.31238 | 1.2998 | 2.5474 | 100.69 | 21.19 | 26.323 | 218.303 | 141.8786 | 183.6713 | 0.006586 | 0.77283 | 0.018373 | 23.7733 | 1.1032 | 0.15398 | 0.10691 | 5.8407 | 5.8407 | 0.001089
2019-03-28 15:03 | 1553799807.7561 | 169 | 0.18805 | 0.019176 | 0.000197 | 91.3305 | 80.1206 | 90.7945 | 21.2612 | 0.008016 | 0.006665 | 225.4039 | 0.25861 | 8.4016 | 0.004948 | 0.006255 | 0.006078 | 5.2e-05 | 0.00119 | 0.31509 | 1.2853 | 2.5474 | 100.69 | 21.19 | 26.337 | 218.1116 | 141.571 | 184.3557 | 0.006589 | 0.77167 | 0.018398 | 23.8423 | 1.1713 | 0.15385 | 0.10524 | 5.7584 | 5.7584 | 0.001103
2019-03-28 15:03 | 1553799808.7561 | 170 | 0.18848 | 0.018907 | 0.000201 | 91.1657 | 80.4536 | 90.9422 | 21.2612 | 0.007923 | 0.006637 | 225.2035 | 0.25903 | 8.4031 | 0.004876 | 0.006284 | 0.006293 | 5.7e-05 | 0.001192 | 0.31741 | 1.2714 | 2.5473 | 100.69 | 21.19 | 26.337 | 215.7871 | 139.7939 | 183.6688 | 0.006589 | 0.77284 | 0.018427 | 23.8437 | 1.1181 | 0.1541 | 0.103 | 5.6456 | 5.6456 | 0.001127
2019-03-28 15:03 | 1553799809.7572 | 171 | 0.18891 | 0.018643 | 0.000204 | 90.8983 | 80.3938 | 91.4189 | 21.2612 | 0.007945 | 0.006577 | 224.9836 | 0.25815 | 8.4004 | 0.004889 | 0.00676 | 0.006695 | 5.7e-05 | 0.001194 | 0.29927 | 1.2845 | 2.5474 | 100.69 | 21.19 | 26.353 | 216.3415 | 140.1002 | 182.1717 | 0.006593 | 0.77538 | 0.018506 | 23.8669 | 1.0984 | 0.15358 | 0.10079 | 5.5491 | 5.5491 | 0.001143
2019-03-28 15:03 | 1553799810.7562 | 172 | 0.18928 | 0.018383 | 0.000207 | 90.6254 | 80.6598 | 91.6597 | 21.2612 | 0.00787 | 0.006577 | 224.8021 | 0.25821 | 8.4019 | 0.004856 | 0.006107 | 0.005974 | 5.5e-05 | 0.001194 | 0.34307 | 1.2869 | 2.5474 | 100.69 | 21.19 | 26.353 | 214.4744 | 139.292 | 182.1725 | 0.006593 | 0.77537 | 0.01853 | 23.8984 | 1.042 | 0.15361 | 0.098897 | 5.4526 | 5.4526 | 0.001164
2019-03-28 15:04 | 1553799811.7566 | 173 | 0.18971 | 0.018113 | 0.000208 | 90.4235 | 80.8694 | 92.1024 | 21.2612 | 0.007863 | 0.00652 | 224.6105 | 0.25883 | 8.4033 | 0.004718 | 0.005916 | 0.005856 | 5.8e-05 | 0.001192 | 0.30127 | 1.2963 | 2.5474 | 100.69 | 21.19 | 26.364 | 214.3039 | 135.8745 | 180.734 | 0.006596 | 0.77783 | 0.018604 | 23.9181 | 0.99883 | 0.15398 | 0.096636 | 5.3502 | 5.3502 | 0.001189
2019-03-28 15:04 | 1553799812.7577 | 174 | 0.18999 | 0.017837 | 0.000209 | 90.2344 | 80.8673 | 92.2102 | 21.2612 | 0.00784 | 0.006521 | 224.413 | 0.25885 | 8.4013 | 0.004686 | 0.005699 | 0.005619 | 5.7e-05 | 0.001185 | 0.30093 | 1.2738 | 2.5473 | 100.69 | 21.19 | 26.364 | 213.72 | 135.0847 | 180.769 | 0.006596 | 0.77777 | 0.018614 | 23.9331 | 0.98032 | 0.15399 | 0.095289 | 5.2791 | 5.2791 | 0.001205
2019-03-28 15:04 | 1553799813.757 | 175 | 0.19025 | 0.017551 | 0.00021 | 90.0157 | 80.4023 | 92.4137 | 21.2612 | 0.007801 | 0.006493 | 224.2135 | 0.25925 | 8.403 | 0.004632 | 0.006069 | 0.005978 | 5.7e-05 | 0.001181 | 0.31526 | 1.3143 | 2.5473 | 100.69 | 21.19 | 26.379 | 212.7511 | 133.7763 | 180.0472 | 0.0066 | 0.77901 | 0.01865 | 23.9404 | 1.0102 | 0.15423 | 0.094118 | 5.2245 | 5.2245 | 0.001219
2019-03-28 15:04 | 1553799814.7563 | 176 | 0.19033 | 0.017269 | 0.000214 | 89.8868 | 79.8628 | 92.7127 | 21.2612 | 0.007763 | 0.006478 | 224.0257 | 0.25847 | 8.4016 | 0.004613 | 0.005826 | 0.005662 | 5.6e-05 | 0.001179 | 0.32781 | 1.302 | 2.5474 | 100.69 | 21.19 | 26.383 | 211.8006 | 133.2974 | 179.6787 | 0.006601 | 0.77964 | 0.018688 | 23.9693 | 1.0576 | 0.15377 | 0.093977 | 5.2272 | 5.2272 | 0.001215
2019-03-28 15:04 | 1553799815.7574 | 177 | 0.19045 | 0.017002 | 0.000218 | 89.7918 | 79.6545 | 92.7831 | 21.2612 | 0.007743 | 0.00641 | 223.8139 | 0.25882 | 8.4012 | 0.00464 | 0.006386 | 0.006332 | 5.8e-05 | 0.001178 | 0.34595 | 1.2837 | 2.5474 | 100.69 | 21.19 | 26.39 | 211.3046 | 133.957 | 177.9801 | 0.006602 | 0.78258 | 0.01873 | 23.9334 | 1.0715 | 0.15397 | 0.093525 | 5.2138 | 5.2138 | 0.00122
2019-03-28 15:04 | 1553799816.7564 | 178 | 0.19076 | 0.016751 | 0.000223 | 89.7995 | 79.4582 | 92.9843 | 21.2612 | 0.007714 | 0.006451 | 223.6225 | 0.25926 | 8.4029 | 0.004694 | 0.00567 | 0.00566 | 5.5e-05 | 0.001175 | 0.32421 | 1.317 | 2.5473 | 100.69 | 21.19 | 26.393 | 210.5813 | 135.2998 | 178.9977 | 0.006603 | 0.78082 | 0.018729 | 23.9864 | 1.0944 | 0.15423 | 0.091975 | 5.1276 | 5.1276 | 0.001242
2019-03-28 15:04 | 1553799817.7574 | 179 | 0.19115 | 0.016517 | 0.000229 | 89.8129 | 79.5648 | 93.1943 | 21.2612 | 0.007717 | 0.006354 | 223.4281 | 0.25931 | 8.402 | 0.004603 | 0.00703 | 0.006849 | 5.4e-05 | 0.00117 | 0.32333 | 1.3069 | 2.5474 | 100.69 | 21.19 | 26.393 | 210.6492 | 133.0513 | 176.5932 | 0.006603 | 0.78499 | 0.0188 | 23.9495 | 1.0837 | 0.15426 | 0.089963 | 5.035 | 5.035 | 0.001265
2019-03-28 15:04 | 1553799818.7565 | 180 | 0.1916 | 0.016292 | 0.000236 | 89.9728 | 80.0058 | 93.5492 | 21.2612 | 0.007692 | 0.0064 | 223.2525 | 0.25908 | 8.4022 | 0.004578 | 0.006376 | 0.006366 | 5.6e-05 | 0.001171 | 0.31669 | 1.3283 | 2.5474 | 100.69 | 21.19 | 26.404 | 210.0365 | 132.4457 | 177.7285 | 0.006606 | 0.78302 | 0.018812 | 24.0253 | 1.0502 | 0.15413 | 0.087546 | 4.9035 | 4.9035 | 0.001298
2019-03-28 15:04 | 1553799819.7565 | 181 | 0.19196 | 0.016067 | 0.000242 | 90.2403 | 80.4573 | 93.7072 | 21.2612 | 0.007619 | 0.006335 | 223.0791 | 0.25871 | 8.4026 | 0.004448 | 0.005602 | 0.00549 | 5.3e-05 | 0.001174 | 0.31797 | 1.3 | 2.5474 | 100.69 | 21.19 | 26.404 | 208.2174 | 129.2313 | 176.0942 | 0.006606 | 0.78587 | 0.018862 | 24.002 | 1.0264 | 0.1539 | 0.085636 | 4.8097 | 4.8097 | 0.001322
2019-03-28 15:04 | 1553799820.7559 | 182 | 0.19237 | 0.015841 | 0.000247 | 90.4381 | 81.0593 | 94.0616 | 21.2612 | 0.007615 | 0.006308 | 222.9067 | 0.25867 | 8.3998 | 0.004373 | 0.006581 | 0.006433 | 5.5e-05 | 0.001173 | 0.33931 | 1.2809 | 2.5473 | 100.69 | 21.2 | 26.415 | 208.1251 | 127.3854 | 175.4425 | 0.006613 | 0.78701 | 0.018912 | 24.0299 | 0.97929 | 0.15389 | 0.083523 | 4.7038 | 4.7038 | 0.001351
2019-03-28 15:04 | 1553799821.7555 | 183 | 0.19248 | 0.015607 | 0.000252 | 90.7356 | 81.7299 | 94.4715 | 21.2612 | 0.007563 | 0.006322 | 222.7348 | 0.25821 | 8.4046 | 0.004454 | 0.005857 | 0.005766 | 5.6e-05 | 0.001176 | 0.31509 | 1.3163 | 2.5472 | 100.69 | 21.2 | 26.418 | 206.8106 | 129.3869 | 175.7753 | 0.006614 | 0.78642 | 0.018946 | 24.0911 | 0.93498 | 0.15361 | 0.08309 | 4.6878 | 4.6878 | 0.001353
2019-03-28 15:04 | 1553799822.7566 | 184 | 0.19281 | 0.015361 | 0.000256 | 91.0579 | 82.6475 | 94.5682 | 21.2612 | 0.00753 | 0.00624 | 222.5759 | 0.25822 | 8.4035 | 0.004399 | 0.006137 | 0.006077 | 5.4e-05 | 0.001175 | 0.34722 | 1.2994 | 2.5473 | 100.69 | 21.2 | 26.425 | 205.9852 | 128.0403 | 173.7226 | 0.006615 | 0.79004 | 0.018999 | 24.0483 | 0.86683 | 0.15362 | 0.081436 | 4.6077 | 4.6077 | 0.001377
2019-03-28 15:04 | 1553799823.756 | 185 | 0.19301 | 0.015115 | 0.00026 | 91.3436 | 83.0885 | 95.0214 | 21.2612 | 0.00752 | 0.006252 | 222.4331 | 0.25897 | 8.4018 | 0.004344 | 0.005813 | 0.005729 | 5.2e-05 | 0.001173 | 0.33373 | 1.292 | 2.5473 | 100.69 | 21.2 | 26.429 | 205.7325 | 126.6933 | 174.0414 | 0.006616 | 0.78947 | 0.019038 | 24.1144 | 0.84724 | 0.15406 | 0.080516 | 4.565 | 4.565 | 0.001394
2019-03-28 15:04 | 1553799824.7566 | 186 | 0.19306 | 0.014858 | 0.000264 | 91.5851 | 83.4455 | 95.3472 | 21.2612 | 0.007452 | 0.006187 | 222.2606 | 0.25867 | 8.4032 | 0.004408 | 0.00573 | 0.005605 | 5.4e-05 | 0.001165 | 0.31885 | 1.2646 | 2.5472 | 100.69 | 21.2 | 26.429 | 204.0461 | 128.2555 | 172.4177 | 0.006616 | 0.79235 | 0.019105 | 24.1118 | 0.83251 | 0.15389 | 0.080538 | 4.5823 | 4.5823 | 0.001387
2019-03-28 15:04 | 1553799825.7567 | 187 | 0.19347 | 0.014615 | 0.000267 | 91.8547 | 83.5247 | 95.7456 | 21.2612 | 0.007408 | 0.006159 | 222.1262 | 0.25972 | 8.4009 | 0.004391 | 0.006099 | 0.005978 | 5.3e-05 | 0.001159 | 0.32477 | 1.3022 | 2.5472 | 100.69 | 21.2 | 26.434 | 202.9406 | 127.8316 | 171.7063 | 0.006618 | 0.79362 | 0.01916 | 24.1429 | 0.85032 | 0.15451 | 0.078379 | 4.4728 | 4.4728 | 0.001427
2019-03-28 15:04 | 1553799826.7568 | 188 | 0.19383 | 0.014377 | 0.000272 | 92.3158 | 83.6032 | 96.0432 | 21.2612 | 0.007354 | 0.006142 | 221.9699 | 0.25814 | 8.4033 | 0.004396 | 0.005919 | 0.005887 | 5.4e-05 | 0.001161 | 0.31542 | 1.3247 | 2.5474 | 100.69 | 21.2 | 26.436 | 201.6013 | 127.9679 | 171.2799 | 0.006618 | 0.79438 | 0.019199 | 24.1688 | 0.8867 | 0.15356 | 0.076522 | 4.376 | 4.376 | 0.001449
2019-03-28 15:04 | 1553799827.7567 | 189 | 0.19417 | 0.014147 | 0.000276 | 92.5698 | 84.2198 | 96.5406 | 21.2612 | 0.007332 | 0.006131 | 221.8469 | 0.25825 | 8.4044 | 0.004319 | 0.006172 | 0.006095 | 5.6e-05 | 0.001162 | 0.31669 | 1.3064 | 2.5474 | 100.69 | 21.2 | 26.438 | 201.049 | 126.0683 | 171.0075 | 0.006619 | 0.79487 | 0.019255 | 24.2238 | 0.84556 | 0.15363 | 0.074778 | 4.289 | 4.289 | 0.001479
2019-03-28 15:04 | 1553799828.7561 | 190 | 0.1944 | 0.013926 | 0.00028 | 92.845 | 84.6794 | 96.6361 | 21.2612 | 0.007309 | 0.006107 | 221.6953 | 0.25959 | 8.4036 | 0.004331 | 0.005989 | 0.005837 | 6e-05 | 0.001168 | 0.30021 | 1.3366 | 2.5474 | 100.69 | 21.2 | 26.438 | 200.4799 | 126.3667 | 170.413 | 0.006619 | 0.79593 | 0.019277 | 24.2196 | 0.82343 | 0.15443 | 0.073671 | 4.2305 | 4.2305 | 0.001508
2019-03-28 15:04 | 1553799829.7558 | 191 | 0.1948 | 0.013692 | 0.000286 | 93.237 | 85.3599 | 97.0351 | 21.2612 | 0.007288 | 0.006034 | 221.5707 | 0.25842 | 8.4041 | 0.004308 | 0.005318 | 0.005256 | 5.2e-05 | 0.001167 | 0.3011 | 1.3705 | 2.5473 | 100.69 | 21.2 | 26.441 | 199.9462 | 125.8031 | 168.5827 | 0.006619 | 0.79923 | 0.019357 | 24.2194 | 0.78952 | 0.15374 | 0.071544 | 4.1255 | 4.1255 | 0.001539
2019-03-28 15:04 | 1553799830.7562 | 192 | 0.19511 | 0.01345 | 0.000294 | 93.4869 | 86.0169 | 97.3328 | 21.2612 | 0.007267 | 0.006003 | 221.4053 | 0.25884 | 8.4017 | 0.004221 | 0.005873 | 0.005734 | 5.5e-05 | 0.001168 | 0.31492 | 1.3715 | 2.5475 | 100.69 | 21.2 | 26.436 | 199.4125 | 123.6764 | 167.7991 | 0.006618 | 0.80065 | 0.019404 | 24.235 | 0.74488 | 0.15398 | 0.069966 | 4.0443 | 4.0443 | 0.001572
2019-03-28 15:04 | 1553799831.7561 | 193 | 0.19539 | 0.013203 | 0.000302 | 93.853 | 86.4524 | 97.4701 | 21.2612 | 0.007201 | 0.00599 | 221.2581 | 0.2591 | 8.4023 | 0.004246 | 0.005596 | 0.005456 | 5.7e-05 | 0.001169 | 0.34578 | 1.3207 | 2.5473 | 100.69 | 21.2 | 26.431 | 197.7714 | 124.291 | 167.488 | 0.006617 | 0.80121 | 0.019424 | 24.2435 | 0.73466 | 0.15414 | 0.068575 | 3.968 | 3.968 | 0.001604
2019-03-28 15:04 | 1553799832.7569 | 194 | 0.19575 | 0.012961 | 0.00031 | 94.0731 | 86.9814 | 97.7264 | 21.2612 | 0.007155 | 0.00597 | 221.097 | 0.25878 | 8.4047 | 0.004251 | 0.006012 | 0.00597 | 5.4e-05 | 0.001172 | 0.31669 | 1.3089 | 2.5474 | 100.69 | 21.2 | 26.431 | 196.6331 | 124.4062 | 166.984 | 0.006617 | 0.80213 | 0.019461 | 24.2615 | 0.70105 | 0.15395 | 0.06674 | 3.8691 | 3.8691 | 0.001643
2019-03-28 15:04 | 1553799833.7563 | 195 | 0.19617 | 0.012714 | 0.000317 | 94.3605 | 87.8864 | 97.9736 | 21.2612 | 0.007138 | 0.005941 | 220.9439 | 0.25867 | 8.4003 | 0.004174 | 0.005778 | 0.005652 | 5.1e-05 | 0.00117 | 0.32875 | 1.2996 | 2.5473 | 100.69 | 21.2 | 26.416 | 196.2032 | 122.5056 | 166.2712 | 0.006613 | 0.80343 | 0.019501 | 24.2725 | 0.63575 | 0.15389 | 0.06455 | 3.7499 | 3.7499 | 0.001695
2019-03-28 15:04 | 1553799834.7559 | 196 | 0.19651 | 0.012472 | 0.000323 | 94.6756 | 88.2458 | 98.0254 | 21.2612 | 0.007076 | 0.005894 | 220.8107 | 0.25933 | 8.4017 | 0.004236 | 0.005352 | 0.005297 | 5.5e-05 | 0.001174 | 0.34274 | 1.2873 | 2.5474 | 100.69 | 21.2 | 26.416 | 194.646 | 124.0268 | 165.0887 | 0.006613 | 0.8056 | 0.019533 | 24.2462 | 0.62907 | 0.15427 | 0.062783 | 3.6531 | 3.6531 | 0.001744
2019-03-28 15:04 | 1553799835.755 | 197 | 0.19678 | 0.012212 | 0.000329 | 94.9985 | 88.8753 | 98.3529 | 21.2612 | 0.00706 | 0.005894 | 220.667 | 0.25933 | 8.4023 | 0.004205 | 0.005734 | 0.005643 | 5.6e-05 | 0.001174 | 0.31166 | 1.3005 | 2.5474 | 100.69 | 21.2 | 26.401 | 194.258 | 123.2717 | 165.0966 | 0.006609 | 0.80559 | 0.019565 | 24.2869 | 0.59594 | 0.15428 | 0.061484 | 3.5833 | 3.5833 | 0.001778
2019-03-28 15:04 | 1553799836.755 | 198 | 0.19718 | 0.01195 | 0.000334 | 95.2681 | 89.5208 | 98.8235 | 21.2612 | 0.007037 | 0.005812 | 220.5384 | 0.2589 | 8.402 | 0.004096 | 0.00615 | 0.00597 | 5.3e-05 | 0.001173 | 0.2896 | 1.3215 | 2.5475 | 100.69 | 21.2 | 26.397 | 193.6712 | 120.6138 | 163.033 | 0.006608 | 0.8094 | 0.019658 | 24.2876 | 0.55656 | 0.15402 | 0.059384 | 3.4774 | 3.4774 | 0.001829
2019-03-28 15:04 | 1553799837.7561 | 199 | 0.19746 | 0.011695 | 0.000338 | 95.5323 | 89.8667 | 98.8913 | 21.2612 | 0.007013 | 0.005773 | 220.3812 | 0.25925 | 8.4051 | 0.00417 | 0.00547 | 0.005371 | 5.6e-05 | 0.001177 | 0.30215 | 1.3168 | 2.5475 | 100.69 | 21.2 | 26.392 | 193.0795 | 122.416 | 162.0627 | 0.006607 | 0.8112 | 0.019687 | 24.2689 | 0.54684 | 0.15423 | 0.058017 | 3.4022 | 3.4022 | 0.001872
2019-03-28 15:04 | 1553799838.7554 | 200 | 0.19772 | 0.011433 | 0.000341 | 95.8691 | 90.6779 | 99.2319 | 21.2612 | 0.00692 | 0.005757 | 220.2716 | 0.2592 | 8.4042 | 0.004169 | 0.005598 | 0.005484 | 5.6e-05 | 0.001184 | 0.35153 | 1.3068 | 2.5473 | 100.69 | 21.19 | 26.383 | 190.7368 | 122.3859 | 161.6654 | 0.006601 | 0.81194 | 0.01973 | 24.2995 | 0.49794 | 0.15419 | 0.056753 | 3.3353 | 3.3353 | 0.001909
2019-03-28 15:04 | 1553799839.7566 | 201 | 0.19804 | 0.011168 | 0.000345 | 96.091 | 90.8288 | 99.6129 | 21.2612 | 0.006837 | 0.005742 | 220.1144 | 0.25925 | 8.4028 | 0.004116 | 0.005226 | 0.005181 | 5.1e-05 | 0.001188 | 0.35081 | 1.2958 | 2.5473 | 100.69 | 21.2 | 26.378 | 188.657 | 121.0954 | 161.2892 | 0.006604 | 0.81265 | 0.019776 | 24.3356 | 0.50375 | 0.15423 | 0.055158 | 3.249 | 3.249 | 0.00196
2019-03-28 15:04 | 1553799840.7552 | 202 | 0.19846 | 0.010913 | 0.000346 | 96.2718 | 90.966 | 99.7941 | 21.2612 | 0.006829 | 0.005704 | 220.0323 | 0.25891 | 8.4016 | 0.004061 | 0.005698 | 0.005606 | 5.7e-05 | 0.001189 | 0.31653 | 1.3356 | 2.5475 | 100.69 | 21.2 | 26.371 | 188.462 | 119.7404 | 160.347 | 0.006602 | 0.81441 | 0.019816 | 24.3313 | 0.50706 | 0.15402 | 0.052999 | 3.128 | 3.128 | 0.002034
2019-03-28 15:04 | 1553799841.7552 | 203 | 0.19874 | 0.010655 | 0.000349 | 96.595 | 91.2792 | 100.1624 | 21.2612 | 0.006803 | 0.005642 | 219.9104 | 0.25903 | 8.4057 | 0.004067 | 0.005425 | 0.005347 | 5.4e-05 | 0.001189 | 0.30718 | 1.3098 | 2.5473 | 100.69 | 21.2 | 26.371 | 187.8119 | 119.8854 | 158.8124 | 0.006602 | 0.81731 | 0.019887 | 24.3329 | 0.50629 | 0.1541 | 0.051621 | 3.0576 | 3.0576 | 0.002081
2019-03-28 15:04 | 1553799842.7552 | 204 | 0.19884 | 0.010406 | 0.000352 | 96.8172 | 92.1836 | 100.7567 | 21.2612 | 0.006753 | 0.005619 | 219.8195 | 0.25925 | 8.4021 | 0.003995 | 0.005477 | 0.005404 | 5.5e-05 | 0.001199 | 0.31509 | 1.3776 | 2.5475 | 100.69 | 21.2 | 26.37 | 186.5688 | 118.1342 | 158.2384 | 0.006602 | 0.81839 | 0.01996 | 24.3888 | 0.43866 | 0.15423 | 0.051292 | 3.049 | 3.049 | 0.002089
2019-03-28 15:04 | 1553799843.7562 | 205 | 0.19916 | 0.010146 | 0.000355 | 97.0864 | 93.0264 | 101.1696 | 21.2612 | 0.006676 | 0.005543 | 219.7208 | 0.25888 | 8.4013 | 0.004048 | 0.00554 | 0.005445 | 5.6e-05 | 0.001195 | 0.32621 | 1.329 | 2.5474 | 100.69 | 21.2 | 26.373 | 184.6342 | 119.4334 | 156.3492 | 0.006602 | 0.82199 | 0.020044 | 24.3852 | 0.38209 | 0.15401 | 0.049711 | 2.9673 | 2.9673 | 0.002144
2019-03-28 15:04 | 1553799844.7557 | 206 | 0.1995 | 0.009891 | 0.000359 | 97.258 | 93.5398 | 101.5906 | 21.2612 | 0.006687 | 0.005582 | 219.6231 | 0.25884 | 8.4011 | 0.003977 | 0.005124 | 0.005067 | 5.5e-05 | 0.001192 | 0.29247 | 1.3503 | 2.5473 | 100.69 | 21.2 | 26.374 | 184.911 | 117.6866 | 157.318 | 0.006603 | 0.82014 | 0.020063 | 24.4634 | 0.34866 | 0.15398 | 0.047969 | 2.8658 | 2.8658 | 0.002219
2019-03-28 15:04 | 1553799845.7553 | 207 | 0.19976 | 0.009631 | 0.000362 | 97.535 | 94.1995 | 102.0343 | 21.2612 | 0.006568 | 0.005571 | 219.5391 | 0.25814 | 8.4029 | 0.003919 | 0.005317 | 0.00525 | 5.4e-05 | 0.001187 | 0.32875 | 1.3317 | 2.5473 | 100.69 | 21.2 | 26.374 | 181.9257 | 116.2867 | 157.0312 | 0.006603 | 0.82069 | 0.020114 | 24.5086 | 0.31124 | 0.15356 | 0.046706 | 2.7971 | 2.7971 | 0.002267
2019-03-28 15:04 | 1553799846.7557 | 208 | 0.20029 | 0.009384 | 0.000365 | 97.7468 | 94.5009 | 102.3737 | 21.2612 | 0.006591 | 0.005531 | 219.4326 | 0.25899 | 8.405 | 0.003869 | 0.004924 | 0.004778 | 5.7e-05 | 0.001192 | 0.30574 | 1.3191 | 2.5474 | 100.69 | 21.2 | 26.373 | 182.5055 | 115.0525 | 156.0472 | 0.006602 | 0.82257 | 0.02017 | 24.5212 | 0.30207 | 0.15407 | 0.043857 | 2.6334 | 2.6334 | 0.002416
2019-03-28 15:04 | 1553799847.7554 | 209 | 0.20066 | 0.009141 | 0.00037 | 98.0101 | 95.165 | 102.5978 | 21.2612 | 0.006581 | 0.005468 | 219.3578 | 0.25925 | 8.3991 | 0.003913 | 0.005491 | 0.005421 | 5.4e-05 | 0.001189 | 0.30486 | 1.2821 | 2.5473 | 100.69 | 21.2 | 26.373 | 182.2498 | 116.1333 | 154.4655 | 0.006602 | 0.82561 | 0.02023 | 24.5028 | 0.26349 | 0.15423 | 0.04192 | 2.524 | 2.524 | 0.002524
2019-03-28 15:04 | 1553799848.7548 | 210 | 0.20093 | 0.008888 | 0.000374 | 98.1513 | 95.6542 | 102.8673 | 21.2612 | 0.006493 | 0.005404 | 219.238 | 0.25876 | 8.4022 | 0.003846 | 0.005091 | 0.004932 | 5.4e-05 | 0.001192 | 0.33339 | 1.2808 | 2.5474 | 100.69 | 21.21 | 26.371 | 180.0652 | 114.4926 | 152.8822 | 0.006606 | 0.82868 | 0.020294 | 24.4895 | 0.2305 | 0.15394 | 0.040596 | 2.4515 | 2.4515 | 0.002593
2019-03-28 15:04 | 1553799849.7545 | 211 | 0.20104 | 0.008641 | 0.000376 | 98.4158 | 96.1234 | 103.3025 | 21.2612 | 0.006454 | 0.0054 | 219.1875 | 0.25885 | 8.4037 | 0.003869 | 0.005316 | 0.005251 | 5.7e-05 | 0.001196 | 0.31957 | 1.3138 | 2.5473 | 100.69 | 21.21 | 26.371 | 179.0733 | 115.0515 | 152.7912 | 0.006606 | 0.82886 | 0.020339 | 24.5386 | 0.21081 | 0.15399 | 0.040254 | 2.4361 | 2.4361 | 0.002611
2019-03-28 15:04 | 1553799850.7555 | 212 | 0.20126 | 0.008389 | 0.000378 | 98.6019 | 96.5745 | 103.9541 | 21.2612 | 0.00641 | 0.005334 | 219.0902 | 0.25856 | 8.4053 | 0.003813 | 0.005841 | 0.005818 | 5.3e-05 | 0.0012 | 0.3027 | 1.3455 | 2.5474 | 100.69 | 21.21 | 26.368 | 177.9946 | 113.6925 | 151.1453 | 0.006605 | 0.83208 | 0.020443 | 24.5683 | 0.18583 | 0.15382 | 0.039244 | 2.3867 | 2.3867 | 0.002662
2019-03-28 15:04 | 1553799851.7555 | 213 | 0.2015 | 0.008145 | 0.000378 | 98.7931 | 96.704 | 104.5022 | 21.2612 | 0.006356 | 0.005301 | 219.0026 | 0.25859 | 8.4067 | 0.003808 | 0.005408 | 0.005363 | 5.3e-05 | 0.001196 | 0.32084 | 1.3927 | 2.5475 | 100.69 | 21.2 | 26.363 | 176.6279 | 113.5833 | 150.3151 | 0.0066 | 0.83371 | 0.020517 | 24.6088 | 0.19118 | 0.15383 | 0.038103 | 2.3254 | 2.3254 | 0.002732
2019-03-28 15:04 | 1553799852.7558 | 214 | 0.20182 | 0.007909 | 0.000378 | 98.9969 | 97.217 | 104.808 | 21.2612 | 0.006394 | 0.005282 | 218.8822 | 0.25886 | 8.4025 | 0.003742 | 0.005298 | 0.005235 | 5.5e-05 | 0.001198 | 0.31868 | 1.3626 | 2.5474 | 100.69 | 21.2 | 26.361 | 177.5856 | 111.965 | 149.8454 | 0.006599 | 0.83463 | 0.020558 | 24.6311 | 0.16228 | 0.154 | 0.03649 | 2.231 | 2.231 | 0.002851
2019-03-28 15:04 | 1553799853.7546 | 215 | 0.20207 | 0.007677 | 0.000379 | 99.0672 | 97.4208 | 105.0974 | 21.2612 | 0.006257 | 0.005255 | 218.8266 | 0.25839 | 8.4022 | 0.003778 | 0.005132 | 0.004986 | 5.3e-05 | 0.001193 | 0.32781 | 1.3656 | 2.5475 | 100.69 | 21.2 | 26.354 | 174.1452 | 112.8511 | 149.1678 | 0.006597 | 0.83597 | 0.020603 | 24.6454 | 0.1499 | 0.15372 | 0.035279 | 2.1612 | 2.1612 | 0.002937
2019-03-28 15:04 | 1553799854.7556 | 216 | 0.20236 | 0.007454 | 0.000379 | 99.3123 | 97.5071 | 105.6971 | 21.2612 | 0.006202 | 0.005224 | 218.7467 | 0.25899 | 8.4015 | 0.003733 | 0.005088 | 0.005113 | 5.7e-05 | 0.001197 | 0.30613 | 1.3337 | 2.5474 | 100.69 | 21.2 | 26.354 | 172.7784 | 111.7502 | 148.4115 | 0.006597 | 0.83747 | 0.02068 | 24.6934 | 0.16408 | 0.15407 | 0.033811 | 2.0786 | 2.0786 | 0.003061
2019-03-28 15:04 | 1553799855.7558 | 217 | 0.20263 | 0.007241 | 0.000381 | 99.3942 | 97.7011 | 105.8945 | 21.2612 | 0.006235 | 0.00518 | 218.6731 | 0.25862 | 8.4034 | 0.00373 | 0.005951 | 0.005785 | 5.6e-05 | 0.001196 | 0.33395 | 1.3006 | 2.5474 | 100.69 | 21.2 | 26.347 | 173.6034 | 111.6879 | 147.3192 | 0.006596 | 0.83965 | 0.020726 | 24.6844 | 0.15368 | 0.15385 | 0.032422 | 1.9971 | 1.9971 | 0.003182
2019-03-28 15:04 | 1553799856.7547 | 218 | 0.20279 | 0.007031 | 0.000382 | 99.5387 | 98.1882 | 106.0802 | 21.2612 | 0.006205 | 0.005114 | 218.6091 | 0.25782 | 8.4022 | 0.003661 | 0.004481 | 0.004513 | 5.6e-05 | 0.0012 | 0.33915 | 1.3191 | 2.5473 | 100.69 | 21.2 | 26.347 | 172.8539 | 109.9945 | 145.6826 | 0.006596 | 0.84293 | 0.020785 | 24.6579 | 0.12218 | 0.15338 | 0.031732 | 1.9597 | 1.9597 | 0.003232
2019-03-28 15:04 | 1553799857.7557 | 219 | 0.20306 | 0.006822 | 0.000382 | 99.6721 | 98.249 | 106.5407 | 21.2612 | 0.006048 | 0.005165 | 218.5322 | 0.2591 | 8.4034 | 0.003658 | 0.005322 | 0.005215 | 5.5e-05 | 0.0012 | 0.32908 | 1.3523 | 2.5475 | 100.69 | 21.2 | 26.337 | 168.9394 | 109.9267 | 146.9546 | 0.006593 | 0.84037 | 0.020798 | 24.7489 | 0.12862 | 0.15414 | 0.030395 | 1.8778 | 1.8778 | 0.00339
2019-03-28 15:04 | 1553799858.7561 | 220 | 0.20325 | 0.006617 | 0.000383 | 99.86 | 98.4199 | 106.9123 | 21.2612 | 0.005997 | 0.005048 | 218.4851 | 0.25872 | 8.4025 | 0.003598 | 0.004968 | 0.004888 | 5.6e-05 | 0.001202 | 0.30862 | 1.3881 | 2.5475 | 100.69 | 21.2 | 26.337 | 167.6703 | 108.4613 | 144.0459 | 0.006593 | 0.84623 | 0.020907 | 24.7061 | 0.12993 | 0.15391 | 0.029491 | 1.831 | 1.831 | 0.003472
2019-03-28 15:04 | 1553799859.7557 | 221 | 0.20342 | 0.006416 | 0.000385 | 99.9473 | 98.7875 | 106.9653 | 21.2612 | 0.006013 | 0.004999 | 218.3778 | 0.25867 | 8.402 | 0.003578 | 0.00553 | 0.00545 | 5.5e-05 | 0.001202 | 0.30879 | 1.328 | 2.5473 | 100.69 | 21.2 | 26.329 | 168.0537 | 107.9735 | 142.827 | 0.006591 | 0.84871 | 0.020943 | 24.6761 | 0.1044 | 0.15388 | 0.02871 | 1.7851 | 1.7851 | 0.00356
2019-03-28 15:04 | 1553799860.7548 | 222 | 0.20366 | 0.006214 | 0.000386 | 100.0081 | 99.0827 | 107.2947 | 21.2612 | 0.005916 | 0.005019 | 218.3177 | 0.25799 | 8.4043 | 0.003574 | 0.005368 | 0.005297 | 5.6e-05 | 0.0012 | 0.35817 | 1.3066 | 2.5473 | 100.69 | 21.2 | 26.329 | 165.6263 | 107.8948 | 143.3202 | 0.006591 | 0.84771 | 0.020963 | 24.7287 | 0.083146 | 0.15348 | 0.027515 | 1.7117 | 1.7117 | 0.003703
2019-03-28 15:04 | 1553799861.7557 | 223 | 0.20389 | 0.006021 | 0.000388 | 100.2121 | 99.2166 | 107.2679 | 21.2612 | 0.005912 | 0.00501 | 218.2436 | 0.2591 | 8.4041 | 0.003462 | 0.004876 | 0.004787 | 5.8e-05 | 0.001199 | 0.34235 | 1.3022 | 2.5473 | 100.69 | 21.2 | 26.319 | 165.5469 | 105.175 | 143.0904 | 0.006589 | 0.84818 | 0.020966 | 24.7188 | 0.0893 | 0.15414 | 0.026332 | 1.6376 | 1.6376 | 0.003887
2019-03-28 15:04 | 1553799862.7552 | 224 | 0.20402 | 0.005834 | 0.000389 | 100.199 | 99.346 | 107.3529 | 21.2612 | 0.00586 | 0.004935 | 218.2133 | 0.25892 | 8.4064 | 0.00356 | 0.005144 | 0.005069 | 5.7e-05 | 0.001197 | 0.33179 | 1.2763 | 2.5473 | 100.69 | 21.2 | 26.318 | 164.2286 | 107.5491 | 141.2581 | 0.006588 | 0.85193 | 0.02102 | 24.6741 | 0.07647 | 0.15403 | 0.025812 | 1.6091 | 1.6091 | 0.003953
2019-03-28 15:04 | 1553799863.7549 | 225 | 0.20409 | 0.005656 | 0.000388 | 100.3298 | 99.4442 | 107.8003 | 21.2612 | 0.005841 | 0.004943 | 218.1246 | 0.2587 | 8.3994 | 0.003534 | 0.005222 | 0.005162 | 5.5e-05 | 0.001198 | 0.34324 | 1.2798 | 2.5473 | 100.69 | 21.2 | 26.316 | 163.7608 | 106.926 | 141.4459 | 0.006588 | 0.85154 | 0.021059 | 24.731 | 0.079299 | 0.1539 | 0.025602 | 1.5988 | 1.5988 | 0.003976
2019-03-28 15:04 | 1553799864.7549 | 226 | 0.20424 | 0.005489 | 0.000387 | 100.3479 | 99.4483 | 107.6519 | 21.2612 | 0.005759 | 0.004889 | 218.0811 | 0.25906 | 8.4012 | 0.003493 | 0.00448 | 0.004416 | 5.8e-05 | 0.001197 | 0.35402 | 1.309 | 2.5473 | 100.69 | 21.2 | 26.313 | 161.726 | 105.9114 | 140.1131 | 0.006587 | 0.85429 | 0.021079 | 24.6742 | 0.080552 | 0.15412 | 0.024903 | 1.5562 | 1.5562 | 0.00409
2019-03-28 15:04 | 1553799865.7549 | 227 | 0.20435 | 0.005328 | 0.000385 | 100.4808 | 99.7248 | 107.8477 | 21.2612 | 0.005762 | 0.004891 | 218.0015 | 0.25931 | 8.4002 | 0.003469 | 0.005084 | 0.00492 | 5.7e-05 | 0.001195 | 0.34274 | 1.3409 | 2.5473 | 100.69 | 21.2 | 26.312 | 161.7836 | 105.3466 | 140.1499 | 0.006587 | 0.85421 | 0.021097 | 24.6978 | 0.067547 | 0.15426 | 0.024451 | 1.5291 | 1.5291 | 0.004166
2019-03-28 15:04 | 1553799866.754 | 228 | 0.20454 | 0.005173 | 0.000382 | 100.4444 | 99.6474 | 108.2246 | 21.2612 | 0.005749 | 0.004828 | 217.9543 | 0.2581 | 8.4033 | 0.003581 | 0.00522 | 0.005204 | 5.5e-05 | 0.00119 | 0.33627 | 1.3444 | 2.5475 | 100.69 | 21.2 | 26.31 | 161.4777 | 108.0672 | 138.5984 | 0.006586 | 0.85743 | 0.021174 | 24.6944 | 0.071262 | 0.15354 | 0.023529 | 1.4763 | 1.4763 | 0.004295
2019-03-28 15:04 | 1553799867.755 | 229 | 0.20468 | 0.005023 | 0.000378 | 100.5761 | 99.8189 | 108.4293 | 21.2612 | 0.00572 | 0.004793 | 217.8998 | 0.25813 | 8.4037 | 0.003479 | 0.004995 | 0.004905 | 5.6e-05 | 0.001187 | 0.31996 | 1.3253 | 2.5473 | 100.69 | 21.2 | 26.31 | 160.7399 | 105.5802 | 137.7354 | 0.006586 | 0.85923 | 0.021216 | 24.6918 | 0.067592 | 0.15356 | 0.022876 | 1.438 | 1.438 | 0.00441
2019-03-28 15:04 | 1553799868.755 | 230 | 0.2049 | 0.004888 | 0.000373 | 100.6339 | 99.9378 | 108.7476 | 21.2612 | 0.005632 | 0.004757 | 217.8529 | 0.25953 | 8.4017 | 0.003446 | 0.005265 | 0.005201 | 5.9e-05 | 0.001185 | 0.34003 | 1.298 | 2.5473 | 100.69 | 21.2 | 26.31 | 158.5438 | 104.7747 | 136.8472 | 0.006586 | 0.86109 | 0.02127 | 24.7013 | 0.06208 | 0.15439 | 0.021707 | 1.3674 | 1.3674 | 0.004663
2019-03-28 15:04 | 1553799869.7541 | 231 | 0.20506 | 0.00475 | 0.000368 | 100.6973 | 99.8023 | 109.1445 | 21.2612 | 0.005583 | 0.004724 | 217.8375 | 0.25887 | 8.4026 | 0.00339 | 0.005449 | 0.005313 | 5.8e-05 | 0.001184 | 0.32117 | 1.295 | 2.5474 | 100.69 | 21.2 | 26.311 | 157.3278 | 103.4343 | 136.0232 | 0.006587 | 0.86283 | 0.02133 | 24.7215 | 0.079858 | 0.154 | 0.020892 | 1.3194 | 1.3194 | 0.004821
2019-03-28 15:04 | 1553799870.7551 | 232 | 0.20513 | 0.004619 | 0.000361 | 100.7652 | 100.056 | 109.2978 | 21.2612 | 0.005579 | 0.004713 | 217.7976 | 0.25931 | 8.4011 | 0.003403 | 0.005321 | 0.005182 | 6e-05 | 0.001177 | 0.31116 | 1.3163 | 2.5473 | 100.69 | 21.2 | 26.313 | 157.2364 | 103.7298 | 135.7536 | 0.006587 | 0.8634 | 0.021352 | 24.7307 | 0.063173 | 0.15427 | 0.02068 | 1.3076 | 1.3076 | 0.004873
2019-03-28 15:05 | 1553799871.7554 | 233 | 0.20526 | 0.0045 | 0.000355 | 100.6562 | 99.8818 | 109.6121 | 21.2612 | 0.005508 | 0.004656 | 217.7787 | 0.25889 | 8.404 | 0.003422 | 0.004972 | 0.004947 | 5.9e-05 | 0.001171 | 0.31797 | 1.2846 | 2.5472 | 100.69 | 21.2 | 26.313 | 155.4691 | 104.203 | 134.3481 | 0.006587 | 0.86637 | 0.02142 | 24.7236 | 0.069084 | 0.15401 | 0.020047 | 1.2714 | 1.2714 | 0.005003
2019-03-28 15:05 | 1553799872.7542 | 234 | 0.20531 | 0.004376 | 0.00035 | 100.7164 | 100.0577 | 109.8301 | 21.2612 | 0.005471 | 0.004653 | 217.7359 | 0.25873 | 8.4009 | 0.003416 | 0.004034 | 0.003924 | 6e-05 | 0.001178 | 0.31974 | 1.3367 | 2.5475 | 100.69 | 21.19 | 26.32 | 154.5439 | 104.06 | 134.2808 | 0.006585 | 0.86652 | 0.021443 | 24.7461 | 0.058684 | 0.15392 | 0.019885 | 1.2626 | 1.2626 | 0.005035
2019-03-28 15:05 | 1553799873.7545 | 235 | 0.20542 | 0.004255 | 0.000344 | 100.6887 | 99.9684 | 109.7691 | 21.2612 | 0.005454 | 0.00455 | 217.7228 | 0.25944 | 8.4051 | 0.003419 | 0.005287 | 0.00522 | 5.6e-05 | 0.001173 | 0.3183 | 1.2994 | 2.5472 | 100.69 | 21.19 | 26.32 | 154.1242 | 104.1213 | 131.746 | 0.006585 | 0.87194 | 0.021504 | 24.6622 | 0.064214 | 0.15434 | 0.019396 | 1.2349 | 1.2349 | 0.005162
2019-03-28 15:05 | 1553799874.7542 | 236 | 0.20556 | 0.004133 | 0.000339 | 100.7018 | 100.0333 | 110.0635 | 21.2612 | 0.005449 | 0.004623 | 217.6646 | 0.25906 | 8.4028 | 0.003318 | 0.004334 | 0.004285 | 5.8e-05 | 0.001175 | 0.32421 | 1.3144 | 2.5474 | 100.69 | 21.19 | 26.327 | 153.9876 | 101.6881 | 133.5459 | 0.006587 | 0.86808 | 0.021485 | 24.75 | 0.059575 | 0.15411 | 0.018692 | 1.1887 | 1.1887 | 0.005354
2019-03-28 15:05 | 1553799875.7553 | 237 | 0.20568 | 0.004021 | 0.000333 | 100.652 | 100.0513 | 110.0705 | 21.2612 | 0.005371 | 0.004536 | 217.6467 | 0.25797 | 8.4033 | 0.003379 | 0.005082 | 0.004976 | 5.9e-05 | 0.001174 | 0.35081 | 1.2929 | 2.5472 | 100.69 | 21.19 | 26.33 | 152.0524 | 103.1576 | 131.3929 | 0.006587 | 0.8727 | 0.021543 | 24.6852 | 0.053546 | 0.15347 | 0.018135 | 1.1563 | 1.1563 | 0.005482
2019-03-28 15:05 | 1553799876.7553 | 238 | 0.20582 | 0.003912 | 0.000327 | 100.6446 | 100.1308 | 110.1285 | 21.2612 | 0.005398 | 0.004555 | 217.6056 | 0.25914 | 8.4046 | 0.003299 | 0.004798 | 0.004801 | 5.8e-05 | 0.001175 | 0.30989 | 1.3061 | 2.5473 | 100.69 | 21.19 | 26.333 | 152.7323 | 101.2285 | 131.8664 | 0.006588 | 0.87168 | 0.021536 | 24.7062 | 0.04578 | 0.15416 | 0.017468 | 1.1131 | 1.1131 | 0.00572
2019-03-28 15:05 | 1553799877.7544 | 239 | 0.20597 | 0.003802 | 0.000321 | 100.7146 | 100.0598 | 110.1987 | 21.2612 | 0.005323 | 0.004471 | 217.5894 | 0.25899 | 8.4027 | 0.003283 | 0.004469 | 0.004376 | 5.6e-05 | 0.001167 | 0.30917 | 1.296 | 2.5473 | 100.69 | 21.19 | 26.335 | 150.8534 | 100.84 | 129.8073 | 0.006589 | 0.87614 | 0.021598 | 24.6511 | 0.058343 | 0.15407 | 0.016709 | 1.0675 | 1.0675 | 0.005961
2019-03-28 15:05 | 1553799878.7553 | 240 | 0.20603 | 0.003698 | 0.000316 | 100.591 | 99.9147 | 110.0782 | 21.2612 | 0.005263 | 0.004458 | 217.5288 | 0.25845 | 8.4068 | 0.00329 | 0.00495 | 0.004925 | 5.6e-05 | 0.001165 | 0.28799 | 1.2651 | 2.5473 | 100.69 | 21.19 | 26.337 | 149.3709 | 100.9948 | 129.4786 | 0.006589 | 0.87685 | 0.021595 | 24.6276 | 0.060335 | 0.15375 | 0.016441 | 1.0502 | 1.0502 | 0.006046
2019-03-28 15:05 | 1553799879.7538 | 241 | 0.20608 | 0.003592 | 0.000311 | 100.5963 | 99.9317 | 110.1057 | 21.2612 | 0.005232 | 0.004431 | 217.4874 | 0.25951 | 8.4028 | 0.003338 | 0.004786 | 0.004666 | 5.9e-05 | 0.001162 | 0.30574 | 1.25 | 2.5473 | 100.69 | 21.19 | 26.339 | 148.5984 | 102.1646 | 128.8152 | 0.00659 | 0.8783 | 0.021615 | 24.6104 | 0.059292 | 0.15438 | 0.016265 | 1.0401 | 1.0401 | 0.00613
2019-03-28 15:05 | 1553799880.7534 | 242 | 0.20614 | 0.003495 | 0.000304 | 100.5067 | 99.8515 | 110.2773 | 21.2612 | 0.00516 | 0.004427 | 217.4294 | 0.25864 | 8.4035 | 0.003274 | 0.004971 | 0.004858 | 5.5e-05 | 0.001159 | 0.3147 | 1.3099 | 2.5475 | 100.69 | 21.19 | 26.339 | 146.8306 | 100.6157 | 128.7297 | 0.00659 | 0.87848 | 0.021634 | 24.6269 | 0.058496 | 0.15386 | 0.016087 | 1.0298 | 1.0298 | 0.00617
2019-03-28 15:05 | 1553799881.7535 | 243 | 0.20623 | 0.003401 | 0.000299 | 100.4891 | 100.0924 | 110.366 | 21.2612 | 0.005129 | 0.004393 | 217.3717 | 0.25814 | 8.4012 | 0.003215 | 0.004846 | 0.004802 | 5.6e-05 | 0.001156 | 0.33339 | 1.3009 | 2.5473 | 100.69 | 21.19 | 26.341 | 146.0445 | 99.1878 | 127.8969 | 0.00659 | 0.88031 | 0.021666 | 24.6113 | 0.035381 | 0.15356 | 0.015635 | 1.0022 | 1.0022 | 0.006328
2019-03-28 15:05 | 1553799882.7535 | 244 | 0.20624 | 0.003317 | 0.000293 | 100.4218 | 99.6579 | 110.6362 | 21.2612 | 0.005162 | 0.004354 | 217.3132 | 0.25843 | 8.4002 | 0.003202 | 0.004828 | 0.004646 | 5.5e-05 | 0.001154 | 0.31404 | 1.3004 | 2.5475 | 100.69 | 21.19 | 26.341 | 146.8797 | 98.8741 | 126.9282 | 0.00659 | 0.88244 | 0.021718 | 24.6116 | 0.068301 | 0.15374 | 0.015671 | 1.0073 | 1.0073 | 0.006304
2019-03-28 15:05 | 1553799883.7538 | 245 | 0.20629 | 0.003238 | 0.00029 | 100.3536 | 99.7839 | 110.7911 | 21.2612 | 0.005104 | 0.004366 | 217.2448 | 0.25808 | 8.3992 | 0.003228 | 0.004727 | 0.004639 | 5.6e-05 | 0.001158 | 0.32139 | 1.3044 | 2.5475 | 100.69 | 21.19 | 26.339 | 145.4188 | 99.5093 | 127.2157 | 0.00659 | 0.88181 | 0.021726 | 24.6377 | 0.050918 | 0.15353 | 0.015497 | 0.99649 | 0.99649 | 0.006363
2019-03-28 15:05 | 1553799884.7546 | 246 | 0.20636 | 0.003173 | 0.000286 | 100.3184 | 99.7516 | 111.0162 | 21.2612 | 0.005034 | 0.004324 | 217.2209 | 0.25883 | 8.4012 | 0.003277 | 0.004991 | 0.004926 | 5.6e-05 | 0.00116 | 0.32853 | 1.3352 | 2.5474 | 100.69 | 21.19 | 26.339 | 143.6929 | 100.6815 | 126.1822 | 0.00659 | 0.88409 | 0.021776 | 24.6309 | 0.050679 | 0.15398 | 0.015185 | 0.97861 | 0.97861 | 0.006498
2019-03-28 15:05 | 1553799885.7539 | 247 | 0.20644 | 0.003108 | 0.000282 | 100.2726 | 99.6242 | 111.3303 | 21.2612 | 0.005018 | 0.004322 | 217.171 | 0.25845 | 8.4001 | 0.003193 | 0.005154 | 0.005027 | 5.5e-05 | 0.001162 | 0.30718 | 1.3206 | 2.5475 | 100.69 | 21.19 | 26.333 | 143.292 | 98.6679 | 126.1471 | 0.006588 | 0.88417 | 0.021808 | 24.6646 | 0.058019 | 0.15375 | 0.014748 | 0.95161 | 0.95161 | 0.006673
2019-03-28 15:05 | 1553799886.7545 | 248 | 0.20655 | 0.003045 | 0.000278 | 100.1787 | 99.5916 | 111.7633 | 21.2612 | 0.004995 | 0.004251 | 217.1468 | 0.25914 | 8.4036 | 0.00317 | 0.004473 | 0.004476 | 5.5e-05 | 0.001164 | 0.29264 | 1.3299 | 2.5475 | 100.69 | 21.19 | 26.333 | 142.7347 | 98.1025 | 124.391 | 0.006588 | 0.88807 | 0.021898 | 24.6581 | 0.052574 | 0.15416 | 0.014211 | 0.92049 | 0.92049 | 0.006917
2019-03-28 15:05 | 1553799887.7537 | 249 | 0.20655 | 0.002988 | 0.000273 | 100.1204 | 99.4262 | 111.8147 | 21.2612 | 0.005001 | 0.004273 | 217.0768 | 0.25793 | 8.4028 | 0.003151 | 0.004362 | 0.004284 | 5.5e-05 | 0.001158 | 0.31382 | 1.3353 | 2.5474 | 100.69 | 21.19 | 26.324 | 142.8864 | 97.6393 | 124.9374 | 0.006586 | 0.88685 | 0.021888 | 24.6807 | 0.062232 | 0.15344 | 0.01428 | 0.92488 | 0.92488 | 0.006852
2019-03-28 15:05 | 1553799888.7547 | 250 | 0.20664 | 0.002924 | 0.000268 | 100.0852 | 99.381 | 111.9471 | 21.2612 | 0.004948 | 0.004253 | 217.0772 | 0.258 | 8.4021 | 0.003168 | 0.004401 | 0.004304 | 5.1e-05 | 0.001159 | 0.31492 | 1.3226 | 2.5475 | 100.69 | 21.18 | 26.319 | 141.5592 | 98.0626 | 124.4388 | 0.006581 | 0.88797 | 0.021915 | 24.6799 | 0.063157 | 0.15349 | 0.01384 | 0.89729 | 0.89729 | 0.007065
2019-03-28 15:05 | 1553799889.7538 | 251 | 0.2068 | 0.002868 | 0.000263 | 100.0006 | 99.2852 | 112.3926 | 21.2612 | 0.004911 | 0.004181 | 217.0156 | 0.25853 | 8.4018 | 0.003181 | 0.004726 | 0.00471 | 5.5e-05 | 0.00116 | 0.30431 | 1.291 | 2.5473 | 100.69 | 21.18 | 26.319 | 140.6635 | 98.3682 | 122.6907 | 0.006581 | 0.89189 | 0.022007 | 24.6745 | 0.064221 | 0.1538 | 0.012939 | 0.84183 | 0.84183 | 0.007545
2019-03-28 15:05 | 1553799890.7547 | 252 | 0.20689 | 0.002807 | 0.000259 | 99.93 | 99.1738 | 112.6034 | 21.2612 | 0.004894 | 0.004204 | 216.98 | 0.25865 | 8.4019 | 0.003126 | 0.004834 | 0.0047 | 5.4e-05 | 0.001152 | 0.33826 | 1.3271 | 2.5473 | 100.69 | 21.18 | 26.307 | 140.2231 | 97.0382 | 123.2614 | 0.006578 | 0.8906 | 0.022012 | 24.7154 | 0.067949 | 0.15387 | 0.012486 | 0.81229 | 0.81229 | 0.007823
2019-03-28 15:05 | 1553799891.7538 | 253 | 0.20699 | 0.002744 | 0.000255 | 99.8448 | 99.2373 | 112.8794 | 21.2612 | 0.004799 | 0.004173 | 216.9689 | 0.25884 | 8.4025 | 0.003128 | 0.004757 | 0.004676 | 5.4e-05 | 0.001146 | 0.31686 | 1.3135 | 2.5474 | 100.69 | 21.18 | 26.307 | 137.875 | 97.0985 | 122.4995 | 0.006578 | 0.89232 | 0.02206 | 24.7219 | 0.054593 | 0.15398 | 0.011966 | 0.77984 | 0.77984 | 0.008155
2019-03-28 15:05 | 1553799892.7548 | 254 | 0.20705 | 0.002677 | 0.000252 | 99.9059 | 98.9496 | 112.9372 | 21.2612 | 0.004763 | 0.004151 | 216.9676 | 0.2585 | 8.403 | 0.00309 | 0.004279 | 0.004219 | 5.6e-05 | 0.00114 | 0.30735 | 1.3215 | 2.5473 | 100.69 | 21.18 | 26.293 | 136.9881 | 96.1774 | 121.9475 | 0.006574 | 0.89357 | 0.022081 | 24.711 | 0.086034 | 0.15378 | 0.011672 | 0.76127 | 0.76127 | 0.008343
2019-03-28 15:05 | 1553799893.7548 | 255 | 0.20715 | 0.002608 | 0.000248 | 99.7418 | 99.0195 | 113.4686 | 21.2612 | 0.004727 | 0.004104 | 216.9682 | 0.25834 | 8.4024 | 0.003106 | 0.004429 | 0.00443 | 5.6e-05 | 0.00114 | 0.34971 | 1.2778 | 2.5474 | 100.69 | 21.18 | 26.287 | 136.0977 | 96.5557 | 120.7948 | 0.006573 | 0.89618 | 0.022165 | 24.7329 | 0.065005 | 0.15368 | 0.011179 | 0.73156 | 0.73156 | 0.008676
2019-03-28 15:05 | 1553799894.7549 | 256 | 0.20729 | 0.002533 | 0.000244 | 99.6748 | 98.848 | 113.7773 | 21.2612 | 0.004722 | 0.004077 | 216.9376 | 0.2586 | 8.4033 | 0.003136 | 0.004412 | 0.004314 | 5.3e-05 | 0.00114 | 0.33373 | 1.3409 | 2.5474 | 100.69 | 21.18 | 26.278 | 135.9722 | 97.2741 | 120.1497 | 0.00657 | 0.89765 | 0.022213 | 24.7462 | 0.074504 | 0.15384 | 0.010414 | 0.68232 | 0.68232 | 0.009312
2019-03-28 15:05 | 1553799895.7532 | 257 | 0.20731 | 0.002463 | 0.00024 | 99.6661 | 99.009 | 113.9533 | 21.2612 | 0.004748 | 0.00406 | 216.9297 | 0.25832 | 8.4035 | 0.003054 | 0.004635 | 0.004514 | 5.4e-05 | 0.001138 | 0.27168 | 1.3485 | 2.5476 | 100.69 | 21.18 | 26.278 | 136.6295 | 95.3118 | 119.7355 | 0.00657 | 0.8986 | 0.022242 | 24.7523 | 0.05917 | 0.15368 | 0.010366 | 0.68017 | 0.68017 | 0.009331
2019-03-28 15:05 | 1553799896.754 | 258 | 0.20746 | 0.002388 | 0.000236 | 99.5318 | 98.5086 | 113.9864 | 21.2612 | 0.004659 | 0.004081 | 216.8989 | 0.25823 | 8.4046 | 0.003093 | 0.004468 | 0.004448 | 5.5e-05 | 0.001141 | 0.29015 | 1.3457 | 2.5475 | 100.69 | 21.18 | 26.275 | 134.4326 | 96.2438 | 120.2299 | 0.00657 | 0.89747 | 0.022232 | 24.7715 | 0.092428 | 0.15362 | 0.009614 | 0.62982 | 0.62982 | 0.010074
2019-03-28 15:05 | 1553799897.754 | 259 | 0.20753 | 0.00231 | 0.000232 | 99.5074 | 98.5626 | 114.6971 | 21.2612 | 0.004662 | 0.004044 | 216.9077 | 0.25936 | 8.4034 | 0.0031 | 0.00416 | 0.004064 | 5.6e-05 | 0.001139 | 0.32261 | 1.3189 | 2.5474 | 100.69 | 21.18 | 26.268 | 134.5103 | 96.413 | 119.3281 | 0.006568 | 0.89953 | 0.022326 | 24.8201 | 0.08533 | 0.15429 | 0.009259 | 0.6089 | 0.6089 | 0.010465
2019-03-28 15:05 | 1553799898.753 | 260 | 0.20762 | 0.002238 | 0.000229 | 99.5014 | 98.5969 | 115.0809 | 21.2612 | 0.004634 | 0.003983 | 216.843 | 0.25911 | 8.4023 | 0.003062 | 0.004559 | 0.004569 | 5.1e-05 | 0.00114 | 0.30237 | 1.3106 | 2.5474 | 100.69 | 21.18 | 26.268 | 133.8066 | 95.5015 | 117.8563 | 0.006568 | 0.90292 | 0.022406 | 24.8149 | 0.081681 | 0.15415 | 0.008798 | 0.58021 | 0.58021 | 0.010972
2019-03-28 15:05 | 1553799899.753 | 261 | 0.20769 | 0.002164 | 0.000225 | 99.4177 | 98.5344 | 114.9445 | 21.2612 | 0.004606 | 0.003946 | 216.8432 | 0.25821 | 8.4019 | 0.003102 | 0.004362 | 0.004165 | 5.7e-05 | 0.001143 | 0.29856 | 1.2695 | 2.5474 | 100.69 | 21.18 | 26.262 | 133.1153 | 96.4521 | 116.9423 | 0.006566 | 0.90503 | 0.022419 | 24.7712 | 0.079827 | 0.15361 | 0.008498 | 0.56051 | 0.56051 | 0.011318
2019-03-28 15:05 | 1553799900.7545 | 262 | 0.20776 | 0.002091 | 0.000221 | 99.4019 | 98.7138 | 115.222 | 21.2612 | 0.004573 | 0.004008 | 216.8374 | 0.25915 | 8.4023 | 0.003095 | 0.004296 | 0.004282 | 5.6e-05 | 0.001142 | 0.33859 | 1.3023 | 2.5473 | 100.69 | 21.18 | 26.262 | 132.3201 | 96.2925 | 118.4441 | 0.006566 | 0.90156 | 0.022403 | 24.8488 | 0.062138 | 0.15417 | 0.008146 | 0.53663 | 0.53663 | 0.011865
2019-03-28 15:05 | 1553799901.7541 | 263 | 0.20783 | 0.002025 | 0.000218 | 99.3799 | 98.4461 | 115.1588 | 21.2612 | 0.004506 | 0.003954 | 216.7993 | 0.25849 | 8.4003 | 0.003073 | 0.004594 | 0.004475 | 5.7e-05 | 0.001138 | 0.335 | 1.2913 | 2.5472 | 100.69 | 21.18 | 26.261 | 130.6737 | 95.7559 | 117.1488 | 0.006566 | 0.90455 | 0.022434 | 24.8009 | 0.084444 | 0.15377 | 0.007833 | 0.51648 | 0.51648 | 0.012296
2019-03-28 15:05 | 1553799902.7542 | 264 | 0.20788 | 0.001957 | 0.000215 | 99.2664 | 98.391 | 115.3593 | 21.2612 | 0.004459 | 0.003955 | 216.7632 | 0.25835 | 8.4046 | 0.00315 | 0.004509 | 0.004457 | 5.3e-05 | 0.001138 | 0.33124 | 1.2986 | 2.5474 | 100.69 | 21.18 | 26.261 | 129.5098 | 97.6198 | 117.1512 | 0.006566 | 0.90455 | 0.022453 | 24.8225 | 0.079227 | 0.15369 | 0.00761 | 0.50207 | 0.50207 | 0.012643
2019-03-28 15:05 | 1553799903.7537 | 265 | 0.20795 | 0.001897 | 0.000212 | 99.2739 | 98.443 | 115.1882 | 21.2612 | 0.004478 | 0.00389 | 216.7391 | 0.25854 | 8.4035 | 0.003022 | 0.004556 | 0.004449 | 5.5e-05 | 0.001138 | 0.31996 | 1.3089 | 2.5474 | 100.69 | 21.18 | 26.259 | 129.97 | 94.5182 | 115.5801 | 0.006566 | 0.9082 | 0.022482 | 24.7541 | 0.075174 | 0.1538 | 0.007253 | 0.47868 | 0.47868 | 0.01327
2019-03-28 15:05 | 1553799904.7542 | 266 | 0.20798 | 0.001835 | 0.000209 | 99.2697 | 98.3601 | 115.0464 | 21.2612 | 0.004477 | 0.0039 | 216.7268 | 0.25858 | 8.4035 | 0.003074 | 0.004926 | 0.004843 | 5.8e-05 | 0.001138 | 0.29175 | 1.2962 | 2.5474 | 100.69 | 21.18 | 26.259 | 129.9491 | 95.7841 | 115.8293 | 0.006566 | 0.90762 | 0.022461 | 24.7468 | 0.082333 | 0.15383 | 0.007196 | 0.47456 | 0.47456 | 0.013388
2019-03-28 15:05 | 1553799905.7536 | 267 | 0.20806 | 0.001782 | 0.000206 | 99.2003 | 98.4029 | 115.2185 | 21.2612 | 0.004424 | 0.003811 | 216.6976 | 0.25853 | 8.4012 | 0.003136 | 0.004405 | 0.004279 | 5.6e-05 | 0.001144 | 0.32493 | 1.321 | 2.5475 | 100.69 | 21.18 | 26.256 | 128.6525 | 97.2811 | 113.6442 | 0.006565 | 0.91275 | 0.022541 | 24.6957 | 0.072186 | 0.1538 | 0.006782 | 0.44839 | 0.44839 | 0.014166
2019-03-28 15:05 | 1553799906.7533 | 268 | 0.20812 | 0.001727 | 0.000202 | 99.1464 | 98.3149 | 115.1579 | 21.2612 | 0.004386 | 0.003823 | 216.6773 | 0.25944 | 8.4046 | 0.003123 | 0.005044 | 0.004969 | 5.7e-05 | 0.001145 | 0.34468 | 1.2914 | 2.5472 | 100.69 | 21.18 | 26.25 | 127.7137 | 96.9701 | 113.9534 | 0.006563 | 0.91202 | 0.022526 | 24.699 | 0.075325 | 0.15434 | 0.0065 | 0.42919 | 0.42919 | 0.014852
2019-03-28 15:05 | 1553799907.7535 | 269 | 0.20815 | 0.001675 | 0.000198 | 99.1038 | 98.3747 | 115.037 | 21.2612 | 0.004314 | 0.003756 | 216.64 | 0.25858 | 8.4029 | 0.003107 | 0.004894 | 0.004909 | 5.5e-05 | 0.001142 | 0.30525 | 1.2944 | 2.5473 | 100.69 | 21.18 | 26.25 | 125.9432 | 96.59 | 112.3067 | 0.006563 | 0.91591 | 0.022562 | 24.6335 | 0.066043 | 0.15383 | 0.006392 | 0.42279 | 0.42279 | 0.015027
2019-03-28 15:05 | 1553799908.7544 | 270 | 0.20823 | 0.001626 | 0.000194 | 99.1541 | 98.0751 | 115.2936 | 21.2612 | 0.004343 | 0.003807 | 216.6399 | 0.25816 | 8.4042 | 0.003087 | 0.00426 | 0.004082 | 5.3e-05 | 0.001144 | 0.35242 | 1.2883 | 2.5474 | 100.69 | 21.18 | 26.239 | 126.6511 | 96.0881 | 113.5449 | 0.006561 | 0.91298 | 0.022551 | 24.7005 | 0.097871 | 0.15358 | 0.005968 | 0.39406 | 0.39406 | 0.016096
2019-03-28 15:05 | 1553799909.7537 | 271 | 0.2083 | 0.001575 | 0.00019 | 99.1535 | 98.0734 | 115.5541 | 21.2612 | 0.004266 | 0.003746 | 216.6243 | 0.25718 | 8.4017 | 0.003119 | 0.004424 | 0.004381 | 5.6e-05 | 0.001138 | 0.30486 | 1.3071 | 2.5474 | 100.69 | 21.18 | 26.239 | 124.768 | 96.8628 | 112.0636 | 0.006561 | 0.91649 | 0.02262 | 24.681 | 0.097972 | 0.153 | 0.005633 | 0.37272 | 0.37272 | 0.016953
2019-03-28 15:05 | 1553799910.7534 | 272 | 0.20832 | 0.00153 | 0.000187 | 99.1223 | 98.3146 | 115.2168 | 21.2612 | 0.004324 | 0.003735 | 216.619 | 0.25829 | 8.4025 | 0.003083 | 0.00436 | 0.004314 | 5.6e-05 | 0.001141 | 0.31349 | 1.3357 | 2.5473 | 100.69 | 21.17 | 26.221 | 126.1822 | 95.9939 | 111.7906 | 0.006552 | 0.91714 | 0.022595 | 24.6362 | 0.073179 | 0.15366 | 0.005526 | 0.3652 | 0.3652 | 0.017377
2019-03-28 15:05 | 1553799911.7534 | 273 | 0.20836 | 0.001484 | 0.000183 | 99.1774 | 98.3088 | 115.4858 | 21.2612 | 0.004208 | 0.003738 | 216.5996 | 0.2581 | 8.4027 | 0.003 | 0.004507 | 0.004408 | 5.4e-05 | 0.001148 | 0.31492 | 1.3129 | 2.5474 | 100.69 | 21.17 | 26.221 | 123.3559 | 94.0047 | 111.8757 | 0.006552 | 0.91694 | 0.022619 | 24.6677 | 0.078675 | 0.15354 | 0.00536 | 0.35451 | 0.35451 | 0.017888
2019-03-28 15:05 | 1553799912.7534 | 274 | 0.20843 | 0.001444 | 0.00018 | 99.1279 | 98.2623 | 115.8628 | 21.2612 | 0.00421 | 0.003722 | 216.6161 | 0.25834 | 8.4041 | 0.003078 | 0.004472 | 0.004351 | 5.3e-05 | 0.00115 | 0.32947 | 1.2791 | 2.5473 | 100.69 | 21.17 | 26.203 | 123.3956 | 95.8867 | 111.4864 | 0.006548 | 0.91787 | 0.022667 | 24.6954 | 0.078451 | 0.15368 | 0.005007 | 0.3314 | 0.3314 | 0.019153
2019-03-28 15:05 | 1553799913.7524 | 275 | 0.20844 | 0.001398 | 0.000177 | 99.1315 | 98.2317 | 116.1534 | 21.2612 | 0.00417 | 0.00367 | 216.5776 | 0.2592 | 8.4029 | 0.00306 | 0.004392 | 0.004315 | 5.5e-05 | 0.001149 | 0.32139 | 1.3259 | 2.5474 | 100.69 | 21.17 | 26.203 | 122.4104 | 95.4513 | 110.2145 | 0.006548 | 0.92091 | 0.022733 | 24.6855 | 0.081557 | 0.15419 | 0.004974 | 0.33025 | 0.33025 | 0.019283
2019-03-28 15:05 | 1553799914.7526 | 276 | 0.2085 | 0.001361 | 0.000173 | 99.192 | 98.2419 | 116.4665 | 21.2612 | 0.004205 | 0.003668 | 216.5516 | 0.25826 | 8.4022 | 0.002965 | 0.003781 | 0.003752 | 5.8e-05 | 0.001155 | 0.35651 | 1.2856 | 2.5475 | 100.69 | 21.17 | 26.185 | 123.267 | 93.1628 | 110.1688 | 0.006543 | 0.92102 | 0.022765 | 24.7172 | 0.086083 | 0.15364 | 0.0047 | 0.31214 | 0.31214 | 0.020328
2019-03-28 15:05 | 1553799915.7525 | 277 | 0.20855 | 0.00132 | 0.000171 | 99.1288 | 98.0316 | 116.6381 | 21.2612 | 0.004135 | 0.003703 | 216.532 | 0.2594 | 8.4026 | 0.00303 | 0.00447 | 0.004346 | 5.7e-05 | 0.001155 | 0.32731 | 1.3013 | 2.5474 | 100.69 | 21.17 | 26.185 | 121.5535 | 94.7129 | 111.0131 | 0.006543 | 0.919 | 0.022757 | 24.7627 | 0.099554 | 0.15432 | 0.004455 | 0.29542 | 0.29542 | 0.021573
2019-03-28 15:05 | 1553799916.7526 | 278 | 0.20856 | 0.001279 | 0.000168 | 99.2301 | 98.3031 | 116.9251 | 21.2612 | 0.004152 | 0.003623 | 216.5155 | 0.25925 | 8.4002 | 0.003023 | 0.005052 | 0.004989 | 5.9e-05 | 0.001157 | 0.30663 | 1.2993 | 2.5474 | 100.69 | 21.17 | 26.176 | 121.984 | 94.5472 | 109.065 | 0.006541 | 0.92368 | 0.022843 | 24.7302 | 0.083954 | 0.15423 | 0.004409 | 0.29349 | 0.29349 | 0.021703
2019-03-28 15:05 | 1553799917.7526 | 279 | 0.2086 | 0.001245 | 0.000165 | 99.2146 | 98.0568 | 116.7963 | 21.2612 | 0.004119 | 0.00363 | 216.5173 | 0.25781 | 8.4026 | 0.003001 | 0.004575 | 0.00448 | 5.6e-05 | 0.001159 | 0.32892 | 1.328 | 2.5475 | 100.69 | 21.17 | 26.169 | 121.1725 | 94.0141 | 109.2533 | 0.006539 | 0.92323 | 0.022825 | 24.7226 | 0.105 | 0.15337 | 0.004254 | 0.28278 | 0.28278 | 0.0224
2019-03-28 15:05 | 1553799918.7526 | 280 | 0.20863 | 0.001215 | 0.000163 | 99.2223 | 98.2796 | 116.8965 | 21.2612 | 0.004016 | 0.003551 | 216.5249 | 0.25807 | 8.405 | 0.002964 | 0.004639 | 0.004544 | 5.2e-05 | 0.001158 | 0.30879 | 1.2899 | 2.5473 | 100.69 | 21.17 | 26.169 | 118.6391 | 93.1252 | 107.3323 | 0.006539 | 0.92789 | 0.022892 | 24.671 | 0.085389 | 0.15353 | 0.004069 | 0.27106 | 0.27106 | 0.023392
2019-03-28 15:05 | 1553799919.7527 | 281 | 0.20863 | 0.001183 | 0.00016 | 99.3266 | 98.3289 | 117.0132 | 21.2612 | 0.004044 | 0.003654 | 216.5275 | 0.25888 | 8.4042 | 0.002981 | 0.004955 | 0.004889 | 5.8e-05 | 0.00116 | 0.32908 | 1.2843 | 2.5473 | 100.69 | 21.16 | 26.17 | 119.3303 | 93.5291 | 109.8227 | 0.006535 | 0.92185 | 0.022829 | 24.764 | 0.0903 | 0.15401 | 0.004135 | 0.27496 | 0.27496 | 0.023133
2019-03-28 15:05 | 1553799920.7529 | 282 | 0.20867 | 0.00115 | 0.000159 | 99.295 | 98.404 | 117.193 | 21.2612 | 0.004033 | 0.003546 | 216.5132 | 0.25795 | 8.4044 | 0.002955 | 0.004264 | 0.004159 | 5.3e-05 | 0.001158 | 0.32676 | 1.271 | 2.5473 | 100.69 | 21.16 | 26.171 | 119.071 | 92.9055 | 107.2165 | 0.006536 | 0.92817 | 0.022924 | 24.6985 | 0.080629 | 0.15345 | 0.00397 | 0.26479 | 0.26479 | 0.023935
2019-03-28 15:05 | 1553799921.7538 | 283 | 0.20869 | 0.001122 | 0.000158 | 99.415 | 98.4478 | 117.312 | 21.2612 | 0.004062 | 0.003566 | 216.5347 | 0.25859 | 8.4009 | 0.002959 | 0.003945 | 0.003781 | 5.5e-05 | 0.001154 | 0.36696 | 1.2608 | 2.5473 | 100.69 | 21.16 | 26.173 | 119.7864 | 92.9979 | 107.6853 | 0.006536 | 0.92703 | 0.022922 | 24.7263 | 0.087453 | 0.15383 | 0.003854 | 0.25686 | 0.25686 | 0.024735
2019-03-28 15:05 | 1553799922.7528 | 284 | 0.20875 | 0.001091 | 0.000156 | 99.4386 | 98.3251 | 117.1296 | 21.2612 | 0.003977 | 0.003494 | 216.491 | 0.25862 | 8.4032 | 0.002945 | 0.004356 | 0.004312 | 5.4e-05 | 0.001154 | 0.30596 | 1.2945 | 2.5473 | 100.69 | 21.16 | 26.177 | 117.7032 | 92.6818 | 105.9532 | 0.006537 | 0.93127 | 0.022956 | 24.6508 | 0.10072 | 0.15385 | 0.003513 | 0.23382 | 0.23382 | 0.027175
2019-03-28 15:05 | 1553799923.7539 | 285 | 0.20879 | 0.001062 | 0.000154 | 99.497 | 98.3489 | 116.7327 | 21.2612 | 0.003982 | 0.003502 | 216.5026 | 0.2586 | 8.4045 | 0.002924 | 0.00383 | 0.00381 | 5.6e-05 | 0.001153 | 0.30431 | 1.2666 | 2.5473 | 100.69 | 21.16 | 26.179 | 117.8087 | 92.1727 | 106.1341 | 0.006538 | 0.93082 | 0.022912 | 24.6149 | 0.10381 | 0.15384 | 0.003347 | 0.22201 | 0.22201 | 0.028619
2019-03-28 15:05 | 1553799924.7538 | 286 | 0.20888 | 0.001033 | 0.000152 | 99.5431 | 98.4801 | 116.8183 | 21.2612 | 0.003977 | 0.00348 | 216.4947 | 0.25887 | 8.4013 | 0.002874 | 0.004363 | 0.004264 | 5.5e-05 | 0.001156 | 0.33682 | 1.2856 | 2.5473 | 100.69 | 21.16 | 26.183 | 117.689 | 90.9615 | 105.5952 | 0.006539 | 0.93215 | 0.022937 | 24.6064 | 0.096037 | 0.154 | 0.002841 | 0.18743 | 0.18743 | 0.033933
2019-03-28 15:05 | 1553799925.7531 | 287 | 0.20888 | 0.001002 | 0.00015 | 99.6027 | 98.5208 | 116.8981 | 21.2612 | 0.003913 | 0.003517 | 216.5079 | 0.25894 | 8.4013 | 0.002919 | 0.004134 | 0.004072 | 5.8e-05 | 0.001153 | 0.32659 | 1.2892 | 2.5473 | 100.69 | 21.16 | 26.187 | 116.1403 | 92.0445 | 106.5104 | 0.00654 | 0.9299 | 0.022917 | 24.6446 | 0.097686 | 0.15404 | 0.002884 | 0.19034 | 0.19034 | 0.033425
2019-03-28 15:05 | 1553799926.7529 | 288 | 0.2089 | 0.000974 | 0.000148 | 99.6408 | 98.3771 | 116.8362 | 21.2612 | 0.003883 | 0.003435 | 216.4805 | 0.25847 | 8.4028 | 0.002883 | 0.004379 | 0.00421 | 5.7e-05 | 0.001146 | 0.33052 | 1.3031 | 2.5475 | 100.69 | 21.16 | 26.187 | 115.4123 | 91.1796 | 104.5168 | 0.00654 | 0.93481 | 0.022971 | 24.5733 | 0.11417 | 0.15376 | 0.002784 | 0.184 | 0.184 | 0.034512
2019-03-28 15:05 | 1553799927.7529 | 289 | 0.20895 | 0.000949 | 0.000146 | 99.7254 | 98.6064 | 117.192 | 21.2612 | 0.003829 | 0.003427 | 216.4787 | 0.25848 | 8.4019 | 0.002882 | 0.004103 | 0.004111 | 5.5e-05 | 0.001144 | 0.30431 | 1.3333 | 2.5475 | 100.69 | 21.16 | 26.199 | 114.0862 | 91.1649 | 104.3319 | 0.006543 | 0.93526 | 0.023012 | 24.6046 | 0.10093 | 0.15377 | 0.002534 | 0.16713 | 0.16713 | 0.037999
2019-03-28 15:05 | 1553799928.753 | 290 | 0.20897 | 0.000923 | 0.000144 | 99.8058 | 98.5008 | 117.407 | 21.2612 | 0.003854 | 0.003418 | 216.4437 | 0.25873 | 8.4043 | 0.002878 | 0.004643 | 0.004592 | 5.5e-05 | 0.001142 | 0.31758 | 1.3194 | 2.5474 | 100.69 | 21.16 | 26.199 | 114.6924 | 91.0625 | 104.1133 | 0.006543 | 0.93581 | 0.02304 | 24.62 | 0.11773 | 0.15392 | 0.002466 | 0.16272 | 0.16272 | 0.039067
2019-03-28 15:05 | 1553799929.753 | 291 | 0.20895 | 0.000895 | 0.000141 | 99.8934 | 98.8205 | 117.5518 | 21.2612 | 0.003785 | 0.003399 | 216.4374 | 0.25886 | 8.4007 | 0.00285 | 0.00433 | 0.004213 | 6e-05 | 0.001139 | 0.31382 | 1.2901 | 2.5473 | 100.69 | 21.16 | 26.21 | 113.0137 | 90.3831 | 103.6356 | 0.006545 | 0.93699 | 0.023068 | 24.6196 | 0.096592 | 0.15399 | 0.002584 | 0.17129 | 0.17129 | 0.037129
2019-03-28 15:05 | 1553799930.7522 | 292 | 0.20895 | 0.000878 | 0.000138 | 100.0009 | 99.0722 | 117.5148 | 21.2612 | 0.003827 | 0.00342 | 216.419 | 0.25898 | 8.4006 | 0.0029 | 0.004557 | 0.004509 | 5.6e-05 | 0.001139 | 0.32421 | 1.2957 | 2.5473 | 100.69 | 21.16 | 26.21 | 114.0442 | 91.5962 | 104.14 | 0.006545 | 0.93574 | 0.023049 | 24.6322 | 0.083458 | 0.15407 | 0.002641 | 0.1753 | 0.1753 | 0.036297
2019-03-28 15:06 | 1553799931.7527 | 293 | 0.20899 | 0.000854 | 0.000135 | 100.0432 | 98.8293 | 117.6018 | 21.2612 | 0.003826 | 0.003408 | 216.4147 | 0.25857 | 8.4016 | 0.002906 | 0.003985 | 0.003946 | 5.8e-05 | 0.001144 | 0.31885 | 1.3294 | 2.5475 | 100.69 | 21.16 | 26.217 | 114.0111 | 91.7442 | 103.8674 | 0.006547 | 0.93642 | 0.023066 | 24.6324 | 0.1092 | 0.15382 | 0.002414 | 0.15983 | 0.15983 | 0.039749
2019-03-28 15:06 | 1553799932.7531 | 294 | 0.20906 | 0.000835 | 0.000132 | 100.1201 | 99.0997 | 117.6324 | 21.2612 | 0.003815 | 0.003378 | 216.434 | 0.25791 | 8.4014 | 0.002864 | 0.004691 | 0.004593 | 6e-05 | 0.001146 | 0.31526 | 1.263 | 2.5474 | 100.69 | 21.16 | 26.217 | 113.7446 | 90.7129 | 103.1223 | 0.006547 | 0.93827 | 0.023092 | 24.6113 | 0.091624 | 0.15343 | 0.002053 | 0.13504 | 0.13504 | 0.046924
2019-03-28 15:06 | 1553799933.7523 | 295 | 0.20903 | 0.000818 | 0.00013 | 100.135 | 99.2509 | 117.8255 | 21.2612 | 0.003802 | 0.003374 | 216.4285 | 0.25805 | 8.4022 | 0.002879 | 0.004822 | 0.004691 | 6e-05 | 0.001147 | 0.3131 | 1.283 | 2.5474 | 100.69 | 21.16 | 26.227 | 113.4223 | 91.0772 | 103.0421 | 0.00655 | 0.93847 | 0.023113 | 24.6288 | 0.07932 | 0.15351 | 0.002255 | 0.14942 | 0.14942 | 0.042432
2019-03-28 15:06 | 1553799934.7524 | 296 | 0.20909 | 0.000805 | 0.000127 | 100.1332 | 99.0545 | 117.792 | 21.2612 | 0.003693 | 0.003407 | 216.4293 | 0.2589 | 8.4023 | 0.00285 | 0.004953 | 0.004799 | 5.7e-05 | 0.001144 | 0.34219 | 1.284 | 2.5474 | 100.69 | 21.16 | 26.231 | 110.7818 | 90.3761 | 103.8461 | 0.006551 | 0.93647 | 0.023086 | 24.6516 | 0.096874 | 0.15402 | 0.001916 | 0.12582 | 0.12582 | 0.050556
2019-03-28 15:06 | 1553799935.7522 | 297 | 0.20904 | 0.000786 | 0.000125 | 100.235 | 99.2091 | 117.6955 | 21.2612 | 0.003728 | 0.00337 | 216.412 | 0.25832 | 8.4039 | 0.002901 | 0.004688 | 0.004675 | 5.8e-05 | 0.001141 | 0.3314 | 1.2969 | 2.5474 | 100.69 | 21.16 | 26.237 | 111.6412 | 91.6198 | 102.929 | 0.006552 | 0.93875 | 0.023104 | 24.6116 | 0.092021 | 0.15368 | 0.002204 | 0.14609 | 0.14609 | 0.043443
2019-03-28 15:06 | 1553799936.7522 | 298 | 0.20907 | 0.000773 | 0.000123 | 100.2523 | 99.2431 | 117.534 | 21.2612 | 0.003757 | 0.003328 | 216.4382 | 0.25885 | 8.4032 | 0.002817 | 0.004274 | 0.004198 | 5.9e-05 | 0.001133 | 0.323 | 1.3043 | 2.5474 | 100.69 | 21.16 | 26.24 | 112.3384 | 89.583 | 101.9117 | 0.006553 | 0.9413 | 0.02312 | 24.5614 | 0.0905 | 0.15399 | 0.002032 | 0.1343 | 0.1343 | 0.047354
2019-03-28 15:06 | 1553799937.7522 | 299 | 0.20907 | 0.000763 | 0.000122 | 100.4072 | 99.398 | 117.4453 | 21.2612 | 0.003721 | 0.003311 | 216.3942 | 0.25894 | 8.4028 | 0.002877 | 0.004209 | 0.004065 | 5.7e-05 | 0.001133 | 0.29446 | 1.3001 | 2.5473 | 100.69 | 21.16 | 26.24 | 111.4645 | 91.0444 | 101.5216 | 0.006553 | 0.94228 | 0.023123 | 24.5393 | 0.090357 | 0.15404 | 0.002062 | 0.1365 | 0.1365 | 0.046606
2019-03-28 15:06 | 1553799938.7524 | 300 | 0.2091 | 0.000752 | 0.00012 | 100.4006 | 99.5906 | 117.2861 | 21.2612 | 0.003673 | 0.003293 | 216.4091 | 0.25819 | 8.402 | 0.002801 | 0.004314 | 0.004237 | 5.8e-05 | 0.001132 | 0.30879 | 1.3143 | 2.5473 | 100.69 | 21.16 | 26.247 | 110.2848 | 89.2009 | 101.0858 | 0.006555 | 0.94338 | 0.023121 | 24.5084 | 0.072456 | 0.1536 | 0.001907 | 0.12582 | 0.12582 | 0.050419
2019-03-28 15:06 | 1553799939.7523 | 301 | 0.20916 | 0.000742 | 0.000118 | 100.4697 | 99.6484 | 117.1406 | 21.2612 | 0.003672 | 0.003299 | 216.371 | 0.25777 | 8.4016 | 0.002907 | 0.004434 | 0.004399 | 5.6e-05 | 0.001131 | 0.31725 | 1.2821 | 2.5473 | 100.69 | 21.16 | 26.247 | 110.2563 | 91.762 | 101.2232 | 0.006555 | 0.94303 | 0.023102 | 24.4977 | 0.073425 | 0.15335 | 0.00153 | 0.099529 | 0.099529 | 0.063632
2019-03-28 15:06 | 1553799940.7533 | 302 | 0.20906 | 0.000731 | 0.000117 | 100.5314 | 99.5539 | 117.2763 | 21.2612 | 0.003642 | 0.003292 | 216.3867 | 0.25843 | 8.4021 | 0.002913 | 0.003794 | 0.003801 | 5.5e-05 | 0.001133 | 0.31581 | 1.2802 | 2.5474 | 100.69 | 21.16 | 26.256 | 109.5412 | 91.8902 | 101.0606 | 0.006557 | 0.94344 | 0.02312 | 24.5066 | 0.087394 | 0.15374 | 0.002153 | 0.14306 | 0.14306 | 0.044381
2019-03-28 15:06 | 1553799941.7534 | 303 | 0.20908 | 0.000723 | 0.000115 | 100.6285 | 99.7527 | 117.1866 | 21.2612 | 0.003676 | 0.003305 | 216.3701 | 0.25827 | 8.4022 | 0.002823 | 0.004543 | 0.004407 | 5.5e-05 | 0.001132 | 0.29717 | 1.3231 | 2.5475 | 100.69 | 21.16 | 26.259 | 110.355 | 89.7415 | 101.3708 | 0.006558 | 0.94266 | 0.023102 | 24.5074 | 0.078191 | 0.15364 | 0.002089 | 0.13858 | 0.13858 | 0.045789
2019-03-28 15:06 | 1553799942.7524 | 304 | 0.20906 | 0.000716 | 0.000114 | 100.5558 | 99.8277 | 117.2993 | 21.2612 | 0.003617 | 0.003297 | 216.3705 | 0.2582 | 8.4042 | 0.00287 | 0.004583 | 0.004538 | 5.5e-05 | 0.001134 | 0.33356 | 1.2903 | 2.5474 | 100.69 | 21.16 | 26.266 | 108.9364 | 90.8584 | 101.1808 | 0.006559 | 0.94314 | 0.023119 | 24.5129 | 0.064999 | 0.1536 | 0.002182 | 0.14528 | 0.14528 | 0.043667
2019-03-28 15:06 | 1553799943.7526 | 305 | 0.20915 | 0.000715 | 0.000113 | 100.6225 | 99.7489 | 117.3434 | 21.2612 | 0.003635 | 0.003237 | 216.3537 | 0.25854 | 8.4023 | 0.00285 | 0.004229 | 0.00418 | 5.3e-05 | 0.001137 | 0.31061 | 1.2923 | 2.5476 | 100.69 | 21.16 | 26.27 | 109.3575 | 90.3932 | 99.7163 | 0.00656 | 0.94684 | 0.023169 | 24.4695 | 0.077996 | 0.1538 | 0.001676 | 0.11037 | 0.11037 | 0.057553
2019-03-28 15:06 | 1553799944.7515 | 306 | 0.20911 | 0.000706 | 0.000112 | 100.5898 | 99.8559 | 117.2737 | 21.2612 | 0.003601 | 0.003291 | 216.3288 | 0.2581 | 8.4029 | 0.002811 | 0.003819 | 0.003614 | 5e-05 | 0.001137 | 0.30431 | 1.3153 | 2.5475 | 100.69 | 21.16 | 26.272 | 108.5322 | 89.4431 | 101.026 | 0.006561 | 0.94353 | 0.023121 | 24.5052 | 0.065492 | 0.15354 | 0.001917 | 0.12697 | 0.12697 | 0.049944
2019-03-28 15:06 | 1553799945.7528 | 307 | 0.20908 | 0.000702 | 0.00011 | 100.7473 | 99.9453 | 117.2716 | 21.2612 | 0.003571 | 0.003213 | 216.3147 | 0.25763 | 8.4021 | 0.002782 | 0.005103 | 0.004992 | 5.4e-05 | 0.001141 | 0.32803 | 1.2877 | 2.5473 | 100.69 | 21.15 | 26.277 | 107.8113 | 88.7487 | 99.1453 | 0.006558 | 0.94829 | 0.023179 | 24.4433 | 0.071493 | 0.15327 | 0.00209 | 0.13939 | 0.13939 | 0.04541
2019-03-28 15:06 | 1553799946.7529 | 308 | 0.20911 | 0.000695 | 0.00011 | 100.775 | 99.859 | 117.5298 | 21.2612 | 0.003571 | 0.00321 | 216.2947 | 0.25778 | 8.4019 | 0.002846 | 0.004726 | 0.004703 | 5.4e-05 | 0.001141 | 0.31404 | 1.3201 | 2.5474 | 100.69 | 21.15 | 26.279 | 107.8246 | 90.2947 | 99.0645 | 0.006559 | 0.9485 | 0.023207 | 24.4675 | 0.081677 | 0.15335 | 0.001949 | 0.1298 | 0.1298 | 0.048795
2019-03-28 15:06 | 1553799947.7526 | 309 | 0.20916 | 0.000691 | 0.000108 | 100.8087 | 100.0421 | 117.3474 | 21.2612 | 0.003566 | 0.003241 | 216.3082 | 0.25763 | 8.4013 | 0.002875 | 0.004115 | 0.00404 | 5.1e-05 | 0.001142 | 0.34379 | 1.305 | 2.5473 | 100.69 | 21.15 | 26.283 | 107.6986 | 90.9779 | 99.8215 | 0.00656 | 0.94658 | 0.023166 | 24.4734 | 0.068278 | 0.15326 | 0.001649 | 0.10868 | 0.10868 | 0.058241
2019-03-28 15:06 | 1553799948.7516 | 310 | 0.2091 | 0.000685 | 0.000107 | 100.8733 | 99.9035 | 117.4317 | 21.2612 | 0.003472 | 0.003255 | 216.2854 | 0.25694 | 8.4016 | 0.002791 | 0.004705 | 0.004638 | 5.3e-05 | 0.001141 | 0.29717 | 1.2994 | 2.5475 | 100.69 | 21.15 | 26.283 | 105.4119 | 88.968 | 100.1472 | 0.00656 | 0.94575 | 0.023164 | 24.4929 | 0.086412 | 0.15285 | 0.001994 | 0.13284 | 0.13284 | 0.047522
2019-03-28 15:06 | 1553799949.7517 | 311 | 0.20913 | 0.000683 | 0.000105 | 100.8858 | 100.2334 | 117.276 | 21.2612 | 0.003562 | 0.003221 | 216.3131 | 0.25817 | 8.4015 | 0.00282 | 0.004022 | 0.003945 | 4.9e-05 | 0.001142 | 0.323 | 1.2717 | 2.5474 | 100.69 | 21.16 | 26.29 | 107.6056 | 89.6673 | 99.3399 | 0.006565 | 0.9478 | 0.023174 | 24.4501 | 0.058035 | 0.15358 | 0.001838 | 0.12207 | 0.12207 | 0.051964
2019-03-28 15:06 | 1553799950.7526 | 312 | 0.2091 | 0.000678 | 0.000105 | 100.9764 | 100.1522 | 117.4983 | 21.2612 | 0.003583 | 0.003294 | 216.2894 | 0.25748 | 8.4016 | 0.002858 | 0.004095 | 0.00399 | 5.4e-05 | 0.00114 | 0.3178 | 1.285 | 2.5473 | 100.69 | 21.16 | 26.295 | 108.1004 | 90.5872 | 101.0983 | 0.006567 | 0.94334 | 0.023141 | 24.531 | 0.073308 | 0.15318 | 0.002006 | 0.13369 | 0.13369 | 0.047319
2019-03-28 15:06 | 1553799951.7528 | 313 | 0.20916 | 0.000675 | 0.000103 | 100.9624 | 100.3006 | 117.5988 | 21.2612 | 0.003511 | 0.003203 | 216.2843 | 0.25929 | 8.4046 | 0.002803 | 0.004892 | 0.004798 | 5.2e-05 | 0.001144 | 0.3115 | 1.3023 | 2.5474 | 100.69 | 21.16 | 26.298 | 106.3505 | 89.2493 | 98.9118 | 0.006567 | 0.94889 | 0.023219 | 24.4697 | 0.058819 | 0.15425 | 0.001683 | 0.11165 | 0.11165 | 0.057058
2019-03-28 15:06 | 1553799952.752 | 314 | 0.20914 | 0.000671 | 0.000101 | 100.9573 | 100.4443 | 117.3495 | 21.2612 | 0.003518 | 0.003182 | 216.2437 | 0.25733 | 8.4013 | 0.002783 | 0.004223 | 0.004228 | 5.4e-05 | 0.001144 | 0.3115 | 1.2885 | 2.5474 | 100.69 | 21.16 | 26.298 | 106.5244 | 88.7619 | 98.3926 | 0.006567 | 0.95022 | 0.023211 | 24.4267 | 0.045564 | 0.15308 | 0.001771 | 0.11782 | 0.11782 | 0.053662
2019-03-28 15:06 | 1553799953.7528 | 315 | 0.20912 | 0.000665 | 0.0001 | 100.9895 | 100.2837 | 117.5103 | 21.2612 | 0.003497 | 0.003227 | 216.2456 | 0.25814 | 8.401 | 0.002823 | 0.003724 | 0.003657 | 5.1e-05 | 0.001144 | 0.32117 | 1.3158 | 2.5475 | 100.69 | 21.16 | 26.303 | 106.0169 | 89.7439 | 99.4716 | 0.006569 | 0.94746 | 0.023193 | 24.4789 | 0.062735 | 0.15356 | 0.001913 | 0.12777 | 0.12777 | 0.049639
2019-03-28 15:06 | 1553799954.7518 | 316 | 0.20914 | 0.000664 | 9.9e-05 | 100.9478 | 100.2741 | 118.0088 | 21.2612 | 0.003457 | 0.00312 | 216.2527 | 0.25822 | 8.4062 | 0.002835 | 0.004424 | 0.004356 | 5.3e-05 | 0.001144 | 0.29878 | 1.3435 | 2.5475 | 100.69 | 21.16 | 26.303 | 105.0477 | 90.0151 | 96.8831 | 0.006569 | 0.95409 | 0.023323 | 24.4454 | 0.059889 | 0.15362 | 0.001813 | 0.12151 | 0.12151 | 0.052214
2019-03-28 15:06 | 1553799955.7519 | 317 | 0.20917 | 0.000659 | 9.8e-05 | 101.0366 | 100.3224 | 117.8997 | 21.2612 | 0.003456 | 0.003238 | 216.259 | 0.25775 | 8.4023 | 0.002871 | 0.004011 | 0.003897 | 5.2e-05 | 0.001145 | 0.2959 | 1.2804 | 2.5474 | 100.69 | 21.16 | 26.308 | 105.0295 | 90.8926 | 99.739 | 0.00657 | 0.94678 | 0.023223 | 24.5282 | 0.063448 | 0.15334 | 0.001659 | 0.11027 | 0.11027 | 0.05743
2019-03-28 15:06 | 1553799956.753 | 318 | 0.20909 | 0.000658 | 9.7e-05 | 101.0157 | 100.406 | 117.7722 | 21.2612 | 0.003436 | 0.003112 | 216.2616 | 0.2586 | 8.4014 | 0.002824 | 0.004165 | 0.00409 | 5.2e-05 | 0.001146 | 0.32101 | 1.277 | 2.5474 | 100.69 | 21.16 | 26.309 | 104.5341 | 89.7557 | 96.6974 | 0.00657 | 0.95457 | 0.023306 | 24.4148 | 0.054154 | 0.15384 | 0.00213 | 0.14379 | 0.14379 | 0.044186
2019-03-28 15:06 | 1553799957.7519 | 319 | 0.20889 | 0.000655 | 9.6e-05 | 101.128 | 100.5187 | 117.9659 | 21.2612 | 0.003474 | 0.003169 | 216.245 | 0.25718 | 8.4031 | 0.002816 | 0.004111 | 0.004082 | 5.3e-05 | 0.001144 | 0.29159 | 1.2795 | 2.5473 | 100.69 | 21.16 | 26.309 | 105.4606 | 89.5556 | 98.0735 | 0.00657 | 0.95103 | 0.023281 | 24.4802 | 0.05405 | 0.153 | 0.003325 | 0.22744 | 0.22744 | 0.027783
2019-03-28 15:06 | 1553799958.7519 | 320 | 0.20922 | 0.00065 | 9.5e-05 | 101.0967 | 100.4556 | 117.7882 | 21.2612 | 0.003425 | 0.003126 | 216.261 | 0.25853 | 8.4028 | 0.002824 | 0.004424 | 0.004375 | 5.4e-05 | 0.001145 | 0.33964 | 1.2966 | 2.5474 | 100.69 | 21.16 | 26.31 | 104.2643 | 89.7487 | 97.05 | 0.00657 | 0.95366 | 0.023296 | 24.428 | 0.056908 | 0.1538 | 0.001351 | 0.08923 | 0.08923 | 0.071184
2019-03-28 15:06 | 1553799959.752 | 321 | 0.20914 | 0.000653 | 9.3e-05 | 101.131 | 100.5079 | 117.976 | 21.2612 | 0.003438 | 0.003169 | 216.2299 | 0.25803 | 8.4025 | 0.002799 | 0.004734 | 0.004587 | 5.4e-05 | 0.001147 | 0.35673 | 1.2936 | 2.5474 | 100.69 | 21.16 | 26.31 | 104.5845 | 89.1595 | 98.0877 | 0.00657 | 0.951 | 0.023282 | 24.4817 | 0.05528 | 0.1535 | 0.001811 | 0.12148 | 0.12148 | 0.052186
2019-03-28 15:06 | 1553799960.752 | 322 | 0.20914 | 0.000645 | 9.3e-05 | 101.1676 | 100.6332 | 118.16 | 21.2612 | 0.003428 | 0.003176 | 216.2199 | 0.25885 | 8.4016 | 0.002846 | 0.00412 | 0.004039 | 5.6e-05 | 0.001151 | 0.29358 | 1.2825 | 2.5473 | 100.69 | 21.16 | 26.31 | 104.3407 | 90.2824 | 98.2452 | 0.00657 | 0.95059 | 0.023295 | 24.506 | 0.047371 | 0.15399 | 0.001832 | 0.12306 | 0.12306 | 0.051681
2019-03-28 15:06 | 1553799961.7514 | 323 | 0.2092 | 0.000643 | 9.1e-05 | 101.0923 | 100.3217 | 117.8523 | 21.2612 | 0.003463 | 0.00317 | 216.1923 | 0.25735 | 8.4024 | 0.00283 | 0.004129 | 0.004105 | 5.3e-05 | 0.001154 | 0.32045 | 1.2798 | 2.5473 | 100.69 | 21.16 | 26.31 | 105.1932 | 89.912 | 98.0986 | 0.00657 | 0.95097 | 0.023269 | 24.4693 | 0.068438 | 0.1531 | 0.001509 | 0.10035 | 0.10035 | 0.063008
2019-03-28 15:06 | 1553799962.7521 | 324 | 0.20915 | 0.000642 | 9.1e-05 | 101.0964 | 100.5225 | 118.1422 | 21.2612 | 0.003381 | 0.003117 | 216.2105 | 0.25848 | 8.4043 | 0.002782 | 0.004462 | 0.004398 | 5.2e-05 | 0.001156 | 0.32925 | 1.3354 | 2.5474 | 100.69 | 21.16 | 26.31 | 103.2056 | 88.7369 | 96.829 | 0.00657 | 0.95423 | 0.023338 | 24.4574 | 0.050924 | 0.15377 | 0.001767 | 0.11887 | 0.11887 | 0.053427
2019-03-28 15:06 | 1553799963.7516 | 325 | 0.20913 | 0.000643 | 9e-05 | 101.044 | 100.41 | 118.329 | 21.2612 | 0.003405 | 0.003133 | 216.2125 | 0.25853 | 8.4022 | 0.002806 | 0.003834 | 0.003709 | 5.4e-05 | 0.001163 | 0.33732 | 1.3162 | 2.5475 | 100.69 | 21.16 | 26.31 | 103.7942 | 89.3126 | 97.2186 | 0.00657 | 0.95323 | 0.023344 | 24.4896 | 0.056296 | 0.1538 | 0.001889 | 0.1275 | 0.1275 | 0.049818
2019-03-28 15:06 | 1553799964.7511 | 326 | 0.20914 | 0.000638 | 9e-05 | 101.1024 | 100.3747 | 117.7408 | 21.2612 | 0.003368 | 0.003089 | 216.2105 | 0.25884 | 8.4032 | 0.002868 | 0.004322 | 0.004261 | 5.5e-05 | 0.001163 | 0.34092 | 1.3288 | 2.5474 | 100.69 | 21.16 | 26.312 | 102.8913 | 90.809 | 96.1554 | 0.006571 | 0.95597 | 0.02332 | 24.3936 | 0.06461 | 0.15398 | 0.001881 | 0.12679 | 0.12679 | 0.050159
2019-03-28 15:06 | 1553799965.7511 | 327 | 0.20913 | 0.000642 | 9e-05 | 101.1015 | 100.3679 | 117.8347 | 21.2612 | 0.003354 | 0.003158 | 216.1918 | 0.25849 | 8.4014 | 0.00288 | 0.004098 | 0.004044 | 5.8e-05 | 0.001161 | 0.323 | 1.3031 | 2.5475 | 100.69 | 21.16 | 26.313 | 102.5644 | 91.1092 | 97.822 | 0.006571 | 0.95168 | 0.023276 | 24.4583 | 0.065137 | 0.15378 | 0.001921 | 0.12937 | 0.12937 | 0.049092
2019-03-28 15:06 | 1553799966.7521 | 328 | 0.20914 | 0.000639 | 8.9e-05 | 101.0309 | 100.3869 | 118.003 | 21.2612 | 0.003334 | 0.003143 | 216.1904 | 0.25864 | 8.4064 | 0.002911 | 0.00445 | 0.004332 | 5e-05 | 0.001158 | 0.34811 | 1.3144 | 2.5473 | 100.69 | 21.16 | 26.314 | 102.0744 | 91.8543 | 97.448 | 0.006571 | 0.95264 | 0.023305 | 24.4634 | 0.057194 | 0.15387 | 0.001832 | 0.1233 | 0.1233 | 0.051538
2019-03-28 15:06 | 1553799967.7522 | 329 | 0.20914 | 0.000634 | 8.8e-05 | 101.0288 | 100.3322 | 117.9281 | 21.2612 | 0.003313 | 0.003082 | 216.1953 | 0.25882 | 8.4049 | 0.00286 | 0.004523 | 0.00445 | 5.7e-05 | 0.001155 | 0.31614 | 1.2941 | 2.5474 | 100.69 | 21.16 | 26.314 | 101.5639 | 90.6231 | 95.9838 | 0.006571 | 0.95642 | 0.023344 | 24.4073 | 0.061884 | 0.15398 | 0.001851 | 0.12494 | 0.12494 | 0.050896
2019-03-28 15:06 | 1553799968.7522 | 330 | 0.20918 | 0.000637 | 8.8e-05 | 101.013 | 100.2996 | 117.868 | 21.2612 | 0.003296 | 0.003105 | 216.1707 | 0.25849 | 8.4029 | 0.002894 | 0.004088 | 0.003974 | 5.5e-05 | 0.001153 | 0.32676 | 1.3219 | 2.5474 | 100.69 | 21.16 | 26.315 | 101.1494 | 91.4478 | 96.5407 | 0.006572 | 0.95498 | 0.02332 | 24.4195 | 0.063395 | 0.15378 | 0.001595 | 0.10682 | 0.10682 | 0.059454
2019-03-28 15:06 | 1553799969.7513 | 331 | 0.20917 | 0.000634 | 8.8e-05 | 100.9999 | 100.3886 | 117.8995 | 21.2612 | 0.003302 | 0.003107 | 216.1651 | 0.25811 | 8.4062 | 0.002817 | 0.004158 | 0.004065 | 5.5e-05 | 0.001154 | 0.33124 | 1.3213 | 2.5474 | 100.69 | 21.16 | 26.315 | 101.3009 | 89.5946 | 96.5847 | 0.006572 | 0.95486 | 0.023322 | 24.4242 | 0.0543 | 0.15355 | 0.001689 | 0.11341 | 0.11341 | 0.055917
2019-03-28 15:06 | 1553799970.7524 | 332 | 0.20916 | 0.000636 | 8.7e-05 | 100.9737 | 100.549 | 117.626 | 21.2612 | 0.00336 | 0.003091 | 216.1504 | 0.25873 | 8.402 | 0.002861 | 0.004353 | 0.00436 | 5.4e-05 | 0.00115 | 0.33539 | 1.2941 | 2.5474 | 100.69 | 21.16 | 26.314 | 102.7058 | 90.6394 | 96.2032 | 0.006571 | 0.95585 | 0.023307 | 24.3833 | 0.037703 | 0.15392 | 0.00172 | 0.11563 | 0.11563 | 0.054976
2019-03-28 15:06 | 1553799971.7513 | 333 | 0.20917 | 0.000637 | 8.6e-05 | 100.9165 | 100.2643 | 117.7453 | 21.2612 | 0.003325 | 0.003027 | 216.1705 | 0.2591 | 8.4047 | 0.002885 | 0.004293 | 0.004151 | 5.8e-05 | 0.00115 | 0.33179 | 1.2912 | 2.5475 | 100.69 | 21.17 | 26.313 | 101.8449 | 91.2343 | 94.6488 | 0.006575 | 0.95989 | 0.023368 | 24.3442 | 0.057997 | 0.15414 | 0.00168 | 0.11311 | 0.11311 | 0.05628
2019-03-28 15:06 | 1553799972.7514 | 334 | 0.20917 | 0.000634 | 8.5e-05 | 100.8263 | 100.2989 | 117.6586 | 21.2612 | 0.003309 | 0.003062 | 216.1487 | 0.25901 | 8.4052 | 0.002851 | 0.004729 | 0.004624 | 5.4e-05 | 0.001158 | 0.32925 | 1.2693 | 2.5474 | 100.69 | 21.17 | 26.312 | 101.4736 | 90.419 | 95.4943 | 0.006575 | 0.95769 | 0.023332 | 24.3632 | 0.046903 | 0.15409 | 0.001666 | 0.11202 | 0.11202 | 0.056807
2019-03-28 15:06 | 1553799973.7524 | 335 | 0.20918 | 0.000632 | 8.4e-05 | 100.8519 | 100.3248 | 117.4373 | 21.2612 | 0.003266 | 0.003047 | 216.1411 | 0.25813 | 8.4022 | 0.002878 | 0.004651 | 0.004547 | 5.2e-05 | 0.001156 | 0.32156 | 1.3007 | 2.5473 | 100.69 | 21.17 | 26.307 | 100.4241 | 91.05 | 95.1405 | 0.006574 | 0.95861 | 0.023322 | 24.3286 | 0.046873 | 0.15356 | 0.001612 | 0.10828 | 0.10828 | 0.058572
2019-03-28 15:06 | 1553799974.7515 | 336 | 0.2092 | 0.000632 | 8.3e-05 | 100.7712 | 100.195 | 117.568 | 21.2612 | 0.003296 | 0.00305 | 216.1322 | 0.25828 | 8.4026 | 0.002793 | 0.004207 | 0.004172 | 5.7e-05 | 0.001149 | 0.32173 | 1.3482 | 2.5474 | 100.69 | 21.17 | 26.307 | 101.1375 | 89.0049 | 95.1929 | 0.006574 | 0.95847 | 0.023333 | 24.3439 | 0.051289 | 0.15365 | 0.001498 | 0.10038 | 0.10038 | 0.063221
2019-03-28 15:06 | 1553799975.7509 | 337 | 0.20921 | 0.000629 | 8.3e-05 | 100.8057 | 100.2327 | 117.6255 | 21.2612 | 0.003276 | 0.003018 | 216.1365 | 0.25867 | 8.4019 | 0.002852 | 0.004156 | 0.004041 | 5.6e-05 | 0.001145 | 0.34164 | 1.3723 | 2.5475 | 100.69 | 21.17 | 26.299 | 100.6721 | 90.4395 | 94.443 | 0.006572 | 0.96042 | 0.023362 | 24.3251 | 0.05099 | 0.15388 | 0.001462 | 0.098019 | 0.098019 | 0.064838
2019-03-28 15:06 | 1553799976.7515 | 338 | 0.20918 | 0.000627 | 8.2e-05 | 100.733 | 100.2021 | 117.7645 | 21.2612 | 0.003243 | 0.003057 | 216.1224 | 0.25901 | 8.4006 | 0.002835 | 0.004193 | 0.004105 | 5.4e-05 | 0.001141 | 0.32045 | 1.3441 | 2.5474 | 100.69 | 21.17 | 26.299 | 99.8575 | 90.0165 | 95.3637 | 0.006572 | 0.95803 | 0.023347 | 24.3699 | 0.047268 | 0.15409 | 0.001614 | 0.10866 | 0.10866 | 0.058565
2019-03-28 15:06 | 1553799977.7506 | 339 | 0.20921 | 0.000622 | 8.1e-05 | 100.6479 | 100.1393 | 117.8673 | 21.2612 | 0.003295 | 0.003053 | 216.1172 | 0.25784 | 8.4028 | 0.002806 | 0.004435 | 0.004428 | 5.8e-05 | 0.001146 | 0.34164 | 1.369 | 2.5474 | 100.69 | 21.17 | 26.284 | 101.1141 | 89.315 | 95.2727 | 0.006568 | 0.95826 | 0.02336 | 24.3775 | 0.045312 | 0.15339 | 0.001475 | 0.098973 | 0.098973 | 0.064006
2019-03-28 15:06 | 1553799978.7516 | 340 | 0.20919 | 0.000624 | 8e-05 | 100.6184 | 100.1464 | 117.738 | 21.2612 | 0.003241 | 0.003043 | 216.106 | 0.25865 | 8.4026 | 0.002763 | 0.004128 | 0.004069 | 5.5e-05 | 0.001145 | 0.3194 | 1.314 | 2.5474 | 100.69 | 21.17 | 26.284 | 99.8106 | 88.298 | 95.0387 | 0.006568 | 0.95887 | 0.023355 | 24.3564 | 0.042054 | 0.15387 | 0.001575 | 0.10605 | 0.10605 | 0.059926
2019-03-28 15:06 | 1553799979.7516 | 341 | 0.20923 | 0.000617 | 8e-05 | 100.5978 | 99.8495 | 117.8025 | 21.2612 | 0.003282 | 0.003032 | 216.1043 | 0.25843 | 8.402 | 0.002815 | 0.004522 | 0.004422 | 5.2e-05 | 0.001143 | 0.31924 | 1.3131 | 2.5472 | 100.69 | 21.17 | 26.273 | 100.815 | 89.5329 | 94.7788 | 0.006565 | 0.95955 | 0.023369 | 24.3545 | 0.066786 | 0.15374 | 0.001371 | 0.091781 | 0.091781 | 0.069182
2019-03-28 15:06 | 1553799980.7516 | 342 | 0.20918 | 0.000617 | 7.9e-05 | 100.5707 | 99.9844 | 117.9553 | 21.2612 | 0.003258 | 0.003023 | 216.1202 | 0.25864 | 8.4002 | 0.002791 | 0.004352 | 0.004255 | 5.5e-05 | 0.001147 | 0.32421 | 1.3396 | 2.5475 | 100.69 | 21.17 | 26.269 | 100.2241 | 88.9565 | 94.5566 | 0.006564 | 0.96013 | 0.023392 | 24.3629 | 0.052302 | 0.15386 | 0.001666 | 0.11274 | 0.11274 | 0.056364
2019-03-28 15:06 | 1553799981.7506 | 343 | 0.20925 | 0.00062 | 7.8e-05 | 100.4939 | 99.9759 | 117.9584 | 21.2612 | 0.003243 | 0.003035 | 216.0972 | 0.25823 | 8.4005 | 0.002816 | 0.004374 | 0.00433 | 5.1e-05 | 0.001151 | 0.32853 | 1.3522 | 2.5474 | 100.69 | 21.17 | 26.263 | 99.8616 | 89.565 | 94.8548 | 0.006563 | 0.95935 | 0.023382 | 24.3731 | 0.046224 | 0.15362 | 0.001244 | 0.083033 | 0.083033 | 0.07641
2019-03-28 15:06 | 1553799982.7507 | 344 | 0.20922 | 0.000614 | 7.6e-05 | 100.4054 | 99.6725 | 117.9258 | 21.2612 | 0.003195 | 0.002986 | 216.0945 | 0.25817 | 8.4002 | 0.00275 | 0.004591 | 0.004539 | 5.3e-05 | 0.001153 | 0.32549 | 1.2992 | 2.5476 | 100.69 | 21.17 | 26.263 | 98.7109 | 87.9745 | 93.6502 | 0.006563 | 0.9625 | 0.023417 | 24.3298 | 0.065531 | 0.15358 | 0.001434 | 0.096646 | 0.096646 | 0.065631
2019-03-28 15:06 | 1553799983.7507 | 345 | 0.20923 | 0.000612 | 7.5e-05 | 100.3158 | 99.7034 | 117.8124 | 21.2612 | 0.003215 | 0.003005 | 216.0617 | 0.25873 | 8.4026 | 0.002777 | 0.005126 | 0.004994 | 5.5e-05 | 0.001146 | 0.32444 | 1.2978 | 2.5473 | 100.69 | 21.17 | 26.259 | 99.1885 | 88.6206 | 94.1111 | 0.006562 | 0.96129 | 0.023392 | 24.3334 | 0.054767 | 0.15392 | 0.001392 | 0.093608 | 0.093608 | 0.06791
2019-03-28 15:06 | 1553799984.7508 | 346 | 0.20921 | 0.000609 | 7.5e-05 | 100.2428 | 99.3929 | 117.9917 | 21.2612 | 0.003203 | 0.003002 | 216.0728 | 0.25812 | 8.4027 | 0.002798 | 0.003751 | 0.003672 | 5e-05 | 0.001152 | 0.31277 | 1.2845 | 2.5472 | 100.69 | 21.17 | 26.255 | 98.8992 | 89.1229 | 94.0455 | 0.006561 | 0.96146 | 0.023411 | 24.3497 | 0.076157 | 0.15356 | 0.001518 | 0.10259 | 0.10259 | 0.06182
2019-03-28 15:06 | 1553799985.7518 | 347 | 0.20918 | 0.000613 | 7.4e-05 | 100.2064 | 99.4636 | 117.8274 | 21.2612 | 0.003202 | 0.002968 | 216.0524 | 0.2587 | 8.4021 | 0.002788 | 0.004166 | 0.004063 | 5e-05 | 0.001153 | 0.34882 | 1.2884 | 2.5474 | 100.69 | 21.17 | 26.253 | 98.8818 | 88.8946 | 93.2349 | 0.00656 | 0.96359 | 0.023421 | 24.3059 | 0.066557 | 0.1539 | 0.001671 | 0.11349 | 0.11349 | 0.056005
2019-03-28 15:06 | 1553799986.7507 | 348 | 0.20921 | 0.000608 | 7.4e-05 | 100.1329 | 99.4765 | 117.8633 | 21.2612 | 0.003232 | 0.002992 | 216.0705 | 0.2592 | 8.4024 | 0.00274 | 0.005108 | 0.005073 | 5.6e-05 | 0.001152 | 0.31669 | 1.3058 | 2.5475 | 100.69 | 21.17 | 26.253 | 99.5971 | 87.741 | 93.7958 | 0.00656 | 0.96212 | 0.023407 | 24.3282 | 0.058827 | 0.1542 | 0.001481 | 0.10002 | 0.10002 | 0.06367
2019-03-28 15:06 | 1553799987.7517 | 349 | 0.20923 | 0.000604 | 7.3e-05 | 100.0438 | 99.273 | 117.9988 | 21.2612 | 0.003175 | 0.002974 | 216.0549 | 0.25856 | 8.4039 | 0.002797 | 0.004399 | 0.004276 | 5.4e-05 | 0.001152 | 0.32676 | 1.2982 | 2.5474 | 100.69 | 21.17 | 26.251 | 98.224 | 89.1079 | 93.3695 | 0.00656 | 0.96324 | 0.023434 | 24.328 | 0.069185 | 0.15382 | 0.001389 | 0.093711 | 0.093711 | 0.067789
2019-03-28 15:06 | 1553799988.7509 | 350 | 0.20922 | 0.000608 | 7.2e-05 | 99.9956 | 99.3735 | 118.2612 | 21.2612 | 0.003175 | 0.00299 | 216.0821 | 0.25782 | 8.4017 | 0.002735 | 0.004567 | 0.004558 | 5.3e-05 | 0.001152 | 0.33052 | 1.2964 | 2.5474 | 100.69 | 21.17 | 26.25 | 98.2327 | 87.6084 | 93.7519 | 0.006559 | 0.96223 | 0.023447 | 24.3678 | 0.055814 | 0.15338 | 0.001446 | 0.09786 | 0.09786 | 0.064731
2019-03-28 15:06 | 1553799989.7512 | 351 | 0.20919 | 0.000603 | 7.1e-05 | 99.8934 | 99.185 | 118.144 | 21.2612 | 0.003143 | 0.003047 | 216.0597 | 0.25854 | 8.4036 | 0.00278 | 0.004088 | 0.004044 | 5.2e-05 | 0.001156 | 0.32604 | 1.2904 | 2.5475 | 100.69 | 21.17 | 26.247 | 97.4599 | 88.6986 | 95.1342 | 0.006559 | 0.95862 | 0.023392 | 24.4015 | 0.063658 | 0.1538 | 0.001611 | 0.1093 | 0.1093 | 0.058115
2019-03-28 15:06 | 1553799990.7499 | 352 | 0.20924 | 0.0006 | 7.1e-05 | 99.8818 | 99.1948 | 118.1898 | 21.2612 | 0.003167 | 0.002999 | 216.051 | 0.25867 | 8.4032 | 0.002793 | 0.004404 | 0.004318 | 5.1e-05 | 0.001156 | 0.3501 | 1.3107 | 2.5475 | 100.69 | 21.17 | 26.247 | 98.0249 | 89.015 | 93.9774 | 0.006559 | 0.96164 | 0.023433 | 24.3679 | 0.061729 | 0.15388 | 0.001331 | 0.089731 | 0.089731 | 0.070827
2019-03-28 15:06 | 1553799991.7502 | 353 | 0.20922 | 0.000599 | 7e-05 | 99.7415 | 99.2193 | 118.4769 | 21.2612 | 0.003209 | 0.002981 | 216.0457 | 0.25862 | 8.4066 | 0.002751 | 0.004821 | 0.004794 | 5.1e-05 | 0.001156 | 0.31564 | 1.3226 | 2.5473 | 100.69 | 21.17 | 26.241 | 99.0568 | 88.0051 | 93.5337 | 0.006557 | 0.96281 | 0.023476 | 24.3827 | 0.046952 | 0.15385 | 0.001433 | 0.097183 | 0.097183 | 0.065384
2019-03-28 15:07 | 1553799992.75 | 354 | 0.20929 | 0.000598 | 7e-05 | 99.5944 | 99.0396 | 118.5595 | 21.2612 | 0.003173 | 0.002968 | 216.043 | 0.25795 | 8.3987 | 0.002674 | 0.003736 | 0.003667 | 5.2e-05 | 0.001154 | 0.31094 | 1.3235 | 2.5473 | 100.69 | 21.17 | 26.241 | 98.1677 | 86.1499 | 93.2308 | 0.006557 | 0.9636 | 0.023494 | 24.3811 | 0.049963 | 0.15345 | 0.001033 | 0.068958 | 0.068958 | 0.091906
2019-03-28 15:07 | 1553799993.7511 | 355 | 0.20924 | 0.000594 | 7e-05 | 99.5964 | 98.9125 | 118.8236 | 21.2612 | 0.00311 | 0.002986 | 216.0599 | 0.25759 | 8.4007 | 0.002758 | 0.004761 | 0.004573 | 5.1e-05 | 0.001153 | 0.32444 | 1.3183 | 2.5473 | 100.69 | 21.17 | 26.234 | 96.6605 | 88.1613 | 93.6526 | 0.006555 | 0.96249 | 0.023506 | 24.4223 | 0.061633 | 0.15324 | 0.001347 | 0.091212 | 0.091212 | 0.069386
2019-03-28 15:07 | 1553799994.7512 | 356 | 0.20924 | 0.000595 | 7e-05 | 99.4922 | 98.8385 | 118.905 | 21.2612 | 0.003147 | 0.002997 | 216.0326 | 0.25821 | 8.4025 | 0.002726 | 0.003813 | 0.003832 | 5e-05 | 0.001148 | 0.30198 | 1.3244 | 2.5475 | 100.69 | 21.17 | 26.23 | 97.5492 | 87.3997 | 93.9335 | 0.006554 | 0.96176 | 0.023505 | 24.44 | 0.058965 | 0.15361 | 0.001364 | 0.092422 | 0.092422 | 0.068641
2019-03-28 15:07 | 1553799995.7511 | 357 | 0.20924 | 0.000596 | 6.9e-05 | 99.452 | 98.5776 | 119.0095 | 21.2612 | 0.003131 | 0.002952 | 216.0446 | 0.25827 | 8.4049 | 0.002751 | 0.004376 | 0.004319 | 5e-05 | 0.001144 | 0.28711 | 1.3211 | 2.5475 | 100.69 | 21.17 | 26.221 | 97.155 | 87.9988 | 92.8477 | 0.006552 | 0.96461 | 0.023551 | 24.4146 | 0.078992 | 0.15364 | 0.001346 | 0.091368 | 0.091368 | 0.069449
2019-03-28 15:07 | 1553799996.7511 | 358 | 0.20924 | 0.000589 | 6.9e-05 | 99.2113 | 98.496 | 119.0806 | 21.2612 | 0.003169 | 0.002947 | 216.0166 | 0.25893 | 8.4039 | 0.002671 | 0.004458 | 0.004393 | 5.1e-05 | 0.001143 | 0.30663 | 1.3214 | 2.5474 | 100.69 | 21.17 | 26.218 | 98.0797 | 86.0734 | 92.7257 | 0.006551 | 0.96493 | 0.023562 | 24.4178 | 0.064722 | 0.15404 | 0.001317 | 0.089355 | 0.089355 | 0.071196
2019-03-28 15:07 | 1553799997.7512 | 359 | 0.20925 | 0.000588 | 6.7e-05 | 99.1994 | 98.2409 | 118.887 | 21.2612 | 0.003082 | 0.002973 | 216.0282 | 0.25846 | 8.4045 | 0.002744 | 0.00412 | 0.004142 | 5.1e-05 | 0.001135 | 0.32532 | 1.3176 | 2.5473 | 100.69 | 21.17 | 26.214 | 95.9753 | 87.8387 | 93.3584 | 0.00655 | 0.96327 | 0.023522 | 24.419 | 0.086847 | 0.15376 | 0.001274 | 0.086255 | 0.086255 | 0.073622
2019-03-28 15:07 | 1553799998.7502 | 360 | 0.20926 | 0.000592 | 6.7e-05 | 99.0329 | 98.3615 | 118.6046 | 21.2612 | 0.003116 | 0.002938 | 216.0114 | 0.25819 | 8.4022 | 0.002648 | 0.003769 | 0.003679 | 4.8e-05 | 0.001139 | 0.34235 | 1.3025 | 2.5475 | 100.69 | 21.17 | 26.207 | 96.7969 | 85.524 | 92.5121 | 0.006549 | 0.9655 | 0.023521 | 24.3618 | 0.060848 | 0.1536 | 0.001208 | 0.081564 | 0.081564 | 0.077775
2019-03-28 15:07 | 1553799999.7515 | 361 | 0.20924 | 0.00059 | 6.7e-05 | 98.9322 | 98.0228 | 118.8773 | 21.2612 | 0.003139 | 0.002974 | 216.0434 | 0.25913 | 8.4004 | 0.002739 | 0.00397 | 0.003861 | 5e-05 | 0.001134 | 0.29048 | 1.2962 | 2.5473 | 100.69 | 21.17 | 26.205 | 97.3446 | 87.7184 | 93.3716 | 0.006548 | 0.96323 | 0.023521 | 24.4185 | 0.082604 | 0.15416 | 0.001333 | 0.090444 | 0.090444 | 0.070393
2019-03-28 15:07 | 1553800000.7512 | 362 | 0.20927 | 0.000588 | 6.7e-05 | 98.7714 | 98.0578 | 118.9456 | 21.2612 | 0.003151 | 0.002937 | 216.0357 | 0.25783 | 8.4022 | 0.002696 | 0.003423 | 0.003322 | 4.9e-05 | 0.001133 | 0.30646 | 1.2901 | 2.5473 | 100.69 | 21.17 | 26.201 | 97.6533 | 86.6699 | 92.4738 | 0.006547 | 0.9656 | 0.023556 | 24.3956 | 0.064858 | 0.15338 | 0.00118 | 0.079737 | 0.079737 | 0.079445
2019-03-28 15:07 | 1553800001.7503 | 363 | 0.20926 | 0.000585 | 6.6e-05 | 98.7008 | 97.7826 | 118.5435 | 21.2612 | 0.003088 | 0.002984 | 216.037 | 0.25838 | 8.4025 | 0.002639 | 0.004058 | 0.00391 | 4.8e-05 | 0.001135 | 0.35402 | 1.3341 | 2.5474 | 100.69 | 21.17 | 26.201 | 96.1146 | 85.2931 | 93.6041 | 0.006547 | 0.96262 | 0.02348 | 24.3919 | 0.083599 | 0.15371 | 0.00123 | 0.08303 | 0.08303 | 0.076458
2019-03-28 15:07 | 1553800002.7514 | 364 | 0.20927 | 0.000587 | 6.6e-05 | 98.5453 | 97.697 | 118.5217 | 21.2612 | 0.003077 | 0.002922 | 216.0327 | 0.25897 | 8.4016 | 0.002665 | 0.003473 | 0.003432 | 5.2e-05 | 0.001134 | 0.32964 | 1.3342 | 2.5473 | 100.69 | 21.17 | 26.198 | 95.8609 | 85.9322 | 92.1076 | 0.006546 | 0.96657 | 0.023526 | 24.3399 | 0.077333 | 0.15406 | 0.001145 | 0.077223 | 0.077223 | 0.082396
2019-03-28 15:07 | 1553800003.7517 | 365 | 0.20931 | 0.000585 | 6.5e-05 | 98.4738 | 97.5781 | 118.536 | 21.2612 | 0.003142 | 0.002933 | 216.0122 | 0.25769 | 8.4028 | 0.002662 | 0.00446 | 0.00436 | 5.2e-05 | 0.001134 | 0.32388 | 1.3131 | 2.5475 | 100.69 | 21.17 | 26.198 | 97.434 | 85.8665 | 92.3909 | 0.006546 | 0.96582 | 0.023518 | 24.3508 | 0.081735 | 0.1533 | 0.000955 | 0.063818 | 0.063818 | 0.099207
2019-03-28 15:07 | 1553800004.7496 | 366 | 0.20929 | 0.000581 | 6.4e-05 | 98.2987 | 97.4928 | 118.3231 | 21.2612 | 0.00312 | 0.002931 | 216.0228 | 0.25906 | 8.4028 | 0.002637 | 0.004417 | 0.004369 | 5.2e-05 | 0.001136 | 0.36282 | 1.3317 | 2.5474 | 100.69 | 21.17 | 26.199 | 96.902 | 85.2594 | 92.3414 | 0.006547 | 0.96595 | 0.023499 | 24.3272 | 0.073633 | 0.15412 | 0.001082 | 0.072787 | 0.072787 | 0.087447
2019-03-28 15:07 | 1553800005.7508 | 367 | 0.20929 | 0.000582 | 6.4e-05 | 98.1915 | 97.2971 | 118.4494 | 21.2612 | 0.003084 | 0.002913 | 216.0088 | 0.25841 | 8.4022 | 0.002683 | 0.004275 | 0.004166 | 5e-05 | 0.001134 | 0.30342 | 1.3109 | 2.5473 | 100.69 | 21.17 | 26.199 | 96.0333 | 86.36 | 91.9076 | 0.006547 | 0.9671 | 0.023525 | 24.3258 | 0.081841 | 0.15373 | 0.001071 | 0.072059 | 0.072059 | 0.088107
2019-03-28 15:07 | 1553800006.7494 | 368 | 0.2093 | 0.00058 | 6.3e-05 | 97.9711 | 97.1208 | 118.2631 | 21.2612 | 0.003114 | 0.002944 | 216.0023 | 0.25961 | 8.4061 | 0.002635 | 0.004676 | 0.004573 | 5.6e-05 | 0.001137 | 0.30757 | 1.3298 | 2.5474 | 100.69 | 21.17 | 26.207 | 96.7498 | 85.2009 | 92.6543 | 0.006549 | 0.96512 | 0.023483 | 24.3315 | 0.077965 | 0.15444 | 0.000991 | 0.066385 | 0.066385 | 0.096083
2019-03-28 15:07 | 1553800007.7505 | 369 | 0.2093 | 0.000582 | 6.3e-05 | 97.8144 | 96.8422 | 118.3546 | 21.2612 | 0.003095 | 0.002939 | 215.9803 | 0.25897 | 8.4029 | 0.002669 | 0.00364 | 0.003542 | 5.6e-05 | 0.001137 | 0.31332 | 1.3273 | 2.5474 | 100.69 | 21.17 | 26.207 | 96.2971 | 86.0376 | 92.5324 | 0.006549 | 0.96544 | 0.023496 | 24.3368 | 0.089342 | 0.15406 | 0.001031 | 0.069264 | 0.069264 | 0.091861
2019-03-28 15:07 | 1553800008.7495 | 370 | 0.2093 | 0.000581 | 6.2e-05 | 97.6264 | 96.8507 | 118.8421 | 21.2612 | 0.003088 | 0.002903 | 215.9702 | 0.25824 | 8.4004 | 0.002643 | 0.004028 | 0.003942 | 4.9e-05 | 0.001132 | 0.29502 | 1.3335 | 2.5475 | 100.69 | 21.17 | 26.219 | 96.1132 | 85.4116 | 91.6666 | 0.006552 | 0.96773 | 0.023572 | 24.358 | 0.071355 | 0.15363 | 0.001009 | 0.067899 | 0.067899 | 0.093444
2019-03-28 15:07 | 1553800009.7505 | 371 | 0.20931 | 0.000575 | 6.2e-05 | 97.4495 | 96.7002 | 118.8062 | 21.2612 | 0.003098 | 0.002929 | 215.9581 | 0.25862 | 8.4007 | 0.002647 | 0.004035 | 0.003863 | 5.3e-05 | 0.001134 | 0.3449 | 1.3199 | 2.5473 | 100.69 | 21.17 | 26.219 | 96.3567 | 85.4944 | 92.2773 | 0.006552 | 0.96612 | 0.023549 | 24.3747 | 0.06904 | 0.15385 | 0.000968 | 0.064949 | 0.064949 | 0.097831
2019-03-28 15:07 | 1553800010.7506 | 372 | 0.20929 | 0.000575 | 6.1e-05 | 97.2652 | 96.3496 | 118.5623 | 21.2612 | 0.003089 | 0.002924 | 215.9579 | 0.25827 | 8.402 | 0.00266 | 0.003519 | 0.003531 | 5.3e-05 | 0.001132 | 0.3235 | 1.37 | 2.5475 | 100.69 | 21.17 | 26.227 | 96.1404 | 85.8064 | 92.1703 | 0.006554 | 0.9664 | 0.023528 | 24.3461 | 0.08459 | 0.15365 | 0.00106 | 0.0715 | 0.0715 | 0.088749
2019-03-28 15:07 | 1553800011.7506 | 373 | 0.20932 | 0.000578 | 6.1e-05 | 97.0257 | 96.1404 | 118.8353 | 21.2612 | 0.003095 | 0.00289 | 215.9453 | 0.25876 | 8.4044 | 0.002675 | 0.003879 | 0.003816 | 5.1e-05 | 0.001123 | 0.33964 | 1.3258 | 2.5475 | 100.69 | 21.17 | 26.237 | 96.2908 | 86.1617 | 91.3471 | 0.006556 | 0.96858 | 0.023582 | 24.3467 | 0.081989 | 0.15394 | 0.000921 | 0.061755 | 0.061755 | 0.10295
2019-03-28 15:07 | 1553800012.7506 | 374 | 0.20932 | 0.000574 | 6e-05 | 96.8276 | 95.7429 | 119.1602 | 21.2612 | 0.003097 | 0.002911 | 215.9282 | 0.25903 | 8.4036 | 0.002637 | 0.003712 | 0.003607 | 5.3e-05 | 0.00112 | 0.31022 | 1.3435 | 2.5473 | 100.69 | 21.17 | 26.241 | 96.3341 | 85.2517 | 91.8644 | 0.006557 | 0.96721 | 0.023597 | 24.3972 | 0.10077 | 0.1541 | 0.000919 | 0.061781 | 0.061781 | 0.10301
2019-03-28 15:07 | 1553800013.7496 | 375 | 0.2093 | 0.000576 | 6e-05 | 96.6078 | 95.7205 | 119.1789 | 21.2612 | 0.003022 | 0.002887 | 215.9427 | 0.25803 | 8.4014 | 0.00268 | 0.004093 | 0.003996 | 5.4e-05 | 0.001115 | 0.33085 | 1.3219 | 2.5474 | 100.69 | 21.17 | 26.249 | 94.5409 | 86.2974 | 91.2632 | 0.006559 | 0.96881 | 0.023619 | 24.379 | 0.082534 | 0.1535 | 0.001023 | 0.069228 | 0.069228 | 0.091576
2019-03-28 15:07 | 1553800014.7511 | 376 | 0.20932 | 0.00057 | 5.9e-05 | 96.349 | 95.2248 | 119.2367 | 21.2612 | 0.003047 | 0.002905 | 215.9224 | 0.25798 | 8.4028 | 0.002713 | 0.003832 | 0.003828 | 5.6e-05 | 0.001113 | 0.34667 | 1.3675 | 2.5474 | 100.69 | 21.17 | 26.249 | 95.1429 | 87.0867 | 91.7042 | 0.006559 | 0.96763 | 0.02361 | 24.3997 | 0.10497 | 0.15347 | 0.000886 | 0.059467 | 0.059467 | 0.10659
2019-03-28 15:07 | 1553800015.7508 | 377 | 0.2093 | 0.000573 | 5.8e-05 | 96.0222 | 95.0237 | 119.3393 | 21.2612 | 0.003038 | 0.002859 | 215.9284 | 0.25969 | 8.402 | 0.002674 | 0.004455 | 0.004251 | 5.4e-05 | 0.001111 | 0.29878 | 1.3331 | 2.5473 | 100.69 | 21.17 | 26.256 | 94.9127 | 86.1406 | 90.5886 | 0.006561 | 0.9706 | 0.023656 | 24.3728 | 0.093501 | 0.15449 | 0.001009 | 0.068402 | 0.068402 | 0.093278
2019-03-28 15:07 | 1553800016.7498 | 378 | 0.20929 | 0.000568 | 5.8e-05 | 95.8206 | 94.7407 | 119.3435 | 21.2612 | 0.003091 | 0.002952 | 215.931 | 0.25866 | 8.4051 | 0.002655 | 0.003347 | 0.003274 | 5.6e-05 | 0.001115 | 0.33915 | 1.3362 | 2.5473 | 100.69 | 21.17 | 26.256 | 96.2001 | 85.6981 | 92.8491 | 0.006561 | 0.96461 | 0.023584 | 24.4489 | 0.10138 | 0.15388 | 0.001063 | 0.072028 | 0.072028 | 0.088231
2019-03-28 15:07 | 1553800017.7508 | 379 | 0.2093 | 0.000571 | 5.7e-05 | 95.5436 | 94.5457 | 119.3275 | 21.2612 | 0.003089 | 0.002913 | 215.9461 | 0.25912 | 8.4009 | 0.002715 | 0.004567 | 0.004471 | 5.5e-05 | 0.00112 | 0.34274 | 1.2986 | 2.5474 | 100.69 | 21.16 | 26.264 | 96.1429 | 87.1406 | 91.9069 | 0.006559 | 0.9671 | 0.023612 | 24.4158 | 0.09391 | 0.15415 | 0.001039 | 0.070435 | 0.070435 | 0.090389
2019-03-28 15:07 | 1553800018.7509 | 380 | 0.20933 | 0.000572 | 5.7e-05 | 95.2567 | 94.2627 | 119.2163 | 21.2612 | 0.003088 | 0.002891 | 215.9357 | 0.25869 | 8.4014 | 0.002733 | 0.004524 | 0.004465 | 5.5e-05 | 0.001121 | 0.31614 | 1.2924 | 2.5473 | 100.69 | 21.16 | 26.264 | 96.1139 | 87.5608 | 91.3712 | 0.006559 | 0.96852 | 0.023619 | 24.3865 | 0.09383 | 0.1539 | 0.000812 | 0.054351 | 0.054351 | 0.11694
2019-03-28 15:07 | 1553800019.7499 | 381 | 0.20926 | 0.00057 | 5.6e-05 | 94.9994 | 93.8207 | 119.3738 | 21.2612 | 0.003039 | 0.002831 | 215.913 | 0.25963 | 8.4004 | 0.002737 | 0.004091 | 0.003984 | 5.5e-05 | 0.001123 | 0.31996 | 1.2719 | 2.5474 | 100.69 | 21.16 | 26.269 | 94.9385 | 87.6592 | 89.9142 | 0.00656 | 0.9724 | 0.023682 | 24.3537 | 0.11167 | 0.15446 | 0.001255 | 0.086151 | 0.086151 | 0.074044
2019-03-28 15:07 | 1553800020.7499 | 382 | 0.2093 | 0.000564 | 5.6e-05 | 94.7343 | 93.4593 | 119.589 | 21.2612 | 0.003021 | 0.002872 | 215.8827 | 0.25795 | 8.407 | 0.002649 | 0.00409 | 0.00406 | 5.7e-05 | 0.001126 | 0.30359 | 1.2744 | 2.5473 | 100.69 | 21.16 | 26.269 | 94.511 | 85.5536 | 90.9107 | 0.00656 | 0.96974 | 0.023671 | 24.4091 | 0.1212 | 0.15346 | 0.001033 | 0.070312 | 0.070312 | 0.090138
2019-03-28 15:07 | 1553800021.7493 | 383 | 0.20931 | 0.000565 | 5.6e-05 | 94.4466 | 93.1307 | 119.7094 | 21.2612 | 0.003031 | 0.002884 | 215.8996 | 0.25839 | 8.401 | 0.002647 | 0.004484 | 0.004433 | 5.4e-05 | 0.001129 | 0.33339 | 1.2753 | 2.5473 | 100.69 | 21.16 | 26.273 | 94.7467 | 85.5076 | 91.2022 | 0.006561 | 0.96897 | 0.023673 | 24.4312 | 0.12549 | 0.15371 | 0.000962 | 0.065263 | 0.065263 | 0.097275
2019-03-28 15:07 | 1553800022.75 | 384 | 0.20932 | 0.00056 | 5.6e-05 | 94.1723 | 92.8162 | 119.3954 | 21.2612 | 0.003002 | 0.002911 | 215.9181 | 0.25867 | 8.4021 | 0.002684 | 0.004003 | 0.00388 | 5.4e-05 | 0.00113 | 0.34954 | 1.3055 | 2.5474 | 100.69 | 21.16 | 26.273 | 94.0541 | 86.3833 | 91.863 | 0.006561 | 0.96721 | 0.023621 | 24.4212 | 0.12974 | 0.15388 | 0.000927 | 0.0626 | 0.0626 | 0.10152
2019-03-28 15:07 | 1553800023.749 | 385 | 0.20932 | 0.000565 | 5.5e-05 | 93.8765 | 92.6487 | 119.1552 | 21.2612 | 0.002999 | 0.002891 | 208.9504 | 0.2578 | 8.4034 | 0.002642 | 0.004413 | 0.004376 | 5.5e-05 | 0.00113 | 0.32803 | 1.3283 | 2.5473 | 100.69 | 21.16 | 26.274 | 93.9671 | 85.3765 | 91.3728 | 0.006561 | 0.96851 | 0.023613 | 24.3803 | 0.11776 | 0.15337 | 0.000891 | 0.060082 | 0.060082 | 0.10542
2019-03-28 15:07 | 1553800024.7491 | 386 | 0.20932 | 0.000562 | 5.4e-05 | 93.5831 | 92.2121 | 119.559 | 21.2612 | 0.003054 | 0.002889 | 172.2336 | 0.25963 | 8.4024 | 0.002711 | 0.003491 | 0.003381 | 5.5e-05 | 0.001129 | 0.31902 | 1.3611 | 2.5474 | 100.69 | 21.16 | 26.28 | 95.3111 | 87.0282 | 91.3119 | 0.006563 | 0.96868 | 0.023655 | 24.4195 | 0.13201 | 0.15446 | 0.000932 | 0.06313 | 0.06313 | 0.10104
2019-03-28 15:07 | 1553800025.7494 | 387 | 0.20931 | 0.000564 | 5.4e-05 | 93.2647 | 91.9003 | 119.5916 | 21.2612 | 0.002978 | 0.002799 | 135.4655 | 0.25906 | 8.4034 | 0.002626 | 0.004038 | 0.004025 | 5.4e-05 | 0.00113 | 0.30558 | 1.3682 | 2.5474 | 100.69 | 21.16 | 26.28 | 93.4765 | 84.9854 | 89.1587 | 0.006563 | 0.97443 | 0.023728 | 24.3506 | 0.13182 | 0.15411 | 0.000969 | 0.066017 | 0.066017 | 0.096413
2019-03-28 15:07 | 1553800026.7501 | 388 | 0.20922 | 0.000557 | 5.3e-05 | 92.9073 | 91.46 | 120.0005 | 21.2612 | 0.002996 | 0.002852 | 98.6923 | 0.25888 | 8.402 | 0.002658 | 0.003822 | 0.003767 | 5.3e-05 | 0.001129 | 0.31382 | 1.3325 | 2.5474 | 100.69 | 21.16 | 26.291 | 93.9046 | 85.7535 | 90.4284 | 0.006566 | 0.97103 | 0.023727 | 24.4349 | 0.14044 | 0.15401 | 0.001546 | 0.10726 | 0.10726 | 0.059303
2019-03-28 15:07 | 1553800027.7501 | 389 | 0.20926 | 0.000557 | 5.3e-05 | 92.6928 | 91.2018 | 120.4798 | 21.2612 | 0.002966 | 0.002843 | 61.9183 | 0.25903 | 8.4006 | 0.002717 | 0.004094 | 0.003981 | 5.1e-05 | 0.001125 | 0.31813 | 1.3626 | 2.5475 | 100.69 | 21.16 | 26.295 | 93.1778 | 87.1774 | 90.2174 | 0.006567 | 0.97159 | 0.023781 | 24.4765 | 0.14505 | 0.1541 | 0.001305 | 0.090339 | 0.090339 | 0.070449
2019-03-28 15:07 | 1553800028.7491 | 390 | 0.2093 | 0.000558 | 5e-05 | 92.3318 | 90.7703 | 120.5596 | 21.2612 | 0.003026 | 0.002863 | 25.1741 | 0.25357 | 8.4015 | 0.00271 | 0.003719 | 0.003724 | 5.6e-05 | 0.001124 | 0.32493 | 1.3021 | 2.5474 | 100.69 | 21.16 | 26.3 | 94.623 | 87.0157 | 90.6903 | 0.006568 | 0.97033 | 0.023774 | 24.5005 | 0.15256 | 0.15085 | 0.001042 | 0.071577 | 0.071577 | 0.087039
2019-03-28 15:07 | 1553800029.7502 | 391 | 0.20935 | 0.00056 | 4.7e-05 | 92.0733 | 90.5903 | 120.623 | 21.2612 | 0.002984 | 0.002855 | -11.5667 | 0.23089 | 8.4024 | 0.002694 | 0.004549 | 0.004323 | 5.3e-05 | 0.00112 | 0.31205 | 1.3518 | 2.5474 | 100.69 | 21.16 | 26.308 | 93.608 | 86.6206 | 90.4918 | 0.00657 | 0.97086 | 0.023786 | 24.5003 | 0.14524 | 0.13736 | 0.000765 | 0.05197 | 0.05197 | 0.10916
2019-03-28 15:07 | 1553800030.7502 | 392 | 0.20936 | 0.000553 | 4.4e-05 | 91.8755 | 90.2037 | 120.5579 | 21.2612 | 0.003015 | 0.002796 | -48.3022 | 0.23009 | 8.4036 | 0.002695 | 0.003191 | 0.003026 | 5.3e-05 | 0.001113 | 0.31708 | 1.3313 | 2.5474 | 100.69 | 21.16 | 26.312 | 94.3516 | 86.6629 | 89.0853 | 0.006571 | 0.97463 | 0.023826 | 24.4463 | 0.16426 | 0.13688 | 0.000714 | 0.048628 | 0.048628 | 0.11625
2019-03-28 15:07 | 1553800031.7502 | 393 | 0.20936 | 0.000559 | 4.1e-05 | 91.6959 | 90.0172 | 120.9079 | 21.2612 | 0.002928 | 0.002796 | -85.0863 | 0.23142 | 8.4018 | 0.002692 | 0.003754 | 0.00351 | 5.3e-05 | 0.001107 | 0.30287 | 1.3102 | 2.5473 | 100.69 | 21.17 | 26.321 | 92.2641 | 86.5902 | 89.094 | 0.006577 | 0.97461 | 0.02386 | 24.482 | 0.16527 | 0.13767 | 0.000729 | 0.049903 | 0.049903 | 0.11394
2019-03-28 15:07 | 1553800032.7503 | 394 | 0.20936 | 0.000552 | 3.7e-05 | 91.3933 | 89.6346 | 121.472 | 21.2612 | 0.002975 | 0.002726 | -121.8682 | 0.23031 | 8.406 | 0.002692 | 0.003676 | 0.003518 | 5.2e-05 | 0.001107 | 0.30829 | 1.3172 | 2.5474 | 100.69 | 21.16 | 26.326 | 93.3842 | 86.5676 | 87.3906 | 0.006574 | 0.97921 | 0.023972 | 24.4813 | 0.1738 | 0.13701 | 0.000727 | 0.050243 | 0.050243 | 0.11263
2019-03-28 15:07 | 1553800033.7493 | 395 | 0.20938 | 0.000554 | 3.4e-05 | 91.2074 | 89.3595 | 121.8309 | 21.2612 | 0.002872 | 0.002729 | -151.6761 | 0.23101 | 8.4049 | 0.002668 | 0.004179 | 0.004011 | 5.2e-05 | 0.001107 | 0.29319 | 1.2933 | 2.5473 | 100.69 | 21.16 | 26.326 | 90.9061 | 86.0094 | 87.4763 | 0.006574 | 0.97898 | 0.024005 | 24.5204 | 0.18309 | 0.13743 | 0.000595 | 0.040952 | 0.040952 | 0.13859
2019-03-28 15:07 | 1553800034.7504 | 396 | 0.20941 | 0.000545 | 3e-05 | 90.8879 | 89.1376 | 122.0294 | 21.2612 | 0.002918 | 0.002723 | -151.7101 | 0.22988 | 8.4038 | 0.002652 | 0.004079 | 0.003885 | 5.2e-05 | 0.001111 | 0.30342 | 1.2816 | 2.5474 | 100.69 | 21.16 | 26.343 | 92.024 | 85.6113 | 87.3367 | 0.006579 | 0.97936 | 0.024029 | 24.5356 | 0.17393 | 0.13675 | 0.000408 | 0.027699 | 0.027699 | 0.2039
2019-03-28 15:07 | 1553800035.7494 | 397 | 0.20941 | 0.00054 | 2.7e-05 | 90.7315 | 88.9997 | 122.3991 | 21.2612 | 0.002889 | 0.002731 | -151.7309 | 0.23031 | 8.4022 | 0.002704 | 0.003686 | 0.003527 | 4.8e-05 | 0.001116 | 0.32853 | 1.2968 | 2.5474 | 100.69 | 21.16 | 26.343 | 91.3179 | 86.8751 | 87.5263 | 0.006579 | 0.97884 | 0.024059 | 24.5792 | 0.17238 | 0.13701 | 0.000427 | 0.029305 | 0.029305 | 0.19309
2019-03-28 15:07 | 1553800036.7494 | 398 | 0.20943 | 0.000534 | 2.4e-05 | 90.5352 | 88.804 | 122.5749 | 21.2612 | 0.00288 | 0.002782 | -151.7238 | 0.23126 | 8.4037 | 0.002738 | 0.004443 | 0.004316 | 5.1e-05 | 0.001114 | 0.29607 | 1.3083 | 2.5475 | 100.69 | 21.16 | 26.364 | 91.1116 | 87.6926 | 88.7362 | 0.006584 | 0.97557 | 0.024036 | 24.638 | 0.17269 | 0.13757 | 0.000312 | 0.021164 | 0.021164 | 0.26846
2019-03-28 15:07 | 1553800037.7485 | 399 | 0.20945 | 0.00053 | 2e-05 | 90.3583 | 88.518 | 122.751 | 21.2612 | 0.002894 | 0.002672 | -151.7272 | 0.23024 | 8.4028 | 0.002723 | 0.00431 | 0.004149 | 5.1e-05 | 0.001112 | 0.30989 | 1.3164 | 2.5474 | 100.69 | 21.16 | 26.364 | 91.4408 | 87.3155 | 86.0936 | 0.006584 | 0.98275 | 0.024142 | 24.5656 | 0.18405 | 0.13697 | 0.000237 | 0.015976 | 0.015976 | 0.35409
2019-03-28 15:07 | 1553800038.7494 | 400 | 0.20944 | 0.000525 | 1.7e-05 | 90.1295 | 88.3977 | 122.9402 | 21.2612 | 0.002818 | 0.002708 | -151.735 | 0.22994 | 8.4015 | 0.002733 | 0.003844 | 0.003666 | 4.9e-05 | 0.001111 | 0.30215 | 1.3173 | 2.5475 | 100.69 | 21.16 | 26.378 | 89.6235 | 87.5545 | 86.9711 | 0.006587 | 0.98035 | 0.024131 | 24.6145 | 0.17354 | 0.13679 | 0.000315 | 0.021825 | 0.021825 | 0.25885
2019-03-28 15:07 | 1553800039.7495 | 401 | 0.20948 | 0.000515 | 1.5e-05 | 89.8868 | 88.1782 | 123.0986 | 21.2612 | 0.002822 | 0.002721 | -151.757 | 0.23075 | 8.4012 | 0.002705 | 0.004208 | 0.004054 | 5.2e-05 | 0.001117 | 0.32947 | 1.3081 | 2.5474 | 100.69 | 21.16 | 26.378 | 89.6979 | 86.897 | 87.2679 | 0.006587 | 0.97954 | 0.024136 | 24.6405 | 0.17165 | 0.13728 | 4.7e-05 | 0.002488 | 0.002488 | 2.2791
2019-03-28 15:07 | 1553800040.7496 | 402 | 0.20946 | 0.000508 | 1.2e-05 | 89.7938 | 88.2064 | 123.2955 | 21.2612 | 0.002812 | 0.002704 | -151.7605 | 0.2299 | 8.4013 | 0.002714 | 0.004182 | 0.003977 | 5e-05 | 0.001122 | 0.34899 | 1.3283 | 2.5474 | 100.69 | 21.16 | 26.399 | 89.472 | 87.1117 | 86.8564 | 0.006593 | 0.98066 | 0.02417 | 24.6461 | 0.15954 | 0.13676 | 0.00018 | 0.012327 | 0.012327 | 0.45823
2019-03-28 15:07 | 1553800041.7497 | 403 | 0.20951 | 0.000504 | 1.1e-05 | 89.6178 | 87.9914 | 123.4225 | 21.2612 | 0.002864 | 0.002635 | -151.7583 | 0.22995 | 8.4054 | 0.002772 | 0.003866 | 0.003709 | 5.1e-05 | 0.00112 | 0.31669 | 1.3102 | 2.5474 | 100.69 | 21.16 | 26.406 | 90.7146 | 88.5152 | 85.2103 | 0.006594 | 0.98517 | 0.024237 | 24.6024 | 0.16382 | 0.1368 | -0.000109 | -0.008611 | -0.008611 | -0.65609
2019-03-28 15:07 | 1553800042.7486 | 404 | 0.20952 | 0.000498 | 9e-06 | 89.602 | 87.9119 | 123.4638 | 21.2612 | 0.002774 | 0.002685 | -151.7436 | 0.23083 | 8.4013 | 0.002722 | 0.003238 | 0.003014 | 5.4e-05 | 0.001121 | 0.3011 | 1.2692 | 2.5473 | 100.69 | 21.16 | 26.419 | 88.5479 | 87.2943 | 86.4115 | 0.006598 | 0.98188 | 0.024201 | 24.6477 | 0.17033 | 0.13732 | -0.000166 | -0.012657 | -0.012657 | -0.44807
2019-03-28 15:07 | 1553800043.7486 | 405 | 0.20951 | 0.000501 | 7e-06 | 89.7977 | 88.3308 | 123.5394 | 21.2612 | 0.00274 | 0.002623 | -151.7368 | 0.23077 | 8.4024 | 0.002756 | 0.004591 | 0.004421 | 4.9e-05 | 0.00112 | 0.33876 | 1.3311 | 2.5474 | 100.69 | 21.16 | 26.424 | 87.7442 | 88.1105 | 84.9145 | 0.006599 | 0.98598 | 0.024259 | 24.6038 | 0.14733 | 0.13729 | -9.6e-05 | -0.007465 | -0.007465 | -0.75958
2019-03-28 15:07 | 1553800044.7491 | 406 | 0.20951 | 0.000496 | 6e-06 | 89.918 | 88.4408 | 123.4911 | 21.2612 | 0.002794 | 0.002613 | -151.7469 | 0.23057 | 8.402 | 0.002778 | 0.004031 | 0.003838 | 5e-05 | 0.001124 | 0.3126 | 1.3442 | 2.5474 | 100.69 | 21.16 | 26.432 | 89.0227 | 88.6416 | 84.6687 | 0.006601 | 0.98666 | 0.024263 | 24.5906 | 0.14816 | 0.13716 | -6.4e-05 | -0.005047 | -0.005047 | -1.1225
2019-03-28 15:07 | 1553800045.7497 | 407 | 0.20947 | 0.000483 | 5e-06 | 90.1209 | 88.6311 | 123.2955 | 21.2612 | 0.002762 | 0.002614 | -151.7292 | 0.23022 | 8.4024 | 0.002714 | 0.004863 | 0.004644 | 4.9e-05 | 0.001121 | 0.31852 | 1.3432 | 2.5475 | 100.69 | 21.16 | 26.434 | 88.266 | 87.0977 | 84.693 | 0.006601 | 0.98659 | 0.024243 | 24.5719 | 0.1491 | 0.13695 | 0.000158 | 0.01121 | 0.01121 | 0.50459
2019-03-28 15:07 | 1553800046.7487 | 408 | 0.20958 | 0.000486 | 4e-06 | 90.3842 | 88.7925 | 123.1169 | 21.2612 | 0.002677 | 0.002594 | -151.7261 | 0.22987 | 8.4022 | 0.002684 | 0.003931 | 0.003761 | 5.1e-05 | 0.001123 | 0.29391 | 1.3135 | 2.5474 | 100.69 | 21.16 | 26.435 | 86.2219 | 86.3881 | 84.2303 | 0.006602 | 0.98787 | 0.024241 | 24.5383 | 0.15892 | 0.13675 | -0.000498 | -0.036597 | -0.036597 | -0.15432
2019-03-28 15:07 | 1553800047.7487 | 409 | 0.20949 | 0.00048 | 4e-06 | 90.6207 | 89.0493 | 123.1061 | 21.2612 | 0.002762 | 0.002614 | -151.7241 | 0.23079 | 8.4049 | 0.002791 | 0.004211 | 0.003974 | 5.3e-05 | 0.001125 | 0.29662 | 1.3349 | 2.5474 | 100.69 | 21.17 | 26.429 | 88.2542 | 88.9555 | 84.7023 | 0.006604 | 0.98657 | 0.024224 | 24.5534 | 0.15646 | 0.1373 | 8.4e-05 | 0.005908 | 0.005908 | 0.95974
2019-03-28 15:07 | 1553800048.7498 | 410 | 0.20944 | 0.000481 | 3e-06 | 90.7928 | 89.4009 | 123.0939 | 21.2612 | 0.002772 | 0.002578 | -151.7382 | 0.23062 | 8.4052 | 0.00272 | 0.004601 | 0.004442 | 5.1e-05 | 0.001127 | 0.31852 | 1.3169 | 2.5473 | 100.69 | 21.17 | 26.429 | 88.4933 | 87.2501 | 83.8316 | 0.006604 | 0.98897 | 0.024252 | 24.5223 | 0.13819 | 0.1372 | 0.00038 | 0.027561 | 0.027561 | 0.20558
2019-03-28 15:07 | 1553800049.7488 | 411 | 0.20955 | 0.000475 | 3e-06 | 91.0555 | 89.6421 | 123.1765 | 21.2612 | 0.002719 | 0.002572 | -151.7228 | 0.23071 | 8.4026 | 0.002723 | 0.004264 | 0.004116 | 5.2e-05 | 0.001127 | 0.31526 | 1.3055 | 2.5473 | 100.69 | 21.17 | 26.416 | 87.2272 | 87.3336 | 83.6902 | 0.006601 | 0.98937 | 0.024265 | 24.5256 | 0.13993 | 0.13725 | -0.000307 | -0.022622 | -0.022622 | -0.25057
2019-03-28 15:07 | 1553800050.75 | 412 | 0.20949 | 0.000474 | 2e-06 | 91.3635 | 89.9333 | 123.3612 | 21.2612 | 0.002659 | 0.002549 | -151.7313 | 0.23106 | 8.4015 | 0.002669 | 0.004289 | 0.00412 | 5.1e-05 | 0.001128 | 0.33052 | 1.3304 | 2.5474 | 100.69 | 21.17 | 26.416 | 85.7792 | 86.0359 | 83.1316 | 0.006601 | 0.99092 | 0.024302 | 24.5248 | 0.14113 | 0.13746 | 0.000101 | 0.007301 | 0.007301 | 0.77758
2019-03-28 15:08 | 1553800051.7479 | 413 | 0.20955 | 0.000478 | 2e-06 | 91.583 | 90.3712 | 123.292 | 21.2612 | 0.002687 | 0.002553 | -151.7083 | 0.23098 | 8.4033 | 0.002691 | 0.00392 | 0.003727 | 5.1e-05 | 0.001124 | 0.3063 | 1.3465 | 2.5474 | 100.68 | 21.17 | 26.404 | 86.4678 | 86.5503 | 83.2379 | 0.006598 | 0.99062 | 0.024292 | 24.5216 | 0.11915 | 0.13741 | -0.000258 | -0.018993 | -0.018993 | -0.29879
2019-03-28 15:08 | 1553800052.7489 | 414 | 0.20954 | 0.000473 | 1e-06 | 91.8061 | 90.4129 | 123.2148 | 21.2612 | 0.002703 | 0.002557 | -151.7339 | 0.23073 | 8.4029 | 0.002638 | 0.004207 | 0.004147 | 5.4e-05 | 0.001121 | 0.30287 | 1.3162 | 2.5474 | 100.68 | 21.17 | 26.404 | 86.8369 | 85.2907 | 83.3401 | 0.006598 | 0.99034 | 0.024281 | 24.5174 | 0.13678 | 0.13726 | -0.000212 | -0.015556 | -0.015556 | -0.36443
2019-03-28 15:08 | 1553800053.7479 | 415 | 0.20952 | 0.000472 | 1e-06 | 91.9842 | 90.7207 | 123.4481 | 21.2612 | 0.00266 | 0.002535 | -151.7239 | 0.23044 | 8.4018 | 0.002627 | 0.003901 | 0.003665 | 5e-05 | 0.001121 | 0.31957 | 1.3242 | 2.5474 | 100.68 | 21.17 | 26.39 | 85.7987 | 85.0222 | 82.808 | 0.006595 | 0.99182 | 0.024322 | 24.5223 | 0.12372 | 0.13709 | -0.0001 | -0.007447 | -0.007447 | -0.76022
2019-03-28 15:08 | 1553800054.7494 | 416 | 0.20952 | 0.000462 | 1e-06 | 92.1269 | 91.1369 | 123.162 | 21.2612 | 0.002698 | 0.002511 | -151.7433 | 0.23117 | 8.4022 | 0.00268 | 0.003963 | 0.003837 | 5.1e-05 | 0.001121 | 0.31653 | 1.3103 | 2.5475 | 100.68 | 21.17 | 26.39 | 86.7144 | 86.2849 | 82.227 | 0.006595 | 0.99344 | 0.024313 | 24.4739 | 0.096641 | 0.13752 | -8.5e-05 | -0.006279 | -0.006279 | -0.90457
2019-03-28 15:08 | 1553800055.749 | 417 | 0.20951 | 0.000467 | 1e-06 | 92.235 | 91.1002 | 123.4824 | 21.2612 | 0.002646 | 0.002491 | -151.7458 | 0.23017 | 8.4035 | 0.002665 | 0.003901 | 0.003743 | 5e-05 | 0.00112 | 0.30829 | 1.2711 | 2.5473 | 100.69 | 21.16 | 26.378 | 85.4771 | 85.9274 | 81.7426 | 0.006587 | 0.99479 | 0.024362 | 24.489 | 0.11073 | 0.13693 | -4.7e-05 | -0.003517 | -0.003517 | -1.6078
2019-03-28 15:08 | 1553800056.7491 | 418 | 0.20951 | 0.000467 | 0 | 92.5064 | 91.3095 | 123.4892 | 21.2612 | 0.002646 | 0.002529 | -151.7618 | 0.22911 | 8.4021 | 0.002623 | 0.004022 | 0.003831 | 5.3e-05 | 0.00113 | 0.30879 | 1.2923 | 2.5475 | 100.69 | 21.16 | 26.378 | 85.4764 | 84.9113 | 82.6621 | 0.006587 | 0.99222 | 0.024331 | 24.5213 | 0.11649 | 0.1363 | -4.9e-05 | -0.003611 | -0.003611 | -1.5587
2019-03-28 15:08 | 1553800057.7491 | 419 | 0.20952 | 0.000465 | 0 | 92.5737 | 91.4875 | 123.3068 | 21.2612 | 0.00262 | 0.002486 | -151.7745 | 0.22967 | 8.402 | 0.002635 | 0.004168 | 0.003997 | 5.2e-05 | 0.00113 | 0.33843 | 1.3111 | 2.5474 | 100.69 | 21.17 | 26.372 | 84.8505 | 85.2033 | 81.614 | 0.00659 | 0.99516 | 0.024349 | 24.4671 | 0.10557 | 0.13663 | -0.000121 | -0.008901 | -0.008901 | -0.634
2019-03-28 15:08 | 1553800058.7494 | 420 | 0.20955 | 0.000462 | 0 | 92.6693 | 91.3302 | 123.5526 | 21.2612 | 0.002632 | 0.002487 | -151.7626 | 0.23077 | 8.4043 | 0.002642 | 0.004135 | 0.003975 | 5.4e-05 | 0.001134 | 0.30453 | 1.3125 | 2.5474 | 100.69 | 21.17 | 26.372 | 85.1439 | 85.378 | 81.6575 | 0.00659 | 0.99503 | 0.024371 | 24.493 | 0.1302 | 0.13728 | -0.000247 | -0.018163 | -0.018163 | -0.31217
2019-03-28 15:08 | 1553800059.7482 | 421 | 0.20953 | 0.000462 | 0 | 92.8039 | 91.8191 | 123.3554 | 21.2612 | 0.002666 | 0.002507 | -151.7637 | 0.23079 | 8.4029 | 0.002648 | 0.003271 | 0.003084 | 5.3e-05 | 0.001135 | 0.3011 | 1.3156 | 2.5474 | 100.69 | 21.16 | 26.371 | 85.9454 | 85.517 | 82.1193 | 0.006586 | 0.99374 | 0.024336 | 24.4894 | 0.09543 | 0.1373 | -0.000158 | -0.011573 | -0.011573 | -0.48997
2019-03-28 15:08 | 1553800060.7491 | 422 | 0.20953 | 0.000462 | 0 | 92.8975 | 91.869 | 123.4 | 21.2612 | 0.002586 | 0.002459 | -151.777 | 0.23079 | 8.4021 | 0.002562 | 0.003569 | 0.003347 | 5.4e-05 | 0.001134 | 0.33035 | 1.289 | 2.5473 | 100.69 | 21.17 | 26.368 | 84.0391 | 83.4569 | 80.9751 | 0.006589 | 0.99695 | 0.02438 | 24.4543 | 0.099576 | 0.1373 | -0.000157 | -0.011496 | -0.011496 | -0.49322
2019-03-28 15:08 | 1553800061.7482 | 423 | 0.20953 | 0.00046 | 0 | 92.9386 | 91.8741 | 123.4404 | 21.2612 | 0.002618 | 0.002453 | -151.794 | 0.2305 | 8.4012 | 0.002599 | 0.003733 | 0.003526 | 5.4e-05 | 0.001134 | 0.29518 | 1.331 | 2.5473 | 100.69 | 21.17 | 26.368 | 84.8046 | 84.3315 | 80.8278 | 0.006589 | 0.99737 | 0.024389 | 24.4532 | 0.10304 | 0.13712 | -0.000151 | -0.011053 | -0.011053 | -0.51237
2019-03-28 15:08 | 1553800062.7482 | 424 | 0.20952 | 0.000458 | 0 | 93.0407 | 91.8867 | 123.2009 | 21.2612 | 0.002617 | 0.002418 | -151.7625 | 0.2304 | 8.4022 | 0.002575 | 0.003918 | 0.003801 | 5.2e-05 | 0.001138 | 0.30127 | 1.3246 | 2.5473 | 100.69 | 21.17 | 26.368 | 84.7827 | 83.7597 | 79.9951 | 0.006589 | 0.99972 | 0.024394 | 24.4007 | 0.11164 | 0.13707 | -6.9e-05 | -0.005094 | -0.005094 | -1.1113
2019-03-28 15:08 | 1553800063.7482 | 425 | 0.20952 | 0.000461 | 0 | 93.0985 | 92.1466 | 123.0819 | 21.2612 | 0.002612 | 0.002463 | -151.7701 | 0.2304 | 8.4023 | 0.002586 | 0.003957 | 0.003705 | 5e-05 | 0.001145 | 0.323 | 1.3789 | 2.5476 | 100.69 | 21.17 | 26.368 | 84.653 | 84.0391 | 81.0713 | 0.006589 | 0.99668 | 0.024345 | 24.4261 | 0.09193 | 0.13706 | -0.000101 | -0.007387 | -0.007387 | -0.76627
2019-03-28 15:08 | 1553800064.7493 | 426 | 0.20951 | 0.000463 | 0 | 93.2459 | 92.2444 | 123.6448 | 21.2612 | 0.002537 | 0.002461 | -151.7333 | 0.23087 | 8.4016 | 0.002553 | 0.003396 | 0.003282 | 5.4e-05 | 0.001141 | 0.33179 | 1.3688 | 2.5475 | 100.69 | 21.17 | 26.369 | 82.84 | 83.241 | 81.0275 | 0.006589 | 0.9968 | 0.024402 | 24.4804 | 0.096592 | 0.13734 | -2e-06 | -0.000105 | -0.000105 | -54.01
2019-03-28 15:08 | 1553800065.7487 | 427 | 0.20952 | 0.000456 | 0 | 93.3004 | 92.4377 | 123.4875 | 21.2612 | 0.002577 | 0.002443 | -151.7368 | 0.23025 | 8.4024 | 0.002539 | 0.003721 | 0.003543 | 5.2e-05 | 0.001141 | 0.32925 | 1.3136 | 2.5473 | 100.69 | 21.17 | 26.367 | 83.8215 | 82.9056 | 80.5901 | 0.006589 | 0.99804 | 0.024402 | 24.4497 | 0.083095 | 0.13698 | -8.7e-05 | -0.0064 | -0.0064 | -0.88399
2019-03-28 15:08 | 1553800066.7483 | 428 | 0.20944 | 0.000455 | 0 | 93.4014 | 92.4778 | 123.4399 | 21.2612 | 0.002615 | 0.002416 | -151.734 | 0.23031 | 8.403 | 0.002593 | 0.003905 | 0.003708 | 5e-05 | 0.001147 | 0.30973 | 1.3339 | 2.5474 | 100.69 | 21.17 | 26.366 | 84.7271 | 84.2098 | 79.935 | 0.006588 | 0.99989 | 0.02442 | 24.4223 | 0.088891 | 0.13701 | 0.000392 | 0.028822 | 0.028822 | 0.19633
2019-03-28 15:08 | 1553800067.7484 | 429 | 0.20949 | 0.000454 | 0 | 93.4598 | 92.3755 | 123.5396 | 21.2612 | 0.002532 | 0.002406 | -151.7083 | 0.23098 | 8.4036 | 0.002667 | 0.004478 | 0.004339 | 5.2e-05 | 0.001155 | 0.28279 | 1.282 | 2.5475 | 100.69 | 21.17 | 26.366 | 82.7316 | 85.9743 | 79.6988 | 0.006588 | 1.0006 | 0.024438 | 24.424 | 0.10438 | 0.13741 | 0.000128 | 0.009385 | 0.009385 | 0.60469
2019-03-28 15:08 | 1553800068.7474 | 430 | 0.20951 | 0.000458 | 0 | 93.5798 | 92.6381 | 123.5596 | 21.2612 | 0.002527 | 0.002424 | -151.7157 | 0.23088 | 8.4019 | 0.002565 | 0.003813 | 0.003659 | 5.3e-05 | 0.001156 | 0.29518 | 1.2863 | 2.5474 | 100.68 | 21.17 | 26.356 | 82.6013 | 83.5146 | 80.1472 | 0.006586 | 0.99929 | 0.024424 | 24.4415 | 0.090465 | 0.13735 | -3e-05 | -0.002186 | -0.002186 | -2.5948
2019-03-28 15:08 | 1553800069.7485 | 431 | 0.20953 | 0.000454 | 0 | 93.6915 | 92.6629 | 123.4127 | 21.2612 | 0.002551 | 0.002399 | -151.7136 | 0.22942 | 8.405 | 0.002622 | 0.004171 | 0.004005 | 5.7e-05 | 0.001164 | 0.31509 | 1.3272 | 2.5472 | 100.68 | 21.17 | 26.356 | 83.1848 | 84.8964 | 79.5383 | 0.006586 | 1.001 | 0.024431 | 24.4059 | 0.09874 | 0.13648 | -0.000112 | -0.008209 | -0.008209 | -0.68663
2019-03-28 15:08 | 1553800070.7485 | 432 | 0.2095 | 0.000456 | -1e-06 | 93.727 | 92.72 | 123.265 | 21.2612 | 0.002544 | 0.002387 | -151.6839 | 0.22982 | 8.4023 | 0.002606 | 0.004 | 0.003752 | 5e-05 | 0.001166 | 0.32493 | 1.3125 | 2.5474 | 100.68 | 21.17 | 26.335 | 83.0249 | 84.5111 | 79.2482 | 0.006581 | 1.0018 | 0.024426 | 24.3813 | 0.096616 | 0.13672 | 1.9e-05 | 0.001465 | 0.001465 | 3.8539
2019-03-28 15:08 | 1553800071.7489 | 433 | 0.20954 | 0.000457 | 0 | 93.7204 | 92.738 | 123.1993 | 21.2612 | 0.002536 | 0.002422 | -151.688 | 0.23048 | 8.4051 | 0.002584 | 0.003869 | 0.003757 | 5.5e-05 | 0.001173 | 0.32316 | 1.3186 | 2.5474 | 100.68 | 21.17 | 26.335 | 82.8251 | 83.9773 | 80.0934 | 0.006581 | 0.99944 | 0.02439 | 24.404 | 0.094254 | 0.13711 | -0.000177 | -0.013 | -0.013 | -0.43558
2019-03-28 15:08 | 1553800072.7476 | 434 | 0.20954 | 0.000458 | -1e-06 | 93.8515 | 92.9432 | 123.4216 | 21.2612 | 0.002527 | 0.00243 | -151.6958 | 0.22995 | 8.4016 | 0.002663 | 0.003944 | 0.003768 | 5e-05 | 0.001173 | 0.32693 | 1.2719 | 2.5474 | 100.68 | 21.17 | 26.302 | 82.6207 | 85.8714 | 80.2796 | 0.006573 | 0.99891 | 0.024406 | 24.4324 | 0.086984 | 0.1368 | -0.00018 | -0.013188 | -0.013188 | -0.4284
2019-03-28 15:08 | 1553800073.7476 | 435 | 0.20953 | 0.000456 | -1e-06 | 93.8795 | 92.7778 | 123.5767 | 21.2612 | 0.002524 | 0.002398 | -151.7161 | 0.22997 | 8.4026 | 0.002539 | 0.003926 | 0.003848 | 5.2e-05 | 0.001177 | 0.32532 | 1.2824 | 2.5473 | 100.68 | 21.17 | 26.302 | 82.5349 | 82.89 | 79.5212 | 0.006573 | 1.0011 | 0.024447 | 24.4215 | 0.10559 | 0.13681 | -0.000154 | -0.011341 | -0.011341 | -0.49822
2019-03-28 15:08 | 1553800074.748 | 436 | 0.20949 | 0.000452 | 0 | 93.8988 | 92.9599 | 123.4256 | 21.2612 | 0.002503 | 0.002331 | -151.7432 | 0.22992 | 8.4043 | 0.002561 | 0.003846 | 0.003655 | 5.5e-05 | 0.00118 | 0.35187 | 1.3167 | 2.5474 | 100.68 | 21.18 | 26.278 | 82.0428 | 83.4384 | 77.9072 | 0.006571 | 1.0057 | 0.024489 | 24.3507 | 0.089894 | 0.13678 | 8.5e-05 | 0.006328 | 0.006328 | 0.89274
2019-03-28 15:08 | 1553800075.7477 | 437 | 0.20952 | 0.000455 | 0 | 93.9119 | 92.9799 | 123.3948 | 21.2612 | 0.002503 | 0.002392 | -151.7303 | 0.22924 | 8.4026 | 0.002524 | 0.003983 | 0.003796 | 5e-05 | 0.001173 | 0.3402 | 1.2917 | 2.5473 | 100.68 | 21.18 | 26.278 | 82.0342 | 82.5296 | 79.3684 | 0.006571 | 1.0015 | 0.024435 | 24.3983 | 0.089214 | 0.13637 | -6.9e-05 | -0.00507 | -0.00507 | -1.1108
2019-03-28 15:08 | 1553800076.7477 | 438 | 0.20956 | 0.000459 | -1e-06 | 93.9322 | 93.0458 | 123.5781 | 21.2612 | 0.002484 | 0.002364 | -151.7264 | 0.23058 | 8.4033 | 0.002562 | 0.003263 | 0.003094 | 5.3e-05 | 0.001168 | 0.31094 | 1.3018 | 2.5474 | 100.68 | 21.17 | 26.252 | 81.5883 | 83.4423 | 78.7049 | 0.006561 | 1.0034 | 0.024476 | 24.3934 | 0.084805 | 0.13717 | -0.000306 | -0.022523 | -0.022523 | -0.25153
2019-03-28 15:08 | 1553800077.7477 | 439 | 0.20948 | 0.000456 | -1e-06 | 94.046 | 93.1256 | 123.8032 | 21.2612 | 0.0025 | 0.002352 | -151.75 | 0.231 | 8.4023 | 0.002569 | 0.003194 | 0.002986 | 5.3e-05 | 0.001156 | 0.32587 | 1.2995 | 2.5472 | 100.68 | 21.18 | 26.224 | 81.9601 | 83.6185 | 78.3969 | 0.006558 | 1.0043 | 0.024509 | 24.4049 | 0.087962 | 0.13742 | 0.000176 | 0.013068 | 0.013068 | 0.4343
2019-03-28 15:08 | 1553800078.7478 | 440 | 0.2095 | 0.000455 | -1e-06 | 94.113 | 93.2666 | 123.822 | 21.2612 | 0.002471 | 0.002353 | -151.765 | 0.23005 | 8.4022 | 0.002527 | 0.003407 | 0.003256 | 5.2e-05 | 0.00115 | 0.3126 | 1.2721 | 2.5473 | 100.68 | 21.18 | 26.224 | 81.2568 | 82.6123 | 78.4397 | 0.006558 | 1.0041 | 0.024509 | 24.4082 | 0.080802 | 0.13686 | 5.7e-05 | 0.004226 | 0.004226 | 1.3376
2019-03-28 15:08 | 1553800079.7478 | 441 | 0.20949 | 0.000456 | -1e-06 | 94.2101 | 93.2649 | 123.8934 | 21.2612 | 0.002497 | 0.002347 | -151.7771 | 0.23083 | 8.4035 | 0.002498 | 0.004008 | 0.003754 | 5.4e-05 | 0.00115 | 0.31614 | 1.2675 | 2.5473 | 100.68 | 21.18 | 26.186 | 81.8907 | 81.9063 | 78.2933 | 0.006548 | 1.0046 | 0.024521 | 24.4102 | 0.090188 | 0.13732 | 9.3e-05 | 0.006928 | 0.006928 | 0.81866
2019-03-28 15:08 | 1553800080.748 | 442 | 0.20952 | 0.000455 | -1e-06 | 94.2988 | 93.411 | 124.0868 | 21.2612 | 0.002407 | 0.002383 | -151.7979 | 0.23093 | 8.4049 | 0.002433 | 0.003539 | 0.003365 | 5.3e-05 | 0.001147 | 0.32925 | 1.3067 | 2.5474 | 100.68 | 21.18 | 26.186 | 79.737 | 80.3484 | 79.1444 | 0.006548 | 1.0021 | 0.024511 | 24.4588 | 0.084613 | 0.13738 | -8.2e-05 | -0.005986 | -0.005986 | -0.94789
2019-03-28 15:08 | 1553800081.7488 | 443 | 0.20951 | 0.000459 | 0 | 94.345 | 93.5894 | 124.1614 | 21.2612 | 0.002458 | 0.0023 | -151.8148 | 0.23046 | 8.4025 | 0.002464 | 0.003879 | 0.003763 | 5e-05 | 0.001152 | 0.35081 | 1.338 | 2.5473 | 100.68 | 21.18 | 26.164 | 80.955 | 81.097 | 77.1629 | 0.006543 | 1.0078 | 0.024588 | 24.3973 | 0.071929 | 0.1371 | -3.8e-05 | -0.002764 | -0.002764 | -2.0489
2019-03-28 15:08 | 1553800082.7478 | 444 | 0.20952 | 0.000454 | -1e-06 | 94.4713 | 93.8105 | 124.2818 | 21.2612 | 0.002469 | 0.002359 | -151.8205 | 0.23111 | 8.4047 | 0.002474 | 0.00377 | 0.00351 | 5.2e-05 | 0.001152 | 0.31686 | 1.3727 | 2.5476 | 100.68 | 21.18 | 26.164 | 81.2061 | 81.3333 | 78.573 | 0.006543 | 1.0038 | 0.02455 | 24.4582 | 0.062782 | 0.13749 | -6e-05 | -0.004424 | -0.004424 | -1.2834
2019-03-28 15:08 | 1553800083.7469 | 445 | 0.2095 | 0.000462 | 0 | 94.6432 | 93.7225 | 124.2243 | 21.2612 | 0.002408 | 0.002296 | -151.7999 | 0.23126 | 8.4035 | 0.002541 | 0.004348 | 0.004184 | 5.2e-05 | 0.001148 | 0.29175 | 1.3107 | 2.5474 | 100.68 | 21.18 | 26.133 | 79.7543 | 82.9429 | 77.0709 | 0.006535 | 1.0081 | 0.024597 | 24.4002 | 0.087432 | 0.13758 | 3.3e-05 | 0.002446 | 0.002446 | 2.3233
2019-03-28 15:08 | 1553800084.7479 | 446 | 0.20952 | 0.000456 | 0 | 94.7885 | 93.9468 | 124.2504 | 21.2612 | 0.002446 | 0.002309 | -151.7953 | 0.23018 | 8.4024 | 0.002484 | 0.003821 | 0.003696 | 4.9e-05 | 0.001144 | 0.35651 | 1.3173 | 2.5473 | 100.68 | 21.18 | 26.133 | 80.6687 | 81.5772 | 77.3608 | 0.006535 | 1.0072 | 0.024589 | 24.4129 | 0.079786 | 0.13694 | -6.3e-05 | -0.004648 | -0.004648 | -1.2167
2019-03-28 15:08 | 1553800085.748 | 447 | 0.20952 | 0.000457 | 0 | 94.8648 | 93.9716 | 124.5515 | 21.2612 | 0.002396 | 0.00229 | -151.8177 | 0.23072 | 8.4012 | 0.002487 | 0.003856 | 0.00361 | 5e-05 | 0.001145 | 0.31758 | 1.3488 | 2.5474 | 100.68 | 21.18 | 26.112 | 79.4532 | 81.6491 | 76.9111 | 0.00653 | 1.0085 | 0.024635 | 24.4268 | 0.084617 | 0.13725 | -7.3e-05 | -0.005397 | -0.005397 | -1.0504
2019-03-28 15:08 | 1553800086.747 | 448 | 0.20951 | 0.000458 | 0 | 94.9925 | 94.3724 | 124.3539 | 21.2612 | 0.002473 | 0.002337 | -151.8197 | 0.23039 | 8.4024 | 0.002507 | 0.003729 | 0.003615 | 5.4e-05 | 0.001142 | 0.32405 | 1.3333 | 2.5474 | 100.68 | 21.18 | 26.112 | 81.3184 | 82.1412 | 78.0492 | 0.00653 | 1.0053 | 0.024576 | 24.447 | 0.058581 | 0.13706 | 8e-06 | 0.000627 | 0.000627 | 9.0277
2019-03-28 15:08 | 1553800087.748 | 449 | 0.20949 | 0.00046 | 0 | 95.0619 | 94.3228 | 124.0532 | 21.2612 | 0.002403 | 0.002309 | -151.8009 | 0.23138 | 8.4033 | 0.002567 | 0.003709 | 0.003501 | 5.2e-05 | 0.001138 | 0.2995 | 1.3643 | 2.5476 | 100.68 | 21.18 | 26.094 | 79.6404 | 83.5796 | 77.3733 | 0.006525 | 1.0072 | 0.024569 | 24.3939 | 0.069816 | 0.13765 | 8.6e-05 | 0.006408 | 0.006408 | 0.8871
2019-03-28 15:08 | 1553800088.7471 | 450 | 0.20947 | 0.000458 | -1e-06 | 95.154 | 94.403 | 123.8582 | 21.2612 | 0.00239 | 0.002327 | -151.7821 | 0.23003 | 8.4047 | 0.002588 | 0.004543 | 0.004392 | 5.1e-05 | 0.00114 | 0.32405 | 1.398 | 2.5474 | 100.68 | 21.18 | 26.094 | 79.316 | 84.0718 | 77.813 | 0.006525 | 1.0059 | 0.024535 | 24.39 | 0.070871 | 0.13685 | 0.000243 | 0.017999 | 0.017999 | 0.314
2019-03-28 15:08 | 1553800089.7481 | 451 | 0.20953 | 0.00046 | 0 | 95.1826 | 94.5277 | 123.9664 | 21.2612 | 0.002441 | 0.002238 | -151.772 | 0.23085 | 8.3999 | 0.002577 | 0.003481 | 0.003337 | 5e-05 | 0.001139 | 0.33716 | 1.3479 | 2.5475 | 100.68 | 21.18 | 26.09 | 80.5425 | 83.8205 | 75.6772 | 0.006524 | 1.0121 | 0.02462 | 24.3263 | 0.061754 | 0.13733 | -0.000145 | -0.01073 | -0.01073 | -0.52861
2019-03-28 15:08 | 1553800090.7474 | 452 | 0.20952 | 0.000461 | 0 | 95.3419 | 94.65 | 124.1241 | 21.2612 | 0.002469 | 0.002308 | -151.7719 | 0.23066 | 8.4005 | 0.002525 | 0.003892 | 0.00368 | 5.7e-05 | 0.001136 | 0.33284 | 1.3208 | 2.5474 | 100.68 | 21.18 | 26.073 | 81.2068 | 82.5693 | 77.3491 | 0.00652 | 1.0073 | 0.024577 | 24.4001 | 0.065151 | 0.13722 | -4.8e-05 | -0.003501 | -0.003501 | -1.6189
2019-03-28 15:08 | 1553800091.7472 | 453 | 0.20944 | 0.000458 | -1e-06 | 95.4125 | 94.5858 | 124.2023 | 21.2612 | 0.002381 | 0.002332 | -151.7584 | 0.23101 | 8.4031 | 0.002541 | 0.00459 | 0.00437 | 5.7e-05 | 0.001138 | 0.29374 | 1.3315 | 2.5474 | 100.68 | 21.19 | 26.067 | 79.1072 | 82.9477 | 77.9228 | 0.006522 | 1.0056 | 0.024565 | 24.4277 | 0.07784 | 0.13743 | 0.000394 | 0.029153 | 0.029153 | 0.1947
2019-03-28 15:08 | 1553800092.7483 | 454 | 0.20953 | 0.000453 | -1e-06 | 95.4372 | 94.6904 | 124.4675 | 21.2612 | 0.002433 | 0.002288 | -151.7666 | 0.23046 | 8.4031 | 0.002606 | 0.00329 | 0.003129 | 5.4e-05 | 0.001131 | 0.29822 | 1.3023 | 2.5471 | 100.68 | 21.19 | 26.059 | 80.3445 | 84.522 | 76.8652 | 0.00652 | 1.0087 | 0.024628 | 24.4169 | 0.070268 | 0.1371 | -0.000105 | -0.007779 | -0.007779 | -0.72785
2019-03-28 15:08 | 1553800093.7483 | 455 | 0.20966 | 0.000457 | 0 | 95.3976 | 94.6908 | 124.7696 | 21.2612 | 0.002389 | 0.002305 | -151.7757 | 0.23042 | 8.4031 | 0.002543 | 0.003274 | 0.003143 | 5.3e-05 | 0.00113 | 0.3361 | 1.3153 | 2.5473 | 100.68 | 21.19 | 26.059 | 79.3028 | 82.9929 | 77.2643 | 0.00652 | 1.0075 | 0.024644 | 24.4605 | 0.066523 | 0.13708 | -0.000893 | -0.066257 | -0.066257 | -0.085444
2019-03-28 15:08 | 1553800094.7486 | 456 | 0.20932 | 0.000456 | -1e-06 | 95.4131 | 94.7893 | 124.5187 | 21.2612 | 0.002367 | 0.00233 | -151.7723 | 0.23076 | 8.4047 | 0.002639 | 0.003708 | 0.003511 | 5.4e-05 | 0.001128 | 0.34938 | 1.2916 | 2.5474 | 100.68 | 21.19 | 26.059 | 78.7688 | 85.2969 | 77.8877 | 0.00652 | 1.0057 | 0.024597 | 24.4576 | 0.058673 | 0.13728 | 0.001115 | 0.082677 | 0.082677 | 0.068576
2019-03-28 15:08 | 1553800095.7473 | 457 | 0.20934 | 0.000457 | 0 | 95.5132 | 94.8103 | 124.3337 | 21.2612 | 0.002416 | 0.002257 | -151.781 | 0.23063 | 8.4027 | 0.002596 | 0.004108 | 0.003908 | 5e-05 | 0.001122 | 0.33251 | 1.3026 | 2.5474 | 100.68 | 21.19 | 26.059 | 79.9499 | 84.2769 | 76.1199 | 0.00652 | 1.0108 | 0.024641 | 24.3778 | 0.066063 | 0.1372 | 0.001019 | 0.075631 | 0.075631 | 0.074923
2019-03-28 15:08 | 1553800096.7473 | 458 | 0.20955 | 0.000456 | 0 | 95.4432 | 94.8042 | 124.345 | 21.2612 | 0.002386 | 0.002268 | -151.7732 | 0.23072 | 8.405 | 0.002607 | 0.003985 | 0.003802 | 5.2e-05 | 0.001119 | 0.32764 | 1.3391 | 2.5474 | 100.68 | 21.19 | 26.062 | 79.2302 | 84.5306 | 76.3926 | 0.006521 | 1.01 | 0.024633 | 24.3884 | 0.060083 | 0.13725 | -0.000276 | -0.020443 | -0.020443 | -0.27729
2019-03-28 15:08 | 1553800097.7474 | 459 | 0.20946 | 0.000451 | 0 | 95.526 | 94.8205 | 124.4245 | 21.2612 | 0.002339 | 0.002301 | -151.7871 | 0.23035 | 8.4032 | 0.002619 | 0.004307 | 0.004132 | 5.2e-05 | 0.001118 | 0.31166 | 1.3259 | 2.5474 | 100.68 | 21.19 | 26.062 | 78.0867 | 84.8168 | 77.1768 | 0.006521 | 1.0078 | 0.024613 | 24.4236 | 0.066301 | 0.13704 | 0.000303 | 0.022445 | 0.022445 | 0.25215
2019-03-28 15:08 | 1553800098.7474 | 460 | 0.20955 | 0.000457 | 0 | 95.5474 | 94.9568 | 124.7426 | 21.2612 | 0.002349 | 0.002267 | -151.7998 | 0.23081 | 8.4034 | 0.002647 | 0.004229 | 0.004006 | 5.1e-05 | 0.001118 | 0.32947 | 1.3145 | 2.5474 | 100.68 | 21.19 | 26.062 | 78.3368 | 85.5014 | 76.3599 | 0.006521 | 1.0101 | 0.024673 | 24.4262 | 0.055467 | 0.13731 | -0.000269 | -0.019966 | -0.019966 | -0.28403
2019-03-28 15:08 | 1553800099.7467 | 461 | 0.20947 | 0.00045 | 0 | 95.4462 | 94.8813 | 124.8933 | 21.2612 | 0.002412 | 0.002253 | -151.8177 | 0.23059 | 8.4042 | 0.002606 | 0.003458 | 0.003345 | 5.2e-05 | 0.00112 | 0.32693 | 1.3382 | 2.5475 | 100.68 | 21.19 | 26.062 | 79.8401 | 84.5063 | 76.0334 | 0.006521 | 1.0111 | 0.0247 | 24.4296 | 0.053089 | 0.13718 | 0.000229 | 0.017047 | 0.017047 | 0.33235
2019-03-28 15:08 | 1553800100.7466 | 462 | 0.20951 | 0.000448 | -1e-06 | 95.4661 | 94.74 | 124.717 | 21.2612 | 0.002378 | 0.002275 | -151.7937 | 0.23089 | 8.4042 | 0.002578 | 0.004315 | 0.004059 | 5.3e-05 | 0.001124 | 0.32997 | 1.3191 | 2.5474 | 100.68 | 21.19 | 26.057 | 79.037 | 83.8393 | 76.5627 | 0.00652 | 1.0095 | 0.024664 | 24.4308 | 0.068292 | 0.13735 | -1.2e-05 | -0.000876 | -0.000876 | -6.4746
2019-03-28 15:08 | 1553800101.7468 | 463 | 0.20951 | 0.000455 | 0 | 95.4369 | 94.6531 | 124.4494 | 21.2612 | 0.002341 | 0.002311 | -151.7791 | 0.2312 | 8.4018 | 0.002537 | 0.003948 | 0.003773 | 5.2e-05 | 0.001129 | 0.33069 | 1.3204 | 2.5473 | 100.68 | 21.18 | 26.056 | 78.1319 | 82.8626 | 77.4296 | 0.006516 | 1.007 | 0.024607 | 24.4348 | 0.073771 | 0.13754 | -2.3e-05 | -0.001683 | -0.001683 | -3.3756
2019-03-28 15:08 | 1553800102.7475 | 464 | 0.20953 | 0.000454 | 0 | 95.4596 | 94.6561 | 124.3034 | 21.2612 | 0.00231 | 0.002251 | -151.7638 | 0.231 | 8.4013 | 0.002548 | 0.004235 | 0.004046 | 5.5e-05 | 0.001131 | 0.33876 | 1.2996 | 2.5474 | 100.68 | 21.18 | 26.056 | 77.389 | 83.117 | 75.989 | 0.006516 | 1.0112 | 0.024643 | 24.3703 | 0.075603 | 0.13742 | -0.000146 | -0.010817 | -0.010817 | -0.52468
2019-03-28 15:08 | 1553800103.7475 | 465 | 0.20953 | 0.000454 | -1e-06 | 95.4646 | 94.882 | 123.815 | 21.2612 | 0.002314 | 0.002244 | -151.755 | 0.23026 | 8.4011 | 0.002552 | 0.004099 | 0.003939 | 5.6e-05 | 0.001136 | 0.32621 | 1.2973 | 2.5474 | 100.68 | 21.18 | 26.06 | 77.4887 | 83.2018 | 75.8191 | 0.006517 | 1.0117 | 0.0246 | 24.3164 | 0.054755 | 0.13698 | -0.000151 | -0.011135 | -0.011135 | -0.50805
2019-03-28 15:08 | 1553800104.7476 | 466 | 0.20952 | 0.000453 | 0 | 95.4769 | 94.7292 | 123.9429 | 21.2612 | 0.002335 | 0.00227 | -151.7639 | 0.22989 | 8.4013 | 0.002581 | 0.004061 | 0.003853 | 5.3e-05 | 0.001138 | 0.34612 | 1.2767 | 2.5473 | 100.68 | 21.18 | 26.06 | 78.0048 | 83.9039 | 76.4356 | 0.006517 | 1.0099 | 0.024591 | 24.3505 | 0.070323 | 0.13676 | -4.4e-05 | -0.003279 | -0.003279 | -1.7226
2019-03-28 15:08 | 1553800105.7477 | 467 | 0.20953 | 0.000453 | 0 | 95.4152 | 94.6085 | 124.2164 | 21.2612 | 0.002331 | 0.002288 | -151.7403 | 0.22989 | 8.4031 | 0.002585 | 0.003876 | 0.00374 | 5.2e-05 | 0.001136 | 0.34827 | 1.2995 | 2.5475 | 100.68 | 21.18 | 26.057 | 77.9019 | 84.0141 | 76.869 | 0.006516 | 1.0086 | 0.024603 | 24.3924 | 0.07594 | 0.13676 | -0.000143 | -0.010616 | -0.010616 | -0.53207
2019-03-28 15:08 | 1553800106.748 | 468 | 0.20953 | 0.000455 | -1e-06 | 95.2874 | 94.6021 | 124.6119 | 21.2612 | 0.002357 | 0.00221 | -151.763 | 0.23041 | 8.4031 | 0.002576 | 0.003997 | 0.003884 | 5.5e-05 | 0.001141 | 0.31581 | 1.3079 | 2.5475 | 100.68 | 21.18 | 26.057 | 78.5317 | 83.7885 | 74.9826 | 0.006516 | 1.0141 | 0.024709 | 24.3653 | 0.064563 | 0.13707 | -0.000141 | -0.01042 | -0.01042 | -0.54327
2019-03-28 15:08 | 1553800107.8089 | 469 | 0.20949 | 0.000455 | 0 | 95.2627 | 94.4496 | 124.5628 | 21.2612 | 0.002304 | 0.002251 | -151.759 | 0.23085 | 8.4041 | 0.002559 | 0.003983 | 0.003876 | 5.5e-05 | 0.00114 | 0.33395 | 1.3197 | 2.5475 | 100.68 | 21.18 | 26.057 | 77.2626 | 83.3884 | 75.9866 | 0.006516 | 1.0112 | 0.024669 | 24.3956 | 0.076676 | 0.13733 | 0.000116 | 0.008647 | 0.008647 | 0.65593
2019-03-28 15:08 | 1553800108.7588 | 470 | 0.20952 | 0.000449 | 0 | 95.1766 | 94.4336 | 124.2093 | 21.2612 | 0.00231 | 0.002225 | -151.7503 | 0.23031 | 8.4031 | 0.00262 | 0.004221 | 0.004046 | 5.5e-05 | 0.001139 | 0.3298 | 1.3127 | 2.5473 | 100.68 | 21.18 | 26.061 | 77.3851 | 84.845 | 75.3607 | 0.006517 | 1.013 | 0.024656 | 24.3391 | 0.070103 | 0.13701 | -6.8e-05 | -0.005012 | -0.005012 | -1.1289
2019-03-28 15:08 | 1553800109.7599 | 471 | 0.2095 | 0.000452 | 0 | 95.1924 | 94.348 | 124.1593 | 21.2612 | 0.002316 | 0.002242 | -151.7415 | 0.23047 | 8.4042 | 0.002604 | 0.003736 | 0.00359 | 5.5e-05 | 0.001141 | 0.31797 | 1.3371 | 2.5472 | 100.68 | 21.18 | 26.065 | 77.5495 | 84.4657 | 75.7552 | 0.006518 | 1.0119 | 0.024637 | 24.348 | 0.079699 | 0.13711 | 4.1e-05 | 0.003047 | 0.003047 | 1.8586
2019-03-28 15:08 | 1553800110.7593 | 472 | 0.20952 | 0.000454 | -1e-06 | 95.1665 | 94.4156 | 124.2901 | 21.2612 | 0.002313 | 0.002224 | -151.762 | 0.23025 | 8.4001 | 0.002535 | 0.003398 | 0.003234 | 5.7e-05 | 0.001149 | 0.33539 | 1.2866 | 2.5474 | 100.68 | 21.18 | 26.079 | 77.4755 | 82.7924 | 75.3302 | 0.006521 | 1.0131 | 0.024665 | 24.346 | 0.070857 | 0.13697 | -8.3e-05 | -0.006162 | -0.006162 | -0.91805
2019-03-28 15:08 | 1553800111.7599 | 473 | 0.2095 | 0.000454 | 0 | 95.1418 | 94.2315 | 124.6464 | 21.2612 | 0.00231 | 0.002216 | -151.7804 | 0.23082 | 8.4022 | 0.002556 | 0.004004 | 0.003769 | 5.3e-05 | 0.00115 | 0.33179 | 1.326 | 2.5473 | 100.68 | 21.18 | 26.079 | 77.4077 | 83.3002 | 75.1455 | 0.006521 | 1.0136 | 0.024707 | 24.3744 | 0.085993 | 0.13731 | 3.7e-05 | 0.00276 | 0.00276 | 2.0545
2019-03-28 15:09 | 1553800112.7589 | 474 | 0.20952 | 0.000453 | -1e-06 | 95.157 | 94.4085 | 124.5849 | 21.2612 | 0.002343 | 0.002216 | -151.7986 | 0.23015 | 8.4018 | 0.002567 | 0.0044 | 0.004303 | 5.5e-05 | 0.00115 | 0.34827 | 1.3376 | 2.5474 | 100.68 | 21.18 | 26.097 | 78.1965 | 83.5733 | 75.1277 | 0.006526 | 1.0137 | 0.024701 | 24.3677 | 0.070636 | 0.13692 | -5.2e-05 | -0.003826 | -0.003826 | -1.4781
2019-03-28 15:09 | 1553800113.7593 | 475 | 0.20951 | 0.000453 | 0 | 95.0801 | 94.4458 | 124.7236 | 21.2612 | 0.002262 | 0.002199 | -151.8241 | 0.23055 | 8.4029 | 0.002575 | 0.003813 | 0.003623 | 5.3e-05 | 0.001152 | 0.34451 | 1.3444 | 2.5476 | 100.68 | 21.18 | 26.097 | 76.2422 | 83.7707 | 74.7387 | 0.006526 | 1.0148 | 0.024729 | 24.3677 | 0.059871 | 0.13716 | 1.1e-05 | 0.000881 | 0.000881 | 6.4298
2019-03-28 15:09 | 1553800114.758 | 476 | 0.20953 | 0.00045 | 0 | 95.0247 | 94.3273 | 124.7806 | 21.2612 | 0.002348 | 0.00219 | -151.8009 | 0.23087 | 8.4011 | 0.00258 | 0.003808 | 0.003633 | 5.3e-05 | 0.001149 | 0.33987 | 1.3456 | 2.5475 | 100.68 | 21.18 | 26.122 | 78.3097 | 83.8938 | 74.504 | 0.006532 | 1.0155 | 0.024743 | 24.365 | 0.065893 | 0.13734 | -9.4e-05 | -0.007011 | -0.007011 | -0.8091
2019-03-28 15:09 | 1553800115.7584 | 477 | 0.20952 | 0.000457 | 0 | 95.0289 | 94.1588 | 124.8731 | 21.2612 | 0.002285 | 0.002234 | -151.8006 | 0.2311 | 8.4018 | 0.002522 | 0.004206 | 0.003979 | 5.3e-05 | 0.00115 | 0.31669 | 1.3417 | 2.5474 | 100.68 | 21.18 | 26.132 | 76.7847 | 82.4921 | 75.5767 | 0.006535 | 1.0124 | 0.024714 | 24.4116 | 0.082278 | 0.13748 | -5.9e-05 | -0.004337 | -0.004337 | -1.3091
2019-03-28 15:09 | 1553800116.7591 | 478 | 0.2095 | 0.000455 | -1e-06 | 95.0345 | 94.2675 | 124.7067 | 21.2612 | 0.002291 | 0.002193 | -151.7859 | 0.22988 | 8.4017 | 0.002553 | 0.003592 | 0.003387 | 5.6e-05 | 0.001149 | 0.3246 | 1.3179 | 2.5474 | 100.68 | 21.18 | 26.137 | 76.9298 | 83.2441 | 74.579 | 0.006536 | 1.0153 | 0.024733 | 24.3604 | 0.07249 | 0.13676 | 7.2e-05 | 0.005369 | 0.005369 | 1.0519
2019-03-28 15:09 | 1553800117.7581 | 479 | 0.20952 | 0.000454 | -1e-06 | 95.0485 | 94.3422 | 124.8424 | 21.2612 | 0.002319 | 0.002225 | -151.7877 | 0.23051 | 8.4008 | 0.002556 | 0.003399 | 0.003254 | 5.3e-05 | 0.00115 | 0.32731 | 1.3298 | 2.5473 | 100.68 | 21.18 | 26.137 | 77.6064 | 83.3119 | 75.349 | 0.006536 | 1.013 | 0.024719 | 24.4006 | 0.066718 | 0.13713 | -6.1e-05 | -0.004511 | -0.004511 | -1.2554
2019-03-28 15:09 | 1553800118.7591 | 480 | 0.20951 | 0.000454 | -1e-06 | 95.0143 | 94.135 | 125.0921 | 21.2612 | 0.002261 | 0.002155 | -151.7883 | 0.23071 | 8.4008 | 0.002589 | 0.003566 | 0.003417 | 5.7e-05 | 0.001148 | 0.34418 | 1.2967 | 2.5475 | 100.68 | 21.18 | 26.141 | 76.2158 | 84.0913 | 73.666 | 0.006537 | 1.018 | 0.024804 | 24.366 | 0.083164 | 0.13725 | 1.6e-05 | 0.001237 | 0.001237 | 4.5808
2019-03-28 15:09 | 1553800119.7585 | 481 | 0.20952 | 0.000453 | 0 | 95.0634 | 94.2957 | 125.4846 | 21.2612 | 0.002314 | 0.002206 | -151.7994 | 0.23062 | 8.4021 | 0.002547 | 0.004156 | 0.003952 | 5.7e-05 | 0.001155 | 0.31957 | 1.2738 | 2.5474 | 100.68 | 21.18 | 26.139 | 77.5028 | 83.0996 | 74.8985 | 0.006536 | 1.0144 | 0.024798 | 24.4475 | 0.072533 | 0.1372 | -3.8e-05 | -0.002819 | -0.002819 | -2.0098
2019-03-28 15:09 | 1553800120.7592 | 482 | 0.2095 | 0.000454 | 0 | 95.0039 | 94.2365 | 125.3219 | 21.2612 | 0.002295 | 0.002169 | -151.7904 | 0.23167 | 8.4059 | 0.002566 | 0.003965 | 0.003847 | 5.5e-05 | 0.001154 | 0.32587 | 1.285 | 2.5473 | 100.68 | 21.18 | 26.131 | 77.0427 | 83.5507 | 74.0012 | 0.006534 | 1.017 | 0.024814 | 24.4001 | 0.072535 | 0.13782 | 2.8e-05 | 0.002095 | 0.002095 | 2.7165
2019-03-28 15:09 | 1553800121.7596 | 483 | 0.20951 | 0.000454 | 0 | 94.9455 | 94.118 | 125.1245 | 21.2612 | 0.002288 | 0.002186 | -151.7701 | 0.22997 | 8.4002 | 0.002484 | 0.004192 | 0.004013 | 5.7e-05 | 0.001155 | 0.34058 | 1.3218 | 2.5473 | 100.68 | 21.18 | 26.131 | 76.8635 | 81.5758 | 74.4137 | 0.006534 | 1.0158 | 0.02478 | 24.3954 | 0.078298 | 0.13681 | 2.2e-05 | 0.001655 | 0.001655 | 3.4136
2019-03-28 15:09 | 1553800122.7596 | 484 | 0.20951 | 0.000457 | -1e-06 | 94.8579 | 94.1805 | 125.0818 | 21.2612 | 0.002276 | 0.002194 | -151.7404 | 0.23176 | 8.4016 | 0.002623 | 0.003403 | 0.0032 | 5.5e-05 | 0.001156 | 0.33356 | 1.3223 | 2.5475 | 100.68 | 21.18 | 26.114 | 76.5791 | 84.9224 | 74.6172 | 0.00653 | 1.0152 | 0.024769 | 24.3984 | 0.064106 | 0.13787 | -2.3e-05 | -0.001661 | -0.001661 | -3.4287
2019-03-28 15:09 | 1553800123.7584 | 485 | 0.20951 | 0.000455 | -1e-06 | 94.7444 | 93.8459 | 125.3569 | 21.2612 | 0.002268 | 0.002185 | -151.7398 | 0.23096 | 8.4028 | 0.002614 | 0.004012 | 0.003822 | 5.7e-05 | 0.001159 | 0.31454 | 1.3585 | 2.5474 | 100.68 | 21.18 | 26.114 | 76.3919 | 84.7086 | 74.4012 | 0.00653 | 1.0158 | 0.024804 | 24.4176 | 0.085236 | 0.1374 | -7e-06 | -0.000447 | -0.000447 | -12.685
2019-03-28 15:09 | 1553800124.7593 | 486 | 0.20953 | 0.000454 | -1e-06 | 94.7376 | 93.9104 | 125.2419 | 21.2612 | 0.002252 | 0.002219 | -151.769 | 0.23114 | 8.4018 | 0.002555 | 0.003998 | 0.003783 | 5.8e-05 | 0.001163 | 0.32316 | 1.344 | 2.5474 | 100.68 | 21.17 | 26.083 | 76.0123 | 83.2752 | 75.2173 | 0.006518 | 1.0134 | 0.024763 | 24.435 | 0.07844 | 0.1375 | -0.000127 | -0.009421 | -0.009421 | -0.60278
2019-03-28 15:09 | 1553800125.7597 | 487 | 0.20949 | 0.000455 | 0 | 94.6938 | 94.0782 | 124.5492 | 21.2612 | 0.002251 | 0.002197 | -151.7945 | 0.23116 | 8.4045 | 0.002577 | 0.003816 | 0.003637 | 5.5e-05 | 0.001163 | 0.32781 | 1.3109 | 2.5474 | 100.68 | 21.17 | 26.074 | 75.9827 | 83.8097 | 74.6794 | 0.006516 | 1.015 | 0.024714 | 24.3485 | 0.058335 | 0.13752 | 9.4e-05 | 0.006988 | 0.006988 | 0.81273
2019-03-28 15:09 | 1553800126.7598 | 488 | 0.20949 | 0.000458 | 0 | 94.6208 | 93.7174 | 124.5053 | 21.2612 | 0.002291 | 0.002184 | -151.8026 | 0.23105 | 8.3996 | 0.002504 | 0.0043 | 0.004111 | 5.5e-05 | 0.001158 | 0.33035 | 1.2966 | 2.5472 | 100.68 | 21.17 | 26.059 | 76.9486 | 82.0662 | 74.3637 | 0.006512 | 1.0159 | 0.024721 | 24.3332 | 0.085807 | 0.13745 | 0.000143 | 0.010697 | 0.010697 | 0.53068
2019-03-28 15:09 | 1553800127.7594 | 489 | 0.2095 | 0.000455 | 0 | 94.5928 | 93.8887 | 124.8074 | 21.2612 | 0.002304 | 0.002196 | -151.8017 | 0.23088 | 8.4032 | 0.002611 | 0.00413 | 0.003966 | 5.4e-05 | 0.001158 | 0.32278 | 1.297 | 2.5474 | 100.68 | 21.17 | 26.052 | 77.2448 | 84.6374 | 74.6592 | 0.006511 | 1.0151 | 0.02474 | 24.3731 | 0.066833 | 0.13735 | 5.2e-05 | 0.003916 | 0.003916 | 1.4486
2019-03-28 15:09 | 1553800128.7599 | 490 | 0.20945 | 0.000453 | 0 | 94.5482 | 93.6566 | 124.7011 | 21.2612 | 0.002199 | 0.002162 | -151.8147 | 0.23057 | 8.4033 | 0.002559 | 0.004005 | 0.003822 | 5.4e-05 | 0.001158 | 0.31492 | 1.3018 | 2.5474 | 100.68 | 21.17 | 26.041 | 74.727 | 83.3783 | 73.8517 | 0.006508 | 1.0174 | 0.024758 | 24.3344 | 0.08474 | 0.13717 | 0.000356 | 0.02659 | 0.02659 | 0.21305
2019-03-28 15:09 | 1553800129.7585 | 491 | 0.20948 | 0.000453 | 0 | 94.4171 | 93.5364 | 124.1053 | 21.2612 | 0.00227 | 0.002185 | -151.7943 | 0.23065 | 8.404 | 0.002631 | 0.00376 | 0.003615 | 5.5e-05 | 0.001162 | 0.32604 | 1.2991 | 2.5474 | 100.68 | 21.17 | 26.036 | 76.4395 | 85.104 | 74.3911 | 0.006507 | 1.0158 | 0.02468 | 24.2951 | 0.083826 | 0.13721 | 0.000179 | 0.013355 | 0.013355 | 0.42434
2019-03-28 15:09 | 1553800130.7575 | 492 | 0.20951 | 0.000455 | -1e-06 | 94.3525 | 93.5523 | 124.2135 | 21.2612 | 0.002246 | 0.002144 | -151.7877 | 0.23099 | 8.4031 | 0.002633 | 0.003559 | 0.003375 | 5.3e-05 | 0.001159 | 0.34164 | 1.2755 | 2.5474 | 100.68 | 21.17 | 26.033 | 75.8665 | 85.1588 | 73.4104 | 0.006506 | 1.0187 | 0.024725 | 24.2713 | 0.076174 | 0.13742 | 2e-06 | 0.000158 | 0.000158 | 35.8506
2019-03-28 15:09 | 1553800131.7579 | 493 | 0.20954 | 0.000456 | -1e-06 | 94.3974 | 93.4433 | 124.2088 | 21.2612 | 0.002242 | 0.002145 | -151.7887 | 0.23088 | 8.3999 | 0.002605 | 0.002987 | 0.002839 | 5.5e-05 | 0.001159 | 0.32068 | 1.3058 | 2.5474 | 100.68 | 21.17 | 26.027 | 75.7605 | 84.4782 | 73.4308 | 0.006504 | 1.0187 | 0.024724 | 24.2716 | 0.09087 | 0.13735 | -0.000153 | -0.01139 | -0.01139 | -0.49802
2019-03-28 15:09 | 1553800132.7589 | 494 | 0.20954 | 0.000455 | -1e-06 | 94.3173 | 93.4939 | 124.3058 | 21.2612 | 0.002194 | 0.002213 | -151.8059 | 0.23095 | 8.4015 | 0.002596 | 0.003621 | 0.003391 | 5.2e-05 | 0.00116 | 0.3298 | 1.3168 | 2.5474 | 100.68 | 21.17 | 26.027 | 74.6047 | 84.2779 | 75.0676 | 0.006504 | 1.0139 | 0.024676 | 24.3383 | 0.07843 | 0.13739 | -0.00019 | -0.014072 | -0.014072 | -0.40324
2019-03-28 15:09 | 1553800133.7587 | 495 | 0.2095 | 0.000457 | 0 | 94.2363 | 93.3764 | 124.3595 | 21.2612 | 0.002237 | 0.002157 | -151.8062 | 0.23077 | 8.404 | 0.002675 | 0.003362 | 0.003221 | 5.6e-05 | 0.001165 | 0.30613 | 1.324 | 2.5473 | 100.68 | 21.17 | 26.023 | 75.6459 | 86.1733 | 73.7215 | 0.006503 | 1.0178 | 0.024729 | 24.2965 | 0.081996 | 0.13728 | 3.9e-05 | 0.002902 | 0.002902 | 1.954
2019-03-28 15:09 | 1553800134.7586 | 496 | 0.20953 | 0.000457 | -1e-06 | 94.1925 | 93.4294 | 124.2778 | 21.2612 | 0.002201 | 0.002147 | -151.7953 | 0.23106 | 8.399 | 0.002598 | 0.003522 | 0.003342 | 5.5e-05 | 0.001166 | 0.35043 | 1.2815 | 2.5475 | 100.68 | 21.17 | 26.023 | 74.7801 | 84.3161 | 73.4798 | 0.006503 | 1.0185 | 0.024729 | 24.28 | 0.072764 | 0.13746 | -0.000111 | -0.008258 | -0.008258 | -0.68745
2019-03-28 15:09 | 1553800135.7593 | 497 | 0.20952 | 0.000455 | -1e-06 | 94.1055 | 93.2551 | 124.7971 | 21.2612 | 0.002182 | 0.002174 | -151.7945 | 0.231 | 8.4008 | 0.002662 | 0.00371 | 0.0035 | 5.7e-05 | 0.001164 | 0.32853 | 1.3097 | 2.5472 | 100.68 | 21.17 | 26.023 | 74.3325 | 85.855 | 74.1275 | 0.006503 | 1.0166 | 0.024758 | 24.3534 | 0.081202 | 0.13742 | -5.2e-05 | -0.003807 | -0.003807 | -1.4909
2019-03-28 15:09 | 1553800136.7588 | 498 | 0.20952 | 0.000454 | 0 | 94.0353 | 93.2065 | 125.0273 | 21.2612 | 0.002223 | 0.002162 | -151.7756 | 0.23114 | 8.4012 | 0.002615 | 0.003521 | 0.003331 | 5.5e-05 | 0.001159 | 0.32084 | 1.3347 | 2.5475 | 100.68 | 21.17 | 26.023 | 75.3084 | 84.7343 | 73.8493 | 0.006503 | 1.0174 | 0.024791 | 24.3661 | 0.079179 | 0.1375 | -7e-05 | -0.0052 | -0.0052 | -1.0921
2019-03-28 15:09 | 1553800137.7588 | 499 | 0.20954 | 0.000455 | -1e-06 | 93.9393 | 93.2245 | 124.829 | 21.2612 | 0.00226 | 0.002163 | -151.7714 | 0.23107 | 8.4048 | 0.002648 | 0.003519 | 0.003302 | 5.1e-05 | 0.001153 | 0.32068 | 1.3778 | 2.5475 | 100.68 | 21.17 | 26.022 | 76.1862 | 85.5177 | 73.8649 | 0.006503 | 1.0174 | 0.02477 | 24.3473 | 0.068324 | 0.13746 | -0.000162 | -0.012026 | -0.012026 | -0.47209
2019-03-28 15:09 | 1553800138.7591 | 500 | 0.20952 | 0.000448 | -1e-06 | 93.8932 | 93.056 | 124.894 | 21.2612 | 0.002206 | 0.002193 | -151.7598 | 0.2317 | 8.4025 | 0.002585 | 0.004134 | 0.00402 | 5e-05 | 0.001147 | 0.3282 | 1.3917 | 2.5477 | 100.68 | 21.17 | 26.022 | 74.8923 | 83.9984 | 74.5953 | 0.006503 | 1.0152 | 0.024751 | 24.3793 | 0.080109 | 0.13784 | -6.7e-05 | -0.004918 | -0.004918 | -1.1575
2019-03-28 15:09 | 1553800139.7577 | 501 | 0.20951 | 0.000458 | -1e-06 | 93.8029 | 92.9405 | 125.3458 | 21.2612 | 0.00225 | 0.002122 | -151.7561 | 0.2309 | 8.4022 | 0.002584 | 0.003451 | 0.003359 | 5.5e-05 | 0.001141 | 0.33179 | 1.3505 | 2.5475 | 100.68 | 21.17 | 26.025 | 75.9609 | 83.9742 | 72.8772 | 0.006504 | 1.0203 | 0.024857 | 24.3629 | 0.082617 | 0.13736 | 1e-05 | 0.000801 | 0.000801 | 7.0855
2019-03-28 15:09 | 1553800140.7588 | 502 | 0.20954 | 0.000453 | 0 | 93.7192 | 92.7774 | 125.1747 | 21.2612 | 0.002252 | 0.002166 | -151.7627 | 0.23192 | 8.4024 | 0.002576 | 0.003293 | 0.003018 | 5.3e-05 | 0.001138 | 0.3235 | 1.3171 | 2.5471 | 100.68 | 21.17 | 26.025 | 75.9928 | 83.7847 | 73.931 | 0.006504 | 1.0172 | 0.024802 | 24.3833 | 0.09034 | 0.13797 | -0.000163 | -0.01212 | -0.01212 | -0.47015
2019-03-28 15:09 | 1553800141.7589 | 503 | 0.2095 | 0.000457 | 0 | 93.6623 | 92.6874 | 125.3421 | 21.2612 | 0.002236 | 0.00221 | -151.7715 | 0.23128 | 8.404 | 0.002589 | 0.004042 | 0.003972 | 5.3e-05 | 0.001135 | 0.31542 | 1.3325 | 2.5473 | 100.68 | 21.17 | 26.028 | 75.6257 | 84.103 | 74.9999 | 0.006505 | 1.0141 | 0.024781 | 24.4372 | 0.093592 | 0.13759 | 7e-05 | 0.005217 | 0.005217 | 1.0892
2019-03-28 15:09 | 1553800142.7582 | 504 | 0.20952 | 0.000456 | -1e-06 | 93.5724 | 92.7081 | 125.2576 | 21.2612 | 0.002229 | 0.002115 | -151.7851 | 0.23061 | 8.402 | 0.002565 | 0.003734 | 0.00348 | 5.1e-05 | 0.001124 | 0.3366 | 1.3338 | 2.5473 | 100.68 | 21.17 | 26.028 | 75.4487 | 83.5146 | 72.7199 | 0.006505 | 1.0207 | 0.024854 | 24.3488 | 0.082998 | 0.13719 | -6.3e-05 | -0.004654 | -0.004654 | -1.2173
2019-03-28 15:09 | 1553800143.758 | 505 | 0.2095 | 0.000459 | 0 | 93.4282 | 92.5763 | 125.071 | 21.2612 | 0.002227 | 0.002174 | -151.7741 | 0.23145 | 8.4018 | 0.002541 | 0.00358 | 0.003386 | 4.8e-05 | 0.001114 | 0.30829 | 1.3035 | 2.5476 | 100.68 | 21.17 | 26.032 | 75.4083 | 82.9499 | 74.1393 | 0.006506 | 1.0166 | 0.024785 | 24.3805 | 0.081934 | 0.13769 | 7e-05 | 0.005256 | 0.005256 | 1.0819
2019-03-28 15:09 | 1553800144.7584 | 506 | 0.20952 | 0.000456 | -1e-06 | 93.2474 | 92.4768 | 125.4393 | 21.2612 | 0.002212 | 0.002113 | -151.7766 | 0.23095 | 8.4022 | 0.002529 | 0.003814 | 0.003629 | 5e-05 | 0.001112 | 0.3115 | 1.2879 | 2.5471 | 100.68 | 21.17 | 26.032 | 75.0318 | 82.6481 | 72.655 | 0.006506 | 1.0209 | 0.024874 | 24.3642 | 0.074231 | 0.13739 | -4.5e-05 | -0.003291 | -0.003291 | -1.7244
2019-03-28 15:09 | 1553800145.7581 | 507 | 0.20949 | 0.000453 | -1e-06 | 93.1408 | 91.9781 | 125.1287 | 21.2612 | 0.002228 | 0.002178 | -151.7575 | 0.23017 | 8.4043 | 0.002566 | 0.004044 | 0.00392 | 5.2e-05 | 0.001116 | 0.3147 | 1.2871 | 2.5471 | 100.68 | 21.17 | 26.032 | 75.4254 | 83.5584 | 74.2328 | 0.006506 | 1.0163 | 0.024787 | 24.3895 | 0.11236 | 0.13693 | 0.000119 | 0.008914 | 0.008914 | 0.63438
2019-03-28 15:09 | 1553800146.7574 | 508 | 0.20949 | 0.00045 | -1e-06 | 92.9746 | 91.923 | 124.8647 | 21.2612 | 0.002181 | 0.002158 | -151.7727 | 0.23119 | 8.4021 | 0.002633 | 0.003957 | 0.003807 | 5.1e-05 | 0.001117 | 0.32278 | 1.2796 | 2.5473 | 100.68 | 21.17 | 26.032 | 74.3022 | 85.1689 | 73.7525 | 0.006506 | 1.0177 | 0.024778 | 24.3469 | 0.10174 | 0.13753 | 0.000141 | 0.010596 | 0.010596 | 0.53608
2019-03-28 15:09 | 1553800147.7581 | 509 | 0.20951 | 0.000447 | -1e-06 | 92.8072 | 91.7895 | 124.6992 | 21.2612 | 0.002221 | 0.002159 | -151.7887 | 0.2316 | 8.403 | 0.002529 | 0.003403 | 0.003257 | 5.3e-05 | 0.001116 | 0.33157 | 1.3031 | 2.5473 | 100.68 | 21.17 | 26.032 | 75.2579 | 82.6683 | 73.7619 | 0.006506 | 1.0177 | 0.024761 | 24.3311 | 0.098625 | 0.13778 | 3.5e-05 | 0.002618 | 0.002618 | 2.1739
2019-03-28 15:09 | 1553800148.7571 | 510 | 0.2095 | 0.00045 | -1e-06 | 92.7193 | 91.5982 | 124.8773 | 21.2612 | 0.002202 | 0.002104 | -151.7977 | 0.23178 | 8.406 | 0.002589 | 0.003936 | 0.003771 | 5.1e-05 | 0.001115 | 0.32444 | 1.3034 | 2.5473 | 100.68 | 21.17 | 26.031 | 74.8034 | 84.0975 | 72.4445 | 0.006505 | 1.0216 | 0.024826 | 24.3022 | 0.10881 | 0.13788 | 5.3e-05 | 0.004035 | 0.004035 | 1.4114
2019-03-28 15:09 | 1553800149.7585 | 511 | 0.20951 | 0.000451 | -1e-06 | 92.5668 | 91.4858 | 124.7564 | 21.2612 | 0.002282 | 0.002155 | -151.8048 | 0.23143 | 8.4031 | 0.002569 | 0.003803 | 0.003635 | 5.2e-05 | 0.001112 | 0.33035 | 1.2996 | 2.5473 | 100.68 | 21.17 | 26.031 | 76.7177 | 83.6099 | 73.6763 | 0.006505 | 1.0179 | 0.02477 | 24.3336 | 0.10507 | 0.13768 | 1.5e-05 | 0.001138 | 0.001138 | 4.9968
2019-03-28 15:09 | 1553800150.7572 | 512 | 0.20951 | 0.00045 | 0 | 92.2583 | 91.1946 | 124.8135 | 21.2612 | 0.00219 | 0.002181 | -151.8143 | 0.23138 | 8.4059 | 0.002572 | 0.003995 | 0.003791 | 5.1e-05 | 0.001112 | 0.33539 | 1.2938 | 2.5473 | 100.68 | 21.17 | 26.028 | 74.5158 | 83.6943 | 74.2974 | 0.006505 | 1.0161 | 0.024753 | 24.361 | 0.10372 | 0.13765 | -3e-06 | -0.000163 | -0.000163 | -34.8241
2019-03-28 15:09 | 1553800151.7585 | 513 | 0.20953 | 0.000452 | -1e-06 | 92.1594 | 90.9494 | 124.7933 | 21.2612 | 0.002265 | 0.002139 | -151.804 | 0.23111 | 8.4023 | 0.002622 | 0.004473 | 0.004206 | 4.9e-05 | 0.001116 | 0.31006 | 1.3003 | 2.5474 | 100.68 | 21.17 | 26.028 | 76.3179 | 84.9005 | 73.2874 | 0.006505 | 1.0191 | 0.024788 | 24.3236 | 0.11822 | 0.13748 | -8.9e-05 | -0.006583 | -0.006583 | -0.86259
2019-03-28 15:09 | 1553800152.7582 | 514 | 0.20952 | 0.000456 | 0 | 91.9452 | 90.644 | 124.7079 | 21.2612 | 0.002214 | 0.002146 | -151.786 | 0.23114 | 8.398 | 0.002638 | 0.003719 | 0.003562 | 5e-05 | 0.00112 | 0.31868 | 1.3494 | 2.5477 | 100.68 | 21.17 | 26.026 | 75.0785 | 85.2837 | 73.4642 | 0.006504 | 1.0186 | 0.024773 | 24.3215 | 0.12749 | 0.13751 | -7.3e-05 | -0.005396 | -0.005396 | -1.0525
2019-03-28 15:09 | 1553800153.7578 | 515 | 0.2095 | 0.000452 | -1e-06 | 91.6998 | 90.6575 | 124.7907 | 21.2612 | 0.002177 | 0.00215 | -151.7819 | 0.23104 | 8.4059 | 0.002708 | 0.003914 | 0.003769 | 5.1e-05 | 0.00112 | 0.32211 | 1.3694 | 2.5475 | 100.68 | 21.17 | 26.026 | 74.1907 | 86.976 | 73.5483 | 0.006504 | 1.0183 | 0.024778 | 24.3325 | 0.10225 | 0.13745 | 9.5e-05 | 0.007169 | 0.007169 | 0.79177
2019-03-28 15:09 | 1553800154.7583 | 516 | 0.20951 | 0.000448 | -1e-06 | 91.6158 | 90.3287 | 124.7841 | 21.2612 | 0.002226 | 0.002111 | -151.7721 | 0.23108 | 8.4044 | 0.00266 | 0.003183 | 0.002987 | 5.1e-05 | 0.001119 | 0.29607 | 1.3479 | 2.5475 | 100.68 | 21.17 | 26.025 | 75.3816 | 85.8081 | 72.6043 | 0.006504 | 1.0211 | 0.024811 | 24.2987 | 0.12655 | 0.13747 | 2.7e-05 | 0.002079 | 0.002079 | 2.7304
2019-03-28 15:09 | 1553800155.7587 | 517 | 0.20952 | 0.00045 | -1e-06 | 91.4532 | 90.2251 | 125.0879 | 21.2612 | 0.002195 | 0.002136 | -151.7747 | 0.23116 | 8.4032 | 0.00265 | 0.004457 | 0.004221 | 5.2e-05 | 0.001121 | 0.33284 | 1.313 | 2.5471 | 100.68 | 21.17 | 26.025 | 74.6388 | 85.5793 | 73.2117 | 0.006504 | 1.0193 | 0.024819 | 24.3496 | 0.12093 | 0.13752 | -3.7e-05 | -0.00274 | -0.00274 | -2.0729
2019-03-28 15:09 | 1553800156.7587 | 518 | 0.2095 | 0.000451 | -1e-06 | 91.312 | 89.9251 | 125.1348 | 21.2612 | 0.002137 | 0.002196 | -151.7654 | 0.23116 | 8.4026 | 0.002634 | 0.003498 | 0.00333 | 4.9e-05 | 0.001116 | 0.29911 | 1.3373 | 2.5472 | 100.68 | 21.17 | 26.029 | 73.2444 | 85.1815 | 74.6669 | 0.006505 | 1.015 | 0.024772 | 24.4053 | 0.1369 | 0.13752 | 5e-05 | 0.003793 | 0.003793 | 1.4973
2019-03-28 15:09 | 1553800157.7574 | 519 | 0.20951 | 0.00045 | -1e-06 | 91.1306 | 89.8225 | 125.255 | 21.2612 | 0.002188 | 0.002093 | -151.7469 | 0.23185 | 8.401 | 0.002623 | 0.003733 | 0.003555 | 5.1e-05 | 0.001123 | 0.33157 | 1.3155 | 2.5474 | 100.68 | 21.17 | 26.029 | 74.4706 | 84.9169 | 72.1857 | 0.006505 | 1.0223 | 0.024873 | 24.3298 | 0.12932 | 0.13793 | 3e-05 | 0.002305 | 0.002305 | 2.4718
2019-03-28 15:09 | 1553800158.7584 | 520 | 0.2095 | 0.000449 | -1e-06 | 90.9391 | 89.5235 | 125.1916 | 21.2612 | 0.002231 | 0.002129 | -151.7398 | 0.23124 | 8.4038 | 0.002624 | 0.002887 | 0.002672 | 5.4e-05 | 0.001127 | 0.31902 | 1.306 | 2.5474 | 100.68 | 21.17 | 26.041 | 75.4994 | 84.9534 | 73.0471 | 0.006508 | 1.0198 | 0.024836 | 24.3539 | 0.14032 | 0.13757 | 6.6e-05 | 0.004978 | 0.004978 | 1.1413
2019-03-28 15:09 | 1553800159.7584 | 521 | 0.20954 | 0.000458 | -1e-06 | 90.7422 | 89.4277 | 125.2381 | 21.2612 | 0.002204 | 0.002102 | -151.7545 | 0.23117 | 8.4027 | 0.002567 | 0.004258 | 0.004113 | 5.2e-05 | 0.001139 | 0.32781 | 1.3097 | 2.5474 | 100.68 | 21.17 | 26.041 | 74.8548 | 83.5678 | 72.3962 | 0.006508 | 1.0217 | 0.024864 | 24.3355 | 0.13051 | 0.13752 | -0.000203 | -0.015127 | -0.015127 | -0.37547
2019-03-28 15:09 | 1553800160.7565 | 522 | 0.2095 | 0.000458 | -1e-06 | 90.5069 | 89.0187 | 125.3327 | 21.2612 | 0.002225 | 0.002091 | -151.7429 | 0.23148 | 8.4026 | 0.002531 | 0.004025 | 0.003913 | 5.3e-05 | 0.001148 | 0.32748 | 1.2636 | 2.5473 | 100.68 | 21.17 | 26.055 | 75.3497 | 82.6965 | 72.1235 | 0.006511 | 1.0225 | 0.024883 | 24.3351 | 0.1483 | 0.13771 | 4.2e-05 | 0.00316 | 0.00316 | 1.7996
2019-03-28 15:09 | 1553800161.7565 | 523 | 0.20949 | 0.000451 | -1e-06 | 90.3368 | 88.858 | 125.4719 | 21.2612 | 0.002192 | 0.002092 | -151.7643 | 0.23141 | 8.4048 | 0.002614 | 0.004495 | 0.00427 | 5.5e-05 | 0.001151 | 0.31902 | 1.2832 | 2.5473 | 100.68 | 21.17 | 26.055 | 74.5593 | 84.6913 | 72.1468 | 0.006511 | 1.0224 | 0.024896 | 24.3494 | 0.14763 | 0.13767 | 0.000125 | 0.009441 | 0.009441 | 0.60223
2019-03-28 15:09 | 1553800162.7575 | 524 | 0.20954 | 0.000455 | 0 | 90.0348 | 88.6763 | 125.2651 | 21.2612 | 0.002152 | 0.0021 | -151.7776 | 0.23106 | 8.4042 | 0.002587 | 0.003724 | 0.00359 | 5.5e-05 | 0.001152 | 0.3246 | 1.2987 | 2.5474 | 100.68 | 21.17 | 26.063 | 73.61 | 84.0422 | 72.3472 | 0.006513 | 1.0218 | 0.024868 | 24.3364 | 0.13599 | 0.13746 | -0.000164 | -0.012288 | -0.012288 | -0.46199
2019-03-28 15:09 | 1553800163.7576 | 525 | 0.20954 | 0.000455 | -1e-06 | 89.8335 | 88.359 | 125.0698 | 21.2612 | 0.002121 | 0.00212 | -151.7872 | 0.2314 | 8.4035 | 0.002567 | 0.003676 | 0.00349 | 5.8e-05 | 0.001156 | 0.32372 | 1.2866 | 2.5473 | 100.68 | 21.17 | 26.064 | 72.846 | 83.5841 | 72.8383 | 0.006514 | 1.0204 | 0.024831 | 24.3347 | 0.14803 | 0.13766 | -0.000185 | -0.013801 | -0.013801 | -0.41196
2019-03-28 15:09 | 1553800164.7566 | 526 | 0.20952 | 0.000458 | -1e-06 | 89.7959 | 88.2003 | 124.833 | 21.2612 | 0.002169 | 0.002127 | -151.7952 | 0.23094 | 8.404 | 0.002616 | 0.003508 | 0.003348 | 5.4e-05 | 0.001156 | 0.31924 | 1.3303 | 2.5474 | 100.68 | 21.17 | 26.065 | 74.006 | 84.7545 | 72.9942 | 0.006514 | 1.0199 | 0.024802 | 24.3172 | 0.16037 | 0.13738 | -2.5e-05 | -0.001832 | -0.001832 | -3.0969
2019-03-28 15:09 | 1553800165.7579 | 527 | 0.20955 | 0.000454 | -1e-06 | 89.7986 | 88.292 | 124.655 | 21.2612 | 0.002192 | 0.00211 | -151.7884 | 0.23126 | 8.403 | 0.002563 | 0.003589 | 0.003473 | 5.3e-05 | 0.001161 | 0.3235 | 1.3058 | 2.5475 | 100.68 | 21.17 | 26.066 | 74.5648 | 83.4843 | 72.588 | 0.006514 | 1.0211 | 0.024799 | 24.2856 | 0.15134 | 0.13757 | -0.00021 | -0.015647 | -0.015647 | -0.36313
2019-03-28 15:09 | 1553800166.7567 | 528 | 0.20954 | 0.000451 | -1e-06 | 89.5993 | 88.0658 | 124.8327 | 21.2612 | 0.002117 | 0.002133 | -151.8118 | 0.23197 | 8.4052 | 0.002582 | 0.003788 | 0.003551 | 5.1e-05 | 0.00116 | 0.30397 | 1.3738 | 2.5475 | 100.68 | 21.17 | 26.066 | 72.7689 | 83.9275 | 73.1454 | 0.006514 | 1.0195 | 0.024797 | 24.3224 | 0.15442 | 0.138 | -0.000146 | -0.010868 | -0.010868 | -0.52442
2019-03-28 15:09 | 1553800167.7566 | 529 | 0.20949 | 0.000459 | -1e-06 | 89.6979 | 88.1133 | 124.7569 | 21.2612 | 0.002187 | 0.002116 | -151.8316 | 0.23062 | 8.402 | 0.002567 | 0.003434 | 0.003228 | 5.3e-05 | 0.00116 | 0.31581 | 1.3429 | 2.5476 | 100.68 | 21.17 | 26.065 | 74.4408 | 83.581 | 72.7352 | 0.006514 | 1.0207 | 0.024804 | 24.3006 | 0.15943 | 0.1372 | 0.000158 | 0.011878 | 0.011878 | 0.47702
2019-03-28 15:09 | 1553800168.7568 | 530 | 0.20943 | 0.000453 | 0 | 89.681 | 88.1432 | 124.5985 | 21.2612 | 0.002176 | 0.002046 | -151.827 | 0.23142 | 8.4056 | 0.002634 | 0.004156 | 0.004028 | 5e-05 | 0.00115 | 0.32117 | 1.392 | 2.5476 | 100.68 | 21.17 | 26.064 | 74.1838 | 85.1853 | 71.0561 | 0.006514 | 1.0257 | 0.024848 | 24.2262 | 0.1547 | 0.13767 | 0.000502 | 0.037602 | 0.037602 | 0.15121
2019-03-28 15:09 | 1553800169.7578 | 531 | 0.2095 | 0.000452 | 0 | 89.6816 | 88.3318 | 124.7661 | 21.2612 | 0.002157 | 0.002118 | -151.8418 | 0.2306 | 8.4017 | 0.002608 | 0.003795 | 0.003598 | 5.3e-05 | 0.001144 | 0.33229 | 1.3282 | 2.5474 | 100.68 | 21.17 | 26.059 | 73.7215 | 84.5494 | 72.7907 | 0.006512 | 1.0205 | 0.024803 | 24.3035 | 0.13564 | 0.13718 | 7.9e-05 | 0.005929 | 0.005929 | 0.95557
2019-03-28 15:09 | 1553800170.7578 | 532 | 0.20939 | 0.000457 | 0 | 89.697 | 88.3776 | 124.91 | 21.2612 | 0.00216 | 0.002097 | -151.8581 | 0.231 | 8.4031 | 0.002588 | 0.00413 | 0.003947 | 5.4e-05 | 0.001151 | 0.34363 | 1.3059 | 2.5474 | 100.68 | 21.17 | 26.059 | 73.7823 | 84.0795 | 72.2792 | 0.006512 | 1.022 | 0.024835 | 24.2995 | 0.13255 | 0.13742 | 0.000716 | 0.053561 | 0.053561 | 0.10596
2019-03-28 15:10 | 1553800171.7568 | 533 | 0.20956 | 0.000454 | 0 | 89.7983 | 88.3875 | 124.7154 | 21.2612 | 0.002126 | 0.002086 | -151.8704 | 0.22998 | 8.4031 | 0.002607 | 0.003783 | 0.003639 | 5.2e-05 | 0.00115 | 0.34003 | 1.3057 | 2.5473 | 100.68 | 21.18 | 26.046 | 72.987 | 84.5306 | 72.0034 | 0.006513 | 1.0229 | 0.024826 | 24.2709 | 0.14164 | 0.13682 | -0.00029 | -0.021655 | -0.021655 | -0.26093
2019-03-28 15:10 | 1553800172.7572 | 534 | 0.20947 | 0.000457 | -1e-06 | 89.9696 | 88.6596 | 125.4078 | 21.2612 | 0.00216 | 0.002062 | -151.8609 | 0.23075 | 8.4047 | 0.002657 | 0.003794 | 0.003626 | 5.3e-05 | 0.001151 | 0.31852 | 1.2795 | 2.5472 | 100.68 | 21.17 | 26.041 | 73.7806 | 85.7253 | 71.449 | 0.006508 | 1.0245 | 0.024915 | 24.3186 | 0.13119 | 0.13728 | 0.000238 | 0.017893 | 0.017893 | 0.31686
2019-03-28 15:10 | 1553800173.7569 | 535 | 0.20954 | 0.000453 | 0 | 90.1974 | 88.9161 | 125.3083 | 21.2612 | 0.002212 | 0.002142 | -151.8607 | 0.23017 | 8.4037 | 0.002666 | 0.003485 | 0.003312 | 5e-05 | 0.001153 | 0.31094 | 1.3047 | 2.5473 | 100.68 | 21.17 | 26.034 | 75.0465 | 85.9524 | 73.3645 | 0.006506 | 1.0188 | 0.024836 | 24.3764 | 0.12798 | 0.13693 | -0.000158 | -0.011793 | -0.011793 | -0.47953
2019-03-28 15:10 | 1553800174.7583 | 536 | 0.20951 | 0.000457 | -1e-06 | 90.3907 | 89.1645 | 125.3388 | 21.2612 | 0.002164 | 0.002111 | -151.8578 | 0.23077 | 8.402 | 0.002646 | 0.00356 | 0.003391 | 5.1e-05 | 0.001151 | 0.35026 | 1.2873 | 2.5473 | 100.68 | 21.17 | 26.029 | 73.8868 | 85.4701 | 72.6105 | 0.006505 | 1.0211 | 0.024866 | 24.3529 | 0.12218 | 0.13729 | -1.7e-05 | -0.001203 | -0.001203 | -4.7128
2019-03-28 15:10 | 1553800175.757 | 537 | 0.20954 | 0.00045 | -1e-06 | 90.5516 | 89.1036 | 125.1069 | 21.2612 | 0.002163 | 0.002081 | -151.867 | 0.23137 | 8.4031 | 0.002648 | 0.003464 | 0.003367 | 5.1e-05 | 0.001149 | 0.31205 | 1.3047 | 2.5473 | 100.68 | 21.17 | 26.029 | 73.8719 | 85.5091 | 71.8856 | 0.006505 | 1.0232 | 0.024869 | 24.3048 | 0.14418 | 0.13764 | -0.000154 | -0.011458 | -0.011458 | -0.49612
2019-03-28 15:10 | 1553800176.757 | 538 | 0.20953 | 0.000451 | -1e-06 | 90.8122 | 89.3805 | 124.7766 | 21.2612 | 0.002125 | 0.00207 | -151.8658 | 0.231 | 8.4028 | 0.00264 | 0.003467 | 0.003203 | 5.1e-05 | 0.001152 | 0.31686 | 1.3422 | 2.5473 | 100.68 | 21.17 | 26.021 | 72.9575 | 85.3383 | 71.6182 | 0.006503 | 1.024 | 0.024846 | 24.2633 | 0.14214 | 0.13742 | -0.0001 | -0.00744 | -0.00744 | -0.76285
2019-03-28 15:10 | 1553800177.7559 | 539 | 0.20952 | 0.000452 | -1e-06 | 91.0088 | 89.7923 | 124.7773 | 21.2612 | 0.002193 | 0.002083 | -151.8539 | 0.23037 | 8.4033 | 0.00264 | 0.00365 | 0.003537 | 5.1e-05 | 0.001154 | 0.31006 | 1.3266 | 2.5474 | 100.68 | 21.17 | 26.021 | 74.5936 | 85.3234 | 71.9409 | 0.006503 | 1.0231 | 0.024834 | 24.2747 | 0.12037 | 0.13705 | -5.9e-05 | -0.004342 | -0.004342 | -1.3034
2019-03-28 15:10 | 1553800178.757 | 540 | 0.20953 | 0.000447 | 0 | 91.1613 | 89.9159 | 124.9344 | 21.2612 | 0.002158 | 0.00211 | -151.8622 | 0.23088 | 8.402 | 0.00262 | 0.003877 | 0.003687 | 4.6e-05 | 0.001158 | 0.3235 | 1.3504 | 2.5473 | 100.68 | 21.17 | 26.017 | 73.7455 | 84.8551 | 72.588 | 0.006502 | 1.0211 | 0.024827 | 24.3127 | 0.12303 | 0.13735 | -0.000117 | -0.008755 | -0.008755 | -0.64789
2019-03-28 15:10 | 1553800179.7575 | 541 | 0.20951 | 0.00045 | 0 | 91.3829 | 90.1102 | 124.4257 | 21.2612 | 0.002196 | 0.002035 | -151.8377 | 0.23103 | 8.4034 | 0.002626 | 0.003731 | 0.00358 | 5.2e-05 | 0.00116 | 0.32731 | 1.3161 | 2.5473 | 100.68 | 21.17 | 26.017 | 74.6585 | 84.9809 | 70.7793 | 0.006502 | 1.0265 | 0.024841 | 24.1997 | 0.12544 | 0.13744 | 4e-06 | 0.000305 | 0.000305 | 18.6187
2019-03-28 15:10 | 1553800180.7571 | 542 | 0.20956 | 0.000452 | 0 | 91.561 | 90.2896 | 124.4515 | 21.2612 | 0.002224 | 0.002024 | -151.8172 | 0.23144 | 8.4013 | 0.002642 | 0.003699 | 0.003546 | 4.6e-05 | 0.001158 | 0.30917 | 1.293 | 2.5473 | 100.68 | 21.17 | 26.013 | 75.3271 | 85.3686 | 70.5337 | 0.006501 | 1.0272 | 0.024853 | 24.1935 | 0.12507 | 0.13769 | -0.000291 | -0.021761 | -0.021761 | -0.26131
2019-03-28 15:10 | 1553800181.7563 | 543 | 0.20952 | 0.00045 | 0 | 91.8446 | 90.5441 | 124.713 | 21.2612 | 0.002094 | 0.002111 | -151.8091 | 0.23023 | 8.4014 | 0.002664 | 0.003678 | 0.003472 | 4.6e-05 | 0.001155 | 0.34755 | 1.2857 | 2.5473 | 100.68 | 21.17 | 26.013 | 72.2122 | 85.9173 | 72.6153 | 0.006501 | 1.0211 | 0.024804 | 24.2922 | 0.12756 | 0.13696 | -6.5e-05 | -0.00483 | -0.00483 | -1.171
2019-03-28 15:10 | 1553800182.7562 | 544 | 0.20955 | 0.000451 | -1e-06 | 91.872 | 90.733 | 124.0194 | 21.2612 | 0.002143 | 0.002101 | -151.8302 | 0.23077 | 8.4001 | 0.002612 | 0.003784 | 0.003618 | 4.7e-05 | 0.001153 | 0.32244 | 1.2923 | 2.5473 | 100.68 | 21.16 | 26.012 | 73.376 | 84.6663 | 72.3729 | 0.006497 | 1.0218 | 0.024743 | 24.216 | 0.11158 | 0.13728 | -0.000193 | -0.014379 | -0.014379 | -0.39432
2019-03-28 15:10 | 1553800183.7572 | 545 | 0.20955 | 0.000451 | -1e-06 | 92.0483 | 90.9324 | 124.057 | 21.2612 | 0.002169 | 0.002054 | -151.7934 | 0.23033 | 8.4034 | 0.002717 | 0.003647 | 0.003441 | 4.6e-05 | 0.001155 | 0.32139 | 1.3098 | 2.5473 | 100.68 | 21.16 | 26.012 | 74.0168 | 87.1836 | 71.2503 | 0.006497 | 1.0251 | 0.024787 | 24.1803 | 0.1091 | 0.13703 | -0.000245 | -0.018254 | -0.018254 | -0.31002
2019-03-28 15:10 | 1553800184.7576 | 546 | 0.20957 | 0.000451 | -1e-06 | 92.1918 | 90.9504 | 124.0985 | 21.2612 | 0.002155 | 0.002056 | -151.8104 | 0.2304 | 8.4035 | 0.002685 | 0.00401 | 0.003845 | 4.5e-05 | 0.001156 | 0.33732 | 1.3162 | 2.5474 | 100.68 | 21.17 | 26.011 | 73.6684 | 86.4184 | 71.2892 | 0.0065 | 1.025 | 0.02479 | 24.1857 | 0.12127 | 0.13706 | -0.000331 | -0.024688 | -0.024688 | -0.22929
2019-03-28 15:10 | 1553800185.7573 | 547 | 0.20955 | 0.000449 | -1e-06 | 92.3512 | 91.3771 | 124.5682 | 21.2612 | 0.002128 | 0.002099 | -151.8185 | 0.23092 | 8.4025 | 0.00269 | 0.004215 | 0.004051 | 4.6e-05 | 0.001156 | 0.3178 | 1.3011 | 2.5474 | 100.68 | 21.17 | 26.011 | 73.0252 | 86.5419 | 72.3378 | 0.0065 | 1.0219 | 0.024799 | 24.2683 | 0.094848 | 0.13737 | -0.000218 | -0.016251 | -0.016251 | -0.34913
2019-03-28 15:10 | 1553800186.7562 | 548 | 0.20956 | 0.00045 | 0 | 92.4659 | 91.4406 | 124.691 | 21.2612 | 0.002143 | 0.002086 | -151.8203 | 0.23044 | 8.4012 | 0.00261 | 0.003374 | 0.003185 | 4.9e-05 | 0.001157 | 0.30541 | 1.3044 | 2.5474 | 100.68 | 21.17 | 26.011 | 73.3864 | 84.6172 | 72.0259 | 0.0065 | 1.0228 | 0.024823 | 24.2693 | 0.099731 | 0.13709 | -0.000263 | -0.019665 | -0.019665 | -0.28792
2019-03-28 15:10 | 1553800187.7573 | 549 | 0.20956 | 0.00045 | -1e-06 | 92.5198 | 91.5724 | 124.9585 | 21.2612 | 0.002136 | 0.002112 | -151.8156 | 0.2314 | 8.4029 | 0.002719 | 0.004346 | 0.004157 | 4.6e-05 | 0.001155 | 0.31437 | 1.378 | 2.5473 | 100.68 | 21.16 | 26.015 | 73.2256 | 87.2351 | 72.648 | 0.006497 | 1.021 | 0.024827 | 24.3172 | 0.09206 | 0.13766 | -0.000277 | -0.020699 | -0.020699 | -0.27468
2019-03-28 15:10 | 1553800188.7574 | 550 | 0.20955 | 0.000454 | 0 | 92.7789 | 91.6482 | 124.833 | 21.2612 | 0.002189 | 0.00205 | -151.8295 | 0.23118 | 8.401 | 0.002644 | 0.00374 | 0.003586 | 4.4e-05 | 0.001152 | 0.31741 | 1.314 | 2.5473 | 100.68 | 21.16 | 26.015 | 74.4978 | 85.4273 | 71.1465 | 0.006497 | 1.0254 | 0.024868 | 24.2522 | 0.10968 | 0.13753 | -0.00024 | -0.01797 | -0.01797 | -0.31607
2019-03-28 15:10 | 1553800189.7567 | 551 | 0.20955 | 0.000448 | 0 | 92.8555 | 91.8119 | 125.0337 | 21.2612 | 0.002109 | 0.002039 | -151.825 | 0.23144 | 8.4041 | 0.002688 | 0.004457 | 0.004244 | 4.5e-05 | 0.001155 | 0.31526 | 1.313 | 2.5474 | 100.68 | 21.16 | 26.016 | 72.5598 | 86.4786 | 70.8884 | 0.006498 | 1.0262 | 0.024898 | 24.2625 | 0.10109 | 0.13769 | -0.000197 | -0.014772 | -0.014772 | -0.38494
2019-03-28 15:10 | 1553800190.7575 | 552 | 0.20956 | 0.00045 | 0 | 92.968 | 92.2532 | 124.7586 | 21.2612 | 0.002139 | 0.002051 | -151.8466 | 0.23137 | 8.4011 | 0.002688 | 0.003791 | 0.003615 | 4.4e-05 | 0.001157 | 0.31238 | 1.3399 | 2.5474 | 100.68 | 21.16 | 26.015 | 73.2842 | 86.4942 | 71.184 | 0.006497 | 1.0253 | 0.02486 | 24.2463 | 0.069039 | 0.13764 | -0.000265 | -0.019875 | -0.019875 | -0.28601
2019-03-28 15:10 | 1553800191.7555 | 553 | 0.20954 | 0.000453 | -1e-06 | 93.0928 | 92.2675 | 124.4778 | 21.2612 | 0.00213 | 0.002077 | -151.8284 | 0.23127 | 8.402 | 0.002636 | 0.004286 | 0.004112 | 4.3e-05 | 0.00116 | 0.30431 | 1.3659 | 2.5475 | 100.68 | 21.16 | 26.014 | 73.0829 | 85.2281 | 71.7912 | 0.006497 | 1.0235 | 0.02481 | 24.2403 | 0.079654 | 0.13758 | -0.000186 | -0.013846 | -0.013846 | -0.41037
2019-03-28 15:10 | 1553800192.7564 | 554 | 0.20952 | 0.000452 | -1e-06 | 93.2099 | 92.1275 | 124.9269 | 21.2612 | 0.002155 | 0.002024 | -151.8399 | 0.23111 | 8.4049 | 0.002629 | 0.003899 | 0.003731 | 4.6e-05 | 0.001164 | 0.33754 | 1.3799 | 2.5474 | 100.68 | 21.16 | 26.014 | 73.6732 | 85.0636 | 70.5361 | 0.006497 | 1.0272 | 0.0249 | 24.2398 | 0.10447 | 0.13749 | -5.9e-05 | -0.00438 | -0.00438 | -1.2965
2019-03-28 15:10 | 1553800193.7555 | 555 | 0.20957 | 0.000453 | -1e-06 | 93.2495 | 92.3232 | 124.9299 | 21.2612 | 0.002106 | 0.002037 | -151.8738 | 0.232 | 8.4031 | 0.002642 | 0.004151 | 0.003978 | 4.6e-05 | 0.001165 | 0.32477 | 1.3311 | 2.5474 | 100.68 | 21.15 | 26.013 | 72.489 | 85.3686 | 70.8331 | 0.006493 | 1.0263 | 0.024889 | 24.2505 | 0.089296 | 0.13802 | -0.000331 | -0.02474 | -0.02474 | -0.2304
2019-03-28 15:10 | 1553800194.7566 | 556 | 0.20956 | 0.000445 | -1e-06 | 93.419 | 92.3504 | 124.7687 | 21.2612 | 0.002176 | 0.002113 | -151.8524 | 0.23145 | 8.3995 | 0.002654 | 0.003793 | 0.003661 | 4.7e-05 | 0.001158 | 0.29911 | 1.2933 | 2.5474 | 100.68 | 21.15 | 26.013 | 74.1859 | 85.6731 | 72.6543 | 0.006493 | 1.0209 | 0.024808 | 24.2989 | 0.1029 | 0.13769 | -0.000253 | -0.018907 | -0.018907 | -0.30077
2019-03-28 15:10 | 1553800195.7566 | 557 | 0.20953 | 0.000449 | 0 | 93.5646 | 92.5736 | 124.5513 | 21.2612 | 0.002145 | 0.002095 | -151.838 | 0.23153 | 8.4046 | 0.002653 | 0.004178 | 0.003984 | 4.7e-05 | 0.001165 | 0.31902 | 1.2818 | 2.5473 | 100.68 | 21.15 | 26.014 | 73.4267 | 85.6443 | 72.2379 | 0.006493 | 1.0222 | 0.024801 | 24.2631 | 0.095247 | 0.13774 | -0.000114 | -0.008485 | -0.008485 | -0.67043
2019-03-28 15:10 | 1553800196.7556 | 558 | 0.20953 | 0.00045 | 0 | 93.6665 | 92.7108 | 125.0614 | 21.2612 | 0.002173 | 0.002076 | -151.8327 | 0.23081 | 8.4008 | 0.002613 | 0.004313 | 0.004191 | 4.9e-05 | 0.001168 | 0.33085 | 1.3328 | 2.5474 | 100.68 | 21.15 | 26.014 | 74.1004 | 84.6687 | 71.7797 | 0.006493 | 1.0235 | 0.024868 | 24.2966 | 0.091729 | 0.13731 | -8e-05 | -0.005976 | -0.005976 | -0.94893
2019-03-28 15:10 | 1553800197.7556 | 559 | 0.20955 | 0.000448 | 0 | 93.758 | 92.842 | 124.8755 | 21.2612 | 0.002148 | 0.002066 | -151.8359 | 0.23161 | 8.4029 | 0.002626 | 0.003967 | 0.003842 | 4.6e-05 | 0.001169 | 0.32101 | 1.3331 | 2.5475 | 100.68 | 21.15 | 26.017 | 73.4969 | 84.9987 | 71.5362 | 0.006494 | 1.0243 | 0.024859 | 24.27 | 0.087814 | 0.13779 | -0.00022 | -0.01642 | -0.01642 | -0.34656
2019-03-28 15:10 | 1553800198.757 | 560 | 0.20954 | 0.000451 | -1e-06 | 93.8896 | 93.1219 | 125.0327 | 21.2612 | 0.002142 | 0.002028 | -151.8149 | 0.23197 | 8.4027 | 0.00263 | 0.004292 | 0.004044 | 4.8e-05 | 0.001165 | 0.33395 | 1.3302 | 2.5476 | 100.68 | 21.15 | 26.017 | 73.3691 | 85.0987 | 70.6125 | 0.006494 | 1.027 | 0.024908 | 24.2527 | 0.073436 | 0.138 | -0.000135 | -0.010052 | -0.010052 | -0.56698
2019-03-28 15:10 | 1553800199.7557 | 561 | 0.20955 | 0.000453 | 0 | 94.0495 | 93.0839 | 125.0092 | 21.2612 | 0.002121 | 0.002029 | -151.8207 | 0.23121 | 8.4011 | 0.002592 | 0.003759 | 0.003638 | 4.8e-05 | 0.001159 | 0.33085 | 1.3447 | 2.5473 | 100.68 | 21.15 | 26.022 | 72.8498 | 84.1677 | 70.6389 | 0.006495 | 1.0269 | 0.024904 | 24.2514 | 0.092316 | 0.13755 | -0.000245 | -0.018338 | -0.018338 | -0.30978
2019-03-28 15:10 | 1553800200.7561 | 562 | 0.20955 | 0.00045 | 0 | 94.1041 | 93.2928 | 124.883 | 21.2612 | 0.002144 | 0.002028 | -151.8135 | 0.23169 | 8.4044 | 0.002544 | 0.004047 | 0.003872 | 4.1e-05 | 0.001157 | 0.30757 | 1.375 | 2.5476 | 100.68 | 21.15 | 26.022 | 73.4197 | 83.0179 | 70.6233 | 0.006495 | 1.027 | 0.024892 | 24.2386 | 0.077444 | 0.13783 | -0.000201 | -0.015078 | -0.015078 | -0.37754
2019-03-28 15:10 | 1553800201.7557 | 563 | 0.20955 | 0.000449 | -1e-06 | 94.2691 | 93.5316 | 125.043 | 21.2612 | 0.002065 | 0.002095 | -151.8412 | 0.23057 | 8.4021 | 0.002608 | 0.003822 | 0.00367 | 4.6e-05 | 0.001156 | 0.31542 | 1.319 | 2.5474 | 100.68 | 21.15 | 26.025 | 71.5177 | 84.5516 | 72.2201 | 0.006496 | 1.0222 | 0.024851 | 24.3104 | 0.070246 | 0.13716 | -0.000191 | -0.014268 | -0.014268 | -0.39704
2019-03-28 15:10 | 1553800202.7568 | 564 | 0.20956 | 0.00045 | -1e-06 | 94.2473 | 93.3658 | 125.1867 | 21.2612 | 0.002113 | 0.002041 | -151.8231 | 0.23034 | 8.4043 | 0.002577 | 0.003561 | 0.00343 | 4.7e-05 | 0.001159 | 0.34003 | 1.2985 | 2.5473 | 100.68 | 21.15 | 26.033 | 72.6598 | 83.8205 | 70.9352 | 0.006498 | 1.026 | 0.024911 | 24.279 | 0.084051 | 0.13703 | -0.000298 | -0.022299 | -0.022299 | -0.25379
2019-03-28 15:10 | 1553800203.7564 | 565 | 0.20955 | 0.000451 | -1e-06 | 94.3513 | 93.463 | 125.3179 | 21.2612 | 0.002102 | 0.00212 | -151.8347 | 0.23161 | 8.402 | 0.002597 | 0.004424 | 0.004274 | 4.7e-05 | 0.001157 | 0.30613 | 1.3246 | 2.5473 | 100.68 | 21.15 | 26.038 | 72.4087 | 84.305 | 72.8289 | 0.006499 | 1.0204 | 0.024856 | 24.3585 | 0.084608 | 0.13779 | -0.000211 | -0.015737 | -0.015737 | -0.36162
2019-03-28 15:10 | 1553800204.7559 | 566 | 0.20947 | 0.000451 | 0 | 94.4734 | 93.5204 | 125.3968 | 21.2612 | 0.002106 | 0.002012 | -151.8243 | 0.23026 | 8.4043 | 0.002572 | 0.003643 | 0.003431 | 5.1e-05 | 0.001159 | 0.29479 | 1.2772 | 2.5473 | 100.68 | 21.14 | 26.049 | 72.5031 | 83.6926 | 70.2297 | 0.006498 | 1.0281 | 0.024958 | 24.2745 | 0.090684 | 0.13698 | 0.00027 | 0.020289 | 0.020289 | 0.27884
2019-03-28 15:10 | 1553800205.7563 | 567 | 0.20942 | 0.000448 | 0 | 94.5458 | 93.8234 | 125.6864 | 21.2612 | 0.002125 | 0.002105 | -151.8474 | 0.23151 | 8.4019 | 0.002563 | 0.003611 | 0.003459 | 4.3e-05 | 0.001159 | 0.31404 | 1.3095 | 2.5475 | 100.68 | 21.14 | 26.057 | 72.9582 | 83.4694 | 72.4757 | 0.0065 | 1.0215 | 0.024905 | 24.3819 | 0.068598 | 0.13772 | 0.000556 | 0.041716 | 0.041716 | 0.13635
2019-03-28 15:10 | 1553800206.756 | 568 | 0.20955 | 0.000452 | 0 | 94.6348 | 93.8092 | 125.3843 | 21.2612 | 0.002105 | 0.002055 | -151.8332 | 0.23163 | 8.401 | 0.00254 | 0.00367 | 0.003502 | 5e-05 | 0.001161 | 0.31454 | 1.3166 | 2.5474 | 100.68 | 21.14 | 26.061 | 72.4594 | 82.9241 | 71.2618 | 0.006501 | 1.0251 | 0.024919 | 24.3097 | 0.07838 | 0.1378 | -0.000242 | -0.018178 | -0.018178 | -0.31307
2019-03-28 15:10 | 1553800207.7563 | 569 | 0.20951 | 0.000455 | -1e-06 | 94.7126 | 93.8642 | 125.5357 | 21.2612 | 0.002139 | 0.002053 | -151.8502 | 0.23186 | 8.4037 | 0.002639 | 0.00408 | 0.003867 | 4.5e-05 | 0.001165 | 0.2975 | 1.3184 | 2.5474 | 100.68 | 21.14 | 26.061 | 73.2905 | 85.3015 | 71.2152 | 0.006501 | 1.0252 | 0.024936 | 24.3228 | 0.080479 | 0.13793 | -1e-06 | -1.8e-05 | -1.8e-05 | -325.3156
2019-03-28 15:10 | 1553800208.7561 | 570 | 0.20952 | 0.00045 | 0 | 94.7835 | 93.9142 | 125.6911 | 21.2612 | 0.002143 | 0.002018 | -151.8499 | 0.23226 | 8.4041 | 0.002657 | 0.00437 | 0.004206 | 5.1e-05 | 0.001165 | 0.30807 | 1.3628 | 2.5475 | 100.68 | 21.14 | 26.073 | 73.3753 | 85.7309 | 70.3919 | 0.006504 | 1.0277 | 0.024981 | 24.3087 | 0.082413 | 0.13817 | -6.4e-05 | -0.004809 | -0.004809 | -1.1865
2019-03-28 15:10 | 1553800209.756 | 571 | 0.20952 | 0.000449 | 0 | 94.9625 | 94.2162 | 125.417 | 21.2612 | 0.002109 | 0.002034 | -151.8565 | 0.23062 | 8.4048 | 0.002546 | 0.003832 | 0.003661 | 4.8e-05 | 0.001173 | 0.32621 | 1.3309 | 2.5474 | 100.68 | 21.14 | 26.073 | 72.5755 | 83.0638 | 70.7581 | 0.006504 | 1.0266 | 0.024941 | 24.2951 | 0.070572 | 0.1372 | -4.3e-05 | -0.003197 | -0.003197 | -1.7724
2019-03-28 15:10 | 1553800210.755 | 572 | 0.20952 | 0.000451 | 0 | 94.9708 | 94.0996 | 125.5247 | 21.2612 | 0.002085 | 0.002048 | -151.8735 | 0.23204 | 8.4031 | 0.002616 | 0.003941 | 0.003728 | 4.6e-05 | 0.001173 | 0.30165 | 1.3204 | 2.5474 | 100.68 | 21.13 | 26.08 | 71.9822 | 84.7521 | 71.1083 | 0.006502 | 1.0255 | 0.024939 | 24.3179 | 0.082426 | 0.13804 | -6.7e-05 | -0.004999 | -0.004999 | -1.1403
2019-03-28 15:10 | 1553800211.7551 | 573 | 0.20953 | 0.000449 | 0 | 95.0405 | 94.3374 | 125.3965 | 21.2612 | 0.002116 | 0.002061 | -151.8462 | 0.23156 | 8.4046 | 0.002607 | 0.003793 | 0.003599 | 4.7e-05 | 0.00118 | 0.33892 | 1.309 | 2.5474 | 100.68 | 21.13 | 26.08 | 72.7384 | 84.5258 | 71.4195 | 0.006502 | 1.0246 | 0.024915 | 24.3165 | 0.066412 | 0.13776 | -8.2e-05 | -0.006131 | -0.006131 | -0.92792
2019-03-28 15:10 | 1553800212.7555 | 574 | 0.20957 | 0.000451 | 0 | 95.1614 | 94.5382 | 125.2123 | 21.2612 | 0.002148 | 0.002027 | -151.835 | 0.23176 | 8.4026 | 0.002584 | 0.003741 | 0.003586 | 4.6e-05 | 0.001182 | 0.34235 | 1.2903 | 2.5474 | 100.68 | 21.13 | 26.087 | 73.5156 | 83.971 | 70.5938 | 0.006503 | 1.0271 | 0.024926 | 24.2695 | 0.058769 | 0.13787 | -0.000351 | -0.026319 | -0.026319 | -0.21635
2019-03-28 15:10 | 1553800213.7552 | 575 | 0.20953 | 0.000449 | 0 | 95.1847 | 94.5766 | 125.09 | 21.2612 | 0.002088 | 0.002046 | -151.8345 | 0.2316 | 8.4025 | 0.002657 | 0.004178 | 0.003983 | 4.8e-05 | 0.001183 | 0.31996 | 1.3356 | 2.5475 | 100.68 | 21.13 | 26.087 | 72.0656 | 85.7253 | 71.0638 | 0.006503 | 1.0257 | 0.024897 | 24.2742 | 0.057322 | 0.13778 | -8.9e-05 | -0.006679 | -0.006679 | -0.85202
2019-03-28 15:10 | 1553800214.7565 | 576 | 0.20951 | 0.000449 | -1e-06 | 95.3634 | 94.6446 | 125.3994 | 21.2612 | 0.002066 | 0.002051 | -151.8487 | 0.23207 | 8.3996 | 0.00258 | 0.003584 | 0.003422 | 4.8e-05 | 0.001181 | 0.30973 | 1.3032 | 2.5474 | 100.68 | 21.13 | 26.09 | 71.5425 | 83.8948 | 71.1628 | 0.006504 | 1.0254 | 0.024924 | 24.3077 | 0.067676 | 0.13806 | 4.7e-05 | 0.003628 | 0.003628 | 1.5716
2019-03-28 15:10 | 1553800215.7561 | 577 | 0.20954 | 0.000452 | -1e-06 | 95.4107 | 94.7322 | 125.1564 | 21.2612 | 0.002123 | 0.001997 | -151.8347 | 0.23228 | 8.4022 | 0.002582 | 0.003962 | 0.003811 | 4.7e-05 | 0.00118 | 0.29894 | 1.3236 | 2.5475 | 100.68 | 21.13 | 26.089 | 72.9029 | 83.9244 | 69.8709 | 0.006504 | 1.0292 | 0.024947 | 24.2385 | 0.063837 | 0.13818 | -0.000162 | -0.012133 | -0.012133 | -0.47037
2019-03-28 15:10 | 1553800216.7552 | 578 | 0.20953 | 0.000447 | -1e-06 | 95.4456 | 94.8841 | 125.0104 | 21.2612 | 0.002097 | 0.002005 | -151.8438 | 0.23144 | 8.4041 | 0.002686 | 0.004052 | 0.003863 | 4.8e-05 | 0.001179 | 0.323 | 1.3224 | 2.5474 | 100.68 | 21.13 | 26.088 | 72.284 | 86.4365 | 70.0574 | 0.006503 | 1.0287 | 0.024926 | 24.231 | 0.052777 | 0.13769 | -7.7e-05 | -0.005744 | -0.005744 | -0.98999
2019-03-28 15:10 | 1553800217.7556 | 579 | 0.20953 | 0.000448 | -1e-06 | 95.5406 | 94.952 | 124.9496 | 21.2612 | 0.002074 | 0.001996 | -151.8356 | 0.2317 | 8.4021 | 0.002625 | 0.003852 | 0.00363 | 5e-05 | 0.001177 | 0.31293 | 1.3571 | 2.5473 | 100.68 | 21.12 | 26.089 | 71.7328 | 84.9729 | 69.8521 | 0.0065 | 1.0293 | 0.024927 | 24.2178 | 0.055274 | 0.13784 | -0.000122 | -0.009084 | -0.009084 | -0.62665
2019-03-28 15:10 | 1553800218.7553 | 580 | 0.20951 | 0.000451 | 0 | 95.5966 | 95.0726 | 124.7142 | 21.2612 | 0.002074 | 0.002 | -151.8563 | 0.23094 | 8.4035 | 0.002601 | 0.003981 | 0.003841 | 5e-05 | 0.001176 | 0.32637 | 1.3364 | 2.5475 | 100.68 | 21.12 | 26.089 | 71.7265 | 84.3801 | 69.9574 | 0.0065 | 1.029 | 0.0249 | 24.1987 | 0.049161 | 0.13739 | 4.8e-05 | 0.00358 | 0.00358 | 1.5847
2019-03-28 15:10 | 1553800219.7554 | 581 | 0.20952 | 0.000448 | 0 | 95.6392 | 94.8749 | 124.9522 | 21.2612 | 0.002116 | 0.002 | -151.8684 | 0.23245 | 8.4031 | 0.002597 | 0.004708 | 0.004579 | 5e-05 | 0.001176 | 0.33428 | 1.3627 | 2.5475 | 100.68 | 21.12 | 26.093 | 72.74 | 84.288 | 69.9473 | 0.006501 | 1.029 | 0.024924 | 24.2214 | 0.071766 | 0.13829 | -2e-06 | -0.000134 | -0.000134 | -42.4858
2019-03-28 15:10 | 1553800220.7557 | 582 | 0.2095 | 0.000452 | -1e-06 | 95.7684 | 95.0859 | 124.9815 | 21.2612 | 0.002123 | 0.002032 | -151.8516 | 0.2313 | 8.4027 | 0.002592 | 0.00368 | 0.003539 | 4.6e-05 | 0.001173 | 0.34235 | 1.3421 | 2.5475 | 100.68 | 21.12 | 26.097 | 72.9053 | 84.1788 | 70.7108 | 0.006502 | 1.0267 | 0.024899 | 24.2512 | 0.06398 | 0.1376 | 8.9e-05 | 0.00676 | 0.00676 | 0.84062
2019-03-28 15:10 | 1553800221.7554 | 583 | 0.20952 | 0.000448 | 0 | 95.8158 | 95.3322 | 125.1252 | 21.2612 | 0.002067 | 0.002044 | -151.8706 | 0.232 | 8.4015 | 0.002577 | 0.00377 | 0.003563 | 4.8e-05 | 0.00117 | 0.33411 | 1.3585 | 2.5474 | 100.68 | 21.12 | 26.099 | 71.5543 | 83.8205 | 70.9968 | 0.006502 | 1.0259 | 0.024903 | 24.2753 | 0.045263 | 0.13802 | -1.6e-05 | -0.001196 | -0.001196 | -4.7668
2019-03-28 15:10 | 1553800222.7548 | 584 | 0.2095 | 0.000449 | -1e-06 | 95.9844 | 95.4681 | 124.9259 | 21.2612 | 0.002099 | 0.002011 | -151.8762 | 0.23122 | 8.4012 | 0.00262 | 0.003672 | 0.003479 | 4.7e-05 | 0.001169 | 0.32211 | 1.3173 | 2.5474 | 100.68 | 21.12 | 26.104 | 72.3323 | 84.8575 | 70.2023 | 0.006503 | 1.0282 | 0.024912 | 24.2279 | 0.048246 | 0.13756 | 5.8e-05 | 0.004424 | 0.004424 | 1.2843
2019-03-28 15:10 | 1553800223.7545 | 585 | 0.20949 | 0.000452 | 0 | 96.0011 | 95.4286 | 125.313 | 21.2612 | 0.002077 | 0.00201 | -151.8575 | 0.23211 | 8.4017 | 0.002568 | 0.00359 | 0.003389 | 4.7e-05 | 0.001178 | 0.34324 | 1.3082 | 2.5473 | 100.68 | 21.12 | 26.107 | 71.7873 | 83.6084 | 70.1766 | 0.006504 | 1.0283 | 0.024951 | 24.2645 | 0.053493 | 0.13809 | 0.00014 | 0.010536 | 0.010536 | 0.54125
2019-03-28 15:10 | 1553800224.7559 | 586 | 0.20954 | 0.000445 | -1e-06 | 96.0472 | 95.4674 | 125.2616 | 21.2612 | 0.002087 | 0.002027 | -151.8505 | 0.23137 | 8.4013 | 0.002577 | 0.004076 | 0.003841 | 4.8e-05 | 0.001178 | 0.31454 | 1.3198 | 2.5474 | 100.68 | 21.12 | 26.107 | 72.043 | 83.8128 | 70.5983 | 0.006504 | 1.027 | 0.024931 | 24.2744 | 0.054164 | 0.13764 | -0.00013 | -0.00968 | -0.00968 | -0.58724
2019-03-28 15:10 | 1553800225.7546 | 587 | 0.20953 | 0.000453 | -1e-06 | 96.2158 | 95.645 | 125.236 | 21.2612 | 0.00209 | 0.00201 | -151.8428 | 0.23148 | 8.4005 | 0.002508 | 0.003745 | 0.00355 | 4.7e-05 | 0.00118 | 0.33826 | 1.3314 | 2.5474 | 100.68 | 21.12 | 26.115 | 72.1086 | 82.1426 | 70.1852 | 0.006506 | 1.0283 | 0.024943 | 24.2573 | 0.053219 | 0.13771 | -0.000128 | -0.009569 | -0.009569 | -0.59436
2019-03-28 15:10 | 1553800226.7546 | 588 | 0.20951 | 0.000448 | 0 | 96.3189 | 95.679 | 125.3381 | 21.2612 | 0.002097 | 0.002053 | -151.8608 | 0.23196 | 8.4003 | 0.002558 | 0.003578 | 0.003438 | 4.9e-05 | 0.001179 | 0.31598 | 1.3046 | 2.5474 | 100.68 | 21.12 | 26.115 | 72.2903 | 83.3526 | 71.2277 | 0.006506 | 1.0252 | 0.024916 | 24.304 | 0.059618 | 0.13799 | 9e-06 | 0.00068 | 0.00068 | 8.3771
2019-03-28 15:10 | 1553800227.7557 | 589 | 0.20955 | 0.000448 | -1e-06 | 96.3847 | 95.9759 | 124.9834 | 21.2612 | 0.002139 | 0.001974 | -151.8463 | 0.23184 | 8.402 | 0.002642 | 0.003604 | 0.003394 | 4.9e-05 | 0.001179 | 0.32836 | 1.2922 | 2.5474 | 100.68 | 21.12 | 26.124 | 73.2811 | 85.3866 | 69.3313 | 0.006508 | 1.0308 | 0.024949 | 24.2027 | 0.038014 | 0.13792 | -0.000196 | -0.014704 | -0.014704 | -0.38739
2019-03-28 15:10 | 1553800228.7547 | 590 | 0.20953 | 0.00045 | -1e-06 | 96.5164 | 96.0255 | 124.8194 | 21.2612 | 0.002124 | 0.002028 | -151.8303 | 0.23214 | 8.4018 | 0.002623 | 0.003656 | 0.003461 | 4.7e-05 | 0.001179 | 0.32853 | 1.3072 | 2.5472 | 100.68 | 21.12 | 26.124 | 72.9294 | 84.9152 | 70.6313 | 0.006508 | 1.0269 | 0.024886 | 24.2327 | 0.045602 | 0.1381 | -7.8e-05 | -0.005829 | -0.005829 | -0.97849
2019-03-28 15:10 | 1553800229.7557 | 591 | 0.20949 | 0.000452 | 0 | 96.5932 | 96.2375 | 125.005 | 21.2612 | 0.002109 | 0.002027 | -151.8186 | 0.23148 | 8.403 | 0.00255 | 0.003361 | 0.003232 | 4.4e-05 | 0.001177 | 0.32693 | 1.2849 | 2.5474 | 100.68 | 21.12 | 26.13 | 72.5678 | 83.1747 | 70.5899 | 0.00651 | 1.0271 | 0.024906 | 24.2492 | 0.032996 | 0.13771 | 0.000178 | 0.013382 | 0.013382 | 0.42499
2019-03-28 15:10 | 1553800230.7548 | 592 | 0.20951 | 0.000452 | -1e-06 | 96.7427 | 96.2348 | 125.0081 | 21.2612 | 0.002078 | 0.002013 | -151.8269 | 0.23199 | 8.4004 | 0.002585 | 0.004172 | 0.00395 | 5e-05 | 0.001177 | 0.31061 | 1.3383 | 2.5474 | 100.68 | 21.12 | 26.13 | 71.8106 | 84.0133 | 70.2578 | 0.00651 | 1.0281 | 0.024918 | 24.2378 | 0.047084 | 0.13801 | 2.8e-05 | 0.002166 | 0.002166 | 2.6312
2019-03-28 15:10 | 1553800231.7548 | 593 | 0.20949 | 0.000446 | 0 | 96.8333 | 96.2246 | 124.929 | 21.2612 | 0.002065 | 0.002066 | -151.8316 | 0.23187 | 8.4039 | 0.002592 | 0.003381 | 0.003217 | 4.9e-05 | 0.001171 | 0.32084 | 1.3435 | 2.5475 | 100.68 | 21.13 | 26.136 | 71.5209 | 84.167 | 71.5271 | 0.006515 | 1.0243 | 0.024864 | 24.2749 | 0.056399 | 0.13794 | 0.000152 | 0.011421 | 0.011421 | 0.4988
2019-03-28 15:11 | 1553800232.7548 | 594 | 0.2095 | 0.00045 | 0 | 96.9581 | 96.421 | 125.3857 | 21.2612 | 0.0021 | 0.002028 | -151.8241 | 0.23187 | 8.4032 | 0.00262 | 0.003246 | 0.003116 | 4.9e-05 | 0.00117 | 0.32587 | 1.3122 | 2.5474 | 100.68 | 21.13 | 26.143 | 72.3448 | 84.8519 | 70.6118 | 0.006517 | 1.027 | 0.024943 | 24.2869 | 0.049685 | 0.13794 | 8.7e-05 | 0.006592 | 0.006592 | 0.86421
2019-03-28 15:11 | 1553800233.7558 | 595 | 0.20951 | 0.000451 | 0 | 97.0787 | 96.6027 | 125.1735 | 21.2612 | 0.002127 | 0.002018 | -151.8275 | 0.23126 | 8.4013 | 0.002648 | 0.003848 | 0.00359 | 5.1e-05 | 0.001167 | 0.30613 | 1.3195 | 2.5474 | 100.68 | 21.13 | 26.15 | 72.9988 | 85.517 | 70.3731 | 0.006519 | 1.0277 | 0.02493 | 24.2579 | 0.043962 | 0.13758 | 4.3e-05 | 0.003228 | 0.003228 | 1.7604
2019-03-28 15:11 | 1553800234.7548 | 596 | 0.20952 | 0.000449 | -1e-06 | 97.0915 | 96.6252 | 125.2656 | 21.2612 | 0.002072 | 0.002047 | -151.844 | 0.23178 | 8.3996 | 0.002629 | 0.004226 | 0.004025 | 4.7e-05 | 0.001167 | 0.30127 | 1.3173 | 2.5474 | 100.68 | 21.13 | 26.152 | 71.679 | 85.0744 | 71.0833 | 0.006519 | 1.0256 | 0.024914 | 24.2919 | 0.043066 | 0.13789 | -8e-06 | -0.000556 | -0.000556 | -10.2372
2019-03-28 15:11 | 1553800235.7548 | 597 | 0.20951 | 0.000448 | 0 | 97.1258 | 96.812 | 125.3508 | 21.2612 | 0.002081 | 0.001954 | -151.8671 | 0.23174 | 8.405 | 0.002644 | 0.004115 | 0.003972 | 4.9e-05 | 0.001166 | 0.32549 | 1.3051 | 2.5472 | 100.68 | 21.13 | 26.156 | 71.9058 | 85.435 | 68.8532 | 0.00652 | 1.0323 | 0.025003 | 24.2213 | 0.028941 | 0.13786 | 1.3e-05 | 0.001002 | 0.001002 | 5.6826
2019-03-28 15:11 | 1553800236.7553 | 598 | 0.20952 | 0.000451 | 0 | 97.2354 | 96.7699 | 125.1679 | 21.2612 | 0.002085 | 0.002002 | -151.8554 | 0.23214 | 8.4044 | 0.002655 | 0.004212 | 0.00401 | 4.8e-05 | 0.001164 | 0.33157 | 1.3893 | 2.5475 | 100.68 | 21.13 | 26.158 | 72.001 | 85.6878 | 69.9886 | 0.006521 | 1.0289 | 0.024944 | 24.2438 | 0.042923 | 0.1381 | -9e-06 | -0.000636 | -0.000636 | -8.9726
2019-03-28 15:11 | 1553800237.7549 | 599 | 0.2095 | 0.000449 | 0 | 97.3045 | 96.9832 | 125.2135 | 21.2612 | 0.002067 | 0.002019 | -151.8865 | 0.23171 | 8.4024 | 0.002657 | 0.003662 | 0.0035 | 5e-05 | 0.001163 | 0.32676 | 1.3694 | 2.5473 | 100.68 | 21.13 | 26.158 | 71.5519 | 85.7323 | 70.3981 | 0.006521 | 1.0276 | 0.024933 | 24.2627 | 0.029579 | 0.13784 | 7.1e-05 | 0.005375 | 0.005375 | 1.0592
2019-03-28 15:11 | 1553800238.7549 | 600 | 0.20953 | 0.000448 | -1e-06 | 97.4537 | 97.1854 | 125.29 | 21.2612 | 0.002096 | 0.001997 | -151.8822 | 0.23095 | 8.4029 | 0.002663 | 0.004176 | 0.004013 | 4.7e-05 | 0.001162 | 0.31708 | 1.3431 | 2.5475 | 100.68 | 21.12 | 26.16 | 72.2653 | 85.8752 | 69.8809 | 0.006517 | 1.0292 | 0.02496 | 24.2518 | 0.024662 | 0.13739 | -8.9e-05 | -0.006681 | -0.006681 | -0.8493
2019-03-28 15:11 | 1553800239.7541 | 601 | 0.20952 | 0.000449 | -1e-06 | 97.5422 | 97.1235 | 125.2336 | 21.2612 | 0.002082 | 0.002021 | -151.8803 | 0.23174 | 8.4003 | 0.002679 | 0.004026 | 0.003853 | 4.8e-05 | 0.001158 | 0.31222 | 1.2808 | 2.5474 | 100.68 | 21.12 | 26.16 | 71.9293 | 86.2702 | 70.4549 | 0.006517 | 1.0275 | 0.024933 | 24.2666 | 0.03847 | 0.13786 | 0 | 3.5e-05 | 3.5e-05 | 161.6595
2019-03-28 15:11 | 1553800240.754 | 602 | 0.20951 | 0.000448 | 0 | 97.5404 | 97.0529 | 125.4437 | 21.2612 | 0.00205 | 0.001998 | -151.8832 | 0.23122 | 8.4021 | 0.002605 | 0.004241 | 0.004101 | 4.9e-05 | 0.00116 | 0.3194 | 1.2977 | 2.5473 | 100.68 | 21.12 | 26.157 | 71.1402 | 84.4955 | 69.8975 | 0.006517 | 1.0291 | 0.024975 | 24.2673 | 0.044816 | 0.13755 | 2.4e-05 | 0.001848 | 0.001848 | 3.0739
2019-03-28 15:11 | 1553800241.7551 | 603 | 0.2095 | 0.000451 | -1e-06 | 97.737 | 97.2852 | 125.2881 | 21.2612 | 0.002042 | 0.002005 | -151.8799 | 0.23144 | 8.4019 | 0.002627 | 0.004174 | 0.004079 | 5.1e-05 | 0.001163 | 0.33323 | 1.3452 | 2.5474 | 100.68 | 21.12 | 26.157 | 70.9492 | 85.0136 | 70.0689 | 0.006517 | 1.0286 | 0.024953 | 24.2583 | 0.041435 | 0.13768 | 8.3e-05 | 0.006291 | 0.006291 | 0.90385
2019-03-28 15:11 | 1553800242.7541 | 604 | 0.20951 | 0.000449 | -1e-06 | 97.7563 | 97.2397 | 125.02 | 21.2612 | 0.002047 | 0.001989 | -151.8926 | 0.23139 | 8.4016 | 0.002607 | 0.0046 | 0.004456 | 4.8e-05 | 0.001168 | 0.33356 | 1.4017 | 2.5474 | 100.68 | 21.12 | 26.154 | 71.0732 | 84.5383 | 69.6853 | 0.006516 | 1.0298 | 0.02494 | 24.2188 | 0.047393 | 0.13765 | 5.3e-05 | 0.004025 | 0.004025 | 1.4125
2019-03-28 15:11 | 1553800243.7552 | 605 | 0.20949 | 0.000445 | -1e-06 | 97.913 | 97.6087 | 125.3076 | 21.2612 | 0.002136 | 0.001998 | -151.9025 | 0.23234 | 8.4014 | 0.002608 | 0.004123 | 0.003885 | 4.6e-05 | 0.001168 | 0.33323 | 1.3524 | 2.5474 | 100.68 | 21.12 | 26.154 | 73.2194 | 84.547 | 69.9083 | 0.006516 | 1.0291 | 0.024961 | 24.2545 | 0.027844 | 0.13822 | 0.000132 | 0.009999 | 0.009999 | 0.57089
2019-03-28 15:11 | 1553800244.7542 | 606 | 0.2095 | 0.000447 | 0 | 97.9401 | 97.6175 | 125.4447 | 21.2612 | 0.002059 | 0.002056 | -151.9104 | 0.23148 | 8.4036 | 0.002681 | 0.004134 | 0.003997 | 5e-05 | 0.00117 | 0.31492 | 1.3473 | 2.5475 | 100.68 | 21.12 | 26.15 | 71.3688 | 86.3169 | 71.2916 | 0.006515 | 1.025 | 0.024924 | 24.3166 | 0.02951 | 0.13771 | 7.1e-05 | 0.005368 | 0.005368 | 1.0594
2019-03-28 15:11 | 1553800245.7552 | 607 | 0.20954 | 0.00045 | 0 | 98.0285 | 97.5974 | 125.5423 | 21.2612 | 0.00203 | 0.001999 | -151.9082 | 0.23034 | 8.4046 | 0.002671 | 0.003799 | 0.003657 | 4.8e-05 | 0.001169 | 0.32444 | 1.3318 | 2.5474 | 100.68 | 21.12 | 26.15 | 70.6772 | 86.0797 | 69.9247 | 0.006515 | 1.0291 | 0.024983 | 24.2778 | 0.039422 | 0.13703 | -0.000142 | -0.010629 | -0.010629 | -0.53245
2019-03-28 15:11 | 1553800246.7556 | 608 | 0.20961 | 0.000455 | -1e-06 | 98.0932 | 97.8125 | 125.6578 | 21.2612 | 0.002061 | 0.002025 | -151.8847 | 0.23126 | 8.4019 | 0.002662 | 0.003624 | 0.003478 | 4.6e-05 | 0.001166 | 0.32587 | 1.3342 | 2.5474 | 100.68 | 21.12 | 26.157 | 71.4132 | 85.8596 | 70.5431 | 0.006517 | 1.0272 | 0.024972 | 24.3108 | 0.025631 | 0.13757 | -0.000553 | -0.041528 | -0.041528 | -0.13682
2019-03-28 15:11 | 1553800247.7542 | 609 | 0.2095 | 0.000449 | 0 | 98.2031 | 97.8193 | 125.8911 | 21.2612 | 0.002039 | 0.002012 | -151.866 | 0.23175 | 8.4032 | 0.002601 | 0.00424 | 0.003974 | 5.1e-05 | 0.001169 | 0.32836 | 1.3142 | 2.5473 | 100.68 | 21.12 | 26.157 | 70.879 | 84.3895 | 70.2297 | 0.006517 | 1.0281 | 0.025007 | 24.3223 | 0.035026 | 0.13787 | 8.9e-05 | 0.006746 | 0.006746 | 0.84411
2019-03-28 15:11 | 1553800248.7543 | 610 | 0.2095 | 0.000447 | 0 | 98.2603 | 98.033 | 126.0866 | 21.2612 | 0.002104 | 0.001994 | -151.8664 | 0.23103 | 8.4025 | 0.002692 | 0.003228 | 0.003089 | 4.9e-05 | 0.001167 | 0.31365 | 1.2912 | 2.5473 | 100.68 | 21.12 | 26.178 | 72.4414 | 86.5864 | 69.8139 | 0.006522 | 1.0294 | 0.025042 | 24.3264 | 0.020715 | 0.13744 | 9e-05 | 0.006839 | 0.006839 | 0.83002
2019-03-28 15:11 | 1553800249.7553 | 611 | 0.20952 | 0.000452 | 0 | 98.3201 | 98.1624 | 125.8923 | 21.2612 | 0.002081 | 0.00203 | -151.883 | 0.23083 | 8.3996 | 0.00265 | 0.003914 | 0.003717 | 4.7e-05 | 0.001167 | 0.31061 | 1.2957 | 2.5473 | 100.68 | 21.12 | 26.178 | 71.8895 | 85.5646 | 70.6608 | 0.006522 | 1.0269 | 0.024991 | 24.3377 | 0.014361 | 0.13732 | 1e-06 | 0.000105 | 0.000105 | 54.0387
2019-03-28 15:11 | 1553800250.7543 | 612 | 0.20951 | 0.000446 | 0 | 98.3666 | 97.9511 | 125.5857 | 21.2612 | 0.002033 | 0.002022 | -151.8706 | 0.23113 | 8.4042 | 0.002671 | 0.003668 | 0.003471 | 4.8e-05 | 0.001172 | 0.33212 | 1.3076 | 2.5474 | 100.68 | 21.12 | 26.198 | 70.7442 | 86.0845 | 70.4774 | 0.006527 | 1.0274 | 0.024968 | 24.3015 | 0.037862 | 0.1375 | 6.2e-05 | 0.004709 | 0.004709 | 1.2059
2019-03-28 15:11 | 1553800251.7554 | 613 | 0.20951 | 0.000453 | 0 | 98.4256 | 98.0921 | 125.5522 | 21.2612 | 0.00206 | 0.002024 | -151.8499 | 0.23101 | 8.4011 | 0.002649 | 0.00385 | 0.003649 | 4.6e-05 | 0.001177 | 0.32139 | 1.3389 | 2.5473 | 100.69 | 21.12 | 26.219 | 71.399 | 85.5403 | 70.5313 | 0.006531 | 1.0272 | 0.024962 | 24.3002 | 0.030358 | 0.13743 | 2.6e-05 | 0.002001 | 0.002001 | 2.837
2019-03-28 15:11 | 1553800252.7554 | 614 | 0.20952 | 0.000448 | 0 | 98.4977 | 98.3734 | 125.2921 | 21.2612 | 0.002094 | 0.001994 | -151.8374 | 0.2311 | 8.4048 | 0.002605 | 0.003716 | 0.003501 | 4.9e-05 | 0.001179 | 0.31061 | 1.3481 | 2.5473 | 100.69 | 21.12 | 26.231 | 72.2114 | 84.4751 | 69.81 | 0.006534 | 1.0294 | 0.024963 | 24.2495 | 0.011293 | 0.13748 | -1.6e-05 | -0.001238 | -0.001238 | -4.5861
2019-03-28 15:11 | 1553800253.7538 | 615 | 0.2095 | 0.00045 | 0 | 98.5936 | 98.3465 | 124.9729 | 21.2612 | 0.002079 | 0.001977 | -151.8574 | 0.23081 | 8.403 | 0.002625 | 0.00374 | 0.003612 | 4.9e-05 | 0.001186 | 0.31277 | 1.3368 | 2.5474 | 100.69 | 21.12 | 26.258 | 71.8561 | 84.9715 | 69.3914 | 0.006541 | 1.0307 | 0.024946 | 24.2038 | 0.022439 | 0.13731 | 8.1e-05 | 0.006126 | 0.006126 | 0.92571
2019-03-28 15:11 | 1553800254.7535 | 616 | 0.20949 | 0.000449 | -1e-06 | 98.7124 | 98.654 | 124.8508 | 21.2612 | 0.002063 | 0.002042 | -151.8528 | 0.23202 | 8.4038 | 0.002616 | 0.003951 | 0.003796 | 4.8e-05 | 0.001191 | 0.32444 | 1.3055 | 2.5474 | 100.69 | 21.12 | 26.271 | 71.4584 | 84.7577 | 70.9586 | 0.006544 | 1.026 | 0.024877 | 24.2473 | 0.005295 | 0.13803 | 0.000124 | 0.009348 | 0.009348 | 0.60985
2019-03-28 15:11 | 1553800255.7546 | 617 | 0.20953 | 0.00045 | -1e-06 | 98.7934 | 98.5551 | 124.9487 | 21.2612 | 0.002071 | 0.002031 | -151.8376 | 0.23119 | 8.4071 | 0.00271 | 0.004028 | 0.003876 | 4.7e-05 | 0.001186 | 0.32549 | 1.2905 | 2.5473 | 100.69 | 21.12 | 26.283 | 71.6634 | 87.001 | 70.6921 | 0.006547 | 1.0268 | 0.024896 | 24.2474 | 0.021601 | 0.13753 | -7e-05 | -0.005242 | -0.005242 | -1.0836
2019-03-28 15:11 | 1553800256.7545 | 618 | 0.20952 | 0.00045 | 0 | 98.9507 | 98.8667 | 125.1883 | 21.2612 | 0.002061 | 0.002027 | -151.8586 | 0.23157 | 8.4031 | 0.002621 | 0.003952 | 0.003782 | 5e-05 | 0.001184 | 0.3194 | 1.3479 | 2.5474 | 100.69 | 21.12 | 26.283 | 71.4202 | 84.8748 | 70.6039 | 0.006547 | 1.027 | 0.024923 | 24.2675 | 0.0076 | 0.13776 | -3.1e-05 | -0.002329 | -0.002329 | -2.4433
2019-03-28 15:11 | 1553800257.7541 | 619 | 0.2095 | 0.000449 | -1e-06 | 99.0826 | 98.9676 | 125.0081 | 21.2612 | 0.00208 | 0.00198 | -151.854 | 0.23126 | 8.4049 | 0.002596 | 0.0038 | 0.003518 | 4.8e-05 | 0.001183 | 0.32781 | 1.3388 | 2.5474 | 100.69 | 21.12 | 26.302 | 71.8731 | 84.2779 | 69.4615 | 0.006552 | 1.0305 | 0.024947 | 24.2097 | 0.010395 | 0.13757 | 7.9e-05 | 0.005966 | 0.005966 | 0.95236
2019-03-28 15:11 | 1553800258.7536 | 620 | 0.20943 | 0.000448 | 0 | 99.2006 | 98.8809 | 124.678 | 21.2612 | 0.002065 | 0.002031 | -151.8714 | 0.23213 | 8.4004 | 0.002568 | 0.003488 | 0.003359 | 5.1e-05 | 0.001184 | 0.31957 | 1.3358 | 2.5474 | 100.69 | 21.12 | 26.308 | 71.5105 | 83.6014 | 70.7005 | 0.006554 | 1.0267 | 0.024869 | 24.2214 | 0.028869 | 0.13809 | 0.000495 | 0.037084 | 0.037084 | 0.15379
2019-03-28 15:11 | 1553800259.7546 | 621 | 0.20945 | 0.000449 | 0 | 99.2679 | 99.0909 | 124.863 | 21.2612 | 0.00201 | 0.002066 | -151.8493 | 0.23158 | 8.4011 | 0.002614 | 0.003814 | 0.003566 | 4.6e-05 | 0.001176 | 0.31708 | 1.322 | 2.5473 | 100.68 | 21.12 | 26.311 | 70.1991 | 84.692 | 71.5317 | 0.006555 | 1.0243 | 0.024858 | 24.2686 | 0.015965 | 0.13776 | 0.000387 | 0.028974 | 0.028974 | 0.19637
2019-03-28 15:11 | 1553800260.755 | 622 | 0.20951 | 0.00045 | 0 | 99.2962 | 99.2821 | 125.0937 | 21.2612 | 0.002122 | 0.001938 | -151.8569 | 0.23134 | 8.4023 | 0.002588 | 0.003929 | 0.00379 | 4.8e-05 | 0.001173 | 0.32493 | 1.2765 | 2.5473 | 100.68 | 21.12 | 26.311 | 72.8756 | 84.0857 | 68.4493 | 0.006555 | 1.0335 | 0.024992 | 24.1822 | 0.001267 | 0.13762 | 1.3e-05 | 0.001037 | 0.001037 | 5.4836
2019-03-28 15:11 | 1553800261.756 | 623 | 0.20949 | 0.000451 | 0 | 99.3939 | 99.3093 | 125.3932 | 21.2612 | 0.002075 | 0.001984 | -151.8799 | 0.23174 | 8.4029 | 0.002582 | 0.003711 | 0.003468 | 4.7e-05 | 0.001175 | 0.32908 | 1.3304 | 2.5474 | 100.68 | 21.13 | 26.306 | 71.753 | 83.9431 | 69.5668 | 0.006558 | 1.0301 | 0.024982 | 24.2507 | 0.007614 | 0.13786 | 0.00015 | 0.011297 | 0.011297 | 0.50401
2019-03-28 15:11 | 1553800262.7547 | 624 | 0.2095 | 0.000452 | 0 | 99.4859 | 99.2804 | 125.7693 | 21.2612 | 0.002063 | 0.00196 | -151.8923 | 0.23184 | 8.4046 | 0.00259 | 0.004173 | 0.003991 | 5.3e-05 | 0.001173 | 0.30182 | 1.3471 | 2.5475 | 100.68 | 21.13 | 26.306 | 71.4723 | 84.1288 | 68.9842 | 0.006558 | 1.0319 | 0.02504 | 24.2664 | 0.018494 | 0.13792 | 9.5e-05 | 0.007203 | 0.007203 | 0.79079
2019-03-28 15:11 | 1553800263.7547 | 625 | 0.20951 | 0.00045 | 0 | 99.6161 | 99.4466 | 126.096 | 21.2612 | 0.002072 | 0.001995 | -151.855 | 0.23153 | 8.4012 | 0.002592 | 0.003739 | 0.003522 | 5.1e-05 | 0.001174 | 0.32189 | 1.3257 | 2.5474 | 100.68 | 21.13 | 26.293 | 71.6734 | 84.1701 | 69.8367 | 0.006555 | 1.0293 | 0.025042 | 24.3281 | 0.015235 | 0.13774 | 1.5e-05 | 0.001114 | 0.001114 | 5.1048
2019-03-28 15:11 | 1553800264.7541 | 626 | 0.2095 | 0.000451 | 0 | 99.6778 | 99.452 | 125.831 | 21.2612 | 0.002036 | 0.002018 | -151.8392 | 0.23097 | 8.4037 | 0.002632 | 0.004476 | 0.004289 | 5.2e-05 | 0.00118 | 0.32604 | 1.3097 | 2.5472 | 100.68 | 21.13 | 26.293 | 70.8206 | 85.1401 | 70.3722 | 0.006555 | 1.0277 | 0.024996 | 24.3215 | 0.02028 | 0.1374 | 0.000108 | 0.008171 | 0.008171 | 0.6945
2019-03-28 15:11 | 1553800265.7547 | 627 | 0.20952 | 0.000453 | 0 | 99.7519 | 99.574 | 125.5116 | 21.2612 | 0.002035 | 0.002016 | -151.8537 | 0.23131 | 8.4032 | 0.002599 | 0.004636 | 0.004472 | 5.3e-05 | 0.001184 | 0.33467 | 1.2815 | 2.5472 | 100.68 | 21.13 | 26.268 | 70.7855 | 84.3402 | 70.3287 | 0.006548 | 1.0279 | 0.024966 | 24.2891 | 0.01597 | 0.13761 | -2.9e-05 | -0.002125 | -0.002125 | -2.674
2019-03-28 15:11 | 1553800266.7537 | 628 | 0.20952 | 0.000451 | 0 | 99.8389 | 99.7075 | 125.1623 | 21.2612 | 0.002035 | 0.002015 | -151.8604 | 0.23142 | 8.4027 | 0.002596 | 0.004419 | 0.00428 | 5.2e-05 | 0.001183 | 0.32012 | 1.2985 | 2.5472 | 100.68 | 21.12 | 26.259 | 70.7973 | 84.2738 | 70.3013 | 0.006542 | 1.0279 | 0.024932 | 24.2543 | 0.011781 | 0.13767 | -7e-06 | -0.000556 | -0.000556 | -10.2229
2019-03-28 15:11 | 1553800267.7538 | 629 | 0.20946 | 0.00045 | -1e-06 | 99.8994 | 99.6491 | 125.5789 | 21.2612 | 0.002074 | 0.002003 | -151.8676 | 0.2312 | 8.4029 | 0.002625 | 0.004121 | 0.003964 | 5.1e-05 | 0.001183 | 0.3027 | 1.3261 | 2.5472 | 100.68 | 21.12 | 26.259 | 71.7148 | 84.9621 | 70.0213 | 0.006542 | 1.0288 | 0.024983 | 24.2847 | 0.022439 | 0.13754 | 0.000358 | 0.027011 | 0.027011 | 0.2103
2019-03-28 15:11 | 1553800268.7539 | 630 | 0.20949 | 0.00045 | 0 | 99.9378 | 99.8831 | 126.123 | 21.2612 | 0.002099 | 0.001983 | -151.866 | 0.23153 | 8.4004 | 0.00262 | 0.003842 | 0.003679 | 4.9e-05 | 0.001185 | 0.32764 | 1.373 | 2.5474 | 100.68 | 21.12 | 26.233 | 72.3167 | 84.8529 | 69.5488 | 0.006536 | 1.0302 | 0.025055 | 24.3205 | 0.004893 | 0.13774 | 0.000179 | 0.013521 | 0.013521 | 0.4207
2019-03-28 15:11 | 1553800269.7531 | 631 | 0.20953 | 0.000451 | 0 | 99.9512 | 99.8101 | 126.1944 | 21.2612 | 0.00206 | 0.001987 | -151.8745 | 0.23163 | 8.4019 | 0.002612 | 0.004509 | 0.004359 | 5.3e-05 | 0.001183 | 0.32045 | 1.3382 | 2.5473 | 100.68 | 21.12 | 26.233 | 71.3844 | 84.6429 | 69.6377 | 0.006536 | 1.0299 | 0.025059 | 24.3306 | 0.012636 | 0.1378 | -7.4e-05 | -0.005586 | -0.005586 | -1.0188
2019-03-28 15:11 | 1553800270.7531 | 632 | 0.20947 | 0.000451 | 0 | 100.0462 | 99.7877 | 125.8547 | 21.2612 | 0.002045 | 0.001973 | -151.8694 | 0.23159 | 8.3995 | 0.002607 | 0.003732 | 0.003603 | 5.2e-05 | 0.001182 | 0.32189 | 1.3341 | 2.5472 | 100.68 | 21.12 | 26.211 | 71.0367 | 84.5338 | 69.3072 | 0.00653 | 1.0309 | 0.025037 | 24.2861 | 0.023144 | 0.13777 | 0.000272 | 0.020534 | 0.020534 | 0.2771
2019-03-28 15:11 | 1553800271.7543 | 633 | 0.20951 | 0.00045 | 0 | 100.1087 | 100.0132 | 125.6937 | 21.2612 | 0.002059 | 0.001965 | -151.8572 | 0.23203 | 8.4027 | 0.00263 | 0.003825 | 0.003632 | 5.4e-05 | 0.001178 | 0.31974 | 1.3713 | 2.5475 | 100.68 | 21.12 | 26.211 | 71.3695 | 85.0876 | 69.1137 | 0.00653 | 1.0315 | 0.025028 | 24.2637 | 0.008538 | 0.13803 | 2.1e-05 | 0.001604 | 0.001604 | 3.5531
2019-03-28 15:11 | 1553800272.7542 | 634 | 0.20952 | 0.000451 | 0 | 100.1016 | 100.0329 | 125.7118 | 21.2612 | 0.002094 | 0.001907 | -151.8494 | 0.23201 | 8.4027 | 0.002627 | 0.003907 | 0.00367 | 5.2e-05 | 0.001177 | 0.33124 | 1.3373 | 2.5473 | 100.68 | 21.12 | 26.199 | 72.1968 | 85.0205 | 67.7237 | 0.006527 | 1.0357 | 0.025081 | 24.2161 | 0.006137 | 0.13802 | 7e-06 | 0.00055 | 0.00055 | 10.3569
2019-03-28 15:11 | 1553800273.7531 | 635 | 0.20951 | 0.000448 | 0 | 100.1549 | 100.0312 | 125.3376 | 21.2612 | 0.002059 | 0.001998 | -151.8773 | 0.23234 | 8.4032 | 0.002669 | 0.003605 | 0.003401 | 5.6e-05 | 0.001171 | 0.33395 | 1.3416 | 2.5472 | 100.68 | 21.12 | 26.199 | 71.3743 | 86.0265 | 69.8903 | 0.006527 | 1.0292 | 0.024964 | 24.2568 | 0.011052 | 0.13822 | 6.1e-05 | 0.004586 | 0.004586 | 1.2448
2019-03-28 15:11 | 1553800274.7535 | 636 | 0.20953 | 0.000452 | 0 | 100.2154 | 99.9935 | 125.478 | 21.2612 | 0.002037 | 0.001978 | -151.8746 | 0.23228 | 8.4015 | 0.002595 | 0.004061 | 0.003873 | 5.5e-05 | 0.001162 | 0.33876 | 1.3381 | 2.5473 | 100.68 | 21.12 | 26.184 | 70.8386 | 84.2512 | 69.4303 | 0.006523 | 1.0305 | 0.024995 | 24.2541 | 0.019822 | 0.13819 | -9.7e-05 | -0.007304 | -0.007304 | -0.78138
2019-03-28 15:11 | 1553800275.7541 | 637 | 0.2095 | 0.000453 | 0 | 100.2493 | 100.3305 | 125.5235 | 21.2612 | 0.002065 | 0.001951 | -151.864 | 0.23259 | 8.3996 | 0.0026 | 0.004339 | 0.004193 | 5.3e-05 | 0.001155 | 0.31509 | 1.3078 | 2.5474 | 100.68 | 21.12 | 26.184 | 71.506 | 84.3575 | 68.7753 | 0.006523 | 1.0325 | 0.025023 | 24.2353 | -0.007243 | 0.13837 | 8.4e-05 | 0.006341 | 0.006341 | 0.9012
2019-03-28 15:11 | 1553800276.7532 | 638 | 0.2095 | 0.000452 | 0 | 100.1829 | 100.1783 | 125.4339 | 21.2612 | 0.002099 | 0.002013 | -151.8407 | 0.23177 | 8.405 | 0.002609 | 0.004423 | 0.004251 | 5.4e-05 | 0.001153 | 0.32947 | 1.327 | 2.5473 | 100.68 | 21.12 | 26.173 | 72.3191 | 84.5758 | 70.2554 | 0.006521 | 1.0281 | 0.024961 | 24.279 | 0.000408 | 0.13788 | 7.8e-05 | 0.005808 | 0.005808 | 0.98048
2019-03-28 15:11 | 1553800277.7532 | 639 | 0.20948 | 0.000453 | -1e-06 | 100.219 | 100.2028 | 125.1454 | 21.2612 | 0.002002 | 0.002018 | -151.8163 | 0.23154 | 8.4011 | 0.002558 | 0.004491 | 0.004333 | 5.3e-05 | 0.001153 | 0.33516 | 1.2888 | 2.5472 | 100.68 | 21.12 | 26.173 | 69.9965 | 83.3672 | 70.3847 | 0.006521 | 1.0277 | 0.024927 | 24.2556 | 0.001441 | 0.13774 | 0.000205 | 0.015457 | 0.015457 | 0.36802
2019-03-28 15:11 | 1553800278.7532 | 640 | 0.20958 | 0.000451 | -1e-06 | 100.205 | 100.1118 | 124.8287 | 21.2612 | 0.002067 | 0.001968 | -151.8262 | 0.23203 | 8.4021 | 0.002645 | 0.003622 | 0.003393 | 5.2e-05 | 0.001155 | 0.34075 | 1.2822 | 2.5473 | 100.68 | 21.13 | 26.162 | 71.5636 | 85.4537 | 69.1839 | 0.006522 | 1.0313 | 0.024939 | 24.1825 | 0.008323 | 0.13803 | -0.000361 | -0.027084 | -0.027084 | -0.21049
2019-03-28 15:11 | 1553800279.7543 | 641 | 0.20958 | 0.00045 | 0 | 100.1499 | 100.178 | 124.7644 | 21.2612 | 0.002085 | 0.001977 | -151.8303 | 0.23289 | 8.4032 | 0.002616 | 0.003763 | 0.003617 | 5.2e-05 | 0.001154 | 0.31421 | 1.3478 | 2.5473 | 100.68 | 21.13 | 26.162 | 71.9863 | 84.7577 | 69.4039 | 0.006522 | 1.0306 | 0.024925 | 24.1841 | -0.002514 | 0.13855 | -0.000352 | -0.026402 | -0.026402 | -0.21673
2019-03-28 15:11 | 1553800280.7543 | 642 | 0.20959 | 0.000446 | 0 | 100.2252 | 100.0975 | 124.9968 | 21.2612 | 0.002067 | 0.001993 | -151.852 | 0.23154 | 8.4019 | 0.00259 | 0.003891 | 0.003714 | 5.1e-05 | 0.001156 | 0.33052 | 1.2885 | 2.5472 | 100.68 | 21.13 | 26.159 | 71.548 | 84.1187 | 69.7718 | 0.006521 | 1.0295 | 0.024935 | 24.2196 | 0.011405 | 0.13774 | -0.000438 | -0.032851 | -0.032851 | -0.17317
2019-03-28 15:11 | 1553800281.7535 | 643 | 0.20954 | 0.000449 | 0 | 100.2136 | 100.3428 | 125.0686 | 21.2612 | 0.002046 | 0.001941 | -151.8481 | 0.23205 | 8.402 | 0.002634 | 0.003148 | 0.002983 | 5.1e-05 | 0.001157 | 0.31797 | 1.3443 | 2.5473 | 100.68 | 21.13 | 26.152 | 71.0451 | 85.193 | 68.5234 | 0.006519 | 1.0333 | 0.024987 | 24.1824 | -0.011522 | 0.13805 | -0.000115 | -0.008604 | -0.008604 | -0.66264
2019-03-28 15:11 | 1553800282.7547 | 644 | 0.20954 | 0.000454 | 0 | 100.2166 | 100.2983 | 125.0959 | 21.2612 | 0.002061 | 0.00193 | -151.8627 | 0.23173 | 8.4018 | 0.002629 | 0.003614 | 0.00339 | 4.8e-05 | 0.001161 | 0.34722 | 1.3019 | 2.5473 | 100.68 | 21.13 | 26.152 | 71.4219 | 85.073 | 68.2565 | 0.006519 | 1.0341 | 0.025 | 24.1756 | -0.007289 | 0.13786 | -0.00017 | -0.01278 | -0.01278 | -0.44549
2019-03-28 15:11 | 1553800283.7544 | 645 | 0.20954 | 0.000452 | 0 | 100.2517 | 100.2643 | 125.1548 | 21.2612 | 0.002095 | 0.001974 | -151.8575 | 0.23228 | 8.4017 | 0.00266 | 0.003033 | 0.002877 | 5.4e-05 | 0.00116 | 0.33411 | 1.3262 | 2.5474 | 100.68 | 21.13 | 26.144 | 72.2278 | 85.8025 | 69.3329 | 0.006517 | 1.0308 | 0.024966 | 24.2194 | -0.001122 | 0.13818 | -0.000169 | -0.012682 | -0.012682 | -0.45002
2019-03-28 15:11 | 1553800284.7535 | 646 | 0.20957 | 0.000449 | 0 | 100.1945 | 100.1495 | 124.836 | 21.2612 | 0.002026 | 0.001954 | -151.8662 | 0.23226 | 8.4012 | 0.002657 | 0.003506 | 0.003341 | 5.3e-05 | 0.001158 | 0.31686 | 1.3325 | 2.5474 | 100.68 | 21.13 | 26.144 | 70.5705 | 85.7434 | 68.8517 | 0.006517 | 1.0323 | 0.024952 | 24.1715 | 0.004024 | 0.13817 | -0.000318 | -0.023924 | -0.023924 | -0.23853
2019-03-28 15:11 | 1553800285.7544 | 647 | 0.20958 | 0.000446 | 0 | 100.2282 | 100.2093 | 124.8487 | 21.2612 | 0.002075 | 0.001937 | -151.8774 | 0.23149 | 8.4025 | 0.002623 | 0.003928 | 0.003627 | 4.8e-05 | 0.001164 | 0.32444 | 1.3132 | 2.5474 | 100.68 | 21.13 | 26.136 | 71.7523 | 84.9277 | 68.4414 | 0.006515 | 1.0335 | 0.024968 | 24.1582 | 0.001689 | 0.13771 | -0.000393 | -0.029561 | -0.029561 | -0.1924
2019-03-28 15:11 | 1553800286.7534 | 648 | 0.20955 | 0.000451 | 0 | 100.2627 | 100.3186 | 124.9351 | 21.2612 | 0.002052 | 0.001996 | -151.8951 | 0.23185 | 8.4011 | 0.002615 | 0.003914 | 0.00384 | 5.1e-05 | 0.001165 | 0.33411 | 1.3394 | 2.5473 | 100.68 | 21.13 | 26.136 | 71.1854 | 84.724 | 69.8413 | 0.006515 | 1.0293 | 0.024926 | 24.216 | -0.004987 | 0.13793 | -0.000177 | -0.013299 | -0.013299 | -0.42831
2019-03-28 15:11 | 1553800287.7525 | 649 | 0.20954 | 0.000449 | -1e-06 | 100.227 | 100.2687 | 124.9822 | 21.2612 | 0.002056 | 0.001954 | -151.9284 | 0.23172 | 8.4047 | 0.002657 | 0.003597 | 0.003393 | 5.2e-05 | 0.001168 | 0.32676 | 1.3159 | 2.5473 | 100.68 | 21.13 | 26.126 | 71.2837 | 85.7333 | 68.8486 | 0.006513 | 1.0323 | 0.024967 | 24.1855 | -0.003722 | 0.13785 | -0.00015 | -0.011209 | -0.011209 | -0.50794
2019-03-28 15:11 | 1553800288.7536 | 650 | 0.20954 | 0.000451 | 0 | 100.2773 | 100.3302 | 124.7079 | 21.2612 | 0.002038 | 0.001887 | -151.9154 | 0.23167 | 8.4019 | 0.002652 | 0.003462 | 0.003252 | 5.2e-05 | 0.001168 | 0.3314 | 1.3389 | 2.5475 | 100.68 | 21.13 | 26.123 | 70.8651 | 85.6176 | 67.2239 | 0.006512 | 1.0372 | 0.024999 | 24.1015 | -0.004715 | 0.13782 | -0.000151 | -0.01136 | -0.01136 | -0.50105
2019-03-28 15:11 | 1553800289.7529 | 651 | 0.20955 | 0.000447 | 0 | 100.3664 | 100.2755 | 125.0829 | 21.2612 | 0.00205 | 0.001962 | -151.9157 | 0.23258 | 8.4041 | 0.002632 | 0.003785 | 0.003532 | 5.1e-05 | 0.00117 | 0.32012 | 1.3219 | 2.5473 | 100.68 | 21.13 | 26.123 | 71.1395 | 85.1391 | 69.0411 | 0.006512 | 1.0317 | 0.02497 | 24.2021 | 0.008103 | 0.13836 | -0.000207 | -0.015566 | -0.015566 | -0.3671
2019-03-28 15:11 | 1553800290.7527 | 652 | 0.20954 | 0.000452 | 0 | 100.3137 | 100.2606 | 125.2376 | 21.2612 | 0.002031 | 0.00192 | -151.9051 | 0.232 | 8.4025 | 0.002623 | 0.003661 | 0.003547 | 5e-05 | 0.001165 | 0.33339 | 1.3336 | 2.5474 | 100.68 | 21.13 | 26.122 | 70.6858 | 84.9113 | 68.0219 | 0.006512 | 1.0348 | 0.025023 | 24.181 | 0.004737 | 0.13802 | -0.000142 | -0.01071 | -0.01071 | -0.53222
2019-03-28 15:11 | 1553800291.7536 | 653 | 0.20958 | 0.000451 | 0 | 100.2997 | 100.3407 | 125.2221 | 21.2612 | 0.002047 | 0.001963 | -151.8994 | 0.23142 | 8.4016 | 0.002645 | 0.004334 | 0.004106 | 5.3e-05 | 0.001163 | 0.31437 | 1.2968 | 2.5473 | 100.68 | 21.13 | 26.122 | 71.0739 | 85.4513 | 69.0613 | 0.006512 | 1.0317 | 0.024983 | 24.2163 | -0.003661 | 0.13767 | -0.000376 | -0.028292 | -0.028292 | -0.20097
2019-03-28 15:12 | 1553800292.7526 | 654 | 0.20953 | 0.000453 | 0 | 100.3101 | 100.3462 | 124.9804 | 21.2612 | 0.002078 | 0.001951 | -151.9009 | 0.23154 | 8.4019 | 0.002694 | 0.003666 | 0.003486 | 5e-05 | 0.001162 | 0.32908 | 1.2881 | 2.5474 | 100.68 | 21.13 | 26.121 | 71.8318 | 86.6355 | 68.7621 | 0.006512 | 1.0326 | 0.02497 | 24.1823 | -0.003216 | 0.13774 | -8.4e-05 | -0.006336 | -0.006336 | -0.89786
2019-03-28 15:12 | 1553800293.7526 | 655 | 0.20954 | 0.00045 | 0 | 100.3149 | 100.1811 | 124.8539 | 21.2612 | 0.002093 | 0.001958 | -151.8959 | 0.23225 | 8.406 | 0.002626 | 0.004313 | 0.004053 | 5e-05 | 0.001169 | 0.33196 | 1.3702 | 2.5474 | 100.68 | 21.13 | 26.124 | 72.174 | 84.9847 | 68.95 | 0.006512 | 1.032 | 0.02495 | 24.1767 | 0.011938 | 0.13817 | -0.000143 | -0.01073 | -0.01073 | -0.5318
2019-03-28 15:12 | 1553800294.7537 | 656 | 0.20954 | 0.000451 | 0 | 100.434 | 100.3268 | 125.2261 | 21.2612 | 0.001998 | 0.001869 | -151.8987 | 0.23202 | 8.403 | 0.002641 | 0.003879 | 0.003731 | 5.1e-05 | 0.00117 | 0.31404 | 1.313 | 2.5474 | 100.68 | 21.13 | 26.126 | 69.8927 | 85.3429 | 66.7994 | 0.006513 | 1.0385 | 0.025066 | 24.1365 | 0.009552 | 0.13803 | -0.00014 | -0.01059 | -0.01059 | -0.5383
2019-03-28 15:12 | 1553800295.7538 | 657 | 0.20955 | 0.000449 | 0 | 100.3881 | 100.2493 | 125.6249 | 21.2612 | 0.002046 | 0.001954 | -151.9053 | 0.23206 | 8.4039 | 0.00271 | 0.003814 | 0.003709 | 5.3e-05 | 0.001171 | 0.30702 | 1.3193 | 2.5472 | 100.68 | 21.13 | 26.127 | 71.0499 | 87.0056 | 68.8361 | 0.006513 | 1.0323 | 0.025031 | 24.2472 | 0.012374 | 0.13805 | -0.000213 | -0.016017 | -0.016017 | -0.35596
2019-03-28 15:12 | 1553800296.7538 | 658 | 0.20952 | 0.000446 | 0 | 100.3706 | 100.392 | 125.5024 | 21.2612 | 0.002018 | 0.001986 | -151.9063 | 0.23216 | 8.4044 | 0.002716 | 0.003796 | 0.003677 | 5.1e-05 | 0.001172 | 0.29623 | 1.3138 | 2.5474 | 100.68 | 21.13 | 26.135 | 70.3878 | 87.1492 | 69.6106 | 0.006515 | 1.03 | 0.024991 | 24.2628 | -0.001913 | 0.13811 | 0 | 2.9e-05 | 2.9e-05 | 195.611
2019-03-28 15:12 | 1553800297.7529 | 659 | 0.20954 | 0.000448 | 0 | 100.4027 | 100.4134 | 125.5571 | 21.2612 | 0.002031 | 0.001967 | -151.9109 | 0.23205 | 8.4017 | 0.002642 | 0.003887 | 0.003756 | 5.4e-05 | 0.001172 | 0.31188 | 1.3047 | 2.5473 | 100.68 | 21.13 | 26.139 | 70.6921 | 85.3741 | 69.1495 | 0.006516 | 1.0314 | 0.025013 | 24.2518 | -0.000954 | 0.13805 | -0.000108 | -0.008146 | -0.008146 | -0.69992
2019-03-28 15:12 | 1553800298.7528 | 660 | 0.20955 | 0.000451 | 0 | 100.4245 | 100.3628 | 125.8479 | 21.2612 | 0.002077 | 0.001987 | -151.9022 | 0.23211 | 8.4026 | 0.002654 | 0.004177 | 0.004024 | 5.5e-05 | 0.001162 | 0.31741 | 1.3198 | 2.5474 | 100.68 | 21.13 | 26.143 | 71.8054 | 85.6568 | 69.6401 | 0.006517 | 1.0299 | 0.025024 | 24.2972 | 0.005493 | 0.13808 | -0.000205 | -0.015438 | -0.015438 | -0.3694
2019-03-28 15:12 | 1553800299.7539 | 661 | 0.20957 | 0.000453 | 0 | 100.4227 | 100.2582 | 125.4869 | 21.2612 | 0.002003 | 0.001952 | -151.8867 | 0.23267 | 8.4018 | 0.00267 | 0.003542 | 0.003402 | 5.3e-05 | 0.001151 | 0.32836 | 1.303 | 2.5473 | 100.68 | 21.13 | 26.154 | 70.0206 | 86.0508 | 68.7847 | 0.00652 | 1.0325 | 0.025019 | 24.2321 | 0.014664 | 0.13841 | -0.000294 | -0.022166 | -0.022166 | -0.25789
2019-03-28 15:12 | 1553800300.753 | 662 | 0.20953 | 0.000451 | 0 | 100.4531 | 100.5147 | 125.44 | 21.2612 | 0.002042 | 0.001984 | -151.894 | 0.23237 | 8.4016 | 0.0026 | 0.003942 | 0.003663 | 5.2e-05 | 0.00115 | 0.30558 | 1.3048 | 2.5476 | 100.68 | 21.13 | 26.158 | 70.9586 | 84.3628 | 69.5536 | 0.006521 | 1.0302 | 0.024987 | 24.2548 | -0.005484 | 0.13824 | -8.8e-05 | -0.006629 | -0.006629 | -0.8612
2019-03-28 15:12 | 1553800301.7519 | 663 | 0.20956 | 0.000451 | 0 | 100.3586 | 100.4851 | 125.4573 | 21.2612 | 0.002042 | 0.00199 | -151.9219 | 0.23179 | 8.4055 | 0.002584 | 0.004582 | 0.004412 | 5.1e-05 | 0.001151 | 0.32278 | 1.3183 | 2.5473 | 100.68 | 21.13 | 26.158 | 70.9509 | 83.9859 | 69.6978 | 0.006521 | 1.0297 | 0.024983 | 24.2615 | -0.011265 | 0.13789 | -0.000233 | -0.017554 | -0.017554 | -0.32443
2019-03-28 15:12 | 1553800302.7529 | 664 | 0.20955 | 0.000455 | 0 | 100.3658 | 100.4943 | 125.5817 | 21.2612 | 0.002015 | 0.001915 | -151.9264 | 0.23204 | 8.4035 | 0.002689 | 0.004352 | 0.004096 | 5.5e-05 | 0.001153 | 0.31492 | 1.3205 | 2.5473 | 100.68 | 21.13 | 26.173 | 70.3061 | 86.4998 | 67.9086 | 0.006525 | 1.0351 | 0.025061 | 24.2101 | -0.011444 | 0.13804 | -0.000226 | -0.017054 | -0.017054 | -0.33429
2019-03-28 15:12 | 1553800303.753 | 665 | 0.20954 | 0.000451 | 0 | 100.4185 | 100.3394 | 125.8702 | 21.2612 | 0.002057 | 0.001959 | -151.9312 | 0.23148 | 8.4022 | 0.002735 | 0.003947 | 0.003816 | 5.2e-05 | 0.001157 | 0.31924 | 1.3266 | 2.5474 | 100.68 | 21.13 | 26.173 | 71.3118 | 87.6209 | 68.9695 | 0.006525 | 1.0319 | 0.025051 | 24.2756 | 0.007052 | 0.13771 | -0.000162 | -0.012175 | -0.012175 | -0.46714
2019-03-28 15:12 | 1553800304.7531 | 666 | 0.20952 | 0.00045 | 0 | 100.4111 | 100.4121 | 125.8585 | 21.2612 | 0.001996 | 0.001936 | -151.9302 | 0.23233 | 8.4051 | 0.002728 | 0.003899 | 0.003705 | 5.5e-05 | 0.001163 | 0.33539 | 1.2942 | 2.5473 | 100.68 | 21.13 | 26.184 | 69.8562 | 87.4474 | 68.4024 | 0.006527 | 1.0336 | 0.02507 | 24.2543 | -9e-05 | 0.13821 | -2.2e-05 | -0.001619 | -0.001619 | -3.5253
2019-03-28 15:12 | 1553800305.7525 | 667 | 0.20954 | 0.000453 | 0 | 100.3265 | 100.2504 | 125.4681 | 21.2612 | 0.002045 | 0.001915 | -151.9131 | 0.23198 | 8.4043 | 0.002642 | 0.003354 | 0.00317 | 5.3e-05 | 0.001165 | 0.32709 | 1.3246 | 2.5474 | 100.68 | 21.13 | 26.184 | 71.0374 | 85.3696 | 67.9103 | 0.006527 | 1.0351 | 0.02505 | 24.1992 | 0.006788 | 0.13801 | -0.000118 | -0.008903 | -0.008903 | -0.64019
2019-03-28 15:12 | 1553800306.7521 | 668 | 0.20952 | 0.00045 | 0 | 100.3539 | 100.2266 | 125.6838 | 21.2612 | 0.002028 | 0.001956 | -151.9089 | 0.23246 | 8.4036 | 0.002677 | 0.004092 | 0.003904 | 5.8e-05 | 0.00117 | 0.32333 | 1.3023 | 2.5473 | 100.68 | 21.13 | 26.195 | 70.6094 | 86.2108 | 68.9015 | 0.00653 | 1.0321 | 0.025035 | 24.2552 | 0.011353 | 0.13829 | -1.5e-05 | -0.001142 | -0.001142 | -4.999
2019-03-28 15:12 | 1553800307.7532 | 669 | 0.20952 | 0.000454 | 0 | 100.3965 | 100.3475 | 125.7808 | 21.2612 | 0.001965 | 0.001959 | -151.889 | 0.23172 | 8.4042 | 0.0026 | 0.004447 | 0.004194 | 5.4e-05 | 0.00117 | 0.335 | 1.3398 | 2.5473 | 100.68 | 21.13 | 26.195 | 69.1183 | 84.3683 | 68.9741 | 0.00653 | 1.0319 | 0.025042 | 24.2671 | 0.004362 | 0.13785 | -4.5e-05 | -0.003423 | -0.003423 | -1.6634
2019-03-28 15:12 | 1553800308.7522 | 670 | 0.20949 | 0.000454 | 0 | 100.4298 | 100.4328 | 126.0831 | 21.2612 | 0.002057 | 0.001979 | -151.8956 | 0.23157 | 8.4015 | 0.002627 | 0.003874 | 0.00377 | 5.9e-05 | 0.00117 | 0.32637 | 1.3503 | 2.5474 | 100.68 | 21.13 | 26.198 | 71.318 | 85.0143 | 69.4341 | 0.006531 | 1.0305 | 0.025055 | 24.3126 | -0.000264 | 0.13776 | 0.000177 | 0.013387 | 0.013387 | 0.425
2019-03-28 15:12 | 1553800309.7522 | 671 | 0.20954 | 0.000451 | 0 | 100.3997 | 100.3649 | 126.1636 | 21.2612 | 0.002069 | 0.001994 | -151.8943 | 0.23152 | 8.4025 | 0.002621 | 0.003953 | 0.003709 | 5.4e-05 | 0.001172 | 0.33157 | 1.3284 | 2.5474 | 100.68 | 21.12 | 26.195 | 71.6095 | 84.8717 | 69.7992 | 0.006526 | 1.0294 | 0.02505 | 24.3333 | 0.003109 | 0.13773 | -0.000143 | -0.010756 | -0.010756 | -0.52882
2019-03-28 15:12 | 1553800310.7533 | 672 | 0.20951 | 0.000454 | 0 | 100.3592 | 100.3513 | 125.9822 | 21.2612 | 0.001996 | 0.001926 | -151.8778 | 0.23156 | 8.4022 | 0.002666 | 0.003736 | 0.003611 | 5.6e-05 | 0.001165 | 0.3126 | 1.3271 | 2.5474 | 100.68 | 21.12 | 26.192 | 69.849 | 85.9625 | 68.1582 | 0.006525 | 1.0344 | 0.025092 | 24.2576 | 0.000711 | 0.13775 | 3.2e-05 | 0.002361 | 0.002361 | 2.4093
2019-03-28 15:12 | 1553800311.7523 | 673 | 0.20949 | 0.000453 | 0 | 100.356 | 100.3176 | 125.7303 | 21.2612 | 0.002105 | 0.001946 | -151.8629 | 0.23212 | 8.4021 | 0.002657 | 0.004121 | 0.003857 | 5.8e-05 | 0.001161 | 0.32139 | 1.2886 | 2.5473 | 100.68 | 21.12 | 26.19 | 72.4743 | 85.7347 | 68.6582 | 0.006525 | 1.0329 | 0.025048 | 24.2511 | 0.003417 | 0.13809 | 0.000176 | 0.013324 | 0.013324 | 0.42802
2019-03-28 15:12 | 1553800312.7537 | 674 | 0.20952 | 0.000452 | 0 | 100.2946 | 100.3655 | 125.8709 | 21.2612 | 0.001988 | 0.001988 | -151.8418 | 0.23268 | 8.4041 | 0.002619 | 0.00428 | 0.004118 | 5.8e-05 | 0.001162 | 0.33467 | 1.3178 | 2.5473 | 100.68 | 21.12 | 26.19 | 69.6642 | 84.8192 | 69.6509 | 0.006525 | 1.0299 | 0.025026 | 24.2998 | -0.006323 | 0.13842 | -4.4e-05 | -0.003297 | -0.003297 | -1.7342
2019-03-28 15:12 | 1553800313.7534 | 675 | 0.20953 | 0.00045 | 0 | 100.3229 | 100.2894 | 125.9061 | 21.2612 | 0.00205 | 0.001953 | -151.8221 | 0.23137 | 8.4015 | 0.002709 | 0.00397 | 0.003715 | 5.6e-05 | 0.00116 | 0.32117 | 1.3156 | 2.5474 | 100.68 | 21.12 | 26.178 | 71.1534 | 86.9993 | 68.8135 | 0.006522 | 1.0324 | 0.02506 | 24.2735 | 0.002985 | 0.13764 | -0.000108 | -0.008121 | -0.008121 | -0.69995
2019-03-28 15:12 | 1553800314.7524 | 676 | 0.20956 | 0.000453 | 0 | 100.3062 | 100.198 | 125.7129 | 21.2612 | 0.002091 | 0.001928 | -151.8216 | 0.23157 | 8.4002 | 0.002589 | 0.004384 | 0.004141 | 5.5e-05 | 0.001158 | 0.33732 | 1.33 | 2.5475 | 100.68 | 21.12 | 26.178 | 72.1281 | 84.1115 | 68.2099 | 0.006522 | 1.0342 | 0.025063 | 24.2335 | 0.009652 | 0.13776 | -0.000238 | -0.017981 | -0.017981 | -0.31642
2019-03-28 15:12 | 1553800315.7527 | 677 | 0.20954 | 0.00045 | 0 | 100.2499 | 100.3553 | 125.843 | 21.2612 | 0.002064 | 0.001935 | -151.8072 | 0.23204 | 8.4043 | 0.002646 | 0.003862 | 0.003779 | 5.6e-05 | 0.001159 | 0.29839 | 1.3303 | 2.5474 | 100.68 | 21.12 | 26.167 | 71.4764 | 85.4655 | 68.3854 | 0.006519 | 1.0337 | 0.02507 | 24.2522 | -0.0094 | 0.13804 | -0.000151 | -0.011355 | -0.011355 | -0.50208
2019-03-28 15:12 | 1553800316.7524 | 678 | 0.20954 | 0.000451 | 1e-06 | 100.2642 | 100.1529 | 125.7317 | 21.2612 | 0.001957 | 0.001971 | -151.8152 | 0.2315 | 8.4046 | 0.002581 | 0.004084 | 0.003914 | 5.3e-05 | 0.001161 | 0.33019 | 1.2917 | 2.5472 | 100.68 | 21.12 | 26.167 | 68.9094 | 83.9063 | 69.2462 | 0.006519 | 1.0311 | 0.025027 | 24.272 | 0.00994 | 0.13772 | -0.000126 | -0.009572 | -0.009572 | -0.59425
2019-03-28 15:12 | 1553800317.7525 | 679 | 0.20955 | 0.000448 | 0 | 100.3655 | 100.2558 | 125.7374 | 21.2612 | 0.002018 | 0.001947 | -151.8139 | 0.23163 | 8.4011 | 0.002645 | 0.003403 | 0.003239 | 5.4e-05 | 0.001162 | 0.32565 | 1.2932 | 2.5473 | 100.68 | 21.12 | 26.154 | 70.3871 | 85.4444 | 68.6652 | 0.006516 | 1.0329 | 0.025049 | 24.252 | 0.009781 | 0.1378 | -0.000176 | -0.013237 | -0.013237 | -0.42993
2019-03-28 15:12 | 1553800318.7525 | 680 | 0.20952 | 0.00045 | 0 | 100.2115 | 100.1753 | 125.4123 | 21.2612 | 0.002054 | 0.001965 | -151.8323 | 0.23159 | 8.4028 | 0.002673 | 0.003744 | 0.003621 | 5.5e-05 | 0.00116 | 0.32621 | 1.3218 | 2.5474 | 100.68 | 21.12 | 26.154 | 71.2472 | 86.1242 | 69.1068 | 0.006516 | 1.0315 | 0.025 | 24.2363 | 0.003234 | 0.13778 | 6e-06 | 0.000464 | 0.000464 | 12.2577
2019-03-28 15:12 | 1553800319.7528 | 681 | 0.20954 | 0.000449 | 0 | 100.219 | 100.1726 | 125.689 | 21.2612 | 0.001988 | 0.001964 | -151.8541 | 0.232 | 8.4077 | 0.002663 | 0.004033 | 0.003878 | 5.8e-05 | 0.001163 | 0.29358 | 1.3554 | 2.5475 | 100.68 | 21.12 | 26.144 | 69.6533 | 85.8738 | 69.0904 | 0.006513 | 1.0316 | 0.025028 | 24.2624 | 0.004141 | 0.13801 | -0.000141 | -0.010616 | -0.010616 | -0.53692
2019-03-28 15:12 | 1553800320.7516 | 682 | 0.20952 | 0.000452 | 0 | 100.2654 | 100.0625 | 126.1385 | 21.2612 | 0.001998 | 0.00193 | -151.8605 | 0.23107 | 8.4024 | 0.00258 | 0.003991 | 0.003814 | 5.6e-05 | 0.001164 | 0.32947 | 1.3275 | 2.5473 | 100.68 | 21.12 | 26.144 | 69.8983 | 83.8962 | 68.2707 | 0.006513 | 1.034 | 0.025103 | 24.2766 | 0.018121 | 0.13746 | -2.7e-05 | -0.002004 | -0.002004 | -2.8323
2019-03-28 15:12 | 1553800321.7528 | 683 | 0.20955 | 0.000448 | 0 | 100.2687 | 100.2194 | 126.2502 | 21.2612 | 0.002013 | 0.001892 | -151.8561 | 0.23175 | 8.3991 | 0.002641 | 0.003272 | 0.003082 | 5.4e-05 | 0.001168 | 0.3147 | 1.3168 | 2.5474 | 100.68 | 21.12 | 26.142 | 70.2506 | 85.3477 | 67.3525 | 0.006513 | 1.0368 | 0.025148 | 24.2547 | 0.004394 | 0.13787 | -0.000208 | -0.015726 | -0.015726 | -0.36207
2019-03-28 15:12 | 1553800322.7516 | 684 | 0.20953 | 0.000449 | 0 | 100.163 | 100.2833 | 126.0141 | 21.2612 | 0.002053 | 0.001961 | -151.8632 | 0.2315 | 8.402 | 0.002597 | 0.00416 | 0.003976 | 5.2e-05 | 0.001169 | 0.31116 | 1.3124 | 2.5474 | 100.68 | 21.12 | 26.142 | 71.2135 | 84.2911 | 69.0053 | 0.006513 | 1.0318 | 0.025064 | 24.2907 | -0.010742 | 0.13772 | -9.1e-05 | -0.006891 | -0.006891 | -0.82542
2019-03-28 15:12 | 1553800323.7516 | 685 |  | 0.00045 |  | 100.2866 | 100.2653 | 125.9796 | 21.2612 | 0.002017 | 0.001915 | -151.8637 | 0.2317 | 8.4021 | 0.002583 | 0.004167 | 0.00393 | 5.4e-05 | 0.001169 | 0.31686 | 1.337 | 2.5474 | 100.68 | 21.12 | 26.142 | 70.352 | 83.9641 | 67.9142 | 0.006513 | 1.0351 | 0.0251 | 24.2487 | 0.001896 | 0.13784
2019-03-28 15:12 | 1553800324.7516 | 686 |  | 0.00045 |  | 100.3226 | 100.2351 | 126.0894 | 21.2612 | 0.002019 | 0.001914 | -151.8628 | 0.23096 | 8.4029 | 0.002614 | 0.004471 | 0.004317 | 5.8e-05 | 0.001178 | 0.31902 | 1.2899 | 2.5474 | 100.68 | 21.12 | 26.145 | 70.4058 | 84.7023 | 67.8735 | 0.006514 | 1.0353 | 0.025113 | 24.2578 | 0.007807 | 0.1374
2019-03-28 15:12 | 1553800325.7516 | 687 |  |  |  | 100.344 | 100.3078 | 125.992 | 21.2612 | 0.002007 | 0.001933 | -151.8922 | 0.23102 | 8.4054 | 0.00259 | 0.004058 | 0.003889 | 5.5e-05 | 0.001181 | 0.31653 | 1.2874 | 2.5474 | 100.68 | 21.12 | 26.147 | 70.1102 | 84.1358 | 68.3409 | 0.006514 | 1.0338 | 0.025086 | 24.265 | 0.003234 | 0.13744
2019-03-28 15:12 | 1553800326.7516 | 688 |  |  |  | 100.3202 | 100.3085 | 126.0465 | 21.2612 | 0.002006 | 0.001949 | -151.8914 | 0.23211 | 8.4025 | 0.002664 | 0.003853 | 0.00371 | 5.7e-05 | 0.001181 | 0.28064 | 1.3055 | 2.5474 | 100.68 | 21.12 | 26.15 | 70.0823 | 85.8947 | 68.7159 | 0.006515 | 1.0327 | 0.025078 | 24.2836 | 0.001049 | 0.13808
2019-03-28 15:12 | 1553800327.7516 | 689 |  |  |  | 100.3021 | 100.3649 | 125.7059 | 21.2612 | 0.002032 | 0.00193 | -151.9017 | 0.23196 | 8.4027 | 0.002627 | 0.004134 | 0.003996 | 5.6e-05 | 0.001186 | 0.30989 | 1.2964 | 2.5472 | 100.68 | 21.12 | 26.15 | 70.7231 | 85.0104 | 68.2572 | 0.006515 | 1.0341 | 0.025061 | 24.2345 | -0.005598 | 0.13799
2019-03-28 15:12 | 1553800328.7516 | 690 |  |  |  | 100.3616 | 100.4036 |  | 21.2612 | 0.002058 | 0.001911 | -151.8786 | 0.23113 | 8.4034 | 0.002646 | 0.003991 | 0.003807 | 5.3e-05 | 0.001181 | 0.31404 | 1.342 | 2.5474 | 100.68 | 21.12 | 26.162 | 71.3399 | 85.4631 | 67.8057 | 0.006518 | 1.0355 |  |  | -0.003739 | 0.1375
2019-03-28 15:12 | 1553800329.7516 | 691 |  |  |  | 100.3089 | 100.3934 |  | 21.2612 | 0.001991 |  | -151.8729 | 0.23225 | 8.4035 | 0.002666 | 0.004329 | 0.004127 | 5.6e-05 | 0.001173 | 0.32516 | 1.3438 | 2.5474 | 100.68 | 21.12 | 26.162 | 69.743 | 85.9649 |  | 0.006518 |  |  |  | -0.00753 | 0.13816
2019-03-28 15:12 | 1553800330.7516 | 692 |  |  |  | 100.3866 | 100.283 |  | 21.2612 | 0.002047 |  |  | 0.23211 | 8.4026 | 0.002578 | 0.003896 | 0.003758 | 5.4e-05 | 0.001168 | 0.31238 | 1.3472 | 2.5473 | 100.68 | 21.12 | 26.179 | 71.0842 | 83.826 |  | 0.006522 |  |  |  | 0.009241 | 0.13808
2019-03-28 15:12 | 1553800331.7516 | 693 |  |  |  | 100.4173 | 100.4141 |  | 21.2612 |  |  |  | 0.23221 | 8.4035 | 0.002605 | 0.004561 | 0.004305 | 5.4e-05 | 0.001167 | 0.31382 | 1.3334 | 2.5473 | 100.68 | 21.12 | 26.179 |  | 84.4782 |  | 0.006522 |  |  |  | 0.000286 | 0.13814
2019-03-28 15:12 | 1553800332.7516 | 694 |  |  |  | 100.3473 | 100.3744 |  | 21.2612 |  |  |  | 0.23061 | 8.4028 | 0.002541 | 0.003715 | 0.003579 | 5.2e-05 | 0.001162 | 0.3183 | 1.3578 | 2.5475 | 100.68 | 21.12 | 26.192 |  | 82.9398 |  | 0.006525 |  |  |  | -0.00241 | 0.13719
2019-03-28 15:12 | 1553800333.7516 | 695 |  |  |  | 100.3378 | 100.3499 |  | 21.2612 |  |  |  | 0.23185 | 8.4027 | 0.002562 | 0.003815 | 0.003665 | 5.5e-05 | 0.001162 | 0.33339 | 1.3743 | 2.5473 | 100.68 | 21.12 | 26.192 |  | 83.4555 |  | 0.006525 |  |  |  | -0.001079 | 0.13793
2019-03-28 15:12 | 1553800334.7516 | 696 |  |  |  | 100.3786 | 100.3863 |  | 21.2612 |  |  |  | 0.23138 | 8.4007 | 0.002557 | 0.004209 | 0.003963 | 5.5e-05 | 0.001158 | 0.34036 | 1.3351 | 2.5475 | 100.68 | 21.11 | 26.198 |  | 83.3401 |  | 0.006523 |  |  |  | -0.000682 | 0.13765
2019-03-28 15:12 | 1553800335.7516

## Post Data
Time Stamp | Time (s) | Mass Loss Rate (g/m2s) | CO yield (kg/kg) | CO2 yield (kg/kg) | Eff HoC (MJ/kg) | Mass (g)
2019-03-28 15:01 | 0 | 0 | 0 | 0 | 0 | 44.2 | -0.005966
2019-03-28 15:01 | 1 | 0 | 0 | 0 | 0 | 44.0342 | -0.001126
2019-03-28 15:01 | 2 | 0 | 0 | 0 | 0 | 43.9871 | 0.0174
2019-03-28 15:01 | 3 | 0 | 0 | 0 | 0 | 43.9472 | 0.032124
2019-03-28 15:01 | 4 | 0 | 0 | 0 | 0 | 43.8209 | 0.28619
2019-03-28 15:01 | 5 | 0 | 0 | 0 | 0 | 43.9639 | 0.73765
2019-03-28 15:01 | 6 | 0 | 0 | 0 | 0 | 43.6684 | 1.1236
2019-03-28 15:01 | 7 | 0 | 0 | 0 | 0 | 43.5725 | 1.5214
2019-03-28 15:01 | 8 | 0 | 0 | 0 | 0 | 43.4902 | 2.0608
2019-03-28 15:01 | 9 | 0 | 0 | 0 | 0 | 43.4006 | 2.4091
2019-03-28 15:01 | 10 | 0 | 0 | 0 | 0 | 43.2696 | 2.6024
2019-03-28 15:01 | 11 | 0 | 0 | 0 | 0 | 43.1816 | 2.8595
2019-03-28 15:01 | 12 | 0 | 0 | 0 | 0 | 43.0681 | 2.9806
2019-03-28 15:01 | 13 | 0.11904 | 0 | 0 | 0 | 42.9708 | 3.0892
2019-03-28 15:01 | 14 | 0.10417 | 0 | 0 | 0 | 42.8362 | 3.1372
2019-03-28 15:01 | 15 | 0.12669 | 0 | 0 | 0 | 42.7542 | 3.1763
2019-03-28 15:01 | 16 | 0.1374 | 0 | 0 | 0 | 42.586 | 3.1636
2019-03-28 15:01 | 17 | 0.10203 | 0 | 0 | 0 | 42.4895 | 3.1679
2019-03-28 15:01 | 18 | 0.13336 | 0 | 0 | 0 | 42.3671 | 3.1478
2019-03-28 15:01 | 19 | 0.13739 | 0 | 0 | 0 | 42.2274 | 3.1309
2019-03-28 15:01 | 20 | 0.15539 | 0 | 0 | 0 | 42.0898 | 3.1134
2019-03-28 15:01 | 21 | 0.16009 | 0 | 0 | 0 | 41.9202 | 3.0989
2019-03-28 15:01 | 22 | 0.15222 | 0 | 0 | 0 | 41.7737 | 3.066
2019-03-28 15:01 | 23 | 0.124 | 0 | 0 | 0 | 41.6197 | 3.0238
2019-03-28 15:01 | 24 | 0.15452 | 0 | 0 | 0 | 41.5126 | 2.9713
2019-03-28 15:01 | 25 | 0.14599 | 0 | 0 | 0 | 41.3192 | 2.9235
2019-03-28 15:01 | 26 | 0.12274 | 0 | 0 | 0 | 41.224 | 2.8786
2019-03-28 15:01 | 27 | 0.15019 | 0 | 0 | 0 | 41.0631 | 2.8422
2019-03-28 15:01 | 28 | 0.11888 | 0 | 0 | 0 | 40.9365 | 2.7997
2019-03-28 15:01 | 29 | 0.10994 | 0 | 0 | 0 | 40.821 | 2.7699
2019-03-28 15:01 | 30 | 0.10779 | 0 | 0 | 0 | 40.7142 | 2.7475
2019-03-28 15:01 | 31 | 0.11101 | 0 | 0 | 0 | 40.6039 | 2.7638
2019-03-28 15:01 | 32 | 0.11048 | 0 | 0 | 0 | 40.4924 | 2.7807
2019-03-28 15:01 | 33 | 0.1281 | 0 | 0 | 0 | 40.3794 | 2.7835
2019-03-28 15:01 | 34 | 0.10977 | 0 | 0 | 0 | 40.2442 | 2.7793
2019-03-28 15:01 | 35 | 0.11299 | 0 | 0 | 0 | 40.1553 | 2.7734
2019-03-28 15:01 | 36 | 0.11833 | 0 | 0 | 0 | 40.0168 | 2.7409
2019-03-28 15:01 | 37 | 0.13327 | 0 | 0 | 0 | 39.9163 | 2.7487
2019-03-28 15:01 | 38 | 0.14566 | 0.020188 | 1.6352 | 18.8825 | 39.7521 | 2.7504
2019-03-28 15:01 | 39 | 0.10899 | 0.026771 | 2.1803 | 25.3098 | 39.6365 | 2.7585
2019-03-28 15:01 | 40 | 0.10759 | 0.02692 | 2.2049 | 25.5003 | 39.5264 | 2.7436
2019-03-28 15:01 | 41 | 0.11181 | 0.025373 | 2.12 | 24.2041 | 39.419 | 2.7063
2019-03-28 15:01 | 42 | 0.11326 | 0.024391 | 2.0829 | 23.7989 | 39.3036 | 2.6954
2019-03-28 15:01 | 43 | 0.1036 | 0.025883 | 2.277 | 25.9331 | 39.195 | 2.6867
2019-03-28 15:01 | 44 | 0.10506 | 0.024898 | 2.2384 | 24.9279 | 39.0936 | 2.6188
2019-03-28 15:01 | 45 | 0.11323 | 0.022254 | 2.06 | 22.5589 | 38.9829 | 2.5543
2019-03-28 15:01 | 46 | 0.12668 | 0.019357 | 1.8278 | 19.8268 | 38.8671 | 2.5118
2019-03-28 15:01 | 47 | 0.088099 | 0.027331 | 2.6047 | 28.5381 | 38.7414 | 2.5142
2019-03-28 15:01 | 48 | 0.087882 | 0.026859 | 2.5871 | 28.5403 | 38.6815 | 2.5082
2019-03-28 15:01 | 49 | 0.12876 | 0.017948 | 1.7463 | 19.3946 | 38.5555 | 2.4973
2019-03-28 15:01 | 50 | 0.12338 | 0.01872 | 1.8111 | 20.1899 | 38.4346 | 2.4909
2019-03-28 15:01 | 51 | 0.1103 | 0.020698 | 2.0113 | 22.5188 | 38.3112 | 2.4838
2019-03-28 15:01 | 52 | 0.1227 | 0.018595 | 1.8009 | 20.2649 | 38.208 | 2.4865
2019-03-28 15:01 | 53 | 0.13161 | 0.017292 | 1.6674 | 19.1273 | 38.0666 | 2.5173
2019-03-28 15:02 | 54 | 0.11841 | 0.019019 | 1.8464 | 21.498 | 37.9501 | 2.5457
2019-03-28 15:02 | 55 | 0.11877 | 0.018929 | 1.8351 | 21.976 | 37.827 | 2.6102
2019-03-28 15:02 | 56 | 0.11344 | 0.019779 | 1.9186 | 23.23 | 37.7124 | 2.6352
2019-03-28 15:02 | 57 | 0.15055 | 0.014888 | 1.4507 | 17.6235 | 37.5907 | 2.6532
2019-03-28 15:02 | 58 | 0.14969 | 0.014977 | 1.4664 | 18.4706 | 37.4202 | 2.7649
2019-03-28 15:02 | 59 | 0.13278 | 0.017022 | 1.661 | 21.151 | 37.296 | 2.8085
2019-03-28 15:02 | 60 | 0.13086 | 0.017258 | 1.6984 | 21.8546 | 37.1517 | 2.8599
2019-03-28 15:02 | 61 | 0.12346 | 0.018136 | 1.8204 | 23.4249 | 37.0353 | 2.8919
2019-03-28 15:02 | 62 | 0.11415 | 0.019555 | 1.9817 | 25.9292 | 36.9049 | 2.9597
2019-03-28 15:02 | 63 | 0.11845 | 0.018892 | 1.9377 | 25.5277 | 36.8033 | 3.0237
2019-03-28 15:02 | 64 | 0.13695 | 0.016242 | 1.6889 | 22.0604 | 36.6651 | 3.0212
2019-03-28 15:02 | 65 | 0.11437 | 0.019457 | 2.0521 | 26.3547 | 36.5384 | 3.0142
2019-03-28 15:02 | 66 | 0.12535 | 0.017756 | 1.8981 | 24.0415 | 36.4295 | 3.0136
2019-03-28 15:02 | 67 | 0.12027 | 0.018472 | 1.9967 | 25.0895 | 36.2907 | 3.0174
2019-03-28 15:02 | 68 | 0.12025 | 0.018593 | 2.0105 | 24.8937 | 36.1879 | 2.9934
2019-03-28 15:02 | 69 | 0.12757 | 0.017489 | 1.9058 | 23.3095 | 36.0486 | 2.9736
2019-03-28 15:02 | 70 | 0.11992 | 0.018427 | 2.0421 | 24.934 | 35.9356 | 2.9902
2019-03-28 15:02 | 71 | 0.13414 | 0.016426 | 1.8331 | 22.9444 | 35.8038 | 3.0778
2019-03-28 15:02 | 72 | 0.14179 | 0.015498 | 1.7427 | 21.9513 | 35.6685 | 3.1124
2019-03-28 15:02 | 73 | 0.14702 | 0.014806 | 1.6953 | 21.1514 | 35.5213 | 3.1097
2019-03-28 15:02 | 74 | 0.14465 | 0.015074 | 1.7403 | 21.5889 | 35.3766 | 3.1229
2019-03-28 15:02 | 75 | 0.14333 | 0.015043 | 1.7667 | 21.7013 | 35.2326 | 3.1104
2019-03-28 15:02 | 76 | 0.1164 | 0.018337 | 2.188 | 26.6385 | 35.0952 | 3.1007
2019-03-28 15:02 | 77 | 0.13059 | 0.016173 | 1.9597 | 23.4444 | 34.9905 | 3.0616
2019-03-28 15:02 | 78 | 0.13183 | 0.015902 | 1.9541 | 22.5068 | 34.8364 | 2.9671
2019-03-28 15:02 | 79 | 0.12186 | 0.016993 | 2.1104 | 24.1414 | 34.729 | 2.942
2019-03-28 15:02 | 80 | 0.14629 | 0.014028 | 1.7576 | 20.1301 | 34.5854 | 2.9448
2019-03-28 15:02 | 81 | 0.13361 | 0.015345 | 1.9161 | 21.8916 | 34.4449 | 2.925
2019-03-28 15:02 | 82 | 0.12151 | 0.016552 | 2.1033 | 24.2272 | 34.3191 | 2.944
2019-03-28 15:02 | 83 | 0.10796 | 0.018596 | 2.355 | 26.8989 | 34.2019 | 2.9041
2019-03-28 15:02 | 84 | 0.10644 | 0.018856 | 2.3783 | 27.2449 | 34.0996 | 2.9
2019-03-28 15:02 | 85 | 0.12291 | 0.016342 | 2.0555 | 23.271 | 33.9843 | 2.8603
2019-03-28 15:02 | 86 | 0.134 | 0.015168 | 1.8807 | 21.1929 | 33.8554 | 2.8398
2019-03-28 15:02 | 87 | 0.11433 | 0.018154 | 2.1949 | 24.8811 | 33.7234 | 2.8447
2019-03-28 15:02 | 88 | 0.12602 | 0.016639 | 1.9809 | 22.5719 | 33.6199 | 2.8446
2019-03-28 15:02 | 89 | 0.13449 | 0.015894 | 1.8438 | 21.2099 | 33.4722 | 2.8525
2019-03-28 15:02 | 90 | 0.10742 | 0.019869 | 2.2922 | 26.6456 | 33.3588 | 2.8622
2019-03-28 15:02 | 91 | 0.11798 | 0.018273 | 2.0826 | 24.272 | 33.2486 | 2.8635
2019-03-28 15:02 | 92 | 0.13855 | 0.01579 | 1.7627 | 20.5797 | 33.12 | 2.8513
2019-03-28 15:02 | 93 | 0.1299 | 0.016982 | 1.858 | 21.5137 | 32.978 | 2.7946
2019-03-28 15:02 | 94 | 0.13947 | 0.016049 | 1.7069 | 19.7365 | 32.8561 | 2.7527
2019-03-28 15:02 | 95 | 0.16233 | 0.014275 | 1.4388 | 16.6824 | 32.6957 | 2.708
2019-03-28 15:02 | 96 | 0.16369 | 0.015078 | 1.3987 | 16.8142 | 32.5352 | 2.7524
2019-03-28 15:02 | 97 | 0.20579 | 0.012942 | 1.0869 | 13.9367 | 32.3592 | 2.868
2019-03-28 15:02 | 98 | 0.2237 | 0.012662 | 0.98331 | 15.558 | 32.1289 | 3.4803
2019-03-28 15:02 | 99 | 0.23453 | 0.012719 | 0.94361 | 18.6098 | 31.9141 | 4.3646
2019-03-28 15:02 | 100 | 0.2505 | 0.012638 | 0.99128 | 21.4266 | 31.6584 | 5.3674
2019-03-28 15:02 | 101 | 0.27845 | 0.012391 | 1.0504 | 25.1065 | 31.4098 | 6.9909
2019-03-28 15:02 | 102 | 0.30949 | 0.01374 | 1.1362 | 25.9857 | 31.1005 | 8.0422
2019-03-28 15:02 | 103 | 0.33351 | 0.016804 | 1.2681 | 25.6517 | 30.7928 | 8.5551
2019-03-28 15:02 | 104 | 0.34929 | 0.020804 | 1.4385 | 27.0617 | 30.436 | 9.4524
2019-03-28 15:02 | 105 | 0.35987 | 0.025016 | 1.627 | 26.9459 | 30.0954 | 9.697
2019-03-28 15:02 | 106 | 0.38998 | 0.027277 | 1.7068 | 25.1822 | 29.713 | 9.8205
2019-03-28 15:02 | 107 | 0.36608 | 0.032601 | 1.9944 | 27.1412 | 29.3274 | 9.9358
2019-03-28 15:02 | 108 | 0.38391 | 0.033439 | 1.991 | 26.1256 | 28.9722 | 10.0299
2019-03-28 15:02 | 109 | 0.38485 | 0.034801 | 2.0254 | 25.9541 | 28.5621 | 9.9885
2019-03-28 15:02 | 110 | 0.40796 | 0.033955 | 1.9359 | 24.3054 | 28.1973 | 9.9156
2019-03-28 15:02 | 111 | 0.43126 | 0.032755 | 1.8411 | 22.6286 | 27.7465 | 9.7589
2019-03-28 15:02 | 112 | 0.41829 | 0.034089 | 1.9014 | 23.1556 | 27.3431 | 9.6857
2019-03-28 15:02 | 113 | 0.4376 | 0.032355 | 1.8122 | 22.0079 | 26.9033 | 9.6306
2019-03-28 15:02 | 114 | 0.44056 | 0.031409 | 1.7823 | 21.8455 | 26.4712 | 9.6242
2019-03-28 15:03 | 115 | 0.44346 | 0.030175 | 1.7492 | 21.7051 | 26.0225 | 9.6254
2019-03-28 15:03 | 116 | 0.44863 | 0.028683 | 1.7114 | 22.003 | 25.5835 | 9.8712
2019-03-28 15:03 | 117 | 0.46705 | 0.026256 | 1.6269 | 21.659 | 25.1232 | 10.1157
2019-03-28 15:03 | 118 | 0.43483 | 0.02677 | 1.7469 | 24.3857 | 24.6602 | 10.6036
2019-03-28 15:03 | 119 | 0.46075 | 0.023439 | 1.6544 | 23.2686 | 24.2407 | 10.7211
2019-03-28 15:03 | 120 | 0.47381 | 0.021007 | 1.6545 | 23.4264 | 23.7412 | 11.0997
2019-03-28 15:03 | 121 | 0.45817 | 0.02014 | 1.7758 | 25.1657 | 23.3001 | 11.5301
2019-03-28 15:03 | 122 | 0.46363 | 0.018641 | 1.8144 | 25.2789 | 22.8209 | 11.7202
2019-03-28 15:03 | 123 | 0.4512 | 0.017664 | 1.9256 | 26.527 | 22.3765 | 11.9689
2019-03-28 15:03 | 124 | 0.44979 | 0.016716 | 1.9712 | 26.8607 | 21.9164 | 12.0817
2019-03-28 15:03 | 125 | 0.44044 | 0.016156 | 2.0678 | 27.4362 | 21.4783 | 12.084
2019-03-28 15:03 | 126 | 0.44597 | 0.014973 | 2.0718 | 27.004 | 21.0325 | 12.0431
2019-03-28 15:03 | 127 | 0.4288 | 0.014725 | 2.1909 | 28.1037 | 20.5909 | 12.0509
2019-03-28 15:03 | 128 | 0.43945 | 0.01369 | 2.1716 | 27.2744 | 20.1693 | 11.9858
2019-03-28 15:03 | 129 | 0.42195 | 0.01367 | 2.2936 | 28.0851 | 19.7182 | 11.8505
2019-03-28 15:03 | 130 | 0.41191 | 0.013448 | 2.3778 | 28.0647 | 19.324 | 11.5602
2019-03-28 15:03 | 131 | 0.41145 | 0.012943 | 2.4007 | 27.5322 | 18.8921 | 11.3281
2019-03-28 15:03 | 132 | 0.39784 | 0.0128 | 2.4846 | 27.9116 | 18.5039 | 11.1042
2019-03-28 15:03 | 133 | 0.39421 | 0.012459 | 2.5098 | 27.6895 | 18.0944 | 10.9154
2019-03-28 15:03 | 134 | 0.38698 | 0.012198 | 2.5451 | 27.3961 | 17.7163 | 10.6017
2019-03-28 15:03 | 135 | 0.37448 | 0.012073 | 2.606 | 27.2615 | 17.322 | 10.2089
2019-03-28 15:03 | 136 | 0.35654 | 0.012106 | 2.707 | 27.9959 | 16.9682 | 9.9816
2019-03-28 15:03 | 137 | 0.36522 | 0.011241 | 2.5768 | 26.8783 | 16.6029 | 9.8164
2019-03-28 15:03 | 138 | 0.34894 | 0.011114 | 2.6281 | 27.4007 | 16.2425 | 9.5612
2019-03-28 15:03 | 139 | 0.36352 | 0.010154 | 2.4687 | 26.0117 | 15.8991 | 9.4559
2019-03-28 15:03 | 140 | 0.34034 | 0.01032 | 2.5811 | 27.1654 | 15.5246 | 9.2455
2019-03-28 15:03 | 141 | 0.3027 | 0.011117 | 2.8256 | 29.72 | 15.222 | 8.9961
2019-03-28 15:03 | 142 | 0.30257 | 0.010661 | 2.7594 | 29.3326 | 14.9097 | 8.8751
2019-03-28 15:03 | 143 | 0.31638 | 0.0097974 | 2.5731 | 27.4408 | 14.6128 | 8.6818
2019-03-28 15:03 | 144 | 0.30667 | 0.0098091 | 2.596 | 27.403 | 14.2822 | 8.4036
2019-03-28 15:03 | 145 | 0.30418 | 0.0096516 | 2.547 | 26.4046 | 13.9982 | 8.0318
2019-03-28 15:03 | 146 | 0.30939 | 0.0094032 | 2.4467 | 25.5248 | 13.6725 | 7.897
2019-03-28 15:03 | 147 | 0.28753 | 0.010025 | 2.5906 | 27.156 | 13.3856 | 7.8082
2019-03-28 15:03 | 148 | 0.2833 | 0.010171 | 2.5849 | 26.991 | 13.0945 | 7.6464
2019-03-28 15:03 | 149 | 0.25134 | 0.011413 | 2.8685 | 30.0477 | 12.8241 | 7.552
2019-03-28 15:03 | 150 | 0.27195 | 0.010508 | 2.6115 | 27.7152 | 12.5796 | 7.537
2019-03-28 15:03 | 151 | 0.27943 | 0.010147 | 2.5086 | 27.0152 | 12.2827 | 7.5489
2019-03-28 15:03 | 152 | 0.26372 | 0.010606 | 2.612 | 28.7452 | 12.0269 | 7.5808
2019-03-28 15:03 | 153 | 0.24652 | 0.011241 | 2.763 | 30.3533 | 11.7561 | 7.4827
2019-03-28 15:03 | 154 | 0.24047 | 0.011468 | 2.7993 | 30.6182 | 11.531 | 7.3628
2019-03-28 15:03 | 155 | 0.24044 | 0.0115 | 2.7722 | 30.3688 | 11.2735 | 7.302
2019-03-28 15:03 | 156 | 0.22972 | 0.012242 | 2.8761 | 30.7469 | 11.052 | 7.0633
2019-03-28 15:03 | 157 | 0.23714 | 0.012051 | 2.7562 | 28.9943 | 10.8095 | 6.8758
2019-03-28 15:03 | 158 | 0.25191 | 0.011643 | 2.5595 | 26.4856 | 10.5755 | 6.6721
2019-03-28 15:03 | 159 | 0.26473 | 0.011418 | 2.3968 | 24.681 | 10.3074 | 6.5338
2019-03-28 15:03 | 160 | 0.22245 | 0.01401 | 2.8093 | 28.7879 | 10.0583 | 6.4038
2019-03-28 15:03 | 161 | 0.2447 | 0.012979 | 2.5138 | 25.6934 | 9.8483 | 6.2873
2019-03-28 15:03 | 162 | 0.25222 | 0.012811 | 2.397 | 24.5524 | 9.5719 | 6.1925
2019-03-28 15:03 | 163 | 0.22199 | 0.014742 | 2.6547 | 27.6921 | 9.3524 | 6.1473
2019-03-28 15:03 | 164 | 0.23571 | 0.014028 | 2.4604 | 25.7474 | 9.1184 | 6.069
2019-03-28 15:03 | 165 | 0.23499 | 0.014252 | 2.4279 | 25.3877 | 8.884 | 5.9658
2019-03-28 15:03 | 166 | 0.21737 | 0.015419 | 2.5767 | 26.9451 | 8.6527 | 5.8571
2019-03-28 15:03 | 167 | 0.21009 | 0.016164 | 2.6245 | 27.7544 | 8.4473 | 5.831
2019-03-28 15:03 | 168 | 0.19799 | 0.01749 | 2.7458 | 29.5 | 8.2333 | 5.8407
2019-03-28 15:03 | 169 | 0.18776 | 0.018657 | 2.8539 | 30.6692 | 8.0499 | 5.7584
2019-03-28 15:03 | 170 | 0.21482 | 0.016664 | 2.4632 | 26.2806 | 7.8495 | 5.6456
2019-03-28 15:03 | 171 | 0.20152 | 0.018107 | 2.6003 | 27.5364 | 7.6295 | 5.5491
2019-03-28 15:03 | 172 | 0.18281 | 0.020279 | 2.83 | 29.8261 | 7.448 | 5.4526
2019-03-28 15:03 | 173 | 0.19519 | 0.019162 | 2.6221 | 27.4109 | 7.2565 | 5.3502
2019-03-28 15:04 | 174 | 0.19999 | 0.018801 | 2.5215 | 26.3968 | 7.059 | 5.2791
2019-03-28 15:04 | 175 | 0.19183 | 0.019733 | 2.5917 | 27.2359 | 6.8594 | 5.2245
2019-03-28 15:04 | 176 | 0.2005 | 0.019278 | 2.4447 | 26.0707 | 6.6717 | 5.2272
2019-03-28 15:04 | 177 | 0.20333 | 0.019409 | 2.3788 | 25.6427 | 6.4599 | 5.2138
2019-03-28 15:04 | 178 | 0.19279 | 0.020939 | 2.4716 | 26.5972 | 6.2685 | 5.1276
2019-03-28 15:04 | 179 | 0.18548 | 0.022434 | 2.5427 | 27.1459 | 6.0741 | 5.035
2019-03-28 15:04 | 180 | 0.17304 | 0.024798 | 2.6901 | 28.3374 | 5.8984 | 4.9035
2019-03-28 15:04 | 181 | 0.17271 | 0.025545 | 2.6651 | 27.8489 | 5.725 | 4.8097
2019-03-28 15:04 | 182 | 0.1731 | 0.026082 | 2.6285 | 27.1729 | 5.5527 | 4.7038
2019-03-28 15:04 | 183 | 0.1667 | 0.027681 | 2.694 | 28.1202 | 5.3808 | 4.6878
2019-03-28 15:04 | 184 | 0.14734 | 0.031906 | 3.0085 | 31.2732 | 5.2219 | 4.6077
2019-03-28 15:04 | 185 | 0.15948 | 0.029998 | 2.7404 | 28.6239 | 5.079 | 4.565
2019-03-28 15:04 | 186 | 0.15404 | 0.031646 | 2.7988 | 29.7464 | 4.9066 | 4.5823
2019-03-28 15:04 | 187 | 0.14501 | 0.034098 | 2.933 | 30.8451 | 4.7722 | 4.4728
2019-03-28 15:04 | 188 | 0.13913 | 0.036278 | 3.0133 | 31.4535 | 4.6158 | 4.376
2019-03-28 15:04 | 189 | 0.13676 | 0.037557 | 3.0251 | 31.3603 | 4.4928 | 4.289
2019-03-28 15:04 | 190 | 0.13707 | 0.038059 | 2.9745 | 30.8629 | 4.3412 | 4.2305
2019-03-28 15:04 | 191 | 0.14425 | 0.037093 | 2.7905 | 28.5991 | 4.2167 | 4.1255
2019-03-28 15:04 | 192 | 0.15856 | 0.034774 | 2.4999 | 25.5068 | 4.0512 | 4.0443
2019-03-28 15:04 | 193 | 0.15327 | 0.036992 | 2.5414 | 25.8893 | 3.904 | 3.968
2019-03-28 15:04 | 194 | 0.15988 | 0.03647 | 2.3961 | 24.1992 | 3.743 | 3.8691
2019-03-28 15:04 | 195 | 0.14165 | 0 | 0 | 0 | 3.5899 | 3.7499
2019-03-28 15:04 | 196 | 0.13804 | 0 | 0 | 0 | 3.4566 | 3.6531
2019-03-28 15:04 | 197 | 0.1346 | 0 | 0 | 0 | 3.313 | 3.5833
2019-03-28 15:04 | 198 | 0.14566 | 0 | 0 | 0 | 3.1844 | 3.4774
2019-03-28 15:04 | 199 | 0.13183 | 0 | 0 | 0 | 3.0271 | 3.4022
2019-03-28 15:04 | 200 | 0.13566 | 0 | 0 | 0 | 2.9176 | 3.3353
2019-03-28 15:04 | 201 | 0.12028 | 0 | 0 | 0 | 2.7604 | 3.249
2019-03-28 15:04 | 202 | 0.098345 | 0 | 0 | 0 | 2.6783 | 3.128
2019-03-28 15:04 | 203 | 0.10912 | 0 | 0 | 0 | 2.5563 | 3.0576
2019-03-28 15:04 | 204 | 0.092258 | 0 | 0 | 0 | 2.4654 | 3.049
2019-03-28 15:04 | 205 | 0.099987 | 0 | 0 | 0 | 2.3668 | 2.9673
2019-03-28 15:04 | 206 | 0.088898 | 0 | 0 | 0 | 2.269 | 2.8658
2019-03-28 15:04 | 207 | 0.096738 | 0 | 0 | 0 | 2.1851 | 2.7971
2019-03-28 15:04 | 208 | 0.088774 | 0 | 0 | 0 | 2.0786 | 2.6334
2019-03-28 15:04 | 209 | 0.10046 | 0 | 0 | 0 | 2.0038 | 2.524
2019-03-28 15:04 | 210 | 0.085027 | 0 | 0 | 0 | 1.8839 | 2.4515
2019-03-28 15:04 | 211 | 0.06889 | 0 | 0 | 0 | 1.8334 | 2.4361
2019-03-28 15:04 | 212 | 0.093624 | 0 | 0 | 0 | 1.7362 | 2.3867
2019-03-28 15:04 | 213 | 0.10861 | 0 | 0 | 0 | 1.6485 | 2.3254
2019-03-28 15:04 | 214 | 0.088699 | 0 | 0 | 0 | 1.5282 | 2.231
2019-03-28 15:04 | 215 | 0.0629 | 0 | 0 | 0 | 1.4725 | 2.1612
2019-03-28 15:04 | 216 | 0.07954 | 0 | 0 | 0 | 1.3926 | 2.0786
2019-03-28 15:04 | 217 | 0.0672 | 0 | 0 | 0 | 1.3191 | 1.9971
2019-03-28 15:04 | 218 | 0.07217 | 0 | 0 | 0 | 1.255 | 1.9597
2019-03-28 15:04 | 219 | 0.058061 | 0 | 0 | 0 | 1.1781 | 1.8778
2019-03-28 15:04 | 220 | 0.078661 | 0 | 0 | 0 | 1.131 | 1.831
2019-03-28 15:04 | 221 | 0.087518 | 0 | 0 | 0 | 1.0237 | 1.7851
2019-03-28 15:04 | 222 | 0.066802 | 0 | 0 | 0 | 0.96366 | 1.7117
2019-03-28 15:04 | 223 | 0.048488 | 0 | 0 | 0 | 0.88956 | 1.6376
2019-03-28 15:04 | 224 | 0.05965 | 0 | 0 | 0 | 0.85928 | 1.6091
2019-03-28 15:04 | 225 | 0.067952 | 0 | 0 | 0 | 0.77051 | 1.5988
2019-03-28 15:04 | 226 | 0.060474 | 0 | 0 | 0 | 0.72708 | 1.5562
2019-03-28 15:04 | 227 | 0.065836 | 0 | 0 | 0 | 0.64742 | 1.5291
2019-03-28 15:04 | 228 | 0.048758 | 0 | 0 | 0 | 0.60023 | 1.4763
2019-03-28 15:04 | 229 | 0.053904 | 0 | 0 | 0 | 0.54576 | 1.438
2019-03-28 15:04 | 230 | 0.028497 | 0 | 0 | 0 | 0.49888 | 1.3674
2019-03-28 15:04 | 231 | 0.026799 | 0 | 0 | 0 | 0.48343 | 1.3194
2019-03-28 15:04 | 232 | 0.029428 | 0 | 0 | 0 | 0.44354 | 1.3076
2019-03-28 15:04 | 233 | 0.031552 | 0 | 0 | 0 | 0.42466 | 1.2714
2019-03-28 15:05 | 234 | 0.026182 | 0 | 0 | 0 | 0.38188 | 1.2626
2019-03-28 15:05 | 235 | 0.036518 | 0 | 0 | 0 | 0.36877 | 1.2349
2019-03-28 15:05 | 236 | 0.039904 | 0 | 0 | 0 | 0.3106 | 1.1887
2019-03-28 15:05 | 237 | 0.028267 | 0 | 0 | 0 | 0.29262 | 1.1563
2019-03-28 15:05 | 238 | 0.026838 | 0 | 0 | 0 | 0.25152 | 1.1131
2019-03-28 15:05 | 239 | 0.037888 | 0 | 0 | 0 | 0.23538 | 1.0675
2019-03-28 15:05 | 240 | 0.053366 | 0 | 0 | 0 | 0.17478 | 1.0502
2019-03-28 15:05 | 241 | 0.048111 | 0 | 0 | 0 | 0.13332 | 1.0401
2019-03-28 15:05 | 242 | 0.059131 | 0 | 0 | 0 | 0.075396 | 1.0298
2019-03-28 15:05 | 243 | 0.057311 | 0 | 0 | 0 | 0.017664 | 1.0022
2019-03-28 15:05 | 244 | 0.067227 | 0 | 0 | 0 | -0.040891 | 1.0073
2019-03-28 15:05 | 245 | 0.044762 | 0 | 0 | 0 | -0.10924 | 0.99649
2019-03-28 15:05 | 246 | 0.035346 | 0 | 0 | 0 | -0.13312 | 0.97861
2019-03-28 15:05 | 247 | 0.035413 | 0 | 0 | 0 | -0.18305 | 0.95161
2019-03-28 15:05 | 248 | 0.050817 | 0 | 0 | 0 | -0.20724 | 0.92049
2019-03-28 15:05 | 249 | 0.033437 | 0 | 0 | 0 | -0.27724 | 0.92488
2019-03-28 15:05 | 250 | 0.026885 | 0 | 0 | 0 | -0.27682 | 0.89729
2019-03-28 15:05 | 251 | 0.055845 | 0 | 0 | 0 | -0.33842 | 0.84183
2019-03-28 15:05 | 252 | 0.022006 | 0 | 0 | 0 | -0.37407 | 0.81229
2019-03-28 15:05 | 253 | 0.0042935 | 0 | 0 | 0 | -0.38513 | 0.77984
2019-03-28 15:05 | 254 | -0.0030811 | 0 | 0 | 0 | -0.38644 | 0.76127
2019-03-28 15:05 | 255 | 0.016716 | 0 | 0 | 0 | -0.3858 | 0.73156
2019-03-28 15:05 | 256 | 0.019986 | 0 | 0 | 0 | -0.41641 | 0.68232
2019-03-28 15:05 | 257 | 0.020751 | 0 | 0 | 0 | -0.42437 | 0.68017
2019-03-28 15:05 | 258 | 0.0067821 | 0 | 0 | 0 | -0.45511 | 0.62982
2019-03-28 15:05 | 259 | 0.030092 | 0 | 0 | 0 | -0.44637 | 0.6089
2019-03-28 15:05 | 260 | 0.037856 | 0 | 0 | 0 | -0.51106 | 0.58021
2019-03-28 15:05 | 261 | -0.0053027 | 0 | 0 | 0 | -0.51085 | 0.56051
2019-03-28 15:05 | 262 | 0.022641 | 0 | 0 | 0 | -0.51666 | 0.53663
2019-03-28 15:05 | 263 | 0.040767 | 0 | 0 | 0 | -0.55478 | 0.51648
2019-03-28 15:05 | 264 | 0.030869 | 0 | 0 | 0 | -0.59082 | 0.50207
2019-03-28 15:05 | 265 | 0.01582 | 0 | 0 | 0 | -0.61491 | 0.47868
2019-03-28 15:05 | 266 | 0.02052 | 0 | 0 | 0 | -0.62725 | 0.47456
2019-03-28 15:05 | 267 | 0.02471 | 0 | 0 | 0 | -0.65643 | 0.44839
2019-03-28 15:05 | 268 | 0.031147 | 0 | 0 | 0 | -0.6767 | 0.42919
2019-03-28 15:05 | 269 | 0.018833 | 0 | 0 | 0 | -0.714 | 0.42279
2019-03-28 15:05 | 270 | 0.0056254 | 0 | 0 | 0 | -0.71412 | 0.39406
2019-03-28 15:05 | 271 | 0.010606 | 0 | 0 | 0 | -0.72974 | 0.37272
2019-03-28 15:05 | 272 | 0.014513 | 0 | 0 | 0 | -0.73509 | 0.3652
2019-03-28 15:05 | 273 | -0.0020012 | 0 | 0 | 0 | -0.75448 | 0.35451
2019-03-28 15:05 | 274 | 0.0090328 | 0 | 0 | 0 | -0.73792 | 0.3314
2019-03-28 15:05 | 275 | 0.037366 | 0 | 0 | 0 | -0.77645 | 0.33025
2019-03-28 15:05 | 276 | 0.022026 | 0 | 0 | 0 | -0.80242 | 0.31214
2019-03-28 15:05 | 277 | 0.019063 | 0 | 0 | 0 | -0.82207 | 0.29542
2019-03-28 15:05 | 278 | 0.0075488 | 0 | 0 | 0 | -0.83855 | 0.29349
2019-03-28 15:05 | 279 | -0.0066429 | 0 | 0 | 0 | -0.83673 | 0.28278
2019-03-28 15:05 | 280 | -0.0070125 | 0 | 0 | 0 | -0.82914 | 0.27106
2019-03-28 15:05 | 281 | 0.0092323 | 0 | 0 | 0 | -0.8265 | 0.27496
2019-03-28 15:05 | 282 | -0.0076027 | 0 | 0 | 0 | -0.84081 | 0.26479
2019-03-28 15:05 | 283 | 0.012727 | 0 | 0 | 0 | -0.81933 | 0.25686
2019-03-28 15:05 | 284 | 0.019898 | 0 | 0 | 0 | -0.86303 | 0.23382
2019-03-28 15:05 | 285 | -0.0047006 | 0 | 0 | 0 | -0.85149 | 0.22201
2019-03-28 15:05 | 286 | -0.0044432 | 0 | 0 | 0 | -0.85933 | 0.18743

## Channels
Index | Channel Name/Description | Legend/ID | Group Name | Source Type | Range/Type | Definition | Offset | Slope | LowCal (Zero) | HighCal (Span) | Calibration TimeStamp | Cal Group | Range Low | Range High | Out of Range | Background | Delay | Scaled Units | SCALED Global Name;RAW Global Name | Notes | Dip Switch Gnd_Ref Pull_Up (Up/Down) | Smooth Interval (s) | Pdf Plot # | Significant Digits
0 | Oxygen Volume Fraction (Dry) | O2 | Oxygen | DAQ Voltage (V) | ±10V default | Dev1/ai33 | -0.062482 | 0.062688 | 0 | 0.2095 | 2019-03-28 11:01 | Group A | 0.1 | 0.22 | Do Nothing | 0.20949 | 12 | Vol fr | O2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
1 | Carbon Dioxide Volume Fraction (Dry) | CO2 | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai32 | -0.025283 | 0.02541 | 0 | 0.0261 | 2019-03-28 11:01 | Group A | -0.01 | 0.1 | Do Nothing | 0.00044025 | 10 | Vol fr | CO2 | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
2 | Carbon Monoxide Volume Fraction (Dry) | CO | CO2_CO | DAQ Voltage (V) | ±10V default | Dev1/ai23 | -0.002743 | 0.002756 | 0 | 0.0011 | 2019-03-28 11:01 | Group A | -0.01 | 0.1 | Do Nothing | -7.8267e-07 | 12 | Vol fr | CO | Servomex 4100 O2 CO2 CO analyzer | NA | 1 |  | 6
3 | Smoke Detector 1 -compensation | Io | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai21 | -18.6589 | 183.6167 | 0 | 100 | 2019-03-28 11:01 | None | 0 | 1000 | Do Nothing | 100.9596 | 0 | % | Io |  | NA | 1 |  | 6
4 | Smoke Detector 2 - main | I | Smoke | DAQ Voltage (V) | ±10V default | Dev1/ai22 | -11.412 | 209.4304 | 0 | 100 | 2019-03-28 11:01 | None | 0 | 1000 | Do Nothing | 101.0938 | 0 | % | I |  | NA | 1 |  | 6
5 | Duct Flow Oriface Plate Differential Pressure | Pe | Duct Pressure | DAQ Voltage (V) | ±10V default | Dev1/ai20 | -16.6878 | 1447 | 0 |  | 2019-03-28 10:45 | None | -10 | 1000 | Do Nothing | 126.5887 | 2 | Pa | Pe | MKS 223 pressure transducer | NA | 10 |  | 6
6 | CJC Connector 0 | CJC0 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai0 | 0 | 100 |  |  |  | None | 10 | 50 | Do Nothing | 21.2612 | 0 | °C | CJC0 |  | NA | 1 |  | 6
7 | TC1 Duct Inlet - near smokemeter | TCraw1 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai4 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0018833 | 4 | V | TCraw1 |  | NA | 1 |  | 6
8 | TC2 Duct Outlet - near oriface plate | TCraw2 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai5 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0017688 | 6 | V | TCraw2 |  | NA | 1 |  | 6
9 | Load Cell for Sample Mass | SampMass | Sample Mass | DAQ Voltage (V) | ±10V default | Dev1/ai16 | -89.824 | 118.16 | 0 | 200 |  | Group B | -10 | 600 | Do Nothing | -151.6736 | 0 | g | SampMass | ATC mass load cell | NA | 10 | 3 | 6
10 | Burner Methane Mass Flow Rate | BurnerFlow | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai18 | 0.23647 | 4.0102 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 0.23006 | 0 | SLPM | BurnerFlow | MKS M100B-29174 mass flow controller (output) | NA | 1 |  | 6
11 | Burner Mass Flow Rate setpoint | BurnerFlowSet | Methane Burner Flow | DAQ Voltage (V) | ±10V default | Dev1/ai19 | 0.201 | 4.0069 | 7 | 10 |  | Group C | -1 | 10 | Do Nothing | 8.4026 | 0 | SLPM | BurnerFlowSet | 10 turn precision potentiometer for setpoint input to MKS M100B | NA | 1 |  | 6
12 | TC1 Cone Heater | TCraw3 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai1 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0025582 | 0 | V | TCraw3 |  | NA | 1 |  | 6
13 | TC2 Cone Heater | TCraw4 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai2 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0038345 | 0 | V | TCraw4 |  | NA | 1 |  | 6
14 | TC3 Cone Heater | TCraw5 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai3 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0036574 | 0 | V | TCraw5 |  | NA | 1 |  | 6
15 | Spare heat Flux Gauge | Hfspare | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai6 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 5.4525e-05 | 0 | V | Hfspare |  | NA | 1 |  | 6
16 | TC spare | TCraw7 | MISC | DAQ Voltage (V) | ±0.2V default | Dev1/ai7 | 0 | 1 |  |  |  | None | -0.1 | 0.1 | Do Nothing | 0.0012911 | 0 | V | TCraw7 |  | NA | 1 |  | 6
17 | Radiometer Heat Flux | HeatFlux | Cone Radiometer Heat Flux | DAQ Voltage (V) | ±0.2V default | Dev1/ai17 | -0.0457 | 5530.2 | 0 |  |  | None | -10 | 200 | Do Nothing | 0.32309 | 0 | kW/m^2 | HeatFlux | Medtherm GTW-7-32-485A serial #78461 | NA | 1 |  | 6
18 | Canopy Hood Flow Alarm | CanopyFlow | CanopyFlow | DAQ Voltage (V) | ±10V default | Dev1/ai34 | -2.5201 | 1 | 0 |  |  | None | -1 | 5 | Do Nothing | 1.328 | 0 | V | CanopyFlow | Setra 2651001WB2BT1C pressure transducer serial #3698689, 1" WC | NA | 1 |  | 6
19 | mark signal | MarkSig | marker | DAQ Voltage (V) | ±10V default | Dev1/ai35 | 0 | 1 |  |  |  | None | -10 | 10 | Do Nothing | 2.5474 | 0 | V | MarkSig | push buttons for marking events | NA | 1 |  | 4
20 | Ambient Pressure | Pamb | vaisala | Global Variable | N/A | 129.6.159.102(Pressure) | 0 | 1 |  |  |  | None | 80 | 110 | Do Nothing | 100.7 | 0 | kPa | Pamb | globdata(r,Pamb) | NA | 1 |  | 6
21 | Ambient Temperature | Tamb | vaisala | Global Variable | N/A | 129.6.159.102(Temperature) | 0 | 1 |  |  |  | None | 0 | 50 | Do Nothing | 21.1682 | 0 | °C | Tamb | globdata(r,Tamb) | NA | 1 |  | 6
22 | Ambient Relative Humidity | RHamb | vaisala | Global Variable | N/A | 129.6.159.102(Humidity) | 0 | 1 |  |  |  | None | 0 | 100 | Do Nothing | 25.8187 | 0 | % | RHamb | globdata(r,RHamb) | NA | 1 |  | 6
23 | Gas Temperature near smoke meter | Tsmk | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw1[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 67.1415 | 4 | °C | Tsmk | Ktemp_its90(scaledata(r,TCraw1),scaledata(r,CJC0)) | NA | 1 |  | 6
24 | Cone Heater Temperature | Tcone | Cone Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw3[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 83.3617 | 0 | °C | Tcone | Ktemp_its90(scaledata(r,TCraw3),scaledata(r,CJC0)) | NA | 1 |  | 6
25 | Exhaust Gas Temperature | Te | Duct Temperature | Math Function | N/A | TC K-type(CJC0[i];TCraw2[i]) | [5 empty] | None | 0 | 2000 | Do Nothing | 64.3879 | 6 | °C | Te | Ktemp_its90(scaledata(r,TCraw2),scaledata(r,CJC0)) | NA | 1 |  | 6
26 | Ambient Water Volume Fraction | H2Oa | MISC | Excel Function | N/A | H2Oa=(0.00000659628*Tamb[i]^4 - 0.000102573*Tamb[i]^3 + 0.0192184*Tamb[i]^2 + 0.234477*Tamb[i]+4.94319)*RHamb[i]/(Pamb[i]*760/101.325)/100 | [5 empty] | None | 0 | 0.05 | Do Nothing | 0.0064502 | 0 | Vol fr | H2Oa | (0.00000659628*scaledata(r,Tamb)^4 - 0.000102573*scaledata(r,Tamb)^3 + 0.0192184*scaledata(r,Tamb)^2 + 0.234477*scaledata(r,Tamb) + 4.94319)*scaledata(r,RHamb)/(scaledata(r,Pamb)*760/101.325)/100 | NA | 1 |  | 6
27 | Exhaust Gas Density | Rhoe | MISC | Excel Function | N/A | Rhoe=760*101325/760*28.97/8314.47/(Te[i]+273.15) | [5 empty] | None | 0 | 2 | Do Nothing | 1.0459 | 6 | kg/m3 | Rhoe | 760*101325/760*28.97/8314.47/(scaledata(r,Te)+273.15) | NA | 1 |  | 6
28 | Duct Mass Flow Rate | MFRe | Duct Mass Flow Rate | Excel Function | N/A | MFRe=Cf[p]*ABS((Pe[i])/(Te[i]+273.15))^0.5 | [5 empty] | None | -10 | 100 | Set to "-1" | 0.02529 | 6 | kg/s | MFRe | Cf(i)*abs((scaledata(r,Pe))/(scaledata(r,Te)+273.15))^0.5 | NA | 1 |  | 6
29 | Duct Volume Flow Rate | VFRe | Duct Volume Flow Rate | Excel Function | N/A | VFRe=MFRe[i]/Rhoe[i]*1000 | [5 empty] | None | -10 | 100 | Set to "-1" | 24.1789 | 6 | L/s | VFRe | scaledata(r,MFRe)/scaledata(r,Rhoe)*1000 | NA | 1 | 4 | 6
30 | Smoke Extinction Coefficient | Ksm | Smoke | Excel Function | N/A | Ksm=1/D[p]*LN(Io[i]/I[i]) | [5 empty] | None | -100 | 1000 | Set to "NaN" |  | 0 | 1/m | Ksm | 1/D(i)*log(scaledata(r,Io)/scaledata(r,I)) | NA | 1 |  | 6
31 | Methane Burner HRR | HRRmethane | Heat Release Rate | Excel Function | N/A | HRRmethane=(BurnerFlow[i]*101325/273.15/8314*16/60*HOCcal[p]) | [5 empty] | None | -1 | 15 | Do Nothing |  | 0 | kW | HRRmethane | scaledata(r,BurnerFlow)*101325/273.15/8314*16/60*50 | NA | 1 | 2 | 6
32 | Oxygen Depletion Factor | Phi | Phi | Excel Function | N/A | Phi=(O2[0]*(1-CO2[i]-CO[i])-(O2[i]*(1-CO2[0])))/(O2[0]*(1-CO2[i]-CO[i]-O2[i])) | [5 empty] | None | -0.1 | 1 | Do Nothing |  | 12 | none | Phi | (XO2o*(1-scaledata(r,CO2)-scaledata(r,CO))-(scaledata(r,O2)*(1-XCO2o)))/(XO2o*(1-scaledata(r,CO2)-scaledata(r,CO)-scaledata(r,O2))) | NA | 1 |  | 6
33 | Rate of Heat Release | RHR | Heat Release Rate | Excel Function | N/A | RHR=(Ef[p]*Phi[i]-(17690-Ef[p])*(1-Phi[i])*CO[i]/O2[i]/2)*((MFRe[i]*(1-H2Oa[i])*O2[0]*32/28.97/(1+Phi[i]*(Alpha[p]-1)))) | [5 empty] | None | -10 | 50 | Set to "-1" |  | 12 | kW | RHR | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1)))) | NA | 1 | 2 | 6
34 | Rate of Heat Release per Unit Area | RHRA | Heat Release Rate per Area | Excel Function | N/A | RHRA=(RHR[i]/As[p]) | [5 empty] | None | -1000 | 5000 | Do Nothing |  | 0 | kW/m2 | RHRA | (Ef(i)*scaledata(r,Phi)-(17690-Ef(i))*(1-scaledata(r,Phi))*scaledata(r,CO)/scaledata(r,O2)/2)*((scaledata(r,MFRe)*(1-scaledata(r,H2Oa))*XO2o*32/28.97/(1+scaledata(r,Phi)*(Alpha(i)-1))))/As(i) | NA | 1 | 1 | 6
35 | New Calibration Factor | newCf | Cfactor | Excel Function | N/A | newCf=HRRmethane[i]*Cf[p]/RHR[i] | [5 empty] | None | 0.02 | 0.06 | Set to "NaN" |  | 0 | none | newCf | scaledata(r,HRRmethane)*Cf(i)/scaledata(r,RHR) | NA | 1 |  | 6
36 | EventDriver | EvMark | marker | Excel Function | N/A | EvMark=IF(MarkSig[i]<=2.5,"Event:Ignition",IF(MarkSig[i]>=2.6,"Event:Start Test",0)) | [5 empty] | None | -5 | 5 | Do Nothing |  | 0 | none | EvMark | NaN