## Sheet1
Detection of gradually increasing gamma radiation levels and Neutron detection
N42.32 Data Sheet Section 6.5
Manufacturer:
Model: |  |  | Serial Number:
| Requirement: | The instrument’s alarm threshold shall not be affected by slowly increasing radiation levels that may be caused when a wearer is slowly approaching or is being approached by a radiation source. The alarm shall activate within 2 seconds after the instrument reaches the test position.
Note: | Comments are required when the requirement is not verified.
| Ambient Conditions: |  | °C |  | %RH |  | mmHg
Test Equipment:
Source Data:
| Gamma Background Field: |  | μR/h |  | Neutron Background Field: |  | (add units)
| Gamma Alarm Threshold: |  | μR/h |  | Neutron Alarm Threshold: |  | (add units)
Instrument Mode of operation
Record in table if instrument alarmed or not within 2 seconds. It is expected a Yes/No (y/n) entry for 10 readings
|  |  | Cs-137 |  | Cf-252
|  |  | Time to alarm ≤ 2s |  | Time to alarm ≤ 2s
|  |  | Yes | No | Yes | No
1
2
3
4
5
6
7
8
9
10
|  |  | Yes | No
Did the instrument alarm within 2 seconds for gamma?
Did the instrument alarm within 2 seconds for neutron?
For Cs-137:
Record maximum time to alarm: |  |  |  |  | seconds
Record minimum time to alarm: |  |  |  |  | seconds
For Cf-252:
Record maximum time to alarm: |  |  |  |  | seconds
Record minimum time to alarm: |  |  |  |  | seconds
Comments:
Performed by: |  |  | Date:
Reviewed by: |  |  | Date: