## Sheet1
Audible Alarm
Data Sheet Section 5.5
Manufacturer:
Instrument:
Model: |  |  |  |  | Serial Number:
Date Performed: |  |  |  |  | Test Location:
| Requirements: | The instrument shall provide an audible alarm to indicate an increase in the radiation level that is greater than the alarm set point. Different techniques may be used to differentiate between the type of radiation detected (e.g., gamma, neutron, over-range). The frequency of an audible alarm signal shall be from 1000 Hz to at least 4000 Hz. Where an intermittent alarm signal is provided, the interval shall not exceed 2 s. The A-weighted alarm signal volume at a distance of 30 cm from the instrument shall be at least 80 dB(A) and shall not exceed 100 dB(A). It shall not be possible to disable both the vibration and audible alarm indications simultaneously, except through the restricted mode. When both alarm signals are off, an indication shall be provided on the display to inform the user of this condition. An earphone connection should be available to enable use of the audible function in a high-noise environment.
|  | Ambient Conditions: |  | °C | %RH |  | mmHg
Test Equipment:
Instrument Mode of operation
Note: | Comments are required when the requirement is not verified.
Measurement Results
|  | Alarm volume |  |  |  | Frequency
|  |  | dB(A) |  |  |  | Hz
|  | 1 |  |  |  | 1
|  | 2 |  |  |  | 2
|  | 3 |  |  |  | 3
|  | 4 |  |  |  | 4
|  | 5 |  |  |  | 5
|  | 6 |  |  |  | 6
|  | 7 |  |  |  | 7
|  | 8 |  |  |  | 8
|  | 9 |  |  |  | 9
|  | 10 |  |  |  | 10
|  | Mean | #DIV/0! | dB(A) |  | Mean | #DIV/0! | Hz
|  | STD | #DIV/0! | dB(A) |  | STD | #DIV/0! | Hz
|  | COV % | #DIV/0! |  |  | COV % | #DIV/0!
Verify
[7 empty] | Yes | No
Is the alarm frequency within 1000 to 4000 Hz?
Where an intermittent alarm is provided, is the interval less than 2 seconds?
Is the alarm volume at a distance of 30 cm within 85 dB(A) and 100 dB(A)?
If the audible alarm can be disabled, does the instrument have a vibration or visual alarm?
Is an earphone connection available?
Alarms are distinguishable for different types of radiation (gamma, neutron, over-range)?
If yes, describe differences:
Does the instrument have preventive measures for disabling all the alarms?
If yes, describe:
Record intermittent alarm interval (seconds):
Comments:
| Performed by: |  |  |  |  | Date:
| Reviewed by: |  |  |  |  | Date: