## Section 5.8
Sections 5.8 User Interface
Test Data and Report
Manufacturer:
| Model: | Serial Number:
Requirement: | NOTE: Portions of this section may not be required based on the ability of the BRD to store data. a. A BRD shall include a display that is readable over the required temperature range of -20 ºC to 50 ºC (tested in temperature test section). b. The display shall be readable under low light levels (<150 lux – overcast day or work area) and high light levels (>10000 lux – full daylight). c. The detector response (gamma-ray exposure and/or count rate and, when provided, neutron) and battery status shall be provided at the user interface and updated at a minimum rate of once per second. d. The identification results, confidence indicator, and ability to access spectra shall be provided at the user interface when identification capabilities are available. e. A directional pointer shall be provided at the user interface when directionality capabilities are available. The pointer should be referenced to the detector position or a compass direction. f. The following shall be automatically displayed at the user interface: • Alarm indication (visual and audible, with alarm type (gamma and/or neutron) and level (user selectable)) • Changes in state of health (e.g., communication status, detector status, GPS status, energy stabilization, battery status) • Over-range indication for gamma and if provided, neutron radiation • Background changes that can affect the overall response capability to radiation changes g. The following functions and controls shall be provided at the user interface: • Ability to reset alarms • Ability to mute audible alarms • Ability to go silent (mute radio emissions, acoustical emissions, and perceptible vibrations) • Ability to display GPS and detector response data on same display • Ability to modify date and time • Ability to reset (reacquire) background h. The following information and control shall be provided through the use of access controls (e.g., password). • Access to and control of operating parameters such as background update modes (e.g., fixed or rolling), alarm thresholds and the radionuclide libraries if the instrument has identification capabilities. • Access to alarm history, if data storage is provided • If data storage is provided, access to count rate time-history profiles including gamma-ray and when available, neutron indication • Access to background radiation information (most recently stored or currently used for operation) • Access to alarm selection criteria • Access to the radionuclide library (when radionuclide identification capabilities are provided) i. When data storage is provided, it is desirable to allow controlled access of the data storage capabilities to enable or disable record storage. If access is provided and data storage is disabled, an indication shall be provided on the user interface indicating that data storage has been disabled.
Note: | Comments are required when the requirement is not verified.
|  |  |  | Yes | No
Test Results - Low Light Levels (<150 lux)
Turn on the instrument and verify that it is working properly (e.g., the battery is charged, the detector is present and working, memory is available, self-check passed)?
Can perform a background reading?
Can calibrate (if necessary)?
Can perform real-time data collection?
Can make an identification (of a single radionuclide) and save the data? (if applicable)
Test Results - High Light Levels (>10000 lux)
Turn on the instrument and verify that it is working properly (e.g., the battery is charged, the detector is present and working, memory is available, self-check passed)?
Can perform a background reading?
Can calibrate (if necessary)?
Can perform real-time data collection?
Can make an identification (of a single radionuclide) and save the data? (if applicable)
Does the display update at a minimum rate of 1 times per second?
If the BRD has identification capability, are all requirements met?
If the BRD has directional capability, are all requirements met?
Are the following automatically displayed at the user interface:
Alarm
Changes in state of health
Over-range
Background changes
Are the following functions and controls provided at the user interface:
Reset alarms
Mute audible alarms
Go silent
Display GPS and detector response (when available)
Modify date and time
Reacquire background
Are the following available through access control:
Control of operating parameters
Alarm history
Gamma count rate profiles
Neutron (when available) count rate profiles
Background radiation information
Alarm selection criteria
Radionuclide library, when applicable
Enable/disable record storage (if storage provided)
Enable/disable indicator (if the capability exists)
Comments:
| Completed by: |  | Date:
| Reviewed by: |  | Date:
Draft 3