## Section 6.12.3
Section 6.12.3 - Single Radionuclide Identification
Test Data and Report
Manufacturer:
Model: |  |  |  | Serial Number:
Requirement: | Except for 252Cf, a BRD with identification capabilities shall be able to identify the radionuclides listed in Table 3 at an exposure rate of 10 µR/h (±20%) within an integration time of 60 s or that stated by the manufacturer. NOTE – The low exposure rate required for test is difficult to measure with a radiation instrument. Therefore, equations are used to calculate the distance required based on source activity. Resources are available from, for example, the Health Physics Society-Relationship Between Radionuclide Gamma Emission and Exposure Rate that provide guidance as to how to determine the distance. The medical radionuclides (99mTc, 67Ga, 131I, and 201Tl) shall be surrounded by 8 cm of polymethyl methacrylate (PMMA) for test purposes as a means to represent in-vivo conditions. The 5 µR/h (±20%) above background exposure rate is measured with the source in the PMMA. NOTE - The actual activities for medical radionuclides given to patients for different procedures are much greater than those required in this standard. For 99mTc stress tests, normal doses are typically between 3 and 5 mCi; bone scans with 99mTc are 20 and 25 mCi; diagnosis using 131I between 10 and 30 µCi; therapy with 131I between 29 and 300 mCi; 201Tl stress tests approximately 4 mCi; and for 67Ga between 2 and 5 mCi.
Note: | Comments are required when the requirement is not verified.
| Ambient Conditions: |  | °C |  | %RH | mmHg
Source Data:
Radionuclide Identification Results
| Ba-133 | Co-57 | Co-60 | Cs-137
Source ID
Date
1
2
3
4
5
6
7
8
9
10
# Correct
| Ga-67 | I-131 | Tc-99m | Tl-201
Source ID
Date
1
2
3
4
5
6
7
8
9
10
# Correct
| Ra-226 | Th-232 | DU | HEU | WGPu
Source ID
Date
1
2
3
4
5
6
7
8
9
10
# Correct
Comments:
| Performed by: |  |  | Date:
| Reviewed by: |  |  | Date:
Draft 3