## Section 5.5
Section 5.5 Power Supply
Test Data and Report
Manufacturer:
| Model: |  | Serial Number:
Requirement: | a. A BRD shall have the ability to support a continuous operating time of 8 h under standard operating conditions without replacing batteries. b. The manufacturer shall state the expected continuous operating time using the recommended batteries and the conditions (functional and environmental) used to determine this time. c. BRDs shall be equipped with a visible direct indicator of battery life. d. Batteries shall be replaceable without interrupting operation. e. The battery compartment shall be accessible without special tools. f. The low-battery indication shall be no lower than the minimum power required for proper operation. g. Provided battery chargers shall meet US electrical standards, and shall be capable of operating from single phase AC power with voltage between 100V – 240V and frequency from 47Hz – 63Hz. h. Rechargeable-type batteries shall be fully recharged within 4 hours. i. A lamp or display shall be available to indicate when the batteries are fully charged. j. BRDs should have the ability to operate from an external power source (AC or DC).
Note: | Comments are required when the requirement is not verified.
|  | Ambient Conditions: | °C |  | %RH |  | mmHg
Test Equipment Used:
[6 empty] | Yes | No
Can the BRD operate for at least 8 hours without replacing batteries?
Manufacturer states the expected continuous operating time using the recommended batteries and the conditions (functional and environmental) used to determine this time?
Is the BRD equipped with a visible direct indicator of battery life?
Are the batteries replaceable without interrupting operation?
Can the backpack operate from AC power?
Can the backpack operate from DC power?
| Operational lifetime is: |  |  | Hours
Conditions:
Is the battery compartment accessible without special tools?
Was the monitor able to function using a DC power supply providing power at a voltage level which activates the low battery or battery replacement indicator?
Does the backpack provide an indicator of battery condition?
Do battery chargers meet US electrical standards?
Rechargeable-type batteries recharge within 4 hours?
Is a lamp or similar device is available to indicate when the batteries are fully charged?
Comments:
| Completed by: |  |  |  |  | Date:
| Reviewed by: |  |  |  |  | Date:
Draft 3