## Software Technical Assessment
Software Technical Assessment
Software Description | WAMRF-005 (Rev. 30), Volume Gravimetric Calibration Workbook | Approval Date | 2014-12-10
Assessed By | Dan Wright | Approved By | Dan Wright
Codes | Assessment | Pass/Fail | Evidence
A. Software Inspection | Spreadsheet is clear and makes sense | Pass | All worksheet are appropriately named and easy to understand
| There are instructions for use | Pass | Instructions worksheet
| Instructions and data input appear on the visible portion of the first worksheet | Pass | The workbook opens to the instruction worksheet
| Data-entry fields are labeled and color coded (it is recommended to avoid red and green) | Pass | Light yellow for blank, light blue for non-blank cells
| The Standard Operating Procedure (SOP) used is clearly specified | Pass | Specified on Uncertainty worksheet
| The number of digits to be rounded to is specified | Pass | Air Density and Water Density are rounded according to SOP 14 and the volume correction and uncertainty are rounded according to NISTIR 6969, GLP 9
| The user is warned/notified whenever data-entry fields are left blank | Pass | In data validation cells only
| Data-entry fields are “blank” when opened, preventing loss of old data and ensuring that old data is not used with the current calculations | Pass | Data entry fields on the workbook template are blank. Instructions tell user to break links before saving after a calibration is performed
| The software opens at the right location within the file | Pass | Opens to the Instruction worksheet
| Unused fields/cells are locked | Pass | Unused cells are locked and worksheets are protected
| Rows/columns that the operator need not see are hidden | N/A | No hidden cells or worksheets
| Unused sheets are removed | Pass | No unused worksheets
| Worksheets are named appropriately | Pass | Worksheets are appropriately named
B. Mathematical Specification | The correct SOP is used | Pass | SOP 14 in embedded
| The formulae and methods chosen from that SOP are specified | Pass | Methods for option A or option B are selectable and formulae are show on the SOP 14 and Uncertainty worksheet
| Sources and references for formulae are specified | Pass | As specified in SOP
| The chosen SOP, its methods, and its formulae, are appropriate to the level of precision/uncertainty | Pass | For Volume I Gravimetric calibrations
C. Code review | The formulae in the fields exactly match the SOP | Pass | Formulae on calculation worksheet match SOP
| Repeated calculations appropriately reference the correct cells | Pass
| Calculations, when tested using standard data or reference test data, show appropriate accuracy | Pass | Tested using SOP example
| Rounding is done at the appropriate locations in the file | Pass | Air Density, Water Density, and end volume results rounded
D. Numerical Stability | Calculations are stable as determined by an evaluation that uses large numbers and small differences | Pass | No issues encountered
| Fields, therefore their content, are categorized as “Number” and not “General” when appropriate, and vice versa | Pass | Formatted correctly throughout workbook
| “Number” cells are locked to a limited number of decimal places; this limit is appropriate to the values being used | Pass | Where appropriate
E. Component Testing | Each macro used is functional | N/A | No Macros
| Each command/button is functional | N/A | No command buttons
| Combinations of interdependent macros are functional | Pass | No issues encountered
| Plotted graphs are accurate | Pass | Graphs on uncertainty worksheet display properly
| Worksheets/reports print properly, if needed to | Pass | Worksheets formatted for printing as necessary
| Conditional (color and non-color) formatting is functional | Pass | On data entry and calculation worksheets
F. Numerical Reference | Look-up tables and lists match the latest calibration report. | Pass | Linked tables are validated and up to date
| Uncertainties match the latest Scope | Pass | Test on uncertainty worksheet will not allow an uncertainty that is better than the latest published scope
| Values that reference another workbook or spreadsheet are dated | N/A | Linked references have been validated and are up to date
| When a master list’s date is updated, the file references (A) an old value, (B) a default value, (C) displays zero or (D) an error message, as desired by the user | N/A
G. Embedded Data Evaluation | Embedded data (conversion factors, reference values, etc) is correct | Pass | All embedded data has been validated
| The evaluation of the embedded data is dated and documented | Pass | Records keep in Software Validation folder on computer hard drive
H. Back-to-Back Testing | Newer spreadsheets and older spreadsheets agree down to the level of intermediate calculations; this evaluation is dated and documented | Pass | Evaluated and filed
I. Analysis With Out Computer Assistance | Hand calculations agree with those generated by the spreadsheet, or if they disagree, the differences are significantly smaller than the reported uncertainty | N/A | No hand calculations were done
J. Security | Equation and calculation cells are protected against inadvertent editing | Pass | Cells are locked and worksheets are password protected
| Cells are locked in place; they cannot be moved/dragged | Pass | Locked and password protected
| Confidentiality of passwords is appropriate | Pass | State Metrologist and Program Manager only have access
| Files are backed up automatically | Pass | Backed up to agency servers weekly
| Additional back-up is available at alternate facilities | Pass | Agency IT backs up servers weekly
| Files on network drives cannot be accidentally deleted | Pass | Agency IT keeps an additional back up at a remote location

## Documentation
Documentation
Date | Description | Initials
2001-06-14 | This workbook was developed by Dan Wright, Washington State Department of Agriculture, Weights and Measures Program, Metrology Laboratory. It follows the procedure documented in NISTIR 7383, SOP 14. NIST WMD has evaluated this spreadsheet and provided useful feedback in its development. | DAW
2004-03-03 | Updated water density formula to the Patterson/Morris equation provided by NIST, WMD. | DAW
2004-05-12 | Changed entire workbook to be more user friendly. Updated uncertainties. | DAW
2004-09-07 | Added new uncertainty analysis. | DAW
2007-02-06 | Corrected errors found during validation. Changed uncertainty analysis to the Kragten spreadsheet. Added tolerances for volumetric vessels. | DAW
2007-03-30 | Added unit column to temperature standards. Reformatted standards tables to make it easier to enter or delete standards. | DAW
2007-09-24 | Rounded air density formula to 6 places and water density formula to 8 places. | DAW
2008-02-20 | Removed 'Standards' worksheet. Linked customer and standards cells to the 'Customer' and 'Standards' workbooks. Reformatted 'Documentation' worksheet. | DAW
2008-11-25 | Modified row 21 of 'Data Entry' worksheet to reflect standard deviation of trials instead of range of trials. | DAW
2008-11-25 | Renamed 'Rmean' to 'PooledSD' in cell I21 on 'Data Entry' worksheet. | DAW
2008-11-25 | Modified row 33 of 'Calculations' worksheet to reflect standard deviation of trials instead of range of trials. | DAW
2008-11-25 | Changed cell 26 of 'Uncertainty Analysis' worksheet to reflect 'PooledSD' instead of 'Rmean'. | DAW
2008-11-25 | Deleted Table 9.1 on 'Tables & Lists' worksheet. | DAW
2008-11-25 | Added NIST standard to procedure on 'Reports Pages 1-3' worksheet. | DAW
2009-01-22 | Converted to Excel 2007, updated links. | DAW
2009-01-22 | Color coded data entry cells. Edited all comments. | DAW
2009-01-22 | Added Standards Used table on Data Entry worksheet. | DAW
2009-01-22 | Reformatted conditional formatting. | DAW
2009-01-22 | Redefined water temperature cells, added water temperature to Calculations sheet for each trial. | DAW
2009-10-29 | Removed tolerance worksheet. Linked to Volume Tolerances workbook. | DAW
2009-10-29 | Added information worksheet. | DAW
2009-10-29 | Validated links and updated data. | DAW
2009-10-29 | Validated workbook, file name "WA2009-10-29 WAMRF-005 (Rev. 18), SOP 14 Workbook V&V.pdf". Electronic files in laboratory computer C:\ drive and agency server H:\ drive and a paper copy in lab files. | DAW
2010-03-19 | Added water purity test to data entry sheet, modified sentence for water quality in report. (AR0166) | DAW
2010-03-19 | Validated workbook, file name "WA2010-03-19 WAMRF-005 (Rev. 19), SOP 14 Workbook V&V.pdf". Electronic files in laboratory computer C:\ drive and agency server H:\ drive and a paper copy in lab files. | DAW
2010-09-02 | Modified data entry cells for internal references verses external references where possible. Deleted unnecessary list on the Tables & Lists worksheet. | DAW
2010-09-02 | Modified Observations area of the Data Entry worksheet to clarify water temperature measurements. | DAW
2010-09-02 | Added Software Technical Assessment worksheet. Embedded latest version of SOP 14 in SOP 14 worksheet. | DAW
2010-09-03 | Validated external links. Updated NIST 105-3 uncertainty requirement on the calculations worksheet. | DAW
2010-09-03 | Validated workbook, file name "WA2010-09-03 WAMRF-005 (Rev. 19), SOP 14 Workbook V&V.pdf". Electronic files in laboratory computer C:\ drive and agency server H:\ drive and a paper copy in lab files. | DAW
2012-02-15 | Added FORMAT_CX and ROUND_UNC visual basic modules for rounding results and uncertainties IAW NISTIR 6969, GLP 9 using the FmtCX() and Rnd2() functions. | DAW
2012-02-15 | Added SI statement and conversion factors on the Report Pages 1-3 spreadsheet.
2012-02-15 | Validated links and ranges. | DAW
2012-02-15 | Validated workbook, file name "WA2012-02-15 WAMRF-005 (Rev. 21), SOP 14 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2012-04-06 | Changed "Certification Statement" to "Accreditation Statement" on Report Pages 1-3 worksheet to be in compliance with NVLAP requirement Annex A, l). | DAW
2012-05-28 | Embedded 2012 SOP | DAW
2012-06-18 | Updated water density formula on Calculations worksheet | DAW
2012-06-18 | Updated air density formula on CIPM Air Density worksheet | DAW
2012-06-18 | Modified Data Entry worksheet observations section and added degrees of freedom and best uncertainty k factor entries | DAW
2012-06-18 | Modified Uncertainty Analysis worksheet and references to water and air density formula uncertainties | DAW
2012-06-18 | Updated k factor based on degrees of freedom IAW SOP 29 (2012 edition) | DAW
2012-06-18 | Validated ranges and links | DAW
2012-06-18 | Validated workbook, file name "WA2012-06-18 WAMRF-005 (Rev. 22), SOP 14 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2012-08-13 | Inserted new WSDA Logo on "Report Pages 1-3" worksheet. | DAW
2013-02-18 | Removed GLP Rounding macros | DAW
2013-02-18 | Entered new rounding technique and corrected formulas | DAW
2013-02-18 | Validated links and range names. | DAW
2013-02-18 | Validated workbook, file name "WA2013-02-18 WAMRF-005 (Rev. 23), SOP 14 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2013-09-18 | Changed all ISBLANK() statements. | DAW
2013-09-18 | Redefined uncertainty components, added the Welch-Satterthwaite formula for effective degrees of freedom. | DAW
2013-09-18 | Revised the Control Data worksheet | DAW
2013-09-18 | Validated links and range names and spell checked. | DAW
2013-09-18 | Split Uncertainty Analysis to two worksheets, one for Option A and one for Option B. | DAW
2013-09-18 | Validated workbook, file name "WA2013-09-18 WAMRF-005 (Rev. 24), SOP 14 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-05-30 | Modified Data Entry sheet Mass Standards Data cells and names | DAW
2014-05-30 | Deleted Uncertainty Analysis Option A and Option B sheets and added new Uncertainty Analysis sheet | DAW
2014-05-30 | Validated range names and spell checked. | DAW
2014-05-30 | Validated workbook, file name "WA2014-05-30 WAMRF-005 (Rev. 25), SOP 14 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-06-05 | Added uncertainty for reading the meniscus. | DAW
2014-06-05 | Validated range names and spell checked. | DAW
2014-06-05 | Validated workbook, file name "WA2014-06-05 WAMRF-005 (Rev. 26), SOP 14 Workbook V&V.pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-06-06 | Corrected uncertainty conversions for Mass and Mass Density uncertainties to mL. | DAW
2014-06-06 | Revised measurement process and uncertainties in-accordance-with SOP 14, dated April 2013.
2014-06-06 | Validated range names and spell checked. | DAW
2014-06-06 | Validated workbook, file name "WA2014-06-06 WAMRF-005 (Rev. 27), SOP 14 Workbook V&V (Option A & B).pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-06-12 | Changed units on Uncertainty worksheet cells L94, L97, L100 & L103 to "mL" instead of "mg" & "g" | DAW
2014-06-12 | Validated range names and spell checked. | DAW
2014-06-12 | Validated workbook, file name "WA2014-06-12 WAMRF-005 (Rev. 28), SOP 14 Workbook V&V (Option A & B).pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-06-30 | Fixed formula in cell K101 on the Uncertainty Analysis worksheet which referenced the BestUk instead of the BestUdf value. | DAW
2014-06-30 | Made minor formatting changes to Calculations and Uncertainty worksheets. | DAW
2014-06-30 | Reworded the Uncertainty statement in the Report. | DAW
2014-06-03 | Validated range names and spell checked. | DAW
2014-06-30 | Validated workbook, file name "WA2014-06-30 WAMRF-005 (Rev. 29), SOP 14 Workbook V&V (Option A & B).pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW
2014-10-29 | Changed statistical df from 100 to 10000. | DAW
2014-12-10 | Deleted viscosity uncertainty, not enough information to analyze this uncertainty. | DAW
2014-12-10 | Corrected reference in effective degrees of freedom on Uncertainty worksheet and deleted extra plus sign. Changed J85^4/L86 to J85^4/L85. | DAW
2014-12-10 | Validated range names and spell checked. | DAW
2014-12-10 | Validated workbook, file name "WA2014-12-10 WAMRF-005 (Rev. 30), SOP 14 Workbook V&V (Option A & B).pdf". Electronic copies in laboratory computer C:\ drive and agency server H:\ drive and paper copy in laboratory files. | DAW

## SOP 14
Note on A1: Unprotect the worksheet, then double click the document to open in Adobe Acrobat

## Instructions
Instructions
No. | Description
1 | The uncertainty analysis used in this workbook is only one way of determining uncertainties for this measurement. Please review and insure this method meets your requirements and expectations. Modify as necessary.
2 | This workbook must be validated by the user before use and after any changes or modifications.
3 | The worksheets are password protected. Password is password. You should protect each sheet with your own unique password. Place a password hint here so you don't forget it. Password hint is: "?"
4 | Validate all ranges after any changes and during annual reviews.
5 | Open as a new workbook and save using the Report Number and customer name.
6 | Enter the appropriate data in the light yellow cells on the Data Entry worksheet.
7 | Modify both report worksheets for your laboratory and audit to insure the report meets requirements of ISO/IEC 17025. (hint: I convert the report worksheets to a .pdf file to keep as a permanent copy).
8 | To view the SOP, unprotect the worksheet then double click on it.
Range Names
Name | Range
a | ='Data Entry'!$J$12
A_tw1after | ='Data Entry'!$I$68
A_tw1before | ='Data Entry'!$I$66
A_tw2after | ='Data Entry'!$I$84
A_tw2before | ='Data Entry'!$I$82
B_tw1after | ='Data Entry'!$I$103
B_tw1before | ='Data Entry'!$I$101
B_tw2after | ='Data Entry'!$I$119
B_tw2before | ='Data Entry'!$I$117
BalUnits | ='Data Entry'!$K$22
BestU | ='Data Entry'!$F$24
BestUdf | ='Data Entry'!$K$24
BestUk | ='Data Entry'!$I$24
Cal_Date | ='Data Entry'!$D$21
CCE | ='Tables & Lists'!$L$5:$M$14
CIPM_Pa1 | ='CIPM Air Density'!$D$7
CIPM_Pa2 | ='CIPM Air Density'!$D$13
Condition | ='Data Entry'!$D$16
Cs_empty | ='Data Entry'!$D$45
Cs_filled | ='Data Entry'!$D$41
CustomerID | ='Data Entry'!$K$4
Description | ='Data Entry'!$C$10
Description_list | ='Tables & Lists'!$A$4:$A$9
df | ='Data Entry'!$K$23
df_Ms_empty | ='Data Entry'!$I$45
df_Ms_filled | ='Data Entry'!$I$41
Final_k | ='Uncertainty Analysis'!$K$102
Final_U | ='Uncertainty Analysis'!$K$103
Final_Veff | ='Uncertainty Analysis'!$K$101
Interval | ='Data Entry'!$K$18
IntervalQ | ='Data Entry'!$F$18
JO_No | ='Data Entry'!$J$16
k | ='Uncertainty Analysis'!$K$96
k_Ms_empty | ='Data Entry'!$H$45
k_Ms_filled | ='Data Entry'!$H$41
Line_w | ='Data Entry'!$K$15
Material | ='Data Entry'!$C$12
Material_List | ='Tables & Lists'!$L$5:$L$14
mcr | ='Uncertainty Analysis'!$B$69
MeasuredU | ='Uncertainty Analysis'!$K$97
MFG | ='Data Entry'!$C$11
Ms_empty | =Calculations!$C$17
Ms_empty_nom | ='Data Entry'!$C$45
Ms_filled | =Calculations!$C$16
Ms_filled_nom | ='Data Entry'!$C$41
Neck_d | ='Data Entry'!$C$15
Nom_Val_List | ='Tables & Lists'!$C$5:$C$12
NomVal | ='Data Entry'!$D$13
NomValueUnit | ='Data Entry'!$A$13
Obs1_Run1 | =Calculations!$B$21
Obs1_Run2 | =Calculations!$F$21
Obs2_Run1 | =Calculations!$B$22
Obs2_Run2 | =Calculations!$F$22
Obs3_Run1 | =Calculations!$B$23
Obs3_Run2 | =Calculations!$F$23
Obs4_Run1 | =Calculations!$B$24
Obs4_Run2 | =Calculations!$F$24
option | ='Data Entry'!$J$25
P_0 | ='Uncertainty Analysis'!$B$74
P_1_Aend | ='Data Entry'!$I$73
P_1_Astart | ='Data Entry'!$I$63
P_1_Bend | ='Data Entry'!$I$106
P_1_Bstart | ='Data Entry'!$I$96
P_2_Aend | ='Data Entry'!$I$89
P_2_Astart | ='Data Entry'!$I$79
P_2_Bend | ='Data Entry'!$I$122
P_2_Bstart | ='Data Entry'!$I$112
P_corr | ='Data Entry'!$D$51
P_Ms_empty | ='Data Entry'!$J$45
P_Ms_filled | ='Data Entry'!$J$41
Pa1 | =Calculations!$C$5
Pa2 | =Calculations!$C$6
Pn | ='Uncertainty Analysis'!$D$70
PO_Number | ='Data Entry'!$I$5
POC_Name | ='Data Entry'!$I$6
POC_Phone | ='Data Entry'!$I$7
PooledSD | ='Data Entry'!$F$23
Pr | ='Uncertainty Analysis'!$B$70
Print_Titles | =Instructions!$11:$12
Pt | ='Uncertainty Analysis'!$B$71
Pw1 | =Calculations!$C$12
Pw2 | =Calculations!$C$13
RefTemp | ='Data Entry'!$J$13
RefTempUnit | ='Data Entry'!$G$13
Reported_df | ='Tables & Lists'!$H$13
Reportedk | ='Tables & Lists'!$F$13
ReportedUnc | ='Tables & Lists'!$D$13
rh_corr | ='Data Entry'!$D$52
RptNo | ='Data Entry'!$I$1
SG_d | ='Data Entry'!$G$15
Slicker_Stmt | ='Tables & Lists'!$A$21
SN | ='Data Entry'!$J$11
Spec | ='Data Entry'!$D$14
Spec_List | ='Tables & Lists'!$J$5:$J$15
t_air_corr | ='Data Entry'!$D$50
t_air1_Aend | ='Data Entry'!$I$72
t_air1_Astart | ='Data Entry'!$I$62
t_air1_Bend | ='Data Entry'!$I$105
t_air1_Bstart | ='Data Entry'!$I$95
t_air2_Aend | ='Data Entry'!$I$88
t_air2_Astart | ='Data Entry'!$I$78
t_air2_Bend | ='Data Entry'!$I$121
t_air2_Bstart | ='Data Entry'!$I$111
t_water_corr | ='Data Entry'!$D$49
Tech | ='Data Entry'!$G$21
Tech_Int | ='Data Entry'!$K$21
ToContain_Stmt | ='Tables & Lists'!$A$20
ToDeliver_Stmt | ='Tables & Lists'!$A$19
Tol | ='Data Entry'!$J$14
tw1 | =Calculations!$C$10
tw2 | =Calculations!$C$11
TypeCal | ='Data Entry'!$J$17
TypeCal_List | ='Tables & Lists'!$A$10:$A$11
U_1_Aend | ='Data Entry'!$I$74
U_1_Astart | ='Data Entry'!$I$64
U_1_Bend | ='Data Entry'!$I$107
U_1_Bstart | ='Data Entry'!$I$97
U_2_Aend | ='Data Entry'!$I$90
U_2_Astart | ='Data Entry'!$I$80
U_2_Bend | ='Data Entry'!$I$123
U_2_Bstart | ='Data Entry'!$I$113
U_Ms_empty | ='Data Entry'!$F$45
U_Ms_filled | ='Data Entry'!$F$41
U_P | ='Data Entry'!$F$51
U_P_k | ='Data Entry'!$H$51
U_P_Ms_empty | ='Data Entry'!$K$45
U_P_Ms_filled | ='Data Entry'!$K$41
U_rh | ='Data Entry'!$F$52
U_rh_k | ='Data Entry'!$H$52
U_t | ='Data Entry'!$F$50
U_t_a_k | ='Data Entry'!$H$50
U_t_w | ='Data Entry'!$F$49
U_t_w_k | ='Data Entry'!$H$49
UPa | ='Uncertainty Analysis'!$B$72
Veff | ='Uncertainty Analysis'!$K$95
VolTable | ='Tables & Lists'!$C$5:$H$12
WODate | ='Data Entry'!$J$10

## Report Pages 1-3
REPORT OF CALIBRATION
Issued To:
Point of Contact:
Purchase Order Number:
Report Number:
Calibration Date
Calibration Due Date
This is to certify that the information contained in this report is true and correct as of the date of calibration.
Dan Wright, Program Specialist 3 |  |  |  |  | Date of Issue
Accreditation LOGO |  | Accredited by the Accreditation Authority for the specific scope of accreditation under lab code XXXXXX-X. This report may not be used to claim product endorsement by Accreditation Authority or any other government agency, and may not be reproduced, except in full, without written approval from the laboratory.
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date
Artifact(s) Description
Test Item: |  |  |  | Date Received:
Specification: |  |  |  | Condition:
Serial Number: |  |  |  | Manufacture:
Material: |  |  |  | Delivery Type:
Calibration Information
Job Order #: |  |  |  | Temperature:
Metrologist: |  |  |  | Humidity:
Procedure: | NISTIR 7383, SOP 14, |  |  | Pressure:
Balance: |  |  |  | Water Temperature:
Laboratory Reference Standards Used
Description |  | Serial Number | Cert. Number |  | Cal Date | Cal Due
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
0 |  | 0 | 0 |  | 0 | 0
Traceability Statement
The artifact(s) described in this report have been compared to the Standards of the State of Washington. The Standards of the State of Washington are traceable to the National Institute of Standards and Technology (NIST) and are part of a comprehensive measurement assurance program for ensuring continued accuracy and measurement traceability within the level of uncertainty reported by this laboratory. The International System of Units (SI) for volume is the cubic meter (m³) (see Conversion Factors on page 3). The report number for this report is the only unique report number to be used in referencing measurement traceability for the artifact(s) described in this report.
Uncertainty Statement
The combined standard uncertainty includes uncertainties for the standard(s), for the measurement process, for reading the meniscus, for the material cubical coefficient of expansion, for water density corrections, for air density corrections, for viscosity and for the environmental equipment used. The combined standard uncertainty is multiplied by a coverage factor, k, to give the expanded uncertainty, which defines an interval with a 95.45 % level of confidence. The expanded uncertainty presented in this report is consistent with NIST Technical Note 1297. A component for the effects of viscosity was not included in the uncertainty budget.
Accreditation Statement
Accredited by the Accreditation Authority for the specified scope of accreditation under lab code XXXXXX-X. This laboratory meets the requirements of ISO/IEC 17025 and ANSI/NCSL Z540-1.
WSDA Weights and Measures Metrology Laboratory
Report Number: | [5 empty] | Calibration Date
Pertinent Information
● In-accordance-with ISO/IEC FDIS 17025, General Requirements for the Competence of Testing and Calibration Laboratories, paragraph 5.10.4.4 ‘A calibration certificate (or calibration label) shall not contain any recommendation on the calibration interval except where this has been agreed with the client. This requirement may be superseded by legal regulations.’
● In-accordance-with Washington Administrative Code (WAC) Chapter 16-663, Service Agents -- Reporting, Test Procedures, Standards And Calibration Of Weighing And Measuring Devices, Section 16-663-130, Adequacy of standards and submission of standards for certification, paragraph 2, ‘… All standards used for servicing, repairing and/or calibrating commercial weighing and measuring devices must be submitted at least every two years for examination and certification…’
● The artifact described above was calibrated using water that meets or exceeds ASTM D-1193 Type III Reagent Water specifications.
● The results listed in this report relate only to the artifacts described and extent of calibrations performed.
● The is considered in-tolerance when the error is equal to or less than the specified tolerance minus the measurement uncertainty.
#VALUE!
Conversion Factors
From NIST Special Publication 811, Guide for the Use of the International System of Units (SI)
Factors in boldface are exact
To convert from |  |  | to |  | multiply by
cubic inch (in³) |  |  | to cubic meter (m³) |  | 1.638 706 4 E-05
fluid ounce (U.S.) (fl oz) |  |  | to cubic meter (m³) |  | 2.957 353 E-05
gallon (U.S.) (gal) |  |  | to cubic meter (m³) |  | 3.875 412 E-03
gill (U.S.) (gi) |  |  | to cubic meter (m³) |  | 1.182 941 E-04
liter (L) |  |  | to cubic meter (m³) |  | 1.0 E-03
pint (U.S. liquid) (liq pt) |  |  | to cubic meter (m³) |  | 4.731 765 E-04
quart (U.S. liquid) (liq qt) |  |  | to cubic meter (m³) |  | 9.463 529 E-04

## Report Page 4
WSDA Weights and Measures Metrology Laboratory
Report Number: |  |  |  |  | Calibration Date
Calibration Results
Nominal Volume | As Found/Left | Reference Temperature | Specification Tolerance | Uncertainty | Cubical Coefficient of Expansion

## Data Entry
Data Entry | [6 empty] | Report Number:
Customer Information
Customer ID
PO#
POC
Phone
Artifact Information
Description | [5 empty] | Date Received
Manufacture | [5 empty] | Serial Number
Material | [5 empty] | Cubical Coefficient of Exp. (/ºC) =
Artifact Specifications | [6 empty] | Tolerance (mL)
Neck Inside Dia. (mm) |  |  | Sight Glass Inside Dia. (mm) |  |  |  | Graduation Line Width (mm)
Artifact Condition | [7 empty] | JO#
Type of Volume Container (To Deliver for Wet Calibration or To Contain for Dry Calibration)
Was a Calibration Interval Requested or is one Required? | [5 empty] | What is the Calibration Interval (months)?
Measurement Information
Date of Calibration |  |  |  |  | Technician |  |  |  | Tech. Int.
Balance ID Code |  |  | Balance: |  |  |  |  | Balance Unit:
Pooled Std Dev of the Trials from Control Chart (mL) | [7 empty] | Degrees of Freedom
Accredited U (e.g. NVLAP, A2LA, etc.) (mL) | [6 empty] | U k factor |  | U df
Measurement Option (A = one point balance calibration; B = two point balance calibration)
Standards Used
Standard Item ID | Description |  |  |  | Serial Number |  | Report Number |  | Calibration Date | Calibration Due
Mass Standard(s) Data
Mass ID | ID Code | Nominal Mass (g) | True Mass Correction (mg) |  | Us (mg) |  | k-factor | Degrees of Freedom | Density (g/cm3) | Uρs (g/cm3) k=1
Environmental Standard(s) Data
Parameter |  | ID Code | Correction |  | U |  | k factor |  | Unit
Temperature, Water | [8 empty] | ºC
Temperature, Air | [8 empty] | ºC
Barometric Pressure | [8 empty] | mm Hg
Humidity | [8 empty] | % RH
Water Conductivity Test
According to NISTIR 7383, SOP 14 and GLP 10, ASTM Type III or IV Reagent Water is recommended as adequate for gravimetric calibration purposes.
Conductivity Meter Indication |  |  |  |  | µS/cm @ 25 ºC
Observations
Option A - One Point Balance Calibration Trial 1 | [7 empty] | Balance Unit =
Starting Conditions Trial 1 |  |  |  | Start Time (hh:mm)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
O1 | Zero the balance, then place standard(s), MS filled, slightly larger than the mass of the filled vessel
O2 | Zero the balance, then place dry/wet-down vessel on balance
Record temperature of water used to fill the vessel (ºC)
O3 | Zero the balance, then place filled vessel on balance
Record temperature of water in vessel (ºC)
Ending Conditions Trial 1 |  |  |  | Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
End Time (hh:mm)
Option A - One Point Balance Calibration Trial 2 | [7 empty] | Balance Unit =
Starting Conditions Trial 2 |  |  |  | Start Time (hh:mm)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
O1 | Zero the balance, then place standard(s), MS filled, slightly larger than the mass of the filled vessel
O2 | Zero the balance, then place dry/wet-down vessel on balance
Record temperature of water used to fill the vessel (ºC)
O3 | Zero the balance, then place filled vessel on balance
Record temperature of water in vessel (ºC)
Ending Conditions Trial 2 |  |  |  | Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
End Time (hh:mm)
Option B - One Point Balance Calibration Trial 1 | [7 empty] | Balance Unit =
Starting Conditions Trial 1 |  |  |  | Start Time (hh:mm)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
O1 | Zero the balance, then place standard(s), MS empty, slightly larger than the mass of the dry/wet-down vessel
O2 | Zero the balance, then place dry/wet-down vessel on balance
O3 | Zero the balance, then place standard(s), MS filled, slightly larger than the mass of the filled vessel
Record temperature of water used to fill the vessel (ºC)
O4 | Zero the balance, then place filled vessel on balance
Record temperature of water in vessel (ºC)
Ending Conditions Trial 1 |  |  |  | Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
End Time (hh:mm)
Option B - One Point Balance Calibration Trial 2 | [7 empty] | Balance Unit =
Starting Conditions Trial 2 |  |  |  | Start Time (hh:mm)
Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
O1 | Zero the balance, then place standard(s), MS empty, slightly larger than the mass of the dry/wet-down vessel
O2 | Zero the balance, then place dry/wet-down vessel on balance
O3 | Zero the balance, then place standard(s), MS filled, slightly larger than the mass of the filled vessel
Record temperature of water used to fill the vessel (ºC)
O4 | Zero the balance, then place filled vessel on balance
Record temperature of water in vessel (ºC)
Ending Conditions Trial 2 |  |  |  | Air Temperature (ºC)
Barometric Pressure (mm Hg)
Relative Humidity (% RH)
End Time (hh:mm)

## Control Data
Control Chart Information
Note: Copy and Paste Appropriate Row(s) to Control Chart
R-Chart Information
Tech. Initials | Date | Test Number | Volume (Trial 1) @ ºC (mL) | Volume (Trial 2) @ ºC (mL) | Start Humidity (% RH) | Finish Humidity (% RH) | Start Air Temp. (ºC) | Finish Air Temp. (ºC) | Water Temp. (Trial 1) (ºC) | Water Temp. (Trial 2) (ºC)

## Calculations
Calculations
Air Density Determinations
Air density results are rounded to 9 digits (g/cm³) according to SOP 14
| Run 1 ρa = |  | g/cm3 | Air Density, SOP 2, Option B, 2012 (CIPM 2007)
| Run 2 ρa = |  | g/cm3
Water Density Determinations
Water density results are rounded to 8 digits (g/cm³) according to SOP 14
| Water temperature Trial 1 = |  | º C | M. Tanaka, G. Girard, R. Davis, A. Peuto, and N. Bignell equation, Metrologia, 38, 301-309 (2001) Water Density Formula for Air Saturated Water
| Water temperature Trial 2 = |  | º C
| Run 1 ρw = |  | g/cm3
| Run 2 ρw = |  | g/cm3
Mass Standards Report Values
| Ms (filled) = |  | g | true mass of standard(s)
Observation Values
Trial 1 |  |  | Trial 2
O1 = |  | g | O1 = |  | g
O2 = |  | g | O2 = |  | g
O3 = |  | g | O3 = |  | g
Volume Determinations
| Trial 1 Vt = |  | mL | trial 1 & 2 volume at measured temperature
| Trial 2 Vt = |  | mL
| Trial 1 V @ �C = |  | mL | trial 1 & 2 volume at reference temperature
| Trial 2 V @ �C = |  | mL
| V @ �C = |  | mL | average of trial 1 & 2 at reference temp.
Uncertainty Evaluation
Criteria: The expanded uncertainty must be less than 1/3 of the stated tolerance. For NIST HB 105-3, the expanded uncertainty must be less than 0.05 % of the measured volume.
| 1/3 of Tolerance = |  | mL
| Expanded Uncertainty = |  | mL
Tolerance Evaluation
Criteria: The measured volume minus the nominal volume shall not differ more than the maximum permissible error minus the expanded uncertainty.
| Volume Error = |  | mL
| Tolerance - U (k≈2)= |  | mL
Measurement Control
| Std. Deviation of the Trials = |  | mL
Results Formatted For Calibration Report
Volume @ �C =
Effective Degrees of Freedom =
k factor
Uncertainty (uc*k) =

## CIPM Air Density
CIPM Air Density
Air Density, SOP 2, Option B, 2012 (CIPM 2007)
Air Density Calculation Run 1
Barometric Pressure (mm Hg)
Temperature (ºC)
Relative Humidity (%)
Calculated Air Density (mg/cm3)
Air Density Calculation Run 2
Barometric Pressure (mm Hg)
Temperature (ºC)
Relative Humidity (%)
Calculated Air Density (mg/cm3)

## Uncertainty Analysis
Uncertainty Analysis
Specify Procedure
NISTIR 7383, SOP 14 Option A, Gravimetric Calibration of Volumetric Ware Using an Electronic Balance
Specify Equation(s)
Measurement Option A - One Point Balance Calibration
where
Vt = | Represents either the "to contain" or "to deliver" volume (depending on whether O3 represents a dry or a "wet down" container at the temperature of measurement) (cm³ or mL)
MS = | Mass standard(s) slightly more than the mass of the filled vessel (i.e., true mass, vacuum mass) (g)
On = | Balance Observations
ρa = | Density of air at the conditions of calibration (g/cm³)
ρs = | Density of mass standards (g/cm³)
ρw = | Density of water at temperature of measurement (g/cm³)
Measurement Option B - Two Point Balance Calibration
where
Vt = | Represents either the "to contain" or "to deliver" volume (depending on whether O3 represents a dry or a "wet down" container at the temperature of measurement) (cm³ or mL)
MS1 = | Mass standard(s) slightly more than the mass of the empty/drained vessel (i.e., true mass, vacuum mass) (g)
MS2 = | Mass standard(s) slightly more than the mass of the filled vessel (i.e., true mass, vacuum mass) (g)
On = | Balance Observations
ρa = | Density of air at the conditions of calibration (g/cm³)
ρs = | Density of mass standards (g/cm³)
ρw = | Density of water at temperature of measurement (g/cm³)
Volume at Reference Temperature
where
Vref = | Volume at reference temperature (cm³ or mL @ ºF or ºC)
α = | Coefficient of cubical expansion of the container being calibrated (see NISTIR 6969, Table 9.10)
t = | Temperature of water (ºC)
tref = | Reference temperature for the calibration (ºF of ºC)
Error in Reading Meniscus
where
Vol = | Volume error in reading the meniscus
= | The value of PI, 3.141 592 653 589 79
r = | The radius of the internal diameter
h = | The height of the graduation line or volume in question
Identify Components
Component | Description |  |  |  |  | Reference
s(p) | Uncertainty associated with the measurement process from the average range of several sets of measurements, normal distribution, 1s |  |  |  |  | NISTIR 6969, SOP 29, Para. 3.2.1.2. Standard deviation from a set of replicate measurements and NISTIR 6969, Statistical Techniques, paragraphs 8.3 & 8.4, Pooled standard deviation from standard deviation for each set of replicates from a control chart (value divided by 1.003 mL/g to convert sp from mL to g), Type A evaluation
u(s) | Uncertainty associated with the weights used, normal distribution, k =1 |  |  |  |  | Calibration Report divided by the k-factor, Type B Evaluation
u(ρS) | Uncertainty associated with the density of the weights used, rectangular distribution, k =1 |  |  |  |  | Use reported or typical from manufactures stated density (typical range is 0.05 g/cm³). (See OIML R111-1:2004(E) B.7.9 Test Method F), Type B evaluation
u(m) | Uncertainty associated with reading the meniscus of the standard and unknown vessel, triangular distribution |  |  |  |  | NISTIR 7383, GMP 3, page 4, equation for volume of a cylinder, Type B evaluation
u(ρa) | Density of air using the CIPM 2007 Air Density Formula, rectangular distribution |  |  |  |  | Uncertainty for air buoyancy correction, NISTIR 6969, SOP 2, Para. 4, the uncertainty in the CIPM 2007 air density equation is 0.000 026 4 mg/cm³ or 0.002 2 % of normal air density, uniform distribution.
u(ρw) | Uncertainty in the density of water measurement, rectangular distribution |  |  |  |  | Metrologia, 46 (2009), 196-198, Density of water: roles of the CIPM and IAPWS standards, paragraph 2, 0.88 ppm (0.00000088 g/cm3), Type B evaluation
u(tw) | Uncertainty of the thermometer used for measuring water temperature, rectangular distribution |  |  |  |  | NISTIR 7383, SOP 14, accurate to ± 0.1 ºC, Calibration Report divided by the k-factor, Type B evaluation
u(CCE) | Uncertainty of the Coefficient of Cubical Expansion for the vessel material, rectangular distribution |  |  |  |  | EURAMET cg-19, paragraph 5.3.6 5 % to 10 % of CCE
u(P) | Uncertainty of the thermometer used for measuring water temperature, rectangular distribution |  |  |  |  | NIST IR 7383, SOP 14, accurate to ± 135 Pa (1 mm Hg), Calibration Report divided by the k-factor, Type B evaluation
u(t) | Uncertainty of the thermometer used for measuring air temperature, rectangular distribution |  |  |  |  | NISTIR 7383, SOP 14, accurate to ± 0.5 ºC, Calibration Report divided by the k-factor, Type B evaluation
u(RH) | Uncertainty of the thermometer used for measuring water temperature, rectangular distribution |  |  |  |  | NISTIR 7383, SOP 14, accurate to ± 10 % RH, Calibration Report divided by the k-factor, Type B evaluation
Uncertainty Intermediate Calculations
Meniscus |  |  | Volume error in Neck (mL) =
Volume error in Sight Glass (mL) =
Mass Density |  |  | Trial 1 Vt (ρSE+uρSE) mL =
Trial 1 Vt (ρSF+uρSF) mL =
Air Density |  |  | Trial 1 Vt (ρa+uρa) mL =
Water Density |  |  | Trial 1 Vt (ρw+uρw) mL =
Water Temperature |  |  | Run 1 ρw (tw + utw) g/cm³ =
Coefficient of Expansion |  |  | Trial 1 V @ �C = (CCE + uCCE) mL =
Barometric Pressure |  |  | P (mm Hg) = |  | P + uP (mm Hg) = |  |  | Run 1 ρa (P + uP) g/cm³ = |  | #VALUE!
Air Temperature |  |  | t (ºC) = |  | t + ut (ºC) = |  |  | Run 1 ρa (t + ut) g/cm³ = |  | #VALUE!
Relative Humidity |  |  | RH (%) = |  | RH + uRH (%) = |  |  | Run 1 ρa (RH + uRH) g/cm³ = |  | #VALUE!
Quantity, Convert, Combined, Expand & Report
Component Description |  | Component Symbol | Estimated Uncertainty | Units | Reported Unc (mL) | Type | Distribution | Factor to Normalize | Standard Unc (mL) | % Contribution | Degrees of Freedom
Process Standard Deviation |  | s(p) | 0 | mL | 0 | A | Normal | 1 | 0 |  | 1
Mass Standards (Empty) |  | u(sE) | 0 | g | 0 | B | Normal | 1 | 0 |  | 1
Standards Density (Empty) |  | u(ρSE) | 0 | g/cm³ | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Mass Standards (Filled) |  | u(sF) | 0 | g | 0 | B | Normal | 1 | 0 |  | 1
Standards Density (Filled) |  | u(ρSF) | 0 | g/cm³ | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Meniscus |  | u(m) | 0 | mL | 0 | B | Triangular | 2.4495 | 0 |  | 1
Air Density |  | u(ρa) | 0 | mg/cm³ | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Water Density |  | u(ρw) | 0 | g/cm³ | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Water Temperature |  | u(tw) | 0 | ºC | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Coefficient of Expansion |  | u(CCE) | 0 | /ºC | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Barometric Pressure |  | u(P) | 0 | mm Hg | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Air Temperature |  | u(t) | 0 | ºC | 0 | B | Rectangular | 1.7321 | 0 |  | 1
Relative Humidity |  | u(RH) | 0 | % RH | 0 | B | Rectangular | 1.7321 | 0 |  | 1
[9 empty] | Combined U (uc) = |  | mL
Effective Degrees of Freedom, Welch-Satterthwaite formula, NIST Technical Note 1297, (B-1) (veff) =
k factor =
[9 empty] | Expanded U (uc*k) = |  | mL
[9 empty] | Accredited Effective Degrees of Freedom (veff) = | 0
[9 empty] | Accredited k factor = | 0
[9 empty] | Accredited U = | 0 | mL
Final Effective Degrees of Freedom (veff) =
Final k factor =
[9 empty] | Final U (uc*k) = |  | mL

## Tables & Lists
Tables & Lists | [6 empty] | Tables & Lists
Pick Lists for Data Validation Cells | Volume Conversion Table (VolTable) & Nominal Value Unit (Nom_Val_List) | [5 empty] | Specification List | Cubical Coefficient of Expansion Table (CCE_Table) & Material List (Materal_List)
Graduated Cylinder | NomValueUnit | Vref | U (k=) | Tolerance | Symbol | Nominal Value (mL) | Specification | Material | cm³/ºC
Volumetric Flask | Nominal Value (cubic inch) | #VALUE! | #VALUE! | #VALUE! | in³ | 0 | ASTM E 1272, Class A | Glass, Borosilicate (TICA) | 1e-05
Volumetric Container | Nominal Value (fluid ounce) | #VALUE! | #VALUE! | #VALUE! | fl oz | 0 | ASTM E 1272, Class B | Glass, Borosilicate (TICB) | 1.5e-05
Slicker Plate Standard | Nominal Value (gallon) | #VALUE! | #VALUE! | #VALUE! | gal | 0 | ASTM E 288, Class A | Glass, Soda-Lime | 2.5e-05
Prover | Nominal Value (gill) | #VALUE! | #VALUE! | #VALUE! | gi | 0 | ASTM E 288, Class A "Wide Mouth" | Plastic, Polycarbonate | 0.00045
Test Measure | Nominal Value (liquid pint) | #VALUE! | #VALUE! | #VALUE! | liq pt | 0 | ASTM E 288, Class B | Plastic, Polypropylene | 0.00024
| Nominal Value (liquid quart) | #VALUE! | #VALUE! | #VALUE! | liq qt | 0 | ASTM E 288, Class B "Wide Mouth" | Plastic, Polystyrene | 0.00021
| Nominal Value (liter) | #VALUE! | #VALUE! | #VALUE! | L | 0 | NIST HB 105-2 | Silica, Fused (Quartz) | 1.6e-06
| Nominal Value (milliliter) | #VALUE! | #VALUE! | #VALUE! | mL | 0 | NIST HB 105-3 | Steel, Prover, Low Carbon | 3.35e-05
| Final Unformatted Uncertainty = |  | mL k = |  | df = |  | Not Applicable | Steel, Stainless | 4.77e-05
[7 empty] | OIML R 43 | Steel, Terneplate | 3.5e-05
User Specified
Calibration Report Statements
● The delivers the stated volume at the reference temperature from the specified graduation line when it is emptied gradually within a 30 (± 5 seconds) pour and a 10-second drain while held at a 10º to 15º angle from vertical.
● The , when cleaned and dried, contains the stated volume at the reference temperature when filled to its nominal graduation line. The must always be cleaned and dried between successive uses for purposes of accuracy.
● The delivers the stated volume at the reference temperature is when it is drained for a 30 (± 5) second period after cessation of the main flow.