## Descriptive Information
Experiment related information | Input
Date of analysis | 2001-10-26
Element | Co
Identification label of the Test Sample | Example
Lot number of the Test Sample | 1
Nominal mass fraction of Test Sample (mg/g) | 9.952
SRM 3100 series number | 3113
SRM lot number | 630
Certified Mass Fraction (mg/g) of SRM 3113, lot 000630 | 9.996
Uncertainty of Certified Mass Fraction (mg/g) | 0.023
Coverage Factor, k, for the SRM Certified Uncertainty | 2.31
SRM Certificates of Analysis can be found at
https://srmors.nist.gov/view_detail.cfm?srm=3103a
Instrument related information | Input
Instrumental comparison method | HP-ICP
Working Response Level | 1000000
Mass of Working Solutions (g) | 60
Internal standard element | Sc
Mass fraction of Internal Standard Stock Solution (mg/g) | 10
Note on B7: For example, SRM 3101a for Arsenic.
Note on B8: Six digit number found on the SRM label and the matching Certificate of Analysis.
Note on B9: As found on the SRM Certificate of Analysis
Note on B10: The expanded uncertainty as found on the Certificate of Analysis
Note on B11: The coverage factor, k, for a 95 % confidence interval is usually given on the Certificate of Analysis of the SRM. If k is not given, a value of 2 should be assumed.
Note on B17: For example, ICP-OES.

## Instrument Sensitivity Test
HP-ICP Comparison of Example (lot 1) to NIST SRM 3113 (lot 000630), 10/26/2001.
F | measure instrument response for the Sensitivity Solutions of the analyte and internal standard
|  | ug/g | HP-ICP response | Sensitivity (response/(µg/g)) | Target working mass fraction | Sensitivity ratio
Co (Analyte) Sensitivity Solution |  | 10 | 1670869 | 167086.9 | 5.9849 | 0.037901
Sc (Int. Std.) Sensitivity Solution |  | 5 | 22042376 | 4408475.2 | 0.22684

## Solution Preparation
HP-ICP Comparison of Example (lot 1) to NIST SRM 3113 (lot 000630), 10/26/2001.
F | Prepare Sc Internal Standard Spike Solution. |  |  |  | g of Sc Internal Standard Stock Solution to dilute to 200g
0.94715
F | Into labeled bottles, prepare spiked Test Samples and SRMs following the target values listed below:
|  | g of Test Sample |  |  |  | g of SRM 3113 |  |  |  | g of IS Spike Solution
|  | Target | Actual |  |  | Target | Actual |  |  | Target | Actual
|  |  | Tare | Gross | Net |  | Tare | Gross | Net |  | Tare | Gross | Net
| SRM 1 |  |  |  |  | 2.5 | 6.275 | 8.882 | 2.607 | 20 | 8.882 | 29.335 | 20.453
| SRM 2 |  |  |  |  | 2.5 | 6.277 | 8.931 | 2.654 | 20 | 8.931 | 28.888 | 19.957
| SRM 3 |  |  |  |  | 2.5 | 6.241 | 8.889 | 2.648 | 20 | 8.889 | 28.457 | 19.568
| SRM 4 |  |  |  |  | 2.5 | 6.162 | 8.536 | 2.374 | 20 | 8.536 | 28.582 | 20.046
| Test Sample 1 | 2.5111 | 6.305 | 8.961 | 2.656 |  |  |  |  | 20 | 8.961 | 29.397 | 20.436
| Test Sample 2 | 2.5111 | 6.23 | 8.884 | 2.654 |  |  |  |  | 20 | 8.884 | 29.281 | 20.397
| Test Sample 3 | 2.5111 | 6.259 | 8.922 | 2.663 |  |  |  |  | 20 | 8.922 | 28.333 | 19.411
| Test Sample 4 | 2.5111 | 6.272 | 8.585 | 2.313 |  |  |  |  | 20 | 8.585 | 28.965 | 20.38
F | Prepare Working Solutions by dilution of spiked solutions.
g of spiked solution to dilute to 60 g
[5 empty] | SRM Samples |  | 0.32331
[5 empty] | Test Samples |  | 0.32347

## Instrument Measurements
HP-ICP Comparison of Example (lot 1) to NIST SRM 3113 (lot 000630), 10/26/2001.
F | Each solution must be measured 5 times in the order specified in the ID column.
Each measurement may include as many as 10 readings, but fewer readings may be made.
Instrument response values for analyte and internal standard must be entered in the table below.
Repeat # | ID | Element | Reading 1 | Reading 2 | Reading 3 | Reading 4 | Reading 5 | Reading 6 | Reading 7 | Reading 8 | Reading 9 | Reading 10 | Average | sd | rsd
1 | SRM 1 | Analyte | 1137374.554 | 1137878.917 | 1145514.296 | 1137828.692 | 1133506.365 | [5 empty] | 1138420.5648 | 4364.8659 | 0.3834%
|  | IS | 1113068.125 | 1113556.491 | 1121265.873 | 1112967.87 | 1109084.133 | [5 empty] | 1113988.4984 | 4446.5824 | 0.3992%
| Test Sample 1 | Analyte | 1141089.778 | 1140492.796 | 1126372.02 | 1129086.325 | 1128281.425 | [5 empty] | 1133064.4688 | 7125.2656 | 0.6288%
|  | IS | 1099376.973 | 1098617.697 | 1085431.302 | 1088186.797 | 1086878.286 | [5 empty] | 1091698.211 | 6739.4088 | 0.6173%
| SRM 2 | Analyte | 1196692.63 | 1194686.245 | 1201822.402 | 1205384.995 | 1206498.142 | [5 empty] | 1201016.8828 | 5209.3571 | 0.4337%
|  | IS | 1123082.638 | 1120905.485 | 1126103.499 | 1130591.566 | 1131440.278 | [5 empty] | 1126424.6932 | 4589.4738 | 0.4074%
| Test Sample 2 | Analyte | 1108934.268 | 1108875.982 | 1113501.381 | 1111789.903 | 1104356.456 | [5 empty] | 1109491.598 | 3479.2186 | 0.3136%
|  | IS | 1068717.365 | 1066963.666 | 1072041.731 | 1069882.601 | 1062620.62 | [5 empty] | 1068045.1966 | 3549.5032 | 0.3323%
| SRM 3 | Analyte | 1208361.011 | 1206795.087 | 1198804.724 | 1214333.384 | 1227517.7 | [5 empty] | 1211162.3812 | 10690.9347 | 0.8827%
|  | IS | 1114122.937 | 1111590.823 | 1103448.798 | 1118955.065 | 1131138 | [5 empty] | 1115851.1246 | 10225.917 | 0.9164%
| Test Sample 3 | Analyte | 1243910.542 | 1241297.1 | 1246063.898 | 1251752.997 | 1251150.312 | [5 empty] | 1246834.9698 | 4544.8873 | 0.3645%
|  | IS | 1137017.309 | 1132851.943 | 1136376.099 | 1143101.725 | 1142512.438 | [5 empty] | 1138371.9028 | 4353.2563 | 0.3824%
| SRM 4 | Analyte | 1038993.798 | 1042956.128 | 1039237.13 | 1037880.466 | 1036580.308 | [5 empty] | 1039129.566 | 2383.8806 | 0.2294%
|  | IS | 1092479.249 | 1096261.569 | 1092818.391 | 1091711.035 | 1089798.888 | [5 empty] | 1092613.8264 | 2350.7921 | 0.2152%
| Test Sample 4 | Analyte | 1077577.542 | 1084556.562 | 1084312.952 | 1075816.606 | 1068783.825 | [5 empty] | 1078209.4974 | 6567.0767 | 0.6091%
|  | IS | 1188758.018 | 1196641.665 | 1196296.824 | 1187227.052 | 1178770.293 | [5 empty] | 1189538.7704 | 7383.0231 | 0.6207%
2 | SRM 1 | Analyte | 1126370.405 | 1132046.9 | 1122450.228 | 1121445.693 | 1119735.13 | [5 empty] | 1124409.6712 | 4915.4771 | 0.4372%
|  | IS | 1102142.816 | 1108018.851 | 1097852.015 | 1096643.76 | 1095655.728 | [5 empty] | 1100062.634 | 5089.8159 | 0.4627%
| Test Sample 1 | Analyte | 1132849.535 | 1135765.701 | 1133252.593 | 1128778.549 | 1133706.815 | [5 empty] | 1132870.6386 | 2548.1419 | 0.2249%
|  | IS | 1092214.668 | 1094429.173 | 1091930.964 | 1088288.644 | 1092676.072 | [5 empty] | 1091907.9042 | 2243.9117 | 0.2055%
| SRM 2 | Analyte | 1207295.28 | 1204485.099 | 1207093.713 | 1199909.698 | 1204279.195 | [5 empty] | 1204612.597 | 2983.1777 | 0.2476%
|  | IS | 1132187.502 | 1129408.535 | 1132447.608 | 1124854.795 | 1129828.912 | [5 empty] | 1129745.4704 | 3053.8607 | 0.2703%
| Test Sample 2 | Analyte | 1111136.638 | 1108478.888 | 1107617.711 | 1103457.976 | 1110336.708 | [5 empty] | 1108205.5842 | 3003.8658 | 0.2711%
|  | IS | 1069968.017 | 1066407.552 | 1065942.208 | 1061815.832 | 1068156.766 | [5 empty] | 1066458.075 | 3042.3813 | 0.2853%
| SRM 3 | Analyte | 1214994.385 | 1215094.076 | 1214792.601 | 1224695.274 | 1221031.496 | [5 empty] | 1218121.5664 | 4519.632 | 0.371%
|  | IS | 1119541.706 | 1118391.338 | 1118592.243 | 1127251.372 | 1123540.988 | [5 empty] | 1121463.5294 | 3846.7846 | 0.343%
| Test Sample 3 | Analyte | 1231031.771 | 1229925.437 | 1248030.882 | 1246312.676 | 1243300.688 | [5 empty] | 1239720.2908 | 8613.5685 | 0.6948%
|  | IS | 1122825.22 | 1123386.184 | 1139291.35 | 1138186.182 | 1135182.604 | [5 empty] | 1131774.308 | 8057.3155 | 0.7119%
| SRM 4 | Analyte | 1051854.309 | 1042755.187 | 1047176.247 | 1044414.235 | 1044916.134 | [5 empty] | 1046223.2224 | 3522.2712 | 0.3367%
|  | IS | 1107376.208 | 1098369.294 | 1102194.529 | 1100711.576 | 1100626.801 | [5 empty] | 1101855.6816 | 3375.0946 | 0.3063%
| Test Sample 4 | Analyte | 1075063.367 | 1073904.317 | 1075612.961 | 1080237.281 | 1075613.832 | [5 empty] | 1076086.3516 | 2423.0849 | 0.2252%
|  | IS | 1186723.377 | 1184810.323 | 1187331.414 | 1191662.632 | 1186742.546 | [5 empty] | 1187454.0584 | 2537.5488 | 0.2137%
3 | SRM 1 | Analyte | 1139182.618 | 1136973.116 | 1129080.134 | 1133510.14 | 1134817.517 | [5 empty] | 1134712.705 | 3816.8241 | 0.3364%
|  | IS | 1115025.262 | 1112309.551 | 1105262.933 | 1109591.882 | 1110396.982 | [5 empty] | 1110517.322 | 3606.0065 | 0.3247%
| Test Sample 1 | Analyte | 1128984.331 | 1132102.238 | 1135263.031 | 1132950.743 | 1135261.453 | [5 empty] | 1132912.3592 | 2604.3749 | 0.2299%
|  | IS | 1088690.608 | 1090603.132 | 1094892.39 | 1092880.124 | 1094830.88 | [5 empty] | 1092379.4268 | 2708.2241 | 0.2479%
| SRM 2 | Analyte | 1197496.24 | 1204177.83 | 1195385.258 | 1196491.057 | 1197195.137 | [5 empty] | 1198149.1044 | 3466.6189 | 0.2893%
|  | IS | 1124042.412 | 1130030.631 | 1121726.474 | 1123078.377 | 1123544.404 | [5 empty] | 1124484.4596 | 3218.0227 | 0.2862%
| Test Sample 2 | Analyte | 1107071.474 | 1108526.663 | 1110790.441 | 1106968.595 | 1102350.884 | [5 empty] | 1107141.6114 | 3091.0455 | 0.2792%
|  | IS | 1066707.794 | 1067254.09 | 1068861.644 | 1064534.051 | 1059803.464 | [5 empty] | 1065432.2086 | 3507.0233 | 0.3292%
| SRM 3 | Analyte | 1214189.494 | 1211017.274 | 1212082.72 | 1216504.246 | 1213029.33 | [5 empty] | 1213364.6128 | 2109.7031 | 0.1739%
|  | IS | 1119554.659 | 1117138.834 | 1117142.475 | 1120763.521 | 1118850.328 | [5 empty] | 1118689.9634 | 1571.1285 | 0.1404%
| Test Sample 3 | Analyte | 1233749.107 | 1222378.737 | 1229222.363 | 1231134.745 | 1237671.444 | [5 empty] | 1230831.2792 | 5688.3158 | 0.4622%
|  | IS | 1128012.718 | 1118345.075 | 1125148.645 | 1127306.087 | 1132848.085 | [5 empty] | 1126332.122 | 5277.886 | 0.4686%
| SRM 4 | Analyte | 1040693.682 | 1050090.964 | 1047076.245 | 1042912.073 | 1045367.586 | [5 empty] | 1045228.11 | 3640.3115 | 0.3483%
|  | IS | 1096859.491 | 1106887.158 | 1103989.458 | 1099119.616 | 1101536.328 | [5 empty] | 1101678.4102 | 3945.3008 | 0.3581%
| Test Sample 4 | Analyte | 1067269.462 | 1071799.706 | 1059994.899 | 1064719.017 | 1067275.486 | [5 empty] | 1066211.714 | 4311.4101 | 0.4044%
|  | IS | 1179873.516 | 1184405.543 | 1170409.358 | 1175143.051 | 1177658.307 | [5 empty] | 1177497.955 | 5224.4032 | 0.4437%
4 | SRM 1 | Analyte | 1140342.815 | 1134213.874 | 1139337.124 | 1141799.138 | 1136522.427 | [5 empty] | 1138443.0756 | 3051.4965 | 0.268%
|  | IS | 1118347.55 | 1112404.399 | 1117038.797 | 1119556.019 | 1113613.201 | [5 empty] | 1116191.9932 | 3069.0475 | 0.275%
| Test Sample 1 | Analyte | 1116463.299 | 1118475.078 | 1123401.885 | 1112500.001 | 1117418.996 | [5 empty] | 1117651.8518 | 3929.0275 | 0.3515%
|  | IS | 1077175.938 | 1079947.016 | 1084519.444 | 1074154.902 | 1078077.849 | [5 empty] | 1078775.0298 | 3832.7726 | 0.3553%
| SRM 2 | Analyte | 1207101.011 | 1201211.668 | 1202425.568 | 1203733.724 | 1203432.198 | [5 empty] | 1203580.8338 | 2201.403 | 0.1829%
|  | IS | 1133854.775 | 1129339.09 | 1130669.555 | 1130929.134 | 1131131.204 | [5 empty] | 1131184.7516 | 1648.4179 | 0.1457%
| Test Sample 2 | Analyte | 1106567.61 | 1108476.508 | 1100038.429 | 1098228.536 | 1101043.44 | [5 empty] | 1102870.9046 | 4415.9527 | 0.4004%
|  | IS | 1067109.359 | 1069821.552 | 1062071.518 | 1059512.967 | 1062575.092 | [5 empty] | 1064218.0976 | 4159.4356 | 0.3908%
| SRM 3 | Analyte | 1215341.11 | 1217456.218 | 1224248.944 | 1218053.684 | 1216551.112 | [5 empty] | 1218330.2136 | 3463.0784 | 0.2842%
|  | IS | 1123379.185 | 1123125.223 | 1130246.066 | 1124587.731 | 1122574.466 | [5 empty] | 1124782.5342 | 3141.6462 | 0.2793%
| Test Sample 3 | Analyte | 1240087.037 | 1237469.358 | 1235861.642 | 1238677.062 | 1232891.698 | [5 empty] | 1236997.3594 | 2772.03 | 0.2241%
|  | IS | 1135925.686 | 1132844.295 | 1132041.213 | 1135359.124 | 1128633.89 | [5 empty] | 1132960.8416 | 2920.782 | 0.2578%
| SRM 4 | Analyte | 1037383.854 | 1034669.461 | 1033107.925 | 1041143.906 | 1040640.32 | [5 empty] | 1037389.0932 | 3549.4339 | 0.3422%
|  | IS | 1094632.003 | 1092109.575 | 1089450.01 | 1097402.695 | 1097704.785 | [5 empty] | 1094259.8136 | 3522.8466 | 0.3219%
| Test Sample 4 | Analyte | 1063956.922 | 1066972.377 | 1060236.352 | 1057725.386 | 1060639.72 | [5 empty] | 1061906.1514 | 3596.6366 | 0.3387%
|  | IS | 1177214.457 | 1180336.206 | 1173888.089 | 1171525.602 | 1173524.667 | [5 empty] | 1175297.8042 | 3479.2195 | 0.296%
5 | SRM 1 | Analyte | 1123252.226 | 1125413.867 | 1125714.162 | 1120485.701 | 1122045.297 | [5 empty] | 1123382.2506 | 2222.5901 | 0.1978%
|  | IS | 1103387.893 | 1105458.894 | 1105258.174 | 1099521.334 | 1099621.871 | [5 empty] | 1102649.6332 | 2923.8271 | 0.2652%
| Test Sample 1 | Analyte | 1128926.161 | 1123803.862 | 1129227.451 | 1127624.563 | 1126920.033 | [5 empty] | 1127300.414 | 2169.802 | 0.1925%
|  | IS | 1092771.06 | 1087134.682 | 1092167.883 | 1091719.844 | 1090053.879 | [5 empty] | 1090769.4696 | 2268.7489 | 0.208%
| SRM 2 | Analyte | 1196393.54 | 1198701.877 | 1188589.214 | 1189956.601 | 1184022.318 | [5 empty] | 1191532.71 | 5970.3203 | 0.5011%
|  | IS | 1124993.858 | 1128210.469 | 1118862.601 | 1119048.385 | 1114960.66 | [5 empty] | 1121215.1946 | 5304.7414 | 0.4731%
| Test Sample 2 | Analyte | 1106270.505 | 1103052.919 | 1102351.327 | 1106868.885 | 1101646.342 | [5 empty] | 1104037.9956 | 2373.4609 | 0.215%
|  | IS | 1069921.317 | 1066096.521 | 1065275.82 | 1068919.332 | 1063380.377 | [5 empty] | 1066718.6734 | 2679.2231 | 0.2512%
| SRM 3 | Analyte | 1215945.612 | 1212526.617 | 1208258.549 | 1212125.136 | 1206698.473 | [5 empty] | 1211110.8774 | 3674.6696 | 0.3034%
|  | IS | 1123944.006 | 1120659.372 | 1116934.245 | 1120558.408 | 1115121.576 | [5 empty] | 1119443.5214 | 3462.3441 | 0.3093%
| Test Sample 3 | Analyte | 1235961.646 | 1239175.928 | 1238174.468 | 1234856.368 | 1231491.247 | [5 empty] | 1235931.9314 | 3017.8586 | 0.2442%
|  | IS | 1132041.021 | 1135162.795 | 1134272.301 | 1132640.059 | 1128411.359 | [5 empty] | 1132505.507 | 2606.6569 | 0.2302%
| SRM 4 | Analyte | 1042308.547 | 1036377.751 | 1038086.948 | 1032761.75 | 1032859.849 | [5 empty] | 1036478.969 | 3984.2053 | 0.3844%
|  | IS | 1099013.191 | 1094885.522 | 1095849.484 | 1091160.414 | 1090154.222 | [5 empty] | 1094212.5666 | 3604.3085 | 0.3294%
| Test Sample 4 | Analyte | 1073005.151 | 1072802.264 | 1071648.96 | 1065712.88 | 1075209.544 | [5 empty] | 1071675.7598 | 3573.5339 | 0.3335%
|  | IS | 1186736.112 | 1186635.355 | 1185830.245 | 1178481.344 | 1190307.953 | [5 empty] | 1185598.2018 | 4337.5048 | 0.3658%

## Ratio Calculations
HP-ICP Comparison of Example (lot 1) to NIST SRM 3113 (lot 000630), 10/26/2001.
Repeat # | ID | Ratio 1 | Ratio 2 | Ratio 3 | Ratio 4 | Ratio 5 | Ratio 6 | Ratio 7 | Ratio 8 | Ratio 9 | Ratio 10 | Average | sd | rsd | n | sd mean
1 | SRM 1 | Ratio 1: 1.0218; Ratio 2: 1.0218; Ratio 3: 1.0216; Ratio 4: 1.0223; Ratio 5: 1.022 | 1.0219 | 0.00026592 | 0.02602% | 5 | 0.00011892
| Test Sample 1 | Ratio 1: 1.0379; Ratio 2: 1.0381; Ratio 3: 1.0377; Ratio 4: 1.0376; Ratio 5: 1.0381 | 1.0379 | 0.00023335 | 0.02248% | 5 | 0.00010436
| SRM 2 | Ratio 1: 1.0655; Ratio 2: 1.0658; Ratio 3: 1.0672; Ratio 4: 1.0662; Ratio 5: 1.0663 | 1.0662 | 0.00064707 | 0.06069% | 5 | 0.00028938
| Test Sample 2 | Ratio 1: 1.0376; Ratio 2: 1.0393; Ratio 3: 1.0387; Ratio 4: 1.0392; Ratio 5: 1.0393 | 1.0388 | 0.00070326 | 0.0677% | 5 | 0.00031451
| SRM 3 | Ratio 1: 1.0846; Ratio 2: 1.0856; Ratio 3: 1.0864; Ratio 4: 1.0852; Ratio 5: 1.0852 | 1.0854 | 0.00067452 | 0.06214% | 5 | 0.00030166
| Test Sample 3 | Ratio 1: 1.094; Ratio 2: 1.0957; Ratio 3: 1.0965; Ratio 4: 1.095; Ratio 5: 1.0951 | 1.0953 | 0.00092827 | 0.08475% | 5 | 0.00041513
| SRM 4 | Ratio 1: 0.95104; Ratio 2: 0.95138; Ratio 3: 0.95097; Ratio 4: 0.95069; Ratio 5: 0.95117 | 0.95105 | 0.00025218 | 0.02652% | 5 | 0.00011278
| Test Sample 4 | Ratio 1: 0.90647; Ratio 2: 0.90633; Ratio 3: 0.90639; Ratio 4: 0.90616; Ratio 5: 0.90669 | 0.90641 | 0.00019606 | 0.02163% | 5 | 8.7679e-05
2 | SRM 1 | Ratio 1: 1.022; Ratio 2: 1.0217; Ratio 3: 1.0224; Ratio 4: 1.0226; Ratio 5: 1.022 | 1.0221 | 0.00037252 | 0.03645% | 5 | 0.00016659
| Test Sample 1 | Ratio 1: 1.0372; Ratio 2: 1.0378; Ratio 3: 1.0378; Ratio 4: 1.0372; Ratio 5: 1.0376 | 1.0375 | 0.00030263 | 0.02917% | 5 | 0.00013534
| SRM 2 | Ratio 1: 1.0663; Ratio 2: 1.0665; Ratio 3: 1.0659; Ratio 4: 1.0667; Ratio 5: 1.0659 | 1.0663 | 0.00036007 | 0.03377% | 5 | 0.00016103
| Test Sample 2 | Ratio 1: 1.0385; Ratio 2: 1.0395; Ratio 3: 1.0391; Ratio 4: 1.0392; Ratio 5: 1.0395 | 1.0391 | 0.00040817 | 0.03928% | 5 | 0.00018254
| SRM 3 | Ratio 1: 1.0853; Ratio 2: 1.0865; Ratio 3: 1.086; Ratio 4: 1.0864; Ratio 5: 1.0868 | 1.0862 | 0.00058682 | 0.05403% | 5 | 0.00026243
| Test Sample 3 | Ratio 1: 1.0964; Ratio 2: 1.0948; Ratio 3: 1.0954; Ratio 4: 1.095; Ratio 5: 1.0952 | 1.0954 | 0.00060048 | 0.05482% | 5 | 0.00026854
| SRM 4 | Ratio 1: 0.94986; Ratio 2: 0.94937; Ratio 3: 0.95008; Ratio 4: 0.94885; Ratio 5: 0.94938 | 0.94951 | 0.00047945 | 0.05049% | 5 | 0.00021442
| Test Sample 4 | Ratio 1: 0.90591; Ratio 2: 0.90639; Ratio 3: 0.90591; Ratio 4: 0.9065; Ratio 5: 0.90636 | 0.90621 | 0.00028247 | 0.03117% | 5 | 0.00012632
3 | SRM 1 | Ratio 1: 1.0217; Ratio 2: 1.0222; Ratio 3: 1.0215; Ratio 4: 1.0216; Ratio 5: 1.022 | 1.0218 | 0.00028125 | 0.02753% | 5 | 0.00012578
3 | Test Sample 1 | Ratio 1: 1.037; Ratio 2: 1.0381; Ratio 3: 1.0369; Ratio 4: 1.0367; Ratio 5: 1.0369 | 1.0371 | 0.00054398 | 0.05245% | 5 | 0.00024328
3 | SRM 2 | Ratio 1: 1.0653; Ratio 2: 1.0656; Ratio 3: 1.0657; Ratio 4: 1.0654; Ratio 5: 1.0656 | 1.0655 | 0.00014462 | 0.01357% | 5 | 6.4674e-05
3 | Test Sample 2 | Ratio 1: 1.0378; Ratio 2: 1.0387; Ratio 3: 1.0392; Ratio 4: 1.0399; Ratio 5: 1.0401 | 1.0391 | 0.00092902 | 0.0894% | 5 | 0.00041547
3 | SRM 3 | Ratio 1: 1.0845; Ratio 2: 1.084; Ratio 3: 1.085; Ratio 4: 1.0854; Ratio 5: 1.0842 | 1.0846 | 0.00057642 | 0.05314% | 5 | 0.00025778
3 | Test Sample 3 | Ratio 1: 1.0937; Ratio 2: 1.093; Ratio 3: 1.0925; Ratio 4: 1.0921; Ratio 5: 1.0925 | 1.0928 | 0.00062751 | 0.05742% | 5 | 0.00028063
3 | SRM 4 | Ratio 1: 0.94879; Ratio 2: 0.94869; Ratio 3: 0.94845; Ratio 4: 0.94886; Ratio 5: 0.94901 | 0.94876 | 0.00020973 | 0.02211% | 5 | 9.3796e-05
3 | Test Sample 4 | Ratio 1: 0.90456; Ratio 2: 0.90493; Ratio 3: 0.90566; Ratio 4: 0.90603; Ratio 5: 0.90627 | 0.90549 | 0.00072628 | 0.08021% | 5 | 0.0003248
4 | SRM 1 | Ratio 1: 1.0197; Ratio 2: 1.0196; Ratio 3: 1.02; Ratio 4: 1.0199; Ratio 5: 1.0206 | 1.0199 | 0.00038435 | 0.03768% | 5 | 0.00017189
4 | Test Sample 1 | Ratio 1: 1.0365; Ratio 2: 1.0357; Ratio 3: 1.0359; Ratio 4: 1.0357; Ratio 5: 1.0365 | 1.036 | 0.00041115 | 0.03969% | 5 | 0.00018387
4 | SRM 2 | Ratio 1: 1.0646; Ratio 2: 1.0636; Ratio 3: 1.0635; Ratio 4: 1.0644; Ratio 5: 1.0639 | 1.064 | 0.00048056 | 0.04517% | 5 | 0.00021491
4 | Test Sample 2 | Ratio 1: 1.037; Ratio 2: 1.0361; Ratio 3: 1.0357; Ratio 4: 1.0365; Ratio 5: 1.0362 | 1.0363 | 0.00046269 | 0.04465% | 5 | 0.00020692
4 | SRM 3 | Ratio 1: 1.0819; Ratio 2: 1.084; Ratio 3: 1.0832; Ratio 4: 1.0831; Ratio 5: 1.0837 | 1.0832 | 0.00081914 | 0.07562% | 5 | 0.00036633
4 | Test Sample 3 | Ratio 1: 1.0917; Ratio 2: 1.0924; Ratio 3: 1.0917; Ratio 4: 1.091; Ratio 5: 1.0924 | 1.0918 | 0.00056885 | 0.0521% | 5 | 0.0002544
4 | SRM 4 | Ratio 1: 0.9477; Ratio 2: 0.9474; Ratio 3: 0.94828; Ratio 4: 0.94873; Ratio 5: 0.94801 | 0.94803 | 0.00051495 | 0.05432% | 5 | 0.00023029
4 | Test Sample 4 | Ratio 1: 0.90379; Ratio 2: 0.90396; Ratio 3: 0.90318; Ratio 4: 0.90286; Ratio 5: 0.90381 | 0.90352 | 0.00047262 | 0.05231% | 5 | 0.00021136
5 | SRM 1 | Ratio 1: 1.018; Ratio 2: 1.0181; Ratio 3: 1.0185; Ratio 4: 1.0191; Ratio 5: 1.0204 | 1.0188 | 0.0009857 | 0.09675% | 5 | 0.00044082
5 | Test Sample 1 | Ratio 1: 1.0331; Ratio 2: 1.0337; Ratio 3: 1.0339; Ratio 4: 1.0329; Ratio 5: 1.0338 | 1.0335 | 0.00047115 | 0.04559% | 5 | 0.00021071
5 | SRM 2 | Ratio 1: 1.0635; Ratio 2: 1.0625; Ratio 3: 1.0623; Ratio 4: 1.0634; Ratio 5: 1.0619 | 1.0627 | 0.00067046 | 0.06309% | 5 | 0.00029984
5 | Test Sample 2 | Ratio 1: 1.034; Ratio 2: 1.0347; Ratio 3: 1.0348; Ratio 4: 1.0355; Ratio 5: 1.036 | 1.035 | 0.00077884 | 0.07525% | 5 | 0.00034831
5 | SRM 3 | Ratio 1: 1.0819; Ratio 2: 1.082; Ratio 3: 1.0818; Ratio 4: 1.0817; Ratio 5: 1.0821 | 1.0819 | 0.00016528 | 0.01528% | 5 | 7.3914e-05
5 | Test Sample 3 | Ratio 1: 1.0918; Ratio 2: 1.0916; Ratio 3: 1.0916; Ratio 4: 1.0902; Ratio 5: 1.0913 | 1.0913 | 0.00062425 | 0.0572% | 5 | 0.00027917
5 | SRM 4 | Ratio 1: 0.9484; Ratio 2: 0.94656; Ratio 3: 0.94729; Ratio 4: 0.94648; Ratio 5: 0.94744 | 0.94724 | 0.00078022 | 0.08237% | 5 | 0.00034892
5 | Test Sample 4 | Ratio 1: 0.90416; Ratio 2: 0.90407; Ratio 3: 0.90371; Ratio 4: 0.90431; Ratio 5: 0.9033 | 0.90391 | 0.00040556 | 0.04487% | 5 | 0.00018137

## Results (no drift correction)
HP-ICP Comparison of Example (lot 1) to NIST SRM 3113 (lot 000630), 10/26/2001.
SRM 3113 Measurements | Average Signal Ratio | sd | rsd | Mass Ratio | Slope of Calibration Curve | Apparent Mass Fraction | sd
|  |  |  | (mg analyte/g IS) | (signal ratio/mass ratio) | mg/g | mg/g
SRM 1 | 1.0209 | 0.0014745 | 0.1444% | 1.2741 | 0.80127 | 9.9946 | 0.014435
SRM 2 | 1.0649 | 0.0015465 | 0.1452% | 1.3293 | 0.80111 | 9.9927 | 0.014511
SRM 3 | 1.0843 | 0.0017334 | 0.1599% | 1.3527 | 0.80156 | 9.9982 | 0.015984
SRM 4 | 0.94892 | 0.001461 | 0.154% | 1.1838 | 0.80158 | 9.9985 | 0.015394
|  |  |  | Average slope | 0.80138
|  |  |  | sd of slope | 0.0002275
Example Measurements | Average Signal Ratio | sd | rsd | Observed Mass Fraction | sd | Observed /Nominal
|  |  |  | mg/g | mg/g
Test Sample 1 | 1.0364 | 0.0017716 | 0.1709% | 9.9508 | 0.01701 | 99.99%
Test Sample 2 | 1.0377 | 0.001916 | 0.1846% | 9.9515 | 0.018375 | 100%
Test Sample 3 | 1.0933 | 0.0019091 | 0.1746% | 9.9445 | 0.017365 | 99.92%
Test Sample 4 | 0.90511 | 0.0013242 | 0.1463% | 9.9515 | 0.01456 | 100%
|  |  | Average mass fraction | 9.9496
|  |  | sd of mass fraction | 0.0034059
Uncertainty Budget
Source | (mg/g) | relative | df | coverage factor, k
SRM 3113 Measurements | 0.0014123 | 0.01419% | 3
Example Measurements | 0.0017029 | 0.01712% | 3
Certified Value of SRM 3113 | 0.0099567 | 0.09961% | 8
Combined Uncertainty | 0.010154 | 0.1021% | 9
Coverage Factor, k |  |  |  | 2.2622
Expanded Uncertainty | 0.022971 | 0.2309%
Determined Traceable Value >> 9.9496 mg/g ± 0.0230 mg/g
Note on G3: The apparent mass fractions of the different calibrants, treating them as samples.

## Results (drift corrected)
HP-ICP Comparison of Example (lot 1) to NIST SRM 3113 (lot 000630), 10/26/2001.
SRM 3113 Measurements | Average Signal Ratio | sd | rsd | Mass Ratio | Slope of Calibration Curve | Apparent Mass Fraction | sd
|  |  |  | (mg analyte/g IS) | (signal ratio/mass ratio) | mg/g | mg/g
SRM 1 | 1.0207 | 0.00023974 | 0.02349% | 1.2741 | 0.80108 | 9.9925 | 0.0023471
SRM 2 | 1.0649 | 0.00017127 | 0.01608% | 1.3293 | 0.80106 | 9.9922 | 0.0016071
SRM 3 | 1.0843 | 0.00029476 | 0.02718% | 1.3527 | 0.80161 | 9.9991 | 0.0027181
SRM 4 | 0.94905 | 0.00053522 | 0.0564% | 1.1838 | 0.8017 | 10.0002 | 0.0056396
|  |  |  | Average slope | 0.80136
|  |  |  | sd of slope | 0.00033994
Example Measurements | Average Signal Ratio | sd | rsd | Observed Mass Fraction | sd | Observed /Nominal
|  |  |  | mg/g | mg/g
Test Sample 1 | 1.0362 | 0.0004128 | 0.03984% | 9.9496 | 0.0039635 | 99.98%
Test Sample 2 | 1.0377 | 0.000534 | 0.05146% | 9.9518 | 0.0051212 | 100%
Test Sample 3 | 1.0934 | 0.00053285 | 0.04873% | 9.9459 | 0.0048468 | 99.94%
Test Sample 4 | 0.90523 | 0.00025012 | 0.02763% | 9.9531 | 0.0027502 | 100%
|  |  | Average mass fraction | 9.9501
|  |  | sd of mass fraction | 0.0031474
Uncertainty Budget
Source | (mg/g) | relative | df | coverage factor, k
SRM 3113 Measurements | 0.0021104 | 0.02121% | 3
Example Measurements | 0.0015737 | 0.01582% | 3
Certified Value of SRM 3113 | 0.0099567 | 0.09961% | 8
Combined Uncertainty | 0.010255 | 0.1031% | 9
Coverage Factor, k |  |  |  | 2.2622
Expanded Uncertainty | 0.023198 | 0.2331%
Determined Traceable Value >> 9.9501 mg/g ± 0.0232 mg/g
Note on G3: The apparent mass fractions of the different calibrants, treating them as samples.

## Correlation Diagnostic
Point | index 1 | index 2 | Analyte | IS | Avg Anal | Avg IS | Ratio | NormAnalyte | NormIS
1 | 1 | 1 | 1137374.554 | 1113068.125 | 569210.2824 | 556994.2492 | 1.0218 | 199.8162 | 199.8348
2 | 1 | 2 | 1137878.917 | 1113556.491 | 569210.2824 | 556994.2492 | 1.0218 | 199.9048 | 199.9224
3 | 1 | 3 | 1145514.296 | 1121265.873 | 569210.2824 | 556994.2492 | 1.0216 | 201.2462 | 201.3065
4 | 1 | 4 | 1137828.692 | 1112967.87 | 569210.2824 | 556994.2492 | 1.0223 | 199.896 | 199.8168
5 | 1 | 5 | 1133506.365 | 1109084.133 | 569210.2824 | 556994.2492 | 1.022 | 199.1367 | 199.1195
6 | 1 | 6 | 0 | 0 | 569210.2824 | 556994.2492 |  | 0 | 0
7 | 1 | 7 | 0 | 0 | 569210.2824 | 556994.2492 |  | 0 | 0
8 | 1 | 8 | 0 | 0 | 569210.2824 | 556994.2492 |  | 0 | 0
9 | 1 | 9 | 0 | 0 | 569210.2824 | 556994.2492 |  | 0 | 0
10 | 1 | 10 | 0 | 0 | 569210.2824 | 556994.2492 |  | 0 | 0
11 | 3 | 1 | 1141089.778 | 1099376.973 | 566532.2344 | 545849.1055 | 1.0379 | 201.4166 | 201.4068
12 | 3 | 2 | 1140492.796 | 1098617.697 | 566532.2344 | 545849.1055 | 1.0381 | 201.3112 | 201.2677
13 | 3 | 3 | 1126372.02 | 1085431.302 | 566532.2344 | 545849.1055 | 1.0377 | 198.8187 | 198.8519
14 | 3 | 4 | 1129086.325 | 1088186.797 | 566532.2344 | 545849.1055 | 1.0376 | 199.2978 | 199.3567
15 | 3 | 5 | 1128281.425 | 1086878.286 | 566532.2344 | 545849.1055 | 1.0381 | 199.1557 | 199.117
16 | 3 | 6 | 0 | 0 | 566532.2344 | 545849.1055 |  | 0 | 0
17 | 3 | 7 | 0 | 0 | 566532.2344 | 545849.1055 |  | 0 | 0
18 | 3 | 8 | 0 | 0 | 566532.2344 | 545849.1055 |  | 0 | 0
19 | 3 | 9 | 0 | 0 | 566532.2344 | 545849.1055 |  | 0 | 0
20 | 3 | 10 | 0 | 0 | 566532.2344 | 545849.1055 |  | 0 | 0
21 | 5 | 1 | 1196692.63 | 1123082.638 | 600508.4414 | 563212.3466 | 1.0655 | 199.2799 | 199.4066
22 | 5 | 2 | 1194686.245 | 1120905.485 | 600508.4414 | 563212.3466 | 1.0658 | 198.9458 | 199.02
23 | 5 | 3 | 1201822.402 | 1126103.499 | 600508.4414 | 563212.3466 | 1.0672 | 200.1341 | 199.943
24 | 5 | 4 | 1205384.995 | 1130591.566 | 600508.4414 | 563212.3466 | 1.0662 | 200.7274 | 200.7398
25 | 5 | 5 | 1206498.142 | 1131440.278 | 600508.4414 | 563212.3466 | 1.0663 | 200.9128 | 200.8905
26 | 5 | 6 | 0 | 0 | 600508.4414 | 563212.3466 |  | 0 | 0
27 | 5 | 7 | 0 | 0 | 600508.4414 | 563212.3466 |  | 0 | 0
28 | 5 | 8 | 0 | 0 | 600508.4414 | 563212.3466 |  | 0 | 0
29 | 5 | 9 | 0 | 0 | 600508.4414 | 563212.3466 |  | 0 | 0
30 | 5 | 10 | 0 | 0 | 600508.4414 | 563212.3466 |  | 0 | 0
31 | 7 | 1 | 1108934.268 | 1068717.365 | 554745.799 | 534022.5983 | 1.0376 | 199.8995 | 200.1259
32 | 7 | 2 | 1108875.982 | 1066963.666 | 554745.799 | 534022.5983 | 1.0393 | 199.889 | 199.7975
33 | 7 | 3 | 1113501.381 | 1072041.731 | 554745.799 | 534022.5983 | 1.0387 | 200.7228 | 200.7484
34 | 7 | 4 | 1111789.903 | 1069882.601 | 554745.799 | 534022.5983 | 1.0392 | 200.4143 | 200.3441
35 | 7 | 5 | 1104356.456 | 1062620.62 | 554745.799 | 534022.5983 | 1.0393 | 199.0743 | 198.9842
36 | 7 | 6 | 0 | 0 | 554745.799 | 534022.5983 |  | 0 | 0
37 | 7 | 7 | 0 | 0 | 554745.799 | 534022.5983 |  | 0 | 0
38 | 7 | 8 | 0 | 0 | 554745.799 | 534022.5983 |  | 0 | 0
39 | 7 | 9 | 0 | 0 | 554745.799 | 534022.5983 |  | 0 | 0
40 | 7 | 10 | 0 | 0 | 554745.799 | 534022.5983 |  | 0 | 0
41 | 9 | 1 | 1208361.011 | 1114122.937 | 605581.1906 | 557925.5623 | 1.0846 | 199.5374 | 199.6902
42 | 9 | 2 | 1206795.087 | 1111590.823 | 605581.1906 | 557925.5623 | 1.0856 | 199.2788 | 199.2364
43 | 9 | 3 | 1198804.724 | 1103448.798 | 605581.1906 | 557925.5623 | 1.0864 | 197.9594 | 197.7771
44 | 9 | 4 | 1214333.384 | 1118955.065 | 605581.1906 | 557925.5623 | 1.0852 | 200.5236 | 200.5563
45 | 9 | 5 | 1227517.7 | 1131138 | 605581.1906 | 557925.5623 | 1.0852 | 202.7008 | 202.7399
46 | 9 | 6 | 0 | 0 | 605581.1906 | 557925.5623 |  | 0 | 0
47 | 9 | 7 | 0 | 0 | 605581.1906 | 557925.5623 |  | 0 | 0
48 | 9 | 8 | 0 | 0 | 605581.1906 | 557925.5623 |  | 0 | 0
49 | 9 | 9 | 0 | 0 | 605581.1906 | 557925.5623 |  | 0 | 0
50 | 9 | 10 | 0 | 0 | 605581.1906 | 557925.5623 |  | 0 | 0
51 | 11 | 1 | 1243910.542 | 1137017.309 | 623417.4849 | 569185.9514 | 1.094 | 199.5309 | 199.762
52 | 11 | 2 | 1241297.1 | 1132851.943 | 623417.4849 | 569185.9514 | 1.0957 | 199.1117 | 199.0302
53 | 11 | 3 | 1246063.898 | 1136376.099 | 623417.4849 | 569185.9514 | 1.0965 | 199.8763 | 199.6494
54 | 11 | 4 | 1251752.997 | 1143101.725 | 623417.4849 | 569185.9514 | 1.095 | 200.7889 | 200.831
55 | 11 | 5 | 1251150.312 | 1142512.438 | 623417.4849 | 569185.9514 | 1.0951 | 200.6922 | 200.7274
56 | 11 | 6 | 0 | 0 | 623417.4849 | 569185.9514 |  | 0 | 0
57 | 11 | 7 | 0 | 0 | 623417.4849 | 569185.9514 |  | 0 | 0
58 | 11 | 8 | 0 | 0 | 623417.4849 | 569185.9514 |  | 0 | 0
59 | 11 | 9 | 0 | 0 | 623417.4849 | 569185.9514 |  | 0 | 0
60 | 11 | 10 | 0 | 0 | 623417.4849 | 569185.9514 |  | 0 | 0
61 | 13 | 1 | 1038993.798 | 1092479.249 | 519564.783 | 546306.9132 | 0.95104 | 199.9739 | 199.9754
62 | 13 | 2 | 1042956.128 | 1096261.569 | 519564.783 | 546306.9132 | 0.95138 | 200.7365 | 200.6677
63 | 13 | 3 | 1039237.13 | 1092818.391 | 519564.783 | 546306.9132 | 0.95097 | 200.0207 | 200.0374
64 | 13 | 4 | 1037880.466 | 1091711.035 | 519564.783 | 546306.9132 | 0.95069 | 199.7596 | 199.8347
65 | 13 | 5 | 1036580.308 | 1089798.888 | 519564.783 | 546306.9132 | 0.95117 | 199.5093 | 199.4847
66 | 13 | 6 | 0 | 0 | 519564.783 | 546306.9132 |  | 0 | 0
67 | 13 | 7 | 0 | 0 | 519564.783 | 546306.9132 |  | 0 | 0
68 | 13 | 8 | 0 | 0 | 519564.783 | 546306.9132 |  | 0 | 0
69 | 13 | 9 | 0 | 0 | 519564.783 | 546306.9132 |  | 0 | 0
70 | 13 | 10 | 0 | 0 | 519564.783 | 546306.9132 |  | 0 | 0
71 | 15 | 1 | 1077577.542 | 1188758.018 | 539104.7487 | 594769.3852 | 0.90647 | 199.8828 | 199.8687
72 | 15 | 2 | 1084556.562 | 1196641.665 | 539104.7487 | 594769.3852 | 0.90633 | 201.1773 | 201.1942
73 | 15 | 3 | 1084312.952 | 1196296.824 | 539104.7487 | 594769.3852 | 0.90639 | 201.1321 | 201.1362
74 | 15 | 4 | 1075816.606 | 1187227.052 | 539104.7487 | 594769.3852 | 0.90616 | 199.5561 | 199.6113
75 | 15 | 5 | 1068783.825 | 1178770.293 | 539104.7487 | 594769.3852 | 0.90669 | 198.2516 | 198.1895
76 | 15 | 6 | 0 | 0 | 539104.7487 | 594769.3852 |  | 0 | 0
77 | 15 | 7 | 0 | 0 | 539104.7487 | 594769.3852 |  | 0 | 0
78 | 15 | 8 | 0 | 0 | 539104.7487 | 594769.3852 |  | 0 | 0
79 | 15 | 9 | 0 | 0 | 539104.7487 | 594769.3852 |  | 0 | 0
80 | 15 | 10 | 0 | 0 | 539104.7487 | 594769.3852 |  | 0 | 0
81 | 17 | 1 | 1126370.405 | 1102142.816 | 562204.8356 | 550031.317 | 1.022 | 200.3488 | 200.3782
82 | 17 | 2 | 1132046.9 | 1108018.851 | 562204.8356 | 550031.317 | 1.0217 | 201.3584 | 201.4465
83 | 17 | 3 | 1122450.228 | 1097852.015 | 562204.8356 | 550031.317 | 1.0224 | 199.6515 | 199.5981
84 | 17 | 4 | 1121445.693 | 1096643.76 | 562204.8356 | 550031.317 | 1.0226 | 199.4728 | 199.3784
85 | 17 | 5 | 1119735.13 | 1095655.728 | 562204.8356 | 550031.317 | 1.022 | 199.1685 | 199.1988
86 | 17 | 6 | 0 | 0 | 562204.8356 | 550031.317 |  | 0 | 0
87 | 17 | 7 | 0 | 0 | 562204.8356 | 550031.317 |  | 0 | 0
88 | 17 | 8 | 0 | 0 | 562204.8356 | 550031.317 |  | 0 | 0
89 | 17 | 9 | 0 | 0 | 562204.8356 | 550031.317 |  | 0 | 0
90 | 17 | 10 | 0 | 0 | 562204.8356 | 550031.317 |  | 0 | 0
91 | 19 | 1 | 1132849.535 | 1092214.668 | 566435.3193 | 545953.9521 | 1.0372 | 199.9963 | 200.0562
92 | 19 | 2 | 1135765.701 | 1094429.173 | 566435.3193 | 545953.9521 | 1.0378 | 200.5111 | 200.4618
93 | 19 | 3 | 1133252.593 | 1091930.964 | 566435.3193 | 545953.9521 | 1.0378 | 200.0674 | 200.0042
94 | 19 | 4 | 1128778.549 | 1088288.644 | 566435.3193 | 545953.9521 | 1.0372 | 199.2776 | 199.3371
95 | 19 | 5 | 1133706.815 | 1092676.072 | 566435.3193 | 545953.9521 | 1.0376 | 200.1476 | 200.1407
96 | 19 | 6 | 0 | 0 | 566435.3193 | 545953.9521 |  | 0 | 0
97 | 19 | 7 | 0 | 0 | 566435.3193 | 545953.9521 |  | 0 | 0
98 | 19 | 8 | 0 | 0 | 566435.3193 | 545953.9521 |  | 0 | 0
99 | 19 | 9 | 0 | 0 | 566435.3193 | 545953.9521 |  | 0 | 0
100 | 19 | 10 | 0 | 0 | 566435.3193 | 545953.9521 |  | 0 | 0
101 | 21 | 1 | 1207295.28 | 1132187.502 | 602306.2985 | 564872.7352 | 1.0663 | 200.4454 | 200.4323
102 | 21 | 2 | 1204485.099 | 1129408.535 | 602306.2985 | 564872.7352 | 1.0665 | 199.9788 | 199.9404
103 | 21 | 3 | 1207093.713 | 1132447.608 | 602306.2985 | 564872.7352 | 1.0659 | 200.4119 | 200.4784
104 | 21 | 4 | 1199909.698 | 1124854.795 | 602306.2985 | 564872.7352 | 1.0667 | 199.2192 | 199.1342
105 | 21 | 5 | 1204279.195 | 1129828.912 | 602306.2985 | 564872.7352 | 1.0659 | 199.9446 | 200.0148
106 | 21 | 6 | 0 | 0 | 602306.2985 | 564872.7352 |  | 0 | 0
107 | 21 | 7 | 0 | 0 | 602306.2985 | 564872.7352 |  | 0 | 0
108 | 21 | 8 | 0 | 0 | 602306.2985 | 564872.7352 |  | 0 | 0
109 | 21 | 9 | 0 | 0 | 602306.2985 | 564872.7352 |  | 0 | 0
110 | 21 | 10 | 0 | 0 | 602306.2985 | 564872.7352 |  | 0 | 0
111 | 23 | 1 | 1111136.638 | 1069968.017 | 554102.7921 | 533229.0375 | 1.0385 | 200.529 | 200.6582
112 | 23 | 2 | 1108478.888 | 1066407.552 | 554102.7921 | 533229.0375 | 1.0395 | 200.0493 | 199.9905
113 | 23 | 3 | 1107617.711 | 1065942.208 | 554102.7921 | 533229.0375 | 1.0391 | 199.8939 | 199.9033
114 | 23 | 4 | 1103457.976 | 1061815.832 | 554102.7921 | 533229.0375 | 1.0392 | 199.1432 | 199.1294
115 | 23 | 5 | 1110336.708 | 1068156.766 | 554102.7921 | 533229.0375 | 1.0395 | 200.3846 | 200.3186
116 | 23 | 6 | 0 | 0 | 554102.7921 | 533229.0375 |  | 0 | 0
117 | 23 | 7 | 0 | 0 | 554102.7921 | 533229.0375 |  | 0 | 0
118 | 23 | 8 | 0 | 0 | 554102.7921 | 533229.0375 |  | 0 | 0
119 | 23 | 9 | 0 | 0 | 554102.7921 | 533229.0375 |  | 0 | 0
120 | 23 | 10 | 0 | 0 | 554102.7921 | 533229.0375 |  | 0 | 0
121 | 25 | 1 | 1214994.385 | 1119541.706 | 609060.7832 | 560731.7647 | 1.0853 | 199.4866 | 199.6573
122 | 25 | 2 | 1215094.076 | 1118391.338 | 609060.7832 | 560731.7647 | 1.0865 | 199.5029 | 199.4521
123 | 25 | 3 | 1214792.601 | 1118592.243 | 609060.7832 | 560731.7647 | 1.086 | 199.4534 | 199.4879
124 | 25 | 4 | 1224695.274 | 1127251.372 | 609060.7832 | 560731.7647 | 1.0864 | 201.0793 | 201.0322
125 | 25 | 5 | 1221031.496 | 1123540.988 | 609060.7832 | 560731.7647 | 1.0868 | 200.4778 | 200.3705
126 | 25 | 6 | 0 | 0 | 609060.7832 | 560731.7647 |  | 0 | 0
127 | 25 | 7 | 0 | 0 | 609060.7832 | 560731.7647 |  | 0 | 0
128 | 25 | 8 | 0 | 0 | 609060.7832 | 560731.7647 |  | 0 | 0
129 | 25 | 9 | 0 | 0 | 609060.7832 | 560731.7647 |  | 0 | 0
130 | 25 | 10 | 0 | 0 | 609060.7832 | 560731.7647 |  | 0 | 0
131 | 27 | 1 | 1231031.771 | 1122825.22 | 619860.1454 | 565887.154 | 1.0964 | 198.5983 | 198.4186
132 | 27 | 2 | 1229925.437 | 1123386.184 | 619860.1454 | 565887.154 | 1.0948 | 198.4198 | 198.5177
133 | 27 | 3 | 1248030.882 | 1139291.35 | 619860.1454 | 565887.154 | 1.0954 | 201.3407 | 201.3284
134 | 27 | 4 | 1246312.676 | 1138186.182 | 619860.1454 | 565887.154 | 1.095 | 201.0635 | 201.1331
135 | 27 | 5 | 1243300.688 | 1135182.604 | 619860.1454 | 565887.154 | 1.0952 | 200.5776 | 200.6023
136 | 27 | 6 | 0 | 0 | 619860.1454 | 565887.154 |  | 0 | 0
137 | 27 | 7 | 0 | 0 | 619860.1454 | 565887.154 |  | 0 | 0
138 | 27 | 8 | 0 | 0 | 619860.1454 | 565887.154 |  | 0 | 0
139 | 27 | 9 | 0 | 0 | 619860.1454 | 565887.154 |  | 0 | 0
140 | 27 | 10 | 0 | 0 | 619860.1454 | 565887.154 |  | 0 | 0
141 | 29 | 1 | 1051854.309 | 1107376.208 | 523111.6112 | 550927.8408 | 0.94986 | 201.0765 | 201.002
142 | 29 | 2 | 1042755.187 | 1098369.294 | 523111.6112 | 550927.8408 | 0.94937 | 199.337 | 199.3672
143 | 29 | 3 | 1047176.247 | 1102194.529 | 523111.6112 | 550927.8408 | 0.95008 | 200.1822 | 200.0615
144 | 29 | 4 | 1044414.235 | 1100711.576 | 523111.6112 | 550927.8408 | 0.94885 | 199.6542 | 199.7923
145 | 29 | 5 | 1044916.134 | 1100626.801 | 523111.6112 | 550927.8408 | 0.94938 | 199.7501 | 199.7769
146 | 29 | 6 | 0 | 0 | 523111.6112 | 550927.8408 |  | 0 | 0
147 | 29 | 7 | 0 | 0 | 523111.6112 | 550927.8408 |  | 0 | 0
148 | 29 | 8 | 0 | 0 | 523111.6112 | 550927.8408 |  | 0 | 0
149 | 29 | 9 | 0 | 0 | 523111.6112 | 550927.8408 |  | 0 | 0
150 | 29 | 10 | 0 | 0 | 523111.6112 | 550927.8408 |  | 0 | 0
151 | 31 | 1 | 1075063.367 | 1186723.377 | 538043.1758 | 593727.0292 | 0.90591 | 199.8099 | 199.8769
152 | 31 | 2 | 1073904.317 | 1184810.323 | 538043.1758 | 593727.0292 | 0.90639 | 199.5944 | 199.5547
153 | 31 | 3 | 1075612.961 | 1187331.414 | 538043.1758 | 593727.0292 | 0.90591 | 199.912 | 199.9793
154 | 31 | 4 | 1080237.281 | 1191662.632 | 538043.1758 | 593727.0292 | 0.9065 | 200.7715 | 200.7088
155 | 31 | 5 | 1075613.832 | 1186742.546 | 538043.1758 | 593727.0292 | 0.90636 | 199.9122 | 199.8802
156 | 31 | 6 | 0 | 0 | 538043.1758 | 593727.0292 |  | 0 | 0
157 | 31 | 7 | 0 | 0 | 538043.1758 | 593727.0292 |  | 0 | 0
158 | 31 | 8 | 0 | 0 | 538043.1758 | 593727.0292 |  | 0 | 0
159 | 31 | 9 | 0 | 0 | 538043.1758 | 593727.0292 |  | 0 | 0
160 | 31 | 10 | 0 | 0 | 538043.1758 | 593727.0292 |  | 0 | 0
161 | 33 | 1 | 1139182.618 | 1115025.262 | 567356.3525 | 555258.661 | 1.0217 | 200.7878 | 200.8119
162 | 33 | 2 | 1136973.116 | 1112309.551 | 567356.3525 | 555258.661 | 1.0222 | 200.3984 | 200.3228
163 | 33 | 3 | 1129080.134 | 1105262.933 | 567356.3525 | 555258.661 | 1.0215 | 199.0072 | 199.0537
164 | 33 | 4 | 1133510.14 | 1109591.882 | 567356.3525 | 555258.661 | 1.0216 | 199.788 | 199.8333
165 | 33 | 5 | 1134817.517 | 1110396.982 | 567356.3525 | 555258.661 | 1.022 | 200.0185 | 199.9783
166 | 33 | 6 | 0 | 0 | 567356.3525 | 555258.661 |  | 0 | 0
167 | 33 | 7 | 0 | 0 | 567356.3525 | 555258.661 |  | 0 | 0
168 | 33 | 8 | 0 | 0 | 567356.3525 | 555258.661 |  | 0 | 0
169 | 33 | 9 | 0 | 0 | 567356.3525 | 555258.661 |  | 0 | 0
170 | 33 | 10 | 0 | 0 | 567356.3525 | 555258.661 |  | 0 | 0
171 | 35 | 1 | 1128984.331 | 1088690.608 | 566456.1796 | 546189.7134 | 1.037 | 199.3066 | 199.3246
172 | 35 | 2 | 1132102.238 | 1090603.132 | 566456.1796 | 546189.7134 | 1.0381 | 199.857 | 199.6748
173 | 35 | 3 | 1135263.031 | 1094892.39 | 566456.1796 | 546189.7134 | 1.0369 | 200.415 | 200.4601
174 | 35 | 4 | 1132950.743 | 1092880.124 | 566456.1796 | 546189.7134 | 1.0367 | 200.0068 | 200.0917
175 | 35 | 5 | 1135261.453 | 1094830.88 | 566456.1796 | 546189.7134 | 1.0369 | 200.4147 | 200.4488
176 | 35 | 6 | 0 | 0 | 566456.1796 | 546189.7134 |  | 0 | 0
177 | 35 | 7 | 0 | 0 | 566456.1796 | 546189.7134 |  | 0 | 0
178 | 35 | 8 | 0 | 0 | 566456.1796 | 546189.7134 |  | 0 | 0
179 | 35 | 9 | 0 | 0 | 566456.1796 | 546189.7134 |  | 0 | 0
180 | 35 | 10 | 0 | 0 | 566456.1796 | 546189.7134 |  | 0 | 0
181 | 37 | 1 | 1197496.24 | 1124042.412 | 599074.5522 | 562242.2298 | 1.0653 | 199.891 | 199.9214
182 | 37 | 2 | 1204177.83 | 1130030.631 | 599074.5522 | 562242.2298 | 1.0656 | 201.0063 | 200.9864
183 | 37 | 3 | 1195385.258 | 1121726.474 | 599074.5522 | 562242.2298 | 1.0657 | 199.5386 | 199.5095
184 | 37 | 4 | 1196491.057 | 1123078.377 | 599074.5522 | 562242.2298 | 1.0654 | 199.7232 | 199.7499
185 | 37 | 5 | 1197195.137 | 1123544.404 | 599074.5522 | 562242.2298 | 1.0656 | 199.8408 | 199.8328
186 | 37 | 6 | 0 | 0 | 599074.5522 | 562242.2298 |  | 0 | 0
187 | 37 | 7 | 0 | 0 | 599074.5522 | 562242.2298 |  | 0 | 0
188 | 37 | 8 | 0 | 0 | 599074.5522 | 562242.2298 |  | 0 | 0
189 | 37 | 9 | 0 | 0 | 599074.5522 | 562242.2298 |  | 0 | 0
190 | 37 | 10 | 0 | 0 | 599074.5522 | 562242.2298 |  | 0 | 0
191 | 39 | 1 | 1107071.474 | 1066707.794 | 553570.8057 | 532716.1043 | 1.0378 | 199.9873 | 200.2394
192 | 39 | 2 | 1108526.663 | 1067254.09 | 553570.8057 | 532716.1043 | 1.0387 | 200.2502 | 200.342
193 | 39 | 3 | 1110790.441 | 1068861.644 | 553570.8057 | 532716.1043 | 1.0392 | 200.6591 | 200.6438
194 | 39 | 4 | 1106968.595 | 1064534.051 | 553570.8057 | 532716.1043 | 1.0399 | 199.9687 | 199.8314
195 | 39 | 5 | 1102350.884 | 1059803.464 | 553570.8057 | 532716.1043 | 1.0401 | 199.1346 | 198.9434
196 | 39 | 6 | 0 | 0 | 553570.8057 | 532716.1043 |  | 0 | 0
197 | 39 | 7 | 0 | 0 | 553570.8057 | 532716.1043 |  | 0 | 0
198 | 39 | 8 | 0 | 0 | 553570.8057 | 532716.1043 |  | 0 | 0
199 | 39 | 9 | 0 | 0 | 553570.8057 | 532716.1043 |  | 0 | 0
200 | 39 | 10 | 0 | 0 | 553570.8057 | 532716.1043 |  | 0 | 0
201 | 41 | 1 | 1214189.494 | 1119554.659 | 606682.3064 | 559344.9817 | 1.0845 | 200.136 | 200.1546
202 | 41 | 2 | 1211017.274 | 1117138.834 | 606682.3064 | 559344.9817 | 1.084 | 199.6131 | 199.7227
203 | 41 | 3 | 1212082.72 | 1117142.475 | 606682.3064 | 559344.9817 | 1.085 | 199.7887 | 199.7233
204 | 41 | 4 | 1216504.246 | 1120763.521 | 606682.3064 | 559344.9817 | 1.0854 | 200.5175 | 200.3707
205 | 41 | 5 | 1213029.33 | 1118850.328 | 606682.3064 | 559344.9817 | 1.0842 | 199.9447 | 200.0287
206 | 41 | 6 | 0 | 0 | 606682.3064 | 559344.9817 |  | 0 | 0
207 | 41 | 7 | 0 | 0 | 606682.3064 | 559344.9817 |  | 0 | 0
208 | 41 | 8 | 0 | 0 | 606682.3064 | 559344.9817 |  | 0 | 0
209 | 41 | 9 | 0 | 0 | 606682.3064 | 559344.9817 |  | 0 | 0
210 | 41 | 10 | 0 | 0 | 606682.3064 | 559344.9817 |  | 0 | 0
211 | 43 | 1 | 1233749.107 | 1128012.718 | 615415.6396 | 563166.061 | 1.0937 | 200.4741 | 200.2984
212 | 43 | 2 | 1222378.737 | 1118345.075 | 615415.6396 | 563166.061 | 1.093 | 198.6265 | 198.5818
213 | 43 | 3 | 1229222.363 | 1125148.645 | 615415.6396 | 563166.061 | 1.0925 | 199.7386 | 199.7899
214 | 43 | 4 | 1231134.745 | 1127306.087 | 615415.6396 | 563166.061 | 1.0921 | 200.0493 | 200.1729
215 | 43 | 5 | 1237671.444 | 1132848.085 | 615415.6396 | 563166.061 | 1.0925 | 201.1115 | 201.157
216 | 43 | 6 | 0 | 0 | 615415.6396 | 563166.061 |  | 0 | 0
217 | 43 | 7 | 0 | 0 | 615415.6396 | 563166.061 |  | 0 | 0
218 | 43 | 8 | 0 | 0 | 615415.6396 | 563166.061 |  | 0 | 0
219 | 43 | 9 | 0 | 0 | 615415.6396 | 563166.061 |  | 0 | 0
220 | 43 | 10 | 0 | 0 | 615415.6396 | 563166.061 |  | 0 | 0
221 | 45 | 1 | 1040693.682 | 1096859.491 | 522614.055 | 550839.2051 | 0.94879 | 199.1324 | 199.1252
222 | 45 | 2 | 1050090.964 | 1106887.158 | 522614.055 | 550839.2051 | 0.94869 | 200.9305 | 200.9456
223 | 45 | 3 | 1047076.245 | 1103989.458 | 522614.055 | 550839.2051 | 0.94845 | 200.3536 | 200.4196
224 | 45 | 4 | 1042912.073 | 1099119.616 | 522614.055 | 550839.2051 | 0.94886 | 199.5568 | 199.5355
225 | 45 | 5 | 1045367.586 | 1101536.328 | 522614.055 | 550839.2051 | 0.94901 | 200.0267 | 199.9742
226 | 45 | 6 | 0 | 0 | 522614.055 | 550839.2051 |  | 0 | 0
227 | 45 | 7 | 0 | 0 | 522614.055 | 550839.2051 |  | 0 | 0
228 | 45 | 8 | 0 | 0 | 522614.055 | 550839.2051 |  | 0 | 0
229 | 45 | 9 | 0 | 0 | 522614.055 | 550839.2051 |  | 0 | 0
230 | 45 | 10 | 0 | 0 | 522614.055 | 550839.2051 |  | 0 | 0
231 | 47 | 1 | 1067269.462 | 1179873.516 | 533105.857 | 588748.9775 | 0.90456 | 200.1984 | 200.4035
232 | 47 | 2 | 1071799.706 | 1184405.543 | 533105.857 | 588748.9775 | 0.90493 | 201.0482 | 201.1733
233 | 47 | 3 | 1059994.899 | 1170409.358 | 533105.857 | 588748.9775 | 0.90566 | 198.8338 | 198.796
234 | 47 | 4 | 1064719.017 | 1175143.051 | 533105.857 | 588748.9775 | 0.90603 | 199.72 | 199.6
235 | 47 | 5 | 1067275.486 | 1177658.307 | 533105.857 | 588748.9775 | 0.90627 | 200.1995 | 200.0272
236 | 47 | 6 | 0 | 0 | 533105.857 | 588748.9775 |  | 0 | 0
237 | 47 | 7 | 0 | 0 | 533105.857 | 588748.9775 |  | 0 | 0
238 | 47 | 8 | 0 | 0 | 533105.857 | 588748.9775 |  | 0 | 0
239 | 47 | 9 | 0 | 0 | 533105.857 | 588748.9775 |  | 0 | 0
240 | 47 | 10 | 0 | 0 | 533105.857 | 588748.9775 |  | 0 | 0
241 | 49 | 1 | 1140342.815 | 1118347.55 | 569221.5378 | 558095.9966 | 1.0197 | 200.3337 | 200.3862
242 | 49 | 2 | 1134213.874 | 1112404.399 | 569221.5378 | 558095.9966 | 1.0196 | 199.257 | 199.3213
243 | 49 | 3 | 1139337.124 | 1117038.797 | 569221.5378 | 558095.9966 | 1.02 | 200.1571 | 200.1517
244 | 49 | 4 | 1141799.138 | 1119556.019 | 569221.5378 | 558095.9966 | 1.0199 | 200.5896 | 200.6028
245 | 49 | 5 | 1136522.427 | 1113613.201 | 569221.5378 | 558095.9966 | 1.0206 | 199.6626 | 199.5379
246 | 49 | 6 | 0 | 0 | 569221.5378 | 558095.9966 |  | 0 | 0
247 | 49 | 7 | 0 | 0 | 569221.5378 | 558095.9966 |  | 0 | 0
248 | 49 | 8 | 0 | 0 | 569221.5378 | 558095.9966 |  | 0 | 0
249 | 49 | 9 | 0 | 0 | 569221.5378 | 558095.9966 |  | 0 | 0
250 | 49 | 10 | 0 | 0 | 569221.5378 | 558095.9966 |  | 0 | 0
251 | 51 | 1 | 1116463.299 | 1077175.938 | 558825.9259 | 539387.5149 | 1.0365 | 199.7873 | 199.7035
252 | 51 | 2 | 1118475.078 | 1079947.016 | 558825.9259 | 539387.5149 | 1.0357 | 200.1473 | 200.2173
253 | 51 | 3 | 1123401.885 | 1084519.444 | 558825.9259 | 539387.5149 | 1.0359 | 201.0289 | 201.065
254 | 51 | 4 | 1112500.001 | 1074154.902 | 558825.9259 | 539387.5149 | 1.0357 | 199.0781 | 199.1434
255 | 51 | 5 | 1117418.996 | 1078077.849 | 558825.9259 | 539387.5149 | 1.0365 | 199.9583 | 199.8707
256 | 51 | 6 | 0 | 0 | 558825.9259 | 539387.5149 |  | 0 | 0
257 | 51 | 7 | 0 | 0 | 558825.9259 | 539387.5149 |  | 0 | 0
258 | 51 | 8 | 0 | 0 | 558825.9259 | 539387.5149 |  | 0 | 0
259 | 51 | 9 | 0 | 0 | 558825.9259 | 539387.5149 |  | 0 | 0
260 | 51 | 10 | 0 | 0 | 558825.9259 | 539387.5149 |  | 0 | 0
261 | 53 | 1 | 1207101.011 | 1133854.775 | 601790.4169 | 565592.3758 | 1.0646 | 200.585 | 200.4721
262 | 53 | 2 | 1201211.668 | 1129339.09 | 601790.4169 | 565592.3758 | 1.0636 | 199.6063 | 199.6737
263 | 53 | 3 | 1202425.568 | 1130669.555 | 601790.4169 | 565592.3758 | 1.0635 | 199.808 | 199.9089
264 | 53 | 4 | 1203733.724 | 1130929.134 | 601790.4169 | 565592.3758 | 1.0644 | 200.0254 | 199.9548
265 | 53 | 5 | 1203432.198 | 1131131.204 | 601790.4169 | 565592.3758 | 1.0639 | 199.9753 | 199.9905
266 | 53 | 6 | 0 | 0 | 601790.4169 | 565592.3758 |  | 0 | 0
267 | 53 | 7 | 0 | 0 | 601790.4169 | 565592.3758 |  | 0 | 0
268 | 53 | 8 | 0 | 0 | 601790.4169 | 565592.3758 |  | 0 | 0
269 | 53 | 9 | 0 | 0 | 601790.4169 | 565592.3758 |  | 0 | 0
270 | 53 | 10 | 0 | 0 | 601790.4169 | 565592.3758 |  | 0 | 0
271 | 55 | 1 | 1106567.61 | 1067109.359 | 551435.4523 | 532109.0488 | 1.037 | 200.6704 | 200.5434
272 | 55 | 2 | 1108476.508 | 1069821.552 | 551435.4523 | 532109.0488 | 1.0361 | 201.0165 | 201.0531
273 | 55 | 3 | 1100038.429 | 1062071.518 | 551435.4523 | 532109.0488 | 1.0357 | 199.4863 | 199.5966
274 | 55 | 4 | 1098228.536 | 1059512.967 | 551435.4523 | 532109.0488 | 1.0365 | 199.1581 | 199.1158
275 | 55 | 5 | 1101043.44 | 1062575.092 | 551435.4523 | 532109.0488 | 1.0362 | 199.6686 | 199.6912
276 | 55 | 6 | 0 | 0 | 551435.4523 | 532109.0488 |  | 0 | 0
277 | 55 | 7 | 0 | 0 | 551435.4523 | 532109.0488 |  | 0 | 0
278 | 55 | 8 | 0 | 0 | 551435.4523 | 532109.0488 |  | 0 | 0
279 | 55 | 9 | 0 | 0 | 551435.4523 | 532109.0488 |  | 0 | 0
280 | 55 | 10 | 0 | 0 | 551435.4523 | 532109.0488 |  | 0 | 0
281 | 57 | 1 | 1215341.11 | 1123379.185 | 609165.1068 | 562391.2671 | 1.0819 | 199.5093 | 199.7505
282 | 57 | 2 | 1217456.218 | 1123125.223 | 609165.1068 | 562391.2671 | 1.084 | 199.8565 | 199.7053
283 | 57 | 3 | 1224248.944 | 1130246.066 | 609165.1068 | 562391.2671 | 1.0832 | 200.9716 | 200.9715
284 | 57 | 4 | 1218053.684 | 1124587.731 | 609165.1068 | 562391.2671 | 1.0831 | 199.9546 | 199.9654
285 | 57 | 5 | 1216551.112 | 1122574.466 | 609165.1068 | 562391.2671 | 1.0837 | 199.7079 | 199.6074
286 | 57 | 6 | 0 | 0 | 609165.1068 | 562391.2671 |  | 0 | 0
287 | 57 | 7 | 0 | 0 | 609165.1068 | 562391.2671 |  | 0 | 0
288 | 57 | 8 | 0 | 0 | 609165.1068 | 562391.2671 |  | 0 | 0
289 | 57 | 9 | 0 | 0 | 609165.1068 | 562391.2671 |  | 0 | 0
290 | 57 | 10 | 0 | 0 | 609165.1068 | 562391.2671 |  | 0 | 0
291 | 59 | 1 | 1240087.037 | 1135925.686 | 618498.6797 | 566480.4208 | 1.0917 | 200.4995 | 200.5234
292 | 59 | 2 | 1237469.358 | 1132844.295 | 618498.6797 | 566480.4208 | 1.0924 | 200.0763 | 199.9794
293 | 59 | 3 | 1235861.642 | 1132041.213 | 618498.6797 | 566480.4208 | 1.0917 | 199.8164 | 199.8377
294 | 59 | 4 | 1238677.062 | 1135359.124 | 618498.6797 | 566480.4208 | 1.091 | 200.2716 | 200.4234
295 | 59 | 5 | 1232891.698 | 1128633.89 | 618498.6797 | 566480.4208 | 1.0924 | 199.3362 | 199.2362
296 | 59 | 6 | 0 | 0 | 618498.6797 | 566480.4208 |  | 0 | 0
297 | 59 | 7 | 0 | 0 | 618498.6797 | 566480.4208 |  | 0 | 0
298 | 59 | 8 | 0 | 0 | 618498.6797 | 566480.4208 |  | 0 | 0
299 | 59 | 9 | 0 | 0 | 618498.6797 | 566480.4208 |  | 0 | 0
300 | 59 | 10 | 0 | 0 | 618498.6797 | 566480.4208 |  | 0 | 0
301 | 61 | 1 | 1037383.854 | 1094632.003 | 518694.5466 | 547129.9068 | 0.9477 | 199.999 | 200.068
302 | 61 | 2 | 1034669.461 | 1092109.575 | 518694.5466 | 547129.9068 | 0.9474 | 199.4757 | 199.607
303 | 61 | 3 | 1033107.925 | 1089450.01 | 518694.5466 | 547129.9068 | 0.94828 | 199.1746 | 199.1209
304 | 61 | 4 | 1041143.906 | 1097402.695 | 518694.5466 | 547129.9068 | 0.94873 | 200.7239 | 200.5744
305 | 61 | 5 | 1040640.32 | 1097704.785 | 518694.5466 | 547129.9068 | 0.94801 | 200.6268 | 200.6296
306 | 61 | 6 | 0 | 0 | 518694.5466 | 547129.9068 |  | 0 | 0
307 | 61 | 7 | 0 | 0 | 518694.5466 | 547129.9068 |  | 0 | 0
308 | 61 | 8 | 0 | 0 | 518694.5466 | 547129.9068 |  | 0 | 0
309 | 61 | 9 | 0 | 0 | 518694.5466 | 547129.9068 |  | 0 | 0
310 | 61 | 10 | 0 | 0 | 518694.5466 | 547129.9068 |  | 0 | 0
311 | 63 | 1 | 1063956.922 | 1177214.457 | 530953.0757 | 587648.9021 | 0.90379 | 200.3862 | 200.3262
312 | 63 | 2 | 1066972.377 | 1180336.206 | 530953.0757 | 587648.9021 | 0.90396 | 200.9542 | 200.8574
313 | 63 | 3 | 1060236.352 | 1173888.089 | 530953.0757 | 587648.9021 | 0.90318 | 199.6855 | 199.7601
314 | 63 | 4 | 1057725.386 | 1171525.602 | 530953.0757 | 587648.9021 | 0.90286 | 199.2126 | 199.3581
315 | 63 | 5 | 1060639.72 | 1173524.667 | 530953.0757 | 587648.9021 | 0.90381 | 199.7615 | 199.6983
316 | 63 | 6 | 0 | 0 | 530953.0757 | 587648.9021 |  | 0 | 0
317 | 63 | 7 | 0 | 0 | 530953.0757 | 587648.9021 |  | 0 | 0
318 | 63 | 8 | 0 | 0 | 530953.0757 | 587648.9021 |  | 0 | 0
319 | 63 | 9 | 0 | 0 | 530953.0757 | 587648.9021 |  | 0 | 0
320 | 63 | 10 | 0 | 0 | 530953.0757 | 587648.9021 |  | 0 | 0
321 | 65 | 1 | 1123252.226 | 1103387.893 | 561691.1253 | 551324.8166 | 1.018 | 199.9769 | 200.1339
322 | 65 | 2 | 1125413.867 | 1105458.894 | 561691.1253 | 551324.8166 | 1.0181 | 200.3617 | 200.5095
323 | 65 | 3 | 1125714.162 | 1105258.174 | 561691.1253 | 551324.8166 | 1.0185 | 200.4152 | 200.4731
324 | 65 | 4 | 1120485.701 | 1099521.334 | 561691.1253 | 551324.8166 | 1.0191 | 199.4843 | 199.4326
325 | 65 | 5 | 1122045.297 | 1099621.871 | 561691.1253 | 551324.8166 | 1.0204 | 199.762 | 199.4508
326 | 65 | 6 | 0 | 0 | 561691.1253 | 551324.8166 |  | 0 | 0
327 | 65 | 7 | 0 | 0 | 561691.1253 | 551324.8166 |  | 0 | 0
328 | 65 | 8 | 0 | 0 | 561691.1253 | 551324.8166 |  | 0 | 0
329 | 65 | 9 | 0 | 0 | 561691.1253 | 551324.8166 |  | 0 | 0
330 | 65 | 10 | 0 | 0 | 561691.1253 | 551324.8166 |  | 0 | 0
331 | 67 | 1 | 1128926.161 | 1092771.06 | 563650.207 | 545384.7348 | 1.0331 | 200.2884 | 200.367
332 | 67 | 2 | 1123803.862 | 1087134.682 | 563650.207 | 545384.7348 | 1.0337 | 199.3797 | 199.3335
333 | 67 | 3 | 1129227.451 | 1092167.883 | 563650.207 | 545384.7348 | 1.0339 | 200.3419 | 200.2564
334 | 67 | 4 | 1127624.563 | 1091719.844 | 563650.207 | 545384.7348 | 1.0329 | 200.0575 | 200.1743
335 | 67 | 5 | 1126920.033 | 1090053.879 | 563650.207 | 545384.7348 | 1.0338 | 199.9325 | 199.8688
336 | 67 | 6 | 0 | 0 | 563650.207 | 545384.7348 |  | 0 | 0
337 | 67 | 7 | 0 | 0 | 563650.207 | 545384.7348 |  | 0 | 0
338 | 67 | 8 | 0 | 0 | 563650.207 | 545384.7348 |  | 0 | 0
339 | 67 | 9 | 0 | 0 | 563650.207 | 545384.7348 |  | 0 | 0
340 | 67 | 10 | 0 | 0 | 563650.207 | 545384.7348 |  | 0 | 0
341 | 69 | 1 | 1196393.54 | 1124993.858 | 595766.355 | 560607.5973 | 1.0635 | 200.8159 | 200.674
342 | 69 | 2 | 1198701.877 | 1128210.469 | 595766.355 | 560607.5973 | 1.0625 | 201.2034 | 201.2478
343 | 69 | 3 | 1188589.214 | 1118862.601 | 595766.355 | 560607.5973 | 1.0623 | 199.5059 | 199.5803
344 | 69 | 4 | 1189956.601 | 1119048.385 | 595766.355 | 560607.5973 | 1.0634 | 199.7354 | 199.6135
345 | 69 | 5 | 1184022.318 | 1114960.66 | 595766.355 | 560607.5973 | 1.0619 | 198.7394 | 198.8843
346 | 69 | 6 | 0 | 0 | 595766.355 | 560607.5973 |  | 0 | 0
347 | 69 | 7 | 0 | 0 | 595766.355 | 560607.5973 |  | 0 | 0
348 | 69 | 8 | 0 | 0 | 595766.355 | 560607.5973 |  | 0 | 0
349 | 69 | 9 | 0 | 0 | 595766.355 | 560607.5973 |  | 0 | 0
350 | 69 | 10 | 0 | 0 | 595766.355 | 560607.5973 |  | 0 | 0
351 | 71 | 1 | 1106270.505 | 1069921.317 | 552018.9978 | 533359.3367 | 1.034 | 200.4044 | 200.6005
352 | 71 | 2 | 1103052.919 | 1066096.521 | 552018.9978 | 533359.3367 | 1.0347 | 199.8216 | 199.8834
353 | 71 | 3 | 1102351.327 | 1065275.82 | 552018.9978 | 533359.3367 | 1.0348 | 199.6945 | 199.7295
354 | 71 | 4 | 1106868.885 | 1068919.332 | 552018.9978 | 533359.3367 | 1.0355 | 200.5128 | 200.4126
355 | 71 | 5 | 1101646.342 | 1063380.377 | 552018.9978 | 533359.3367 | 1.036 | 199.5667 | 199.3741
356 | 71 | 6 | 0 | 0 | 552018.9978 | 533359.3367 |  | 0 | 0
357 | 71 | 7 | 0 | 0 | 552018.9978 | 533359.3367 |  | 0 | 0
358 | 71 | 8 | 0 | 0 | 552018.9978 | 533359.3367 |  | 0 | 0
359 | 71 | 9 | 0 | 0 | 552018.9978 | 533359.3367 |  | 0 | 0
360 | 71 | 10 | 0 | 0 | 552018.9978 | 533359.3367 |  | 0 | 0
361 | 73 | 1 | 1215945.612 | 1123944.006 | 605555.4387 | 559721.7607 | 1.0819 | 200.7984 | 200.8041
362 | 73 | 2 | 1212526.617 | 1120659.372 | 605555.4387 | 559721.7607 | 1.082 | 200.2338 | 200.2172
363 | 73 | 3 | 1208258.549 | 1116934.245 | 605555.4387 | 559721.7607 | 1.0818 | 199.529 | 199.5517
364 | 73 | 4 | 1212125.136 | 1120558.408 | 605555.4387 | 559721.7607 | 1.0817 | 200.1675 | 200.1992
365 | 73 | 5 | 1206698.473 | 1115121.576 | 605555.4387 | 559721.7607 | 1.0821 | 199.2713 | 199.2278
366 | 73 | 6 | 0 | 0 | 605555.4387 | 559721.7607 |  | 0 | 0
367 | 73 | 7 | 0 | 0 | 605555.4387 | 559721.7607 |  | 0 | 0
368 | 73 | 8 | 0 | 0 | 605555.4387 | 559721.7607 |  | 0 | 0
369 | 73 | 9 | 0 | 0 | 605555.4387 | 559721.7607 |  | 0 | 0
370 | 73 | 10 | 0 | 0 | 605555.4387 | 559721.7607 |  | 0 | 0
371 | 75 | 1 | 1235961.646 | 1132041.021 | 617965.9657 | 566252.7535 | 1.0918 | 200.0048 | 199.918
372 | 75 | 2 | 1239175.928 | 1135162.795 | 617965.9657 | 566252.7535 | 1.0916 | 200.5249 | 200.4693
373 | 75 | 3 | 1238174.468 | 1134272.301 | 617965.9657 | 566252.7535 | 1.0916 | 200.3629 | 200.312
374 | 75 | 4 | 1234856.368 | 1132640.059 | 617965.9657 | 566252.7535 | 1.0902 | 199.826 | 200.0238
375 | 75 | 5 | 1231491.247 | 1128411.359 | 617965.9657 | 566252.7535 | 1.0913 | 199.2814 | 199.277
376 | 75 | 6 | 0 | 0 | 617965.9657 | 566252.7535 |  | 0 | 0
377 | 75 | 7 | 0 | 0 | 617965.9657 | 566252.7535 |  | 0 | 0
378 | 75 | 8 | 0 | 0 | 617965.9657 | 566252.7535 |  | 0 | 0
379 | 75 | 9 | 0 | 0 | 617965.9657 | 566252.7535 |  | 0 | 0
380 | 75 | 10 | 0 | 0 | 617965.9657 | 566252.7535 |  | 0 | 0
381 | 77 | 1 | 1042308.547 | 1099013.191 | 518239.4845 | 547106.2833 | 0.9484 | 201.1249 | 200.8775
382 | 77 | 2 | 1036377.751 | 1094885.522 | 518239.4845 | 547106.2833 | 0.94656 | 199.9805 | 200.123
383 | 77 | 3 | 1038086.948 | 1095849.484 | 518239.4845 | 547106.2833 | 0.94729 | 200.3103 | 200.2992
384 | 77 | 4 | 1032761.75 | 1091160.414 | 518239.4845 | 547106.2833 | 0.94648 | 199.2827 | 199.4421
385 | 77 | 5 | 1032859.849 | 1090154.222 | 518239.4845 | 547106.2833 | 0.94744 | 199.3017 | 199.2582
386 | 77 | 6 | 0 | 0 | 518239.4845 | 547106.2833 |  | 0 | 0
387 | 77 | 7 | 0 | 0 | 518239.4845 | 547106.2833 |  | 0 | 0
388 | 77 | 8 | 0 | 0 | 518239.4845 | 547106.2833 |  | 0 | 0
389 | 77 | 9 | 0 | 0 | 518239.4845 | 547106.2833 |  | 0 | 0
390 | 77 | 10 | 0 | 0 | 518239.4845 | 547106.2833 |  | 0 | 0
391 | 79 | 1 | 1073005.151 | 1186736.112 | 535837.8799 | 592799.1009 | 0.90416 | 200.2481 | 200.192
392 | 79 | 2 | 1072802.264 | 1186635.355 | 535837.8799 | 592799.1009 | 0.90407 | 200.2102 | 200.175
393 | 79 | 3 | 1071648.96 | 1185830.245 | 535837.8799 | 592799.1009 | 0.90371 | 199.995 | 200.0391
394 | 79 | 4 | 1065712.88 | 1178481.344 | 535837.8799 | 592799.1009 | 0.90431 | 198.8872 | 198.7994
395 | 79 | 5 | 1075209.544 | 1190307.953 | 535837.8799 | 592799.1009 | 0.9033 | 200.6595 | 200.7945
396 | 79 | 6 | 0 | 0 | 535837.8799 | 592799.1009 |  | 0 | 0
397 | 79 | 7 | 0 | 0 | 535837.8799 | 592799.1009 |  | 0 | 0
398 | 79 | 8 | 0 | 0 | 535837.8799 | 592799.1009 |  | 0 | 0
399 | 79 | 9 | 0 | 0 | 535837.8799 | 592799.1009 |  | 0 | 0
400 | 79 | 10 | 0 | 0 | 535837.8799 | 592799.1009 |  | 0 | 0

## ttable (hidden)
t | nu
1.9794 | 123
1.9796 | 122
1.9798 | 121
1.9799 | 120
1.9801 | 119
1.9803 | 118
1.9804 | 117
1.9806 | 116
1.9808 | 115
1.981 | 114
1.9812 | 113
1.9814 | 112
1.9816 | 111
1.9818 | 110
1.982 | 109
1.9822 | 108
1.9824 | 107
1.9826 | 106
1.9828 | 105
1.983 | 104
1.9833 | 103
1.9835 | 102
1.9837 | 101
1.984 | 100
1.9842 | 99
1.9845 | 98
1.9847 | 97
1.985 | 96
1.9853 | 95
1.9855 | 94
1.9858 | 93
1.9861 | 92
1.9864 | 91
1.9867 | 90
1.987 | 89
1.9873 | 88
1.9876 | 87
1.9879 | 86
1.9883 | 85
1.9886 | 84
1.989 | 83
1.9893 | 82
1.9897 | 81
1.9901 | 80
1.9905 | 79
1.9908 | 78
1.9913 | 77
1.9917 | 76
1.9921 | 75
1.9925 | 74
1.993 | 73
1.9935 | 72
1.9939 | 71
1.9944 | 70
1.9949 | 69
1.9955 | 68
1.996 | 67
1.9966 | 66
1.9971 | 65
1.9977 | 64
1.9983 | 63
1.999 | 62
1.9996 | 61
2.0003 | 60
2.001 | 59
2.0017 | 58
2.0025 | 57
2.0032 | 56
2.004 | 55
2.0049 | 54
2.0057 | 53
2.0066 | 52
2.0076 | 51
2.0086 | 50
2.0096 | 49
2.0106 | 48
2.0117 | 47
2.0129 | 46
2.0141 | 45
2.0154 | 44
2.0167 | 43
2.0181 | 42
2.0195 | 41
2.0211 | 40
2.0227 | 39
2.0244 | 38
2.0262 | 37
2.0281 | 36
2.0301 | 35
2.0322 | 34
2.0345 | 33
2.0369 | 32
2.0395 | 31
2.0423 | 30
2.0452 | 29
2.0484 | 28
2.0518 | 27
2.0555 | 26
2.0595 | 25
2.0639 | 24
2.0687 | 23
2.0739 | 22
2.0796 | 21
2.086 | 20
2.093 | 19
2.1009 | 18
2.1098 | 17
2.1199 | 16
2.1314 | 15
2.1448 | 14
2.1604 | 13
2.1788 | 12
2.201 | 11
2.2281 | 10
2.2622 | 9
2.306 | 8
2.3646 | 7
2.4469 | 6
2.5706 | 5
2.7764 | 4
3.1824 | 3
4.3027 | 2
12.7062 | 1

## Drift Correction Module (hidden)
linest
[10 empty] | x6 | x5 | x4 | x3 | x2 | x | b
[10 empty] | 4.318e-11 | -4.6218e-09 | 1.918e-07 | -3.7494e-06 | 2.6867e-05 | 8.6414e-07 | 1.0011
Repeat # | ID | Average | Residual Ratio | x | x^2 | x^3 | x^4 | x^5 | x^6 | Drift Line |  | Drift Corrected Average
1 | SRM 1 | 1.0219 | 1.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1.0011 |  | 1.0208
0 | Test Sample 1 | 1.0379 | 1.0014 | 2 | 4 | 8 | 16 | 32 | 64 | 1.0012 |  | 1.0367
0 | SRM 2 | 1.0662 | 1.0012 | 3 | 9 | 27 | 81 | 243 | 729 | 1.0012 |  | 1.0649
0 | Test Sample 2 | 1.0388 | 1.0011 | 4 | 16 | 64 | 256 | 1024 | 4096 | 1.0013 |  | 1.0374
0 | SRM 3 | 1.0854 | 1.0011 | 5 | 25 | 125 | 625 | 3125 | 15625 | 1.0014 |  | 1.0839
0 | Test Sample 3 | 1.0953 | 1.0018 | 6 | 36 | 216 | 1296 | 7776 | 46656 | 1.0015 |  | 1.0937
0 | SRM 4 | 0.95105 | 1.0022 | 7 | 49 | 343 | 2401 | 16807 | 117649 | 1.0015 |  | 0.94962
0 | Test Sample 4 | 0.90641 | 1.0014 | 8 | 64 | 512 | 4096 | 32768 | 262144 | 1.0015 |  | 0.90502
2 | SRM 1 | 1.0221 | 1.0012 | 9 | 81 | 729 | 6561 | 59049 | 531441 | 1.0015 |  | 1.0206
0 | Test Sample 1 | 1.0375 | 1.0011 | 10 | 100 | 1000 | 10000 | 100000 | 1000000 | 1.0015 |  | 1.0359
0 | SRM 2 | 1.0663 | 1.0012 | 11 | 121 | 1331 | 14641 | 161051 | 1771561 | 1.0015 |  | 1.0647
0 | Test Sample 2 | 1.0391 | 1.0014 | 12 | 144 | 1728 | 20736 | 248832 | 2985984 | 1.0014 |  | 1.0377
0 | SRM 3 | 1.0862 | 1.0018 | 13 | 169 | 2197 | 28561 | 371293 | 4826809 | 1.0014 |  | 1.0847
0 | Test Sample 3 | 1.0954 | 1.0019 | 14 | 196 | 2744 | 38416 | 537824 | 7529536 | 1.0013 |  | 1.094
0 | SRM 4 | 0.94951 | 1.0006 | 15 | 225 | 3375 | 50625 | 759375 | 11390625 | 1.0012 |  | 0.94839
0 | Test Sample 4 | 0.90621 | 1.0012 | 16 | 256 | 4096 | 65536 | 1048576 | 16777216 | 1.0011 |  | 0.90525
3 | SRM 1 | 1.0218 | 1.0009 | 17 | 289 | 4913 | 83521 | 1419857 | 24137569 | 1.0009 |  | 1.0208
3 | Test Sample 1 | 1.0371 | 1.0007 | 18 | 324 | 5832 | 104976 | 1889568 | 34012224 | 1.0008 |  | 1.0363
3 | SRM 2 | 1.0655 | 1.0005 | 19 | 361 | 6859 | 130321 | 2476099 | 47045881 | 1.0007 |  | 1.0648
3 | Test Sample 2 | 1.0391 | 1.0014 | 20 | 400 | 8000 | 160000 | 3200000 | 64000000 | 1.0005 |  | 1.0386
3 | SRM 3 | 1.0846 | 1.0003 | 21 | 441 | 9261 | 194481 | 4084101 | 85766121 | 1.0004 |  | 1.0842
3 | Test Sample 3 | 1.0928 | 0.99951 | 22 | 484 | 10648 | 234256 | 5153632 | 113379904 | 1.0002 |  | 1.0926
3 | SRM 4 | 0.94876 | 0.99984 | 23 | 529 | 12167 | 279841 | 6436343 | 148035889 | 1 |  | 0.94875
3 | Test Sample 4 | 0.90549 | 1.0004 | 24 | 576 | 13824 | 331776 | 7962624 | 191102976 | 0.99983 |  | 0.90564
4 | SRM 1 | 1.0199 | 0.99904 | 25 | 625 | 15625 | 390625 | 9765625 | 244140625 | 0.99964 |  | 1.0203
4 | Test Sample 1 | 1.036 | 0.99964 | 26 | 676 | 17576 | 456976 | 11881376 | 308915776 | 0.99944 |  | 1.0366
4 | SRM 2 | 1.064 | 0.99911 | 27 | 729 | 19683 | 531441 | 14348907 | 387420489 | 0.99923 |  | 1.0648
4 | Test Sample 2 | 1.0363 | 0.99869 | 28 | 784 | 21952 | 614656 | 17210368 | 481890304 | 0.99902 |  | 1.0373
4 | SRM 3 | 1.0832 | 0.999 | 29 | 841 | 24389 | 707281 | 20511149 | 594823321 | 0.9988 |  | 1.0845
4 | Test Sample 3 | 1.0918 | 0.99864 | 30 | 900 | 27000 | 810000 | 24300000 | 729000000 | 0.99858 |  | 1.0934
4 | SRM 4 | 0.94803 | 0.99906 | 31 | 961 | 29791 | 923521 | 28629151 | 887503681 | 0.99837 |  | 0.94958
4 | Test Sample 4 | 0.90352 | 0.99824 | 32 | 1024 | 32768 | 1048576 | 33554432 | 1073741824 | 0.99816 |  | 0.90518
5 | SRM 1 | 1.0188 | 0.99793 | 33 | 1089 | 35937 | 1185921 | 39135393 | 1291467969 | 0.99798 |  | 1.0209
5 | Test Sample 1 | 1.0335 | 0.99719 | 34 | 1156 | 39304 | 1336336 | 45435424 | 1544804416 | 0.99783 |  | 1.0357
5 | SRM 2 | 1.0627 | 0.99791 | 35 | 1225 | 42875 | 1500625 | 52521875 | 1838265625 | 0.99772 |  | 1.0651
5 | Test Sample 2 | 1.035 | 0.9974 | 36 | 1296 | 46656 | 1679616 | 60466176 | 2176782336 | 0.99768 |  | 1.0374
5 | SRM 3 | 1.0819 | 0.99781 | 37 | 1369 | 50653 | 1874161 | 69343957 | 2565726409 | 0.99774 |  | 1.0843
5 | Test Sample 3 | 1.0913 | 0.99818 | 38 | 1444 | 54872 | 2085136 | 79235168 | 3010936384 | 0.99791 |  | 1.0936
5 | SRM 4 | 0.94724 | 0.99823 | 39 | 1521 | 59319 | 2313441 | 90224199 | 3518743761 | 0.99823 |  | 0.94891
5 | Test Sample 4 | 0.90391 | 0.99868 | 40 | 1600 | 64000 | 2560000 | 102400000 | 4096000000 | 0.99875 |  | 0.90505